## Title
Varian TrueBeam
10FFF Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables
Open field
Open Field Depth Doses
Open Field Profiles at 2.4cm
Open Field Profiles at 5cm
Open Field Profiles at 10cm
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
2. Plots
Open Field

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 40.5 | 40.9 | 42.4 | 43.7 | 44.9 | 50.6 | 53.8 | 54.5
0.1 | 44.7 | 45.4 | 46.7 | 47.8 | 48.92 | 54.55 | 57.4333 | 58.275
0.2 | 52.35 | 52.775 | 53.925 | 54.875 | 55.45 | 61.125 | 63.3 | 64.025
0.3 | 61.4 | 61 | 62.1 | 63.5 | 64.075 | 68.54 | 70.5333 | 71
0.4 | 68.85 | 68.25 | 69.1 | 70.4 | 71.1 | 75.06 | 76.62 | 76.75
0.5 | 74.325 | 74.4 | 74.95 | 75.55 | 76.7 | 79.8 | 81.2 | 81.75
0.6 | 79.025 | 78.9 | 79.5 | 80.32 | 80.9667 | 83.75 | 85 | 85.2
0.7 | 83.075 | 83 | 83.45 | 83.8 | 84.66 | 87 | 88.1 | 88.2667
0.8 | 86.375 | 86.05 | 86.5 | 87.1 | 87.55 | 89.75 | 90.7 | 90.8
0.9 | 89.225 | 89 | 89.25 | 89.6 | 90 | 91.9 | 92.7 | 92.775
1 | 91.56 | 90.96 | 91.4 | 91.7 | 92.05 | 93.7 | 94.4 | 94.4
1.1 | 93.35 | 92.825 | 93.15 | 93.5 | 93.9 | 95.2 | 95.68 | 95.7
1.2 | 95.02 | 94.45 | 94.65 | 94.9 | 95.25 | 96.4 | 96.76 | 96.8
1.3 | 96.2 | 95.775 | 95.975 | 96.15 | 96.475 | 97.3 | 97.625 | 97.7
1.4 | 97.1 | 96.8 | 97 | 97.2 | 97.4 | 98.05 | 98.35 | 98.35
1.5 | 98.1 | 97.8 | 97.9 | 98.025 | 98.075 | 98.6667 | 98.925 | 98.9
1.6 | 98.75 | 98.45 | 98.525 | 98.65 | 98.75 | 99.25 | 99.325 | 99.3
1.7 | 99.26 | 99 | 99 | 99.2 | 99.3 | 99.5 | 99.6333 | 99.6
1.8 | 99.575 | 99.475 | 99.44 | 99.5667 | 99.5667 | 99.825 | 99.9 | 99.85
1.9 | 99.8 | 99.625 | 99.75 | 99.8 | 99.8 | 99.9 | 99.95 | 99.9333
2 | 99.9 | 99.85 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100
2.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
2.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
2.3 | 100 | 100 | 100 | 100 | 100 | 99.92 | 99.86 | 99.9
2.4 | 99.8333 | 99.875 | 100 | 99.9 | 99.9 | 99.825 | 99.675 | 99.7
2.5 | 99.6 | 99.7 | 99.825 | 99.825 | 99.8 | 99.625 | 99.5 | 99.5
2.6 | 99.3 | 99.475 | 99.6 | 99.575 | 99.625 | 99.375 | 99.225 | 99.325
2.7 | 99.05 | 99.225 | 99.375 | 99.325 | 99.3 | 99.125 | 99 | 99.08
2.8 | 98.725 | 98.975 | 99.025 | 99 | 99.025 | 98.875 | 98.7 | 98.85
2.9 | 98.375 | 98.625 | 98.76 | 98.8 | 98.775 | 98.54 | 98.4 | 98.55
3 | 97.925 | 98.2 | 98.45 | 98.5 | 98.45 | 98.2 | 98.15 | 98.2
3.1 | 97.575 | 97.85 | 98.1 | 98.1 | 98.2 | 97.95 | 97.8 | 97.95
3.2 | 97.15 | 97.5 | 97.75 | 97.8 | 97.75 | 97.5 | 97.46 | 97.6
3.3 | 96.6 | 97 | 97.4 | 97.42 | 97.5 | 97.25 | 97.1 | 97.2333
3.4 | 96.3 | 96.6667 | 96.9 | 97.025 | 97.1 | 96.9 | 96.825 | 96.95
3.5 | 95.8 | 96.1 | 96.5 | 96.625 | 96.65 | 96.56 | 96.475 | 96.6
3.6 | 95.35 | 95.75 | 96.05 | 96.225 | 96.25 | 96.15 | 96.15 | 96.1667
3.7 | 94.8 | 95.3 | 95.64 | 95.8 | 95.85 | 95.775 | 95.75 | 95.8333
3.8 | 94.4 | 94.7667 | 95.35 | 95.35 | 95.425 | 95.45 | 95.35 | 95.4667
3.9 | 93.9 | 94.35 | 94.85 | 94.9333 | 95.05 | 95.06 | 94.95 | 95.05
4 | 93.45 | 93.85 | 94.35 | 94.55 | 94.65 | 94.55 | 94.6 | 94.6667
4.1 | 92.85 | 93.3667 | 93.95 | 94.1 | 94.2 | 94.1667 | 94.25 | 94.3333
4.2 | 92.35 | 92.9 | 93.45 | 93.7 | 93.8 | 93.88 | 93.8 | 94
4.3 | 91.85 | 92.4333 | 92.98 | 93.2 | 93.4 | 93.4 | 93.5 | 93.5
4.4 | 91.35 | 91.8667 | 92.6 | 92.8333 | 92.9 | 93.05 | 93.1 | 93.2
4.5 | 90.8667 | 91.5 | 92.0667 | 92.375 | 92.55 | 92.6 | 92.72 | 92.8
4.6 | 90.42 | 90.9667 | 91.65 | 91.95 | 92.06 | 92.25 | 92.3333 | 92.425
4.7 | 89.8333 | 90.5 | 91.2333 | 91.5 | 91.6333 | 91.8667 | 91.925 | 92.0333
4.8 | 89.5 | 90.05 | 90.75 | 91.08 | 91.25 | 91.5 | 91.525 | 91.65
4.9 | 88.8 | 89.55 | 90.3 | 90.66 | 90.825 | 91.125 | 91.2 | 91.325
5 | 88.45 | 89.0667 | 89.85 | 90.25 | 90.375 | 90.675 | 90.75 | 90.95
5.1 | 87.9 | 88.6 | 89.35 | 89.8 | 89.98 | 90.3 | 90.4 | 90.6
5.2 | 87.4667 | 88.1 | 88.95 | 89.32 | 89.54 | 89.9333 | 90.05 | 90.2
5.3 | 86.975 | 87.65 | 88.4 | 88.95 | 89.2 | 89.5 | 89.55 | 89.75
5.4 | 86.5 | 87.2 | 88 | 88.4667 | 88.7 | 89.1 | 89.25 | 89.4333
5.5 | 86.1 | 86.75 | 87.6 | 88.1 | 88.3 | 88.825 | 88.95 | 89.0333
5.6 | 85.58 | 86.3333 | 87.225 | 87.6 | 87.95 | 88.36 | 88.6 | 88.7
5.7 | 85.05 | 85.8333 | 86.7333 | 87.2 | 87.525 | 87.95 | 88.1 | 88.34
5.8 | 84.65 | 85.28 | 86.35 | 86.8333 | 87.15 | 87.6 | 87.7 | 87.85
5.9 | 84.12 | 84.9 | 85.75 | 86.3 | 86.74 | 87.3 | 87.4 | 87.5
6 | 83.6 | 84.4333 | 85.4 | 85.85 | 86.275 | 86.7 | 86.975 | 87.1
6.1 | 83.2 | 83.975 | 85 | 85.55 | 85.925 | 86.45 | 86.6 | 86.8
6.2 | 82.8 | 83.525 | 84.6 | 85.15 | 85.46 | 86.1 | 86.325 | 86.425
6.3 | 82.25 | 82.98 | 84.14 | 84.7 | 85.15 | 85.75 | 85.975 | 86.075
6.4 | 81.7667 | 82.5 | 83.75 | 84.3 | 84.8 | 85.25 | 85.625 | 85.725
6.5 | 81.35 | 82.1667 | 83.25 | 83.8 | 84.32 | 84.98 | 85.1 | 85.375
6.6 | 80.85 | 81.7 | 82.85 | 83.4 | 83.85 | 84.65 | 84.85 | 84.95
6.7 | 80.4667 | 81.3 | 82.35 | 83 | 83.4333 | 84.2 | 84.5 | 84.55
6.8 | 80 | 80.8 | 82 | 82.7 | 83.0667 | 83.825 | 84.1 | 84.2
6.9 | 79.5667 | 80.3667 | 81.5 | 82.3 | 82.75 | 83.45 | 83.7 | 83.9
7 | 79.15 | 80 | 81.15 | 81.875 | 82.2667 | 83.125 | 83.35 | 83.5
7.1 | 78.725 | 79.5 | 80.75 | 81.425 | 81.9 | 82.74 | 83.025 | 83.2
7.2 | 78.25 | 79.1 | 80.225 | 81.1 | 81.46 | 82.35 | 82.7 | 82.8
7.3 | 77.775 | 78.66 | 79.85 | 80.65 | 81.15 | 82 | 82.25 | 82.425
7.4 | 77.35 | 78.3 | 79.45 | 80.2 | 80.75 | 81.6 | 81.875 | 82.0667
7.5 | 76.9 | 77.84 | 79 | 79.8 | 80.35 | 81.2 | 81.5 | 81.7
7.6 | 76.5 | 77.375 | 78.64 | 79.45 | 80 | 80.9 | 81.2 | 81.4
7.7 | 76.1 | 76.95 | 78.25 | 79.025 | 79.6 | 80.45 | 80.8667 | 81.05
7.8 | 75.65 | 76.55 | 77.85 | 78.575 | 79.15 | 80.125 | 80.5 | 80.6
7.9 | 75.225 | 76.1 | 77.4 | 78.25 | 78.75 | 79.775 | 80.1667 | 80.3
8 | 74.775 | 75.7 | 77.0667 | 77.8 | 78.4333 | 79.425 | 79.7667 | 79.975
8.1 | 74.425 | 75.3 | 76.55 | 77.4 | 78 | 79.1 | 79.4 | 79.55
8.2 | 73.975 | 74.8 | 76.2333 | 77.1 | 77.6 | 78.75 | 79.1 | 79.175
8.3 | 73.6 | 74.52 | 75.85 | 76.7333 | 77.2 | 78.3 | 78.7 | 78.9
8.4 | 73.1 | 74.05 | 75.4 | 76.275 | 76.9 | 77.975 | 78.45 | 78.6
8.5 | 72.7333 | 73.675 | 75.05 | 75.925 | 76.5333 | 77.625 | 78.1 | 78.18
8.6 | 72.25 | 73.275 | 74.65 | 75.56 | 76.1667 | 77.2667 | 77.7333 | 77.9
8.7 | 71.9 | 72.9 | 74.25 | 75.15 | 75.775 | 77 | 77.35 | 77.55
8.8 | 71.5 | 72.5 | 73.8667 | 74.8 | 75.425 | 76.5667 | 76.95 | 77.2
8.9 | 71.1 | 72.1333 | 73.525 | 74.4 | 75.0667 | 76.2 | 76.6667 | 76.86
9 | 70.65 | 71.75 | 73 | 73.94 | 74.7 | 75.9 | 76.3 | 76.525
9.1 | 70.25 | 71.3 | 72.7 | 73.6667 | 74.3 | 75.5 | 76.025 | 76.15
9.2 | 69.86 | 70.9 | 72.3 | 73.2 | 73.95 | 75.175 | 75.675 | 75.85
9.3 | 69.45 | 70.4667 | 71.8667 | 72.875 | 73.6 | 74.85 | 75.35 | 75.5
9.4 | 69.1667 | 70.1333 | 71.525 | 72.525 | 73.2 | 74.475 | 75 | 75.2667
9.5 | 68.76 | 69.7 | 71.15 | 72.175 | 72.925 | 74.15 | 74.55 | 74.9
9.6 | 68.4 | 69.35 | 70.825 | 71.88 | 72.55 | 73.8667 | 74.3 | 74.5
9.7 | 67.9 | 68.9 | 70.36 | 71.5 | 72.125 | 73.55 | 73.95 | 74.1
9.8 | 67.675 | 68.5 | 70.05 | 71.025 | 71.75 | 73.2 | 73.6 | 73.875
9.9 | 67.35 | 68.2 | 69.7 | 70.7 | 71.38 | 72.8667 | 73.3 | 73.525
10 | 66.825 | 67.85 | 69.3 | 70.4 | 71.05 | 72.5 | 73 | 73.15
10.1 | 66.475 | 67.45 | 68.9 | 70.02 | 70.7 | 72.1 | 72.675 | 72.825
10.2 | 66.2 | 67.05 | 68.56 | 69.625 | 70.35 | 71.85 | 72.325 | 72.5
10.3 | 65.75 | 66.65 | 68.225 | 69.275 | 69.94 | 71.55 | 71.9667 | 72.2
10.4 | 65.375 | 66.28 | 67.875 | 68.9333 | 69.675 | 71.15 | 71.7 | 71.9
10.5 | 65 | 66 | 67.525 | 68.55 | 69.25 | 70.9 | 71.34 | 71.5667
10.6 | 64.625 | 65.6 | 67.075 | 68.225 | 69 | 70.55 | 71.1 | 71.2
10.7 | 64.26 | 65.3 | 66.8 | 67.9 | 68.7 | 70.18 | 70.7 | 70.9
10.8 | 63.84 | 64.9 | 66.475 | 67.52 | 68.2 | 69.85 | 70.3 | 70.55
10.9 | 63.6 | 64.475 | 66.125 | 67.15 | 67.95 | 69.6 | 70.125 | 70.3
11 | 63.2 | 64.15 | 65.675 | 66.775 | 67.6 | 69.275 | 69.775 | 69.9667
11.1 | 62.9 | 63.8 | 65.35 | 66.525 | 67.25 | 68.875 | 69.45 | 69.6667
11.2 | 62.46 | 63.45 | 65.075 | 66.075 | 66.925 | 68.575 | 69.1 | 69.4
11.3 | 62.15 | 63.1 | 64.7 | 65.8 | 66.575 | 68.25 | 68.75 | 69.05
11.4 | 61.8333 | 62.85 | 64.3 | 65.5 | 66.25 | 67.9 | 68.4 | 68.74
11.5 | 61.475 | 62.55 | 64.05 | 65.15 | 65.9333 | 67.65 | 68.2 | 68.45
11.6 | 61.15 | 62.15 | 63.7 | 64.8 | 65.625 | 67.3 | 67.8 | 68.12
11.7 | 60.775 | 61.75 | 63.35 | 64.5 | 65.275 | 66.95 | 67.55 | 67.8
11.8 | 60.5333 | 61.3667 | 63.02 | 64.1 | 64.9667 | 66.7 | 67.2667 | 67.55
11.9 | 60.0667 | 61.175 | 62.66 | 63.74 | 64.65 | 66.36 | 66.9 | 67.2333
12 | 59.825 | 60.84 | 62.35 | 63.475 | 64.25 | 66.05 | 66.675 | 66.9
12.1 | 59.4667 | 60.4 | 61.95 | 63.05 | 63.925 | 65.78 | 66.35 | 66.65
12.2 | 59.1 | 60.15 | 61.625 | 62.75 | 63.66 | 65.5 | 66.075 | 66.3
12.3 | 58.825 | 59.7 | 61.3667 | 62.5 | 63.3 | 65.1 | 65.7 | 66.0667
12.4 | 58.5 | 59.45 | 61.0333 | 62.2 | 63.05 | 64.8 | 65.5333 | 65.7
12.5 | 58.2 | 59.1667 | 60.7 | 61.85 | 62.7667 | 64.52 | 65.2 | 65.4333
12.6 | 57.8 | 58.78 | 60.3667 | 61.55 | 62.4333 | 64.25 | 64.9 | 65.0667
12.7 | 57.5 | 58.45 | 60.1 | 61.2333 | 62.05 | 63.925 | 64.575 | 64.8
12.8 | 57.25 | 58.1 | 59.7 | 60.9333 | 61.8 | 63.675 | 64.3 | 64.575
12.9 | 56.9 | 57.8 | 59.425 | 60.575 | 61.475 | 63.425 | 64 | 64.325
13 | 56.675 | 57.5 | 59.125 | 60.28 | 61.1667 | 63.075 | 63.7 | 63.875
13.1 | 56.35 | 57.25 | 58.8 | 60 | 60.9 | 62.85 | 63.4667 | 63.58
13.2 | 56 | 56.88 | 58.5 | 59.75 | 60.55 | 62.5 | 63.1 | 63.3
13.3 | 55.6 | 56.5333 | 58.1 | 59.4 | 60.2 | 62.2 | 62.85 | 62.9
13.4 | 55.3667 | 56.275 | 57.96 | 59.1 | 59.975 | 61.85 | 62.6 | 62.6
13.5 | 55.025 | 56 | 57.5667 | 58.75 | 59.675 | 61.65 | 62.275 | 62.32
13.6 | 54.775 | 55.7 | 57.275 | 58.425 | 59.325 | 61.275 | 62.025 | 62.075
13.7 | 54.45 | 55.3 | 57.025 | 58.15 | 58.975 | 61 | 61.675 | 61.8
13.8 | 54.1333 | 54.975 | 56.675 | 57.7 | 58.725 | 60.7 | 61.425 | 61.4667
13.9 | 53.9 | 54.725 | 56.3667 | 57.4667 | 58.5 | 60.4333 | 61.2 | 61.225
14 | 53.5 | 54.475 | 56.05 | 57.18 | 58.15 | 60.15 | 60.85 | 60.95
14.1 | 53.25 | 54.125 | 55.7667 | 56.95 | 57.9 | 59.8 | 60.5667 | 60.625
14.2 | 53 | 53.875 | 55.44 | 56.725 | 57.55 | 59.575 | 60.275 | 60.34
14.3 | 52.65 | 53.6 | 55.175 | 56.375 | 57.275 | 59.325 | 60 | 60.1
14.4 | 52.42 | 53.3 | 54.925 | 56.125 | 56.9667 | 59.075 | 59.76 | 59.8
14.5 | 52.1 | 53 | 54.575 | 55.85 | 56.7 | 58.74 | 59.475 | 59.6
14.6 | 51.825 | 52.6333 | 54.34 | 55.45 | 56.3667 | 58.5 | 59.225 | 59.3
14.7 | 51.5 | 52.35 | 54 | 55.175 | 56.15 | 58.25 | 58.95 | 59.0667
14.8 | 51.3 | 52.12 | 53.8 | 54.94 | 55.9 | 58 | 58.7 | 58.8
14.9 | 51.05 | 51.875 | 53.5 | 54.7 | 55.575 | 57.65 | 58.4 | 58.5
15 | 50.74 | 51.625 | 53.2 | 54.36 | 55.325 | 57.4 | 58.1667 | 58.2333
15.1 | 50.375 | 51.375 | 52.9 | 54.125 | 54.975 | 57.125 | 57.85 | 58
15.2 | 50.125 | 51.1333 | 52.675 | 53.85 | 54.725 | 56.8 | 57.725 | 57.7
15.3 | 49.875 | 50.8 | 52.425 | 53.575 | 54.4 | 56.6 | 57.375 | 57.4
15.4 | 49.625 | 50.5333 | 52.075 | 53.325 | 54.2333 | 56.35 | 57.125 | 57.1667
15.5 | 49.375 | 50.25 | 51.825 | 52.975 | 53.9 | 56 | 56.85 | 56.9
15.6 | 49 | 49.975 | 51.58 | 52.7 | 53.66 | 55.775 | 56.55 | 56.675
15.7 | 48.7 | 49.7 | 51.25 | 52.45 | 53.425 | 55.525 | 56.25 | 56.4333
15.8 | 48.55 | 49.4333 | 51 | 52.2 | 53.075 | 55.275 | 56.05 | 56.175
15.9 | 48.25 | 49.1667 | 50.75 | 51.875 | 52.8333 | 55.025 | 55.8 | 55.95
16 | 47.975 | 48.95 | 50.5 | 51.725 | 52.575 | 54.775 | 55.575 | 55.7
16.1 | 47.625 | 48.55 | 50.175 | 51.5333 | 52.35 | 54.45 | 55.325 | 55.45
16.2 | 47.5 | 48.4 | 49.9 | 51.275 | 52.05 | 54.25 | 55.05 | 55.2
16.3 | 47.25 | 48.16 | 49.6667 | 51.025 | 51.8 | 53.9 | 54.725 | 54.9333
16.4 | 47 | 47.9 | 49.4 | 50.8 | 51.525 | 53.7667 | 54.5 | 54.6333
16.5 | 46.76 | 47.575 | 49.15 | 50.5 | 51.275 | 53.475 | 54.26 | 54.475
16.6 | 46.525 | 47.325 | 48.8667 | 50.225 | 51.025 | 53.225 | 54 | 54.125
16.7 | 46.175 | 47.1667 | 48.65 | 49.925 | 50.775 | 52.975 | 53.7 | 53.975
16.8 | 45.925 | 46.85 | 48.4 | 49.68 | 50.5 | 52.75 | 53.575 | 53.7
16.9 | 45.75 | 46.62 | 48.175 | 49.5 | 50.35 | 52.5 | 53.225 | 53.4
17 | 45.5 | 46.4333 | 47.925 | 49.3 | 50 | 52.25 | 53.06 | 53.15
17.1 | 45.3 | 46.15 | 47.7 | 49 | 49.8 | 52 | 52.75 | 53
17.2 | 45.05 | 45.975 | 47.45 | 48.675 | 49.5 | 51.775 | 52.5 | 52.7
17.3 | 44.8 | 45.625 | 47.1667 | 48.475 | 49.3 | 51.5 | 52.375 | 52.5
17.4 | 44.575 | 45.375 | 46.9 | 48.275 | 49 | 51.3 | 52.1 | 52.225
17.5 | 44.25 | 45.15 | 46.65 | 48.025 | 48.725 | 51.1 | 51.875 | 51.975
17.6 | 44.075 | 44.9 | 46.4 | 47.775 | 48.475 | 50.775 | 51.625 | 51.725
17.7 | 43.84 | 44.65 | 46.16 | 47.525 | 48.3 | 50.525 | 51.375 | 51.55
17.8 | 43.5 | 44.4 | 45.925 | 47.3 | 47.975 | 50.275 | 51.14 | 51.3
17.9 | 43.3 | 44.16 | 45.675 | 47.075 | 47.74 | 50.125 | 50.9 | 50.95
18 | 43.1667 | 43.925 | 45.425 | 46.8667 | 47.5 | 49.9 | 50.65 | 50.8
18.1 | 42.8 | 43.725 | 45.175 | 46.6 | 47.4 | 49.6667 | 50.475 | 50.4
18.2 | 42.62 | 43.475 | 44.95 | 46.35 | 47.1333 | 49.3333 | 50.2 | 50.2
18.3 | 42.375 | 43.25 | 44.75 | 46.05 | 46.85 | 49.125 | 50.025 | 49.975
18.4 | 42.2 | 43 | 44.5 | 45.8 | 46.6 | 48.975 | 49.8333 | 49.725
18.5 | 41.925 | 42.75 | 44.3 | 45.65 | 46.375 | 48.725 | 49.5 | 49.5
18.6 | 41.7 | 42.6 | 44 | 45.425 | 46.2 | 48.4333 | 49.24 | 49.28
18.7 | 41.45 | 42.3 | 43.825 | 45.15 | 45.95 | 48.2 | 49.1 | 49
18.8 | 41.3 | 42.15 | 43.575 | 44.9 | 45.7 | 48.05 | 48.85 | 48.85
18.9 | 41.06 | 41.875 | 43.4 | 44.6 | 45.45 | 47.8 | 48.62 | 48.625
19 | 40.825 | 41.6 | 43.2 | 44.45 | 45.3 | 47.55 | 48.45 | 48.46
19.1 | 40.6 | 41.3667 | 42.9333 | 44.2 | 45.06 | 47.3 | 48.125 | 48.225
19.2 | 40.35 | 41.2 | 42.6667 | 44.06 | 44.825 | 47.05 | 47.9 | 47.9667
19.3 | 40.2 | 41 | 42.4 | 43.825 | 44.575 | 46.875 | 47.6333 | 47.8
19.4 | 39.975 | 40.8 | 42.25 | 43.575 | 44.325 | 46.6333 | 47.5 | 47.6
19.5 | 39.75 | 40.575 | 42.1 | 43.325 | 44.1 | 46.475 | 47.35 | 47.4
19.6 | 39.5 | 40.325 | 41.82 | 43.1 | 43.925 | 46.25 | 47.0333 | 47.1
19.7 | 39.3 | 40.175 | 41.525 | 42.9 | 43.7 | 46 | 46.8667 | 46.92
19.8 | 39.1 | 39.925 | 41.4 | 42.7 | 43.5 | 45.8 | 46.6 | 46.7
19.9 | 38.95 | 39.8 | 41.24 | 42.45 | 43.3 | 45.6 | 46.425 | 46.525
20 | 38.7 | 39.525 | 41 | 42.3 | 43.05 | 45.375 | 46.18 | 46.3
20.1 | 38.5 | 39.3 | 40.75 | 42.075 | 42.8 | 45.125 | 45.95 | 46.05
20.2 | 38.325 | 39.1 | 40.6 | 41.825 | 42.7 | 44.9 | 45.8 | 45.8
20.3 | 38.1 | 38.9 | 40.36 | 41.675 | 42.42 | 44.725 | 45.55 | 45.6
20.4 | 37.85 | 38.7 | 40.125 | 41.44 | 42.275 | 44.475 | 45.36 | 45.425
20.5 | 37.7 | 38.48 | 39.875 | 41.2 | 42 | 44.35 | 45.2 | 45.175
20.6 | 37.5 | 38.3 | 39.7 | 41 | 41.8 | 44.1 | 44.95 | 44.925
20.7 | 37.3 | 38.1 | 39.5 | 40.875 | 41.55 | 43.9333 | 44.8 | 44.8
20.8 | 37.175 | 37.85 | 39.3 | 40.625 | 41.42 | 43.65 | 44.575 | 44.6
20.9 | 36.925 | 37.6 | 39.1 | 40.375 | 41.2 | 43.42 | 44.325 | 44.4
21 | 36.7 | 37.475 | 38.9 | 40.125 | 41 | 43.2 | 44.15 | 44.2
21.1 | 36.5 | 37.3 | 38.66 | 39.975 | 40.725 | 43.04 | 43.9 | 44
21.2 | 36.35 | 37.125 | 38.525 | 39.75 | 40.575 | 42.9 | 43.75 | 43.775
21.3 | 36.175 | 36.875 | 38.275 | 39.5 | 40.35 | 42.65 | 43.5 | 43.6
21.4 | 36 | 36.725 | 38.125 | 39.35 | 40.2 | 42.5 | 43.375 | 43.4
21.5 | 35.8 | 36.48 | 37.98 | 39.2 | 39.95 | 42.275 | 43.125 | 43.15
21.6 | 35.575 | 36.35 | 37.75 | 38.96 | 39.8 | 42.02 | 42.875 | 43
21.7 | 35.4 | 36.125 | 37.55 | 38.725 | 39.6 | 41.8667 | 42.75 | 42.8667
21.8 | 35.2 | 35.975 | 37.3 | 38.575 | 39.4 | 41.6667 | 42.6 | 42.6
21.9 | 35.05 | 35.775 | 37.1 | 38.4 | 39.2 | 41.48 | 42.35 | 42.4
22 | 34.9 | 35.525 | 37 | 38.2 | 38.95 | 41.3 | 42.12 | 42.275
22.1 | 34.65 | 35.4 | 36.775 | 38 | 38.8 | 41.1 | 41.975 | 42.025
22.2 | 34.5 | 35.25 | 36.65 | 37.8 | 38.6 | 40.9 | 41.725 | 41.875
22.3 | 34.375 | 35 | 36.4 | 37.6667 | 38.35 | 40.72 | 41.575 | 41.7
22.4 | 34.125 | 34.8 | 36.15 | 37.5 | 38.2 | 40.575 | 41.4 | 41.4
22.5 | 33.975 | 34.6 | 36.02 | 37.275 | 38 | 40.325 | 41.2 | 41.3
22.6 | 33.825 | 34.46 | 35.8 | 37.1 | 37.825 | 40.175 | 41 | 41.1
22.7 | 33.6 | 34.325 | 35.7 | 36.875 | 37.68 | 39.925 | 40.75 | 40.9
22.8 | 33.5 | 34.1 | 35.5 | 36.725 | 37.55 | 39.775 | 40.625 | 40.7
22.9 | 33.25 | 33.925 | 35.325 | 36.475 | 37.3 | 39.6 | 40.475 | 40.5
23 | 33.1 | 33.8 | 35.075 | 36.325 | 37.1667 | 39.4333 | 40.3 | 40.35
23.1 | 32.875 | 33.6 | 34.9 | 36.1 | 36.9 | 39.2333 | 40.1 | 40.1667
23.2 | 32.725 | 33.45 | 34.75 | 36 | 36.8 | 39.025 | 39.9 | 39.975
23.3 | 32.58 | 33.22 | 34.6 | 35.8 | 36.6 | 38.8333 | 39.7 | 39.825
23.4 | 32.35 | 33.075 | 34.375 | 35.66 | 36.375 | 38.65 | 39.6 | 39.6
23.5 | 32.2 | 32.94 | 34.2 | 35.425 | 36.2 | 38.55 | 39.35 | 39.45
23.6 | 32.05 | 32.7 | 34.1 | 35.275 | 36 | 38.3 | 39.2 | 39.3
23.7 | 31.9 | 32.55 | 33.9 | 35.1 | 35.85 | 38.15 | 38.975 | 39.1
23.8 | 31.8 | 32.4 | 33.7 | 34.9 | 35.6 | 37.9 | 38.825 | 38.9
23.9 | 31.575 | 32.25 | 33.5333 | 34.75 | 35.5 | 37.75 | 38.6 | 38.725
24 | 31.375 | 32.1 | 33.3667 | 34.575 | 35.3 | 37.625 | 38.425 | 38.575
24.1 | 31.225 | 31.875 | 33.22 | 34.3 | 35.175 | 37.475 | 38.3 | 38.4
24.2 | 31.1 | 31.8 | 33 | 34.2 | 35 | 37.3 | 38.1 | 38.2
24.3 | 31 | 31.575 | 32.8 | 34.1 | 34.8 | 37.1 | 37.9 | 38.08
24.4 | 30.75 | 31.4 | 32.7 | 33.9 | 34.625 | 36.9 | 37.8 | 37.9
24.5 | 30.65 | 31.25 | 32.575 | 33.75 | 34.48 | 36.8 | 37.6333 | 37.7
24.6 | 30.4 | 31.1 | 32.4 | 33.5 | 34.3 | 36.55 | 37.475 | 37.55
24.7 | 30.275 | 30.95 | 32.175 | 33.36 | 34.1 | 36.4 | 37.3 | 37.3
24.8 | 30.125 | 30.8 | 32.1 | 33.2 | 33.975 | 36.22 | 37.1 | 37.15
24.9 | 29.975 | 30.625 | 31.9 | 33.025 | 33.8 | 36.075 | 36.95 | 37.02
25 | 29.825 | 30.4 | 31.8 | 32.875 | 33.675 | 35.925 | 36.8 | 36.875
25.1 | 29.7 | 30.3 | 31.6 | 32.725 | 33.425 | 35.7 | 36.55 | 36.7
25.2 | 29.55 | 30.125 | 31.4 | 32.5 | 33.275 | 35.6 | 36.42 | 36.6
25.3 | 29.4 | 30 | 31.3 | 32.4 | 33.15 | 35.4 | 36.275 | 36.35
25.4 | 29.2333 | 29.9 | 31.1 | 32.3 | 33 | 35.2 | 36.1 | 36.2
25.5 | 29.1 | 29.7 | 31 | 32.05 | 32.85 | 35.05 | 35.975 | 36.0667
25.6 | 28.9 | 29.5667 | 30.8 | 31.92 | 32.7 | 34.9 | 35.775 | 35.9
25.7 | 28.8 | 29.4 | 30.6667 | 31.775 | 32.5 | 34.7 | 35.6 | 35.7
25.8 | 28.625 | 29.275 | 30.475 | 31.6 | 32.325 | 34.52 | 35.4 | 35.6
25.9 | 28.5 | 29.14 | 30.3667 | 31.475 | 32.175 | 34.4 | 35.3 | 35.4
26 | 28.3 | 29 | 30.2 | 31.3 | 32 | 34.275 | 35.1 | 35.225
26.1 | 28.2 | 28.8 | 30.0333 | 31.2 | 31.9 | 34.025 | 34.95 | 35.1
26.2 | 28 | 28.7 | 29.9 | 31 | 31.7 | 33.9 | 34.7667 | 34.9
26.3 | 27.9333 | 28.5 | 29.75 | 30.85 | 31.6 | 33.75 | 34.625 | 34.78
26.4 | 27.8 | 28.4 | 29.6 | 30.7 | 31.42 | 33.6 | 34.475 | 34.6
26.5 | 27.6 | 28.2 | 29.4 | 30.575 | 31.275 | 33.46 | 34.34 | 34.45
26.6 | 27.5 | 28.1 | 29.3 | 30.4 | 31.125 | 33.325 | 34.2 | 34.3
26.7 | 27.375 | 27.9 | 29.2 | 30.2 | 30.975 | 33.175 | 34 | 34.1
26.8 | 27.2 | 27.8667 | 28.975 | 30.1 | 30.85 | 33.025 | 33.82 | 34
26.9 | 27.075 | 27.675 | 28.825 | 30 | 30.6 | 32.875 | 33.7 | 33.8
27 | 26.95 | 27.525 | 28.7667 | 29.8 | 30.5 | 32.7 | 33.525 | 33.675
27.1 | 26.7667 | 27.4 | 28.6 | 29.6667 | 30.4 | 32.6 | 33.4 | 33.6
27.2 | 26.6 | 27.3 | 28.45 | 29.5 | 30.25 | 32.4 | 33.24 | 33.375
27.3 | 26.5 | 27.1333 | 28.3 | 29.375 | 30.1 | 32.2667 | 33.1 | 33.24
27.4 | 26.4 | 27 | 28.1 | 29.2 | 29.975 | 32.1 | 32.95 | 33.05
27.5 | 26.275 | 26.85 | 28 | 29.08 | 29.8 | 31.975 | 32.8 | 32.92
27.6 | 26.125 | 26.7 | 27.9 | 28.9 | 29.675 | 31.825 | 32.6 | 32.8
27.7 | 26 | 26.525 | 27.8 | 28.8 | 29.5 | 31.68 | 32.525 | 32.625
27.8 | 25.9 | 26.475 | 27.6 | 28.675 | 29.3667 | 31.5 | 32.375 | 32.5
27.9 | 25.7 | 26.3 | 27.475 | 28.5 | 29.2 | 31.35 | 32.225 | 32.325
28 | 25.6 | 26.2 | 27.325 | 28.375 | 29.1 | 31.2 | 32.1 | 32.2
28.1 | 25.5 | 26.0333 | 27.2 | 28.2 | 28.9 | 31.1 | 31.9 | 32.04
28.2 | 25.4 | 25.95 | 27.05 | 28.1 | 28.75 | 30.9 | 31.75 | 31.9
28.3 | 25.2 | 25.8 | 26.9 | 28 | 28.675 | 30.8 | 31.65 | 31.8
28.4 | 25.1 | 25.6 | 26.8 | 27.8 | 28.525 | 30.675 | 31.5 | 31.6
28.5 | 24.975 | 25.5 | 26.6 | 27.7 | 28.375 | 30.525 | 31.325 | 31.475
28.6 | 24.8 | 25.4 | 26.5 | 27.5 | 28.225 | 30.3 | 31.2 | 31.325
28.7 | 24.7 | 25.275 | 26.4 | 27.4 | 28.1 | 30.25 | 31.025 | 31.2
28.8 | 24.6 | 25.125 | 26.2 | 27.3 | 28 | 30.1 | 30.9 | 31.1
28.9 | 24.475 | 25 | 26.15 | 27.125 | 27.8667 | 29.9667 | 30.8 | 30.9
29 | 24.325 | 24.9 | 26 | 27 | 27.75 | 29.8 | 30.6 | 30.8
29.1 | 24.275 | 24.725 | 25.9 | 26.8667 | 27.6 | 29.7 | 30.575 | 30.68
29.2 | 24.1 | 24.625 | 25.7333 | 26.7667 | 27.475 | 29.575 | 30.4 | 30.55
29.3 | 24 | 24.55 | 25.6 | 26.6 | 27.3 | 29.425 | 30.2 | 30.4
29.4 | 23.9 | 24.4 | 25.5 | 26.5 | 27.2 | 29.275 | 30.14 | 30.3
29.5 | 23.76 | 24.3 | 25.3333 | 26.4 | 27 | 29.14 | 30 | 30.1
29.6 | 23.625 | 24.1 | 25.2 | 26.275 | 26.9 | 29.0667 | 29.8333 | 30
29.7 | 23.5 | 24.05 | 25.1 | 26.1 | 26.8 | 28.9 | 29.7 | 29.86
29.8 | 23.4 | 23.9 | 25 | 25.975 | 26.7 | 28.7 | 29.5 | 29.7
29.9 | 23.275 | 23.8 | 24.9 | 25.9 | 26.56 | 28.6 | 29.4 | 29.525
30 | 23.15 | 23.6333 | 24.725 | 25.8 | 26.4 | 28.5 | 29.3 | 29.4

## Open Field Profiles at 2.4cm
Lateral Profiles
Open field
SSD=100cm
Depth=2.4cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.4 | [7 empty] | 1.1
-28.3 | [7 empty] | 1.1
-28.2 | [7 empty] | 1.1
-28.1 | [7 empty] | 1.1
-28 | [7 empty] | 1.16
-27.9 | [7 empty] | 1.2
-27.8 | [7 empty] | 1.2
-27.7 | [7 empty] | 1.2
-27.6 | [7 empty] | 1.2
-27.5 | [7 empty] | 1.2
-27.4 | [7 empty] | 1.2
-27.3 | [7 empty] | 1.26
-27.2 | [7 empty] | 1.3
-27.1 | [7 empty] | 1.3
-27 | [7 empty] | 1.3
-26.9 | [7 empty] | 1.3
-26.8 | [7 empty] | 1.3
-26.7 | [7 empty] | 1.36
-26.6 | [7 empty] | 1.4
-26.5 | [7 empty] | 1.4
-26.4 | [7 empty] | 1.4
-26.3 | [7 empty] | 1.4
-26.2 | [7 empty] | 1.46
-26.1 | [7 empty] | 1.5
-26 | [7 empty] | 1.5
-25.9 | [7 empty] | 1.5
-25.8 | [7 empty] | 1.5
-25.7 | [7 empty] | 1.56
-25.6 | [7 empty] | 1.6
-25.5 | [7 empty] | 1.6
-25.4 | [7 empty] | 1.6
-25.3 | [7 empty] | 1.66
-25.2 | [7 empty] | 1.7
-25.1 | [7 empty] | 1.7
-25 | [7 empty] | 1.7
-24.9 | [7 empty] | 1.76
-24.8 | [7 empty] | 1.8
-24.7 | [7 empty] | 1.8
-24.6 | [7 empty] | 1.86
-24.5 | [7 empty] | 1.9
-24.4 | [7 empty] | 1.9
-24.3 | [7 empty] | 1.96
-24.2 | [7 empty] | 2
-24.1 | [7 empty] | 2
-24 | [7 empty] | 2.06
-23.9 | [7 empty] | 2.1
-23.8 | [7 empty] | 2.1
-23.7 | [7 empty] | 2.16
-23.6 | [7 empty] | 2.2
-23.5 | [7 empty] | 2.26
-23.4 | [7 empty] | 2.3
-23.3 | [7 empty] | 2.3
-23.2 | [7 empty] | 2.36
-23.1 | [7 empty] | 2.4
-23 | [7 empty] | 2.46
-22.9 | [7 empty] | 2.56
-22.8 | [7 empty] | 2.6
-22.7 | [7 empty] | 2.66
-22.6 | [7 empty] | 2.7
-22.5 | [7 empty] | 2.76
-22.4 | [7 empty] | 2.86
-22.3 | [7 empty] | 2.9
-22.2 | [7 empty] | 2.96
-22.1 | [7 empty] | 3.06
-22 | [7 empty] | 3.16
-21.9 | [7 empty] | 3.26
-21.8 | [7 empty] | 3.36
-21.7 | [7 empty] | 3.52
-21.6 | [7 empty] | 3.72
-21.5 | [7 empty] | 3.92
-21.4 | [7 empty] | 4.12
-21.3 | [7 empty] | 4.44
-21.2 | [7 empty] | 4.9
-21.1 | [7 empty] | 5.52
-21 | [7 empty] | 6.4
-20.9 | [7 empty] | 7.76
-20.8 | [7 empty] | 9.9
-20.7 | [7 empty] | 12.94
-20.6 | [7 empty] | 16.64
-20.5 | [7 empty] | 20.78
-20.4 | [7 empty] | 25.02
-20.3 | [6 empty] | 1.8 | 28.98
-20.2 | [6 empty] | 1.8 | 32.42
-20.1 | [6 empty] | 1.8 | 35.08
-20 | [6 empty] | 1.9 | 36.84
-19.9 | [6 empty] | 1.9 | 38
-19.8 | [6 empty] | 1.9 | 38.76
-19.7 | [6 empty] | 1.96 | 39.3
-19.6 | [6 empty] | 2 | 39.74
-19.5 | [6 empty] | 2 | 40.14
-19.4 | [6 empty] | 2 | 40.48
-19.3 | [6 empty] | 2.1 | 40.78
-19.2 | [6 empty] | 2.1 | 41.02
-19.1 | [6 empty] | 2.1 | 41.28
-19 | [6 empty] | 2.2 | 41.52
-18.9 | [6 empty] | 2.2 | 41.78
-18.8 | [6 empty] | 2.26 | 42.02
-18.7 | [6 empty] | 2.3 | 42.22
-18.6 | [6 empty] | 2.36 | 42.42
-18.5 | [6 empty] | 2.4 | 42.62
-18.4 | [6 empty] | 2.4 | 42.82
-18.3 | [6 empty] | 2.5 | 43.02
-18.2 | [6 empty] | 2.5 | 43.22
-18.1 | [6 empty] | 2.56 | 43.42
-18 | [6 empty] | 2.6 | 43.62
-17.9 | [6 empty] | 2.7 | 43.82
-17.8 | [6 empty] | 2.7 | 44.02
-17.7 | [6 empty] | 2.8 | 44.22
-17.6 | [6 empty] | 2.86 | 44.48
-17.5 | [6 empty] | 2.9 | 44.72
-17.4 | [6 empty] | 3 | 44.92
-17.3 | [6 empty] | 3.06 | 45.12
-17.2 | [6 empty] | 3.16 | 45.32
-17.1 | [6 empty] | 3.2 | 45.52
-17 | [6 empty] | 3.3 | 45.72
-16.9 | [6 empty] | 3.4 | 45.92
-16.8 | [6 empty] | 3.5 | 46.12
-16.7 | [6 empty] | 3.66 | 46.32
-16.6 | [6 empty] | 3.76 | 46.58
-16.5 | [6 empty] | 3.96 | 46.76
-16.4 | [6 empty] | 4.16 | 46.98
-16.3 | [6 empty] | 4.46 | 47.22
-16.2 | [6 empty] | 4.82 | 47.42
-16.1 | [6 empty] | 5.28 | 47.62
-16 | [6 empty] | 5.88 | 47.82
-15.9 | [6 empty] | 6.8 | 48.08
-15.8 | [6 empty] | 8.24 | 48.26
-15.7 | [6 empty] | 10.54 | 48.48
-15.6 | [6 empty] | 14 | 48.72
-15.5 | [6 empty] | 18.38 | 48.92
-15.4 | [6 empty] | 23.42 | 49.12
-15.3 | [6 empty] | 28.72 | 49.38
-15.2 | [5 empty] | 1.6 | 33.8 | 49.62
-15.1 | [5 empty] | 1.6 | 38.46 | 49.82
-15 | [5 empty] | 1.6 | 42.06 | 50.02
-14.9 | [5 empty] | 1.64 | 44.6 | 50.28
-14.8 | [5 empty] | 1.7 | 46.32 | 50.52
-14.7 | [5 empty] | 1.7 | 47.4 | 50.72
-14.6 | [5 empty] | 1.7 | 48.24 | 50.98
-14.5 | [5 empty] | 1.8 | 48.88 | 51.22
-14.4 | [5 empty] | 1.8 | 49.42 | 51.48
-14.3 | [5 empty] | 1.8 | 49.88 | 51.72
-14.2 | [5 empty] | 1.9 | 50.26 | 51.92
-14.1 | [5 empty] | 1.9 | 50.62 | 52.18
-14 | [5 empty] | 1.9 | 51.02 | 52.42
-13.9 | [5 empty] | 2 | 51.32 | 52.68
-13.8 | [5 empty] | 2 | 51.62 | 52.92
-13.7 | [5 empty] | 2.04 | 51.92 | 53.18
-13.6 | [5 empty] | 2.1 | 52.22 | 53.42
-13.5 | [5 empty] | 2.1 | 52.52 | 53.62
-13.4 | [5 empty] | 2.2 | 52.76 | 53.88
-13.3 | [5 empty] | 2.2 | 53.12 | 54.18
-13.2 | [5 empty] | 2.24 | 53.36 | 54.42
-13.1 | [5 empty] | 2.3 | 53.66 | 54.68
-13 | [5 empty] | 2.34 | 53.96 | 54.92
-12.9 | [5 empty] | 2.4 | 54.22 | 55.18
-12.8 | [5 empty] | 2.5 | 54.52 | 55.42
-12.7 | [5 empty] | 2.5 | 54.86 | 55.68
-12.6 | [5 empty] | 2.6 | 55.06 | 55.98
-12.5 | [5 empty] | 2.64 | 55.36 | 56.22
-12.4 | [5 empty] | 2.7 | 55.66 | 56.48
-12.3 | [5 empty] | 2.8 | 55.92 | 56.78
-12.2 | [5 empty] | 2.9 | 56.16 | 57.08
-12.1 | [5 empty] | 2.94 | 56.52 | 57.32
-12 | [5 empty] | 3.04 | 56.82 | 57.58
-11.9 | [5 empty] | 3.14 | 57.06 | 57.82
-11.8 | [5 empty] | 3.24 | 57.42 | 58.08
-11.7 | [5 empty] | 3.4 | 57.66 | 58.38
-11.6 | [5 empty] | 3.54 | 57.92 | 58.68
-11.5 | [5 empty] | 3.74 | 58.22 | 58.98
-11.4 | [5 empty] | 3.94 | 58.56 | 59.22
-11.3 | [5 empty] | 4.18 | 58.86 | 59.54
-11.2 | [5 empty] | 4.48 | 59.16 | 59.88
-11.1 | [5 empty] | 4.88 | 59.46 | 60.12
-11 | [5 empty] | 5.42 | 59.76 | 60.38
-10.9 | [5 empty] | 6.2 | 60.06 | 60.68
-10.8 | [5 empty] | 7.24 | 60.32 | 60.98
-10.7 | [5 empty] | 8.84 | 60.66 | 61.28
-10.6 | [5 empty] | 11.52 | 60.96 | 61.58
-10.5 | [5 empty] | 15.84 | 61.36 | 61.88
-10.4 | [5 empty] | 21.44 | 61.62 | 62.18
-10.3 | [5 empty] | 28.1 | 61.92 | 62.48
-10.2 | [5 empty] | 35.14 | 62.26 | 62.84
-10.1 |  |  |  |  | 1 | 42.08 | 62.52 | 63.18
-10 |  |  |  |  | 1 | 48.22 | 62.86 | 63.48
-9.9 |  |  |  |  | 1 | 53.24 | 63.22 | 63.78
-9.8 |  |  |  |  | 1.1 | 56.62 | 63.52 | 64.08
-9.7 |  |  |  |  | 1.1 | 58.76 | 63.86 | 64.38
-9.6 |  |  |  |  | 1.1 | 60.18 | 64.22 | 64.74
-9.5 |  |  |  |  | 1.1 | 61.3 | 64.62 | 65.08
-9.4 |  |  |  |  | 1.2 | 62.16 | 64.92 | 65.38
-9.3 |  |  |  |  | 1.2 | 62.82 | 65.22 | 65.74
-9.2 |  |  |  |  | 1.2 | 63.52 | 65.62 | 66.08
-9.1 |  |  |  |  | 1.2 | 64.08 | 65.92 | 66.38
-9 |  |  |  | 0.8 | 1.3 | 64.62 | 66.26 | 66.74
-8.9 |  |  |  | 0.9 | 1.3 | 65.08 | 66.62 | 67.08
-8.8 |  |  |  | 0.9 | 1.3 | 65.58 | 66.92 | 67.38
-8.7 |  |  |  | 0.9 | 1.3 | 65.98 | 67.36 | 67.74
-8.6 |  |  |  | 0.9 | 1.4 | 66.48 | 67.72 | 68.08
-8.5 |  |  |  | 0.9 | 1.4 | 66.88 | 68.02 | 68.44
-8.4 |  |  |  | 0.9 | 1.4 | 67.28 | 68.32 | 68.84
-8.3 |  |  |  | 1 | 1.5 | 67.74 | 68.82 | 69.18
-8.2 |  |  |  | 1 | 1.5 | 68.08 | 69.12 | 69.48
-8.1 |  |  |  | 1 | 1.5 | 68.58 | 69.52 | 69.84
-8 |  |  | 0.6 | 1 | 1.6 | 68.98 | 69.92 | 70.24
-7.9 |  |  | 0.6 | 1.06 | 1.6 | 69.38 | 70.22 | 70.58
-7.8 |  |  | 0.66 | 1.1 | 1.7 | 69.78 | 70.62 | 70.94
-7.7 |  |  | 0.7 | 1.1 | 1.7 | 70.18 | 71.02 | 71.34
-7.6 |  |  | 0.7 | 1.1 | 1.8 | 70.58 | 71.46 | 71.68
-7.5 |  |  | 0.7 | 1.1 | 1.8 | 70.98 | 71.82 | 72.04
-7.4 |  |  | 0.7 | 1.2 | 1.9 | 71.44 | 72.16 | 72.44
-7.3 |  |  | 0.7 | 1.2 | 2 | 71.88 | 72.52 | 72.84
-7.2 |  |  | 0.7 | 1.2 | 2 | 72.28 | 72.92 | 73.24
-7.1 |  |  | 0.76 | 1.3 | 2.1 | 72.68 | 73.32 | 73.64
-7 |  | 0.4 | 0.8 | 1.3 | 2.2 | 73.08 | 73.68 | 74.04
-6.9 |  | 0.4 | 0.8 | 1.36 | 2.3 | 73.52 | 74.12 | 74.38
-6.8 |  | 0.4 | 0.8 | 1.4 | 2.4 | 73.98 | 74.56 | 74.8
-6.7 |  | 0.4 | 0.8 | 1.4 | 2.5 | 74.32 | 74.92 | 75.24
-6.6 |  | 0.42 | 0.86 | 1.5 | 2.7 | 74.78 | 75.32 | 75.64
-6.5 | 0.3 | 0.5 | 0.9 | 1.5 | 2.8 | 75.18 | 75.72 | 75.98
-6.4 | 0.3 | 0.5 | 0.9 | 1.6 | 3 | 75.64 | 76.12 | 76.4
-6.3 | 0.3 | 0.5 | 0.9 | 1.66 | 3.2 | 76.08 | 76.52 | 76.84
-6.2 | 0.3 | 0.5 | 0.9 | 1.7 | 3.5 | 76.48 | 76.92 | 77.24
-6.1 | 0.3 | 0.5 | 1 | 1.8 | 3.9 | 76.92 | 77.32 | 77.64
-6 | 0.34 | 0.5 | 1 | 1.86 | 4.4 | 77.38 | 77.76 | 78.04
-5.9 | 0.4 | 0.5 | 1 | 1.96 | 5 | 77.78 | 78.22 | 78.44
-5.8 | 0.4 | 0.5 | 1.1 | 2.06 | 5.9 | 78.28 | 78.62 | 78.9
-5.7 | 0.4 | 0.5 | 1.1 | 2.16 | 7.2 | 78.68 | 79.12 | 79.34
-5.6 | 0.4 | 0.6 | 1.1 | 2.3 | 9 | 79.18 | 79.52 | 79.74
-5.5 | 0.4 | 0.6 | 1.2 | 2.46 | 12.1 | 79.58 | 79.92 | 80.14
-5.4 | 0.4 | 0.6 | 1.2 | 2.66 | 17.2 | 80.02 | 80.38 | 80.54
-5.3 | 0.4 | 0.6 | 1.3 | 2.86 | 24.5 | 80.54 | 80.82 | 81.06
-5.2 | 0.4 | 0.6 | 1.3 | 3.12 | 33.3 | 80.92 | 81.28 | 81.54
-5.1 | 0.4 | 0.62 | 1.4 | 3.52 | 42.9 | 81.38 | 81.82 | 81.94
-5 | 0.4 | 0.7 | 1.46 | 3.98 | 52.6 | 81.82 | 82.22 | 82.4
-4.9 | 0.4 | 0.7 | 1.5 | 4.58 | 61.3 | 82.32 | 82.66 | 82.84
-4.8 | 0.44 | 0.7 | 1.6 | 5.4 | 68.7 | 82.78 | 83.06 | 83.3
-4.7 | 0.5 | 0.7 | 1.7 | 6.62 | 73.8 | 83.34 | 83.52 | 83.74
-4.6 | 0.5 | 0.72 | 1.86 | 8.36 | 77 | 83.78 | 84.02 | 84.2
-4.5 | 0.5 | 0.8 | 1.96 | 11.12 | 79.2 | 84.18 | 84.48 | 84.64
-4.4 | 0.5 | 0.8 | 2.16 | 15.84 | 80.8 | 84.72 | 84.98 | 85.1
-4.3 | 0.5 | 0.9 | 2.36 | 23 | 82.1 | 85.12 | 85.38 | 85.54
-4.2 | 0.5 | 0.9 | 2.62 | 31.86 | 83 | 85.62 | 85.82 | 85.94
-4.1 | 0.6 | 0.92 | 2.92 | 41.88 | 83.9 | 86.12 | 86.32 | 86.4
-4 | 0.6 | 1 | 3.38 | 52.22 | 84.7 | 86.52 | 86.72 | 86.9
-3.9 | 0.6 | 1.1 | 3.98 | 61.86 | 85.5 | 86.98 | 87.22 | 87.34
-3.8 | 0.6 | 1.12 | 4.74 | 70.32 | 86.2 | 87.48 | 87.72 | 87.8
-3.7 | 0.7 | 1.22 | 5.86 | 76.6 | 86.8 | 88.02 | 88.12 | 88.3
-3.6 | 0.7 | 1.32 | 7.44 | 80.56 | 87.4 | 88.48 | 88.62 | 88.74
-3.5 | 0.74 | 1.42 | 9.9 | 83.16 | 87.9 | 88.88 | 89.02 | 89.2
-3.4 | 0.8 | 1.6 | 14.02 | 84.92 | 88.6 | 89.38 | 89.48 | 89.7
-3.3 | 0.84 | 1.82 | 20.58 | 86.32 | 89.1 | 89.82 | 89.96 | 90.14
-3.2 | 0.94 | 2.02 | 29.34 | 87.5 | 89.6 | 90.32 | 90.48 | 90.6
-3.1 | 1 | 2.32 | 39.84 | 88.46 | 90.2 | 90.78 | 90.92 | 91.04
-3 | 1.14 | 2.74 | 50.74 | 89.24 | 90.7 | 91.22 | 91.42 | 91.44
-2.9 | 1.24 | 3.26 | 61.38 | 90.04 | 91.1 | 91.68 | 91.82 | 91.9
-2.8 | 1.44 | 4 | 71.06 | 90.74 | 91.7 | 92.08 | 92.22 | 92.34
-2.7 | 1.64 | 5.04 | 78.78 | 91.34 | 92.2 | 92.58 | 92.68 | 92.8
-2.6 | 1.88 | 6.6 | 83.7 | 91.98 | 92.7 | 93.08 | 93.12 | 93.24
-2.5 | 2.28 | 8.92 | 86.88 | 92.58 | 93.2 | 93.48 | 93.58 | 93.64
-2.4 | 2.82 | 12.72 | 89.1 | 93.12 | 93.7 | 93.92 | 93.92 | 94.04
-2.3 | 3.5 | 19.02 | 90.72 | 93.68 | 94.1 | 94.38 | 94.42 | 94.5
-2.2 | 4.44 | 27.8 | 91.96 | 94.12 | 94.6 | 94.74 | 94.82 | 94.88
-2.1 | 5.96 | 38.46 | 93.04 | 94.62 | 95 | 95.18 | 95.32 | 95.3
-2 | 8.2 | 50.16 | 93.8 | 95.08 | 95.4 | 95.54 | 95.62 | 95.68
-1.9 | 11.96 | 61.72 | 94.54 | 95.58 | 95.9 | 95.98 | 95.96 | 96.04
-1.8 | 18.06 | 72.2 | 95.08 | 96.02 | 96.3 | 96.28 | 96.38 | 96.44
-1.7 | 27.1 | 80.6 | 95.58 | 96.32 | 96.7 | 96.64 | 96.76 | 96.84
-1.6 | 37.58 | 86.76 | 96.28 | 96.66 | 96.9 | 97.08 | 97.06 | 97.18
-1.5 | 49.64 | 90.58 | 96.72 | 97.12 | 97.3 | 97.38 | 97.32 | 97.48
-1.4 | 61.54 | 93 | 97.22 | 97.48 | 97.7 | 97.74 | 97.66 | 97.78
-1.3 | 72.28 | 94.64 | 97.56 | 97.92 | 97.9 | 97.98 | 98.02 | 98.08
-1.2 | 81.7 | 95.82 | 98.02 | 98.16 | 98.3 | 98.24 | 98.26 | 98.32
-1.1 | 88.4 | 96.78 | 98.32 | 98.42 | 98.5 | 98.58 | 98.52 | 98.58
-1 | 92.3 | 97.46 | 98.56 | 98.66 | 98.8 | 98.78 | 98.86 | 98.82
-0.9 | 94.8 | 98.04 | 98.92 | 98.92 | 99.1 | 99.04 | 99.06 | 99.02
-0.8 | 96.44 | 98.46 | 99.16 | 99.1 | 99.3 | 99.24 | 99.22 | 99.22
-0.7 | 97.56 | 98.92 | 99.36 | 99.26 | 99.4 | 99.44 | 99.46 | 99.42
-0.6 | 98.36 | 99.22 | 99.56 | 99.56 | 99.5 | 99.64 | 99.56 | 99.62
-0.5 | 98.98 | 99.52 | 99.7 | 99.66 | 99.7 | 99.74 | 99.7 | 99.76
-0.4 | 99.38 | 99.7 | 99.8 | 99.8 | 99.8 | 99.84 | 99.8 | 99.86
-0.3 | 99.74 | 99.82 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9
-0.2 | 99.9 | 99.92 | 99.9 | 99.96 | 100 | 100 | 100 | 99.96
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 99.9 | 99.9 | 99.94 | 100 | 100 | 99.94 | 99.94
0.3 | 99.76 | 99.8 | 99.9 | 99.84 | 99.9 | 99.9 | 99.9 | 99.9
0.4 | 99.42 | 99.7 | 99.74 | 99.74 | 99.8 | 99.8 | 99.8 | 99.84
0.5 | 99.02 | 99.56 | 99.7 | 99.64 | 99.7 | 99.7 | 99.7 | 99.74
0.6 | 98.4 | 99.26 | 99.44 | 99.54 | 99.6 | 99.56 | 99.6 | 99.58
0.7 | 97.6 | 98.96 | 99.24 | 99.28 | 99.3 | 99.36 | 99.44 | 99.38
0.8 | 96.48 | 98.54 | 99.14 | 99.1 | 99.2 | 99.16 | 99.24 | 99.18
0.9 | 94.86 | 98.06 | 98.78 | 98.94 | 99 | 98.96 | 99.08 | 99.04
1 | 92.4 | 97.54 | 98.54 | 98.64 | 98.8 | 98.76 | 98.84 | 98.8364
1.1 | 88.52 | 96.92 | 98.28 | 98.44 | 98.5 | 98.56 | 98.64 | 98.55
1.2 | 81.96 | 95.98 | 97.98 | 98.14 | 98.1 | 98.26 | 98.38 | 98.3
1.3 | 72.56 | 94.76 | 97.58 | 97.78 | 97.8 | 98.02 | 98.08 | 98.05
1.4 | 61.86 | 93.2 | 97.18 | 97.44 | 97.5 | 97.62 | 97.78 | 97.75
1.5 | 49.96 | 90.9 | 96.72 | 97.08 | 97.2 | 97.36 | 97.48 | 97.45
1.6 | 37.86 | 87.2 | 96.28 | 96.68 | 96.9 | 97.06 | 97.08 | 97.15
1.7 | 27.34 | 81.26 | 95.78 | 96.34 | 96.5 | 96.72 | 96.78 | 96.8
1.8 | 18.28 | 73.14 | 95.22 | 95.92 | 96.1 | 96.32 | 96.38 | 96.4
1.9 | 12.08 | 62.86 | 94.58 | 95.48 | 95.6 | 95.92 | 96.02 | 96
2 | 8.3 | 51.24 | 93.86 | 94.98 | 95.2 | 95.52 | 95.58 | 95.65
2.1 | 5.98 | 39.66 | 93 | 94.58 | 94.9 | 95.12 | 95.28 | 95.3
2.2 | 4.5 | 28.74 | 92.1 | 94.08 | 94.4 | 94.72 | 94.88 | 94.9
2.3 | 3.54 | 19.84 | 90.88 | 93.62 | 94 | 94.26 | 94.48 | 94.5
2.4 | 2.78 | 13.14 | 89.26 | 93.08 | 93.5 | 93.92 | 94.02 | 94.1
2.5 | 2.32 | 9.22 | 87.18 | 92.52 | 93 | 93.48 | 93.58 | 93.65
2.6 | 1.92 | 6.72 | 84.06 | 91.92 | 92.6 | 92.98 | 93.18 | 93.2
2.7 | 1.66 | 5.18 | 79.1 | 91.36 | 92.1 | 92.48 | 92.68 | 92.8
2.8 | 1.46 | 4.1 | 71.84 | 90.72 | 91.5 | 92.02 | 92.28 | 92.35
2.9 | 1.26 | 3.34 | 62.6 | 90.06 | 91 | 91.56 | 91.86 | 91.9
3 | 1.1 | 2.76 | 51.78 | 89.36 | 90.5 | 91.12 | 91.38 | 91.5
3.1 | 1 | 2.36 | 40.86 | 88.5 | 90 | 90.72 | 90.94 | 91.05
3.2 | 0.9 | 2.06 | 30.46 | 87.44 | 89.5 | 90.22 | 90.52 | 90.55
3.3 | 0.86 | 1.78 | 21.32 | 86.34 | 88.9 | 89.78 | 89.98 | 90.1
3.4 | 0.8 | 1.58 | 14.58 | 84.88 | 88.4 | 89.32 | 89.54 | 89.7
3.5 | 0.76 | 1.4 | 10.28 | 83.1 | 87.8 | 88.88 | 89.08 | 89.25
3.6 | 0.7 | 1.3 | 7.7 | 80.44 | 87.2 | 88.28 | 88.68 | 88.75
3.7 | 0.7 | 1.2 | 5.98 | 76.3 | 86.6 | 87.82 | 88.18 | 88.25
3.8 | 0.6 | 1.1 | 4.9 | 69.92 | 86 | 87.38 | 87.68 | 87.8
3.9 | 0.6 | 1.08 | 4.06 | 61.44 | 85.3 | 86.92 | 87.18 | 87.35
4 | 0.6 | 1 | 3.48 | 51.74 | 84.5 | 86.48 | 86.78 | 86.85
4.1 | 0.6 | 0.98 | 2.98 | 41.5 | 83.7 | 85.92 | 86.32 | 86.4
4.2 | 0.5 | 0.9 | 2.68 | 31.52 | 82.9 | 85.52 | 85.92 | 85.95
4.3 | 0.5 | 0.88 | 2.44 | 22.58 | 81.8 | 85.08 | 85.38 | 85.5
4.4 | 0.5 | 0.8 | 2.14 | 15.66 | 80.6 | 84.52 | 84.92 | 85.05
4.5 | 0.5 | 0.8 | 2.04 | 11.02 | 78.9 | 84.12 | 84.48 | 84.6
4.6 | 0.5 | 0.78 | 1.84 | 8.24 | 76.8 | 83.62 | 84.02 | 84.15
4.7 | 0.5 | 0.7 | 1.74 | 6.58 | 73.4 | 83.18 | 83.58 | 83.7
4.8 | 0.46 | 0.7 | 1.64 | 5.44 | 68.2 | 82.72 | 83.18 | 83.25
4.9 | 0.4 | 0.7 | 1.54 | 4.56 | 60.7 | 82.22 | 82.68 | 82.8
5 | 0.4 | 0.7 | 1.5 | 3.98 | 51.9 | 81.72 | 82.18 | 82.35
5.1 | 0.4 | 0.6 | 1.4 | 3.48 | 42.3 | 81.38 | 81.78 | 81.9
5.2 | 0.4 | 0.6 | 1.34 | 3.14 | 32.8 | 80.82 | 81.32 | 81.5
5.3 | 0.4 | 0.6 | 1.3 | 2.84 | 24.1 | 80.42 | 80.84 | 81.05
5.4 | 0.4 | 0.6 | 1.2 | 2.64 | 16.9 | 79.92 | 80.48 | 80.6
5.5 | 0.4 | 0.6 | 1.2 | 2.44 | 12 | 79.46 | 80.02 | 80.2
5.6 | 0.4 | 0.6 | 1.14 | 2.3 | 8.9 | 79.12 | 79.52 | 79.8
5.7 | 0.4 | 0.5 | 1.1 | 2.14 | 7.1 | 78.68 | 79.08 | 79.35
5.8 | 0.4 | 0.5 | 1.1 | 2.04 | 5.9 | 78.12 | 78.68 | 78.9
5.9 | 0.4 | 0.5 | 1 | 1.94 | 5 | 77.72 | 78.28 | 78.5
6 | 0.36 | 0.5 | 1 | 1.84 | 4.3 | 77.32 | 77.88 | 78.1
6.1 | 0.3 | 0.5 | 1 | 1.74 | 3.9 | 76.82 | 77.48 | 77.7
6.2 | 0.3 | 0.5 | 0.94 | 1.7 | 3.5 | 76.42 | 77.08 | 77.25
6.3 | 0.3 | 0.5 | 0.9 | 1.64 | 3.2 | 75.98 | 76.58 | 76.8
6.4 | 0.3 | 0.5 | 0.9 | 1.54 | 2.9 | 75.52 | 76.18 | 76.4
6.5 |  | 0.5 | 0.9 | 1.5 | 2.8 | 75.12 | 75.78 | 76
6.6 |  | 0.48 | 0.84 | 1.44 | 2.6 | 74.68 | 75.38 | 75.6
6.7 |  | 0.4 | 0.8 | 1.4 | 2.5 | 74.28 | 74.98 | 75.2
6.8 |  | 0.4 | 0.8 | 1.4 | 2.4 | 73.88 | 74.58 | 74.8
6.9 |  | 0.4 | 0.8 | 1.34 | 2.3 | 73.42 | 74.18 | 74.4
7 |  |  | 0.8 | 1.3 | 2.2 | 72.92 | 73.72 | 74
7.1 |  |  | 0.74 | 1.3 | 2.1 | 72.62 | 73.38 | 73.6
7.2 |  |  | 0.7 | 1.24 | 2 | 72.22 | 72.98 | 73.2
7.3 |  |  | 0.7 | 1.2 | 1.9 | 71.82 | 72.58 | 72.85
7.4 |  |  | 0.7 | 1.2 | 1.9 | 71.32 | 72.18 | 72.45
7.5 |  |  | 0.7 | 1.14 | 1.8 | 70.92 | 71.84 | 72.05
7.6 |  |  | 0.7 | 1.1 | 1.7 | 70.58 | 71.44 | 71.7
7.7 |  |  | 0.7 | 1.1 | 1.7 | 70.12 | 71.08 | 71.3
7.8 |  |  | 0.64 | 1.1 | 1.7 | 69.72 | 70.64 | 70.95
7.9 |  |  | 0.6 | 1 | 1.6 | 69.22 | 70.28 | 70.6
8 |  |  | 0.6 | 1 | 1.6 | 68.82 | 69.88 | 70.2
8.1 |  |  |  | 1 | 1.5 | 68.42 | 69.48 | 69.85
8.2 |  |  |  | 1 | 1.5 | 68.02 | 69.12 | 69.5
8.3 |  |  |  | 1 | 1.4 | 67.62 | 68.78 | 69.15
8.4 |  |  |  | 0.9 | 1.4 | 67.32 | 68.48 | 68.8
8.5 |  |  |  | 0.9 | 1.4 | 66.88 | 68.12 | 68.4
8.6 |  |  |  | 0.9 | 1.4 | 66.38 | 67.68 | 68.05
8.7 |  |  |  | 0.9 | 1.3 | 65.96 | 67.38 | 67.75
8.8 |  |  |  | 0.9 | 1.3 | 65.52 | 66.98 | 67.4
8.9 |  |  |  | 0.8 | 1.3 | 65.02 | 66.68 | 67.05
9 |  |  |  | 0.8 | 1.2 | 64.52 | 66.28 | 66.7
9.1 |  |  |  |  | 1.2 | 64.08 | 65.94 | 66.35
9.2 |  |  |  |  | 1.2 | 63.48 | 65.58 | 66.05
9.3 |  |  |  |  | 1.2 | 62.88 | 65.28 | 65.75
9.4 |  |  |  |  | 1.1 | 62.14 | 64.98 | 65.45
9.5 |  |  |  |  | 1.1 | 61.3 | 64.58 | 65.1
9.6 |  |  |  |  | 1.1 | 60.26 | 64.24 | 64.75
9.7 |  |  |  |  | 1.1 | 58.88 | 63.92 | 64.45
9.8 |  |  |  |  | 1.1 | 56.78 | 63.58 | 64.1
9.9 |  |  |  |  | 1 | 53.5 | 63.24 | 63.75
10 |  |  |  |  | 1 | 48.68 | 62.94 | 63.45
10.1 |  |  |  |  | 1 | 42.62 | 62.62 | 63.15
10.2 | [5 empty] | 35.76 | 62.34 | 62.85
10.3 | [5 empty] | 28.7 | 61.98 | 62.55
10.4 | [5 empty] | 22.02 | 61.68 | 62.2
10.5 | [5 empty] | 16.26 | 61.34 | 61.9
10.6 | [5 empty] | 11.88 | 61.08 | 61.6
10.7 | [5 empty] | 9 | 60.74 | 61.3
10.8 | [5 empty] | 7.36 | 60.38 | 61.05
10.9 | [5 empty] | 6.24 | 60.14 | 60.75
11 | [5 empty] | 5.48 | 59.74 | 60.45
11.1 | [5 empty] | 4.92 | 59.54 | 60.15
11.2 | [5 empty] | 4.52 | 59.24 | 59.85
11.3 | [5 empty] | 4.16 | 58.88 | 59.55
11.4 | [5 empty] | 3.96 | 58.54 | 59.25
11.5 | [5 empty] | 3.76 | 58.24 | 58.95
11.6 | [5 empty] | 3.56 | 57.98 | 58.7
11.7 | [5 empty] | 3.4 | 57.68 | 58.45
11.8 | [5 empty] | 3.26 | 57.38 | 58.15
11.9 | [5 empty] | 3.16 | 57.08 | 57.85
12 | [5 empty] | 3.06 | 56.84 | 57.55
12.1 | [5 empty] | 2.96 | 56.54 | 57.3
12.2 | [5 empty] | 2.86 | 56.24 | 57.05
12.3 | [5 empty] | 2.8 | 55.94 | 56.75
12.4 | [5 empty] | 2.7 | 55.64 | 56.5
12.5 | [5 empty] | 2.6 | 55.38 | 56.25
12.6 | [5 empty] | 2.56 | 55.08 | 55.95
12.7 | [5 empty] | 2.5 | 54.84 | 55.7
12.8 | [5 empty] | 2.46 | 54.54 | 55.45
12.9 | [5 empty] | 2.4 | 54.24 | 55.2
13 | [5 empty] | 2.36 | 53.98 | 54.95
13.1 | [5 empty] | 2.3 | 53.68 | 54.65
13.2 | [5 empty] | 2.26 | 53.38 | 54.4
13.3 | [5 empty] | 2.2 | 53.14 | 54.15
13.4 | [5 empty] | 2.16 | 52.84 | 53.9
13.5 | [5 empty] | 2.1 | 52.54 | 53.65
13.6 | [5 empty] | 2.06 | 52.24 | 53.4
13.7 | [5 empty] | 2 | 51.94 | 53.15
13.8 | [5 empty] | 2 | 51.64 | 52.9
13.9 | [5 empty] | 1.96 | 51.34 | 52.65
14 | [5 empty] | 1.9 | 51.08 | 52.4
14.1 | [5 empty] | 1.9 | 50.68 | 52.2
14.2 | [5 empty] | 1.86 | 50.28 | 51.95
14.3 | [5 empty] | 1.8 | 49.94 | 51.7
14.4 | [5 empty] | 1.8 | 49.52 | 51.45
14.5 | [5 empty] | 1.76 | 48.98 | 51.2
14.6 | [5 empty] | 1.7 | 48.36 | 51
14.7 | [5 empty] | 1.7 | 47.66 | 50.75
14.8 | [5 empty] | 1.7 | 46.64 | 50.5
14.9 | [5 empty] | 1.6 | 45.26 | 50.3
15 | [5 empty] | 1.6 | 43.02 | 50.05
15.1 | [5 empty] | 1.6 | 39.76 | 49.8
15.2 | [5 empty] | 1.6 | 35.46 | 49.6
15.3 | [6 empty] | 30.34 | 49.4
15.4 | [6 empty] | 25.04 | 49.15
15.5 | [6 empty] | 19.8 | 48.9
15.6 | [6 empty] | 15.08 | 48.7
15.7 | [6 empty] | 11.34 | 48.5
15.8 | [6 empty] | 8.74 | 48.3
15.9 | [6 empty] | 7.08 | 48.05
16 | [6 empty] | 6.06 | 47.8
16.1 | [6 empty] | 5.42 | 47.6
16.2 | [6 empty] | 4.88 | 47.4
16.3 | [6 empty] | 4.48 | 47.2
16.4 | [6 empty] | 4.24 | 47
16.5 | [6 empty] | 4.04 | 46.8
16.6 | [6 empty] | 3.84 | 46.6
16.7 | [6 empty] | 3.64 | 46.35
16.8 | [6 empty] | 3.54 | 46.1
16.9 | [6 empty] | 3.4 | 45.9
17 | [6 empty] | 3.3 | 45.7
17.1 | [6 empty] | 3.2 | 45.5
17.2 | [6 empty] | 3.14 | 45.3
17.3 | [6 empty] | 3.04 | 45.1
17.4 | [6 empty] | 3 | 44.9
17.5 | [6 empty] | 2.9 | 44.7
17.6 | [6 empty] | 2.84 | 44.5
17.7 | [6 empty] | 2.8 | 44.3
17.8 | [6 empty] | 2.7 | 44.1
17.9 | [6 empty] | 2.64 | 43.9
18 | [6 empty] | 2.6 | 43.65
18.1 | [6 empty] | 2.54 | 43.4
18.2 | [6 empty] | 2.5 | 43.2
18.3 | [6 empty] | 2.44 | 43
18.4 | [6 empty] | 2.4 | 42.8
18.5 | [6 empty] | 2.4 | 42.6
18.6 | [6 empty] | 2.3 | 42.4
18.7 | [6 empty] | 2.3 | 42.2
18.8 | [6 empty] | 2.2 | 42
18.9 | [6 empty] | 2.2 | 41.75
19 | [6 empty] | 2.2 | 41.5
19.1 | [6 empty] | 2.1 | 41.3
19.2 | [6 empty] | 2.1 | 41.05
19.3 | [6 empty] | 2.1 | 40.75
19.4 | [6 empty] | 2 | 40.45
19.5 | [6 empty] | 2 | 40.15
19.6 | [6 empty] | 2 | 39.8
19.7 | [6 empty] | 1.9 | 39.35
19.8 | [6 empty] | 1.9 | 38.75
19.9 | [6 empty] | 1.9 | 37.95
20 | [6 empty] | 1.84 | 36.8
20.1 | [6 empty] | 1.8 | 35.05
20.2 | [6 empty] | 1.8 | 32.45
20.3 | [6 empty] | 1.8 | 29
20.4 | [7 empty] | 25
20.5 | [7 empty] | 20.75
20.6 | [7 empty] | 16.6
20.7 | [7 empty] | 12.9
20.8 | [7 empty] | 9.9
20.9 | [7 empty] | 7.75
21 | [7 empty] | 6.4
21.1 | [7 empty] | 5.5
21.2 | [7 empty] | 4.85
21.3 | [7 empty] | 4.45
21.4 | [7 empty] | 4.15
21.5 | [7 empty] | 3.9
21.6 | [7 empty] | 3.7
21.7 | [7 empty] | 3.55
21.8 | [7 empty] | 3.4
21.9 | [7 empty] | 3.25
22 | [7 empty] | 3.15
22.1 | [7 empty] | 3.05
22.2 | [7 empty] | 3
22.3 | [7 empty] | 2.95
22.4 | [7 empty] | 2.85
22.5 | [7 empty] | 2.75
22.6 | [7 empty] | 2.7
22.7 | [7 empty] | 2.65
22.8 | [7 empty] | 2.6
22.9 | [7 empty] | 2.55
23 | [7 empty] | 2.45
23.1 | [7 empty] | 2.4
23.2 | [7 empty] | 2.4
23.3 | [7 empty] | 2.35
23.4 | [7 empty] | 2.3
23.5 | [7 empty] | 2.25
23.6 | [7 empty] | 2.2
23.7 | [7 empty] | 2.15
23.8 | [7 empty] | 2.1
23.9 | [7 empty] | 2.1
24 | [7 empty] | 2.05
24.1 | [7 empty] | 2
24.2 | [7 empty] | 2
24.3 | [7 empty] | 1.95
24.4 | [7 empty] | 1.9
24.5 | [7 empty] | 1.9
24.6 | [7 empty] | 1.85
24.7 | [7 empty] | 1.8
24.8 | [7 empty] | 1.8
24.9 | [7 empty] | 1.75
25 | [7 empty] | 1.7
25.1 | [7 empty] | 1.7
25.2 | [7 empty] | 1.7
25.3 | [7 empty] | 1.65
25.4 | [7 empty] | 1.6
25.5 | [7 empty] | 1.6
25.6 | [7 empty] | 1.6
25.7 | [7 empty] | 1.6
25.8 | [7 empty] | 1.55
25.9 | [7 empty] | 1.5
26 | [7 empty] | 1.5
26.1 | [7 empty] | 1.5
26.2 | [7 empty] | 1.45
26.3 | [7 empty] | 1.4
26.4 | [7 empty] | 1.4
26.5 | [7 empty] | 1.4
26.6 | [7 empty] | 1.4
26.7 | [7 empty] | 1.35
26.8 | [7 empty] | 1.3
26.9 | [7 empty] | 1.3
27 | [7 empty] | 1.3
27.1 | [7 empty] | 1.3
27.2 | [7 empty] | 1.3
27.3 | [7 empty] | 1.3
27.4 | [7 empty] | 1.25
27.5 | [7 empty] | 1.2
27.6 | [7 empty] | 1.2
27.7 | [7 empty] | 1.2
27.8 | [7 empty] | 1.2
27.9 | [7 empty] | 1.2
28 | [7 empty] | 1.2
28.1 | [7 empty] | 1.15
28.2 | [7 empty] | 1.1
28.3 | [7 empty] | 1.1
28.4 | [7 empty] | 1.1

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 1.3
-28.8 | [7 empty] | 1.3
-28.7 | [7 empty] | 1.3
-28.6 | [7 empty] | 1.3
-28.5 | [7 empty] | 1.3
-28.4 | [7 empty] | 1.37
-28.3 | [7 empty] | 1.4
-28.2 | [7 empty] | 1.4
-28.1 | [7 empty] | 1.4
-28 | [7 empty] | 1.4
-27.9 | [7 empty] | 1.4
-27.8 | [7 empty] | 1.47
-27.7 | [7 empty] | 1.5
-27.6 | [7 empty] | 1.5
-27.5 | [7 empty] | 1.5
-27.4 | [7 empty] | 1.5
-27.3 | [7 empty] | 1.5
-27.2 | [7 empty] | 1.57
-27.1 | [7 empty] | 1.6
-27 | [7 empty] | 1.6
-26.9 | [7 empty] | 1.6
-26.8 | [7 empty] | 1.67
-26.7 | [7 empty] | 1.7
-26.6 | [7 empty] | 1.7
-26.5 | [7 empty] | 1.7
-26.4 | [7 empty] | 1.77
-26.3 | [7 empty] | 1.8
-26.2 | [7 empty] | 1.8
-26.1 | [7 empty] | 1.8
-26 | [7 empty] | 1.8
-25.9 | [7 empty] | 1.87
-25.8 | [7 empty] | 1.9
-25.7 | [7 empty] | 1.9
-25.6 | [7 empty] | 1.97
-25.5 | [7 empty] | 2
-25.4 | [7 empty] | 2
-25.3 | [7 empty] | 2.07
-25.2 | [7 empty] | 2.1
-25.1 | [7 empty] | 2.1
-25 | [7 empty] | 2.17
-24.9 | [7 empty] | 2.2
-24.8 | [7 empty] | 2.2
-24.7 | [7 empty] | 2.27
-24.6 | [7 empty] | 2.3
-24.5 | [7 empty] | 2.3
-24.4 | [7 empty] | 2.37
-24.3 | [7 empty] | 2.4
-24.2 | [7 empty] | 2.47
-24.1 | [7 empty] | 2.5
-24 | [7 empty] | 2.57
-23.9 | [7 empty] | 2.6
-23.8 | [7 empty] | 2.67
-23.7 | [7 empty] | 2.7
-23.6 | [7 empty] | 2.77
-23.5 | [7 empty] | 2.8
-23.4 | [7 empty] | 2.87
-23.3 | [7 empty] | 2.97
-23.2 | [7 empty] | 3
-23.1 | [7 empty] | 3.07
-23 | [7 empty] | 3.17
-22.9 | [7 empty] | 3.27
-22.8 | [7 empty] | 3.37
-22.7 | [7 empty] | 3.47
-22.6 | [7 empty] | 3.57
-22.5 | [7 empty] | 3.67
-22.4 | [7 empty] | 3.77
-22.3 | [7 empty] | 3.94
-22.2 | [7 empty] | 4.07
-22.1 | [7 empty] | 4.24
-22 | [7 empty] | 4.51
-21.9 | [7 empty] | 4.81
-21.8 | [7 empty] | 5.18
-21.7 | [7 empty] | 5.65
-21.6 | [7 empty] | 6.29
-21.5 | [7 empty] | 7.27
-21.4 | [7 empty] | 8.72
-21.3 | [7 empty] | 10.88
-21.2 | [7 empty] | 13.77
-21.1 | [7 empty] | 17.29
-21 | [7 empty] | 21.13
-20.9 | [7 empty] | 25.03
-20.8 | [7 empty] | 28.72
-20.7 | [7 empty] | 31.83
-20.6 | [6 empty] | 2 | 34.17
-20.5 | [6 empty] | 2.1 | 35.78
-20.4 | [6 empty] | 2.1 | 36.9
-20.3 | [6 empty] | 2.16 | 37.69
-20.2 | [6 empty] | 2.2 | 38.25
-20.1 | [6 empty] | 2.2 | 38.75
-20 | [6 empty] | 2.2 | 39.18
-19.9 | [6 empty] | 2.3 | 39.51
-19.8 | [6 empty] | 2.3 | 39.88
-19.7 | [6 empty] | 2.36 | 40.14
-19.6 | [6 empty] | 2.4 | 40.41
-19.5 | [6 empty] | 2.4 | 40.71
-19.4 | [6 empty] | 2.5 | 40.94
-19.3 | [6 empty] | 2.5 | 41.14
-19.2 | [6 empty] | 2.6 | 41.41
-19.1 | [6 empty] | 2.6 | 41.64
-19 | [6 empty] | 2.66 | 41.84
-18.9 | [6 empty] | 2.7 | 42.11
-18.8 | [6 empty] | 2.76 | 42.34
-18.7 | [6 empty] | 2.8 | 42.54
-18.6 | [6 empty] | 2.9 | 42.74
-18.5 | [6 empty] | 2.9 | 42.94
-18.4 | [6 empty] | 3 | 43.14
-18.3 | [6 empty] | 3.06 | 43.41
-18.2 | [6 empty] | 3.1 | 43.64
-18.1 | [6 empty] | 3.2 | 43.84
-18 | [6 empty] | 3.3 | 44.04
-17.9 | [6 empty] | 3.36 | 44.24
-17.8 | [6 empty] | 3.46 | 44.44
-17.7 | [6 empty] | 3.5 | 44.64
-17.6 | [6 empty] | 3.6 | 44.84
-17.5 | [6 empty] | 3.7 | 45.04
-17.4 | [6 empty] | 3.86 | 45.24
-17.3 | [6 empty] | 3.96 | 45.44
-17.2 | [6 empty] | 4.1 | 45.71
-17.1 | [6 empty] | 4.26 | 45.94
-17 | [6 empty] | 4.46 | 46.14
-16.9 | [6 empty] | 4.66 | 46.34
-16.8 | [6 empty] | 4.92 | 46.54
-16.7 | [6 empty] | 5.22 | 46.74
-16.6 | [6 empty] | 5.62 | 47.01
-16.5 | [6 empty] | 6.18 | 47.24
-16.4 | [6 empty] | 6.84 | 47.44
-16.3 | [6 empty] | 7.8 | 47.64
-16.2 | [6 empty] | 9.34 | 47.84
-16.1 | [6 empty] | 11.58 | 48.11
-16 | [6 empty] | 14.84 | 48.34
-15.9 | [6 empty] | 19.02 | 48.54
-15.8 | [6 empty] | 23.7 | 48.74
-15.7 | [6 empty] | 28.6 | 48.94
-15.6 | [6 empty] | 33.38 | 49.14
-15.5 | [6 empty] | 37.6 | 49.41
-15.4 | [5 empty] | 1.8 | 40.94 | 49.64
-15.3 | [5 empty] | 1.9 | 43.34 | 49.84
-15.2 | [5 empty] | 1.9 | 45.02 | 50.04
-15.1 | [5 empty] | 1.9 | 46.2 | 50.31
-15 | [5 empty] | 2 | 47.04 | 50.54
-14.9 | [5 empty] | 2 | 47.68 | 50.74
-14.8 | [5 empty] | 2 | 48.32 | 51.01
-14.7 | [5 empty] | 2.1 | 48.82 | 51.24
-14.6 | [5 empty] | 2.1 | 49.28 | 51.51
-14.5 | [5 empty] | 2.1 | 49.72 | 51.74
-14.4 | [5 empty] | 2.2 | 50.06 | 51.94
-14.3 | [5 empty] | 2.2 | 50.46 | 52.14
-14.2 | [5 empty] | 2.3 | 50.76 | 52.41
-14.1 | [5 empty] | 2.3 | 51.12 | 52.64
-14 | [5 empty] | 2.36 | 51.46 | 52.91
-13.9 | [5 empty] | 2.4 | 51.76 | 53.14
-13.8 | [5 empty] | 2.46 | 52.06 | 53.41
-13.7 | [5 empty] | 2.5 | 52.36 | 53.64
-13.6 | [5 empty] | 2.6 | 52.66 | 53.91
-13.5 | [5 empty] | 2.6 | 52.96 | 54.14
-13.4 | [5 empty] | 2.7 | 53.26 | 54.41
-13.3 | [5 empty] | 2.76 | 53.56 | 54.71
-13.2 | [5 empty] | 2.8 | 53.82 | 54.94
-13.1 | [5 empty] | 2.9 | 54.12 | 55.14
-13 | [5 empty] | 2.96 | 54.42 | 55.41
-12.9 | [5 empty] | 3 | 54.66 | 55.71
-12.8 | [5 empty] | 3.1 | 54.96 | 55.94
-12.7 | [5 empty] | 3.2 | 55.26 | 56.21
-12.6 | [5 empty] | 3.3 | 55.56 | 56.51
-12.5 | [5 empty] | 3.4 | 55.86 | 56.74
-12.4 | [5 empty] | 3.5 | 56.12 | 57.01
-12.3 | [5 empty] | 3.6 | 56.46 | 57.24
-12.2 | [5 empty] | 3.76 | 56.72 | 57.51
-12.1 | [5 empty] | 3.86 | 57.02 | 57.81
-12 | [5 empty] | 4.06 | 57.32 | 58.11
-11.9 | [5 empty] | 4.26 | 57.66 | 58.41
-11.8 | [5 empty] | 4.46 | 57.86 | 58.71
-11.7 | [5 empty] | 4.66 | 58.16 | 58.94
-11.6 | [5 empty] | 4.96 | 58.52 | 59.21
-11.5 | [5 empty] | 5.32 | 58.76 | 59.51
-11.4 | [5 empty] | 5.78 | 59.12 | 59.81
-11.3 | [5 empty] | 6.32 | 59.36 | 60.11
-11.2 | [5 empty] | 7.04 | 59.72 | 60.41
-11.1 | [5 empty] | 8.1 | 60.02 | 60.64
-11 | [5 empty] | 9.54 | 60.32 | 60.98
-10.9 | [5 empty] | 11.84 | 60.66 | 61.24
-10.8 | [5 empty] | 15.46 | 60.96 | 61.51
-10.7 | [5 empty] | 20.48 | 61.22 | 61.88
-10.6 | [5 empty] | 26.46 | 61.52 | 62.21
-10.5 | [5 empty] | 32.98 | 61.96 | 62.51
-10.4 | [5 empty] | 39.52 | 62.16 | 62.74
-10.3 | [5 empty] | 45.48 | 62.52 | 63.08
-10.2 |  |  |  |  | 1.2 | 50.62 | 62.86 | 63.41
-10.1 |  |  |  |  | 1.2 | 54.36 | 63.22 | 63.71
-10 |  |  |  |  | 1.3 | 56.9 | 63.52 | 64.01
-9.9 |  |  |  |  | 1.3 | 58.52 | 63.86 | 64.31
-9.8 |  |  |  |  | 1.3 | 59.7 | 64.22 | 64.68
-9.7 |  |  |  |  | 1.3 | 60.7 | 64.56 | 65.01
-9.6 |  |  |  |  | 1.4 | 61.54 | 64.82 | 65.31
-9.5 |  |  |  |  | 1.4 | 62.28 | 65.18 | 65.61
-9.4 |  |  |  |  | 1.4 | 62.88 | 65.52 | 65.98
-9.3 |  |  |  |  | 1.5 | 63.48 | 65.82 | 66.31
-9.2 |  |  |  |  | 1.5 | 64.08 | 66.16 | 66.61
-9.1 |  |  |  | 1 | 1.5 | 64.62 | 66.52 | 66.98
-9 |  |  |  | 1 | 1.54 | 65.02 | 66.86 | 67.31
-8.9 |  |  |  | 1.1 | 1.6 | 65.58 | 67.22 | 67.61
-8.8 |  |  |  | 1.1 | 1.6 | 66.02 | 67.56 | 67.98
-8.7 |  |  |  | 1.1 | 1.7 | 66.42 | 67.92 | 68.31
-8.6 |  |  |  | 1.1 | 1.7 | 66.86 | 68.22 | 68.68
-8.5 |  |  |  | 1.2 | 1.8 | 67.32 | 68.62 | 69.01
-8.4 |  |  |  | 1.2 | 1.8 | 67.72 | 69.02 | 69.38
-8.3 |  |  |  | 1.2 | 1.84 | 68.12 | 69.36 | 69.78
-8.2 |  |  |  | 1.2 | 1.9 | 68.58 | 69.78 | 70.11
-8.1 |  |  | 0.8 | 1.24 | 1.94 | 68.98 | 70.12 | 70.48
-8 |  |  | 0.8 | 1.3 | 2 | 69.42 | 70.48 | 70.81
-7.9 |  |  | 0.8 | 1.3 | 2.1 | 69.76 | 70.92 | 71.18
-7.8 |  |  | 0.8 | 1.4 | 2.2 | 70.28 | 71.22 | 71.58
-7.7 |  |  | 0.8 | 1.4 | 2.24 | 70.72 | 71.62 | 71.98
-7.6 |  |  | 0.86 | 1.4 | 2.3 | 71.12 | 71.96 | 72.38
-7.5 |  |  | 0.9 | 1.5 | 2.4 | 71.48 | 72.42 | 72.71
-7.4 |  |  | 0.9 | 1.5 | 2.5 | 71.92 | 72.72 | 73.08
-7.3 |  |  | 0.9 | 1.54 | 2.6 | 72.32 | 73.12 | 73.41
-7.2 |  |  | 0.9 | 1.6 | 2.7 | 72.72 | 73.52 | 73.78
-7.1 |  |  | 1 | 1.64 | 2.84 | 73.12 | 74.02 | 74.18
-7 |  | 0.5 | 1 | 1.7 | 2.94 | 73.52 | 74.32 | 74.58
-6.9 |  | 0.56 | 1 | 1.8 | 3.14 | 74.02 | 74.72 | 74.98
-6.8 |  | 0.6 | 1 | 1.8 | 3.24 | 74.42 | 75.12 | 75.38
-6.7 |  | 0.6 | 1.1 | 1.9 | 3.44 | 74.82 | 75.52 | 75.78
-6.6 |  | 0.6 | 1.1 | 2 | 3.68 | 75.26 | 75.92 | 76.18
-6.5 | 0.4 | 0.6 | 1.1 | 2.04 | 3.98 | 75.72 | 76.42 | 76.58
-6.4 | 0.4 | 0.6 | 1.2 | 2.1 | 4.28 | 76.12 | 76.72 | 76.98
-6.3 | 0.4 | 0.6 | 1.2 | 2.2 | 4.68 | 76.58 | 77.12 | 77.38
-6.2 | 0.46 | 0.6 | 1.2 | 2.3 | 5.18 | 77.02 | 77.58 | 77.85
-6.1 | 0.5 | 0.7 | 1.3 | 2.44 | 5.76 | 77.42 | 78.02 | 78.28
-6 | 0.5 | 0.7 | 1.3 | 2.54 | 6.6 | 77.88 | 78.42 | 78.68
-5.9 | 0.5 | 0.7 | 1.4 | 2.7 | 7.64 | 78.32 | 78.82 | 79.08
-5.8 | 0.5 | 0.7 | 1.4 | 2.84 | 9.16 | 78.82 | 79.26 | 79.48
-5.7 | 0.5 | 0.7 | 1.5 | 3.04 | 11.36 | 79.22 | 79.68 | 79.95
-5.6 | 0.5 | 0.76 | 1.56 | 3.24 | 14.92 | 79.68 | 80.12 | 80.38
-5.5 | 0.5 | 0.8 | 1.6 | 3.54 | 20.42 | 80.08 | 80.52 | 80.78
-5.4 | 0.5 | 0.8 | 1.7 | 3.84 | 27.7 | 80.58 | 80.92 | 81.18
-5.3 | 0.5 | 0.8 | 1.76 | 4.18 | 36.2 | 81.02 | 81.38 | 81.65
-5.2 | 0.6 | 0.86 | 1.86 | 4.68 | 45.26 | 81.42 | 81.82 | 82.08
-5.1 | 0.6 | 0.9 | 1.96 | 5.22 | 54.18 | 81.92 | 82.32 | 82.48
-5 | 0.6 | 0.9 | 2.06 | 5.96 | 62.18 | 82.32 | 82.76 | 82.95
-4.9 | 0.6 | 0.96 | 2.2 | 6.9 | 68.58 | 82.88 | 83.12 | 83.38
-4.8 | 0.6 | 1 | 2.36 | 8.22 | 73.18 | 83.28 | 83.58 | 83.78
-4.7 | 0.66 | 1 | 2.5 | 10.14 | 76.14 | 83.72 | 84.12 | 84.25
-4.6 | 0.7 | 1.1 | 2.66 | 13.06 | 78.22 | 84.18 | 84.52 | 84.75
-4.5 | 0.7 | 1.1 | 2.96 | 17.76 | 79.78 | 84.68 | 84.92 | 85.18
-4.4 | 0.7 | 1.16 | 3.22 | 24.56 | 81.04 | 85.12 | 85.48 | 85.65
-4.3 | 0.76 | 1.26 | 3.52 | 33.1 | 82.1 | 85.58 | 85.92 | 86.08
-4.2 | 0.8 | 1.3 | 3.98 | 42.56 | 83 | 86.08 | 86.36 | 86.48
-4.1 | 0.8 | 1.36 | 4.48 | 52.32 | 83.96 | 86.52 | 86.82 | 86.95
-4 | 0.86 | 1.46 | 5.14 | 61.42 | 84.72 | 87.02 | 87.32 | 87.38
-3.9 | 0.9 | 1.56 | 6 | 69.36 | 85.42 | 87.48 | 87.72 | 87.85
-3.8 | 0.96 | 1.7 | 7.16 | 75.32 | 86.12 | 87.88 | 88.18 | 88.28
-3.7 | 1 | 1.8 | 8.8 | 79.3 | 86.82 | 88.36 | 88.56 | 88.75
-3.6 | 1.06 | 1.96 | 11.08 | 81.94 | 87.42 | 88.82 | 89.02 | 89.18
-3.5 | 1.1 | 2.16 | 14.78 | 83.82 | 87.92 | 89.32 | 89.52 | 89.65
-3.4 | 1.2 | 2.42 | 20.66 | 85.28 | 88.58 | 89.76 | 89.88 | 90.08
-3.3 | 1.3 | 2.72 | 28.74 | 86.54 | 89.12 | 90.18 | 90.36 | 90.48
-3.2 | 1.46 | 3.02 | 38.26 | 87.64 | 89.72 | 90.72 | 90.78 | 90.95
-3.1 | 1.56 | 3.52 | 48.62 | 88.56 | 90.18 | 91.12 | 91.26 | 91.38
-3 | 1.76 | 4.08 | 58.96 | 89.36 | 90.72 | 91.58 | 91.72 | 91.85
-2.9 | 1.96 | 4.84 | 68.34 | 90.12 | 91.32 | 92.02 | 92.12 | 92.28
-2.8 | 2.22 | 5.86 | 76.04 | 90.8 | 91.78 | 92.48 | 92.62 | 92.75
-2.7 | 2.52 | 7.28 | 81.44 | 91.52 | 92.22 | 92.92 | 93.02 | 93.11
-2.6 | 2.92 | 9.32 | 85.2 | 92.12 | 92.78 | 93.32 | 93.42 | 93.55
-2.5 | 3.48 | 12.4 | 87.6 | 92.62 | 93.18 | 93.72 | 93.82 | 93.98
-2.4 | 4.2 | 17.56 | 89.38 | 93.22 | 93.66 | 94.22 | 94.22 | 94.38
-2.3 | 5.16 | 24.92 | 90.76 | 93.82 | 94.12 | 94.62 | 94.62 | 94.78
-2.2 | 6.48 | 34.42 | 91.96 | 94.28 | 94.58 | 94.98 | 95.02 | 95.18
-2.1 | 8.36 | 45.16 | 92.88 | 94.72 | 94.98 | 95.42 | 95.42 | 95.58
-2 | 11.24 | 56.3 | 93.64 | 95.18 | 95.38 | 95.76 | 95.82 | 95.98
-1.9 | 16.04 | 66.9 | 94.44 | 95.68 | 95.78 | 96.12 | 96.12 | 96.31
-1.8 | 23.32 | 75.92 | 95.18 | 96.02 | 96.18 | 96.52 | 96.52 | 96.68
-1.7 | 32.76 | 83 | 95.62 | 96.44 | 96.58 | 96.92 | 96.92 | 97.01
-1.6 | 43.78 | 87.74 | 96.22 | 96.82 | 96.98 | 97.16 | 97.22 | 97.31
-1.5 | 55.16 | 90.88 | 96.62 | 97.28 | 97.34 | 97.52 | 97.52 | 97.61
-1.4 | 66.28 | 92.94 | 97.02 | 97.58 | 97.64 | 97.76 | 97.86 | 97.91
-1.3 | 76 | 94.46 | 97.48 | 97.94 | 97.94 | 98.12 | 98.06 | 98.21
-1.2 | 83.84 | 95.6 | 97.92 | 98.14 | 98.22 | 98.36 | 98.42 | 98.51
-1.1 | 89.32 | 96.6 | 98.26 | 98.44 | 98.54 | 98.62 | 98.66 | 98.74
-1 | 92.74 | 97.28 | 98.56 | 98.64 | 98.74 | 98.86 | 98.86 | 98.94
-0.9 | 94.84 | 97.88 | 98.82 | 98.94 | 98.94 | 99.1 | 99 | 99.14
-0.8 | 96.36 | 98.38 | 99.12 | 99.14 | 99.14 | 99.32 | 99.2 | 99.34
-0.7 | 97.4 | 98.82 | 99.36 | 99.34 | 99.28 | 99.46 | 99.42 | 99.54
-0.6 | 98.34 | 99.12 | 99.56 | 99.48 | 99.54 | 99.56 | 99.56 | 99.67
-0.5 | 98.92 | 99.46 | 99.7 | 99.7 | 99.7 | 99.7 | 99.7 | 99.77
-0.4 | 99.42 | 99.66 | 99.86 | 99.8 | 99.8 | 99.8 | 99.8 | 99.87
-0.3 | 99.76 | 99.76 | 99.9 | 99.84 | 99.9 | 99.9 | 99.9 | 99.9
-0.2 | 99.9 | 99.9 | 100 | 99.9 | 99.94 | 99.9 | 99.9 | 99.97
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 99.96 | 100 | 100 | 99.93
0.3 | 99.74 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9
0.4 | 99.38 | 99.74 | 99.84 | 99.8 | 99.76 | 99.8 | 99.8 | 99.83
0.5 | 98.88 | 99.54 | 99.74 | 99.7 | 99.66 | 99.7 | 99.7 | 99.8
0.6 | 98.26 | 99.24 | 99.54 | 99.56 | 99.5 | 99.54 | 99.64 | 99.66
0.7 | 97.4 | 98.98 | 99.34 | 99.46 | 99.36 | 99.44 | 99.38 | 99.53
0.8 | 96.24 | 98.48 | 99.08 | 99.16 | 99.2 | 99.18 | 99.24 | 99.36
0.9 | 94.76 | 98.08 | 98.84 | 98.96 | 99.06 | 99.04 | 99.04 | 99.16
1 | 92.52 | 97.52 | 98.54 | 98.7 | 98.76 | 98.84 | 98.84 | 98.96
1.1 | 88.94 | 96.76 | 98.34 | 98.52 | 98.46 | 98.54 | 98.58 | 98.76
1.2 | 83.26 | 95.9 | 97.88 | 98.18 | 98.26 | 98.38 | 98.34 | 98.49
1.3 | 75.2 | 94.8 | 97.58 | 97.86 | 97.96 | 98.08 | 98.08 | 98.19
1.4 | 65.38 | 93.32 | 97.18 | 97.56 | 97.76 | 97.84 | 97.84 | 97.89
1.5 | 54.18 | 91.28 | 96.72 | 97.16 | 97.36 | 97.48 | 97.54 | 97.59
1.6 | 42.82 | 88.26 | 96.22 | 96.82 | 97.02 | 97.08 | 97.24 | 97.36
1.7 | 31.9 | 83.62 | 95.62 | 96.46 | 96.62 | 96.84 | 96.88 | 97.02
1.8 | 22.62 | 76.7 | 95.02 | 95.98 | 96.22 | 96.48 | 96.48 | 96.69
1.9 | 15.56 | 67.72 | 94.42 | 95.56 | 95.88 | 96.14 | 96.18 | 96.32
2 | 10.92 | 57.16 | 93.7 | 95.08 | 95.42 | 95.78 | 95.78 | 95.92
2.1 | 8.14 | 45.94 | 92.86 | 94.62 | 95.02 | 95.38 | 95.44 | 95.52
2.2 | 6.32 | 35.12 | 92 | 94.12 | 94.72 | 94.98 | 95.08 | 95.12
2.3 | 5.04 | 25.5 | 90.84 | 93.62 | 94.22 | 94.58 | 94.64 | 94.72
2.4 | 4.16 | 17.92 | 89.42 | 93.08 | 93.72 | 94.18 | 94.28 | 94.39
2.5 | 3.42 | 12.64 | 87.5 | 92.58 | 93.22 | 93.72 | 93.88 | 94.02
2.6 | 2.92 | 9.42 | 85.28 | 91.98 | 92.72 | 93.28 | 93.38 | 93.55
2.7 | 2.48 | 7.42 | 82.02 | 91.38 | 92.28 | 92.88 | 92.98 | 93.12
2.8 | 2.18 | 5.94 | 76.62 | 90.68 | 91.82 | 92.48 | 92.58 | 92.72
2.9 | 1.94 | 4.9 | 68.92 | 89.98 | 91.32 | 91.98 | 92.12 | 92.25
3 | 1.74 | 4.12 | 59.76 | 89.24 | 90.78 | 91.58 | 91.68 | 91.82
3.1 | 1.54 | 3.52 | 49.56 | 88.44 | 90.32 | 91.12 | 91.24 | 91.42
3.2 | 1.44 | 3.08 | 39.1 | 87.54 | 89.74 | 90.68 | 90.82 | 90.95
3.3 | 1.3 | 2.74 | 29.44 | 86.4 | 89.18 | 90.22 | 90.32 | 90.52
3.4 | 1.2 | 2.44 | 21.18 | 85.22 | 88.68 | 89.72 | 89.98 | 90.05
3.5 | 1.1 | 2.24 | 15.1 | 83.64 | 88.08 | 89.22 | 89.48 | 89.62
3.6 | 1.04 | 2 | 11.24 | 81.7 | 87.42 | 88.72 | 89.08 | 89.22
3.7 | 1 | 1.84 | 8.86 | 79.22 | 86.78 | 88.38 | 88.58 | 88.75
3.8 | 0.94 | 1.7 | 7.18 | 75.16 | 86.14 | 87.92 | 88.08 | 88.32
3.9 | 0.9 | 1.54 | 6 | 69.34 | 85.44 | 87.48 | 87.62 | 87.85
4 | 0.84 | 1.44 | 5.12 | 61.34 | 84.74 | 86.92 | 87.18 | 87.42
4.1 | 0.8 | 1.4 | 4.42 | 52.24 | 83.94 | 86.48 | 86.72 | 86.95
4.2 | 0.8 | 1.3 | 3.92 | 42.54 | 83.1 | 86.08 | 86.28 | 86.52
4.3 | 0.74 | 1.24 | 3.48 | 33.06 | 82.1 | 85.48 | 85.88 | 86.12
4.4 | 0.7 | 1.2 | 3.18 | 24.48 | 81.06 | 85.08 | 85.38 | 85.65
4.5 | 0.7 | 1.1 | 2.94 | 17.74 | 79.82 | 84.62 | 84.92 | 85.15
4.6 | 0.7 | 1.1 | 2.64 | 13.08 | 78.24 | 84.18 | 84.58 | 84.72
4.7 | 0.64 | 1 | 2.5 | 10.16 | 76.22 | 83.78 | 84.12 | 84.25
4.8 | 0.6 | 1 | 2.34 | 8.28 | 73.42 | 83.24 | 83.62 | 83.82
4.9 | 0.6 | 0.94 | 2.2 | 6.96 | 69.06 | 82.78 | 83.18 | 83.42
5 | 0.6 | 0.9 | 2.04 | 5.94 | 62.82 | 82.32 | 82.78 | 82.95
5.1 | 0.6 | 0.9 | 1.94 | 5.18 | 54.82 | 81.88 | 82.28 | 82.52
5.2 | 0.6 | 0.9 | 1.84 | 4.68 | 46.06 | 81.42 | 81.92 | 82.12
5.3 | 0.5 | 0.8 | 1.74 | 4.22 | 37.04 | 80.98 | 81.48 | 81.65
5.4 | 0.5 | 0.8 | 1.7 | 3.82 | 28.5 | 80.58 | 80.98 | 81.22
5.5 | 0.5 | 0.8 | 1.6 | 3.52 | 21.04 | 80.08 | 80.52 | 80.82
5.6 | 0.5 | 0.8 | 1.54 | 3.26 | 15.32 | 79.58 | 80.08 | 80.35
5.7 | 0.5 | 0.7 | 1.5 | 3.06 | 11.64 | 79.18 | 79.72 | 79.92
5.8 | 0.5 | 0.7 | 1.4 | 2.86 | 9.34 | 78.78 | 79.28 | 79.52
5.9 | 0.5 | 0.7 | 1.4 | 2.66 | 7.76 | 78.28 | 78.88 | 79.05
6 | 0.5 | 0.7 | 1.3 | 2.56 | 6.7 | 77.82 | 78.48 | 78.62
6.1 | 0.5 | 0.7 | 1.3 | 2.46 | 5.88 | 77.38 | 78.02 | 78.22
6.2 | 0.44 | 0.6 | 1.24 | 2.3 | 5.28 | 76.98 | 77.68 | 77.82
6.3 | 0.4 | 0.6 | 1.2 | 2.2 | 4.72 | 76.58 | 77.18 | 77.42
6.4 | 0.4 | 0.6 | 1.2 | 2.1 | 4.32 | 76.12 | 76.78 | 77.02
6.5 | 0.4 | 0.6 | 1.1 | 2.06 | 4.02 | 75.68 | 76.38 | 76.62
6.6 |  | 0.6 | 1.1 | 1.96 | 3.76 | 75.28 | 75.88 | 76.22
6.7 |  | 0.6 | 1.1 | 1.9 | 3.46 | 74.78 | 75.54 | 75.82
6.8 |  | 0.6 | 1 | 1.8 | 3.26 | 74.38 | 75.18 | 75.42
6.9 |  | 0.6 | 1 | 1.76 | 3.1 | 73.98 | 74.78 | 74.95
7 |  | 0.54 | 1 | 1.7 | 2.96 | 73.58 | 74.32 | 74.59
7.1 |  |  | 1 | 1.66 | 2.8 | 73.18 | 73.98 | 74.22
7.2 |  |  | 0.9 | 1.6 | 2.7 | 72.72 | 73.48 | 73.82
7.3 |  |  | 0.9 | 1.56 | 2.56 | 72.28 | 73.14 | 73.42
7.4 |  |  | 0.9 | 1.5 | 2.46 | 71.88 | 72.78 | 73.09
7.5 |  |  | 0.9 | 1.5 | 2.4 | 71.42 | 72.44 | 72.72
7.6 |  |  | 0.9 | 1.4 | 2.3 | 71.08 | 72.02 | 72.39
7.7 |  |  | 0.8 | 1.4 | 2.2 | 70.68 | 71.58 | 71.95
7.8 |  |  | 0.8 | 1.36 | 2.16 | 70.24 | 71.28 | 71.52
7.9 |  |  | 0.8 | 1.3 | 2.1 | 69.78 | 70.88 | 71.19
8 |  |  | 0.8 | 1.3 | 2 | 69.38 | 70.48 | 70.82
8.1 |  |  | 0.8 | 1.26 | 1.96 | 68.98 | 70.08 | 70.42
8.2 |  |  |  | 1.2 | 1.9 | 68.58 | 69.78 | 70.09
8.3 |  |  |  | 1.2 | 1.86 | 68.08 | 69.44 | 69.79
8.4 |  |  |  | 1.2 | 1.8 | 67.68 | 69.02 | 69.42
8.5 |  |  |  | 1.16 | 1.76 | 67.22 | 68.68 | 69.02
8.6 |  |  |  | 1.1 | 1.7 | 66.82 | 68.34 | 68.69
8.7 |  |  |  | 1.1 | 1.7 | 66.38 | 67.94 | 68.32
8.8 |  |  |  | 1.1 | 1.6 | 65.88 | 67.58 | 67.92
8.9 |  |  |  | 1.06 | 1.6 | 65.48 | 67.24 | 67.59
9 |  |  |  | 1 | 1.56 | 64.98 | 66.84 | 67.29
9.1 |  |  |  | 1 | 1.5 | 64.48 | 66.54 | 66.92
9.2 |  |  |  |  | 1.5 | 63.92 | 66.18 | 66.59
9.3 |  |  |  |  | 1.4 | 63.42 | 65.84 | 66.29
9.4 |  |  |  |  | 1.4 | 62.82 | 65.54 | 65.99
9.5 |  |  |  |  | 1.4 | 62.16 | 65.14 | 65.62
9.6 |  |  |  |  | 1.4 | 61.46 | 64.78 | 65.29
9.7 |  |  |  |  | 1.3 | 60.6 | 64.44 | 64.99
9.8 |  |  |  |  | 1.3 | 59.6 | 64.18 | 64.62
9.9 |  |  |  |  | 1.3 | 58.38 | 63.88 | 64.36
10 |  |  |  |  | 1.3 | 56.7 | 63.52 | 64.02
10.1 |  |  |  |  | 1.2 | 54.16 | 63.18 | 63.69
10.2 | [5 empty] | 50.38 | 62.84 | 63.39
10.3 | [5 empty] | 45.22 | 62.54 | 63.09
10.4 | [5 empty] | 39.18 | 62.28 | 62.79
10.5 | [5 empty] | 32.72 | 61.88 | 62.49
10.6 | [5 empty] | 26.18 | 61.58 | 62.19
10.7 | [5 empty] | 20.16 | 61.24 | 61.89
10.8 | [5 empty] | 15.28 | 60.98 | 61.52
10.9 | [5 empty] | 11.7 | 60.64 | 61.26
11 | [5 empty] | 9.46 | 60.38 | 60.99
11.1 | [5 empty] | 8.04 | 60.04 | 60.69
11.2 | [5 empty] | 7 | 59.68 | 60.39
11.3 | [5 empty] | 6.32 | 59.34 | 60.09
11.4 | [5 empty] | 5.72 | 59.14 | 59.79
11.5 | [5 empty] | 5.28 | 58.84 | 59.49
11.6 | [5 empty] | 4.94 | 58.48 | 59.19
11.7 | [5 empty] | 4.64 | 58.18 | 58.96
11.8 | [5 empty] | 4.44 | 57.94 | 58.69
11.9 | [5 empty] | 4.24 | 57.58 | 58.39
12 | [5 empty] | 4.04 | 57.24 | 58.09
12.1 | [5 empty] | 3.84 | 56.98 | 57.79
12.2 | [5 empty] | 3.74 | 56.74 | 57.49
12.3 | [5 empty] | 3.6 | 56.38 | 57.26
12.4 | [5 empty] | 3.5 | 56.14 | 56.99
12.5 | [5 empty] | 3.34 | 55.84 | 56.76
12.6 | [5 empty] | 3.3 | 55.48 | 56.49
12.7 | [5 empty] | 3.2 | 55.24 | 56.26
12.8 | [5 empty] | 3.1 | 54.94 | 55.99
12.9 | [5 empty] | 3 | 54.68 | 55.69
13 | [5 empty] | 2.9 | 54.38 | 55.39
13.1 | [5 empty] | 2.84 | 54.08 | 55.16
13.2 | [5 empty] | 2.8 | 53.84 | 54.96
13.3 | [5 empty] | 2.7 | 53.54 | 54.69
13.4 | [5 empty] | 2.64 | 53.24 | 54.39
13.5 | [5 empty] | 2.6 | 52.98 | 54.16
13.6 | [5 empty] | 2.54 | 52.68 | 53.89
13.7 | [5 empty] | 2.5 | 52.34 | 53.66
13.8 | [5 empty] | 2.44 | 52.08 | 53.39
13.9 | [5 empty] | 2.4 | 51.78 | 53.16
14 | [5 empty] | 2.34 | 51.44 | 52.89
14.1 | [5 empty] | 2.3 | 51.14 | 52.66
14.2 | [5 empty] | 2.2 | 50.78 | 52.39
14.3 | [5 empty] | 2.2 | 50.48 | 52.16
14.4 | [5 empty] | 2.2 | 50.08 | 51.96
14.5 | [5 empty] | 2.1 | 49.74 | 51.76
14.6 | [5 empty] | 2.1 | 49.28 | 51.49
14.7 | [5 empty] | 2.04 | 48.82 | 51.26
14.8 | [5 empty] | 2 | 48.32 | 50.99
14.9 | [5 empty] | 2 | 47.82 | 50.76
15 | [5 empty] | 1.94 | 47.12 | 50.49
15.1 | [5 empty] | 1.9 | 46.3 | 50.33
15.2 | [5 empty] | 1.9 | 45.24 | 50.09
15.3 | [5 empty] | 1.9 | 43.66 | 49.86
15.4 | [5 empty] | 1.8 | 41.42 | 49.59
15.5 | [6 empty] | 38.16 | 49.43
15.6 | [6 empty] | 33.98 | 49.19
15.7 | [6 empty] | 29.2 | 48.96
15.8 | [6 empty] | 24.26 | 48.76
15.9 | [6 empty] | 19.42 | 48.56
16 | [6 empty] | 15.2 | 48.29
16.1 | [6 empty] | 11.8 | 48.06
16.2 | [6 empty] | 9.4 | 47.86
16.3 | [6 empty] | 7.94 | 47.66
16.4 | [6 empty] | 6.86 | 47.46
16.5 | [6 empty] | 6.22 | 47.19
16.6 | [6 empty] | 5.68 | 46.96
16.7 | [6 empty] | 5.28 | 46.76
16.8 | [6 empty] | 4.94 | 46.56
16.9 | [6 empty] | 4.64 | 46.36
17 | [6 empty] | 4.44 | 46.09
17.1 | [6 empty] | 4.24 | 45.86
17.2 | [6 empty] | 4.14 | 45.66
17.3 | [6 empty] | 3.94 | 45.46
17.4 | [6 empty] | 3.84 | 45.26
17.5 | [6 empty] | 3.74 | 45.06
17.6 | [6 empty] | 3.6 | 44.86
17.7 | [6 empty] | 3.5 | 44.66
17.8 | [6 empty] | 3.44 | 44.46
17.9 | [6 empty] | 3.34 | 44.19
18 | [6 empty] | 3.24 | 44.03
18.1 | [6 empty] | 3.2 | 43.79
18.2 | [6 empty] | 3.1 | 43.63
18.3 | [6 empty] | 3.04 | 43.39
18.4 | [6 empty] | 3 | 43.16
18.5 | [6 empty] | 2.9 | 42.96
18.6 | [6 empty] | 2.84 | 42.76
18.7 | [6 empty] | 2.8 | 42.56
18.8 | [6 empty] | 2.74 | 42.36
18.9 | [6 empty] | 2.7 | 42.09
19 | [6 empty] | 2.6 | 41.86
19.1 | [6 empty] | 2.6 | 41.66
19.2 | [6 empty] | 2.54 | 41.39
19.3 | [6 empty] | 2.5 | 41.16
19.4 | [6 empty] | 2.44 | 40.96
19.5 | [6 empty] | 2.4 | 40.69
19.6 | [6 empty] | 2.4 | 40.46
19.7 | [6 empty] | 2.3 | 40.19
19.8 | [6 empty] | 2.3 | 39.89
19.9 | [6 empty] | 2.3 | 39.52
20 | [6 empty] | 2.2 | 39.19
20.1 | [6 empty] | 2.2 | 38.75
20.2 | [6 empty] | 2.14 | 38.25
20.3 | [6 empty] | 2.1 | 37.68
20.4 | [6 empty] | 2.1 | 36.94
20.5 | [6 empty] | 2.1 | 35.79
20.6 | [6 empty] | 2 | 34.14
20.7 | [7 empty] | 31.78
20.8 | [7 empty] | 28.69
20.9 | [7 empty] | 25.04
21 | [7 empty] | 21.1
21.1 | [7 empty] | 17.24
21.2 | [7 empty] | 13.79
21.3 | [7 empty] | 10.84
21.4 | [7 empty] | 8.67
21.5 | [7 empty] | 7.26
21.6 | [7 empty] | 6.34
21.7 | [7 empty] | 5.68
21.8 | [7 empty] | 5.15
21.9 | [7 empty] | 4.79
22 | [7 empty] | 4.49
22.1 | [7 empty] | 4.26
22.2 | [7 empty] | 4.06
22.3 | [7 empty] | 3.93
22.4 | [7 empty] | 3.76
22.5 | [7 empty] | 3.63
22.6 | [7 empty] | 3.53
22.7 | [7 empty] | 3.43
22.8 | [7 empty] | 3.33
22.9 | [7 empty] | 3.23
23 | [7 empty] | 3.13
23.1 | [7 empty] | 3.1
23.2 | [7 empty] | 3.03
23.3 | [7 empty] | 2.93
23.4 | [7 empty] | 2.9
23.5 | [7 empty] | 2.83
23.6 | [7 empty] | 2.73
23.7 | [7 empty] | 2.7
23.8 | [7 empty] | 2.63
23.9 | [7 empty] | 2.6
24 | [7 empty] | 2.53
24.1 | [7 empty] | 2.5
24.2 | [7 empty] | 2.43
24.3 | [7 empty] | 2.4
24.4 | [7 empty] | 2.4
24.5 | [7 empty] | 2.33
24.6 | [7 empty] | 2.3
24.7 | [7 empty] | 2.23
24.8 | [7 empty] | 2.2
24.9 | [7 empty] | 2.2
25 | [7 empty] | 2.13
25.1 | [7 empty] | 2.1
25.2 | [7 empty] | 2.1
25.3 | [7 empty] | 2.03
25.4 | [7 empty] | 2
25.5 | [7 empty] | 2
25.6 | [7 empty] | 1.93
25.7 | [7 empty] | 1.9
25.8 | [7 empty] | 1.9
25.9 | [7 empty] | 1.9
26 | [7 empty] | 1.83
26.1 | [7 empty] | 1.8
26.2 | [7 empty] | 1.8
26.3 | [7 empty] | 1.8
26.4 | [7 empty] | 1.73
26.5 | [7 empty] | 1.7
26.6 | [7 empty] | 1.7
26.7 | [7 empty] | 1.7
26.8 | [7 empty] | 1.63
26.9 | [7 empty] | 1.6
27 | [7 empty] | 1.6
27.1 | [7 empty] | 1.6
27.2 | [7 empty] | 1.6
27.3 | [7 empty] | 1.53
27.4 | [7 empty] | 1.5
27.5 | [7 empty] | 1.5
27.6 | [7 empty] | 1.5
27.7 | [7 empty] | 1.5
27.8 | [7 empty] | 1.43
27.9 | [7 empty] | 1.4
28 | [7 empty] | 1.4
28.1 | [7 empty] | 1.4
28.2 | [7 empty] | 1.4
28.3 | [7 empty] | 1.4
28.4 | [7 empty] | 1.33
28.5 | [7 empty] | 1.3
28.6 | [7 empty] | 1.3
28.7 | [7 empty] | 1.3
28.8 | [7 empty] | 1.3
28.9 | [7 empty] | 1.3

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 1.8
-29.7 | [7 empty] | 1.9
-29.6 | [7 empty] | 1.9
-29.5 | [7 empty] | 1.9
-29.4 | [7 empty] | 1.9
-29.3 | [7 empty] | 2
-29.2 | [7 empty] | 2
-29.1 | [7 empty] | 2
-29 | [7 empty] | 2
-28.9 | [7 empty] | 2.1
-28.8 | [7 empty] | 2.1
-28.7 | [7 empty] | 2.1
-28.6 | [7 empty] | 2.2
-28.5 | [7 empty] | 2.2
-28.4 | [7 empty] | 2.2
-28.3 | [7 empty] | 2.2
-28.2 | [7 empty] | 2.3
-28.1 | [7 empty] | 2.3
-28 | [7 empty] | 2.3
-27.9 | [7 empty] | 2.3
-27.8 | [7 empty] | 2.4
-27.7 | [7 empty] | 2.4
-27.6 | [7 empty] | 2.5
-27.5 | [7 empty] | 2.5
-27.4 | [7 empty] | 2.5
-27.3 | [7 empty] | 2.6
-27.2 | [7 empty] | 2.6
-27.1 | [7 empty] | 2.6
-27 | [7 empty] | 2.7
-26.9 | [7 empty] | 2.7
-26.8 | [7 empty] | 2.7
-26.7 | [7 empty] | 2.8
-26.6 | [7 empty] | 2.8
-26.5 | [7 empty] | 2.9
-26.4 | [7 empty] | 2.9
-26.3 | [7 empty] | 3
-26.2 | [7 empty] | 3
-26.1 | [7 empty] | 3.1
-26 | [7 empty] | 3.1
-25.9 | [7 empty] | 3.2
-25.8 | [7 empty] | 3.2
-25.7 | [7 empty] | 3.3
-25.6 | [7 empty] | 3.3
-25.5 | [7 empty] | 3.4
-25.4 | [7 empty] | 3.5
-25.3 | [7 empty] | 3.5
-25.2 | [7 empty] | 3.6
-25.1 | [7 empty] | 3.6
-25 | [7 empty] | 3.7
-24.9 | [7 empty] | 3.8
-24.8 | [7 empty] | 3.9
-24.7 | [7 empty] | 3.9
-24.6 | [7 empty] | 4
-24.5 | [7 empty] | 4.1
-24.4 | [7 empty] | 4.2
-24.3 | [7 empty] | 4.3
-24.2 | [7 empty] | 4.4
-24.1 | [7 empty] | 4.5
-24 | [7 empty] | 4.6
-23.9 | [7 empty] | 4.7
-23.8 | [7 empty] | 4.8
-23.7 | [7 empty] | 4.9
-23.6 | [7 empty] | 5
-23.5 | [7 empty] | 5.2
-23.4 | [7 empty] | 5.3
-23.3 | [7 empty] | 5.5
-23.2 | [7 empty] | 5.7
-23.1 | [7 empty] | 5.9
-23 | [7 empty] | 6.2
-22.9 | [7 empty] | 6.5
-22.8 | [7 empty] | 6.9
-22.7 | [7 empty] | 7.4
-22.6 | [7 empty] | 8
-22.5 | [7 empty] | 9
-22.4 | [7 empty] | 10.3
-22.3 | [7 empty] | 12.4
-22.2 | [7 empty] | 15
-22.1 | [7 empty] | 18.3
-22 | [7 empty] | 21.7
-21.9 | [7 empty] | 25.3
-21.8 | [7 empty] | 28.5
-21.7 | [7 empty] | 31.3
-21.6 | [7 empty] | 33.4
-21.5 | [7 empty] | 34.9
-21.4 | [6 empty] | 2.92 | 35.9
-21.3 | [6 empty] | 3 | 36.7
-21.2 | [6 empty] | 3 | 37.3
-21.1 | [6 empty] | 3.1 | 37.8
-21 | [6 empty] | 3.1 | 38.2
-20.9 | [6 empty] | 3.2 | 38.5
-20.8 | [6 empty] | 3.2 | 38.9
-20.7 | [6 empty] | 3.3 | 39.2
-20.6 | [6 empty] | 3.32 | 39.5
-20.5 | [6 empty] | 3.4 | 39.8
-20.4 | [6 empty] | 3.5 | 40.1
-20.3 | [6 empty] | 3.5 | 40.3
-20.2 | [6 empty] | 3.6 | 40.6
-20.1 | [6 empty] | 3.62 | 40.8
-20 | [6 empty] | 3.72 | 41.1
-19.9 | [6 empty] | 3.8 | 41.3
-19.8 | [6 empty] | 3.82 | 41.5
-19.7 | [6 empty] | 3.92 | 41.8
-19.6 | [6 empty] | 4.02 | 42
-19.5 | [6 empty] | 4.1 | 42.3
-19.4 | [6 empty] | 4.2 | 42.5
-19.3 | [6 empty] | 4.3 | 42.7
-19.2 | [6 empty] | 4.32 | 42.9
-19.1 | [6 empty] | 4.42 | 43.1
-19 | [6 empty] | 4.52 | 43.4
-18.9 | [6 empty] | 4.62 | 43.6
-18.8 | [6 empty] | 4.72 | 43.8
-18.7 | [6 empty] | 4.82 | 44
-18.6 | [6 empty] | 4.92 | 44.2
-18.5 | [6 empty] | 5.1 | 44.5
-18.4 | [6 empty] | 5.2 | 44.7
-18.3 | [6 empty] | 5.32 | 44.9
-18.2 | [6 empty] | 5.5 | 45.1
-18.1 | [6 empty] | 5.62 | 45.4
-18 | [6 empty] | 5.82 | 45.6
-17.9 | [6 empty] | 6 | 45.8
-17.8 | [6 empty] | 6.14 | 46
-17.7 | [6 empty] | 6.42 | 46.2
-17.6 | [6 empty] | 6.64 | 46.4
-17.5 | [6 empty] | 7.02 | 46.6
-17.4 | [6 empty] | 7.44 | 46.9
-17.3 | [6 empty] | 7.86 | 47.1
-17.2 | [6 empty] | 8.48 | 47.3
-17.1 | [6 empty] | 9.3 | 47.5
-17 | [6 empty] | 10.54 | 47.7
-16.9 | [6 empty] | 12.24 | 48
-16.8 | [6 empty] | 14.8 | 48.2
-16.7 | [6 empty] | 18.08 | 48.4
-16.6 | [6 empty] | 22.12 | 48.7
-16.5 | [6 empty] | 26.34 | 48.9
-16.4 | [6 empty] | 30.74 | 49.1
-16.3 | [6 empty] | 34.78 | 49.3
-16.2 | [6 empty] | 38.3 | 49.5
-16.1 | [6 empty] | 41.04 | 49.8
-16 | [5 empty] | 2.7 | 43.06 | 50
-15.9 | [5 empty] | 2.8 | 44.42 | 50.2
-15.8 | [5 empty] | 2.8 | 45.4 | 50.4
-15.7 | [5 empty] | 2.9 | 46.26 | 50.7
-15.6 | [5 empty] | 2.9 | 46.96 | 50.9
-15.5 | [5 empty] | 3 | 47.46 | 51.1
-15.4 | [5 empty] | 3.1 | 47.96 | 51.3
-15.3 | [5 empty] | 3.1 | 48.44 | 51.6
-15.2 | [5 empty] | 3.1 | 48.84 | 51.9
-15.1 | [5 empty] | 3.2 | 49.32 | 52.1
-15 | [5 empty] | 3.3 | 49.64 | 52.3
-14.9 | [5 empty] | 3.4 | 50.04 | 52.5
-14.8 | [5 empty] | 3.4 | 50.34 | 52.8
-14.7 | [5 empty] | 3.5 | 50.72 | 53
-14.6 | [5 empty] | 3.6 | 51.02 | 53.3
-14.5 | [5 empty] | 3.6 | 51.34 | 53.5
-14.4 | [5 empty] | 3.7 | 51.64 | 53.8
-14.3 | [5 empty] | 3.8 | 52.02 | 54
-14.2 | [5 empty] | 3.9 | 52.32 | 54.2
-14.1 | [5 empty] | 4 | 52.62 | 54.5
-14 | [5 empty] | 4.1 | 52.94 | 54.7
-13.9 | [5 empty] | 4.1 | 53.22 | 55
-13.8 | [5 empty] | 4.2 | 53.52 | 55.3
-13.7 | [5 empty] | 4.3 | 53.82 | 55.5
-13.6 | [5 empty] | 4.4 | 54.04 | 55.8
-13.5 | [5 empty] | 4.5 | 54.42 | 56
-13.4 | [5 empty] | 4.6 | 54.72 | 56.3
-13.3 | [5 empty] | 4.8 | 55.02 | 56.5
-13.2 | [5 empty] | 4.9 | 55.32 | 56.8
-13.1 | [5 empty] | 5 | 55.62 | 57
-13 | [5 empty] | 5.2 | 55.92 | 57.3
-12.9 | [5 empty] | 5.3 | 56.22 | 57.6
-12.8 | [5 empty] | 5.5 | 56.54 | 57.8
-12.7 | [5 empty] | 5.7 | 56.82 | 58.1
-12.6 | [5 empty] | 5.8 | 57.12 | 58.4
-12.5 | [5 empty] | 6 | 57.34 | 58.7
-12.4 | [5 empty] | 6.3 | 57.74 | 58.9
-12.3 | [5 empty] | 6.5 | 58.02 | 59.2
-12.2 | [5 empty] | 6.8 | 58.34 | 59.5
-12.1 | [5 empty] | 7.2 | 58.62 | 59.7
-12 | [5 empty] | 7.6 | 58.92 | 60
-11.9 | [5 empty] | 8.1 | 59.24 | 60.3
-11.8 | [5 empty] | 8.7 | 59.52 | 60.6
-11.7 | [5 empty] | 9.4 | 59.82 | 60.8
-11.6 | [5 empty] | 10.5 | 60.14 | 61.2
-11.5 | [5 empty] | 12 | 60.44 | 61.4
-11.4 | [5 empty] | 14.2 | 60.82 | 61.8
-11.3 | [5 empty] | 17.5 | 61.04 | 62
-11.2 | [5 empty] | 22 | 61.42 | 62.3
-11.1 | [5 empty] | 27.5 | 61.72 | 62.6
-11 | [5 empty] | 33.5 | 62.02 | 62.9
-10.9 | [5 empty] | 39.4 | 62.34 | 63.2
-10.8 | [5 empty] | 45.1 | 62.64 | 63.5
-10.7 | [5 empty] | 49.6 | 62.94 | 63.8
-10.6 | [5 empty] | 53.1 | 63.32 | 64.1
-10.5 | [5 empty] | 55.5 | 63.64 | 64.4
-10.4 |  |  |  |  | 1.98 | 57.2 | 63.92 | 64.7
-10.3 |  |  |  |  | 2 | 58.4 | 64.24 | 65
-10.2 |  |  |  |  | 2 | 59.5 | 64.54 | 65.3
-10.1 |  |  |  |  | 2.1 | 60.3 | 64.92 | 65.6
-10 |  |  |  |  | 2.1 | 61 | 65.24 | 65.9
-9.9 |  |  |  |  | 2.2 | 61.7 | 65.62 | 66.2
-9.8 |  |  |  |  | 2.2 | 62.3 | 65.86 | 66.6
-9.7 |  |  |  |  | 2.3 | 62.9 | 66.24 | 66.9
-9.6 |  |  |  |  | 2.3 | 63.4 | 66.54 | 67.2
-9.5 |  |  |  |  | 2.4 | 64 | 66.92 | 67.5
-9.4 |  |  |  |  | 2.48 | 64.5 | 67.24 | 67.9
-9.3 |  |  |  | 1.6 | 2.5 | 64.9 | 67.62 | 68.2
-9.2 |  |  |  | 1.7 | 2.58 | 65.4 | 67.94 | 68.5
-9.1 |  |  |  | 1.7 | 2.6 | 65.8 | 68.32 | 68.8
-9 |  |  |  | 1.7 | 2.7 | 66.3 | 68.64 | 69.2
-8.9 |  |  |  | 1.8 | 2.78 | 66.8 | 68.94 | 69.5
-8.8 |  |  |  | 1.8 | 2.8 | 67.3 | 69.42 | 69.9
-8.7 |  |  |  | 1.9 | 2.9 | 67.7 | 69.72 | 70.2
-8.6 |  |  |  | 1.9 | 3 | 68.2 | 70.04 | 70.5
-8.5 |  |  |  | 1.98 | 3.1 | 68.6 | 70.44 | 70.9
-8.4 |  |  |  | 2 | 3.2 | 69 | 70.74 | 71.2
-8.3 |  |  |  | 2.08 | 3.28 | 69.5 | 71.14 | 71.6
-8.2 |  |  | 1.26 | 2.1 | 3.38 | 69.8 | 71.54 | 71.9
-8.1 |  |  | 1.3 | 2.2 | 3.48 | 70.3 | 71.92 | 72.3
-8 |  |  | 1.3 | 2.28 | 3.58 | 70.7 | 72.24 | 72.7
-7.9 |  |  | 1.36 | 2.3 | 3.7 | 71.1 | 72.56 | 73
-7.8 |  |  | 1.4 | 2.38 | 3.88 | 71.5 | 72.94 | 73.4
-7.7 |  |  | 1.4 | 2.4 | 3.98 | 72 | 73.34 | 73.7
-7.6 |  |  | 1.46 | 2.5 | 4.08 | 72.4 | 73.72 | 74.1
-7.5 |  |  | 1.5 | 2.6 | 4.28 | 72.8 | 74.06 | 74.5
-7.4 |  |  | 1.5 | 2.68 | 4.48 | 73.2 | 74.54 | 74.9
-7.3 |  |  | 1.6 | 2.78 | 4.6 | 73.6 | 74.84 | 75.2
-7.2 |  |  | 1.6 | 2.8 | 4.86 | 74 | 75.24 | 75.6
-7.1 |  | 0.88 | 1.66 | 2.9 | 5.08 | 74.5 | 75.64 | 75.9
-7 |  | 0.9 | 1.7 | 3 | 5.28 | 74.9 | 76.04 | 76.3
-6.9 |  | 0.9 | 1.8 | 3.18 | 5.58 | 75.3 | 76.42 | 76.7
-6.8 |  | 0.9 | 1.8 | 3.28 | 5.86 | 75.7 | 76.74 | 77.1
-6.7 |  | 0.9 | 1.9 | 3.38 | 6.18 | 76.2 | 77.16 | 77.5
-6.6 | 0.6 | 0.98 | 1.9 | 3.5 | 6.66 | 76.6 | 77.54 | 77.9
-6.5 | 0.62 | 1 | 2 | 3.68 | 7.14 | 77 | 77.94 | 78.3
-6.4 | 0.7 | 1 | 2.1 | 3.78 | 7.74 | 77.5 | 78.34 | 78.7
-6.3 | 0.7 | 1 | 2.1 | 3.98 | 8.52 | 77.9 | 78.66 | 79.1
-6.2 | 0.7 | 1.1 | 2.2 | 4.1 | 9.5 | 78.3 | 79.14 | 79.5
-6.1 | 0.7 | 1.1 | 2.3 | 4.28 | 10.84 | 78.7 | 79.54 | 79.9
-6 | 0.7 | 1.1 | 2.36 | 4.56 | 12.6 | 79.2 | 79.96 | 80.3
-5.9 | 0.72 | 1.2 | 2.46 | 4.78 | 15.46 | 79.6 | 80.34 | 80.7
-5.8 | 0.8 | 1.2 | 2.56 | 5.06 | 19.72 | 80.1 | 80.82 | 81
-5.7 | 0.8 | 1.2 | 2.66 | 5.36 | 25.84 | 80.5 | 81.14 | 81.5
-5.6 | 0.8 | 1.3 | 2.76 | 5.76 | 33.3 | 80.9 | 81.54 | 81.8
-5.5 | 0.8 | 1.3 | 2.86 | 6.16 | 41.56 | 81.4 | 81.96 | 82.2
-5.4 | 0.9 | 1.38 | 2.96 | 6.64 | 49.98 | 81.8 | 82.44 | 82.6
-5.3 | 0.9 | 1.4 | 3.1 | 7.32 | 57.92 | 82.3 | 82.84 | 83.1
-5.2 | 0.9 | 1.5 | 3.26 | 8.12 | 64.86 | 82.7 | 83.26 | 83.5
-5.1 | 0.92 | 1.5 | 3.46 | 9.18 | 69.94 | 83.2 | 83.74 | 83.9
-5 | 1 | 1.6 | 3.6 | 10.54 | 73.5 | 83.6 | 84.14 | 84.3
-4.9 | 1 | 1.6 | 3.82 | 12.48 | 75.9 | 84.1 | 84.54 | 84.7
-4.8 | 1.02 | 1.7 | 4.06 | 15.36 | 77.74 | 84.5 | 84.96 | 85.1
-4.7 | 1.1 | 1.78 | 4.32 | 19.9 | 79.26 | 84.9 | 85.42 | 85.6
-4.6 | 1.1 | 1.88 | 4.62 | 26.3 | 80.48 | 85.4 | 85.84 | 86
-4.5 | 1.2 | 1.9 | 4.92 | 34.26 | 81.4 | 85.9 | 86.24 | 86.4
-4.4 | 1.2 | 2 | 5.42 | 43.26 | 82.4 | 86.3 | 86.64 | 86.8
-4.3 | 1.22 | 2.1 | 5.98 | 52.3 | 83.22 | 86.7 | 87.06 | 87.2
-4.2 | 1.3 | 2.28 | 6.64 | 60.88 | 84.1 | 87.2 | 87.54 | 87.7
-4.1 | 1.4 | 2.38 | 7.44 | 68.4 | 84.74 | 87.7 | 87.94 | 88.1
-4 | 1.42 | 2.5 | 8.56 | 74.1 | 85.44 | 88.1 | 88.44 | 88.5
-3.9 | 1.5 | 2.68 | 9.98 | 77.98 | 86.12 | 88.6 | 88.84 | 88.9
-3.8 | 1.6 | 2.88 | 12.02 | 80.66 | 86.76 | 88.9 | 89.24 | 89.4
-3.7 | 1.7 | 3.08 | 15.08 | 82.46 | 87.4 | 89.4 | 89.64 | 89.8
-3.6 | 1.72 | 3.28 | 19.78 | 84.04 | 87.96 | 89.8 | 90.04 | 90.2
-3.5 | 1.9 | 3.58 | 26.54 | 85.28 | 88.54 | 90.3 | 90.44 | 90.7
-3.4 | 2 | 3.96 | 34.9 | 86.4 | 89.16 | 90.8 | 90.94 | 91
-3.3 | 2.12 | 4.36 | 44.4 | 87.42 | 89.56 | 91.1 | 91.34 | 91.5
-3.2 | 2.32 | 4.86 | 54.3 | 88.3 | 90.14 | 91.5 | 91.72 | 91.8
-3.1 | 2.52 | 5.52 | 63.64 | 88.96 | 90.64 | 91.9 | 92.06 | 92.3
-3 | 2.72 | 6.4 | 71.72 | 89.72 | 91.16 | 92.3 | 92.54 | 92.6
-2.9 | 3.02 | 7.56 | 78.12 | 90.42 | 91.64 | 92.8 | 92.94 | 93.1
-2.8 | 3.34 | 9.12 | 82.42 | 91.06 | 92.24 | 93.2 | 93.26 | 93.5
-2.7 | 3.76 | 11.26 | 85.28 | 91.66 | 92.68 | 93.6 | 93.72 | 93.8
-2.6 | 4.36 | 14.42 | 87.34 | 92.14 | 93.14 | 94.1 | 94.12 | 94.2
-2.5 | 5 | 19.6 | 88.98 | 92.76 | 93.66 | 94.4 | 94.46 | 94.6
-2.4 | 5.92 | 26.88 | 90.46 | 93.44 | 93.98 | 94.8 | 94.84 | 95
-2.3 | 7.16 | 36.12 | 91.5 | 93.96 | 94.44 | 95.1 | 95.22 | 95.3
-2.2 | 8.94 | 46.54 | 92.44 | 94.38 | 94.96 | 95.5 | 95.54 | 95.7
-2.1 | 11.64 | 57.22 | 93.18 | 94.84 | 95.36 | 95.9 | 95.84 | 96
-2 | 15.64 | 67.24 | 93.98 | 95.36 | 95.76 | 96.3 | 96.24 | 96.4
-1.9 | 21.78 | 75.88 | 94.68 | 95.82 | 96.16 | 96.6 | 96.62 | 96.7
-1.8 | 30.16 | 82.6 | 95.28 | 96.18 | 96.48 | 96.9 | 96.84 | 97
-1.7 | 40.4 | 87.3 | 95.78 | 96.46 | 96.86 | 97.2 | 97.14 | 97.3
-1.6 | 51.52 | 90.24 | 96.32 | 96.86 | 97.26 | 97.5 | 97.52 | 97.5
-1.5 | 62.36 | 92.4 | 96.82 | 97.26 | 97.46 | 97.8 | 97.74 | 97.8
-1.4 | 72.5 | 93.96 | 97.22 | 97.68 | 97.94 | 98.1 | 98.02 | 98.1
-1.3 | 80.8 | 95.1 | 97.56 | 98.04 | 98.18 | 98.3 | 98.3 | 98.4
-1.2 | 87.08 | 96.1 | 97.92 | 98.28 | 98.38 | 98.6 | 98.5 | 98.6
-1.1 | 91.18 | 96.92 | 98.32 | 98.48 | 98.58 | 98.8 | 98.64 | 98.8
-1 | 93.68 | 97.64 | 98.56 | 98.76 | 98.88 | 99 | 98.92 | 99
-0.9 | 95.36 | 98.14 | 98.86 | 99.08 | 99.1 | 99.2 | 99.12 | 99.1
-0.8 | 96.72 | 98.56 | 99.12 | 99.28 | 99.28 | 99.4 | 99.22 | 99.3
-0.7 | 97.68 | 98.98 | 99.36 | 99.38 | 99.58 | 99.6 | 99.42 | 99.5
-0.6 | 98.46 | 99.28 | 99.56 | 99.5 | 99.68 | 99.7 | 99.62 | 99.6
-0.5 | 98.94 | 99.48 | 99.66 | 99.68 | 99.78 | 99.8 | 99.72 | 99.7
-0.4 | 99.34 | 99.68 | 99.86 | 99.78 | 99.88 | 99.9 | 99.82 | 99.8
-0.3 | 99.64 | 99.88 | 99.9 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9
-0.2 | 99.9 | 99.9 | 99.96 | 99.9 | 100 | 99.9 | 99.92 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.88 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.68 | 99.9 | 99.9 | 99.92 | 99.9 | 99.9 | 100 | 99.9
0.4 | 99.44 | 99.8 | 99.84 | 99.82 | 99.9 | 99.9 | 99.9 | 99.8
0.5 | 98.96 | 99.52 | 99.64 | 99.7 | 99.8 | 99.8 | 99.8 | 99.7
0.6 | 98.52 | 99.32 | 99.44 | 99.52 | 99.62 | 99.7 | 99.7 | 99.6
0.7 | 97.8 | 99.04 | 99.34 | 99.4 | 99.5 | 99.5 | 99.58 | 99.5
0.8 | 96.78 | 98.64 | 99.14 | 99.12 | 99.4 | 99.3 | 99.38 | 99.3
0.9 | 95.54 | 98.26 | 98.88 | 98.92 | 99.2 | 99.2 | 99.2 | 99.1
1 | 93.8 | 97.66 | 98.54 | 98.72 | 98.94 | 99 | 99.08 | 99
1.1 | 91.38 | 96.98 | 98.28 | 98.62 | 98.64 | 98.8 | 98.88 | 98.8
1.2 | 87.46 | 96.2 | 97.88 | 98.32 | 98.42 | 98.5 | 98.68 | 98.6
1.3 | 81.34 | 95.22 | 97.58 | 98.04 | 98.12 | 98.3 | 98.46 | 98.4
1.4 | 73.08 | 93.96 | 97.18 | 97.64 | 97.82 | 98.1 | 98.16 | 98.1
1.5 | 63.14 | 92.38 | 96.68 | 97.32 | 97.62 | 97.8 | 97.88 | 97.8
1.6 | 52.26 | 90.24 | 96.24 | 96.94 | 97.16 | 97.5 | 97.58 | 97.5
1.7 | 41.12 | 87.06 | 95.78 | 96.64 | 96.84 | 97.1 | 97.26 | 97.3
1.8 | 30.78 | 82.16 | 95.32 | 96.22 | 96.54 | 96.8 | 96.88 | 97
1.9 | 22.38 | 75.36 | 94.72 | 95.76 | 96.04 | 96.5 | 96.58 | 96.7
2 | 16 | 66.62 | 93.96 | 95.32 | 95.72 | 96.2 | 96.36 | 96.4
2.1 | 11.8 | 56.44 | 93.26 | 94.92 | 95.34 | 95.9 | 95.96 | 96
2.2 | 9.1 | 45.68 | 92.52 | 94.52 | 94.84 | 95.4 | 95.56 | 95.7
2.3 | 7.26 | 35.5 | 91.6 | 93.94 | 94.44 | 95.1 | 95.26 | 95.3
2.4 | 6 | 26.26 | 90.48 | 93.44 | 94.04 | 94.8 | 94.96 | 95
2.5 | 5.12 | 19.14 | 89.12 | 92.88 | 93.54 | 94.3 | 94.56 | 94.6
2.6 | 4.34 | 14.12 | 87.4 | 92.36 | 93.14 | 93.9 | 94.16 | 94.2
2.7 | 3.84 | 11.06 | 85.34 | 91.76 | 92.66 | 93.6 | 93.78 | 93.8
2.8 | 3.36 | 8.9 | 82.42 | 91.18 | 92.14 | 93.1 | 93.36 | 93.5
2.9 | 2.98 | 7.44 | 78 | 90.56 | 91.66 | 92.7 | 92.98 | 93.1
3 | 2.76 | 6.3 | 71.78 | 89.86 | 91.16 | 92.4 | 92.64 | 92.6
3.1 | 2.48 | 5.48 | 63.66 | 89.16 | 90.66 | 92 | 92.18 | 92.3
3.2 | 2.28 | 4.78 | 54.16 | 88.4 | 90.16 | 91.5 | 91.86 | 91.8
3.3 | 2.1 | 4.26 | 44.3 | 87.4 | 89.64 | 91 | 91.36 | 91.5
3.4 | 2 | 3.84 | 34.8 | 86.5 | 89.14 | 90.6 | 90.96 | 91
3.5 | 1.88 | 3.54 | 26.3 | 85.32 | 88.54 | 90.2 | 90.48 | 90.7
3.6 | 1.78 | 3.32 | 19.58 | 84.02 | 87.94 | 89.7 | 90.16 | 90.2
3.7 | 1.68 | 3.02 | 15.06 | 82.54 | 87.36 | 89.2 | 89.76 | 89.8
3.8 | 1.58 | 2.82 | 12.02 | 80.62 | 86.68 | 88.8 | 89.26 | 89.4
3.9 | 1.5 | 2.62 | 9.96 | 77.82 | 86.06 | 88.4 | 88.86 | 88.9
4 | 1.48 | 2.52 | 8.54 | 73.86 | 85.46 | 88 | 88.38 | 88.5
4.1 | 1.38 | 2.32 | 7.5 | 68.12 | 84.76 | 87.5 | 88.04 | 88.1
4.2 | 1.3 | 2.22 | 6.66 | 60.7 | 83.98 | 87.1 | 87.56 | 87.7
4.3 | 1.28 | 2.12 | 6.02 | 52.08 | 83.18 | 86.7 | 87.08 | 87.2
4.4 | 1.2 | 2.02 | 5.48 | 43.02 | 82.28 | 86.2 | 86.66 | 86.8
4.5 | 1.18 | 1.92 | 4.98 | 34.08 | 81.4 | 85.8 | 86.28 | 86.4
4.6 | 1.1 | 1.82 | 4.68 | 26.32 | 80.3 | 85.3 | 85.94 | 86
4.7 | 1.1 | 1.8 | 4.34 | 19.88 | 79.12 | 84.9 | 85.38 | 85.6
4.8 | 1 | 1.7 | 4.08 | 15.38 | 77.66 | 84.4 | 85.06 | 85.1
4.9 | 1 | 1.62 | 3.84 | 12.44 | 75.88 | 84 | 84.66 | 84.7
5 | 1 | 1.6 | 3.64 | 10.48 | 73.46 | 83.6 | 84.24 | 84.3
5.1 | 0.98 | 1.5 | 3.44 | 9.12 | 69.98 | 83.1 | 83.76 | 83.9
5.2 | 0.9 | 1.5 | 3.3 | 8.18 | 64.88 | 82.7 | 83.28 | 83.5
5.3 | 0.9 | 1.4 | 3.14 | 7.36 | 58.12 | 82.3 | 82.96 | 83.1
5.4 | 0.88 | 1.4 | 3 | 6.76 | 50.22 | 81.8 | 82.54 | 82.6
5.5 | 0.8 | 1.3 | 2.9 | 6.16 | 41.78 | 81.3 | 82.14 | 82.2
5.6 | 0.8 | 1.3 | 2.74 | 5.74 | 33.42 | 80.9 | 81.68 | 81.8
5.7 | 0.8 | 1.2 | 2.64 | 5.34 | 26 | 80.5 | 81.26 | 81.5
5.8 | 0.8 | 1.2 | 2.54 | 5.04 | 19.96 | 80 | 80.86 | 81
5.9 | 0.7 | 1.2 | 2.44 | 4.82 | 15.68 | 79.6 | 80.46 | 80.7
6 | 0.7 | 1.1 | 2.34 | 4.52 | 12.76 | 79.1 | 80.04 | 80.3
6.1 | 0.7 | 1.1 | 2.3 | 4.32 | 10.88 | 78.7 | 79.58 | 79.9
6.2 | 0.7 | 1.1 | 2.2 | 4.12 | 9.52 | 78.2 | 79.26 | 79.5
6.3 | 0.7 | 1 | 2.14 | 3.92 | 8.58 | 77.9 | 78.84 | 79.1
6.4 | 0.7 | 1 | 2.1 | 3.82 | 7.78 | 77.4 | 78.38 | 78.7
6.5 | 0.68 | 1 | 2 | 3.62 | 7.16 | 77 | 77.96 | 78.3
6.6 |  | 0.92 | 1.94 | 3.52 | 6.64 | 76.5 | 77.56 | 77.9
6.7 |  | 0.9 | 1.9 | 3.4 | 6.24 | 76.2 | 77.24 | 77.5
6.8 |  | 0.9 | 1.8 | 3.22 | 5.84 | 75.7 | 76.86 | 77.1
6.9 |  | 0.9 | 1.8 | 3.12 | 5.54 | 75.3 | 76.46 | 76.7
7 |  | 0.9 | 1.7 | 3.02 | 5.32 | 74.8 | 76.06 | 76.3
7.1 |  | 0.82 | 1.7 | 2.92 | 5.02 | 74.4 | 75.66 | 75.9
7.2 |  |  | 1.6 | 2.9 | 4.82 | 74 | 75.26 | 75.6
7.3 |  |  | 1.6 | 2.72 | 4.62 | 73.6 | 74.86 | 75.2
7.4 |  |  | 1.54 | 2.7 | 4.42 | 73.1 | 74.48 | 74.9
7.5 |  |  | 1.5 | 2.6 | 4.24 | 72.7 | 74.16 | 74.5
7.6 |  |  | 1.44 | 2.5 | 4.12 | 72.3 | 73.76 | 74.1
7.7 |  |  | 1.4 | 2.42 | 4 | 71.9 | 73.36 | 73.7
7.8 |  |  | 1.4 | 2.4 | 3.82 | 71.4 | 72.98 | 73.4
7.9 |  |  | 1.34 | 2.3 | 3.72 | 71 | 72.58 | 73
8 |  |  | 1.3 | 2.22 | 3.62 | 70.6 | 72.26 | 72.7
8.1 |  |  | 1.3 | 2.2 | 3.5 | 70.2 | 71.88 | 72.3
8.2 |  |  | 1.3 | 2.1 | 3.4 | 69.7 | 71.56 | 71.9
8.3 |  |  |  | 2.1 | 3.3 | 69.3 | 71.18 | 71.6
8.4 |  |  |  | 2 | 3.2 | 68.9 | 70.84 | 71.2
8.5 |  |  |  | 2 | 3.1 | 68.4 | 70.36 | 70.9
8.6 |  |  |  | 1.9 | 3 | 68 | 70.06 | 70.5
8.7 |  |  |  | 1.9 | 2.92 | 67.6 | 69.76 | 70.2
8.8 |  |  |  | 1.82 | 2.9 | 67.1 | 69.38 | 69.9
8.9 |  |  |  | 1.8 | 2.8 | 66.7 | 69.06 | 69.5
9 |  |  |  | 1.72 | 2.7 | 66.3 | 68.68 | 69.2
9.1 |  |  |  | 1.7 | 2.62 | 65.8 | 68.36 | 68.8
9.2 |  |  |  | 1.7 | 2.6 | 65.4 | 67.96 | 68.5
9.3 |  |  |  | 1.6 | 2.5 | 64.8 | 67.58 | 68.2
9.4 |  |  |  |  | 2.42 | 64.4 | 67.36 | 67.9
9.5 |  |  |  |  | 2.4 | 64 | 66.96 | 67.5
9.6 |  |  |  |  | 2.32 | 63.4 | 66.56 | 67.2
9.7 |  |  |  |  | 2.3 | 62.9 | 66.28 | 66.9
9.8 |  |  |  |  | 2.2 | 62.3 | 65.88 | 66.6
9.9 |  |  |  |  | 2.2 | 61.6 | 65.58 | 66.2
10 |  |  |  |  | 2.1 | 61 | 65.26 | 65.9
10.1 |  |  |  |  | 2.1 | 60.2 | 64.88 | 65.6
10.2 |  |  |  |  | 2.02 | 59.3 | 64.66 | 65.3
10.3 |  |  |  |  | 2 | 58.3 | 64.38 | 65
10.4 |  |  |  |  | 1.92 | 57.1 | 63.96 | 64.7
10.5 | [5 empty] | 55.4 | 63.66 | 64.4
10.6 | [5 empty] | 52.9 | 63.28 | 64.1
10.7 | [5 empty] | 49.4 | 62.98 | 63.8
10.8 | [5 empty] | 44.6 | 62.68 | 63.5
10.9 | [5 empty] | 39 | 62.36 | 63.2
11 | [5 empty] | 33.1 | 62.06 | 62.9
11.1 | [5 empty] | 27.1 | 61.68 | 62.6
11.2 | [5 empty] | 21.7 | 61.38 | 62.3
11.3 | [5 empty] | 17.2 | 61.08 | 62
11.4 | [5 empty] | 14 | 60.84 | 61.8
11.5 | [5 empty] | 11.8 | 60.48 | 61.4
11.6 | [5 empty] | 10.4 | 60.16 | 61.2
11.7 | [5 empty] | 9.4 | 59.88 | 60.8
11.8 | [5 empty] | 8.6 | 59.56 | 60.6
11.9 | [5 empty] | 8 | 59.26 | 60.3
12 | [5 empty] | 7.5 | 58.96 | 60
12.1 | [5 empty] | 7.1 | 58.66 | 59.7
12.2 | [5 empty] | 6.8 | 58.36 | 59.5
12.3 | [5 empty] | 6.5 | 57.96 | 59.2
12.4 | [5 empty] | 6.2 | 57.66 | 58.9
12.5 | [5 empty] | 6 | 57.38 | 58.7
12.6 | [5 empty] | 5.8 | 57.08 | 58.4
12.7 | [5 empty] | 5.6 | 56.78 | 58.1
12.8 | [5 empty] | 5.5 | 56.56 | 57.8
12.9 | [5 empty] | 5.3 | 56.26 | 57.6
13 | [5 empty] | 5.2 | 55.88 | 57.3
13.1 | [5 empty] | 5 | 55.58 | 57
13.2 | [5 empty] | 4.9 | 55.36 | 56.8
13.3 | [5 empty] | 4.7 | 55.06 | 56.5
13.4 | [5 empty] | 4.6 | 54.68 | 56.3
13.5 | [5 empty] | 4.5 | 54.38 | 56
13.6 | [5 empty] | 4.4 | 54.16 | 55.8
13.7 | [5 empty] | 4.3 | 53.86 | 55.5
13.8 | [5 empty] | 4.2 | 53.56 | 55.3
13.9 | [5 empty] | 4.1 | 53.26 | 55
14 | [5 empty] | 4 | 52.96 | 54.7
14.1 | [5 empty] | 3.9 | 52.66 | 54.5
14.2 | [5 empty] | 3.8 | 52.28 | 54.2
14.3 | [5 empty] | 3.8 | 52.06 | 54
14.4 | [5 empty] | 3.7 | 51.68 | 53.8
14.5 | [5 empty] | 3.6 | 51.46 | 53.5
14.6 | [5 empty] | 3.5 | 51.06 | 53.3
14.7 | [5 empty] | 3.4 | 50.76 | 53
14.8 | [5 empty] | 3.4 | 50.44 | 52.8
14.9 | [5 empty] | 3.3 | 50.06 | 52.5
15 | [5 empty] | 3.3 | 49.68 | 52.3
15.1 | [5 empty] | 3.2 | 49.36 | 52.1
15.2 | [5 empty] | 3.1 | 48.96 | 51.9
15.3 | [5 empty] | 3.1 | 48.56 | 51.6
15.4 | [5 empty] | 3 | 48.06 | 51.3
15.5 | [5 empty] | 3 | 47.56 | 51.1
15.6 | [5 empty] | 2.9 | 47.02 | 50.9
15.7 | [5 empty] | 2.8 | 46.42 | 50.7
15.8 | [5 empty] | 2.8 | 45.62 | 50.4
15.9 | [5 empty] | 2.8 | 44.7 | 50.2
16 | [6 empty] | 43.44 | 50
16.1 | [6 empty] | 41.68 | 49.8
16.2 | [6 empty] | 39.08 | 49.5
16.3 | [6 empty] | 35.72 | 49.3
16.4 | [6 empty] | 31.66 | 49.1
16.5 | [6 empty] | 27.26 | 48.9
16.6 | [6 empty] | 22.86 | 48.7
16.7 | [6 empty] | 18.72 | 48.4
16.8 | [6 empty] | 15.2 | 48.2
16.9 | [6 empty] | 12.5 | 48
17 | [6 empty] | 10.66 | 47.7
17.1 | [6 empty] | 9.42 | 47.5
17.2 | [6 empty] | 8.62 | 47.3
17.3 | [6 empty] | 7.94 | 47.1
17.4 | [6 empty] | 7.46 | 46.9
17.5 | [6 empty] | 7.06 | 46.6
17.6 | [6 empty] | 6.76 | 46.4
17.7 | [6 empty] | 6.46 | 46.2
17.8 | [6 empty] | 6.18 | 46
17.9 | [6 empty] | 5.98 | 45.8
18 | [6 empty] | 5.78 | 45.6
18.1 | [6 empty] | 5.68 | 45.4
18.2 | [6 empty] | 5.48 | 45.1
18.3 | [6 empty] | 5.38 | 44.9
18.4 | [6 empty] | 5.2 | 44.7
18.5 | [6 empty] | 5.08 | 44.5
18.6 | [6 empty] | 4.98 | 44.2
18.7 | [6 empty] | 4.88 | 44
18.8 | [6 empty] | 4.78 | 43.8
18.9 | [6 empty] | 4.6 | 43.6
19 | [6 empty] | 4.5 | 43.4
19.1 | [6 empty] | 4.48 | 43.1
19.2 | [6 empty] | 4.3 | 42.9
19.3 | [6 empty] | 4.28 | 42.7
19.4 | [6 empty] | 4.18 | 42.5
19.5 | [6 empty] | 4.1 | 42.3
19.6 | [6 empty] | 4 | 42
19.7 | [6 empty] | 3.9 | 41.8
19.8 | [6 empty] | 3.88 | 41.5
19.9 | [6 empty] | 3.8 | 41.3
20 | [6 empty] | 3.7 | 41.1
20.1 | [6 empty] | 3.68 | 40.8
20.2 | [6 empty] | 3.6 | 40.6
20.3 | [6 empty] | 3.5 | 40.3
20.4 | [6 empty] | 3.48 | 40.1
20.5 | [6 empty] | 3.4 | 39.8
20.6 | [6 empty] | 3.3 | 39.5
20.7 | [6 empty] | 3.3 | 39.2
20.8 | [6 empty] | 3.2 | 38.9
20.9 | [6 empty] | 3.2 | 38.5
21 | [6 empty] | 3.1 | 38.2
21.1 | [6 empty] | 3.1 | 37.8
21.2 | [6 empty] | 3 | 37.3
21.3 | [6 empty] | 3 | 36.7
21.4 | [6 empty] | 2.9 | 35.9
21.5 | [7 empty] | 34.9
21.6 | [7 empty] | 33.4
21.7 | [7 empty] | 31.3
21.8 | [7 empty] | 28.5
21.9 | [7 empty] | 25.3
22 | [7 empty] | 21.7
22.1 | [7 empty] | 18.3
22.2 | [7 empty] | 15
22.3 | [7 empty] | 12.4
22.4 | [7 empty] | 10.3
22.5 | [7 empty] | 9
22.6 | [7 empty] | 8
22.7 | [7 empty] | 7.4
22.8 | [7 empty] | 6.9
22.9 | [7 empty] | 6.5
23 | [7 empty] | 6.2
23.1 | [7 empty] | 5.9
23.2 | [7 empty] | 5.7
23.3 | [7 empty] | 5.5
23.4 | [7 empty] | 5.3
23.5 | [7 empty] | 5.2
23.6 | [7 empty] | 5
23.7 | [7 empty] | 4.9
23.8 | [7 empty] | 4.8
23.9 | [7 empty] | 4.7
24 | [7 empty] | 4.6
24.1 | [7 empty] | 4.5
24.2 | [7 empty] | 4.4
24.3 | [7 empty] | 4.3
24.4 | [7 empty] | 4.2
24.5 | [7 empty] | 4.1
24.6 | [7 empty] | 4
24.7 | [7 empty] | 3.9
24.8 | [7 empty] | 3.9
24.9 | [7 empty] | 3.8
25 | [7 empty] | 3.7
25.1 | [7 empty] | 3.6
25.2 | [7 empty] | 3.6
25.3 | [7 empty] | 3.5
25.4 | [7 empty] | 3.5
25.5 | [7 empty] | 3.4
25.6 | [7 empty] | 3.3
25.7 | [7 empty] | 3.3
25.8 | [7 empty] | 3.2
25.9 | [7 empty] | 3.2
26 | [7 empty] | 3.1
26.1 | [7 empty] | 3.1
26.2 | [7 empty] | 3
26.3 | [7 empty] | 3
26.4 | [7 empty] | 2.9
26.5 | [7 empty] | 2.9
26.6 | [7 empty] | 2.8
26.7 | [7 empty] | 2.8
26.8 | [7 empty] | 2.7
26.9 | [7 empty] | 2.7
27 | [7 empty] | 2.7
27.1 | [7 empty] | 2.6
27.2 | [7 empty] | 2.6
27.3 | [7 empty] | 2.6
27.4 | [7 empty] | 2.5
27.5 | [7 empty] | 2.5
27.6 | [7 empty] | 2.5
27.7 | [7 empty] | 2.4
27.8 | [7 empty] | 2.4
27.9 | [7 empty] | 2.3
28 | [7 empty] | 2.3
28.1 | [7 empty] | 2.3
28.2 | [7 empty] | 2.3
28.3 | [7 empty] | 2.2
28.4 | [7 empty] | 2.2
28.5 | [7 empty] | 2.2
28.6 | [7 empty] | 2.2
28.7 | [7 empty] | 2.1
28.8 | [7 empty] | 2.1
28.9 | [7 empty] | 2.1
29 | [7 empty] | 2
29.1 | [7 empty] | 2
29.2 | [7 empty] | 2
29.3 | [7 empty] | 2
29.4 | [7 empty] | 1.9
29.5 | [7 empty] | 1.9
29.6 | [7 empty] | 1.9
29.7 | [7 empty] | 1.9
29.8 | [7 empty] | 1.8

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 3.81
-29.7 | [7 empty] | 3.9
-29.6 | [7 empty] | 3.91
-29.5 | [7 empty] | 4.01
-29.4 | [7 empty] | 4.1
-29.3 | [7 empty] | 4.11
-29.2 | [7 empty] | 4.2
-29.1 | [7 empty] | 4.21
-29 | [7 empty] | 4.3
-28.9 | [7 empty] | 4.31
-28.8 | [7 empty] | 4.4
-28.7 | [7 empty] | 4.41
-28.6 | [7 empty] | 4.51
-28.5 | [7 empty] | 4.6
-28.4 | [7 empty] | 4.61
-28.3 | [7 empty] | 4.71
-28.2 | [7 empty] | 4.8
-28.1 | [7 empty] | 4.81
-28 | [7 empty] | 4.91
-27.9 | [7 empty] | 5.01
-27.8 | [7 empty] | 5.1
-27.7 | [7 empty] | 5.11
-27.6 | [7 empty] | 5.21
-27.5 | [7 empty] | 5.31
-27.4 | [7 empty] | 5.4
-27.3 | [7 empty] | 5.41
-27.2 | [7 empty] | 5.51
-27.1 | [7 empty] | 5.61
-27 | [7 empty] | 5.71
-26.9 | [7 empty] | 5.81
-26.8 | [7 empty] | 5.91
-26.7 | [7 empty] | 6.01
-26.6 | [7 empty] | 6.11
-26.5 | [7 empty] | 6.21
-26.4 | [7 empty] | 6.31
-26.3 | [7 empty] | 6.41
-26.2 | [7 empty] | 6.52
-26.1 | [7 empty] | 6.71
-26 | [7 empty] | 6.81
-25.9 | [7 empty] | 6.91
-25.8 | [7 empty] | 7.02
-25.7 | [7 empty] | 7.21
-25.6 | [7 empty] | 7.32
-25.5 | [7 empty] | 7.52
-25.4 | [7 empty] | 7.71
-25.3 | [7 empty] | 7.82
-25.2 | [7 empty] | 8.03
-25.1 | [7 empty] | 8.33
-25 | [7 empty] | 8.63
-24.9 | [7 empty] | 8.94
-24.8 | [7 empty] | 9.35
-24.7 | [7 empty] | 9.86
-24.6 | [7 empty] | 10.49
-24.5 | [7 empty] | 11.43
-24.4 | [7 empty] | 12.77
-24.3 | [7 empty] | 14.52
-24.2 | [7 empty] | 16.77
-24.1 | [7 empty] | 19.49
-24 | [7 empty] | 22.4
-23.9 | [7 empty] | 25.37
-23.8 | [7 empty] | 28.04
-23.7 | [7 empty] | 30.39
-23.6 | [7 empty] | 32.24
-23.5 | [7 empty] | 33.6
-23.4 | [7 empty] | 34.58
-23.3 | [7 empty] | 35.36
-23.2 | [7 empty] | 35.95
-23.1 | [7 empty] | 36.44
-23 | [7 empty] | 36.84
-22.9 | [7 empty] | 37.24
-22.8 | [7 empty] | 37.63
-22.7 | [7 empty] | 37.93
-22.6 | [6 empty] | 4.96 | 38.23
-22.5 | [6 empty] | 5 | 38.53
-22.4 | [6 empty] | 5.1 | 38.82
-22.3 | [6 empty] | 5.16 | 39.03
-22.2 | [6 empty] | 5.26 | 39.33
-22.1 | [6 empty] | 5.3 | 39.62
-22 | [6 empty] | 5.46 | 39.82
-21.9 | [6 empty] | 5.5 | 40.03
-21.8 | [6 empty] | 5.6 | 40.32
-21.7 | [6 empty] | 5.7 | 40.53
-21.6 | [6 empty] | 5.8 | 40.82
-21.5 | [6 empty] | 5.9 | 41.02
-21.4 | [6 empty] | 6 | 41.23
-21.3 | [6 empty] | 6.06 | 41.52
-21.2 | [6 empty] | 6.16 | 41.72
-21.1 | [6 empty] | 6.26 | 41.93
-21 | [6 empty] | 6.4 | 42.22
-20.9 | [6 empty] | 6.5 | 42.42
-20.8 | [6 empty] | 6.6 | 42.62
-20.7 | [6 empty] | 6.7 | 42.83
-20.6 | [6 empty] | 6.86 | 43.12
-20.5 | [6 empty] | 6.96 | 43.32
-20.4 | [6 empty] | 7.1 | 43.52
-20.3 | [6 empty] | 7.26 | 43.72
-20.2 | [6 empty] | 7.36 | 43.92
-20.1 | [6 empty] | 7.5 | 44.13
-20 | [6 empty] | 7.66 | 44.42
-19.9 | [6 empty] | 7.8 | 44.62
-19.8 | [6 empty] | 7.96 | 44.83
-19.7 | [6 empty] | 8.16 | 45.12
-19.6 | [6 empty] | 8.36 | 45.32
-19.5 | [6 empty] | 8.5 | 45.52
-19.4 | [6 empty] | 8.76 | 45.72
-19.3 | [6 empty] | 8.96 | 45.92
-19.2 | [6 empty] | 9.22 | 46.13
-19.1 | [6 empty] | 9.52 | 46.42
-19 | [6 empty] | 9.82 | 46.62
-18.9 | [6 empty] | 10.28 | 46.83
-18.8 | [6 empty] | 10.78 | 47.12
-18.7 | [6 empty] | 11.38 | 47.32
-18.6 | [6 empty] | 12.2 | 47.52
-18.5 | [6 empty] | 13.32 | 47.73
-18.4 | [6 empty] | 15 | 48.01
-18.3 | [6 empty] | 17.28 | 48.12
-18.2 | [6 empty] | 20.16 | 48.33
-18.1 | [6 empty] | 23.58 | 48.62
-18 | [6 empty] | 27.28 | 48.82
-17.9 | [6 empty] | 30.98 | 49.03
-17.8 | [6 empty] | 34.52 | 49.32
-17.7 | [6 empty] | 37.54 | 49.53
-17.6 | [6 empty] | 39.96 | 49.82
-17.5 | [6 empty] | 41.72 | 50.02
-17.4 | [6 empty] | 43.06 | 50.22
-17.3 | [6 empty] | 44.04 | 50.42
-17.2 | [6 empty] | 44.84 | 50.63
-17.1 | [6 empty] | 45.48 | 50.92
-17 | [6 empty] | 46.08 | 51.12
-16.9 | [5 empty] | 4.7 | 46.62 | 51.33
-16.8 | [5 empty] | 4.8 | 47.02 | 51.62
-16.7 | [5 empty] | 4.8 | 47.42 | 51.82
-16.6 | [5 empty] | 4.9 | 47.86 | 52.03
-16.5 | [5 empty] | 5 | 48.22 | 52.32
-16.4 | [5 empty] | 5.06 | 48.62 | 52.52
-16.3 | [5 empty] | 5.16 | 48.96 | 52.72
-16.2 | [5 empty] | 5.26 | 49.32 | 52.93
-16.1 | [5 empty] | 5.36 | 49.66 | 53.22
-16 | [5 empty] | 5.4 | 49.96 | 53.42
-15.9 | [5 empty] | 5.5 | 50.22 | 53.63
-15.8 | [5 empty] | 5.6 | 50.62 | 53.92
-15.7 | [5 empty] | 5.76 | 50.96 | 54.13
-15.6 | [5 empty] | 5.8 | 51.22 | 54.42
-15.5 | [5 empty] | 5.96 | 51.56 | 54.62
-15.4 | [5 empty] | 6.06 | 51.82 | 54.83
-15.3 | [5 empty] | 6.16 | 52.16 | 55.12
-15.2 | [5 empty] | 6.3 | 52.46 | 55.32
-15.1 | [5 empty] | 6.4 | 52.76 | 55.53
-15 | [5 empty] | 6.56 | 53.06 | 55.82
-14.9 | [5 empty] | 6.66 | 53.36 | 56.03
-14.8 | [5 empty] | 6.76 | 53.66 | 56.32
-14.7 | [5 empty] | 6.9 | 53.96 | 56.53
-14.6 | [5 empty] | 7.1 | 54.32 | 56.82
-14.5 | [5 empty] | 7.2 | 54.56 | 57.03
-14.4 | [5 empty] | 7.36 | 54.86 | 57.32
-14.3 | [5 empty] | 7.56 | 55.16 | 57.53
-14.2 | [5 empty] | 7.7 | 55.46 | 57.83
-14.1 | [5 empty] | 7.86 | 55.76 | 58.12
-14 | [5 empty] | 8.06 | 56.12 | 58.33
-13.9 | [5 empty] | 8.26 | 56.42 | 58.62
-13.8 | [5 empty] | 8.46 | 56.76 | 58.82
-13.7 | [5 empty] | 8.66 | 56.96 | 59.03
-13.6 | [5 empty] | 8.86 | 57.26 | 59.32
-13.5 | [5 empty] | 9.12 | 57.56 | 59.54
-13.4 | [5 empty] | 9.36 | 57.86 | 59.92
-13.3 | [5 empty] | 9.66 | 58.16 | 60.12
-13.2 | [5 empty] | 9.96 | 58.46 | 60.33
-13.1 | [5 empty] | 10.32 | 58.76 | 60.63
-13 | [5 empty] | 10.78 | 59.02 | 60.93
-12.9 | [5 empty] | 11.28 | 59.36 | 61.22
-12.8 | [5 empty] | 11.94 | 59.66 | 61.43
-12.7 | [5 empty] | 12.74 | 59.96 | 61.72
-12.6 | [5 empty] | 13.76 | 60.22 | 61.94
-12.5 | [5 empty] | 15.18 | 60.56 | 62.32
-12.4 | [5 empty] | 17.28 | 60.86 | 62.53
-12.3 | [5 empty] | 20.22 | 61.16 | 62.83
-12.2 | [5 empty] | 24.22 | 61.46 | 63.12
-12.1 | [5 empty] | 28.84 | 61.76 | 63.33
-12 | [5 empty] | 33.96 | 62.12 | 63.63
-11.9 | [5 empty] | 39.06 | 62.42 | 63.93
-11.8 | [5 empty] | 43.98 | 62.66 | 64.23
-11.7 | [5 empty] | 48.2 | 63.02 | 64.53
-11.6 | [5 empty] | 51.44 | 63.32 | 64.82
-11.5 | [5 empty] | 53.7 | 63.66 | 65.03
-11.4 | [5 empty] | 55.48 | 63.96 | 65.33
-11.3 | [5 empty] | 56.7 | 64.26 | 65.63
-11.2 | [5 empty] | 57.74 | 64.56 | 65.93
-11.1 | [5 empty] | 58.64 | 64.96 | 66.22
-11 | [5 empty] | 59.38 | 65.22 | 66.44
-10.9 |  |  |  |  | 3.48 | 60.08 | 65.52 | 66.83
-10.8 |  |  |  |  | 3.5 | 60.68 | 65.86 | 67.12
-10.7 |  |  |  |  | 3.6 | 61.28 | 66.12 | 67.33
-10.6 |  |  |  |  | 3.7 | 61.82 | 66.52 | 67.63
-10.5 |  |  |  |  | 3.78 | 62.32 | 66.76 | 67.94
-10.4 |  |  |  |  | 3.8 | 62.82 | 67.12 | 68.32
-10.3 |  |  |  |  | 3.9 | 63.32 | 67.52 | 68.54
-10.2 |  |  |  |  | 4 | 63.72 | 67.76 | 68.92
-10.1 |  |  |  |  | 4.08 | 64.28 | 68.06 | 69.14
-10 |  |  |  |  | 4.18 | 64.72 | 68.46 | 69.53
-9.9 |  |  |  |  | 4.28 | 65.12 | 68.76 | 69.83
-9.8 |  |  |  |  | 4.38 | 65.58 | 69.02 | 70.13
-9.7 |  |  |  | 2.9 | 4.48 | 65.98 | 69.36 | 70.43
-9.6 |  |  |  | 3 | 4.58 | 66.42 | 69.72 | 70.73
-9.5 |  |  |  | 3.08 | 4.68 | 66.82 | 70.02 | 71.03
-9.4 |  |  |  | 3.1 | 4.78 | 67.32 | 70.46 | 71.33
-9.3 |  |  |  | 3.2 | 4.88 | 67.72 | 70.72 | 71.64
-9.2 |  |  |  | 3.3 | 4.98 | 68.12 | 71.12 | 72.03
-9.1 |  |  |  | 3.3 | 5.1 | 68.66 | 71.46 | 72.33
-9 |  |  |  | 3.4 | 5.28 | 69.08 | 71.72 | 72.63
-8.9 |  |  |  | 3.5 | 5.38 | 69.48 | 72.06 | 72.94
-8.8 |  |  |  | 3.58 | 5.5 | 69.82 | 72.42 | 73.33
-8.7 |  |  |  | 3.68 | 5.68 | 70.18 | 72.82 | 73.64
-8.6 |  |  | 2.2 | 3.78 | 5.8 | 70.62 | 73.12 | 74.02
-8.5 |  |  | 2.3 | 3.8 | 5.98 | 71.02 | 73.46 | 74.24
-8.4 |  |  | 2.3 | 3.9 | 6.1 | 71.42 | 73.82 | 74.63
-8.3 |  |  | 2.4 | 4 | 6.28 | 71.86 | 74.22 | 74.94
-8.2 |  |  | 2.4 | 4.1 | 6.48 | 72.28 | 74.62 | 75.33
-8.1 |  |  | 2.5 | 4.2 | 6.6 | 72.72 | 74.86 | 75.64
-8 |  |  | 2.5 | 4.3 | 6.88 | 73.12 | 75.22 | 76.03
-7.9 |  |  | 2.6 | 4.48 | 7.06 | 73.52 | 75.68 | 76.34
-7.8 |  |  | 2.7 | 4.58 | 7.28 | 73.98 | 76.06 | 76.73
-7.7 |  |  | 2.7 | 4.68 | 7.48 | 74.42 | 76.32 | 77.03
-7.6 |  |  | 2.8 | 4.78 | 7.76 | 74.82 | 76.66 | 77.33
-7.5 |  |  | 2.9 | 4.98 | 8.06 | 75.16 | 77.02 | 77.64
-7.4 |  |  | 2.9 | 5.08 | 8.36 | 75.62 | 77.42 | 78.04
-7.3 |  | 1.5 | 3 | 5.28 | 8.68 | 76.02 | 77.76 | 78.43
-7.2 |  | 1.5 | 3.1 | 5.38 | 9.14 | 76.42 | 78.12 | 78.74
-7.1 |  | 1.6 | 3.2 | 5.58 | 9.56 | 76.82 | 78.46 | 79.13
-7 |  | 1.6 | 3.3 | 5.7 | 10.14 | 77.22 | 78.82 | 79.44
-6.9 |  | 1.6 | 3.3 | 5.88 | 10.82 | 77.62 | 79.22 | 79.83
-6.8 |  | 1.7 | 3.4 | 6.08 | 11.62 | 78.12 | 79.56 | 80.15
-6.7 | 1.1 | 1.7 | 3.5 | 6.28 | 12.68 | 78.52 | 79.92 | 80.63
-6.6 | 1.1 | 1.8 | 3.7 | 6.48 | 14.04 | 78.86 | 80.32 | 80.94
-6.5 | 1.1 | 1.8 | 3.8 | 6.7 | 15.9 | 79.32 | 80.68 | 81.33
-6.4 | 1.1 | 1.86 | 3.9 | 6.98 | 18.68 | 79.72 | 81.12 | 81.63
-6.3 | 1.2 | 1.9 | 4 | 7.26 | 22.66 | 80.12 | 81.42 | 81.94
-6.2 | 1.2 | 1.96 | 4.1 | 7.58 | 28.2 | 80.52 | 81.82 | 82.34
-6.1 | 1.2 | 2 | 4.2 | 7.96 | 34.8 | 80.92 | 82.18 | 82.74
-6 | 1.3 | 2.06 | 4.4 | 8.36 | 42.34 | 81.32 | 82.62 | 83.14
-5.9 | 1.3 | 2.1 | 4.5 | 8.84 | 49.76 | 81.78 | 82.96 | 83.53
-5.8 | 1.3 | 2.2 | 4.7 | 9.44 | 57.1 | 82.22 | 83.36 | 83.83
-5.7 | 1.4 | 2.26 | 4.8 | 10.12 | 63.24 | 82.62 | 83.72 | 84.14
-5.6 | 1.4 | 2.3 | 5 | 10.92 | 68.06 | 83.06 | 84.12 | 84.54
-5.5 | 1.4 | 2.4 | 5.2 | 12.08 | 71.68 | 83.42 | 84.52 | 84.94
-5.4 | 1.5 | 2.46 | 5.4 | 13.44 | 74.2 | 83.82 | 84.86 | 85.34
-5.3 | 1.5 | 2.5 | 5.6 | 15.48 | 76.14 | 84.32 | 85.32 | 85.73
-5.2 | 1.6 | 2.6 | 5.8 | 18.38 | 77.64 | 84.72 | 85.62 | 86.04
-5.1 | 1.6 | 2.7 | 6.1 | 22.72 | 78.88 | 85.12 | 86.08 | 86.43
-5 | 1.7 | 2.8 | 6.4 | 28.54 | 80 | 85.52 | 86.46 | 86.75
-4.9 | 1.7 | 2.9 | 6.7 | 35.84 | 80.92 | 85.98 | 86.82 | 87.23
-4.8 | 1.8 | 2.96 | 7.2 | 43.8 | 81.72 | 86.42 | 87.12 | 87.54
-4.7 | 1.8 | 3.06 | 7.7 | 52.26 | 82.6 | 86.88 | 87.52 | 87.93
-4.6 | 1.9 | 3.2 | 8.3 | 60.14 | 83.42 | 87.26 | 87.98 | 88.24
-4.5 | 1.9 | 3.36 | 8.9 | 67.16 | 84.12 | 87.62 | 88.36 | 88.64
-4.4 | 2 | 3.46 | 9.9 | 72.6 | 84.66 | 88.02 | 88.72 | 89.04
-4.3 | 2.1 | 3.66 | 11 | 76.46 | 85.34 | 88.42 | 89.12 | 89.43
-4.2 | 2.2 | 3.76 | 12.5 | 79.26 | 85.84 | 88.82 | 89.52 | 89.74
-4.1 | 2.3 | 3.96 | 14.6 | 81.22 | 86.54 | 89.28 | 89.86 | 90.13
-4 | 2.4 | 4.16 | 17.6 | 82.78 | 87.04 | 89.62 | 90.22 | 90.45
-3.9 | 2.5 | 4.46 | 22.2 | 84.08 | 87.56 | 90.02 | 90.56 | 90.93
-3.8 | 2.6 | 4.72 | 28.5 | 85.3 | 88.06 | 90.38 | 91.02 | 91.24
-3.7 | 2.7 | 5.02 | 36.1 | 86.32 | 88.74 | 90.82 | 91.32 | 91.63
-3.6 | 2.8 | 5.42 | 45 | 87.22 | 89.16 | 91.22 | 91.72 | 91.93
-3.5 | 3 | 5.88 | 54.1 | 87.94 | 89.66 | 91.52 | 92.12 | 92.25
-3.4 | 3.2 | 6.38 | 63 | 88.72 | 90.24 | 91.92 | 92.52 | 92.73
-3.3 | 3.4 | 7.14 | 70.6 | 89.34 | 90.66 | 92.32 | 92.76 | 93.04
-3.2 | 3.6 | 8.1 | 76.7 | 90.04 | 91.24 | 92.72 | 93.06 | 93.43
-3.1 | 3.9 | 9.22 | 81 | 90.64 | 91.76 | 93.12 | 93.46 | 93.73
-3 | 4.3 | 10.84 | 84.1 | 91.24 | 92.16 | 93.52 | 93.82 | 94.04
-2.9 | 4.7 | 13.04 | 86.3 | 91.76 | 92.66 | 93.86 | 94.16 | 94.43
-2.8 | 5.2 | 16.24 | 88.1 | 92.34 | 93.14 | 94.16 | 94.42 | 94.74
-2.7 | 5.9 | 20.98 | 89.5 | 92.84 | 93.46 | 94.56 | 94.92 | 95.12
-2.6 | 6.8 | 27.88 | 90.7 | 93.26 | 93.86 | 94.92 | 95.22 | 95.34
-2.5 | 8 | 36.32 | 91.7 | 93.84 | 94.36 | 95.18 | 95.52 | 95.73
-2.4 | 9.7 | 46 | 92.3 | 94.36 | 94.76 | 95.56 | 95.92 | 96.02
-2.3 | 12 | 56 | 93.1 | 94.76 | 95.24 | 95.82 | 96.16 | 96.24
-2.2 | 15.4 | 65.4 | 93.8 | 95.16 | 95.58 | 96.22 | 96.46 | 96.62
-2.1 | 20.3 | 73.9 | 94.4 | 95.56 | 95.88 | 96.56 | 96.72 | 96.83
-2 | 27.5 | 80.58 | 95 | 95.98 | 96.26 | 96.76 | 97.12 | 97.14
-1.9 | 36.4 | 85.36 | 95.5 | 96.36 | 96.64 | 97.06 | 97.36 | 97.52
-1.8 | 46.5 | 88.68 | 96 | 96.86 | 96.96 | 97.36 | 97.46 | 97.72
-1.7 | 56.8 | 91.04 | 96.5 | 97.24 | 97.28 | 97.56 | 97.82 | 97.92
-1.6 | 66.8 | 92.78 | 96.8 | 97.48 | 97.58 | 97.92 | 98.06 | 98.12
-1.5 | 75.7 | 94.16 | 97.2 | 97.78 | 97.78 | 98.16 | 98.26 | 98.33
-1.4 | 82.8 | 95.2 | 97.6 | 98.08 | 98.06 | 98.3 | 98.52 | 98.61
-1.3 | 87.9 | 96.04 | 98 | 98.38 | 98.46 | 98.56 | 98.76 | 98.72
-1.2 | 91.2 | 96.78 | 98.3 | 98.58 | 98.58 | 98.86 | 98.9 | 98.92
-1.1 | 93.6 | 97.38 | 98.7 | 98.78 | 98.88 | 98.96 | 99.06 | 99.12
-1 | 95.2 | 97.88 | 98.9 | 99 | 99.06 | 99.1 | 99.26 | 99.31
-0.9 | 96.4 | 98.36 | 99.1 | 99.18 | 99.28 | 99.26 | 99.36 | 99.41
-0.8 | 97.5 | 98.72 | 99.3 | 99.38 | 99.38 | 99.4 | 99.5 | 99.51
-0.7 | 98.2 | 99.02 | 99.5 | 99.48 | 99.5 | 99.56 | 99.6 | 99.61
-0.6 | 98.7 | 99.26 | 99.7 | 99.68 | 99.6 | 99.7 | 99.7 | 99.71
-0.5 | 99.2 | 99.46 | 99.8 | 99.78 | 99.7 | 99.76 | 99.76 | 99.81
-0.4 | 99.5 | 99.66 | 99.9 | 99.88 | 99.9 | 99.86 | 99.86 | 99.9
-0.3 | 99.8 | 99.76 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.91
-0.2 | 99.9 | 99.86 | 100 | 100 | 100 | 99.9 | 100 | 100
-0.1 | 100 | 99.96 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 99.94 | 100 | 100 | 100 | 100 | 99.94 | 100
0.3 | 99.8 | 99.84 | 99.9 | 99.92 | 99.92 | 99.94 | 99.9 | 99.99
0.4 | 99.4 | 99.74 | 99.9 | 99.9 | 99.9 | 99.9 | 99.84 | 99.89
0.5 | 99.1 | 99.54 | 99.8 | 99.8 | 99.8 | 99.84 | 99.8 | 99.79
0.6 | 98.7 | 99.4 | 99.6 | 99.7 | 99.7 | 99.74 | 99.7 | 99.69
0.7 | 98.2 | 99.14 | 99.5 | 99.6 | 99.6 | 99.64 | 99.6 | 99.6
0.8 | 97.5 | 98.84 | 99.4 | 99.52 | 99.42 | 99.44 | 99.44 | 99.58
0.9 | 96.4 | 98.48 | 99 | 99.22 | 99.24 | 99.3 | 99.34 | 99.39
1 | 95.1 | 98.08 | 98.9 | 98.94 | 99.02 | 99.14 | 99.24 | 99.28
1.1 | 93.5 | 97.58 | 98.7 | 98.72 | 98.82 | 99.04 | 99.14 | 99.09
1.2 | 91.2 | 97.02 | 98.3 | 98.52 | 98.62 | 98.84 | 98.94 | 98.98
1.3 | 87.8 | 96.36 | 97.9 | 98.24 | 98.42 | 98.64 | 98.64 | 98.78
1.4 | 82.7 | 95.5 | 97.6 | 97.94 | 98.22 | 98.44 | 98.44 | 98.58
1.5 | 75.5 | 94.4 | 97.3 | 97.64 | 97.84 | 98.08 | 98.34 | 98.38
1.6 | 66.5 | 93.18 | 96.8 | 97.32 | 97.62 | 97.84 | 98.04 | 98.18
1.7 | 56.5 | 91.46 | 96.5 | 97.04 | 97.32 | 97.64 | 97.78 | 97.97
1.8 | 46.2 | 89.12 | 96 | 96.66 | 97.04 | 97.44 | 97.54 | 97.68
1.9 | 36.2 | 85.86 | 95.6 | 96.24 | 96.72 | 97.18 | 97.34 | 97.47
2 | 27.3 | 81.24 | 94.9 | 95.92 | 96.34 | 96.88 | 96.94 | 97.17
2.1 | 20.1 | 74.78 | 94.4 | 95.44 | 95.92 | 96.54 | 96.74 | 96.87
2.2 | 15.3 | 66.3 | 93.8 | 95.06 | 95.64 | 96.24 | 96.48 | 96.57
2.3 | 11.9 | 56.8 | 93.1 | 94.64 | 95.16 | 95.94 | 96.14 | 96.27
2.4 | 9.6 | 46.68 | 92.5 | 94.24 | 94.74 | 95.64 | 95.94 | 95.97
2.5 | 8 | 37.04 | 91.6 | 93.74 | 94.44 | 95.38 | 95.54 | 95.67
2.6 | 6.8 | 28.46 | 90.5 | 93.16 | 93.94 | 94.98 | 95.28 | 95.37
2.7 | 5.9 | 21.54 | 89.4 | 92.72 | 93.56 | 94.64 | 94.84 | 95.06
2.8 | 5.2 | 16.54 | 88 | 92.16 | 93.06 | 94.28 | 94.48 | 94.67
2.9 | 4.7 | 13.26 | 86.3 | 91.66 | 92.72 | 93.88 | 94.18 | 94.37
3 | 4.3 | 11 | 84.3 | 91.14 | 92.16 | 93.52 | 93.84 | 94.07
3.1 | 3.9 | 9.34 | 81.2 | 90.56 | 91.74 | 93.14 | 93.54 | 93.77
3.2 | 3.6 | 8.1 | 76.9 | 89.94 | 91.34 | 92.78 | 93.18 | 93.46
3.3 | 3.4 | 7.16 | 70.9 | 89.18 | 90.76 | 92.38 | 92.78 | 93.06
3.4 | 3.2 | 6.46 | 63.3 | 88.66 | 90.26 | 91.98 | 92.48 | 92.66
3.5 | 3 | 5.88 | 54.5 | 87.88 | 89.74 | 91.68 | 92.18 | 92.27
3.6 | 2.8 | 5.48 | 45.3 | 87.06 | 89.24 | 91.28 | 91.68 | 91.96
3.7 | 2.7 | 5.08 | 36.5 | 86.18 | 88.72 | 90.88 | 91.28 | 91.57
3.8 | 2.6 | 4.68 | 28.6 | 85.12 | 88.16 | 90.44 | 90.94 | 91.27
3.9 | 2.5 | 4.44 | 22.3 | 84.02 | 87.58 | 90.04 | 90.58 | 90.95
4 | 2.4 | 4.24 | 17.7 | 82.74 | 87.06 | 89.68 | 90.18 | 90.46
4.1 | 2.3 | 4.04 | 14.6 | 81.18 | 86.56 | 89.28 | 89.88 | 90.07
4.2 | 2.2 | 3.84 | 12.5 | 79.12 | 85.88 | 88.94 | 89.54 | 89.77
4.3 | 2.1 | 3.64 | 11 | 76.3 | 85.34 | 88.48 | 89.18 | 89.45
4.4 | 2 | 3.54 | 9.8 | 72.42 | 84.76 | 88.08 | 88.68 | 88.97
4.5 | 1.9 | 3.34 | 8.9 | 67.1 | 83.98 | 87.68 | 88.32 | 88.66
4.6 | 1.9 | 3.2 | 8.2 | 60.24 | 83.36 | 87.28 | 87.98 | 88.27
4.7 | 1.8 | 3.14 | 7.6 | 52.32 | 82.56 | 86.88 | 87.64 | 87.96
4.8 | 1.8 | 3 | 7.2 | 44.04 | 81.78 | 86.48 | 87.12 | 87.56
4.9 | 1.7 | 2.9 | 6.7 | 35.88 | 80.9 | 86.04 | 86.84 | 87.16
5 | 1.7 | 2.8 | 6.4 | 28.68 | 80.08 | 85.58 | 86.42 | 86.77
5.1 | 1.6 | 2.7 | 6.1 | 22.74 | 78.92 | 85.18 | 86.08 | 86.46
5.2 | 1.6 | 2.64 | 5.8 | 18.48 | 77.82 | 84.68 | 85.62 | 86.06
5.3 | 1.5 | 2.54 | 5.6 | 15.56 | 76.26 | 84.34 | 85.22 | 85.67
5.4 | 1.5 | 2.44 | 5.4 | 13.5 | 74.4 | 83.88 | 84.88 | 85.36
5.5 | 1.4 | 2.4 | 5.1 | 12.04 | 71.98 | 83.48 | 84.48 | 84.96
5.6 | 1.4 | 2.3 | 5 | 11 | 68.58 | 83.18 | 84.14 | 84.56
5.7 | 1.4 | 2.24 | 4.8 | 10.08 | 63.72 | 82.68 | 83.78 | 84.17
5.8 | 1.3 | 2.2 | 4.6 | 9.36 | 57.64 | 82.24 | 83.28 | 83.86
5.9 | 1.3 | 2.1 | 4.5 | 8.84 | 50.44 | 81.82 | 82.94 | 83.47
6 | 1.3 | 2.04 | 4.3 | 8.34 | 42.96 | 81.48 | 82.64 | 83.15
6.1 | 1.2 | 2 | 4.2 | 7.94 | 35.52 | 81.08 | 82.28 | 82.67
6.2 | 1.2 | 1.94 | 4.1 | 7.62 | 28.74 | 80.62 | 81.84 | 82.36
6.3 | 1.2 | 1.9 | 4 | 7.24 | 23.3 | 80.18 | 81.48 | 81.96
6.4 | 1.1 | 1.84 | 3.8 | 6.94 | 19.08 | 79.78 | 81.12 | 81.57
6.5 | 1.1 | 1.8 | 3.7 | 6.74 | 16.16 | 79.38 | 80.74 | 81.26
6.6 | 1.1 | 1.8 | 3.6 | 6.52 | 14.26 | 79.02 | 80.28 | 80.87
6.7 | 1.1 | 1.7 | 3.6 | 6.32 | 12.82 | 78.48 | 79.98 | 80.56
6.8 |  | 1.7 | 3.4 | 6.12 | 11.72 | 78.08 | 79.62 | 80.16
6.9 |  | 1.6 | 3.4 | 5.92 | 10.88 | 77.72 | 79.24 | 79.77
7 |  | 1.6 | 3.3 | 5.72 | 10.26 | 77.22 | 78.88 | 79.46
7.1 |  | 1.6 | 3.2 | 5.52 | 9.64 | 76.88 | 78.48 | 79.07
7.2 |  | 1.5 | 3.1 | 5.32 | 9.14 | 76.48 | 78.18 | 78.76
7.3 |  | 1.5 | 3 | 5.22 | 8.74 | 76.08 | 77.78 | 78.37
7.4 |  |  | 2.9 | 5.1 | 8.42 | 75.68 | 77.44 | 78.06
7.5 |  |  | 2.9 | 4.92 | 8.12 | 75.28 | 77.08 | 77.67
7.6 |  |  | 2.8 | 4.82 | 7.82 | 74.78 | 76.74 | 77.37
7.7 |  |  | 2.7 | 4.7 | 7.52 | 74.38 | 76.38 | 77.07
7.8 |  |  | 2.7 | 4.6 | 7.32 | 74.04 | 75.92 | 76.75
7.9 |  |  | 2.6 | 4.42 | 7.02 | 73.68 | 75.54 | 76.27
8 |  |  | 2.6 | 4.32 | 6.82 | 73.18 | 75.28 | 75.97
8.1 |  |  | 2.5 | 4.22 | 6.62 | 72.74 | 74.88 | 75.66
8.2 |  |  | 2.4 | 4.12 | 6.42 | 72.32 | 74.52 | 75.27
8.3 |  |  | 2.4 | 4.02 | 6.3 | 71.94 | 74.14 | 74.97
8.4 |  |  | 2.3 | 3.92 | 6.12 | 71.58 | 73.84 | 74.66
8.5 |  |  | 2.3 | 3.82 | 6 | 71.18 | 73.48 | 74.27
8.6 |  |  |  | 3.72 | 5.82 | 70.72 | 73.14 | 73.97
8.7 |  |  |  | 3.62 | 5.62 | 70.28 | 72.78 | 73.67
8.8 |  |  |  | 3.6 | 5.52 | 69.88 | 72.48 | 73.36
8.9 |  |  |  | 3.5 | 5.42 | 69.48 | 72.08 | 72.97
9 |  |  |  | 3.4 | 5.22 | 69.02 | 71.78 | 72.67
9.1 |  |  |  | 3.32 | 5.12 | 68.58 | 71.44 | 72.36
9.2 |  |  |  | 3.3 | 5 | 68.24 | 71.08 | 71.97
9.3 |  |  |  | 3.2 | 4.9 | 67.82 | 70.68 | 71.67
9.4 |  |  |  | 3.1 | 4.8 | 67.38 | 70.44 | 71.37
9.5 |  |  |  | 3.1 | 4.7 | 66.88 | 70.14 | 71.07
9.6 |  |  |  | 3 | 4.6 | 66.48 | 69.74 | 70.77
9.7 |  |  |  | 2.92 | 4.5 | 66.08 | 69.38 | 70.46
9.8 |  |  |  |  | 4.32 | 65.58 | 69.14 | 70.08
9.9 |  |  |  |  | 4.22 | 65.22 | 68.78 | 69.86
10 |  |  |  |  | 4.2 | 64.78 | 68.44 | 69.47
10.1 |  |  |  |  | 4.1 | 64.28 | 68.04 | 69.17
10.2 |  |  |  |  | 4.02 | 63.82 | 67.78 | 68.87
10.3 |  |  |  |  | 3.9 | 63.32 | 67.44 | 68.57
10.4 |  |  |  |  | 3.8 | 62.88 | 67.14 | 68.27
10.5 |  |  |  |  | 3.72 | 62.32 | 66.84 | 67.97
10.6 |  |  |  |  | 3.62 | 61.82 | 66.48 | 67.67
10.7 |  |  |  |  | 3.6 | 61.28 | 66.14 | 67.37
10.8 |  |  |  |  | 3.52 | 60.72 | 65.88 | 67.07
10.9 |  |  |  |  | 3.5 | 60.12 | 65.54 | 66.77
11 | [5 empty] | 59.46 | 65.18 | 66.47
11.1 | [5 empty] | 58.66 | 64.84 | 66.17
11.2 | [5 empty] | 57.8 | 64.54 | 65.88
11.3 | [5 empty] | 56.8 | 64.24 | 65.66
11.4 | [5 empty] | 55.52 | 63.94 | 65.27
11.5 | [5 empty] | 53.86 | 63.64 | 64.98
11.6 | [5 empty] | 51.62 | 63.34 | 64.77
11.7 | [5 empty] | 48.42 | 62.98 | 64.47
11.8 | [5 empty] | 44.38 | 62.64 | 64.18
11.9 | [5 empty] | 39.6 | 62.38 | 63.97
12 | [5 empty] | 34.38 | 62.08 | 63.67
12.1 | [5 empty] | 29.2 | 61.84 | 63.37
12.2 | [5 empty] | 24.52 | 61.54 | 63.08
12.3 | [5 empty] | 20.52 | 61.18 | 62.87
12.4 | [5 empty] | 17.52 | 60.88 | 62.57
12.5 | [5 empty] | 15.36 | 60.48 | 62.27
12.6 | [5 empty] | 13.84 | 60.24 | 61.97
12.7 | [5 empty] | 12.8 | 59.94 | 61.68
12.8 | [5 empty] | 11.96 | 59.62 | 61.47
12.9 | [5 empty] | 11.32 | 59.38 | 61.18
13 | [5 empty] | 10.82 | 59.04 | 60.96
13.1 | [5 empty] | 10.38 | 58.74 | 60.58
13.2 | [5 empty] | 9.98 | 58.44 | 60.37
13.3 | [5 empty] | 9.68 | 58.08 | 60.08
13.4 | [5 empty] | 9.38 | 57.84 | 59.87
13.5 | [5 empty] | 9.14 | 57.58 | 59.57
13.6 | [5 empty] | 8.88 | 57.24 | 59.28
13.7 | [5 empty] | 8.64 | 56.98 | 59.08
13.8 | [5 empty] | 8.44 | 56.58 | 58.87
13.9 | [5 empty] | 8.24 | 56.34 | 58.58
14 | [5 empty] | 8.04 | 56.04 | 58.37
14.1 | [5 empty] | 7.84 | 55.74 | 58.07
14.2 | [5 empty] | 7.74 | 55.48 | 57.78
14.3 | [5 empty] | 7.54 | 55.18 | 57.57
14.4 | [5 empty] | 7.44 | 54.94 | 57.28
14.5 | [5 empty] | 7.24 | 54.58 | 57.07
14.6 | [5 empty] | 7.14 | 54.24 | 56.78
14.7 | [5 empty] | 6.94 | 53.98 | 56.57
14.8 | [5 empty] | 6.84 | 53.74 | 56.28
14.9 | [5 empty] | 6.7 | 53.38 | 56.08
15 | [5 empty] | 6.54 | 53.08 | 55.87
15.1 | [5 empty] | 6.44 | 52.74 | 55.58
15.2 | [5 empty] | 6.3 | 52.44 | 55.37
15.3 | [5 empty] | 6.2 | 52.18 | 55.08
15.4 | [5 empty] | 6.1 | 51.84 | 54.88
15.5 | [5 empty] | 6 | 51.54 | 54.67
15.6 | [5 empty] | 5.84 | 51.24 | 54.38
15.7 | [5 empty] | 5.74 | 50.94 | 54.17
15.8 | [5 empty] | 5.64 | 50.58 | 53.88
15.9 | [5 empty] | 5.54 | 50.28 | 53.67
16 | [5 empty] | 5.4 | 49.94 | 53.38
16.1 | [5 empty] | 5.34 | 49.64 | 53.18
16.2 | [5 empty] | 5.24 | 49.28 | 52.98
16.3 | [5 empty] | 5.2 | 48.98 | 52.78
16.4 | [5 empty] | 5.1 | 48.58 | 52.57
16.5 | [5 empty] | 5 | 48.24 | 52.28
16.6 | [5 empty] | 4.9 | 47.88 | 52.08
16.7 | [5 empty] | 4.84 | 47.48 | 51.87
16.8 | [5 empty] | 4.74 | 47.04 | 51.58
16.9 | [5 empty] | 4.7 | 46.58 | 51.38
17 | [6 empty] | 46.12 | 51.17
17.1 | [6 empty] | 45.52 | 50.88
17.2 | [6 empty] | 44.92 | 50.67
17.3 | [6 empty] | 44.1 | 50.38
17.4 | [6 empty] | 43.2 | 50.18
17.5 | [6 empty] | 41.88 | 49.98
17.6 | [6 empty] | 40.2 | 49.78
17.7 | [6 empty] | 37.92 | 49.58
17.8 | [6 empty] | 34.94 | 49.37
17.9 | [6 empty] | 31.46 | 49.08
18 | [6 empty] | 27.72 | 48.87
18.1 | [6 empty] | 23.92 | 48.58
18.2 | [6 empty] | 20.44 | 48.38
18.3 | [6 empty] | 17.46 | 48.18
18.4 | [6 empty] | 15.1 | 47.98
18.5 | [6 empty] | 13.42 | 47.77
18.6 | [6 empty] | 12.3 | 47.48
18.7 | [6 empty] | 11.46 | 47.28
18.8 | [6 empty] | 10.82 | 47.08
18.9 | [6 empty] | 10.28 | 46.88
19 | [6 empty] | 9.88 | 46.67
19.1 | [6 empty] | 9.54 | 46.38
19.2 | [6 empty] | 9.24 | 46.18
19.3 | [6 empty] | 8.94 | 45.97
19.4 | [6 empty] | 8.74 | 45.68
19.5 | [6 empty] | 8.54 | 45.48
19.6 | [6 empty] | 8.34 | 45.28
19.7 | [6 empty] | 8.14 | 45.08
19.8 | [6 empty] | 7.94 | 44.88
19.9 | [6 empty] | 7.84 | 44.67
20 | [6 empty] | 7.64 | 44.38
20.1 | [6 empty] | 7.5 | 44.18
20.2 | [6 empty] | 7.34 | 43.98
20.3 | [6 empty] | 7.24 | 43.78
20.4 | [6 empty] | 7.1 | 43.57
20.5 | [6 empty] | 6.94 | 43.28
20.6 | [6 empty] | 6.84 | 43.08
20.7 | [6 empty] | 6.7 | 42.87
20.8 | [6 empty] | 6.6 | 42.58
20.9 | [6 empty] | 6.5 | 42.38
21 | [6 empty] | 6.4 | 42.18
21.1 | [6 empty] | 6.24 | 41.97
21.2 | [6 empty] | 6.14 | 41.68
21.3 | [6 empty] | 6.04 | 41.48
21.4 | [6 empty] | 6 | 41.28
21.5 | [6 empty] | 5.84 | 41.07
21.6 | [6 empty] | 5.74 | 40.78
21.7 | [6 empty] | 5.7 | 40.57
21.8 | [6 empty] | 5.6 | 40.28
21.9 | [6 empty] | 5.5 | 40.07
22 | [6 empty] | 5.4 | 39.78
22.1 | [6 empty] | 5.3 | 39.58
22.2 | [6 empty] | 5.24 | 39.37
22.3 | [6 empty] | 5.14 | 39.07
22.4 | [6 empty] | 5.1 | 38.77
22.5 | [6 empty] | 5 | 38.48
22.6 | [6 empty] | 4.94 | 38.27
22.7 | [6 empty] | 4.84 | 37.96
22.8 | [6 empty] | 4.8 | 37.57
22.9 | [7 empty] | 37.26
23 | [7 empty] | 36.86
23.1 | [7 empty] | 36.45
23.2 | [7 empty] | 35.94
23.3 | [7 empty] | 35.33
23.4 | [7 empty] | 34.61
23.5 | [7 empty] | 33.67
23.6 | [7 empty] | 32.31
23.7 | [7 empty] | 30.38
23.8 | [7 empty] | 28.12
23.9 | [7 empty] | 25.31
24 | [7 empty] | 22.41
24.1 | [7 empty] | 19.53
24.2 | [7 empty] | 16.86
24.3 | [7 empty] | 14.52
24.4 | [7 empty] | 12.76
24.5 | [7 empty] | 11.41
24.6 | [7 empty] | 10.53
24.7 | [7 empty] | 9.85
24.8 | [7 empty] | 9.36
24.9 | [7 empty] | 8.96
25 | [7 empty] | 8.57
25.1 | [7 empty] | 8.28
25.2 | [7 empty] | 8.08
25.3 | [7 empty] | 7.88
25.4 | [7 empty] | 7.68
25.5 | [7 empty] | 7.49
25.6 | [7 empty] | 7.38
25.7 | [7 empty] | 7.19
25.8 | [7 empty] | 7.08
25.9 | [7 empty] | 6.89
26 | [7 empty] | 6.79
26.1 | [7 empty] | 6.69
26.2 | [7 empty] | 6.59
26.3 | [7 empty] | 6.48
26.4 | [7 empty] | 6.29
26.5 | [7 empty] | 6.19
26.6 | [7 empty] | 6.09
26.7 | [7 empty] | 5.99
26.8 | [7 empty] | 5.89
26.9 | [7 empty] | 5.79
27 | [7 empty] | 5.7
27.1 | [7 empty] | 5.69
27.2 | [7 empty] | 5.59
27.3 | [7 empty] | 5.49
27.4 | [7 empty] | 5.39
27.5 | [7 empty] | 5.29
27.6 | [7 empty] | 5.19
27.7 | [7 empty] | 5.1
27.8 | [7 empty] | 5.09
27.9 | [7 empty] | 4.99
28 | [7 empty] | 4.89
28.1 | [7 empty] | 4.8
28.2 | [7 empty] | 4.79
28.3 | [7 empty] | 4.7
28.4 | [7 empty] | 4.69
28.5 | [7 empty] | 4.59
28.6 | [7 empty] | 4.5
28.7 | [7 empty] | 4.49
28.8 | [7 empty] | 4.39
28.9 | [7 empty] | 4.3
29 | [7 empty] | 4.29
29.1 | [7 empty] | 4.2
29.2 | [7 empty] | 4.19
29.3 | [7 empty] | 4.1
29.4 | [7 empty] | 4.09
29.5 | [7 empty] | 4
29.6 | [7 empty] | 3.99
29.7 | [7 empty] | 3.89

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 6.61
-29.7 | [7 empty] | 6.71
-29.6 | [7 empty] | 6.81
-29.5 | [7 empty] | 6.91
-29.4 | [7 empty] | 7.01
-29.3 | [7 empty] | 7.11
-29.2 | [7 empty] | 7.2
-29.1 | [7 empty] | 7.22
-29 | [7 empty] | 7.4
-28.9 | [7 empty] | 7.42
-28.8 | [7 empty] | 7.6
-28.7 | [7 empty] | 7.62
-28.6 | [7 empty] | 7.81
-28.5 | [7 empty] | 7.91
-28.4 | [7 empty] | 8.01
-28.3 | [7 empty] | 8.11
-28.2 | [7 empty] | 8.21
-28.1 | [7 empty] | 8.31
-28 | [7 empty] | 8.42
-27.9 | [7 empty] | 8.61
-27.8 | [7 empty] | 8.72
-27.7 | [7 empty] | 8.91
-27.6 | [7 empty] | 9.02
-27.5 | [7 empty] | 9.22
-27.4 | [6 empty] | 4.4 | 9.42
-27.3 | [6 empty] | 4.48 | 9.62
-27.2 | [6 empty] | 4.5 | 9.82
-27.1 | [6 empty] | 4.58 | 10.03
-27 | [6 empty] | 4.6 | 10.34
-26.9 | [6 empty] | 4.68 | 10.74
-26.8 | [6 empty] | 4.7 | 11.15
-26.7 | [6 empty] | 4.78 | 11.67
-26.6 | [6 empty] | 4.88 | 12.38
-26.5 | [6 empty] | 4.9 | 13.23
-26.4 | [6 empty] | 4.98 | 14.56
-26.3 | [6 empty] | 5 | 16.21
-26.2 | [6 empty] | 5.08 | 18.34
-26.1 | [6 empty] | 5.18 | 20.75
-26 | [6 empty] | 5.28 | 23.26
-25.9 | [6 empty] | 5.3 | 25.84
-25.8 | [6 empty] | 5.38 | 28.2
-25.7 | [6 empty] | 5.4 | 30.17
-25.6 | [6 empty] | 5.48 | 31.82
-25.5 | [6 empty] | 5.58 | 33
-25.4 | [6 empty] | 5.6 | 33.97
-25.3 | [6 empty] | 5.68 | 34.65
-25.2 | [6 empty] | 5.78 | 35.15
-25.1 | [6 empty] | 5.88 | 35.64
-25 | [6 empty] | 5.9 | 36.04
-24.9 | [6 empty] | 5.98 | 36.44
-24.8 | [6 empty] | 6.08 | 36.83
-24.7 | [6 empty] | 6.18 | 37.13
-24.6 | [6 empty] | 6.2 | 37.43
-24.5 | [6 empty] | 6.28 | 37.72
-24.4 | [6 empty] | 6.38 | 37.93
-24.3 | [6 empty] | 6.48 | 38.22
-24.2 | [6 empty] | 6.58 | 38.43
-24.1 | [6 empty] | 6.6 | 38.72
-24 | [6 empty] | 6.68 | 38.93
-23.9 | [6 empty] | 6.78 | 39.22
-23.8 | [6 empty] | 6.88 | 39.42
-23.7 | [6 empty] | 6.98 | 39.62
-23.6 | [6 empty] | 7.08 | 39.83
-23.5 | [6 empty] | 7.18 | 40.12
-23.4 | [6 empty] | 7.28 | 40.32
-23.3 | [6 empty] | 7.38 | 40.53
-23.2 | [6 empty] | 7.48 | 40.82
-23.1 | [6 empty] | 7.58 | 41.02
-23 | [6 empty] | 7.68 | 41.22
-22.9 | [6 empty] | 7.78 | 41.42
-22.8 | [6 empty] | 7.96 | 41.62
-22.7 | [6 empty] | 8.08 | 41.82
-22.6 | [6 empty] | 8.18 | 42.03
-22.5 | [6 empty] | 8.28 | 42.32
-22.4 | [6 empty] | 8.38 | 42.52
-22.3 | [6 empty] | 8.56 | 42.72
-22.2 | [6 empty] | 8.68 | 42.93
-22.1 | [6 empty] | 8.78 | 43.22
-22 | [6 empty] | 8.96 | 43.42
-21.9 | [6 empty] | 9.08 | 43.62
-21.8 | [6 empty] | 9.18 | 43.82
-21.7 | [6 empty] | 9.36 | 44.02
-21.6 | [6 empty] | 9.48 | 44.23
-21.5 | [6 empty] | 9.66 | 44.52
-21.4 | [6 empty] | 9.86 | 44.72
-21.3 | [6 empty] | 9.98 | 44.92
-21.2 | [6 empty] | 10.16 | 45.12
-21.1 | [6 empty] | 10.36 | 45.32
-21 | [6 empty] | 10.56 | 45.52
-20.9 | [6 empty] | 10.76 | 45.72
-20.8 | [6 empty] | 11.04 | 45.93
-20.7 | [6 empty] | 11.34 | 46.22
-20.6 | [6 empty] | 11.64 | 46.42
-20.5 | [6 empty] | 12.02 | 46.62
-20.4 | [6 empty] | 12.5 | 46.82
-20.3 | [6 empty] | 13 | 47.02
-20.2 | [6 empty] | 13.66 | 47.22
-20.1 | [6 empty] | 14.6 | 47.43
-20 | [6 empty] | 15.84 | 47.72
-19.9 | [6 empty] | 17.54 | 47.92
-19.8 | [6 empty] | 19.74 | 48.12
-19.7 | [6 empty] | 22.6 | 48.32
-19.6 | [6 empty] | 25.76 | 48.52
-19.5 | [6 empty] | 29.12 | 48.72
-19.4 | [6 empty] | 32.44 | 48.93
-19.3 | [6 empty] | 35.42 | 49.22
-19.2 | [6 empty] | 38 | 49.42
-19.1 | [6 empty] | 39.94 | 49.62
-19 | [6 empty] | 41.5 | 49.82
-18.9 | [6 empty] | 42.68 | 50.02
-18.8 | [6 empty] | 43.54 | 50.23
-18.7 | [6 empty] | 44.18 | 50.52
-18.6 | [6 empty] | 44.86 | 50.72
-18.5 | [6 empty] | 45.32 | 50.92
-18.4 | [6 empty] | 45.8 | 51.12
-18.3 | [6 empty] | 46.22 | 51.33
-18.2 | [6 empty] | 46.62 | 51.62
-18.1 | [6 empty] | 47.02 | 51.82
-18 | [5 empty] | 6.22 | 47.42 | 52.02
-17.9 | [5 empty] | 6.32 | 47.74 | 52.23
-17.8 | [5 empty] | 6.42 | 48.12 | 52.52
-17.7 | [5 empty] | 6.5 | 48.44 | 52.72
-17.6 | [5 empty] | 6.6 | 48.74 | 52.92
-17.5 | [5 empty] | 6.72 | 49.04 | 53.13
-17.4 | [5 empty] | 6.82 | 49.42 | 53.42
-17.3 | [5 empty] | 6.92 | 49.74 | 53.62
-17.2 | [5 empty] | 7.02 | 50.04 | 53.82
-17.1 | [5 empty] | 7.12 | 50.34 | 54.03
-17 | [5 empty] | 7.22 | 50.64 | 54.32
-16.9 | [5 empty] | 7.32 | 50.86 | 54.53
-16.8 | [5 empty] | 7.5 | 51.14 | 54.82
-16.7 | [5 empty] | 7.6 | 51.44 | 55.02
-16.6 | [5 empty] | 7.72 | 51.82 | 55.22
-16.5 | [5 empty] | 7.82 | 52.14 | 55.42
-16.4 | [5 empty] | 7.92 | 52.36 | 55.63
-16.3 | [5 empty] | 8.1 | 52.64 | 55.92
-16.2 | [5 empty] | 8.22 | 52.94 | 56.13
-16.1 | [5 empty] | 8.32 | 53.24 | 56.42
-16 | [5 empty] | 8.52 | 53.54 | 56.62
-15.9 | [5 empty] | 8.62 | 53.84 | 56.83
-15.8 | [5 empty] | 8.82 | 54.06 | 57.12
-15.7 | [5 empty] | 9 | 54.42 | 57.32
-15.6 | [5 empty] | 9.12 | 54.66 | 57.53
-15.5 | [5 empty] | 9.3 | 54.94 | 57.82
-15.4 | [5 empty] | 9.42 | 55.24 | 58.02
-15.3 | [5 empty] | 9.6 | 55.54 | 58.23
-15.2 | [5 empty] | 9.8 | 55.84 | 58.52
-15.1 | [5 empty] | 9.92 | 56.14 | 58.73
-15 | [5 empty] | 10.12 | 56.44 | 59.02
-14.9 | [5 empty] | 10.32 | 56.74 | 59.22
-14.8 | [5 empty] | 10.52 | 56.96 | 59.43
-14.7 | [5 empty] | 10.74 | 57.32 | 59.72
-14.6 | [5 empty] | 11.02 | 57.56 | 59.93
-14.5 | [5 empty] | 11.32 | 57.84 | 60.22
-14.4 | [5 empty] | 11.52 | 58.14 | 60.43
-14.3 | [5 empty] | 11.84 | 58.44 | 60.72
-14.2 | [5 empty] | 12.14 | 58.74 | 60.93
-14.1 | [5 empty] | 12.54 | 58.96 | 61.23
-14 | [5 empty] | 13.06 | 59.24 | 61.52
-13.9 | [5 empty] | 13.56 | 59.54 | 61.72
-13.8 | [5 empty] | 14.26 | 59.84 | 61.93
-13.7 | [5 empty] | 15.08 | 60.14 | 62.22
-13.6 | [5 empty] | 16.04 | 60.44 | 62.43
-13.5 | [5 empty] | 17.58 | 60.74 | 62.73
-13.4 | [5 empty] | 19.54 | 61.04 | 63.02
-13.3 | [5 empty] | 22.34 | 61.34 | 63.23
-13.2 | [5 empty] | 26 | 61.64 | 63.53
-13.1 | [5 empty] | 30.16 | 61.94 | 63.82
-13 | [5 empty] | 34.86 | 62.24 | 64.02
-12.9 | [5 empty] | 39.56 | 62.46 | 64.23
-12.8 | [5 empty] | 43.9 | 62.74 | 64.53
-12.7 | [5 empty] | 47.62 | 63.04 | 64.83
-12.6 | [5 empty] | 50.56 | 63.34 | 65.12
-12.5 | [5 empty] | 52.9 | 63.72 | 65.33
-12.4 | [5 empty] | 54.54 | 63.96 | 65.62
-12.3 | [5 empty] | 55.82 | 64.24 | 65.83
-12.2 | [5 empty] | 56.86 | 64.54 | 66.13
-12.1 | [5 empty] | 57.66 | 64.84 | 66.43
-12 | [5 empty] | 58.46 | 65.22 | 66.73
-11.9 | [5 empty] | 59.06 | 65.54 | 67.03
-11.8 | [5 empty] | 59.66 | 65.84 | 67.32
-11.7 | [5 empty] | 60.26 | 66.14 | 67.53
-11.6 | [5 empty] | 60.76 | 66.36 | 67.82
-11.5 |  |  |  |  | 4.7 | 61.26 | 66.72 | 68.04
-11.4 |  |  |  |  | 4.74 | 61.74 | 67.04 | 68.43
-11.3 |  |  |  |  | 4.84 | 62.24 | 67.26 | 68.73
-11.2 |  |  |  |  | 4.94 | 62.66 | 67.62 | 69.02
-11.1 |  |  |  |  | 5.04 | 63.14 | 67.94 | 69.23
-11 |  |  |  |  | 5.14 | 63.54 | 68.32 | 69.53
-10.9 |  |  |  |  | 5.3 | 64.04 | 68.56 | 69.82
-10.8 |  |  |  |  | 5.34 | 64.44 | 68.84 | 70.04
-10.7 |  |  |  |  | 5.44 | 64.84 | 69.14 | 70.43
-10.6 |  |  |  |  | 5.54 | 65.34 | 69.52 | 70.73
-10.5 |  |  |  |  | 5.6 | 65.66 | 69.76 | 71.02
-10.4 |  |  |  |  | 5.8 | 66.14 | 70.04 | 71.23
-10.3 |  |  |  |  | 5.9 | 66.46 | 70.42 | 71.53
-10.2 |  |  |  | 4 | 5.94 | 66.94 | 70.74 | 71.83
-10.1 |  |  |  | 4.1 | 6.14 | 67.26 | 71.04 | 72.13
-10 |  |  |  | 4.12 | 6.24 | 67.74 | 71.3182 | 72.43
-9.9 |  |  |  | 4.22 | 6.34 | 68.14 | 71.61 | 72.73
-9.8 |  |  |  | 4.32 | 6.5 | 68.54 | 71.91 | 73.03
-9.7 |  |  |  | 4.4 | 6.64 | 68.86 | 72.28 | 73.33
-9.6 |  |  |  | 4.5 | 6.74 | 69.34 | 72.61 | 73.63
-9.5 |  |  |  | 4.6 | 6.94 | 69.74 | 72.91 | 73.93
-9.4 |  |  |  | 4.7 | 7.04 | 70.12 | 73.21 | 74.23
-9.3 |  |  |  | 4.8 | 7.24 | 70.54 | 73.58 | 74.53
-9.2 |  |  |  | 4.82 | 7.34 | 70.94 | 73.91 | 74.83
-9.1 |  |  |  | 4.92 | 7.54 | 71.32 | 74.21 | 75.14
-9 |  |  |  | 5.1 | 7.74 | 71.74 | 74.51 | 75.53
-8.9 |  |  |  | 5.2 | 7.84 | 72.12 | 74.88 | 75.82
-8.8 |  |  | 3.1 | 5.3 | 8.04 | 72.44 | 75.21 | 76.04
-8.7 |  |  | 3.2 | 5.4 | 8.24 | 72.94 | 75.51 | 76.43
-8.6 |  |  | 3.28 | 5.5 | 8.48 | 73.24 | 75.88 | 76.73
-8.5 |  |  | 3.3 | 5.6 | 8.74 | 73.66 | 76.21 | 77.03
-8.4 |  |  | 3.4 | 5.72 | 8.94 | 74.22 | 76.51 | 77.33
-8.3 |  |  | 3.48 | 5.9 | 9.14 | 74.46 | 76.81 | 77.63
-8.2 |  |  | 3.58 | 6.02 | 9.38 | 74.84 | 77.18 | 77.94
-8.1 |  |  | 3.6 | 6.2 | 9.64 | 75.32 | 77.58 | 78.33
-8 |  |  | 3.7 | 6.32 | 9.98 | 75.74 | 77.91 | 78.63
-7.9 |  |  | 3.8 | 6.52 | 10.28 | 76.04 | 78.21 | 78.93
-7.8 |  |  | 3.9 | 6.62 | 10.68 | 76.46 | 78.51 | 79.24
-7.7 |  |  | 3.98 | 6.72 | 11.12 | 76.82 | 78.88 | 79.63
-7.6 |  |  | 4 | 6.92 | 11.62 | 77.16 | 79.21 | 79.94
-7.5 |  |  | 4.18 | 7.12 | 12.22 | 77.56 | 79.51 | 80.32
-7.4 |  | 2.1 | 4.2 | 7.3 | 12.86 | 78.02 | 79.88 | 80.53
-7.3 |  | 2.16 | 4.38 | 7.42 | 13.8 | 78.34 | 80.21 | 80.83
-7.2 |  | 2.2 | 4.48 | 7.62 | 14.94 | 78.72 | 80.58 | 81.14
-7.1 |  | 2.26 | 4.58 | 7.82 | 16.42 | 79.04 | 80.91 | 81.53
-7 |  | 2.3 | 4.68 | 8.14 | 18.38 | 79.46 | 81.28 | 81.83
-6.9 |  | 2.36 | 4.78 | 8.34 | 21.22 | 79.94 | 81.68 | 82.14
-6.8 | 1.5 | 2.4 | 4.9 | 8.62 | 25.22 | 80.34 | 82.01 | 82.53
-6.7 | 1.5 | 2.46 | 5 | 8.92 | 30.42 | 80.64 | 82.31 | 82.84
-6.6 | 1.5 | 2.5 | 5.18 | 9.24 | 36.72 | 81.04 | 82.61 | 83.23
-6.5 | 1.6 | 2.6 | 5.28 | 9.56 | 43.56 | 81.54 | 82.98 | 83.53
-6.4 | 1.6 | 2.6 | 5.4 | 10.04 | 50.5 | 81.92 | 83.38 | 83.84
-6.3 | 1.6 | 2.7 | 5.58 | 10.56 | 57.18 | 82.34 | 83.71 | 84.23
-6.2 | 1.7 | 2.76 | 5.68 | 11.16 | 62.98 | 82.64 | 84.08 | 84.54
-6.1 | 1.7 | 2.8 | 5.88 | 11.88 | 67.56 | 83.04 | 84.41 | 84.92
-6 | 1.72 | 2.9 | 6.06 | 12.8 | 70.96 | 83.52 | 84.78 | 85.14
-5.9 | 1.8 | 2.96 | 6.28 | 13.92 | 73.44 | 83.84 | 85.11 | 85.53
-5.8 | 1.82 | 3.06 | 6.48 | 15.38 | 75.32 | 84.24 | 85.48 | 85.84
-5.7 | 1.9 | 3.1 | 6.66 | 17.36 | 76.84 | 84.64 | 85.81 | 86.23
-5.6 | 1.9 | 3.2 | 6.86 | 20.26 | 78.1 | 84.94 | 86.18 | 86.54
-5.5 | 2 | 3.26 | 7.08 | 24.38 | 79.06 | 85.34 | 86.51 | 86.94
-5.4 | 2 | 3.36 | 7.38 | 29.84 | 79.96 | 85.72 | 86.88 | 87.33
-5.3 | 2.1 | 3.46 | 7.76 | 36.62 | 80.86 | 86.14 | 87.21 | 87.63
-5.2 | 2.12 | 3.56 | 8.14 | 44.1 | 81.56 | 86.46 | 87.58 | 87.94
-5.1 | 2.2 | 3.66 | 8.56 | 51.98 | 82.36 | 86.94 | 87.91 | 88.33
-5 | 2.3 | 3.8 | 9.04 | 59.48 | 83.08 | 87.34 | 88.21 | 88.64
-4.9 | 2.3 | 3.9 | 9.64 | 65.98 | 83.56 | 87.66 | 88.58 | 89.03
-4.8 | 2.4 | 4.06 | 10.42 | 71.24 | 84.22 | 88.12 | 88.98 | 89.33
-4.7 | 2.5 | 4.16 | 11.4 | 75.22 | 84.88 | 88.44 | 89.31 | 89.63
-4.6 | 2.52 | 4.36 | 12.48 | 78.12 | 85.48 | 88.84 | 89.61 | 89.94
-4.5 | 2.6 | 4.46 | 14.12 | 80.16 | 86.02 | 89.14 | 89.98 | 90.33
-4.4 | 2.7 | 4.66 | 16.26 | 81.66 | 86.52 | 89.42 | 90.38 | 90.63
-4.3 | 2.8 | 4.86 | 19.34 | 83.04 | 87.12 | 89.76 | 90.71 | 90.93
-4.2 | 2.9 | 5.06 | 23.7 | 84.22 | 87.58 | 90.34 | 91.01 | 91.24
-4.1 | 3 | 5.36 | 29.66 | 85.28 | 88.18 | 90.64 | 91.38 | 91.63
-4 | 3.12 | 5.62 | 36.9 | 86.08 | 88.58 | 90.96 | 91.64 | 91.93
-3.9 | 3.22 | 6.02 | 45.08 | 86.88 | 89.08 | 91.34 | 91.98 | 92.23
-3.8 | 3.34 | 6.42 | 53.84 | 87.68 | 89.52 | 91.74 | 92.31 | 92.54
-3.7 | 3.52 | 6.88 | 62.02 | 88.38 | 90.08 | 92.14 | 92.61 | 92.93
-3.6 | 3.72 | 7.54 | 69.32 | 88.96 | 90.58 | 92.52 | 92.98 | 93.23
-3.5 | 3.92 | 8.24 | 75.34 | 89.54 | 91.02 | 92.82 | 93.38 | 93.53
-3.4 | 4.22 | 9.26 | 79.72 | 90.08 | 91.48 | 93.14 | 93.64 | 93.84
-3.3 | 4.44 | 10.52 | 82.9 | 90.74 | 91.82 | 93.52 | 93.91 | 94.22
-3.2 | 4.84 | 12.2 | 85.2 | 91.26 | 92.38 | 93.82 | 94.28 | 94.43
-3.1 | 5.24 | 14.58 | 86.86 | 91.68 | 92.78 | 94.14 | 94.61 | 94.73
-3 | 5.76 | 17.9 | 88.26 | 92.22 | 93.08 | 94.44 | 94.91 | 95.03
-2.9 | 6.38 | 22.78 | 89.6 | 92.66 | 93.58 | 94.74 | 95.14 | 95.33
-2.8 | 7.3 | 29.32 | 90.6 | 93.12 | 93.94 | 95.12 | 95.48 | 95.63
-2.7 | 8.32 | 37.34 | 91.52 | 93.44 | 94.32 | 95.44 | 95.81 | 95.92
-2.6 | 9.78 | 46.58 | 92.32 | 93.96 | 94.74 | 95.72 | 96.11 | 96.13
-2.5 | 11.78 | 56.08 | 92.94 | 94.44 | 95.04 | 95.94 | 96.34 | 96.43
-2.4 | 14.74 | 64.98 | 93.54 | 94.84 | 95.44 | 96.24 | 96.61 | 96.72
-2.3 | 18.94 | 73.02 | 94.14 | 95.24 | 95.68 | 96.44 | 96.91 | 96.92
-2.2 | 24.82 | 79.54 | 94.66 | 95.64 | 96.08 | 96.82 | 97.14 | 97.12
-2.1 | 32.5 | 84.36 | 95.22 | 95.94 | 96.38 | 97.04 | 97.41 | 97.33
-2 | 41.86 | 87.74 | 95.66 | 96.34 | 96.68 | 97.34 | 97.57 | 97.62
-1.9 | 51.72 | 90.26 | 96.18 | 96.64 | 96.98 | 97.62 | 97.81 | 97.83
-1.8 | 61.76 | 92.08 | 96.46 | 97.04 | 97.28 | 97.82 | 98.04 | 98.12
-1.7 | 71.02 | 93.36 | 96.94 | 97.42 | 97.58 | 98.02 | 98.24 | 98.31
-1.6 | 78.54 | 94.6 | 97.26 | 97.62 | 97.9 | 98.22 | 98.44 | 98.42
-1.5 | 84.4 | 95.48 | 97.58 | 97.92 | 98.08 | 98.34 | 98.64 | 98.62
-1.4 | 88.8 | 96.24 | 97.96 | 98.12 | 98.24 | 98.72 | 98.77 | 98.81
-1.3 | 91.62 | 96.88 | 98.08 | 98.44 | 98.54 | 98.9 | 98.94 | 98.92
-1.2 | 93.66 | 97.48 | 98.46 | 98.7 | 98.7 | 99.02 | 99.14 | 99.12
-1.1 | 95.32 | 97.92 | 98.68 | 98.92 | 98.94 | 99.22 | 99.27 | 99.3
-1 | 96.4 | 98.32 | 98.88 | 99.1 | 99.14 | 99.32 | 99.37 | 99.32
-0.9 | 97.38 | 98.66 | 99.08 | 99.22 | 99.3 | 99.52 | 99.54 | 99.51
-0.8 | 98.06 | 98.96 | 99.28 | 99.32 | 99.4 | 99.6 | 99.6 | 99.61
-0.7 | 98.62 | 99.26 | 99.46 | 99.52 | 99.5 | 99.7 | 99.67 | 99.71
-0.6 | 98.96 | 99.46 | 99.6 | 99.7 | 99.6 | 99.8 | 99.77 | 99.8
-0.5 | 99.42 | 99.56 | 99.7 | 99.8 | 99.74 | 99.9 | 99.87 | 99.81
-0.4 | 99.7 | 99.76 | 99.88 | 99.9 | 99.8 | 99.9 | 99.9 | 99.91
-0.3 | 99.82 | 99.86 | 99.9 | 99.92 | 99.9 | 100 | 99.9 | 100
-0.2 | 99.9 | 99.9 | 100 | 100 | 99.94 | 100 | 99.97 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 99.98 | 99.96 | 100 | 100 | 100
0.2 | 99.9 | 100 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100
0.3 | 99.8 | 99.94 | 99.9 | 99.9 | 99.9 | 99.98 | 99.93 | 99.99
0.4 | 99.68 | 99.84 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9
0.5 | 99.36 | 99.64 | 99.8 | 99.78 | 99.8 | 99.88 | 99.9 | 99.89
0.6 | 98.96 | 99.54 | 99.62 | 99.68 | 99.66 | 99.78 | 99.83 | 99.79
0.7 | 98.56 | 99.34 | 99.42 | 99.5 | 99.56 | 99.7 | 99.73 | 99.69
0.8 | 97.96 | 99.08 | 99.22 | 99.38 | 99.36 | 99.6 | 99.63 | 99.59
0.9 | 97.4 | 98.78 | 99.12 | 99.18 | 99.26 | 99.48 | 99.53 | 99.49
1 | 96.4 | 98.38 | 98.92 | 99.08 | 99.1 | 99.36 | 99.43 | 99.39
1.1 | 95.24 | 98.14 | 98.72 | 98.88 | 98.9 | 99.2 | 99.33 | 99.29
1.2 | 93.6 | 97.68 | 98.42 | 98.68 | 98.72 | 99.08 | 99.16 | 99.18
1.3 | 91.54 | 97.12 | 98.3 | 98.48 | 98.56 | 98.88 | 98.96 | 98.99
1.4 | 88.6 | 96.46 | 97.94 | 98.18 | 98.32 | 98.66 | 98.83 | 98.88
1.5 | 84.16 | 95.72 | 97.64 | 97.96 | 98.1 | 98.46 | 98.66 | 98.68
1.6 | 78.2 | 94.9 | 97.24 | 97.6 | 97.92 | 98.18 | 98.46 | 98.48
1.7 | 70.52 | 93.74 | 96.94 | 97.38 | 97.62 | 98.08 | 98.26 | 98.28
1.8 | 61.32 | 92.32 | 96.54 | 97.08 | 97.2 | 97.88 | 98.06 | 98.08
1.9 | 51.2 | 90.44 | 96.22 | 96.86 | 96.96 | 97.66 | 97.86 | 97.87
2 | 41.34 | 87.9 | 95.84 | 96.54 | 96.76 | 97.28 | 97.59 | 97.57
2.1 | 32.16 | 84.4 | 95.34 | 96.08 | 96.42 | 97.08 | 97.36 | 97.28
2.2 | 24.48 | 79.46 | 94.76 | 95.84 | 96.06 | 96.86 | 97.16 | 97.08
2.3 | 18.62 | 72.88 | 94.24 | 95.36 | 95.82 | 96.56 | 96.96 | 96.88
2.4 | 14.5 | 64.78 | 93.66 | 94.96 | 95.42 | 96.26 | 96.62 | 96.68
2.5 | 11.68 | 55.72 | 93.06 | 94.56 | 94.96 | 95.98 | 96.36 | 96.47
2.6 | 9.74 | 46.26 | 92.36 | 94.14 | 94.62 | 95.74 | 96.16 | 96.17
2.7 | 8.28 | 36.96 | 91.66 | 93.64 | 94.16 | 95.38 | 95.82 | 95.87
2.8 | 7.2 | 28.98 | 90.78 | 93.18 | 93.82 | 95.06 | 95.49 | 95.58
2.9 | 6.34 | 22.62 | 89.72 | 92.86 | 93.42 | 94.66 | 95.19 | 95.36
3 | 5.74 | 17.8 | 88.54 | 92.24 | 93.06 | 94.38 | 94.96 | 94.98
3.1 | 5.16 | 14.42 | 87.22 | 91.74 | 92.64 | 94.18 | 94.62 | 94.77
3.2 | 4.76 | 12.2 | 85.48 | 91.18 | 92.22 | 93.76 | 94.29 | 94.48
3.3 | 4.38 | 10.48 | 83.06 | 90.82 | 91.82 | 93.54 | 93.99 | 94.26
3.4 | 4.16 | 9.24 | 80.04 | 90.22 | 91.38 | 93.08 | 93.69 | 93.86
3.5 | 3.88 | 8.26 | 75.86 | 89.64 | 90.88 | 92.78 | 93.39 | 93.48
3.6 | 3.68 | 7.52 | 70.04 | 89.14 | 90.46 | 92.46 | 93.02 | 93.26
3.7 | 3.58 | 6.92 | 62.94 | 88.5 | 89.88 | 92.08 | 92.69 | 92.87
3.8 | 3.38 | 6.38 | 54.56 | 87.72 | 89.42 | 91.66 | 92.32 | 92.57
3.9 | 3.28 | 5.98 | 46.04 | 87.02 | 89.08 | 91.38 | 91.99 | 92.27
4 | 3.08 | 5.68 | 37.72 | 86.22 | 88.58 | 90.98 | 91.69 | 91.97
4.1 | 2.98 | 5.34 | 30.2 | 85.3 | 88.02 | 90.64 | 91.39 | 91.66
4.2 | 2.88 | 5.04 | 24.16 | 84.3 | 87.48 | 90.26 | 91.09 | 91.27
4.3 | 2.8 | 4.9 | 19.6 | 83.18 | 86.92 | 89.88 | 90.72 | 90.97
4.4 | 2.7 | 4.64 | 16.46 | 81.86 | 86.32 | 89.58 | 90.39 | 90.66
4.5 | 2.6 | 4.54 | 14.4 | 80.2 | 85.94 | 89.16 | 90.02 | 90.27
4.6 | 2.5 | 4.34 | 12.74 | 78.16 | 85.32 | 88.78 | 89.69 | 89.97
4.7 | 2.48 | 4.24 | 11.52 | 75.36 | 84.8 | 88.36 | 89.39 | 89.66
4.8 | 2.4 | 4.04 | 10.48 | 71.52 | 84.08 | 87.98 | 88.95 | 89.27
4.9 | 2.3 | 3.94 | 9.78 | 66.4 | 83.54 | 87.66 | 88.59 | 88.96
5 | 2.28 | 3.8 | 9.26 | 59.96 | 82.88 | 87.34 | 88.29 | 88.57
5.1 | 2.2 | 3.7 | 8.74 | 52.54 | 82.24 | 86.96 | 87.92 | 88.27
5.2 | 2.1 | 3.6 | 8.16 | 44.9 | 81.54 | 86.48 | 87.59 | 87.97
5.3 | 2.1 | 3.44 | 7.84 | 37.26 | 80.8 | 86.1 | 87.22 | 87.66
5.4 | 2 | 3.4 | 7.52 | 30.38 | 79.94 | 85.86 | 86.89 | 87.26
5.5 | 2 | 3.3 | 7.22 | 24.84 | 79 | 85.36 | 86.52 | 86.87
5.6 | 1.9 | 3.2 | 6.94 | 20.56 | 77.86 | 84.96 | 86.19 | 86.56
5.7 | 1.9 | 3.1 | 6.72 | 17.66 | 76.62 | 84.58 | 85.89 | 86.17
5.8 | 1.88 | 3.04 | 6.52 | 15.54 | 75.08 | 84.28 | 85.52 | 85.87
5.9 | 1.8 | 2.94 | 6.32 | 14.08 | 73.16 | 83.84 | 85.12 | 85.56
6 | 1.78 | 2.9 | 6.12 | 12.9 | 70.6 | 83.36 | 84.79 | 85.17
6.1 | 1.7 | 2.84 | 5.92 | 12.02 | 67.06 | 82.98 | 84.49 | 84.87
6.2 | 1.7 | 2.74 | 5.8 | 11.32 | 62.38 | 82.66 | 84.12 | 84.56
6.3 | 1.6 | 2.7 | 5.6 | 10.64 | 56.58 | 82.26 | 83.72 | 84.17
6.4 | 1.6 | 2.6 | 5.5 | 10.14 | 49.9 | 81.84 | 83.39 | 83.86
6.5 | 1.58 | 2.54 | 5.32 | 9.66 | 42.76 | 81.46 | 83.02 | 83.47
6.6 | 1.5 | 2.5 | 5.14 | 9.26 | 35.88 | 80.98 | 82.62 | 83.17
6.7 | 1.5 | 2.44 | 5.02 | 8.98 | 29.68 | 80.56 | 82.36 | 82.86
6.8 | 1.48 | 2.4 | 4.92 | 8.68 | 24.62 | 80.26 | 82.09 | 82.47
6.9 |  | 2.34 | 4.82 | 8.38 | 20.76 | 79.94 | 81.72 | 82.16
7 |  | 2.3 | 4.7 | 8.18 | 18.06 | 79.56 | 81.32 | 81.77
7.1 |  | 2.2 | 4.6 | 7.96 | 16.18 | 79.16 | 80.92 | 81.47
7.2 |  | 2.2 | 4.5 | 7.68 | 14.7 | 78.74 | 80.59 | 81.17
7.3 |  | 2.14 | 4.4 | 7.48 | 13.7 | 78.36 | 80.29 | 80.87
7.4 |  | 2.1 | 4.3 | 7.28 | 12.84 | 77.96 | 79.92 | 80.57
7.5 |  |  | 4.2 | 7.08 | 12.14 | 77.58 | 79.59 | 80.27
7.6 |  |  | 4.1 | 6.98 | 11.58 | 77.26 | 79.22 | 79.96
7.7 |  |  | 4 | 6.78 | 11.08 | 76.84 | 78.89 | 79.56
7.8 |  |  | 3.9 | 6.6 | 10.62 | 76.38 | 78.59 | 79.17
7.9 |  |  | 3.8 | 6.48 | 10.22 | 76.06 | 78.29 | 78.87
8 |  |  | 3.72 | 6.3 | 9.92 | 75.72 | 77.92 | 78.57
8.1 |  |  | 3.7 | 6.18 | 9.62 | 75.18 | 77.59 | 78.27
8.2 |  |  | 3.6 | 6 | 9.36 | 74.76 | 77.22 | 77.97
8.3 |  |  | 3.5 | 5.88 | 9.12 | 74.46 | 76.89 | 77.67
8.4 |  |  | 3.4 | 5.78 | 8.86 | 74.04 | 76.59 | 77.36
8.5 |  |  | 3.4 | 5.68 | 8.66 | 73.66 | 76.29 | 76.97
8.6 |  |  | 3.3 | 5.5 | 8.46 | 73.26 | 75.92 | 76.67
8.7 |  |  | 3.22 | 5.4 | 8.26 | 72.94 | 75.59 | 76.37
8.8 |  |  |  | 5.3 | 8.06 | 72.48 | 75.22 | 76.07
8.9 |  |  |  | 5.2 | 7.86 | 72.16 | 74.89 | 75.77
9 |  |  |  | 5.08 | 7.66 | 71.66 | 74.59 | 75.47
9.1 |  |  |  | 4.98 | 7.5 | 71.28 | 74.29 | 75.17
9.2 |  |  |  | 4.88 | 7.36 | 70.86 | 73.99 | 74.86
9.3 |  |  |  | 4.78 | 7.2 | 70.56 | 73.62 | 74.47
9.4 |  |  |  | 4.7 | 7.06 | 70.16 | 73.22 | 74.18
9.5 |  |  |  | 4.58 | 6.9 | 69.76 | 72.96 | 73.96
9.6 |  |  |  | 4.48 | 6.76 | 69.36 | 72.62 | 73.58
9.7 |  |  |  | 4.4 | 6.6 | 68.94 | 72.29 | 73.37
9.8 |  |  |  | 4.3 | 6.5 | 68.46 | 71.99 | 73.06
9.9 |  |  |  | 4.28 | 6.36 | 68.16 | 71.62 | 72.68
10 |  |  |  | 4.18 | 6.26 | 67.76 | 71.29 | 72.46
10.1 |  |  |  |  | 6.16 | 67.36 | 70.99 | 72.07
10.2 |  |  |  |  | 6.06 | 66.96 | 70.69 | 71.78
10.3 |  |  |  |  | 5.86 | 66.46 | 70.39 | 71.57
10.4 |  |  |  |  | 5.76 | 66.16 | 70.02 | 71.27
10.5 |  |  |  |  | 5.66 | 65.76 | 69.76 | 70.97
10.6 |  |  |  |  | 5.56 | 65.34 | 69.49 | 70.67
10.7 |  |  |  |  | 5.46 | 64.88 | 69.12 | 70.37
10.8 |  |  |  |  | 5.36 | 64.46 | 68.86 | 70.07
10.9 |  |  |  |  | 5.26 | 64.06 | 68.59 | 69.78
11 |  |  |  |  | 5.16 | 63.56 | 68.29 | 69.57
11.1 |  |  |  |  | 5.06 | 63.16 | 67.92 | 69.27
11.2 |  |  |  |  | 4.96 | 62.66 | 67.59 | 68.97
11.3 |  |  |  |  | 4.9 | 62.24 | 67.29 | 68.67
11.4 | [5 empty] | 61.76 | 66.99 | 68.37
11.5 | [5 empty] | 61.34 | 66.69 | 68.07
11.6 | [5 empty] | 60.84 | 66.39 | 67.78
11.7 | [5 empty] | 60.34 | 66.09 | 67.57
11.8 | [5 empty] | 59.76 | 65.79 | 67.27
11.9 | [5 empty] | 59.14 | 65.49 | 66.97
12 | [5 empty] | 58.5 | 65.19 | 66.67
12.1 | [5 empty] | 57.72 | 64.89 | 66.38
12.2 | [5 empty] | 56.8 | 64.59 | 66.16
12.3 | [5 empty] | 55.88 | 64.29 | 65.78
12.4 | [5 empty] | 54.56 | 63.99 | 65.57
12.5 | [5 empty] | 52.92 | 63.69 | 65.28
12.6 | [5 empty] | 50.82 | 63.32 | 65.07
12.7 | [5 empty] | 47.86 | 63.06 | 64.77
12.8 | [5 empty] | 44.08 | 62.79 | 64.48
12.9 | [5 empty] | 39.72 | 62.49 | 64.27
13 | [5 empty] | 35.04 | 62.19 | 63.98
13.1 | [5 empty] | 30.44 | 61.89 | 63.77
13.2 | [5 empty] | 26.2 | 61.59 | 63.48
13.3 | [5 empty] | 22.5 | 61.29 | 63.27
13.4 | [5 empty] | 19.68 | 60.99 | 62.98
13.5 | [5 empty] | 17.64 | 60.69 | 62.77
13.6 | [5 empty] | 16.18 | 60.46 | 62.47
13.7 | [5 empty] | 15.1 | 60.19 | 62.18
13.8 | [5 empty] | 14.24 | 59.89 | 61.97
13.9 | [5 empty] | 13.64 | 59.59 | 61.68
14 | [5 empty] | 13.06 | 59.29 | 61.48
14.1 | [5 empty] | 12.58 | 58.99 | 61.26
14.2 | [5 empty] | 12.26 | 58.76 | 60.89
14.3 | [5 empty] | 11.86 | 58.42 | 60.77
14.4 | [5 empty] | 11.56 | 58.16 | 60.47
14.5 | [5 empty] | 11.28 | 57.89 | 60.18
14.6 | [5 empty] | 10.98 | 57.59 | 59.97
14.7 | [5 empty] | 10.78 | 57.29 | 59.68
14.8 | [5 empty] | 10.58 | 56.99 | 59.47
14.9 | [5 empty] | 10.38 | 56.69 | 59.18
15 | [5 empty] | 10.18 | 56.39 | 58.97
15.1 | [5 empty] | 9.98 | 56.16 | 58.68
15.2 | [5 empty] | 9.78 | 55.82 | 58.48
15.3 | [5 empty] | 9.6 | 55.49 | 58.27
15.4 | [5 empty] | 9.48 | 55.26 | 57.98
15.5 | [5 empty] | 9.28 | 54.99 | 57.77
15.6 | [5 empty] | 9.1 | 54.69 | 57.48
15.7 | [5 empty] | 8.98 | 54.39 | 57.28
15.8 | [5 empty] | 8.86 | 54.09 | 57.08
15.9 | [5 empty] | 8.68 | 53.79 | 56.87
16 | [5 empty] | 8.48 | 53.56 | 56.58
16.1 | [5 empty] | 8.38 | 53.22 | 56.38
16.2 | [5 empty] | 8.2 | 52.96 | 56.17
16.3 | [5 empty] | 8.08 | 52.69 | 55.88
16.4 | [5 empty] | 7.98 | 52.39 | 55.68
16.5 | [5 empty] | 7.88 | 52.09 | 55.47
16.6 | [5 empty] | 7.7 | 51.79 | 55.18
16.7 | [5 empty] | 7.58 | 51.49 | 54.98
16.8 | [5 empty] | 7.48 | 51.19 | 54.78
16.9 | [5 empty] | 7.3 | 50.89 | 54.57
17 | [5 empty] | 7.2 | 50.59 | 54.28
17.1 | [5 empty] | 7.1 | 50.29 | 54.07
17.2 | [5 empty] | 6.98 | 49.99 | 53.78
17.3 | [5 empty] | 6.9 | 49.69 | 53.58
17.4 | [5 empty] | 6.8 | 49.39 | 53.38
17.5 | [5 empty] | 6.68 | 49.09 | 53.17
17.6 | [5 empty] | 6.58 | 48.79 | 52.88
17.7 | [5 empty] | 6.5 | 48.49 | 52.68
17.8 | [5 empty] | 6.38 | 48.12 | 52.48
17.9 | [5 empty] | 6.3 | 47.79 | 52.28
18 | [6 empty] | 47.42 | 52.08
18.1 | [6 empty] | 47.02 | 51.87
18.2 | [6 empty] | 46.69 | 51.58
18.3 | [6 empty] | 46.25 | 51.37
18.4 | [6 empty] | 45.82 | 51.08
18.5 | [6 empty] | 45.35 | 50.88
18.6 | [6 empty] | 44.85 | 50.68
18.7 | [6 empty] | 44.21 | 50.48
18.8 | [6 empty] | 43.51 | 50.28
18.9 | [6 empty] | 42.67 | 50.08
19 | [6 empty] | 41.49 | 49.88
19.1 | [6 empty] | 39.98 | 49.67
19.2 | [6 empty] | 37.96 | 49.38
19.3 | [6 empty] | 35.41 | 49.18
19.4 | [6 empty] | 32.43 | 48.98
19.5 | [6 empty] | 29.12 | 48.77
19.6 | [6 empty] | 25.79 | 48.49
19.7 | [6 empty] | 22.56 | 48.38
19.8 | [6 empty] | 19.78 | 48.17
19.9 | [6 empty] | 17.53 | 47.88
20 | [6 empty] | 15.85 | 47.68
20.1 | [6 empty] | 14.63 | 47.48
20.2 | [6 empty] | 13.67 | 47.28
20.3 | [6 empty] | 12.98 | 47.07
20.4 | [6 empty] | 12.45 | 46.78
20.5 | [6 empty] | 12.02 | 46.58
20.6 | [6 empty] | 11.62 | 46.38
20.7 | [6 empty] | 11.29 | 46.18
20.8 | [6 empty] | 11.06 | 45.98
20.9 | [6 empty] | 10.79 | 45.78
21 | [6 empty] | 10.56 | 45.57
21.1 | [6 empty] | 10.36 | 45.28
21.2 | [6 empty] | 10.16 | 45.08
21.3 | [6 empty] | 9.96 | 44.88
21.4 | [6 empty] | 9.83 | 44.68
21.5 | [6 empty] | 9.66 | 44.48
21.6 | [6 empty] | 9.53 | 44.27
21.7 | [6 empty] | 9.36 | 43.99
21.8 | [6 empty] | 9.23 | 43.87
21.9 | [6 empty] | 9.06 | 43.58
22 | [6 empty] | 8.93 | 43.38
22.1 | [6 empty] | 8.76 | 43.18
22.2 | [6 empty] | 8.63 | 42.98
22.3 | [6 empty] | 8.53 | 42.78
22.4 | [6 empty] | 8.43 | 42.57
22.5 | [6 empty] | 8.26 | 42.28
22.6 | [6 empty] | 8.2 | 42.08
22.7 | [6 empty] | 8.06 | 41.87
22.8 | [6 empty] | 7.93 | 41.58
22.9 | [6 empty] | 7.83 | 41.38
23 | [6 empty] | 7.73 | 41.18
23.1 | [6 empty] | 7.63 | 40.98
23.2 | [6 empty] | 7.53 | 40.78
23.3 | [6 empty] | 7.43 | 40.57
23.4 | [6 empty] | 7.33 | 40.28
23.5 | [6 empty] | 7.23 | 40.08
23.6 | [6 empty] | 7.06 | 39.88
23.7 | [6 empty] | 7 | 39.68
23.8 | [6 empty] | 6.93 | 39.47
23.9 | [6 empty] | 6.83 | 39.17
24 | [6 empty] | 6.73 | 38.88
24.1 | [6 empty] | 6.63 | 38.68
24.2 | [6 empty] | 6.53 | 38.47
24.3 | [6 empty] | 6.43 | 38.18
24.4 | [6 empty] | 6.4 | 37.97
24.5 | [6 empty] | 6.33 | 37.67
24.6 | [6 empty] | 6.23 | 37.38
24.7 | [6 empty] | 6.13 | 37.16
24.8 | [6 empty] | 6.1 | 36.77
24.9 | [6 empty] | 6.03 | 36.46
25 | [6 empty] | 5.93 | 36.06
25.1 | [6 empty] | 5.83 | 35.65
25.2 | [6 empty] | 5.73 | 35.15
25.3 | [6 empty] | 5.7 | 34.63
25.4 | [6 empty] | 5.63 | 33.91
25.5 | [6 empty] | 5.53 | 32.99
25.6 | [6 empty] | 5.5 | 31.84
25.7 | [6 empty] | 5.43 | 30.2
25.8 | [6 empty] | 5.33 | 28.16
25.9 | [6 empty] | 5.3 | 25.75
26 | [6 empty] | 5.23 | 23.24
26.1 | [6 empty] | 5.13 | 20.66
26.2 | [6 empty] | 5.1 | 18.29
26.3 | [6 empty] | 5.03 | 16.22
26.4 | [6 empty] | 5 | 14.47
26.5 | [6 empty] | 4.93 | 13.21
26.6 | [6 empty] | 4.83 | 12.33
26.7 | [6 empty] | 4.8 | 11.64
26.8 | [6 empty] | 4.73 | 11.06
26.9 | [6 empty] | 4.63 | 10.67
27 | [6 empty] | 4.6 | 10.37
27.1 | [6 empty] | 4.53 | 10.07
27.2 | [6 empty] | 4.5 | 9.78
27.3 | [6 empty] | 4.43 | 9.58
27.4 | [6 empty] | 4.4 | 9.39
27.5 | [7 empty] | 9.28
27.6 | [7 empty] | 9.08
27.7 | [7 empty] | 8.88
27.8 | [7 empty] | 8.69
27.9 | [7 empty] | 8.59
28 | [7 empty] | 8.49
28.1 | [7 empty] | 8.38
28.2 | [7 empty] | 8.19
28.3 | [7 empty] | 8.09
28.4 | [7 empty] | 7.99
28.5 | [7 empty] | 7.89
28.6 | [7 empty] | 7.79
28.7 | [7 empty] | 7.69
28.8 | [7 empty] | 7.59
28.9 | [7 empty] | 7.49
29 | [7 empty] | 7.39
29.1 | [7 empty] | 7.29
29.2 | [7 empty] | 7.19
29.3 | [7 empty] | 7.09
29.4 | [7 empty] | 6.99
29.5 | [7 empty] | 6.89
29.6 | [7 empty] | 6.79
29.7 | [7 empty] | 6.69

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2.4 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-45.1 |  |  |  | 1.2
-45 |  |  |  | 1.2
-44.9 |  |  |  | 1.2
-44.8 |  |  |  | 1.2
-44.7 |  |  |  | 1.2
-44.6 |  |  |  | 1.2
-44.5 |  |  |  | 1.2
-44.4 |  |  |  | 1.2
-44.3 |  |  |  | 1.219
-44.2 |  |  |  | 1.3
-44.1 |  |  |  | 1.3
-44 |  |  |  | 1.3
-43.9 |  |  |  | 1.3
-43.8 |  |  |  | 1.3
-43.7 |  |  |  | 1.3
-43.6 |  |  |  | 1.3
-43.5 |  |  |  | 1.3
-43.4 |  |  |  | 1.3
-43.3 |  |  |  | 1.3176
-43.2 |  |  |  | 1.4
-43.1 |  |  |  | 1.4
-43 |  |  |  | 1.4
-42.9 |  |  |  | 1.4
-42.8 |  |  |  | 1.4
-42.7 |  |  |  | 1.4
-42.6 |  |  |  | 1.4
-42.5 |  |  |  | 1.4
-42.4 |  |  |  | 1.4
-42.3 |  |  |  | 1.4134
-42.2 |  |  |  | 1.5
-42.1 |  |  |  | 1.5 | 2.5
-42 |  |  |  | 1.5 | 2.5
-41.9 |  |  |  | 1.5 | 2.5746
-41.8 |  |  |  | 1.5 | 2.6
-41.7 |  |  |  | 1.5 | 2.6
-41.6 |  |  |  | 1.5 | 2.6
-41.5 |  |  |  | 1.5215 | 2.6073
-41.4 |  |  |  | 1.6 | 2.7
-41.3 |  |  |  | 1.6 | 2.7
-41.2 |  |  |  | 1.6 | 2.7103
-41.1 |  |  |  | 1.6 | 2.8
-41 |  |  |  | 1.6 | 2.8
-40.9 |  | 0.6 |  | 1.6 | 2.8
-40.8 |  | 0.6 |  | 1.6126 | 2.825
-40.7 |  | 0.6 |  | 1.7 | 2.9
-40.6 |  | 0.6 |  | 1.7 | 2.9
-40.5 |  | 0.6 |  | 1.7 | 2.9
-40.4 |  | 0.6 |  | 1.7 | 2.9458
-40.3 |  | 0.6 | 0.9 | 1.7 | 3
-40.2 | 0.5 | 0.6 | 0.9 | 1.7204 | 3
-40.1 | 0.5 | 0.6 | 0.9 | 1.8 | 3
-40 | 0.5 | 0.6 | 0.9 | 1.8 | 3.1
-39.9 | 0.5 | 0.6 | 0.9 | 1.8 | 3.1
-39.8 | 0.5 | 0.6 | 0.9 | 1.8 | 3.1245
-39.7 | 0.5 | 0.6 | 0.9 | 1.8 | 3.2
-39.6 | 0.5 | 0.6 | 0.9 | 1.8122 | 3.2
-39.5 | 0.5 | 0.6 | 0.9 | 1.9 | 3.2276
-39.4 | 0.5 | 0.6 | 0.9 | 1.9 | 3.3
-39.3 | 0.5 | 0.6 | 0.9 | 1.9 | 3.3
-39.2 | 0.55919 | 0.6 | 0.90202 | 1.9 | 3.319
-39.1 | 0.6 | 0.7 | 1 | 1.9035 | 3.4
-39 | 0.6 | 0.7 | 1 | 2 | 3.4
-38.9 | 0.6 | 0.7 | 1 | 2 | 3.4726
-38.8 | 0.6 | 0.7 | 1 | 2 | 3.5
-38.7 | 0.6 | 0.7 | 1 | 2 | 3.5071
-38.6 | 0.6 | 0.7 | 1 | 2.0367 | 3.6
-38.5 | 0.6 | 0.7 | 1 | 2.1 | 3.6
-38.4 | 0.6 | 0.7 | 1 | 2.1 | 3.6118
-38.3 | 0.6 | 0.7 | 1 | 2.1 | 3.7
-38.2 | 0.6 | 0.7 | 1 | 2.1 | 3.7
-38.1 | 0.6 | 0.7 | 1.0989 | 2.132 | 3.7654
-38 | 0.6 | 0.7 | 1.1 | 2.2 | 3.8
-37.9 | 0.6 | 0.7 | 1.1 | 2.2 | 3.9
-37.8 | 0.6 | 0.7 | 1.1 | 2.2 | 3.9
-37.7 | 0.6 | 0.7 | 1.1 | 2.2459 | 3.9368
-37.6 | 0.6 | 0.7 | 1.1 | 2.3 | 4
-37.5 | 0.6 | 0.7 | 1.1 | 2.3 | 4.0725
-37.4 | 0.6 | 0.7 | 1.1 | 2.3 | 4.1
-37.3 | 0.6 | 0.77488 | 1.1 | 2.3641 | 4.1082
-37.2 | 0.6 | 0.8 | 1.1071 | 2.4 | 4.2
-37.1 | 0.6 | 0.8 | 1.1954 | 2.4 | 4.2439
-37 | 0.7 | 0.8 | 1.2 | 2.4 | 4.3
-36.9 | 0.7 | 0.8 | 1.2 | 2.4539 | 4.3796
-36.8 | 0.7 | 0.8 | 1.2 | 2.5 | 4.4
-36.7 | 0.7 | 0.8 | 1.2 | 2.5 | 4.4845
-36.6 | 0.7 | 0.8 | 1.2 | 2.5499 | 4.5856
-36.5 | 0.7 | 0.8 | 1.2 | 2.6 | 4.6
-36.4 | 0.7 | 0.8 | 1.2 | 2.6 | 4.7
-36.3 | 0.7 | 0.8 | 1.2743 | 2.6534 | 4.7026
-36.2 | 0.7 | 0.8 | 1.3 | 2.7 | 4.7909
-36.1 | 0.7 | 0.8 | 1.3 | 2.7 | 4.8906
-36 | 0.7 | 0.8 | 1.3 | 2.7631 | 4.9
-35.9 | 0.7 | 0.89263 | 1.3 | 2.8 | 5.0081
-35.8 | 0.7 | 0.9 | 1.3 | 2.8 | 5.1
-35.7 | 0.7 | 0.9 | 1.3 | 2.8662 | 5.1232
-35.6 | 0.7 | 0.9 | 1.395 | 2.9 | 5.2242
-35.5 | 0.70964 | 0.9 | 1.4 | 2.9 | 5.3253
-35.4 | 0.79805 | 0.9 | 1.4 | 2.964 | 5.4
-35.3 | 0.8 | 0.9 | 1.4 | 3 | 5.4274
-35.2 | 0.8 | 0.9 | 1.4 | 3 | 5.5284
-35.1 | 0.8 | 0.9 | 1.4 | 3.0842 | 5.6343
-35 | 0.8 | 0.9 | 1.4 | 3.1018 | 5.7521
-34.9 | 0.8 | 0.9 | 1.5 | 3.2 | 5.8
-34.8 | 0.8 | 0.90384 | 1.5 | 3.2038 | 5.861
-34.7 | 0.8 | 1 | 1.5 | 3.3 | 5.9723
-34.6 | 0.8 | 1 | 1.5 | 3.3 | 6.0773
-34.5 | 0.8 | 1 | 1.5 | 3.3069 | 6.237
-34.4 | 0.8 | 1 | 1.6 | 3.4 | 6.3651
-34.3 | 0.82323 | 1 | 1.6 | 3.4 | 6.4661
-34.2 | 0.9 | 1 | 1.6 | 3.5 | 6.6341
-34.1 | 0.9 | 1 | 1.6 | 3.5109 | 6.8362
-34 | 0.9 | 1 | 1.6 | 3.6119 | 6.9691
-33.9 | 0.9 | 1.0273 | 1.6985 | 3.7 | 7.1637
-33.8 | 0.9 | 1.1 | 1.7 | 3.714 | 7.4242
-33.7 | 0.9 | 1.1 | 1.7 | 3.8 | 7.6684
-33.6 | 0.9 | 1.1 | 1.7 | 3.814 | 8.0069
-33.5 | 0.9 | 1.1 | 1.7313 | 3.9027 | 8.4127
-33.4 | 0.9 | 1.1 | 1.8 | 4 | 8.9085
-33.3 | 0.9 | 1.1 | 1.8 | 4.0047 | 9.5002
-33.2 | 0.94869 | 1.1 | 1.8 | 4.1058 | 10.0461
-33.1 | 1 | 1.1354 | 1.8 | 4.2 | 10.8609
-33 | 1 | 1.2 | 1.8925 | 4.2258 | 11.7856
-32.9 | 1 | 1.2 | 1.9 | 4.3374 | 12.9056
-32.8 | 1 | 1.2 | 1.9 | 4.4384 | 14.1385
-32.7 | 1 | 1.2 | 1.9966 | 4.5394 | 15.5523
-32.6 | 1 | 1.2 | 2 | 4.6 | 16.9907
-32.5 | 1 | 1.2 | 2 | 4.6272 | 18.6069
-32.4 | 1.0568 | 1.2425 | 2.0329 | 4.7329 | 20.223
-32.3 | 1.1 | 1.3 | 2.1 | 4.8435 | 21.8093
-32.2 | 1.1 | 1.3 | 2.1 | 4.9445 | 23.1756
-32.1 | 1.1 | 1.3 | 2.1456 | 5.0456 | 24.2781
-32 | 1.1 | 1.3 | 2.2 | 5.1466 | 25.2578
-31.9 | 1.1 | 1.3 | 2.2 | 5.2832 | 26.0036
-31.8 | 1.1 | 1.3 | 2.2628 | 5.4343 | 26.58
-31.7 | 1.1638 | 1.3782 | 2.3 | 5.5707 | 27.0051
-31.6 | 1.2 | 1.4 | 2.3 | 5.7364 | 27.4661
-31.5 | 1.2 | 1.4 | 2.3659 | 5.8747 | 27.7691
-31.4 | 1.2 | 1.4 | 2.4 | 6.1152 | 28.0721
-31.3 | 1.2 | 1.4793 | 2.4793 | 6.4034 | 28.3439
-31.2 | 1.2 | 1.5 | 2.5 | 6.6782 | 28.5331
-31.1 | 1.2557 | 1.5 | 2.5 | 7.0083 | 28.7828
-31 | 1.3 | 1.5 | 2.6 | 7.4405 | 28.9989
-30.9 | 1.3 | 1.5 | 2.6173 | 7.9752 | 29.2154
-30.8 | 1.3 | 1.5764 | 2.7 | 8.6114 | 29.4174
-30.7 | 1.3 | 1.6 | 2.7312 | 9.3457 | 29.5598
-30.6 | 1.3 | 1.6 | 2.8 | 10.1788 | 29.7418
-30.5 | 1.3475 | 1.6 | 2.8388 | 11.2267 | 29.9774
-30.4 | 1.4 | 1.6913 | 2.9566 | 12.4786 | 30.0913
-30.3 | 1.4 | 1.7 | 3 | 13.9136 | 30.3133
-30.2 | 1.4 | 1.7 | 3.0756 | 15.4454 | 30.5154
-30.1 | 1.4373 | 1.7087 | 3.1 | 17.0564 | 30.7076
-30 | 1.5 | 1.8 | 3.2113 | 18.8745 | 30.796
-29.9 | 1.5 | 1.8 | 3.3 | 20.6927 | 31.025
-29.8 | 1.5 | 1.8403 | 3.3546 | 22.4758 | 31.2261
-29.7 | 1.5362 | 1.9 | 3.4649 | 23.9782 | 31.3541
-29.6 | 1.6 | 1.9 | 3.5 | 25.3236 | 31.5393
-29.5 | 1.6 | 1.9433 | 3.6006 | 26.2377 | 31.7345
-29.4 | 1.6 | 2 | 3.7184 | 27.058 | 31.8851
-29.3 | 1.6696 | 2.0397 | 3.8363 | 27.5839 | 32.0766
-29.2 | 1.7 | 2.1 | 3.9541 | 28.0715 | 32.2786
-29.1 | 1.7 | 2.1331 | 4.072 | 28.4567 | 32.3903
-29 | 1.7 | 2.2 | 4.1899 | 28.7597 | 32.5826
-28.9 | 1.778 | 2.2352 | 4.3846 | 29.0549 | 32.7615
-28.8 | 1.8 | 2.3 | 4.5867 | 29.3152 | 32.958
-28.7 | 1.8801 | 2.3229 | 4.8258 | 29.5258 | 33.0934
-28.6 | 1.9 | 2.4239 | 5.1336 | 29.8192 | 33.2419
-28.5 | 1.9 | 2.5 | 5.5164 | 30.0212 | 33.4106
-28.4 | 1.9974 | 2.526 | 6.0347 | 30.2233 | 33.5233
-28.3 | 2 | 2.627 | 6.6239 | 30.4223 | 33.7297
-28.2 | 2.0852 | 2.7326 | 7.296 | 30.5989 | 33.9326
-28.1 | 2.1 | 2.8 | 8.1755 | 30.8008 | 34.0605
-28 | 2.1872 | 2.8517 | 9.2198 | 31.0028 | 34.2275
-27.9 | 2.2 | 2.9596 | 10.6515 | 31.2048 | 34.4192
-27.8 | 2.2892 | 3.0708 | 12.2378 | 31.4069 | 34.6212
-27.7 | 2.3 | 3.2771 | 14.1943 | 31.6044 | 34.7616
-27.6 | 2.377 | 3.4109 | 16.3148 | 31.7127 | 34.9462
-27.5 | 2.4909 | 3.6197 | 18.5763 | 31.9364 | 35.0682
-27.4 | 2.5 | 3.9227 | 20.9578 | 32.1152 | 35.2293
-27.3 | 2.61 | 4.2343 | 23.0792 | 32.3086 | 35.4533
-27.2 | 2.7209 | 4.648 | 24.9675 | 32.4192 | 35.5945
-27.1 | 2.8124 | 5.1742 | 26.4483 | 32.6165 | 35.7925
-27 | 2.9517 | 5.8813 | 27.6964 | 32.7736 | 35.8973
-26.9 | 3.1537 | 6.588 | 28.5723 | 32.996 | 36.1126
-26.8 | 3.3558 | 7.6922 | 29.2103 | 33.1986 | 36.2318
-26.7 | 3.5867 | 9.1391 | 29.7016 | 33.3008 | 36.4578
-26.6 | 3.9396 | 10.8585 | 30.1042 | 33.5033 | 36.6349
-26.5 | 4.4263 | 12.8956 | 30.4097 | 33.7292 | 36.8055
-26.4 | 5.0237 | 15.1529 | 30.8069 | 33.9069 | 36.9706
-26.3 | 5.7309 | 17.5588 | 31.0104 | 34.1044 | 37.1769
-26.2 | 6.6687 | 20.1851 | 31.2591 | 34.2109 | 37.2911
-26.1 | 7.9124 | 22.7412 | 31.5363 | 34.4113 | 37.4842
-26 | 9.4879 | 25.0573 | 31.7174 | 34.588 | 37.5931
-25.9 | 11.2963 | 26.9187 | 31.9455 | 34.7885 | 37.7648
-25.8 | 13.6861 | 28.2848 | 32.1475 | 34.8953 | 37.9889
-25.7 | 16.3991 | 29.4227 | 32.3289 | 35.0925 | 38.1925
-25.6 | 18.967 | 30.2289 | 32.4802 | 35.2945 | 38.2973
-25.5 | 21.7503 | 30.7075 | 32.6822 | 35.386 | 38.4966
-25.4 | 24.4936 | 31.0975 | 32.8842 | 35.5699 | 38.6986
-25.3 | 26.7911 | 31.451 | 33.0863 | 35.7719 | 38.8754
-25.2 | 28.6655 | 31.7598 | 33.2883 | 35.9739 | 38.987
-25.1 | 30.0437 | 32.0081 | 33.479 | 36.154 | 39.176
-25 | 30.9126 | 32.2462 | 33.6372 | 36.2747 | 39.3782
-24.9 | 31.4728 | 32.4415 | 33.7943 | 36.4515 | 39.5576
-24.8 | 31.9071 | 32.6964 | 33.9964 | 36.6535 | 39.7535
-24.7 | 32.2333 | 32.8984 | 34.1492 | 36.7778 | 39.8778
-24.6 | 32.5364 | 33.0586 | 34.3172 | 36.9379 | 40.0576
-24.5 | 32.7596 | 33.2529 | 34.5529 | 37.1311 | 40.2397
-24.4 | 33.0424 | 33.4616 | 34.6808 | 37.3331 | 40.3776
-24.3 | 33.2432 | 33.5818 | 34.891 | 37.5352 | 40.591
-24.2 | 33.4372 | 33.7658 | 35.0228 | 37.7201 | 40.6972
-24.1 | 33.5696 | 33.9958 | 35.2248 | 37.9107 | 40.8964
-24 | 33.7412 | 34.1984 | 35.4157 | 38.0564 | 41.0984
-23.9 | 33.9672 | 34.3004 | 35.5336 | 38.2147 | 41.3004
-23.8 | 34.087 | 34.5012 | 35.7028 | 38.3584 | 41.5024
-23.7 | 34.276 | 34.6052 | 35.9188 | 38.504 | 41.7019
-23.6 | 34.4557 | 34.8176 | 36.0604 | 38.6923 | 41.7904
-23.5 | 34.5757 | 34.9372 | 36.2075 | 38.8943 | 41.98
-23.4 | 34.779 | 35.1229 | 36.3964 | 39.0964 | 42.1392
-23.3 | 34.892 | 35.2488 | 36.5984 | 39.286 | 42.3412
-23.2 | 35.1073 | 35.4128 | 36.8004 | 39.4359 | 42.5216
-23.1 | 35.2166 | 35.6024 | 36.9448 | 39.5737 | 42.6528
-23 | 35.4188 | 35.7522 | 37.0758 | 39.7758 | 42.8884
-22.9 | 35.6121 | 35.9065 | 37.2389 | 39.9681 | 43.0532
-22.8 | 35.73 | 36.095 | 37.3932 | 40.1513 | 43.2085
-22.7 | 35.882 | 36.241 | 37.6105 | 40.3267 | 43.4289
-22.6 | 36.049 | 36.384 | 37.7563 | 40.4727 | 43.5823
-22.5 | 36.2337 | 36.5861 | 37.9003 | 40.7144 | 43.7717
-22.4 | 36.401 | 36.7514 | 38.0881 | 40.8309 | 44.0011
-22.3 | 36.4894 | 36.9384 | 38.2451 | 41.0329 | 44.1037
-22.2 | 36.6778 | 37.087 | 38.3921 | 41.2153 | 44.2805
-22.1 | 36.8085 | 37.2798 | 38.5941 | 41.3085 | 44.4798
-22 | 37.0122 | 37.3909 | 38.7842 | 41.5158 | 44.6818
-21.9 | 37.2205 | 37.5838 | 38.9348 | 41.8063 | 44.8125
-21.8 | 37.3502 | 37.6813 | 39.0717 | 41.9218 | 45.0145
-21.7 | 37.5369 | 37.8594 | 39.2737 | 42.2072 | 45.2193
-21.6 | 37.6758 | 38.0614 | 39.4663 | 42.2956 | 45.4471
-21.5 | 37.834 | 38.1953 | 39.6491 | 42.396 | 45.6215
-21.4 | 37.9511 | 38.3698 | 39.8256 | 42.6912 | 45.7396
-21.3 | 38.1531 | 38.4837 | 39.9698 | 42.8962 | 46.031
-21.2 | 38.3241 | 38.6695 | 40.2133 | 42.9867 | 46.1733
-21.1 | 38.4287 | 38.8715 | 40.3335 | 43.1715 | 46.3572
-21 | 38.6358 | 39.0519 | 40.5346 | 43.3735 | 46.5594
-20.9 | 38.8307 | 39.2471 | 40.705 | 43.5756 | 46.7307
-20.8 | 38.9634 | 39.3869 | 40.8603 | 43.7554 | 46.874
-20.7 | 39.1327 | 39.5097 | 41.0226 | 43.9511 | 47.1097
-20.6 | 39.3012 | 39.7226 | 41.2246 | 44.1531 | 47.3246
-20.5 | 39.5347 | 39.881 | 41.4267 | 44.3552 | 47.5267
-20.4 | 39.6625 | 40.0584 | 41.6125 | 44.4917 | 47.7287
-20.3 | 39.8524 | 40.2524 | 41.751 | 44.6879 | 47.8762
-20.2 | 40.0378 | 40.3941 | 41.904 | 44.8661 | 48.0882
-20.1 | 40.1776 | 40.5919 | 42.1061 | 45.0632 | 48.3119
-20 | 40.3398 | 40.7939 | 42.3082 | 45.2653 | 48.4596
-19.9 | 40.4953 | 40.9714 | 42.5103 | 45.4673 | 48.6953
-19.8 | 40.6982 | 41.0846 | 42.6562 | 45.6693 | 48.8655
-19.7 | 40.85 | 41.2715 | 42.8715 | 45.8 | 49.0667
-19.6 | 41.0018 | 41.4735 | 43.0582 | 46.0305 | 49.302
-19.5 | 41.1897 | 41.5768 | 43.2184 | 46.2285 | 49.5035
-19.4 | 41.4071 | 41.8233 | 43.4404 | 46.4061 | 49.6804
-19.3 | 41.6224 | 41.9755 | 43.6762 | 46.6081 | 49.9081
-19.2 | 41.7622 | 42.1529 | 43.8088 | 46.8118 | 50.1101
-19.1 | 41.9475 | 42.3549 | 43.9857 | 47.0408 | 50.3106
-19 | 42.1374 | 42.5833 | 44.1857 | 47.2375 | 50.4874
-18.9 | 42.2653 | 42.6938 | 44.4094 | 47.4162 | 50.6877
-18.8 | 42.4325 | 42.9182 | 44.5547 | 47.6182 | 50.8897
-18.7 | 42.6571 | 43.1177 | 44.7452 | 47.8202 | 51.0917
-18.6 | 42.7826 | 43.2945 | 44.9513 | 48.0224 | 51.2937
-18.5 | 42.9673 | 43.4958 | 45.1673 | 48.2214 | 51.4367
-18.4 | 43.1974 | 43.6309 | 45.3693 | 48.3981 | 51.7116
-18.3 | 43.4142 | 43.8333 | 45.4999 | 48.5998 | 51.7999
-18.2 | 43.5305 | 43.9877 | 45.7018 | 48.9018 | 52.1018
-18.1 | 43.7284 | 44.1897 | 45.9038 | 49.1019 | 52.3045
-18 | 43.9068 | 44.3917 | 46.1059 | 49.2059 | 52.5345
-17.9 | 44.1366 | 44.5547 | 46.3079 | 49.4079 | 52.732
-17.8 | 44.3386 | 44.7864 | 46.5116 | 49.6099 | 52.9139
-17.7 | 44.5406 | 45.0263 | 46.7355 | 49.8182 | 53.1403
-17.6 | 44.7426 | 45.2283 | 46.9141 | 50.1141 | 53.3309
-17.5 | 44.9391 | 45.4303 | 47.1162 | 50.3094 | 53.6098
-17.4 | 45.1182 | 45.6158 | 47.3182 | 50.4545 | 53.833
-17.3 | 45.3236 | 45.8059 | 47.5202 | 50.6774 | 53.973
-17.2 | 45.5296 | 46.0079 | 47.7222 | 50.8694 | 54.2507
-17.1 | 45.6814 | 46.2099 | 47.9212 | 51.0527 | 54.4527
-17 | 45.8834 | 46.398 | 48.0979 | 51.2821 | 54.6479
-16.9 | 46.0855 | 46.5855 | 48.2998 | 51.5568 | 54.8424
-16.8 | 46.2765 | 46.7875 | 48.5353 | 51.7686 | 55.1303
-16.7 | 46.4609 | 46.9895 | 48.771 | 52.0044 | 55.3323
-16.6 | 46.6735 | 47.2067 | 48.9915 | 52.22 | 55.5343
-16.5 | 46.8935 | 47.422 | 49.1935 | 52.422 | 55.7477
-16.4 | 47.0956 | 47.6086 | 49.4116 | 52.6086 | 56.0099
-16.3 | 47.2976 | 47.8139 | 49.5631 | 52.8464 | 56.2179
-16.2 | 47.4996 | 48.0283 | 49.7496 | 53.0996 | 56.507
-16.1 | 47.6889 | 48.2303 | 50.0382 | 53.3016 | 56.616
-16 | 47.8876 | 48.4543 | 50.2608 | 53.4907 | 56.818
-15.9 | 48.1057 | 48.6628 | 50.5138 | 53.6772 | 57.03
-15.8 | 48.3615 | 48.8925 | 50.7364 | 53.9188 | 57.3257
-15.7 | 48.564 | 49.0956 | 50.9384 | 54.1812 | 57.5307
-15.6 | 48.7637 | 49.2976 | 51.1404 | 54.4093 | 57.7248
-15.5 | 48.9995 | 49.4996 | 51.3661 | 54.6568 | 57.9853
-15.4 | 49.2019 | 49.7019 | 51.5514 | 54.8882 | 58.2146
-15.3 | 49.4321 | 49.9422 | 51.8197 | 55.1608 | 58.4709
-15.2 | 49.6341 | 50.2057 | 52.0485 | 55.3628 | 58.7066
-15.1 | 49.8542 | 50.4077 | 52.2505 | 55.5851 | 58.9423
-15 | 50.1446 | 50.6097 | 52.5288 | 55.8191 | 59.1145
-14.9 | 50.3689 | 50.8176 | 52.6902 | 55.9808 | 59.4103
-14.8 | 50.5709 | 51.116 | 53.0137 | 56.3371 | 59.5871
-14.7 | 50.7958 | 51.3259 | 53.2184 | 56.5236 | 59.8809
-14.6 | 51.0465 | 51.4984 | 53.4542 | 56.8178 | 60.1178
-14.5 | 51.2485 | 51.7485 | 53.6899 | 57.0198 | 60.3231
-14.4 | 51.4505 | 51.9758 | 53.9077 | 57.2327 | 60.5505
-14.3 | 51.6788 | 52.2306 | 54.1097 | 57.5209 | 60.7788
-14.2 | 51.9636 | 52.4245 | 54.3954 | 57.6977 | 61.0818
-14.1 | 52.1993 | 52.6851 | 54.6422 | 57.9991 | 61.366
-14 | 52.435 | 52.9143 | 54.8263 | 58.1452 | 61.5871
-13.9 | 52.6708 | 53.1909 | 55.0767 | 58.4482 | 61.8338
-13.8 | 52.877 | 53.4626 | 55.3797 | 58.7341 | 62.1065
-13.7 | 53.1087 | 53.697 | 55.6218 | 58.9474 | 62.342
-13.6 | 53.3445 | 53.9584 | 55.8626 | 59.2142 | 62.5525
-13.5 | 53.5802 | 54.1469 | 56.0974 | 59.5115 | 62.7802
-13.4 | 53.7994 | 54.4238 | 56.3739 | 59.7203 | 63.0776
-13.3 | 54.0521 | 54.7014 | 56.6299 | 60.0136 | 63.2871
-13.2 | 54.3552 | 54.9034 | 56.8319 | 60.1905 | 63.6178
-13.1 | 54.6055 | 55.1582 | 57.1009 | 60.4867 | 63.826
-13 | 54.8253 | 55.4253 | 57.3785 | 60.6931 | 64.091
-12.9 | 55.0611 | 55.707 | 57.6277 | 60.9927 | 64.3417
-12.8 | 55.353 | 55.9225 | 57.8838 | 61.2589 | 64.6476
-12.7 | 55.6241 | 56.159 | 58.1703 | 61.5561 | 64.8488
-12.6 | 55.9164 | 56.4268 | 58.4349 | 61.8018 | 65.1682
-12.5 | 56.1463 | 56.7194 | 58.6706 | 62.1032 | 65.3891
-12.4 | 56.4224 | 57.0224 | 59.0055 | 62.3082 | 65.6367
-12.3 | 56.6972 | 57.2503 | 59.2112 | 62.6098 | 65.9397
-12.2 | 56.9236 | 57.4776 | 59.5111 | 62.8748 | 66.2427
-12.1 | 57.1885 | 57.8134 | 59.7467 | 63.1742 | 66.4972
-12 | 57.4915 | 58.063 | 59.9824 | 63.4773 | 66.7157
-11.9 | 57.7703 | 58.344 | 60.2772 | 63.6636 | 66.9703
-11.8 | 58.0548 | 58.5807 | 60.5538 | 63.9406 | 67.2548
-11.7 | 58.3053 | 58.9152 | 60.8841 | 64.2436 | 67.5579
-11.6 | 58.5879 | 59.1919 | 61.1533 | 64.5467 | 67.8609
-11.5 | 58.881 | 59.4914 | 61.4497 | 64.8497 | 68.1093
-11.4 | 59.1113 | 59.74 | 61.7527 | 65.0685 | 68.367
-11.3 | 59.3985 | 59.9863 | 61.9492 | 65.3239 | 68.6488
-11.2 | 59.7161 | 60.273 | 62.2519 | 65.6161 | 68.9013
-11.1 | 60.0191 | 60.6055 | 62.6048 | 65.9191 | 69.137
-11 | 60.3221 | 60.9221 | 62.9061 | 66.2221 | 69.4221
-10.9 | 60.5501 | 61.1501 | 63.1536 | 66.4971 | 69.7252
-10.8 | 60.8282 | 61.4663 | 63.4567 | 66.7855 | 70.0661
-10.7 | 61.1698 | 61.8198 | 63.7697 | 67.0885 | 70.3238
-10.6 | 61.4915 | 62.0562 | 64.0705 | 67.3915 | 70.6705
-10.5 | 61.7945 | 62.3373 | 64.3088 | 67.6945 | 70.8846
-10.4 | 62.0976 | 62.7204 | 64.6304 | 67.9728 | 71.1353
-10.3 | 62.3754 | 62.9622 | 64.984 | 68.2576 | 71.4531
-10.2 | 62.6608 | 63.2876 | 65.3036 | 68.5608 | 71.7786
-10.1 | 62.9639 | 63.641 | 65.6067 | 68.8434 | 72.0612
-10 | 63.2947 | 63.9631 | 65.8631 | 69.1653 | 72.424
-9.9 | 63.6482 | 64.2482 | 66.1982 | 69.4847 | 72.7112
-9.8 | 63.9015 | 64.6018 | 66.5018 | 69.7138 | 72.9872
-9.7 | 64.2053 | 64.8915 | 66.8474 | 69.9903 | 73.2903
-9.6 | 64.5505 | 65.2089 | 67.1505 | 70.3244 | 73.5933
-9.5 | 64.8714 | 65.5249 | 67.4624 | 70.6964 | 73.8562
-9.4 | 65.2566 | 65.828 | 67.8137 | 71.016 | 74.066
-9.3 | 65.5596 | 66.208 | 68.1223 | 71.3596 | 74.4195
-9.2 | 65.8731 | 66.5048 | 68.5132 | 71.6418 | 74.7731
-9.1 | 66.2228 | 66.859 | 68.7267 | 71.8876 | 75.1266
-9 | 66.5259 | 67.1973 | 69.0802 | 72.2802 | 75.3601
-8.9 | 66.8386 | 67.5671 | 69.4814 | 72.6289 | 75.6837
-8.8 | 67.2373 | 67.9373 | 69.7891 | 72.932 | 75.9497
-8.7 | 67.5908 | 68.2493 | 70.1895 | 73.235 | 76.3909
-8.6 | 67.8952 | 68.5524 | 70.5381 | 73.5444 | 76.6963
-8.5 | 68.2977 | 68.8979 | 70.8548 | 73.8979 | 76.9697
-8.4 | 68.7011 | 69.287 | 71.2386 | 74.2017 | 77.2511
-8.3 | 68.9663 | 69.6051 | 71.5043 | 74.6038 | 77.5772
-8.2 | 69.2645 | 70.0115 | 71.8086 | 74.869 | 77.936
-8.1 | 69.6568 | 70.3819 | 72.1711 | 75.1676 | 78.2216
-8 | 70.0608 | 70.6849 | 72.5563 | 75.5608 | 78.4992
-7.9 | 70.4193 | 71.0506 | 72.8924 | 75.9649 | 78.8193
-7.8 | 70.8118 | 71.4228 | 73.3228 | 76.2767 | 79.1975
-7.7 | 71.1764 | 71.7764 | 73.6764 | 76.6264 | 79.5512
-7.6 | 71.5299 | 72.1628 | 74.0299 | 76.9399 | 79.8799
-7.5 | 71.8716 | 72.5669 | 74.3835 | 77.324 | 80.1557
-7.4 | 72.2161 | 72.9032 | 74.7494 | 77.646 | 80.446
-7.3 | 72.6406 | 73.275 | 75.1777 | 77.9906 | 80.7491
-7.2 | 73.0219 | 73.6442 | 75.5076 | 78.3093 | 81.1362
-7.1 | 73.426 | 74.026 | 75.8977 | 78.6123 | 81.4552
-7 | 73.83 | 74.43 | 76.3017 | 78.9513 | 81.7582
-6.9 | 74.2341 | 74.8341 | 76.6536 | 79.5225 | 82.1048
-6.8 | 74.6084 | 75.2381 | 77.0262 | 79.8313 | 82.4389
-6.7 | 74.9619 | 75.6119 | 77.4388 | 80.1245 | 82.7119
-6.6 | 75.3463 | 75.9891 | 77.7418 | 80.4275 | 83.0418
-6.5 | 75.7503 | 76.4415 | 78.1113 | 80.769 | 83.4028
-6.4 | 76.1544 | 76.8479 | 78.5634 | 81.218 | 83.6558
-6.3 | 76.5585 | 77.2013 | 78.9761 | 81.5938 | 84.0156
-6.2 | 77.0062 | 77.6054 | 79.3482 | 81.8969 | 84.3397
-6.1 | 77.4237 | 78.0094 | 79.7332 | 82.311 | 84.7237
-6 | 77.7868 | 78.4135 | 80.1135 | 82.6868 | 85.0458
-5.9 | 78.2434 | 78.8538 | 80.5175 | 83.0747 | 85.3904
-5.8 | 78.6787 | 79.2787 | 80.9585 | 83.4787 | 85.7091
-5.7 | 79.0475 | 79.7535 | 81.3475 | 83.8121 | 86.0828
-5.6 | 79.5012 | 80.151 | 81.744 | 84.0883 | 86.4152
-5.5 | 79.9061 | 80.5338 | 82.1046 | 84.4338 | 86.7909
-5.4 | 80.383 | 80.9775 | 82.4581 | 84.8973 | 87.0721
-5.3 | 80.8705 | 81.3991 | 82.9452 | 85.2411 | 87.3338
-5.2 | 81.2652 | 81.8789 | 83.3174 | 85.6031 | 87.7031
-5.1 | 81.6697 | 82.2688 | 83.7268 | 85.9304 | 88.0688
-5 | 82.1398 | 82.7176 | 84.2192 | 86.389 | 88.5192
-4.9 | 82.5574 | 83.2227 | 84.5345 | 86.7069 | 88.8207
-4.8 | 82.9908 | 83.6622 | 84.9908 | 87.0294 | 89.074
-4.7 | 83.4185 | 84.033 | 85.4106 | 87.3426 | 89.352
-4.6 | 83.8865 | 84.4132 | 85.8561 | 87.8093 | 89.7561
-4.5 | 84.2573 | 84.8535 | 86.2751 | 88.2315 | 90.1201
-4.4 | 84.7624 | 85.3481 | 86.707 | 88.6356 | 90.4436
-4.3 | 85.2674 | 85.8296 | 87.1166 | 88.9229 | 90.7833
-4.2 | 85.7677 | 86.2437 | 87.5677 | 89.3582 | 91.1439
-4.1 | 86.2192 | 86.7347 | 87.918 | 89.718 | 91.5596
-4 | 86.6348 | 87.1079 | 88.3438 | 90.0579 | 91.7809
-3.9 | 87.0877 | 87.4988 | 88.7152 | 90.3812 | 92.0702
-3.8 | 87.5866 | 87.9785 | 89.1928 | 90.7457 | 92.4557
-3.7 | 88.0354 | 88.5476 | 89.6587 | 91.1497 | 92.7587
-3.6 | 88.4744 | 88.9294 | 89.9721 | 91.5538 | 93.0721
-3.5 | 88.9256 | 89.4008 | 90.4094 | 91.8692 | 93.422
-3.4 | 89.3131 | 89.8723 | 90.8792 | 92.1536 | 93.7194
-3.3 | 89.8182 | 90.3284 | 91.2327 | 92.5436 | 93.9709
-3.2 | 90.3079 | 90.7947 | 91.6661 | 92.9363 | 94.274
-3.1 | 90.7455 | 91.2199 | 92.117 | 93.3313 | 94.577
-3 | 91.162 | 91.6048 | 92.4934 | 93.6934 | 94.9067
-2.9 | 91.6671 | 92.0469 | 92.847 | 93.9974 | 95.2646
-2.8 | 92.1577 | 92.4436 | 93.2006 | 94.4004 | 95.5003
-2.7 | 92.5721 | 92.9387 | 93.6046 | 94.6656 | 95.753
-2.6 | 93.0435 | 93.3801 | 94.0087 | 94.9637 | 96.0051
-2.5 | 93.4612 | 93.7842 | 94.4128 | 95.3556 | 96.2408
-2.4 | 93.8147 | 94.1647 | 94.8126 | 95.6698 | 96.5647
-2.3 | 94.2066 | 94.5495 | 95.1209 | 96.0183 | 96.8728
-2.2 | 94.6883 | 94.9535 | 95.5187 | 96.2553 | 97.1164
-2.1 | 95.0754 | 95.3254 | 95.8169 | 96.536 | 97.3815
-2 | 95.4616 | 95.7045 | 96.1053 | 96.8092 | 97.588
-1.9 | 95.7993 | 96.0825 | 96.5325 | 97.0993 | 97.79
-1.8 | 96.1698 | 96.436 | 96.824 | 97.4023 | 98.0103
-1.7 | 96.5396 | 96.7896 | 97.1738 | 97.6369 | 98.1264
-1.6 | 96.8655 | 97.1288 | 97.4389 | 97.8389 | 98.3466
-1.5 | 97.2248 | 97.3645 | 97.685 | 98.0234 | 98.54
-1.4 | 97.5717 | 97.6717 | 97.9717 | 98.2717 | 98.8144
-1.3 | 97.8747 | 97.9879 | 98.2747 | 98.5165 | 98.9179
-1.2 | 98.2074 | 98.2798 | 98.5185 | 98.7382 | 99.0307
-1.1 | 98.4739 | 98.5805 | 98.7582 | 98.9491 | 99.2109
-1 | 98.6572 | 98.7763 | 98.947 | 99.0756 | 99.4
-0.9 | 98.9297 | 98.996 | 99.096 | 99.234 | 99.4727
-0.8 | 99.1552 | 99.198 | 99.2572 | 99.3633 | 99.5905
-0.7 | 99.3572 | 99.35 | 99.4501 | 99.4643 | 99.6929
-0.6 | 99.4929 | 99.5021 | 99.5796 | 99.5572 | 99.7939
-0.5 | 99.5949 | 99.6456 | 99.6941 | 99.6521 | 99.8949
-0.4 | 99.7119 | 99.7388 | 99.7959 | 99.8062 | 99.9
-0.3 | 99.8298 | 99.8348 | 99.8848 | 99.9 | 99.9848
-0.2 | 99.9 | 99.9265 | 99.9 | 99.9482 | 100
-0.1 | 99.9325 | 100 | 99.9866 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 99.9321 | 100 | 100 | 99.9704 | 100
0.2 | 99.9 | 99.9265 | 99.9 | 99.9 | 100
0.3 | 99.8131 | 99.8348 | 99.8848 | 99.8684 | 100
0.4 | 99.7 | 99.7388 | 99.7959 | 99.7714 | 99.9102
0.5 | 99.6613 | 99.6456 | 99.6941 | 99.6806 | 99.9
0.6 | 99.4762 | 99.5021 | 99.5796 | 99.5796 | 99.8197
0.7 | 99.3286 | 99.35 | 99.4501 | 99.4786 | 99.7215
0.8 | 99.1266 | 99.198 | 99.2572 | 99.3804 | 99.6072
0.9 | 98.9246 | 98.996 | 99.096 | 99.2817 | 99.4893
1 | 98.6838 | 98.7763 | 98.947 | 99.0898 | 99.3715
1.1 | 98.4205 | 98.5805 | 98.7582 | 98.9777 | 99.2073
1.2 | 98.1859 | 98.2798 | 98.5185 | 98.7756 | 99.045
1.3 | 97.889 | 97.9879 | 98.2747 | 98.5604 | 98.9346
1.4 | 97.5859 | 97.6717 | 97.9717 | 98.2716 | 98.7751
1.5 | 97.2467 | 97.3645 | 97.685 | 98.0695 | 98.5257
1.6 | 96.8636 | 97.1288 | 97.4389 | 97.8513 | 98.2818
1.7 | 96.5196 | 96.7896 | 97.1738 | 97.5655 | 98.1299
1.8 | 96.186 | 96.436 | 96.824 | 97.368 | 97.9492
1.9 | 95.8564 | 96.0825 | 96.5325 | 97.18 | 97.8108
2 | 95.5045 | 95.7045 | 96.1053 | 96.788 | 97.6165
2.1 | 95.0505 | 95.3254 | 95.8169 | 96.5754 | 97.3565
2.2 | 94.5964 | 94.9535 | 95.5187 | 96.2218 | 97.1411
2.3 | 94.1923 | 94.5495 | 95.1209 | 96.0104 | 96.8183
2.4 | 93.753 | 94.1647 | 94.8126 | 95.7126 | 96.6084
2.5 | 93.3612 | 93.7842 | 94.4128 | 95.3128 | 96.3096
2.6 | 93.0076 | 93.3801 | 94.0087 | 95.0432 | 96.0208
2.7 | 92.5058 | 92.9387 | 93.6046 | 94.7463 | 95.7452
2.8 | 92.1293 | 92.4436 | 93.2006 | 94.4433 | 95.5431
2.9 | 91.6337 | 92.0469 | 92.847 | 93.9671 | 95.2969
3 | 91.1924 | 91.6048 | 92.4934 | 93.6372 | 94.9434
3.1 | 90.7855 | 91.2199 | 92.117 | 93.3342 | 94.6199
3.2 | 90.2829 | 90.7947 | 91.6661 | 92.9415 | 94.3168
3.3 | 89.8409 | 90.3284 | 91.2327 | 92.5827 | 94.0138
3.4 | 89.3417 | 89.8723 | 90.8792 | 92.2393 | 93.7107
3.5 | 88.9293 | 89.4008 | 90.4094 | 91.915 | 93.4077
3.6 | 88.4315 | 88.9294 | 89.9721 | 91.5109 | 93.0395
3.7 | 88.0685 | 88.5476 | 89.6587 | 91.1302 | 92.7016
3.8 | 87.6199 | 87.9785 | 89.1928 | 90.7937 | 92.4237
3.9 | 87.1163 | 87.4988 | 88.7152 | 90.3669 | 92.0845
4 | 86.6112 | 87.1079 | 88.3438 | 90.0639 | 91.7902
4.1 | 86.1062 | 86.7347 | 87.918 | 89.7478 | 91.5323
4.2 | 85.7008 | 86.2437 | 87.5677 | 89.3508 | 91.2008
4.3 | 85.296 | 85.8296 | 87.1166 | 88.9968 | 90.8254
4.4 | 84.7909 | 85.3481 | 86.707 | 88.5927 | 90.466
4.5 | 84.2887 | 84.8535 | 86.2751 | 88.1859 | 90.1173
4.6 | 83.8832 | 84.4132 | 85.8561 | 87.6899 | 89.7132
4.7 | 83.4412 | 84.033 | 85.4106 | 87.4247 | 89.333
4.8 | 83.0337 | 83.6622 | 84.9908 | 87.0824 | 89.0196
4.9 | 82.537 | 83.2227 | 84.5345 | 86.7724 | 88.7839
5 | 82.1255 | 82.7176 | 84.2192 | 86.3605 | 88.4723
5.1 | 81.686 | 82.2688 | 83.7268 | 85.8732 | 88.0786
5.2 | 81.3152 | 81.8789 | 83.3174 | 85.5602 | 87.7152
5.3 | 80.8521 | 81.3991 | 82.9452 | 85.1744 | 87.41
5.4 | 80.4102 | 80.9775 | 82.4581 | 84.8687 | 87.107
5.5 | 79.9394 | 80.5338 | 82.1046 | 84.4284 | 86.7481
5.6 | 79.468 | 80.151 | 81.744 | 84.158 | 86.358
5.7 | 79.0257 | 79.7535 | 81.3475 | 83.84 | 86.04
5.8 | 78.6662 | 79.2787 | 80.9585 | 83.4359 | 85.6579
5.9 | 78.2719 | 78.8538 | 80.5175 | 83.0404 | 85.3917
6 | 77.8135 | 78.4135 | 80.1135 | 82.6849 | 85.0849
6.1 | 77.3618 | 78.0094 | 79.7332 | 82.2856 | 84.6856
6.2 | 76.954 | 77.6054 | 79.3482 | 81.9729 | 84.3797
6.3 | 76.5682 | 77.2013 | 78.9761 | 81.5367 | 84.0261
6.4 | 76.1401 | 76.8479 | 78.5634 | 81.2226 | 83.7194
6.5 | 75.769 | 76.4415 | 78.1113 | 80.8646 | 83.3794
6.6 | 75.3891 | 75.9891 | 77.7418 | 80.4704 | 83.1034
6.7 | 74.9851 | 75.6119 | 77.4388 | 80.1619 | 82.7413
6.8 | 74.581 | 75.2381 | 77.0262 | 79.8084 | 82.4584
6.9 | 74.2048 | 74.8341 | 76.6536 | 79.5041 | 82.1048
7 | 73.8157 | 74.43 | 76.3017 | 78.9022 | 81.7513
7.1 | 73.4477 | 74.026 | 75.8977 | 78.5701 | 81.4974
7.2 | 73.0647 | 73.6442 | 75.5076 | 78.3441 | 81.0942
7.3 | 72.6208 | 73.275 | 75.1777 | 77.9919 | 80.7406
7.4 | 72.2371 | 72.9032 | 74.7494 | 77.6852 | 80.5212
7.5 | 71.9287 | 72.5669 | 74.3835 | 77.2811 | 80.1835
7.6 | 71.4866 | 72.1628 | 74.0299 | 76.885 | 79.8299
7.7 | 71.2264 | 71.7764 | 73.6764 | 76.6198 | 79.5073
7.8 | 70.8547 | 71.4228 | 73.3228 | 76.2261 | 79.2228
7.9 | 70.4506 | 71.0506 | 72.8924 | 75.9165 | 78.8506
8 | 70.0465 | 70.6849 | 72.5563 | 75.518 | 78.4849
8.1 | 69.6622 | 70.3819 | 72.1711 | 75.2466 | 78.1819
8.2 | 69.3217 | 70.0115 | 71.8086 | 74.9502 | 77.8788
8.3 | 68.9915 | 69.6051 | 71.5043 | 74.5629 | 77.5913
8.4 | 68.6156 | 69.287 | 71.2386 | 74.2761 | 77.2874
8.5 | 68.298 | 68.8979 | 70.8548 | 73.912 | 76.9417
8.6 | 67.9365 | 68.5524 | 70.5381 | 73.5809 | 76.6944
8.7 | 67.5493 | 68.2493 | 70.1895 | 73.2038 | 76.3409
8.8 | 67.2284 | 67.9373 | 69.7891 | 72.903 | 75.9873
8.9 | 66.8337 | 67.5671 | 69.4814 | 72.6146 | 75.6338
9 | 66.483 | 67.1973 | 69.0802 | 72.2488 | 75.383
9.1 | 66.18 | 66.859 | 68.7267 | 71.9085 | 75.0766
9.2 | 65.8769 | 66.5048 | 68.5132 | 71.6055 | 74.7487
9.3 | 65.5695 | 66.208 | 68.1223 | 71.3271 | 74.4695
9.4 | 65.216 | 65.828 | 67.8137 | 71.0423 | 74.1423
9.5 | 64.9106 | 65.5249 | 67.4624 | 70.6856 | 73.8392
9.6 | 64.5648 | 65.2089 | 67.1505 | 70.3362 | 73.5362
9.7 | 64.2617 | 64.8915 | 66.8474 | 70.0053 | 73.254
9.8 | 63.9518 | 64.6018 | 66.5018 | 69.7153 | 72.9729
9.9 | 63.6321 | 64.2482 | 66.1982 | 69.4699 | 72.6699
10 | 63.3446 | 63.9631 | 65.8631 | 69.1835 | 72.3669
10.1 | 63.0209 | 63.641 | 65.6067 | 68.8756 | 72.0639
10.2 | 62.7179 | 63.2876 | 65.3036 | 68.5358 | 71.7782
10.3 | 62.3531 | 62.9622 | 64.984 | 68.284 | 71.5006
10.4 | 62.0118 | 62.7204 | 64.6304 | 67.9403 | 71.1976
10.5 | 61.7326 | 62.3373 | 64.3088 | 67.6582 | 70.8769
10.6 | 61.4485 | 62.0562 | 64.0705 | 67.4342 | 70.599
10.7 | 61.1455 | 61.8198 | 63.7697 | 67.1398 | 70.3799
10.8 | 60.8424 | 61.4663 | 63.4567 | 66.8612 | 70.0662
10.9 | 60.5394 | 61.1501 | 63.1536 | 66.4679 | 69.7127
11 | 60.2568 | 60.9221 | 62.9061 | 66.1648 | 69.5053
11.1 | 59.9761 | 60.6055 | 62.6048 | 65.8777 | 69.1703
11.2 | 59.673 | 60.273 | 62.2519 | 65.7011 | 68.9161
11.3 | 59.3989 | 59.9863 | 61.9492 | 65.2988 | 68.613
11.4 | 59.1242 | 59.74 | 61.7527 | 64.9972 | 68.34
11.5 | 58.8212 | 59.4914 | 61.4497 | 64.7927 | 68.107
11.6 | 58.5585 | 59.1919 | 61.1533 | 64.4931 | 67.8159
11.7 | 58.2721 | 58.9152 | 60.8841 | 64.2841 | 67.5343
11.8 | 57.9981 | 58.5807 | 60.5538 | 63.9307 | 67.2404
11.9 | 57.7091 | 58.344 | 60.2772 | 63.7156 | 67.0376
12 | 57.4707 | 58.063 | 59.9824 | 63.4203 | 66.7491
12.1 | 57.2276 | 57.8134 | 59.7467 | 63.1173 | 66.4885
12.2 | 56.9427 | 57.4776 | 59.5111 | 62.8499 | 66.2236
12.3 | 56.6397 | 57.2503 | 59.2112 | 62.5539 | 65.9629
12.4 | 56.3911 | 57.0224 | 59.0055 | 62.3339 | 65.6483
12.5 | 56.1544 | 56.7194 | 58.6706 | 62.0479 | 65.4059
12.6 | 55.8733 | 56.4268 | 58.4349 | 61.7448 | 65.1015
12.7 | 55.6135 | 56.159 | 58.1703 | 61.5279 | 64.8851
12.8 | 55.3838 | 55.9225 | 57.8838 | 61.2388 | 64.6245
12.9 | 55.1277 | 55.707 | 57.6277 | 60.9358 | 64.3611
13 | 54.892 | 55.4253 | 57.3785 | 60.7218 | 64.0755
13.1 | 54.6563 | 55.1582 | 57.1009 | 60.4634 | 63.8483
13.2 | 54.3981 | 54.9034 | 56.8319 | 60.1694 | 63.5694
13.3 | 54.1091 | 54.7014 | 56.6299 | 59.9442 | 63.3182
13.4 | 53.8707 | 54.4238 | 56.3739 | 59.6633 | 63.1135
13.5 | 53.6469 | 54.1469 | 56.0974 | 59.4601 | 62.8173
13.6 | 53.4111 | 53.9584 | 55.8626 | 59.2 | 62.581
13.7 | 53.1755 | 53.697 | 55.6218 | 58.9646 | 62.3632
13.8 | 52.9224 | 53.4626 | 55.3797 | 58.6939 | 62.0397
13.9 | 52.6373 | 53.1909 | 55.0767 | 58.3909 | 61.8178
14 | 52.4017 | 52.9143 | 54.8263 | 58.1762 | 61.6158
14.1 | 52.166 | 52.6851 | 54.6422 | 57.966 | 61.3489
14.2 | 51.9302 | 52.4245 | 54.3954 | 57.7545 | 61.0245
14.3 | 51.693 | 52.2306 | 54.1097 | 57.4788 | 60.7959
14.4 | 51.3942 | 51.9758 | 53.9077 | 57.2505 | 60.5922
14.5 | 51.2173 | 51.7485 | 53.6899 | 57.0011 | 60.3565
14.6 | 51.0178 | 51.4984 | 53.4542 | 56.7749 | 60.1465
14.7 | 50.8158 | 51.3259 | 53.2184 | 56.5094 | 59.8667
14.8 | 50.5206 | 51.116 | 53.0137 | 56.2709 | 59.5636
14.9 | 50.3117 | 50.8176 | 52.6902 | 56.0279 | 59.3404
15 | 50.0812 | 50.6097 | 52.5288 | 55.7954 | 59.0573
15.1 | 49.8818 | 50.4077 | 52.2505 | 55.5933 | 58.8879
15.2 | 49.705 | 50.2057 | 52.0485 | 55.337 | 58.7399
15.3 | 49.5036 | 49.9422 | 51.8197 | 55.1446 | 58.477
15.4 | 49.2594 | 49.7019 | 51.5514 | 54.9521 | 58.216
15.5 | 48.9661 | 49.4996 | 51.3661 | 54.6708 | 57.9708
15.6 | 48.7652 | 49.2976 | 51.1404 | 54.4117 | 57.7117
15.7 | 48.5947 | 49.0956 | 50.9384 | 54.1645 | 57.4947
15.8 | 48.3792 | 48.8925 | 50.7364 | 53.9193 | 57.2792
15.9 | 48.1772 | 48.6628 | 50.5138 | 53.7343 | 57.0772
16 | 47.9542 | 48.4543 | 50.2608 | 53.5323 | 56.8752
16.1 | 47.7444 | 48.2303 | 50.0382 | 53.3303 | 56.6519
16.2 | 47.5424 | 48.0283 | 49.7496 | 53.0924 | 56.4424
16.3 | 47.3604 | 47.8139 | 49.5631 | 52.8032 | 56.2404
16.4 | 47.1671 | 47.6086 | 49.4116 | 52.5527 | 55.988
16.5 | 46.8637 | 47.422 | 49.1935 | 52.3754 | 55.7651
16.6 | 46.6343 | 47.2067 | 48.9915 | 52.2487 | 55.4945
16.7 | 46.5162 | 46.9895 | 48.771 | 52.0066 | 55.3304
16.8 | 46.3303 | 46.7875 | 48.5353 | 51.751 | 55.0882
16.9 | 46.1497 | 46.5855 | 48.2998 | 51.5282 | 54.8747
17 | 45.8821 | 46.398 | 48.0979 | 51.298 | 54.6834
17.1 | 45.6527 | 46.2099 | 47.9212 | 51.1099 | 54.4221
17.2 | 45.513 | 46.0079 | 47.7222 | 50.9079 | 54.1944
17.3 | 45.2915 | 45.8059 | 47.5202 | 50.6588 | 54.0059
17.4 | 45.0895 | 45.6158 | 47.3182 | 50.4439 | 53.8038
17.5 | 44.8875 | 45.4303 | 47.1162 | 50.3261 | 53.5711
17.6 | 44.6873 | 45.2283 | 46.9141 | 50.0639 | 53.3283
17.7 | 44.5106 | 45.0263 | 46.7355 | 49.8406 | 53.157
17.8 | 44.3099 | 44.7864 | 46.5116 | 49.6587 | 52.9783
17.9 | 44.1079 | 44.5547 | 46.3079 | 49.3976 | 52.6976
18 | 43.9774 | 44.3917 | 46.1059 | 49.2201 | 52.5315
18.1 | 43.7754 | 44.1897 | 45.9038 | 49.0325 | 52.3379
18.2 | 43.5767 | 43.9877 | 45.7018 | 48.8305 | 52.0775
18.3 | 43.3998 | 43.8333 | 45.4999 | 48.6285 | 51.857
18.4 | 43.1821 | 43.6309 | 45.3693 | 48.4481 | 51.6549
18.5 | 43.0284 | 43.4958 | 45.1673 | 48.2529 | 51.4529
18.6 | 42.8195 | 43.2945 | 44.9513 | 47.9764 | 51.2261
18.7 | 42.6237 | 43.1177 | 44.7452 | 47.8244 | 51.0202
18.8 | 42.452 | 42.9182 | 44.5547 | 47.666 | 50.841
18.9 | 42.302 | 42.6938 | 44.4094 | 47.4733 | 50.7224
19 | 42.1 | 42.5833 | 44.1857 | 47.207 | 50.4642
19.1 | 41.898 | 42.3549 | 43.9857 | 46.9693 | 50.3204
19.2 | 41.7393 | 42.1529 | 43.8088 | 46.7673 | 50.0883
19.3 | 41.5653 | 41.9755 | 43.6762 | 46.6214 | 49.8429
19.4 | 41.4316 | 41.8233 | 43.4404 | 46.4553 | 49.6347
19.5 | 41.2612 | 41.5768 | 43.2184 | 46.2612 | 49.5163
19.6 | 41.0592 | 41.4735 | 43.0582 | 46.0592 | 49.2458
19.7 | 40.8572 | 41.2715 | 42.8715 | 45.8333 | 49.1142
19.8 | 40.7366 | 41.0846 | 42.6562 | 45.698 | 48.8642
19.9 | 40.562 | 40.9714 | 42.5103 | 45.4939 | 48.6643
20 | 40.3511 | 40.7939 | 42.3082 | 45.1973 | 48.4929
20.1 | 40.1491 | 40.5919 | 42.1061 | 45.1089 | 48.2572
20.2 | 40.0235 | 40.3941 | 41.904 | 44.9184 | 48.1107
20.3 | 39.8191 | 40.2524 | 41.751 | 44.7164 | 47.8858
20.4 | 39.7072 | 40.0584 | 41.6125 | 44.5143 | 47.6572
20.5 | 39.5123 | 39.881 | 41.4267 | 44.2919 | 47.4552
20.6 | 39.3103 | 39.7226 | 41.2246 | 44.1761 | 47.2531
20.7 | 39.1322 | 39.5097 | 41.0226 | 43.9109 | 47.043
20.8 | 39.0174 | 39.3869 | 40.8603 | 43.7063 | 46.8537
20.9 | 38.8327 | 39.2471 | 40.705 | 43.6021 | 46.6899
21 | 38.7013 | 39.0519 | 40.5346 | 43.4022 | 46.5019
21.1 | 38.5002 | 38.8715 | 40.3335 | 43.2251 | 46.3572
21.2 | 38.2984 | 38.6695 | 40.2133 | 43.0267 | 46.2123
21.3 | 38.1216 | 38.4837 | 39.9698 | 42.8246 | 45.9954
21.4 | 38.0113 | 38.3698 | 39.8256 | 42.6226 | 45.7798
21.5 | 37.8206 | 38.1953 | 39.6491 | 42.4206 | 45.5778
21.6 | 37.6412 | 38.0614 | 39.4663 | 42.3206 | 45.3758
21.7 | 37.5225 | 37.8594 | 39.2737 | 42.1451 | 45.2263
21.8 | 37.317 | 37.6813 | 39.0717 | 41.943 | 45.043
21.9 | 37.2063 | 37.5838 | 38.9348 | 41.761 | 44.861
22 | 37.0818 | 37.3909 | 38.7842 | 41.5675 | 44.7228
22.1 | 36.8764 | 37.2798 | 38.5941 | 41.3655 | 44.537
22.2 | 36.6746 | 37.087 | 38.3921 | 41.1634 | 44.3349
22.3 | 36.5473 | 36.9384 | 38.2451 | 40.9788 | 44.1329
22.4 | 36.368 | 36.7514 | 38.0881 | 40.844 | 43.9309
22.5 | 36.2004 | 36.5861 | 37.9003 | 40.6861 | 43.8001
22.6 | 36.0323 | 36.384 | 37.7563 | 40.484 | 43.5979
22.7 | 35.9267 | 36.241 | 37.6105 | 40.341 | 43.3811
22.8 | 35.7266 | 36.095 | 37.3932 | 40.195 | 43.237
22.9 | 35.5921 | 35.9065 | 37.2389 | 40.0065 | 43.0349
23 | 35.3951 | 35.7522 | 37.0758 | 39.7567 | 42.8609
23.1 | 35.2881 | 35.6024 | 36.9448 | 39.5512 | 42.6861
23.2 | 35.0837 | 35.4128 | 36.8004 | 39.4564 | 42.4878
23.3 | 34.8778 | 35.2488 | 36.5984 | 39.3269 | 42.2984
23.4 | 34.7729 | 35.1229 | 36.3964 | 39.1248 | 42.1395
23.5 | 34.6228 | 34.9372 | 36.2075 | 38.9324 | 42.0217
23.6 | 34.4208 | 34.8176 | 36.0604 | 38.7493 | 41.8352
23.7 | 34.2594 | 34.6052 | 35.9188 | 38.5473 | 41.6331
23.8 | 34.1271 | 34.5012 | 35.7028 | 38.3453 | 41.4311
23.9 | 33.9338 | 34.3004 | 35.5336 | 38.1716 | 41.2291
24 | 33.7564 | 34.1984 | 35.4157 | 38.0486 | 41.0486
24.1 | 33.6107 | 33.9958 | 35.2248 | 37.8679 | 40.9268
24.2 | 33.4087 | 33.7658 | 35.0228 | 37.6658 | 40.7267
24.3 | 33.2067 | 33.5818 | 34.891 | 37.4636 | 40.5208
24.4 | 33.0046 | 33.4616 | 34.6808 | 37.2832 | 40.3952
24.5 | 32.8025 | 33.2529 | 34.5529 | 37.1883 | 40.2824
24.6 | 32.5693 | 33.0586 | 34.3172 | 36.9863 | 39.9931
24.7 | 32.2906 | 32.8984 | 34.1492 | 36.7842 | 39.8862
24.8 | 31.9501 | 32.6964 | 33.9964 | 36.5922 | 39.7094
24.9 | 31.5076 | 32.4415 | 33.7943 | 36.5038 | 39.5087
25 | 30.884 | 32.2462 | 33.6372 | 36.3067 | 39.4033
25.1 | 29.9745 | 32.0081 | 33.479 | 36.1046 | 39.2046
25.2 | 28.7364 | 31.7598 | 33.2883 | 36.0013 | 39.0272
25.3 | 26.9061 | 31.451 | 33.0863 | 35.8006 | 38.8291
25.4 | 24.6384 | 31.0975 | 32.8842 | 35.5988 | 38.7135
25.5 | 21.9515 | 30.7075 | 32.6822 | 35.4219 | 38.5251
25.6 | 19.0919 | 30.2289 | 32.4802 | 35.3115 | 38.323
25.7 | 16.1949 | 29.4227 | 32.3289 | 35.121 | 38.1433
25.8 | 13.5997 | 28.2848 | 32.1475 | 34.919 | 38.0111
25.9 | 11.4312 | 26.9187 | 31.9455 | 34.717 | 37.817
26 | 9.4051 | 25.0573 | 31.7174 | 34.538 | 37.6149
26.1 | 7.9909 | 22.7412 | 31.5363 | 34.4207 | 37.4129
26.2 | 6.7259 | 20.1851 | 31.2591 | 34.2394 | 37.2109
26.3 | 5.7859 | 17.5588 | 31.0104 | 34.0578 | 37.1044
26.4 | 4.9485 | 15.1529 | 30.8069 | 33.864 | 36.9746
26.5 | 4.3833 | 12.8956 | 30.4097 | 33.7309 | 36.7695
26.6 | 3.9066 | 10.8585 | 30.1042 | 33.5365 | 36.6516
26.7 | 3.6385 | 9.1391 | 29.7016 | 33.3293 | 36.4865
26.8 | 3.3409 | 7.6922 | 29.2103 | 33.1273 | 36.2651
26.9 | 3.1253 | 6.588 | 28.5723 | 33.0126 | 36.1269
27 | 3.0116 | 5.8813 | 27.6964 | 32.8232 | 35.9271
27.1 | 2.7963 | 5.1742 | 26.4483 | 32.6435 | 35.7212
27.2 | 2.6953 | 4.648 | 24.9675 | 32.4668 | 35.6096
27.3 | 2.6 | 4.2343 | 23.0792 | 32.3371 | 35.3942
27.4 | 2.579 | 3.9227 | 20.9578 | 32.1721 | 35.2865
27.5 | 2.4743 | 3.6197 | 18.5763 | 31.9052 | 35.0844
27.6 | 2.4 | 3.4109 | 16.3148 | 31.734 | 34.8898
27.7 | 2.3616 | 3.2771 | 14.1943 | 31.5438 | 34.7519
27.8 | 2.2606 | 3.0708 | 12.2378 | 31.4178 | 34.5781
27.9 | 2.2 | 2.9596 | 10.6515 | 31.2335 | 34.4645
28 | 2.1637 | 2.8517 | 9.2198 | 31.0315 | 34.1872
28.1 | 2.1 | 2.8 | 8.1755 | 30.8508 | 34.0838
28.2 | 2.0708 | 2.7326 | 7.296 | 30.656 | 33.9319
28.3 | 2 | 2.627 | 6.6239 | 30.4539 | 33.7481
28.4 | 1.9688 | 2.526 | 6.0347 | 30.2519 | 33.5804
28.5 | 1.9 | 2.5 | 5.5164 | 29.9748 | 33.4392
28.6 | 1.9 | 2.4239 | 5.1336 | 29.7668 | 33.2764
28.7 | 1.8576 | 2.3229 | 4.8258 | 29.5743 | 33.1267
28.8 | 1.8 | 2.3 | 4.5867 | 29.3085 | 32.9177
28.9 | 1.8 | 2.2352 | 4.3846 | 29.0703 | 32.7131
29 | 1.7627 | 2.2 | 4.1899 | 28.7366 | 32.6056
29.1 | 1.7 | 2.1331 | 4.072 | 28.4244 | 32.4091
29.2 | 1.7 | 2.1 | 3.9541 | 28.0855 | 32.2313
29.3 | 1.6454 | 2.0397 | 3.8363 | 27.6268 | 32.0336
29.4 | 1.6 | 2 | 3.7184 | 27.0104 | 31.9018
29.5 | 1.6 | 1.9433 | 3.6006 | 26.2467 | 31.7679
29.6 | 1.6 | 1.9 | 3.5 | 25.2274 | 31.5322
29.7 | 1.5127 | 1.9 | 3.4649 | 23.9519 | 31.3482
29.8 | 1.5 | 1.8403 | 3.3546 | 22.4997 | 31.1806
29.9 | 1.5 | 1.8 | 3.3 | 20.868 | 31.0292
30 | 1.5 | 1.8 | 3.2113 | 19.0274 | 30.8481
30.1 | 1.423 | 1.7087 | 3.1 | 17.1709 | 30.6459
30.2 | 1.4 | 1.7 | 3.0756 | 15.4165 | 30.4438
30.3 | 1.4 | 1.7 | 3 | 13.8288 | 30.3209
30.4 | 1.4 | 1.6913 | 2.9566 | 12.4942 | 30.1398
30.5 | 1.4 | 1.6 | 2.8388 | 11.3222 | 29.9108
30.6 | 1.3322 | 1.6 | 2.8 | 10.2646 | 29.8036
30.7 | 1.3 | 1.6 | 2.7312 | 9.3636 | 29.5727
30.8 | 1.3 | 1.5764 | 2.7 | 8.6111 | 29.3461
30.9 | 1.3 | 1.5 | 2.6173 | 7.9606 | 29.172
31 | 1.3 | 1.5 | 2.6 | 7.448 | 29.0323
31.1 | 1.2362 | 1.5 | 2.5 | 7.0226 | 28.8113
31.2 | 1.2 | 1.5 | 2.5 | 6.6639 | 28.5942
31.3 | 1.2 | 1.4793 | 2.4793 | 6.3609 | 28.3182
31.4 | 1.2 | 1.4 | 2.4 | 6.1053 | 28.0152
31.5 | 1.2 | 1.4 | 2.3659 | 5.9032 | 27.7747
31.6 | 1.2 | 1.4 | 2.3 | 5.7135 | 27.4772
31.7 | 1.1496 | 1.3782 | 2.3 | 5.5638 | 27.0984
31.8 | 1.1 | 1.3 | 2.2628 | 5.4257 | 26.6429
31.9 | 1.1 | 1.3 | 2.2 | 5.2618 | 26.033
32 | 1.1 | 1.3 | 2.2 | 5.1608 | 25.2191
32.1 | 1.1 | 1.3 | 2.1456 | 5.0598 | 24.31
32.2 | 1.1 | 1.3 | 2.1 | 4.964 | 23.1231
32.3 | 1.1 | 1.3 | 2.1 | 4.8675 | 21.7087
32.4 | 1.0372 | 1.2425 | 2.0329 | 4.7568 | 20.2951
32.5 | 1 | 1.2 | 2 | 4.6558 | 18.7642
32.6 | 1 | 1.2 | 2 | 4.5547 | 17.0876
32.7 | 1 | 1.2 | 1.9966 | 4.4537 | 15.4909
32.8 | 1 | 1.2 | 1.9 | 4.4 | 14.1097
32.9 | 1 | 1.2 | 1.9 | 4.3517 | 12.8776
33 | 1 | 1.2 | 1.8925 | 4.2425 | 11.8501
33.1 | 1 | 1.1354 | 1.8 | 4.2 | 10.9182
33.2 | 1 | 1.1 | 1.8 | 4.1201 | 10.1404
33.3 | 0.93333 | 1.1 | 1.8 | 4.0191 | 9.5
33.4 | 0.9 | 1.1 | 1.8 | 4 | 8.8939
33.5 | 0.9 | 1.1 | 1.7313 | 3.9313 | 8.3796
33.6 | 0.9 | 1.1 | 1.7 | 3.8303 | 8.0482
33.7 | 0.9 | 1.1 | 1.7 | 3.8 | 7.7025
33.8 | 0.9 | 1.1 | 1.7 | 3.7283 | 7.3994
33.9 | 0.9 | 1.0273 | 1.6985 | 3.7 | 7.1969
34 | 0.9 | 1 | 1.6 | 3.6405 | 6.9668
34.1 | 0.9 | 1 | 1.6 | 3.5395 | 6.7646
34.2 | 0.9 | 1 | 1.6 | 3.5 | 6.5837
34.3 | 0.80889 | 1 | 1.6 | 3.4271 | 6.4946
34.4 | 0.8 | 1 | 1.6 | 3.4 | 6.3873
34.5 | 0.8 | 1 | 1.5 | 3.3212 | 6.1926
34.6 | 0.8 | 1 | 1.5 | 3.3 | 6.0916
34.7 | 0.8 | 1 | 1.5 | 3.2191 | 5.989
34.8 | 0.8 | 0.90384 | 1.5 | 3.2 | 5.8753
34.9 | 0.8 | 0.9 | 1.5 | 3.2 | 5.7742
35 | 0.8 | 0.9 | 1.4 | 3.1187 | 5.7
35.1 | 0.8 | 0.9 | 1.4 | 3.1 | 5.6509
35.2 | 0.8 | 0.9 | 1.4 | 3 | 5.5426
35.3 | 0.8 | 0.9 | 1.4 | 3 | 5.4416
35.4 | 0.78343 | 0.9 | 1.4 | 2.9974 | 5.3406
35.5 | 0.7 | 0.9 | 1.4 | 2.9 | 5.3
35.6 | 0.7 | 0.9 | 1.395 | 2.9 | 5.2462
35.7 | 0.7 | 0.9 | 1.3 | 2.8804 | 5.1518
35.8 | 0.7 | 0.9 | 1.3 | 2.8 | 5.1
35.9 | 0.7 | 0.89263 | 1.3 | 2.7784 | 5.0356
36 | 0.7 | 0.8 | 1.3 | 2.7 | 4.9236
36.1 | 0.7 | 0.8 | 1.3 | 2.7 | 4.8192
36.2 | 0.7 | 0.8 | 1.3 | 2.7 | 4.8
36.3 | 0.7 | 0.8 | 1.2743 | 2.6743 | 4.7172
36.4 | 0.7 | 0.8 | 1.2 | 2.6 | 4.7
36.5 | 0.7 | 0.8 | 1.2 | 2.6 | 4.6011
36.6 | 0.7 | 0.8 | 1.2 | 2.5713 | 4.5
36.7 | 0.7 | 0.8 | 1.2 | 2.5 | 4.4988
36.8 | 0.7 | 0.8 | 1.2 | 2.5 | 4.4
36.9 | 0.7 | 0.8 | 1.2 | 2.4722 | 4.3962
37 | 0.68152 | 0.8 | 1.2 | 2.4 | 4.3
37.1 | 0.6 | 0.8 | 1.1954 | 2.4 | 4.2605
37.2 | 0.6 | 0.8 | 1.1071 | 2.4 | 4.2
37.3 | 0.6 | 0.77488 | 1.1 | 2.3785 | 4.1249
37.4 | 0.6 | 0.7 | 1.1 | 2.3 | 4.1
37.5 | 0.6 | 0.7 | 1.1 | 2.3 | 4
37.6 | 0.6 | 0.7 | 1.1 | 2.3 | 4
37.7 | 0.6 | 0.7 | 1.1 | 2.2743 | 3.9701
37.8 | 0.6 | 0.7 | 1.1 | 2.2 | 3.9
37.9 | 0.6 | 0.7 | 1.1 | 2.2 | 3.9
38 | 0.6 | 0.7 | 1.1 | 2.2 | 3.8
38.1 | 0.6 | 0.7 | 1.0989 | 2.1561 | 3.7821
38.2 | 0.6 | 0.7 | 1 | 2.1 | 3.7
38.3 | 0.6 | 0.7 | 1 | 2.1 | 3.7
38.4 | 0.6 | 0.7 | 1 | 2.1 | 3.6285
38.5 | 0.6 | 0.7 | 1 | 2.052 | 3.6
38.6 | 0.6 | 0.7 | 1 | 2 | 3.6
38.7 | 0.6 | 0.7 | 1 | 2 | 3.525
38.8 | 0.6 | 0.7 | 1 | 2 | 3.5
38.9 | 0.6 | 0.7 | 1 | 2 | 3.4893
39 | 0.6 | 0.7 | 1 | 2 | 3.4
39.1 | 0.6 | 0.7 | 1 | 1.9316 | 3.3535
39.2 | 0.54485 | 0.6 | 0.90202 | 1.9 | 3.3
39.3 | 0.5 | 0.6 | 0.9 | 1.9 | 3.3
39.4 | 0.5 | 0.6 | 0.9 | 1.9 | 3.3
39.5 | 0.5 | 0.6 | 0.9 | 1.9 | 3.2418
39.6 | 0.5 | 0.6 | 0.9 | 1.8266 | 3.2
39.7 | 0.5 | 0.6 | 0.9 | 1.8 | 3.2
39.8 | 0.5 | 0.6 | 0.9 | 1.8 | 3.153
39.9 | 0.5 | 0.6 | 0.9 | 1.8 | 3.1
40 | 0.5 | 0.6 | 0.9 | 1.8 | 3.1
40.1 | 0.5 | 0.6 | 0.9 | 1.8 | 3.025
40.2 | 0.5 | 0.6 | 0.9 | 1.7347 | 3
40.3 |  | 0.6 | 0.9 | 1.7 | 3
40.4 |  | 0.6 |  | 1.7 | 2.9715
40.5 |  | 0.6 |  | 1.7 | 2.9
40.6 |  | 0.6 |  | 1.7 | 2.9
40.7 |  | 0.6 |  | 1.7 | 2.8439
40.8 |  | 0.6 |  | 1.6286 | 2.8
40.9 |  | 0.6 |  | 1.6 | 2.8
41 |  |  |  | 1.6 | 2.8
41.1 |  |  |  | 1.6 | 2.8
41.2 |  |  |  | 1.6 | 2.7245
41.3 |  |  |  | 1.6 | 2.7
41.4 |  |  |  | 1.6 | 2.7
41.5 |  |  |  | 1.5359 | 2.6314
41.6 |  |  |  | 1.5 | 2.6
41.7 |  |  |  | 1.5 | 2.6
41.8 |  |  |  | 1.5 | 2.6
41.9 |  |  |  | 1.5 | 2.6
42 |  |  |  | 1.5 | 2.5021
42.1 |  |  |  | 1.5 | 2.5
42.2 |  |  |  | 1.5
42.3 |  |  |  | 1.4277
42.4 |  |  |  | 1.4
42.5 |  |  |  | 1.4
42.6 |  |  |  | 1.4
42.7 |  |  |  | 1.4
42.8 |  |  |  | 1.4
42.9 |  |  |  | 1.4
43 |  |  |  | 1.4
43.1 |  |  |  | 1.4
43.2 |  |  |  | 1.3328
43.3 |  |  |  | 1.3
43.4 |  |  |  | 1.3
43.5 |  |  |  | 1.3
43.6 |  |  |  | 1.3
43.7 |  |  |  | 1.3
43.8 |  |  |  | 1.3
43.9 |  |  |  | 1.3
44 |  |  |  | 1.3
44.1 |  |  |  | 1.3
44.2 |  |  |  | 1.3
44.3 |  |  |  | 1.2439
44.4 |  |  |  | 1.2
44.5 |  |  |  | 1.2
44.6 |  |  |  | 1.2
44.7 |  |  |  | 1.2
44.8 |  |  |  | 1.2
44.9 |  |  |  | 1.2
45 |  |  |  | 1.2
45.1 |  |  |  | 1.2

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.921 | 0.932 | 0.937 | 0.943 | 0.946 | 0.95 | 0.951 | 0.952 | 0.952
| 4 | 0.934 | 0.947 | 0.954 | 0.961 | 0.965 | 0.969 | 0.97 | 0.972 | 0.972
| 5 | 0.94 | 0.955 | 0.962 | 0.97 | 0.976 | 0.98 | 0.982 | 0.983 | 0.984
| 7 | 0.948 | 0.963 | 0.972 | 0.982 | 0.989 | 0.994 | 0.997 | 0.999 | 1
| 10 | 0.953 | 0.97 | 0.98 | 0.991 | 1 | 1.007 | 1.011 | 1.014 | 1.014
| 15 | 0.956 | 0.975 | 0.985 | 0.998 | 1.008 | 1.018 | 1.022 | 1.026 | 1.028
| 20 | 0.958 | 0.977 | 0.988 | 1.002 | 1.013 | 1.023 | 1.029 | 1.034 | 1.035
| 30 | 0.961 | 0.98 | 0.992 | 1.006 | 1.018 | 1.029 | 1.037 | 1.043 | 1.045
| 40 | 0.962 | 0.982 | 0.993 | 1.008 | 1.02 | 1.032 | 1.04 | 1.047 | 1.048