## Title
Varian TrueBeam
10MV Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables |  | 2. Plots
| Open field |  | Open Field
| Open Field Depth Doses |  | Wedge 15
| Open Field Profiles at 2.4cm |  | Wedge 30 Plots
| Open Field Profiles at 5cm |  | Wedge 45 Plots
| Open Field Profiles at 10cm |  | Wedge 60 Plots
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
Wedge 15
Wedge 15 Depth Doses
SSD80 Wedge 15 Depth Doses
SSD120 Wedge 15 Depth Doses
Wedge 15 Lateral Profiles
Wedge 15 Longitudinal Profiles
Wedge 15 Output Factors
Wedge 30
Wedge 30 Depth Doses
SSD80 Wedge 30 Depth Doses
SSD120 Wedge 30 Depth Doses
Wedge 30 Lateral Profiles
Wedge 30 Longitudinal Profiles
Wedge 30 Output Factors
Wedge 45
Wedge 45 Depth Doses
SSD80 Wedge 45 Depth Doses
Wedge 45 Longitudinal Profiles
Wedge 45 Lateral Profiles
Wedge 45 Longitudinal Profiles
Wedge 45 Output Factors
Wedge 60
Wedge 60 Depth Doses
SSD80 Wedge 60 Depth Doses
SSD120 Wedge 60 Depth Doses
Wedge 60 Lateral Profiles
Wedge 60 Longitudinal Profiles
Wedge 60 Output Factors

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 30 | 30.7 | 33 | 35.5 | 37.6 | 49.4 | 57.5 | 62.2
0.1 | 34.2 | 35.15 | 37.275 | 39.25 | 41.9 | 53.425 | 61.6 | 66.14
0.2 | 41.9333 | 42.68 | 44.4667 | 46.125 | 49.05 | 59.875 | 67.35 | 71.55
0.3 | 51.3 | 51.45 | 53.9 | 55.5 | 57.48 | 67.5 | 74.1 | 77.48
0.4 | 59.3 | 59.5 | 61.7 | 62.95 | 64.8 | 73.9 | 78.925 | 82.05
0.5 | 66.6 | 66.5667 | 67.8 | 69.4 | 71.2 | 79.2 | 83.575 | 85.8
0.6 | 71.78 | 71.9 | 73.4 | 74.5333 | 76.1 | 83.1667 | 87.125 | 88.7
0.7 | 76.8 | 76.6 | 78.1 | 79.16 | 80.6 | 86.6 | 90.025 | 91.3
0.8 | 80.9 | 80.625 | 81.9 | 82.7 | 84.2 | 89.4 | 92.1667 | 93.35
0.9 | 84.225 | 84.1333 | 84.94 | 86.175 | 87.2 | 91.7 | 94 | 94.925
1 | 87.225 | 87.0333 | 87.775 | 88.8 | 89.7 | 93.6 | 95.5333 | 96.18
1.1 | 90.025 | 89.6 | 90.175 | 90.9 | 91.8 | 95.2 | 96.7 | 97.3
1.2 | 91.85 | 91.5667 | 92.2 | 92.75 | 93.45 | 96.34 | 97.6 | 98.1
1.3 | 93.8 | 93.5 | 93.85 | 94.4 | 95.025 | 97.425 | 98.35 | 98.74
1.4 | 95.35 | 95.0333 | 95.35 | 95.65 | 96.3 | 98.075 | 98.9 | 99.18
1.5 | 96.5 | 96.2 | 96.4 | 96.76 | 97.15 | 98.8 | 99.4667 | 99.62
1.6 | 97.5333 | 97.1667 | 97.475 | 97.7 | 98.075 | 99.18 | 99.75 | 99.82
1.7 | 98.3 | 98 | 98.15 | 98.375 | 98.6333 | 99.6 | 99.9 | 99.92
1.8 | 98.8 | 98.6 | 98.8 | 98.975 | 99.16 | 99.8 | 100 | 100
1.9 | 99.25 | 99.175 | 99.3 | 99.35 | 99.55 | 99.9 | 100 | 100
2 | 99.6 | 99.575 | 99.64 | 99.6333 | 99.8 | 100 | 100 | 100
2.1 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 100 | 100 | 99.9
2.2 | 99.9333 | 99.9 | 99.94 | 99.925 | 100 | 99.975 | 99.9 | 99.78
2.3 | 100 | 100 | 100 | 100 | 100 | 99.825 | 99.8 | 99.6
2.4 | 99.9 | 100 | 100 | 100 | 100 | 99.675 | 99.55 | 99.5
2.5 | 99.9 | 99.9 | 100 | 99.9 | 99.85 | 99.55 | 99.4 | 99.375
2.6 | 99.725 | 99.8 | 99.9 | 99.8 | 99.675 | 99.3 | 99.075 | 99.16
2.7 | 99.7 | 99.7 | 99.7 | 99.7 | 99.5 | 99.15 | 98.875 | 98.86
2.8 | 99.3667 | 99.5 | 99.375 | 99.4 | 99.4 | 98.8 | 98.5667 | 98.5833
2.9 | 99.1 | 99.2 | 99.225 | 99.12 | 99.1 | 98.65 | 98.3333 | 98.4
3 | 98.6667 | 98.9 | 99.05 | 98.8667 | 98.8 | 98.3 | 98.1 | 98.1
3.1 | 98.34 | 98.7667 | 98.6 | 98.525 | 98.5 | 97.975 | 97.75 | 97.76
3.2 | 97.9667 | 98.26 | 98.3 | 98.275 | 98.14 | 97.625 | 97.4 | 97.45
3.3 | 97.7 | 97.8667 | 98.02 | 98.025 | 97.875 | 97.3333 | 97.1 | 97.25
3.4 | 97.25 | 97.62 | 97.64 | 97.65 | 97.45 | 97.125 | 96.9 | 96.92
3.5 | 96.9 | 97.05 | 97.26 | 97.225 | 97.15 | 96.675 | 96.475 | 96.625
3.6 | 96.4 | 96.65 | 96.85 | 96.925 | 96.725 | 96.4 | 96.2 | 96.275
3.7 | 96.05 | 96.35 | 96.48 | 96.525 | 96.45 | 96 | 95.85 | 95.94
3.8 | 95.6 | 95.875 | 96.1 | 96 | 96.15 | 95.65 | 95.525 | 95.6
3.9 | 95.1333 | 95.475 | 95.66 | 95.8 | 95.6 | 95.3 | 95.175 | 95.34
4 | 94.7 | 95.1 | 95.28 | 95.35 | 95.35 | 94.9333 | 94.85 | 94.98
4.1 | 94.2667 | 94.6 | 94.8667 | 94.8 | 94.84 | 94.6 | 94.575 | 94.7
4.2 | 93.8 | 94.15 | 94.4 | 94.5333 | 94.475 | 94.3 | 94.125 | 94.28
4.3 | 93.2667 | 93.7 | 93.9667 | 94.175 | 94.15 | 93.975 | 93.9 | 94
4.4 | 92.8667 | 93.3 | 93.6 | 93.75 | 93.6333 | 93.55 | 93.6 | 93.7
4.5 | 92.35 | 92.8333 | 93.2333 | 93.375 | 93.275 | 93.25 | 93.15 | 93.3667
4.6 | 91.98 | 92.45 | 92.7667 | 92.85 | 93 | 92.925 | 92.84 | 93
4.7 | 91.4333 | 91.9667 | 92.3667 | 92.3667 | 92.5333 | 92.5667 | 92.5 | 92.775
4.8 | 91 | 91.475 | 91.9 | 92.15 | 92.2 | 92.2 | 92.05 | 92.42
4.9 | 90.6 | 91.15 | 91.6 | 91.8 | 91.7 | 91.75 | 91.775 | 92
5 | 90 | 90.6 | 91.2 | 91.4 | 91.3 | 91.5 | 91.55 | 91.6571
5.1 | 89.54 | 90.2 | 90.65 | 91.02 | 90.84 | 91.225 | 91.1 | 91.4
5.2 | 89.175 | 89.8 | 90.25 | 90.625 | 90.65 | 90.775 | 90.75 | 91.1333
5.3 | 88.75 | 89.375 | 89.85 | 90.225 | 90.25 | 90.5 | 90.44 | 90.8
5.4 | 88.25 | 88.875 | 89.425 | 89.65 | 89.75 | 89.9 | 90.175 | 90.45
5.5 | 87.85 | 88.5 | 89.1 | 89.4333 | 89.44 | 89.66 | 89.925 | 90.06
5.6 | 87.4 | 88.1 | 88.65 | 89 | 89.0333 | 89.35 | 89.475 | 89.7667
5.7 | 87.0333 | 87.6 | 88.1 | 88.64 | 88.7 | 88.975 | 89.14 | 89.4571
5.8 | 86.55 | 87.25 | 87.75 | 88.2 | 88.35 | 88.725 | 88.9 | 89.12
5.9 | 86.12 | 86.86 | 87.375 | 87.85 | 87.98 | 88.275 | 88.6 | 88.85
6 | 85.7 | 86.32 | 87 | 87.45 | 87.6 | 88 | 88.2667 | 88.55
6.1 | 85.325 | 85.85 | 86.7 | 87.05 | 87.1 | 87.7 | 87.9 | 88.2
6.2 | 84.8 | 85.45 | 86.3 | 86.76 | 86.9 | 87.4 | 87.575 | 87.9
6.3 | 84.5 | 85.05 | 85.9 | 86.2 | 86.45 | 87.0333 | 87.325 | 87.52
6.4 | 83.85 | 84.66 | 85.3 | 85.9 | 85.975 | 86.725 | 86.95 | 87.2
6.5 | 83.525 | 84.3 | 84.9667 | 85.56 | 85.75 | 86.25 | 86.6 | 86.96
6.6 | 83.125 | 83.8 | 84.7 | 85.075 | 85.4 | 85.975 | 86.35 | 86.6
6.7 | 82.7 | 83.4 | 84.275 | 84.7 | 85 | 85.675 | 86 | 86.25
6.8 | 82.3 | 82.95 | 83.85 | 84.35 | 84.6 | 85.25 | 85.675 | 85.95
6.9 | 81.9 | 82.575 | 83.4 | 84.0333 | 84.325 | 85 | 85.325 | 85.6667
7 | 81.4 | 82.15 | 82.9 | 83.65 | 83.875 | 84.55 | 85.075 | 85.4
7.1 | 81.175 | 81.8 | 82.6 | 83.4 | 83.525 | 84.2 | 84.74 | 84.9833
7.2 | 80.75 | 81.5 | 82.25 | 82.9 | 83.1 | 83.9 | 84.38 | 84.8
7.3 | 80.175 | 80.9 | 82.025 | 82.52 | 82.75 | 83.62 | 84.05 | 84.3667
7.4 | 79.825 | 80.5 | 81.65 | 82.15 | 82.3 | 83.275 | 83.7 | 84.12
7.5 | 79.45 | 80.125 | 81.15 | 81.825 | 82.1 | 83 | 83.45 | 83.84
7.6 | 78.9 | 79.75 | 80.775 | 81.475 | 81.8 | 82.7 | 83.225 | 83.6
7.7 | 78.55 | 79.325 | 80.38 | 81.125 | 81.55 | 82.25 | 82.8333 | 83.32
7.8 | 78.225 | 79 | 80 | 80.7667 | 81.05 | 82 | 82.575 | 83.06
7.9 | 77.8667 | 78.6 | 79.6 | 80.4 | 80.65 | 81.65 | 82.225 | 82.7
8 | 77.35 | 78.2 | 79.22 | 80 | 80.3 | 81.42 | 81.9 | 82.3333
8.1 | 77.0667 | 77.8667 | 78.875 | 79.6 | 80.05 | 81.05 | 81.5 | 82.0667
8.2 | 76.7667 | 77.4 | 78.55 | 79.25 | 79.7 | 80.65 | 81.3 | 81.7429
8.3 | 76.25 | 77.075 | 78.175 | 78.925 | 79.3333 | 80.375 | 81.05 | 81.4
8.4 | 75.95 | 76.6333 | 77.925 | 78.475 | 78.875 | 80.08 | 80.64 | 81.175
8.5 | 75.4 | 76.1 | 77.4 | 78.14 | 78.55 | 79.8 | 80.3 | 80.8
8.6 | 75.05 | 75.9 | 77.06 | 77.9 | 78.275 | 79.45 | 80.05 | 80.6143
8.7 | 74.65 | 75.62 | 76.8 | 77.5333 | 77.86 | 79.1 | 79.75 | 80.275
8.8 | 74.25 | 75.15 | 76.35 | 77.2 | 77.45 | 78.75 | 79.5 | 79.9143
8.9 | 74 | 74.75 | 75.95 | 76.8 | 77.2 | 78.475 | 79.2 | 79.6
9 | 73.55 | 74.35 | 75.575 | 76.4 | 76.825 | 78.225 | 78.8 | 79.3333
9.1 | 73.1 | 73.98 | 75.15 | 76 | 76.55 | 77.875 | 78.55 | 79.08
9.2 | 72.8 | 73.575 | 74.85 | 75.75 | 76.1 | 77.58 | 78.225 | 78.7
9.3 | 72.5 | 73.2 | 74.55 | 75.425 | 75.8 | 77.22 | 77.9 | 78.5167
9.4 | 72.2 | 72.8 | 74.2333 | 74.9667 | 75.55 | 76.925 | 77.65 | 78.14
9.5 | 71.8667 | 72.5 | 73.8 | 74.575 | 75.2 | 76.575 | 77.3 | 77.8
9.6 | 71.4 | 72 | 73.4 | 74.3 | 74.85 | 76.325 | 76.95 | 77.575
9.7 | 71.05 | 71.675 | 73.1 | 74.0333 | 74.62 | 75.875 | 76.72 | 77.2667
9.8 | 70.6333 | 71.35 | 72.6 | 73.6 | 74.15 | 75.64 | 76.46 | 76.9333
9.9 | 70.2667 | 71.06 | 72.35 | 73.3 | 73.675 | 75.3 | 76.125 | 76.7
10 | 69.9 | 70.7 | 71.9333 | 72.95 | 73.475 | 75.05 | 75.875 | 76.4667
10.1 | 69.65 | 70.3 | 71.6 | 72.55 | 73.125 | 74.7 | 75.55 | 76.0833
10.2 | 69.2333 | 70 | 71.3 | 72.18 | 72.78 | 74.45 | 75.26 | 75.875
10.3 | 68.8 | 69.62 | 70.9 | 71.9 | 72.5 | 74.125 | 74.9333 | 75.56
10.4 | 68.5 | 69.3333 | 70.625 | 71.45 | 72.1 | 73.875 | 74.7 | 75.2667
10.5 | 68.075 | 68.85 | 70.175 | 71.2 | 71.7 | 73.5667 | 74.3 | 75.06
10.6 | 67.825 | 68.525 | 69.925 | 70.8333 | 71.45 | 73.1667 | 74.1 | 74.76
10.7 | 67.375 | 68.175 | 69.575 | 70.52 | 71.2 | 72.9 | 73.7667 | 74.4
10.8 | 67 | 67.9 | 69.24 | 70.225 | 70.85 | 72.65 | 73.5 | 74.1
10.9 | 66.7 | 67.5 | 69 | 69.85 | 70.4667 | 72.3 | 73.15 | 73.7833
11 | 66.45 | 67.3 | 68.65 | 69.54 | 70.1333 | 72.06 | 72.92 | 73.4833
11.1 | 66.1 | 66.8 | 68.2 | 69.3 | 69.9 | 71.7667 | 72.65 | 73.26
11.2 | 65.75 | 66.6 | 67.85 | 68.9 | 69.55 | 71.4667 | 72.325 | 72.95
11.3 | 65.375 | 66.2 | 67.6 | 68.65 | 69.2 | 71.2 | 72.05 | 72.64
11.4 | 65.05 | 65.975 | 67.275 | 68.26 | 68.975 | 70.9 | 71.7 | 72.4
11.5 | 64.58 | 65.575 | 67.025 | 68 | 68.65 | 70.6333 | 71.55 | 72.1
11.6 | 64.34 | 65.25 | 66.5667 | 67.675 | 68.3 | 70.3 | 71.15 | 71.875
11.7 | 64.1 | 65.025 | 66.3 | 67.25 | 68.075 | 70 | 71 | 71.55
11.8 | 63.75 | 64.6 | 65.95 | 66.975 | 67.75 | 69.725 | 70.675 | 71.35
11.9 | 63.5 | 64.2 | 65.675 | 66.65 | 67.36 | 69.475 | 70.35 | 71.08
12 | 63.1 | 63.9 | 65.4 | 66.4 | 67.05 | 69.125 | 70.0667 | 70.8333
12.1 | 62.84 | 63.55 | 65.125 | 66.0667 | 66.7 | 68.775 | 69.775 | 70.5
12.2 | 62.375 | 63.2667 | 64.85 | 65.7 | 66.45 | 68.625 | 69.6 | 70.22
12.3 | 62.08 | 63 | 64.425 | 65.42 | 66.2 | 68.2 | 69.2333 | 70.0143
12.4 | 61.8 | 62.575 | 64.0667 | 65.125 | 65.875 | 67.95 | 69 | 69.6857
12.5 | 61.45 | 62.275 | 63.775 | 64.95 | 65.525 | 67.625 | 68.7333 | 69.3833
12.6 | 61.1 | 62.05 | 63.5 | 64.625 | 65.275 | 67.375 | 68.475 | 69.2167
12.7 | 60.825 | 61.7 | 63.15 | 64.28 | 64.95 | 67.125 | 68.225 | 68.925
12.8 | 60.575 | 61.4 | 62.7667 | 64.05 | 64.8 | 66.85 | 68.05 | 68.6833
12.9 | 60.125 | 61.1333 | 62.4 | 63.7 | 64.4 | 66.525 | 67.575 | 68.4
13 | 59.98 | 60.825 | 62.26 | 63.3333 | 64.1 | 66.4 | 67.4333 | 68.1
13.1 | 59.725 | 60.4667 | 61.925 | 63.025 | 63.8 | 65.95 | 67.18 | 67.9
13.2 | 59.36 | 60.1667 | 61.675 | 62.7667 | 63.5333 | 65.7 | 66.9 | 67.575
13.3 | 59.05 | 59.85 | 61.35 | 62.46 | 63.2667 | 65.45 | 66.6 | 67.325
13.4 | 58.6 | 59.6333 | 60.975 | 62.08 | 62.9667 | 65.225 | 66.25 | 67.1167
13.5 | 58.26 | 59.3 | 60.74 | 61.9 | 62.7 | 64.95 | 66.02 | 66.9
13.6 | 58.1 | 58.95 | 60.5 | 61.6 | 62.3 | 64.625 | 65.775 | 66.6
13.7 | 57.85 | 58.7 | 60.2 | 61.3 | 62.1 | 64.375 | 65.525 | 66.3667
13.8 | 57.55 | 58.375 | 59.8667 | 61.025 | 61.8 | 64.025 | 65.28 | 66.16
13.9 | 57.1333 | 58.025 | 59.55 | 60.65 | 61.52 | 63.8 | 65 | 65.85
14 | 56.925 | 57.675 | 59.3 | 60.3667 | 61.25 | 63.55 | 64.8 | 65.5667
14.1 | 56.7 | 57.4667 | 58.9667 | 60.1 | 60.925 | 63.32 | 64.5 | 65.35
14.2 | 56.4 | 57.2 | 58.725 | 59.85 | 60.675 | 63.075 | 64.2 | 65.12
14.3 | 56.2 | 56.95 | 58.4333 | 59.6 | 60.4 | 62.825 | 63.9667 | 64.8333
14.4 | 55.7333 | 56.6333 | 58.175 | 59.3 | 60.1 | 62.475 | 63.675 | 64.6
14.5 | 55.3667 | 56.35 | 57.85 | 58.9 | 59.85 | 62.35 | 63.44 | 64.3333
14.6 | 55.1 | 56.05 | 57.5333 | 58.7 | 59.5 | 62 | 63.2 | 64
14.7 | 54.925 | 55.8 | 57.2 | 58.45 | 59.32 | 61.75 | 62.95 | 63.8
14.8 | 54.58 | 55.4 | 56.96 | 58.15 | 58.95 | 61.5 | 62.7 | 63.6167
14.9 | 54.45 | 55.2 | 56.65 | 57.825 | 58.75 | 61.225 | 62.4 | 63.4
15 | 54.1 | 54.85 | 56.375 | 57.6 | 58.425 | 60.975 | 62.15 | 63.0333
15.1 | 53.8 | 54.7 | 56.125 | 57.4 | 58.175 | 60.65 | 61.9 | 62.85
15.2 | 53.5 | 54.32 | 55.875 | 57 | 57.88 | 60.375 | 61.68 | 62.58
15.3 | 53.375 | 54.125 | 55.75 | 56.8 | 57.6 | 60.15 | 61.4 | 62.2833
15.4 | 53.05 | 53.875 | 55.4 | 56.5 | 57.35 | 59.9 | 61.25 | 62.1
15.5 | 52.775 | 53.625 | 55.05 | 56.26 | 57.1 | 59.65 | 60.9 | 61.8167
15.6 | 52.45 | 53.275 | 54.7 | 56 | 56.8 | 59.4 | 60.66 | 61.6
15.7 | 52.2 | 53 | 54.45 | 55.7 | 56.52 | 59.2 | 60.45 | 61.2833
15.8 | 51.95 | 52.78 | 54.225 | 55.48 | 56.325 | 58.975 | 60.175 | 61.1
15.9 | 51.8 | 52.55 | 53.95 | 55.25 | 56.1 | 58.625 | 59.9333 | 60.9
16 | 51.5 | 52.35 | 53.75 | 55 | 55.8 | 58.45 | 59.68 | 60.68
16.1 | 51.25 | 52.04 | 53.5 | 54.6667 | 55.55 | 58.2 | 59.5 | 60.42
16.2 | 50.975 | 51.72 | 53.1 | 54.4667 | 55.3 | 58 | 59.2333 | 60.14
16.3 | 50.725 | 51.525 | 52.95 | 54.225 | 55.0667 | 57.775 | 58.9667 | 59.96
16.4 | 50.475 | 51.175 | 52.7 | 54 | 54.9 | 57.525 | 58.7 | 59.6714
16.5 | 50.15 | 50.98 | 52.475 | 53.6 | 54.5 | 57.18 | 58.5667 | 59.4333
16.6 | 49.9 | 50.675 | 52.2 | 53.4 | 54.25 | 57 | 58.3 | 59.2
16.7 | 49.64 | 50.425 | 51.9 | 53.2 | 54 | 56.775 | 58.075 | 58.9714
16.8 | 49.46 | 50.2 | 51.65 | 52.96 | 53.8 | 56.5 | 57.825 | 58.8
16.9 | 49.1 | 49.95 | 51.4 | 52.7 | 53.5667 | 56.16 | 57.6 | 58.56
17 | 48.9 | 49.7 | 51.125 | 52.45 | 53.325 | 56.0667 | 57.45 | 58.3333
17.1 | 48.575 | 49.5 | 51 | 52.1333 | 53.1 | 55.725 | 57.18 | 58.08
17.2 | 48.4 | 49.225 | 50.65 | 51.9 | 52.7667 | 55.575 | 56.9 | 57.9
17.3 | 48.2 | 49 | 50.4 | 51.65 | 52.6 | 55.325 | 56.65 | 57.6
17.4 | 48 | 48.7 | 50.275 | 51.4 | 52.375 | 55.1 | 56.5 | 57.4
17.5 | 47.7 | 48.5333 | 50.025 | 51.225 | 52.05 | 54.825 | 56.25 | 57.1833
17.6 | 47.5 | 48.325 | 49.675 | 50.95 | 51.78 | 54.65 | 56 | 57
17.7 | 47.1667 | 48.05 | 49.425 | 50.7 | 51.45 | 54.4 | 55.76 | 56.76
17.8 | 47 | 47.775 | 49.2 | 50.425 | 51.3 | 54.2 | 55.525 | 56.4857
17.9 | 46.8 | 47.48 | 49.04 | 50.2 | 51.15 | 53.9 | 55.4 | 56.2667
18 | 46.475 | 47.25 | 48.8 | 49.95 | 50.8333 | 53.75 | 55.125 | 56.0667
18.1 | 46.225 | 47 | 48.55 | 49.7 | 50.7 | 53.475 | 54.9 | 55.8
18.2 | 46.08 | 46.75 | 48.3 | 49.4333 | 50.4 | 53.225 | 54.625 | 55.65
18.3 | 45.85 | 46.55 | 48.05 | 49.3 | 50.225 | 53.1 | 54.5 | 55.46
18.4 | 45.7 | 46.35 | 47.9 | 49.0667 | 50 | 52.8 | 54.25 | 55.1833
18.5 | 45.45 | 46.125 | 47.65 | 48.78 | 49.75 | 52.7 | 54.1 | 55.0167
18.6 | 45.2 | 45.95 | 47.325 | 48.525 | 49.5 | 52.3667 | 53.775 | 54.825
18.7 | 45 | 45.74 | 47.175 | 48.4 | 49.35 | 52.175 | 53.55 | 54.62
18.8 | 44.76 | 45.5 | 46.85 | 48.1333 | 49.125 | 51.925 | 53.375 | 54.4
18.9 | 44.525 | 45.1333 | 46.5333 | 47.9 | 48.875 | 51.675 | 53.125 | 54.1286
19 | 44.3 | 44.85 | 46.4 | 47.725 | 48.54 | 51.44 | 52.9 | 53.96
19.1 | 44.1 | 44.675 | 46.225 | 47.5 | 48.4 | 51.25 | 52.7 | 53.78
19.2 | 43.85 | 44.525 | 46 | 47.2667 | 48.15 | 51.04 | 52.42 | 53.48
19.3 | 43.6 | 44.3 | 45.75 | 47.0333 | 47.9 | 50.9 | 52.275 | 53.3
19.4 | 43.45 | 44.125 | 45.5 | 46.8 | 47.75 | 50.55 | 52 | 53.0333
19.5 | 43.15 | 43.9 | 45.3333 | 46.6 | 47.52 | 50.4 | 51.875 | 52.8333
19.6 | 42.975 | 43.6 | 45.075 | 46.35 | 47.2 | 50.2 | 51.625 | 52.6
19.7 | 42.75 | 43.45 | 44.9 | 46.1333 | 47 | 50 | 51.34 | 52.425
19.8 | 42.525 | 43.22 | 44.6333 | 45.875 | 46.825 | 49.725 | 51.2 | 52.25
19.9 | 42.3 | 42.975 | 44.5 | 45.625 | 46.575 | 49.45 | 51 | 51.9833
20 | 42.05 | 42.825 | 44.2 | 45.48 | 46.4 | 49.3 | 50.8 | 51.8
20.1 | 41.9 | 42.5 | 44.025 | 45.25 | 46.15 | 49.075 | 50.55 | 51.5667
20.2 | 41.7 | 42.4 | 43.7667 | 45 | 46 | 48.8 | 50.4 | 51.3667
20.3 | 41.5 | 42.175 | 43.6 | 44.8 | 45.7 | 48.675 | 50.175 | 51.2
20.4 | 41.25 | 42 | 43.4 | 44.62 | 45.56 | 48.5 | 49.925 | 50.975
20.5 | 41.025 | 41.8 | 43.2 | 44.375 | 45.25 | 48.25 | 49.7 | 50.76
20.6 | 40.825 | 41.6 | 42.925 | 44.2 | 45.1 | 48.15 | 49.525 | 50.54
20.7 | 40.575 | 41.4 | 42.8 | 43.9667 | 44.8333 | 47.9 | 49.3667 | 50.3
20.8 | 40.4 | 41.1667 | 42.55 | 43.8 | 44.7 | 47.6 | 49.1667 | 50.1667
20.9 | 40.25 | 40.925 | 42.32 | 43.6 | 44.5 | 47.475 | 49 | 49.9286
21 | 40 | 40.675 | 42.175 | 43.375 | 44.2333 | 47.225 | 48.725 | 49.8
21.1 | 39.86 | 40.525 | 42 | 43.2 | 44.1 | 47.04 | 48.6 | 49.575
21.2 | 39.7 | 40.3 | 41.7 | 42.975 | 43.9 | 46.8 | 48.45 | 49.3667
21.3 | 39.475 | 40.175 | 41.55 | 42.8 | 43.7 | 46.6 | 48.15 | 49.175
21.4 | 39.325 | 40 | 41.4 | 42.5333 | 43.5 | 46.4 | 47.975 | 48.94
21.5 | 39.14 | 39.8 | 41.1 | 42.3 | 43.3 | 46.175 | 47.7 | 48.7333
21.6 | 39 | 39.55 | 40.95 | 42.2 | 43.1 | 46.025 | 47.6 | 48.5
21.7 | 38.75 | 39.38 | 40.7 | 41.95 | 42.9 | 45.9 | 47.4 | 48.3
21.8 | 38.5 | 39.2 | 40.5 | 41.8 | 42.65 | 45.65 | 47.3 | 48.2
21.9 | 38.3 | 39 | 40.38 | 41.5333 | 42.5 | 45.4 | 47 | 47.9333
22 | 38.125 | 38.875 | 40.25 | 41.3 | 42.275 | 45.3 | 46.85 | 47.75
22.1 | 37.975 | 38.6 | 40.0667 | 41.2 | 42.1 | 45.1 | 46.7 | 47.6
22.2 | 37.825 | 38.4 | 39.8 | 40.9 | 41.9 | 44.9 | 46.45 | 47.42
22.3 | 37.58 | 38.3 | 39.6 | 40.725 | 41.725 | 44.7 | 46.2 | 47.1667
22.4 | 37.35 | 38.1 | 39.5 | 40.5 | 41.5 | 44.425 | 46 | 47.0667
22.5 | 37.2 | 37.9 | 39.2333 | 40.35 | 41.35 | 44.35 | 45.8 | 46.88
22.6 | 37.05 | 37.7 | 39.075 | 40.2 | 41.1333 | 44.1 | 45.675 | 46.66
22.7 | 36.8 | 37.5 | 38.825 | 40 | 41 | 43.95 | 45.5 | 46.3833
22.8 | 36.65 | 37.3 | 38.7 | 39.8333 | 40.825 | 43.7 | 45.3 | 46.2
22.9 | 36.4333 | 37.1333 | 38.5 | 39.625 | 40.575 | 43.55 | 45.1 | 46.0833
23 | 36.3 | 36.975 | 38.3 | 39.475 | 40.4 | 43.325 | 44.9 | 45.9167
23.1 | 36.15 | 36.74 | 38.05 | 39.3 | 40.18 | 43.175 | 44.7 | 45.7
23.2 | 35.9 | 36.6 | 38 | 39.1667 | 40 | 43 | 44.525 | 45.6
23.3 | 35.75 | 36.4 | 37.775 | 38.9 | 39.8 | 42.8 | 44.4 | 45.3833
23.4 | 35.6 | 36.2 | 37.625 | 38.66 | 39.6 | 42.65 | 44.225 | 45.2167
23.5 | 35.475 | 36.1 | 37.4 | 38.525 | 39.475 | 42.4 | 43.975 | 45.02
23.6 | 35.3 | 35.85 | 37.225 | 38.375 | 39.3 | 42.3 | 43.8333 | 44.85
23.7 | 35.1 | 35.7 | 37.1 | 38.15 | 39.075 | 42.1 | 43.6 | 44.6833
23.8 | 35 | 35.525 | 36.85 | 38 | 38.9 | 41.96 | 43.475 | 44.5
23.9 | 34.8 | 35.35 | 36.7 | 37.8333 | 38.7 | 41.725 | 43.3 | 44.32
24 | 34.65 | 35.2 | 36.5 | 37.7 | 38.6 | 41.5 | 43.175 | 44.12
24.1 | 34.4 | 35.0333 | 36.3667 | 37.5333 | 38.4 | 41.375 | 42.925 | 44
24.2 | 34.25 | 34.8 | 36.15 | 37.3 | 38.3 | 41.225 | 42.78 | 43.8333
24.3 | 34.1 | 34.6667 | 35.975 | 37.2 | 38.025 | 41.0667 | 42.6 | 43.5833
24.4 | 33.95 | 34.525 | 35.825 | 37.02 | 37.875 | 40.9 | 42.4 | 43.3833
24.5 | 33.775 | 34.425 | 35.6667 | 36.7333 | 37.725 | 40.675 | 42.25 | 43.3
24.6 | 33.6 | 34.2 | 35.5 | 36.6333 | 37.5667 | 40.425 | 42.1 | 43.1
24.7 | 33.48 | 34 | 35.3 | 36.42 | 37.3 | 40.275 | 41.875 | 42.9167
24.8 | 33.34 | 33.8 | 35.1 | 36.2 | 37.2 | 40.025 | 41.7 | 42.74
24.9 | 33.1 | 33.7 | 34.92 | 36.1 | 37.025 | 40 | 41.5 | 42.6
25 | 32.95 | 33.5 | 34.8 | 35.9 | 36.875 | 39.8 | 41.375 | 42.36
25.1 | 32.8 | 33.3 | 34.7 | 35.7333 | 36.6 | 39.6 | 41.25 | 42.15
25.2 | 32.6 | 33.175 | 34.5 | 35.6 | 36.5 | 39.45 | 41.0667 | 42.02
25.3 | 32.4 | 33.1 | 34.3333 | 35.45 | 36.3 | 39.275 | 40.9 | 41.85
25.4 | 32.3 | 32.825 | 34.1667 | 35.3 | 36.2 | 39.12 | 40.725 | 41.66
25.5 | 32.1 | 32.675 | 34.05 | 35.075 | 36 | 38.975 | 40.5 | 41.5
25.6 | 31.925 | 32.525 | 33.9 | 34.9 | 35.8 | 38.8 | 40.325 | 41.3667
25.7 | 31.85 | 32.4 | 33.675 | 34.8 | 35.7 | 38.675 | 40.225 | 41.28
25.8 | 31.7 | 32.3 | 33.525 | 34.625 | 35.5 | 38.44 | 40.025 | 41.0286
25.9 | 31.6 | 32.1 | 33.38 | 34.4667 | 35.325 | 38.28 | 39.9 | 40.825
26 | 31.4 | 31.95 | 33.25 | 34.3333 | 35.2 | 38.15 | 39.7 | 40.65
26.1 | 31.3 | 31.8 | 33 | 34.1 | 35.1 | 38 | 39.5 | 40.5
26.2 | 31.12 | 31.6 | 32.9 | 34 | 34.9 | 37.875 | 39.4 | 40.4
26.3 | 30.975 | 31.425 | 32.7 | 33.8 | 34.7 | 37.65 | 39.3 | 40.22
26.4 | 30.825 | 31.3 | 32.55 | 33.64 | 34.6 | 37.5 | 39.1 | 40.04
26.5 | 30.7 | 31.2 | 32.425 | 33.575 | 34.4 | 37.3667 | 38.925 | 39.8667
26.6 | 30.55 | 31.1 | 32.275 | 33.35 | 34.25 | 37.2 | 38.7 | 39.7
26.7 | 30.4 | 30.9 | 32.125 | 33.2 | 34.1 | 37 | 38.5 | 39.575
26.8 | 30.2 | 30.75 | 32 | 32.95 | 33.95 | 36.9 | 38.4 | 39.44
26.9 | 30.1 | 30.62 | 31.8 | 32.9 | 33.8 | 36.65 | 38.26 | 39.2
27 | 29.9 | 30.5 | 31.68 | 32.7 | 33.675 | 36.5 | 38.1 | 39.1333
27.1 | 29.8 | 30.3 | 31.5 | 32.575 | 33.525 | 36.4 | 37.975 | 38.9
27.2 | 29.7 | 30.175 | 31.4 | 32.5 | 33.3 | 36.275 | 37.7667 | 38.8333
27.3 | 29.5 | 30.025 | 31.2 | 32.275 | 33.2 | 36.02 | 37.6 | 38.6
27.4 | 29.4 | 29.9 | 31.1 | 32.125 | 32.98 | 35.9 | 37.5 | 38.4167
27.5 | 29.25 | 29.7 | 31 | 32.08 | 32.9 | 35.75 | 37.375 | 38.3
27.6 | 29.1 | 29.65 | 30.8 | 31.85 | 32.7 | 35.6 | 37.2 | 38.16
27.7 | 28.975 | 29.525 | 30.725 | 31.6667 | 32.6 | 35.45 | 37.05 | 38
27.8 | 28.825 | 29.3 | 30.5 | 31.525 | 32.425 | 35.3 | 36.9 | 37.84
27.9 | 28.675 | 29.2 | 30.35 | 31.375 | 32.275 | 35.175 | 36.7 | 37.74
28 | 28.525 | 29.025 | 30.2 | 31.225 | 32.15 | 35 | 36.6 | 37.5333
28.1 | 28.4 | 28.88 | 30 | 31.1 | 32 | 34.8333 | 36.375 | 37.44
28.2 | 28.3 | 28.75 | 29.9 | 30.925 | 31.85 | 34.7 | 36.3 | 37.2
28.3 | 28.1 | 28.6 | 29.825 | 30.8 | 31.7 | 34.55 | 36.1 | 37.1167
28.4 | 28 | 28.5 | 29.675 | 30.7 | 31.6 | 34.4 | 35.95 | 37
28.5 | 27.9 | 28.4 | 29.5 | 30.55 | 31.5 | 34.26 | 35.8 | 36.8167
28.6 | 27.8 | 28.275 | 29.4 | 30.4 | 31.25 | 34.125 | 35.7 | 36.7
28.7 | 27.625 | 28.025 | 29.3 | 30.3 | 31.1 | 33.975 | 35.575 | 36.5667
28.8 | 27.48 | 27.9 | 29.1 | 30.1 | 31 | 33.8667 | 35.4 | 36.4143
28.9 | 27.35 | 27.8 | 28.95 | 30.05 | 30.875 | 33.74 | 35.2667 | 36.24
29 | 27.2 | 27.7 | 28.8 | 29.9 | 30.7 | 33.5 | 35.1333 | 36.1
29.1 | 27.05 | 27.55 | 28.66 | 29.8 | 30.575 | 33.35 | 35 | 35.95
29.2 | 26.9 | 27.425 | 28.6 | 29.6 | 30.45 | 33.2 | 34.85 | 35.775
29.3 | 26.8 | 27.275 | 28.425 | 29.4 | 30.3 | 33.1 | 34.7 | 35.6667
29.4 | 26.7 | 27.125 | 28.34 | 29.35 | 30.1667 | 33 | 34.56 | 35.5
29.5 | 26.6 | 27 | 28.1 | 29.1 | 30 | 32.8 | 34.4 | 35.3
29.6 | 26.44 | 26.94 | 28 | 29 | 29.9 | 32.74 | 34.275 | 35.28
29.7 | 26.3 | 26.8 | 27.9 | 28.8 | 29.775 | 32.58 | 34.1 | 35.05
29.8 | 26.15 | 26.65 | 27.8 | 28.75 | 29.625 | 32.45 | 33.975 | 34.85
29.9 | 26 | 26.5 | 27.62 | 28.6 | 29.5 | 32.3 | 33.8 | 34.7833
30 | 25.95 | 26.4 | 27.5 | 28.5 | 29.35 | 32.1 | 33.65 | 34.7

## Open Field Profiles at 2.4cm
Lateral Profiles
Open field
SSD=100cm
Depth=2.4 cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.4 | [7 empty] | 4.05
-25.3 | [7 empty] | 4.15
-25.2 | [7 empty] | 4.25
-25.1 | [7 empty] | 4.35
-25 | [7 empty] | 4.45
-24.9 | [7 empty] | 4.55
-24.8 | [7 empty] | 4.65
-24.7 | [7 empty] | 4.75
-24.6 | [7 empty] | 4.9
-24.5 | [7 empty] | 5.05
-24.4 | [7 empty] | 5.15
-24.3 | [7 empty] | 5.25
-24.2 | [7 empty] | 5.4
-24.1 | [7 empty] | 5.55
-24 | [7 empty] | 5.7
-23.9 | [7 empty] | 5.85
-23.8 | [7 empty] | 6
-23.7 | [7 empty] | 6.2
-23.6 | [7 empty] | 6.35
-23.5 | [7 empty] | 6.5
-23.4 | [7 empty] | 6.7
-23.3 | [7 empty] | 6.9
-23.2 | [7 empty] | 7.1
-23.1 | [7 empty] | 7.3
-23 | [7 empty] | 7.5
-22.9 | [7 empty] | 7.7
-22.8 | [7 empty] | 7.9
-22.7 | [7 empty] | 8.1
-22.6 | [7 empty] | 8.3
-22.5 | [7 empty] | 8.55
-22.4 | [7 empty] | 8.8
-22.3 | [7 empty] | 9
-22.2 | [7 empty] | 9.25
-22.1 | [7 empty] | 9.55
-22 | [7 empty] | 9.85
-21.9 | [7 empty] | 10.15
-21.8 | [7 empty] | 10.5
-21.7 | [7 empty] | 10.9
-21.6 | [7 empty] | 11.35
-21.5 | [7 empty] | 11.9
-21.4 | [7 empty] | 12.55
-21.3 | [7 empty] | 13.4
-21.2 | [7 empty] | 14.55
-21.1 | [7 empty] | 16.2
-21 | [7 empty] | 18.7
-20.9 | [7 empty] | 22.65
-20.8 | [7 empty] | 28.9
-20.7 | [7 empty] | 37.45
-20.6 | [7 empty] | 47.6
-20.5 | [7 empty] | 58.65
-20.4 | [7 empty] | 69.65
-20.3 | [6 empty] | 4.3 | 79.7
-20.2 | [6 empty] | 4.4 | 88
-20.1 | [6 empty] | 4.5 | 93.95
-20 | [6 empty] | 4.6 | 97.65
-19.9 | [6 empty] | 4.7 | 99.85
-19.8 | [6 empty] | 4.8 | 101.3
-19.7 | [6 empty] | 4.9 | 102.3
-19.6 | [6 empty] | 5 | 103.05
-19.5 | [6 empty] | 5.1 | 103.65
-19.4 | [6 empty] | 5.2 | 104.1
-19.3 | [6 empty] | 5.3 | 104.5
-19.2 | [6 empty] | 5.4 | 104.75
-19.1 | [6 empty] | 5.5 | 104.9
-19 | [6 empty] | 5.7 | 105.1
-18.9 | [6 empty] | 5.8 | 105.3
-18.8 | [6 empty] | 5.9 | 105.45
-18.7 | [6 empty] | 6.1 | 105.55
-18.6 | [6 empty] | 6.2 | 105.65
-18.5 | [6 empty] | 6.3 | 105.75
-18.4 | [6 empty] | 6.5 | 105.85
-18.3 | [6 empty] | 6.7 | 105.9
-18.2 | [6 empty] | 6.8 | 105.95
-18.1 | [6 empty] | 7 | 106
-18 | [6 empty] | 7.1 | 106.05
-17.9 | [6 empty] | 7.3 | 106.1
-17.8 | [6 empty] | 7.5 | 106.1
-17.7 | [6 empty] | 7.7 | 106.1
-17.6 | [6 empty] | 7.9 | 106.1
-17.5 | [6 empty] | 8 | 106.1
-17.4 | [6 empty] | 8.2 | 106.1
-17.3 | [6 empty] | 8.4 | 106.1
-17.2 | [6 empty] | 8.7 | 106.1
-17.1 | [6 empty] | 8.9 | 106.1
-17 | [6 empty] | 9.1 | 106.1
-16.9 | [6 empty] | 9.4 | 106.05
-16.8 | [6 empty] | 9.7 | 106
-16.7 | [6 empty] | 10 | 106
-16.6 | [6 empty] | 10.3 | 106
-16.5 | [6 empty] | 10.8 | 105.95
-16.4 | [6 empty] | 11.3 | 105.9
-16.3 | [6 empty] | 11.9 | 105.9
-16.2 | [6 empty] | 12.7 | 105.85
-16.1 | [6 empty] | 13.7 | 105.8
-16 | [6 empty] | 15.2 | 105.8
-15.9 | [6 empty] | 17.3 | 105.75
-15.8 | [6 empty] | 20.6 | 105.7
-15.7 | [6 empty] | 25.9 | 105.65
-15.6 | [6 empty] | 33.9 | 105.6
-15.5 | [6 empty] | 43.6 | 105.6
-15.4 | [6 empty] | 54.4 | 105.6
-15.3 | [6 empty] | 65.5 | 105.55
-15.2 | [5 empty] | 3.2 | 75.8 | 105.5
-15.1 | [5 empty] | 3.2 | 84.7 | 105.5
-15 | [5 empty] | 3.3 | 91.1 | 105.5
-14.9 | [5 empty] | 3.4 | 95.2 | 105.5
-14.8 | [5 empty] | 3.44 | 97.7 | 105.45
-14.7 | [5 empty] | 3.54 | 99.1 | 105.4
-14.6 | [5 empty] | 3.6 | 100.2 | 105.4
-14.5 | [5 empty] | 3.7 | 100.9 | 105.4
-14.4 | [5 empty] | 3.8 | 101.6 | 105.4
-14.3 | [5 empty] | 3.9 | 102 | 105.35
-14.2 | [5 empty] | 4 | 102.4 | 105.3
-14.1 | [5 empty] | 4.1 | 102.6 | 105.3
-14 | [5 empty] | 4.2 | 102.8 | 105.3
-13.9 | [5 empty] | 4.3 | 103 | 105.25
-13.8 | [5 empty] | 4.4 | 103.1 | 105.2
-13.7 | [5 empty] | 4.54 | 103.2 | 105.2
-13.6 | [5 empty] | 4.64 | 103.3 | 105.2
-13.5 | [5 empty] | 4.74 | 103.4 | 105.2
-13.4 | [5 empty] | 4.94 | 103.5 | 105.2
-13.3 | [5 empty] | 5.04 | 103.5 | 105.15
-13.2 | [5 empty] | 5.14 | 103.6 | 105.1
-13.1 | [5 empty] | 5.24 | 103.6 | 105.1
-13 | [5 empty] | 5.44 | 103.7 | 105.1
-12.9 | [5 empty] | 5.54 | 103.7 | 105.1
-12.8 | [5 empty] | 5.74 | 103.8 | 105.1
-12.7 | [5 empty] | 5.9 | 103.8 | 105.1
-12.6 | [5 empty] | 6.04 | 103.8 | 105.1
-12.5 | [5 empty] | 6.24 | 103.9 | 105.1
-12.4 | [5 empty] | 6.44 | 104 | 105.05
-12.3 | [5 empty] | 6.64 | 104 | 105
-12.2 | [5 empty] | 6.8 | 104.1 | 105
-12.1 | [5 empty] | 7.04 | 104.1 | 105
-12 | [5 empty] | 7.24 | 104.1 | 105
-11.9 | [5 empty] | 7.44 | 104.1 | 105
-11.8 | [5 empty] | 7.74 | 104.1 | 105
-11.7 | [5 empty] | 7.98 | 104.1 | 104.95
-11.6 | [5 empty] | 8.34 | 104.2 | 104.9
-11.5 | [5 empty] | 8.68 | 104.2 | 104.9
-11.4 | [5 empty] | 9.12 | 104.2 | 104.9
-11.3 | [5 empty] | 9.62 | 104.2 | 104.9
-11.2 | [5 empty] | 10.22 | 104.2 | 104.9
-11.1 | [5 empty] | 10.96 | 104.2 | 104.85
-11 | [5 empty] | 12.04 | 104.2 | 104.8
-10.9 | [5 empty] | 13.42 | 104.2 | 104.8
-10.8 | [5 empty] | 15.38 | 104.2 | 104.8
-10.7 | [5 empty] | 18.46 | 104.3 | 104.8
-10.6 | [5 empty] | 23.38 | 104.3 | 104.75
-10.5 | [5 empty] | 30.72 | 104.3 | 104.7
-10.4 | [5 empty] | 40 | 104.3 | 104.7
-10.3 | [5 empty] | 50.6 | 104.3 | 104.7
-10.2 | [5 empty] | 61.6 | 104.3 | 104.7
-10.1 |  |  |  |  | 1.7 | 72.02 | 104.3 | 104.7
-10 |  |  |  |  | 1.8 | 81.14 | 104.3 | 104.65
-9.9 |  |  |  |  | 1.8 | 88.18 | 104.2 | 104.6
-9.8 |  |  |  |  | 1.9 | 92.78 | 104.2 | 104.6
-9.7 |  |  |  |  | 1.9 | 95.54 | 104.2 | 104.55
-9.6 |  |  |  |  | 1.9 | 97.38 | 104.3 | 104.5
-9.5 |  |  |  |  | 2 | 98.54 | 104.2 | 104.5
-9.4 |  |  |  |  | 2 | 99.42 | 104.2 | 104.45
-9.3 |  |  |  |  | 2.1 | 100.18 | 104.2 | 104.4
-9.2 |  |  |  |  | 2.1 | 100.58 | 104.2 | 104.35
-9.1 |  |  |  |  | 2.2 | 100.98 | 104.2 | 104.3
-9 |  |  |  | 1.4 | 2.3 | 101.34 | 104.2 | 104.3
-8.9 |  |  |  | 1.42 | 2.3 | 101.58 | 104.2 | 104.3
-8.8 |  |  |  | 1.5 | 2.4 | 101.84 | 104.2 | 104.25
-8.7 |  |  |  | 1.5 | 2.5 | 102.04 | 104.1 | 104.2
-8.6 |  |  |  | 1.52 | 2.5 | 102.14 | 104.2 | 104.2
-8.5 |  |  |  | 1.6 | 2.6 | 102.24 | 104.1 | 104.2
-8.4 |  |  |  | 1.6 | 2.7 | 102.34 | 104.1 | 104.15
-8.3 |  |  |  | 1.7 | 2.8 | 102.5 | 104.1 | 104.1
-8.2 |  |  |  | 1.7 | 2.8 | 102.54 | 104.1 | 104.1
-8.1 |  |  |  | 1.72 | 3 | 102.6 | 104.1 | 104.1
-8 |  |  | 1.1 | 1.8 | 3 | 102.6 | 104.1 | 104.05
-7.9 |  |  | 1.1 | 1.82 | 3.1 | 102.7 | 104.1 | 104
-7.8 |  |  | 1.1 | 1.9 | 3.2 | 102.8 | 104 | 104
-7.7 |  |  | 1.1 | 1.92 | 3.4 | 102.8 | 104.1 | 104
-7.6 |  |  | 1.2 | 2 | 3.5 | 102.9 | 104 | 104
-7.5 |  |  | 1.2 | 2.1 | 3.6 | 102.9 | 104 | 103.95
-7.4 |  |  | 1.2 | 2.12 | 3.7 | 102.94 | 104 | 103.9
-7.3 |  |  | 1.2 | 2.2 | 3.9 | 103 | 104 | 103.9
-7.2 |  |  | 1.3 | 2.3 | 4 | 103 | 104 | 103.85
-7.1 |  |  | 1.3 | 2.32 | 4.2 | 103 | 103.9 | 103.8
-7 |  | 0.7 | 1.3 | 2.4 | 4.3 | 103 | 103.9 | 103.75
-6.9 |  | 0.8 | 1.4 | 2.5 | 4.5 | 103 | 103.9 | 103.7
-6.8 |  | 0.8 | 1.4 | 2.6 | 4.7 | 103 | 103.9 | 103.65
-6.7 |  | 0.8 | 1.5 | 2.7 | 4.9 | 103.06 | 103.8 | 103.6
-6.6 |  | 0.8 | 1.5 | 2.8 | 5.1 | 103.06 | 103.8 | 103.6
-6.5 | 0.6 | 0.8 | 1.6 | 2.92 | 5.4 | 103.1 | 103.8 | 103.55
-6.4 | 0.6 | 0.86 | 1.6 | 3 | 5.7 | 103.1 | 103.7 | 103.5
-6.3 | 0.6 | 0.9 | 1.7 | 3.12 | 6.1 | 103.1 | 103.7 | 103.45
-6.2 | 0.6 | 0.9 | 1.7 | 3.22 | 6.6 | 103.1 | 103.7 | 103.4
-6.1 | 0.6 | 0.9 | 1.8 | 3.4 | 7.2 | 103.04 | 103.6 | 103.4
-6 | 0.6 | 0.9 | 1.8 | 3.52 | 7.9 | 103 | 103.6 | 103.35
-5.9 | 0.6 | 0.9 | 1.9 | 3.7 | 8.8 | 103 | 103.6 | 103.3
-5.8 | 0.7 | 0.96 | 2 | 3.82 | 10.1 | 103 | 103.5 | 103.3
-5.7 | 0.7 | 1 | 2.1 | 4.02 | 11.9 | 103 | 103.5 | 103.25
-5.6 | 0.7 | 1 | 2.1 | 4.24 | 14.5 | 103 | 103.4 | 103.15
-5.5 | 0.7 | 1 | 2.2 | 4.52 | 18.6 | 103 | 103.4 | 103.1
-5.4 | 0.7 | 1.06 | 2.3 | 4.82 | 25.1 | 103 | 103.3 | 103.05
-5.3 | 0.7 | 1.1 | 2.4 | 5.14 | 34 | 103 | 103.3 | 103
-5.2 | 0.7 | 1.1 | 2.5 | 5.56 | 44.4 | 102.9 | 103.3 | 103
-5.1 | 0.8 | 1.16 | 2.6 | 6.06 | 55.6 | 102.9 | 103.2 | 102.95
-5 | 0.8 | 1.2 | 2.7 | 6.76 | 66.5 | 102.8 | 103.2 | 102.85
-4.9 | 0.8 | 1.26 | 2.9 | 7.6 | 76.3 | 102.76 | 103.1 | 102.8
-4.8 | 0.8 | 1.3 | 3 | 8.74 | 84.3 | 102.7 | 103.1 | 102.75
-4.7 | 0.8 | 1.3 | 3.1 | 10.3 | 89.9 | 102.7 | 103 | 102.7
-4.6 | 0.9 | 1.4 | 3.3 | 12.48 | 93.4 | 102.66 | 103 | 102.65
-4.5 | 0.9 | 1.46 | 3.6 | 15.88 | 95.3 | 102.6 | 102.9 | 102.55
-4.4 | 0.9 | 1.5 | 3.8 | 21.3 | 96.8 | 102.6 | 102.8 | 102.45
-4.3 | 0.9 | 1.56 | 4.1 | 29.1 | 97.8 | 102.5 | 102.8 | 102.4
-4.2 | 1 | 1.6 | 4.4 | 38.84 | 98.6 | 102.5 | 102.7 | 102.35
-4.1 | 1 | 1.7 | 4.9 | 49.62 | 99.1 | 102.4 | 102.6 | 102.25
-4 | 1.1 | 1.8 | 5.5 | 60.88 | 99.6 | 102.36 | 102.6 | 102.15
-3.9 | 1.1 | 1.9 | 6.2 | 71.22 | 99.8 | 102.26 | 102.5 | 102.05
-3.8 | 1.1 | 2 | 7.2 | 80.46 | 100 | 102.2 | 102.5 | 101.95
-3.7 | 1.2 | 2.16 | 8.6 | 87.22 | 100.1 | 102.1 | 102.4 | 101.9
-3.6 | 1.2 | 2.26 | 10.6 | 91.68 | 100.3 | 102.06 | 102.3 | 101.85
-3.5 | 1.3 | 2.46 | 13.6 | 94.2 | 100.4 | 102 | 102.2 | 101.8
-3.4 | 1.4 | 2.6 | 18.3 | 96.06 | 100.5 | 101.96 | 102.2 | 101.75
-3.3 | 1.5 | 2.86 | 25.3 | 97.1 | 100.6 | 101.9 | 102.1 | 101.65
-3.2 | 1.6 | 3.22 | 34.8 | 98.08 | 100.7 | 101.86 | 102 | 101.55
-3.1 | 1.7 | 3.62 | 45.8 | 98.72 | 100.7 | 101.8 | 101.9 | 101.5
-3 | 1.8 | 4.08 | 57.3 | 99.04 | 100.8 | 101.7 | 101.8 | 101.45
-2.9 | 2 | 4.74 | 68.2 | 99.52 | 100.8 | 101.66 | 101.8 | 101.4
-2.8 | 2.2 | 5.7 | 78 | 99.72 | 100.8 | 101.6 | 101.7 | 101.35
-2.7 | 2.5 | 7.02 | 85.6 | 99.84 | 100.8 | 101.56 | 101.6 | 101.25
-2.6 | 2.8 | 8.86 | 90.6 | 100.1 | 100.8 | 101.5 | 101.6 | 101.2
-2.5 | 3.3 | 11.62 | 93.6 | 100.2 | 100.8 | 101.4 | 101.5 | 101.15
-2.4 | 3.9 | 15.8 | 95.6 | 100.22 | 100.7 | 101.3 | 101.5 | 101.05
-2.3 | 4.7 | 22.5 | 97 | 100.3 | 100.7 | 101.26 | 101.4 | 101
-2.2 | 5.9 | 31.8 | 97.8 | 100.32 | 100.7 | 101.2 | 101.3 | 100.95
-2.1 | 7.6 | 42.68 | 98.4 | 100.4 | 100.8 | 101.1 | 101.3 | 100.85
-2 | 10.3 | 54.38 | 98.8 | 100.4 | 100.7 | 101.1 | 101.2 | 100.8
-1.9 | 14.3 | 65.76 | 99.2 | 100.4 | 100.7 | 101 | 101.1 | 100.8
-1.8 | 20.4 | 75.98 | 99.5 | 100.4 | 100.7 | 101 | 101 | 100.75
-1.7 | 29.5 | 84.32 | 99.8 | 100.4 | 100.6 | 100.9 | 101 | 100.7
-1.6 | 40.3 | 89.92 | 99.9 | 100.4 | 100.6 | 100.86 | 100.9 | 100.65
-1.5 | 52.1 | 93.34 | 100 | 100.4 | 100.5 | 100.76 | 100.8 | 100.55
-1.4 | 63.7 | 95.38 | 100 | 100.4 | 100.5 | 100.7 | 100.7 | 100.45
-1.3 | 74.5 | 96.92 | 100.1 | 100.4 | 100.4 | 100.6 | 100.7 | 100.4
-1.2 | 83 | 97.82 | 100.1 | 100.3 | 100.3 | 100.6 | 100.6 | 100.4
-1.1 | 89.3 | 98.48 | 100.1 | 100.3 | 100.3 | 100.5 | 100.5 | 100.35
-1 | 93.2 | 98.98 | 100.1 | 100.3 | 100.3 | 100.4 | 100.5 | 100.3
-0.9 | 95.5 | 99.26 | 100.1 | 100.28 | 100.2 | 100.36 | 100.4 | 100.25
-0.8 | 96.9 | 99.52 | 100.1 | 100.2 | 100.2 | 100.26 | 100.3 | 100.2
-0.7 | 98 | 99.66 | 100.1 | 100.18 | 100.1 | 100.2 | 100.3 | 100.15
-0.6 | 98.7 | 99.76 | 100.1 | 100.1 | 100.1 | 100.2 | 100.2 | 100.1
-0.5 | 99.2 | 99.86 | 100 | 100 | 100.1 | 100.2 | 100.2 | 100.1
-0.4 | 99.6 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.05
-0.3 | 99.7 | 99.96 | 100 | 100 | 100.1 | 100.06 | 100.1 | 100
-0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100.1 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 99.9 | 100 | 100 | 99.9 | 100 | 100 | 100
0.3 | 99.7 | 99.84 | 100 | 100 | 99.9 | 100 | 100.1 | 100
0.4 | 99.5 | 99.8 | 99.9 | 100 | 99.9 | 100.04 | 100.1 | 100.05
0.5 | 99.2 | 99.74 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100.1
0.6 | 98.7 | 99.64 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100.1
0.7 | 98 | 99.54 | 99.9 | 100 | 100 | 100.1 | 100.2 | 100.15
0.8 | 96.9 | 99.34 | 99.9 | 100 | 100 | 100.1 | 100.2 | 100.2
0.9 | 95.3 | 99.08 | 99.9 | 100 | 100.1 | 100.1 | 100.2 | 100.25
1 | 93.1 | 98.78 | 99.9 | 100.1 | 100.1 | 100.2 | 100.2 | 100.3
1.1 | 89.3 | 98.28 | 99.9 | 100.1 | 100.1 | 100.2 | 100.3 | 100.35
1.2 | 82.9 | 97.46 | 99.8 | 100.1 | 100.1 | 100.3 | 100.3 | 100.4
1.3 | 74.3 | 96.64 | 99.8 | 100.12 | 100.2 | 100.3 | 100.4 | 100.4
1.4 | 63.5 | 95.12 | 99.7 | 100.2 | 100.3 | 100.4 | 100.5 | 100.45
1.5 | 51.9 | 92.86 | 99.7 | 100.2 | 100.3 | 100.44 | 100.5 | 100.55
1.6 | 40.2 | 89.22 | 99.6 | 100.2 | 100.3 | 100.5 | 100.6 | 100.65
1.7 | 29.3 | 83.62 | 99.4 | 100.2 | 100.3 | 100.6 | 100.7 | 100.7
1.8 | 20.4 | 75.28 | 99.2 | 100.1 | 100.4 | 100.6 | 100.7 | 100.75
1.9 | 14.2 | 64.94 | 99 | 100.1 | 100.4 | 100.7 | 100.8 | 100.8
2 | 10.2 | 53.58 | 98.6 | 100.1 | 100.4 | 100.7 | 100.8 | 100.8
2.1 | 7.6 | 42.02 | 98.1 | 100.1 | 100.4 | 100.8 | 100.8 | 100.85
2.2 | 5.9 | 31.24 | 97.4 | 100 | 100.4 | 100.8 | 100.9 | 100.95
2.3 | 4.7 | 22 | 96.6 | 100 | 100.4 | 100.8 | 100.9 | 101
2.4 | 3.9 | 15.4 | 95.3 | 99.9 | 100.4 | 100.9 | 101 | 101.05
2.5 | 3.2 | 11.38 | 93.2 | 99.88 | 100.4 | 100.94 | 101 | 101.1
2.6 | 2.8 | 8.64 | 90.2 | 99.78 | 100.4 | 101 | 101.1 | 101.15
2.7 | 2.5 | 6.88 | 85.1 | 99.58 | 100.5 | 101.04 | 101.2 | 101.25
2.8 | 2.2 | 5.6 | 77.6 | 99.38 | 100.5 | 101.2 | 101.2 | 101.35
2.9 | 2 | 4.66 | 68.1 | 99.16 | 100.5 | 101.2 | 101.3 | 101.4
3 | 1.8 | 4.02 | 57 | 98.66 | 100.4 | 101.2 | 101.4 | 101.45
3.1 | 1.7 | 3.58 | 45.5 | 98.26 | 100.4 | 101.3 | 101.5 | 101.5
3.2 | 1.5 | 3.18 | 34.8 | 97.72 | 100.3 | 101.3 | 101.5 | 101.55
3.3 | 1.5 | 2.84 | 25.3 | 96.84 | 100.2 | 101.4 | 101.5 | 101.65
3.4 | 1.4 | 2.64 | 18.1 | 95.76 | 100.1 | 101.4 | 101.6 | 101.75
3.5 | 1.3 | 2.44 | 13.4 | 93.98 | 100 | 101.5 | 101.7 | 101.8
3.6 | 1.3 | 2.24 | 10.5 | 91.48 | 99.9 | 101.54 | 101.7 | 101.85
3.7 | 1.2 | 2.14 | 8.6 | 87.42 | 99.7 | 101.6 | 101.8 | 101.9
3.8 | 1.1 | 1.94 | 7.2 | 80.7 | 99.5 | 101.7 | 101.9 | 101.95
3.9 | 1.1 | 1.84 | 6.2 | 71.66 | 99.3 | 101.7 | 102 | 102.05
4 | 1.1 | 1.8 | 5.4 | 61.4 | 99.1 | 101.8 | 102 | 102.15
4.1 | 1 | 1.7 | 4.9 | 50.36 | 98.6 | 101.8 | 102.1 | 102.25
4.2 | 1 | 1.64 | 4.4 | 39.36 | 98.1 | 101.84 | 102.1 | 102.35
4.3 | 0.9 | 1.54 | 4.1 | 29.44 | 97.3 | 101.9 | 102.2 | 102.4
4.4 | 0.9 | 1.5 | 3.7 | 21.48 | 96.3 | 102 | 102.2 | 102.45
4.5 | 0.9 | 1.44 | 3.5 | 16 | 94.9 | 102.04 | 102.3 | 102.55
4.6 | 0.9 | 1.4 | 3.3 | 12.54 | 92.6 | 102.1 | 102.3 | 102.6
4.7 | 0.8 | 1.3 | 3.1 | 10.24 | 89.1 | 102.14 | 102.4 | 102.65
4.8 | 0.8 | 1.3 | 3 | 8.68 | 83.5 | 102.2 | 102.5 | 102.75
4.9 | 0.8 | 1.2 | 2.8 | 7.6 | 75.3 | 102.2 | 102.5 | 102.8
5 | 0.8 | 1.2 | 2.7 | 6.74 | 65.3 | 102.3 | 102.6 | 102.85
5.1 | 0.8 | 1.2 | 2.6 | 6.14 | 54.6 | 102.3 | 102.7 | 102.95
5.2 | 0.7 | 1.1 | 2.5 | 5.56 | 43.4 | 102.3 | 102.7 | 103
5.3 | 0.7 | 1.1 | 2.4 | 5.16 | 33.1 | 102.4 | 102.8 | 103
5.4 | 0.7 | 1.04 | 2.3 | 4.78 | 24.4 | 102.4 | 102.8 | 103.05
5.5 | 0.7 | 1 | 2.2 | 4.48 | 18.1 | 102.4 | 102.9 | 103.1
5.6 | 0.7 | 1 | 2.1 | 4.26 | 14.2 | 102.4 | 102.9 | 103.15
5.7 | 0.7 | 1 | 2 | 4 | 11.6 | 102.4 | 102.9 | 103.25
5.8 | 0.7 | 0.9 | 1.9 | 3.88 | 10 | 102.4 | 103 | 103.3
5.9 | 0.6 | 0.9 | 1.9 | 3.68 | 8.7 | 102.5 | 103 | 103.3
6 | 0.6 | 0.9 | 1.8 | 3.5 | 7.8 | 102.5 | 103 | 103.35
6.1 | 0.6 | 0.9 | 1.7 | 3.38 | 7.1 | 102.5 | 103 | 103.4
6.2 | 0.6 | 0.9 | 1.7 | 3.28 | 6.5 | 102.5 | 103.1 | 103.4
6.3 | 0.6 | 0.9 | 1.7 | 3.1 | 6.1 | 102.5 | 103.1 | 103.45
6.4 | 0.6 | 0.84 | 1.6 | 3 | 5.7 | 102.5 | 103.2 | 103.5
6.5 | 0.6 | 0.8 | 1.5 | 2.88 | 5.4 | 102.5 | 103.2 | 103.55
6.6 |  | 0.8 | 1.5 | 2.8 | 5.1 | 102.5 | 103.2 | 103.6
6.7 |  | 0.8 | 1.4 | 2.7 | 4.9 | 102.5 | 103.2 | 103.6
6.8 |  | 0.8 | 1.4 | 2.6 | 4.6 | 102.5 | 103.3 | 103.65
6.9 |  | 0.8 | 1.4 | 2.5 | 4.4 | 102.5 | 103.4 | 103.7
7 |  | 0.8 | 1.3 | 2.4 | 4.2 | 102.46 | 103.4 | 103.75
7.1 |  |  | 1.3 | 2.38 | 4.1 | 102.4 | 103.4 | 103.8
7.2 |  |  | 1.3 | 2.28 | 4 | 102.4 | 103.4 | 103.85
7.3 |  |  | 1.2 | 2.2 | 3.8 | 102.4 | 103.5 | 103.9
7.4 |  |  | 1.2 | 2.1 | 3.7 | 102.4 | 103.5 | 103.9
7.5 |  |  | 1.2 | 2.08 | 3.6 | 102.4 | 103.5 | 103.95
7.6 |  |  | 1.1 | 2 | 3.5 | 102.36 | 103.5 | 104
7.7 |  |  | 1.1 | 1.98 | 3.3 | 102.3 | 103.6 | 104
7.8 |  |  | 1.1 | 1.9 | 3.2 | 102.26 | 103.6 | 104
7.9 |  |  | 1.1 | 1.88 | 3.1 | 102.2 | 103.6 | 104
8 |  |  | 1.1 | 1.8 | 3 | 102.16 | 103.6 | 104.05
8.1 |  |  |  | 1.7 | 2.9 | 102.06 | 103.6 | 104.1
8.2 |  |  |  | 1.7 | 2.8 | 102 | 103.6 | 104.1
8.3 |  |  |  | 1.68 | 2.7 | 101.96 | 103.7 | 104.1
8.4 |  |  |  | 1.6 | 2.7 | 101.86 | 103.7 | 104.15
8.5 |  |  |  | 1.6 | 2.6 | 101.76 | 103.7 | 104.2
8.6 |  |  |  | 1.5 | 2.5 | 101.62 | 103.7 | 104.2
8.7 |  |  |  | 1.5 | 2.4 | 101.46 | 103.7 | 104.2
8.8 |  |  |  | 1.5 | 2.4 | 101.26 | 103.7 | 104.25
8.9 |  |  |  | 1.4 | 2.3 | 101.06 | 103.7 | 104.3
9 |  |  |  | 1.4 | 2.2 | 100.82 | 103.8 | 104.3
9.1 |  |  |  |  | 2.2 | 100.46 | 103.8 | 104.3
9.2 |  |  |  |  | 2.1 | 100.12 | 103.8 | 104.35
9.3 |  |  |  |  | 2.1 | 99.64 | 103.8 | 104.4
9.4 |  |  |  |  | 2 | 98.94 | 103.8 | 104.45
9.5 |  |  |  |  | 2 | 98.14 | 103.8 | 104.5
9.6 |  |  |  |  | 1.9 | 96.84 | 103.8 | 104.5
9.7 |  |  |  |  | 1.9 | 94.96 | 103.8 | 104.55
9.8 |  |  |  |  | 1.8 | 92.24 | 103.8 | 104.6
9.9 |  |  |  |  | 1.8 | 87.6 | 103.8 | 104.6
10 |  |  |  |  | 1.7 | 80.42 | 103.8 | 104.65
10.1 | [5 empty] | 71.34 | 103.8 | 104.7
10.2 | [5 empty] | 60.9 | 103.8 | 104.7
10.3 | [5 empty] | 49.84 | 103.8 | 104.7
10.4 | [5 empty] | 39.34 | 103.8 | 104.7
10.5 | [5 empty] | 30.16 | 103.8 | 104.7
10.6 | [5 empty] | 22.94 | 103.8 | 104.7
10.7 | [5 empty] | 18.18 | 103.8 | 104.75
10.8 | [5 empty] | 15.16 | 103.8 | 104.8
10.9 | [5 empty] | 13.22 | 103.8 | 104.8
11 | [5 empty] | 11.9 | 103.8 | 104.8
11.1 | [5 empty] | 10.9 | 103.8 | 104.85
11.2 | [5 empty] | 10.08 | 103.8 | 104.9
11.3 | [5 empty] | 9.48 | 103.8 | 104.9
11.4 | [5 empty] | 9.02 | 103.8 | 104.9
11.5 | [5 empty] | 8.62 | 103.8 | 104.9
11.6 | [5 empty] | 8.22 | 103.8 | 104.9
11.7 | [5 empty] | 7.92 | 103.8 | 104.95
11.8 | [5 empty] | 7.62 | 103.8 | 105
11.9 | [5 empty] | 7.36 | 103.7 | 105
12 | [5 empty] | 7.16 | 103.7 | 105
12.1 | [5 empty] | 6.96 | 103.7 | 105
12.2 | [5 empty] | 6.66 | 103.6 | 105
12.3 | [5 empty] | 6.56 | 103.6 | 105
12.4 | [5 empty] | 6.36 | 103.6 | 105.05
12.5 | [5 empty] | 6.16 | 103.5 | 105.1
12.6 | [5 empty] | 5.96 | 103.4 | 105.1
12.7 | [5 empty] | 5.76 | 103.4 | 105.1
12.8 | [5 empty] | 5.66 | 103.4 | 105.1
12.9 | [5 empty] | 5.46 | 103.3 | 105.1
13 | [5 empty] | 5.36 | 103.3 | 105.1
13.1 | [5 empty] | 5.2 | 103.2 | 105.1
13.2 | [5 empty] | 5.1 | 103.1 | 105.1
13.3 | [5 empty] | 4.96 | 103.1 | 105.15
13.4 | [5 empty] | 4.86 | 103.1 | 105.2
13.5 | [5 empty] | 4.7 | 103 | 105.2
13.6 | [5 empty] | 4.56 | 102.9 | 105.2
13.7 | [5 empty] | 4.46 | 102.8 | 105.2
13.8 | [5 empty] | 4.36 | 102.7 | 105.2
13.9 | [5 empty] | 4.2 | 102.6 | 105.25
14 | [5 empty] | 4.16 | 102.4 | 105.3
14.1 | [5 empty] | 4 | 102.2 | 105.3
14.2 | [5 empty] | 3.96 | 101.9 | 105.3
14.3 | [5 empty] | 3.86 | 101.6 | 105.35
14.4 | [5 empty] | 3.7 | 101.2 | 105.4
14.5 | [5 empty] | 3.66 | 100.8 | 105.4
14.6 | [5 empty] | 3.56 | 99.9 | 105.4
14.7 | [5 empty] | 3.5 | 98.9 | 105.4
14.8 | [5 empty] | 3.4 | 97.4 | 105.45
14.9 | [5 empty] | 3.3 | 95.1 | 105.5
15 | [5 empty] | 3.2 | 91.2 | 105.5
15.1 | [5 empty] | 3.2 | 84.8 | 105.5
15.2 | [5 empty] | 3.1 | 76.2 | 105.5
15.3 | [6 empty] | 65.9 | 105.55
15.4 | [6 empty] | 55 | 105.6
15.5 | [6 empty] | 44.1 | 105.6
15.6 | [6 empty] | 34.5 | 105.6
15.7 | [6 empty] | 26.3 | 105.65
15.8 | [6 empty] | 20.7 | 105.7
15.9 | [6 empty] | 17.3 | 105.75
16 | [6 empty] | 15.2 | 105.8
16.1 | [6 empty] | 13.7 | 105.8
16.2 | [6 empty] | 12.6 | 105.85
16.3 | [6 empty] | 11.8 | 105.9
16.4 | [6 empty] | 11.2 | 105.9
16.5 | [6 empty] | 10.7 | 105.95
16.6 | [6 empty] | 10.3 | 106
16.7 | [6 empty] | 9.9 | 106
16.8 | [6 empty] | 9.6 | 106
16.9 | [6 empty] | 9.3 | 106.05
17 | [6 empty] | 9 | 106.1
17.1 | [6 empty] | 8.8 | 106.1
17.2 | [6 empty] | 8.6 | 106.1
17.3 | [6 empty] | 8.4 | 106.1
17.4 | [6 empty] | 8.2 | 106.1
17.5 | [6 empty] | 7.9 | 106.1
17.6 | [6 empty] | 7.8 | 106.1
17.7 | [6 empty] | 7.6 | 106.1
17.8 | [6 empty] | 7.4 | 106.1
17.9 | [6 empty] | 7.2 | 106.1
18 | [6 empty] | 7.1 | 106.05
18.1 | [6 empty] | 6.9 | 106
18.2 | [6 empty] | 6.7 | 105.95
18.3 | [6 empty] | 6.6 | 105.9
18.4 | [6 empty] | 6.4 | 105.85
18.5 | [6 empty] | 6.3 | 105.75
18.6 | [6 empty] | 6.1 | 105.65
18.7 | [6 empty] | 6 | 105.55
18.8 | [6 empty] | 5.9 | 105.45
18.9 | [6 empty] | 5.7 | 105.3
19 | [6 empty] | 5.6 | 105.1
19.1 | [6 empty] | 5.5 | 104.9
19.2 | [6 empty] | 5.4 | 104.75
19.3 | [6 empty] | 5.2 | 104.5
19.4 | [6 empty] | 5.1 | 104.1
19.5 | [6 empty] | 5 | 103.65
19.6 | [6 empty] | 4.9 | 103.05
19.7 | [6 empty] | 4.8 | 102.3
19.8 | [6 empty] | 4.7 | 101.3
19.9 | [6 empty] | 4.6 | 99.9
20 | [6 empty] | 4.5 | 97.7
20.1 | [6 empty] | 4.4 | 94.05
20.2 | [6 empty] | 4.3 | 88.2
20.3 | [6 empty] | 4.2 | 79.9
20.4 | [7 empty] | 69.9
20.5 | [7 empty] | 58.95
20.6 | [7 empty] | 47.9
20.7 | [7 empty] | 37.7
20.8 | [7 empty] | 29.05
20.9 | [7 empty] | 22.7
21 | [7 empty] | 18.7
21.1 | [7 empty] | 16.25
21.2 | [7 empty] | 14.6
21.3 | [7 empty] | 13.4
21.4 | [7 empty] | 12.55
21.5 | [7 empty] | 11.9
21.6 | [7 empty] | 11.35
21.7 | [7 empty] | 10.9
21.8 | [7 empty] | 10.5
21.9 | [7 empty] | 10.15
22 | [7 empty] | 9.85
22.1 | [7 empty] | 9.55
22.2 | [7 empty] | 9.3
22.3 | [7 empty] | 9.05
22.4 | [7 empty] | 8.8
22.5 | [7 empty] | 8.55
22.6 | [7 empty] | 8.3
22.7 | [7 empty] | 8.1
22.8 | [7 empty] | 7.9
22.9 | [7 empty] | 7.7
23 | [7 empty] | 7.5
23.1 | [7 empty] | 7.3
23.2 | [7 empty] | 7.1
23.3 | [7 empty] | 6.9
23.4 | [7 empty] | 6.7
23.5 | [7 empty] | 6.5
23.6 | [7 empty] | 6.35
23.7 | [7 empty] | 6.2
23.8 | [7 empty] | 6
23.9 | [7 empty] | 5.85
24 | [7 empty] | 5.7
24.1 | [7 empty] | 5.55
24.2 | [7 empty] | 5.4
24.3 | [7 empty] | 5.25
24.4 | [7 empty] | 5.15
24.5 | [7 empty] | 5.05
24.6 | [7 empty] | 4.9
24.7 | [7 empty] | 4.75
24.8 | [7 empty] | 4.65
24.9 | [7 empty] | 4.55
25 | [7 empty] | 4.45
25.1 | [7 empty] | 4.35
25.2 | [7 empty] | 4.25
25.3 | [7 empty] | 4.15
25.4 | [7 empty] | 4.05

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.9 | [7 empty] | 4.78
-25.8 | [7 empty] | 4.88
-25.7 | [7 empty] | 4.98
-25.6 | [7 empty] | 5.08
-25.5 | [7 empty] | 5.18
-25.4 | [7 empty] | 5.28
-25.3 | [7 empty] | 5.38
-25.2 | [7 empty] | 5.56
-25.1 | [7 empty] | 5.68
-25 | [7 empty] | 5.78
-24.9 | [7 empty] | 5.96
-24.8 | [7 empty] | 6.08
-24.7 | [7 empty] | 6.18
-24.6 | [7 empty] | 6.36
-24.5 | [7 empty] | 6.48
-24.4 | [7 empty] | 6.66
-24.3 | [7 empty] | 6.86
-24.2 | [7 empty] | 7.06
-24.1 | [7 empty] | 7.18
-24 | [7 empty] | 7.36
-23.9 | [7 empty] | 7.56
-23.8 | [7 empty] | 7.76
-23.7 | [7 empty] | 7.96
-23.6 | [7 empty] | 8.16
-23.5 | [7 empty] | 8.36
-23.4 | [7 empty] | 8.56
-23.3 | [7 empty] | 8.84
-23.2 | [7 empty] | 9.06
-23.1 | [7 empty] | 9.26
-23 | [7 empty] | 9.54
-22.9 | [7 empty] | 9.76
-22.8 | [7 empty] | 10.04
-22.7 | [7 empty] | 10.34
-22.6 | [7 empty] | 10.64
-22.5 | [7 empty] | 11.02
-22.4 | [7 empty] | 11.34
-22.3 | [7 empty] | 11.72
-22.2 | [7 empty] | 12.2
-22.1 | [7 empty] | 12.7
-22 | [7 empty] | 13.36
-21.9 | [7 empty] | 14.22
-21.8 | [7 empty] | 15.2
-21.7 | [7 empty] | 16.52
-21.6 | [7 empty] | 18.4
-21.5 | [7 empty] | 21.2
-21.4 | [7 empty] | 25.64
-21.3 | [7 empty] | 32.28
-21.2 | [7 empty] | 40.82
-21.1 | [7 empty] | 50.52
-21 | [7 empty] | 60.74
-20.9 | [7 empty] | 70.56
-20.8 | [7 empty] | 79.38
-20.7 | [6 empty] | 4.6 | 86.46
-20.6 | [6 empty] | 4.7 | 91.24
-20.5 | [6 empty] | 4.8 | 94.18
-20.4 | [6 empty] | 4.92 | 96.14
-20.3 | [6 empty] | 5 | 97.54
-20.2 | [6 empty] | 5.1 | 98.6
-20.1 | [6 empty] | 5.2 | 99.36
-20 | [6 empty] | 5.32 | 99.98
-19.9 | [6 empty] | 5.42 | 100.58
-19.8 | [6 empty] | 5.52 | 101.02
-19.7 | [6 empty] | 5.7 | 101.34
-19.6 | [6 empty] | 5.8 | 101.64
-19.5 | [6 empty] | 5.92 | 101.94
-19.4 | [6 empty] | 6.02 | 102.16
-19.3 | [6 empty] | 6.2 | 102.36
-19.2 | [6 empty] | 6.32 | 102.56
-19.1 | [6 empty] | 6.5 | 102.76
-19 | [6 empty] | 6.6 | 102.88
-18.9 | [6 empty] | 6.72 | 102.98
-18.8 | [6 empty] | 6.92 | 103.08
-18.7 | [6 empty] | 7.1 | 103.18
-18.6 | [6 empty] | 7.22 | 103.28
-18.5 | [6 empty] | 7.42 | 103.38
-18.4 | [6 empty] | 7.62 | 103.48
-18.3 | [6 empty] | 7.74 | 103.58
-18.2 | [6 empty] | 7.92 | 103.6
-18.1 | [6 empty] | 8.12 | 103.6
-18 | [6 empty] | 8.32 | 103.68
-17.9 | [6 empty] | 8.54 | 103.7
-17.8 | [6 empty] | 8.82 | 103.7
-17.7 | [6 empty] | 9.02 | 103.78
-17.6 | [6 empty] | 9.22 | 103.8
-17.5 | [6 empty] | 9.52 | 103.8
-17.4 | [6 empty] | 9.74 | 103.8
-17.3 | [6 empty] | 10.04 | 103.8
-17.2 | [6 empty] | 10.34 | 103.8
-17.1 | [6 empty] | 10.74 | 103.8
-17 | [6 empty] | 11.14 | 103.8
-16.9 | [6 empty] | 11.56 | 103.8
-16.8 | [6 empty] | 12.16 | 103.8
-16.7 | [6 empty] | 12.86 | 103.8
-16.6 | [6 empty] | 13.7 | 103.8
-16.5 | [6 empty] | 14.74 | 103.8
-16.4 | [6 empty] | 16.28 | 103.88
-16.3 | [6 empty] | 18.36 | 103.82
-16.2 | [6 empty] | 21.52 | 103.8
-16.1 | [6 empty] | 26.54 | 103.8
-16 | [6 empty] | 33.84 | 103.8
-15.9 | [6 empty] | 42.74 | 103.72
-15.8 | [6 empty] | 52.52 | 103.7
-15.7 | [6 empty] | 62.8 | 103.7
-15.6 | [6 empty] | 72.48 | 103.7
-15.5 | [5 empty] | 3.3 | 80.62 | 103.7
-15.4 | [5 empty] | 3.4 | 86.98 | 103.7
-15.3 | [5 empty] | 3.5 | 91.22 | 103.7
-15.2 | [5 empty] | 3.5 | 93.88 | 103.7
-15.1 | [5 empty] | 3.6 | 95.62 | 103.7
-15 | [5 empty] | 3.7 | 96.8 | 103.7
-14.9 | [5 empty] | 3.8 | 97.7 | 103.7
-14.8 | [5 empty] | 3.9 | 98.46 | 103.7
-14.7 | [5 empty] | 4 | 98.96 | 103.7
-14.6 | [5 empty] | 4.1 | 99.46 | 103.7
-14.5 | [5 empty] | 4.2 | 99.82 | 103.7
-14.4 | [5 empty] | 4.3 | 100.14 | 103.7
-14.3 | [5 empty] | 4.4 | 100.5 | 103.7
-14.2 | [5 empty] | 4.5 | 100.62 | 103.7
-14.1 | [5 empty] | 4.6 | 100.82 | 103.7
-14 | [5 empty] | 4.7 | 101.02 | 103.7
-13.9 | [5 empty] | 4.8 | 101.12 | 103.7
-13.8 | [5 empty] | 5 | 101.32 | 103.7
-13.7 | [5 empty] | 5 | 101.4 | 103.62
-13.6 | [5 empty] | 5.2 | 101.42 | 103.6
-13.5 | [5 empty] | 5.4 | 101.62 | 103.6
-13.4 | [5 empty] | 5.5 | 101.72 | 103.6
-13.3 | [5 empty] | 5.6 | 101.8 | 103.6
-13.2 | [5 empty] | 5.8 | 101.9 | 103.6
-13.1 | [5 empty] | 6 | 102 | 103.68
-13 | [5 empty] | 6.1 | 102 | 103.7
-12.9 | [5 empty] | 6.3 | 102.1 | 103.7
-12.8 | [5 empty] | 6.5 | 102.2 | 103.7
-12.7 | [5 empty] | 6.7 | 102.22 | 103.7
-12.6 | [5 empty] | 6.9 | 102.3 | 103.7
-12.5 | [5 empty] | 7.1 | 102.32 | 103.7
-12.4 | [5 empty] | 7.3 | 102.4 | 103.7
-12.3 | [5 empty] | 7.6 | 102.5 | 103.7
-12.2 | [5 empty] | 7.9 | 102.6 | 103.7
-12.1 | [5 empty] | 8.1 | 102.6 | 103.7
-12 | [5 empty] | 8.4 | 102.6 | 103.7
-11.9 | [5 empty] | 8.7 | 102.62 | 103.7
-11.8 | [5 empty] | 9.2 | 102.7 | 103.7
-11.7 | [5 empty] | 9.6 | 102.7 | 103.62
-11.6 | [5 empty] | 10.1 | 102.7 | 103.6
-11.5 | [5 empty] | 10.8 | 102.72 | 103.6
-11.4 | [5 empty] | 11.6 | 102.8 | 103.6
-11.3 | [5 empty] | 12.5 | 102.8 | 103.6
-11.2 | [5 empty] | 13.8 | 102.9 | 103.6
-11.1 | [5 empty] | 15.5 | 102.9 | 103.6
-11 | [5 empty] | 18.1 | 102.9 | 103.6
-10.9 | [5 empty] | 22.1 | 102.9 | 103.6
-10.8 | [5 empty] | 28.2 | 103 | 103.6
-10.7 | [5 empty] | 36.2 | 103 | 103.6
-10.6 | [5 empty] | 45.8 | 103.02 | 103.6
-10.5 | [5 empty] | 56.2 | 103.1 | 103.6
-10.4 | [5 empty] | 66.2 | 103.1 | 103.6
-10.3 | [5 empty] | 75.4 | 103.1 | 103.6
-10.2 |  |  |  |  | 1.8 | 83.1 | 103.1 | 103.52
-10.1 |  |  |  |  | 1.88 | 88.4 | 103.1 | 103.5
-10 |  |  |  |  | 1.9 | 91.9 | 103.1 | 103.5
-9.9 |  |  |  |  | 1.9 | 94.1 | 103.1 | 103.5
-9.8 |  |  |  |  | 2 | 95.6 | 103.1 | 103.5
-9.7 |  |  |  |  | 2 | 96.7 | 103.1 | 103.42
-9.6 |  |  |  |  | 2.1 | 97.6 | 103.1 | 103.4
-9.5 |  |  |  |  | 2.1 | 98.2 | 103.1 | 103.4
-9.4 |  |  |  |  | 2.2 | 98.9 | 103.1 | 103.32
-9.3 |  |  |  |  | 2.2 | 99.3 | 103.1 | 103.3
-9.2 |  |  |  |  | 2.3 | 99.6 | 103.1 | 103.3
-9.1 |  |  |  | 1.5 | 2.38 | 99.9 | 103.1 | 103.3
-9 |  |  |  | 1.5 | 2.48 | 100.2 | 103.1 | 103.3
-8.9 |  |  |  | 1.58 | 2.5 | 100.4 | 103.1 | 103.3
-8.8 |  |  |  | 1.6 | 2.6 | 100.7 | 103.2 | 103.3
-8.7 |  |  |  | 1.6 | 2.68 | 100.8 | 103.2 | 103.3
-8.6 |  |  |  | 1.68 | 2.78 | 101 | 103.2 | 103.3
-8.5 |  |  |  | 1.7 | 2.8 | 101.1 | 103.2 | 103.3
-8.4 |  |  |  | 1.7 | 2.9 | 101.3 | 103.2 | 103.3
-8.3 |  |  |  | 1.8 | 3 | 101.4 | 103.2 | 103.22
-8.2 |  |  |  | 1.8 | 3.1 | 101.5 | 103.2 | 103.2
-8.1 |  |  | 1.1 | 1.9 | 3.2 | 101.6 | 103.2 | 103.2
-8 |  |  | 1.1 | 1.9 | 3.3 | 101.6 | 103.2 | 103.2
-7.9 |  |  | 1.18 | 2 | 3.48 | 101.7 | 103.2 | 103.2
-7.8 |  |  | 1.2 | 2.08 | 3.58 | 101.8 | 103.2 | 103.12
-7.7 |  |  | 1.2 | 2.1 | 3.68 | 101.9 | 103.2 | 103.1
-7.6 |  |  | 1.2 | 2.2 | 3.8 | 102 | 103.2 | 103.1
-7.5 |  |  | 1.3 | 2.28 | 3.98 | 102 | 103.18 | 103.1
-7.4 |  |  | 1.3 | 2.3 | 4.18 | 102.1 | 103.1 | 103.02
-7.3 |  |  | 1.3 | 2.4 | 4.28 | 102.1 | 103.1 | 103
-7.2 |  |  | 1.4 | 2.5 | 4.48 | 102.1 | 103.1 | 103
-7.1 |  |  | 1.4 | 2.6 | 4.68 | 102.2 | 103 | 103
-7 |  | 0.8 | 1.48 | 2.68 | 4.88 | 102.2 | 103 | 103
-6.9 |  | 0.8 | 1.5 | 2.78 | 5.08 | 102.2 | 103 | 102.92
-6.8 |  | 0.8 | 1.5 | 2.88 | 5.38 | 102.2 | 103 | 102.9
-6.7 |  | 0.8 | 1.58 | 2.98 | 5.68 | 102.2 | 103 | 102.9
-6.6 |  | 0.88 | 1.6 | 3.08 | 6.06 | 102.3 | 103 | 102.9
-6.5 | 0.6 | 0.9 | 1.7 | 3.2 | 6.46 | 102.3 | 103 | 102.82
-6.4 | 0.6 | 0.9 | 1.78 | 3.38 | 6.94 | 102.3 | 103 | 102.8
-6.3 | 0.6 | 0.9 | 1.8 | 3.48 | 7.44 | 102.2 | 103 | 102.8
-6.2 | 0.6 | 0.9 | 1.88 | 3.68 | 8.22 | 102.3 | 102.9 | 102.72
-6.1 | 0.6 | 0.9 | 1.9 | 3.88 | 9.02 | 102.3 | 102.9 | 102.7
-6 | 0.7 | 1 | 2 | 3.98 | 10.18 | 102.3 | 102.9 | 102.7
-5.9 | 0.7 | 1 | 2.1 | 4.26 | 11.64 | 102.3 | 102.9 | 102.62
-5.8 | 0.7 | 1 | 2.2 | 4.48 | 13.76 | 102.3 | 102.8 | 102.6
-5.7 | 0.7 | 1 | 2.28 | 4.76 | 16.86 | 102.3 | 102.8 | 102.52
-5.6 | 0.7 | 1.1 | 2.38 | 5.06 | 21.56 | 102.3 | 102.8 | 102.5
-5.5 | 0.72 | 1.1 | 2.48 | 5.38 | 28.42 | 102.3 | 102.8 | 102.5
-5.4 | 0.8 | 1.18 | 2.58 | 5.76 | 37.36 | 102.3 | 102.7 | 102.42
-5.3 | 0.8 | 1.2 | 2.7 | 6.34 | 47.54 | 102.3 | 102.7 | 102.4
-5.2 | 0.8 | 1.2 | 2.8 | 6.94 | 58.12 | 102.3 | 102.7 | 102.32
-5.1 | 0.8 | 1.3 | 2.98 | 7.72 | 68.12 | 102.3 | 102.7 | 102.3
-5 | 0.9 | 1.3 | 3.08 | 8.7 | 77.16 | 102.3 | 102.6 | 102.3
-4.9 | 0.9 | 1.4 | 3.28 | 9.96 | 84.08 | 102.2 | 102.5 | 102.22
-4.8 | 0.9 | 1.4 | 3.48 | 11.62 | 88.8 | 102.2 | 102.5 | 102.2
-4.7 | 0.92 | 1.5 | 3.68 | 14.02 | 91.84 | 102.1 | 102.4 | 102.12
-4.6 | 1 | 1.5 | 3.98 | 17.58 | 93.84 | 102.1 | 102.4 | 102.1
-4.5 | 1 | 1.6 | 4.26 | 23.06 | 95.46 | 102.1 | 102.38 | 102.02
-4.4 | 1 | 1.7 | 4.58 | 30.68 | 96.5 | 102 | 102.3 | 101.92
-4.3 | 1.1 | 1.78 | 5.04 | 40.1 | 97.32 | 102 | 102.2 | 101.9
-4.2 | 1.1 | 1.88 | 5.56 | 50.34 | 97.92 | 101.9 | 102.2 | 101.82
-4.1 | 1.12 | 1.98 | 6.22 | 60.96 | 98.46 | 101.9 | 102.18 | 101.8
-4 | 1.2 | 2.08 | 7.02 | 70.82 | 98.86 | 101.9 | 102.1 | 101.8
-3.9 | 1.3 | 2.2 | 8.08 | 79.22 | 99.16 | 101.8 | 102.1 | 101.72
-3.8 | 1.3 | 2.36 | 9.54 | 85.64 | 99.38 | 101.8 | 102 | 101.7
-3.7 | 1.4 | 2.48 | 11.4 | 89.84 | 99.5 | 101.7 | 101.9 | 101.62
-3.6 | 1.5 | 2.68 | 14.36 | 92.66 | 99.78 | 101.7 | 101.88 | 101.6
-3.5 | 1.52 | 2.98 | 18.6 | 94.44 | 99.9 | 101.7 | 101.8 | 101.52
-3.4 | 1.62 | 3.18 | 25.16 | 95.78 | 100 | 101.6 | 101.8 | 101.42
-3.3 | 1.72 | 3.56 | 33.68 | 96.8 | 100.1 | 101.5 | 101.78 | 101.4
-3.2 | 1.92 | 4.04 | 43.84 | 97.54 | 100.28 | 101.5 | 101.7 | 101.32
-3.1 | 2.1 | 4.54 | 54.62 | 98.04 | 100.3 | 101.4 | 101.6 | 101.22
-3 | 2.22 | 5.22 | 65.28 | 98.46 | 100.3 | 101.4 | 101.58 | 101.2
-2.9 | 2.52 | 6.1 | 74.8 | 98.86 | 100.4 | 101.3 | 101.5 | 101.12
-2.8 | 2.82 | 7.36 | 82.68 | 99.18 | 100.4 | 101.3 | 101.4 | 101.1
-2.7 | 3.14 | 9.02 | 88.12 | 99.38 | 100.5 | 101.2 | 101.4 | 101.1
-2.6 | 3.66 | 11.42 | 91.6 | 99.58 | 100.5 | 101.2 | 101.3 | 101.02
-2.5 | 4.26 | 14.9 | 93.88 | 99.78 | 100.58 | 101.2 | 101.3 | 100.92
-2.4 | 5 | 20.28 | 95.38 | 99.8 | 100.6 | 101.2 | 101.28 | 100.9
-2.3 | 6.04 | 27.96 | 96.58 | 99.9 | 100.6 | 101.1 | 101.2 | 100.82
-2.2 | 7.58 | 37.76 | 97.42 | 100 | 100.6 | 101 | 101.08 | 100.8
-2.1 | 9.76 | 48.86 | 98.14 | 100.1 | 100.6 | 101 | 101 | 100.8
-2 | 12.8 | 59.9 | 98.48 | 100.1 | 100.6 | 100.9 | 101 | 100.72
-1.9 | 17.62 | 70.38 | 98.88 | 100.1 | 100.6 | 100.8 | 100.9 | 100.7
-1.8 | 24.68 | 79.28 | 99.18 | 100.1 | 100.5 | 100.8 | 100.88 | 100.62
-1.7 | 34.14 | 85.92 | 99.46 | 100.2 | 100.5 | 100.7 | 100.8 | 100.6
-1.6 | 44.82 | 90.42 | 99.58 | 100.2 | 100.5 | 100.7 | 100.78 | 100.52
-1.5 | 56.12 | 93.16 | 99.78 | 100.2 | 100.4 | 100.7 | 100.7 | 100.42
-1.4 | 67.2 | 95.2 | 99.8 | 100.2 | 100.4 | 100.6 | 100.6 | 100.4
-1.3 | 76.74 | 96.4 | 99.9 | 100.2 | 100.4 | 100.5 | 100.5 | 100.32
-1.2 | 84.5 | 97.32 | 99.98 | 100.2 | 100.32 | 100.4 | 100.4 | 100.3
-1.1 | 89.88 | 98.04 | 100 | 100.2 | 100.3 | 100.4 | 100.4 | 100.22
-1 | 93.12 | 98.64 | 100 | 100.2 | 100.3 | 100.3 | 100.38 | 100.2
-0.9 | 95.36 | 99.08 | 100.1 | 100.2 | 100.22 | 100.3 | 100.3 | 100.2
-0.8 | 96.94 | 99.28 | 100.1 | 100.1 | 100.2 | 100.3 | 100.28 | 100.12
-0.7 | 97.98 | 99.4 | 100.1 | 100.1 | 100.1 | 100.3 | 100.2 | 100.1
-0.6 | 98.54 | 99.58 | 100.1 | 100.1 | 100.1 | 100.2 | 100.2 | 100.1
-0.5 | 99.06 | 99.78 | 100.1 | 100 | 100.02 | 100.1 | 100.2 | 100.1
-0.4 | 99.5 | 99.8 | 100.1 | 100 | 100 | 100.1 | 100.1 | 100.02
-0.3 | 99.64 | 99.9 | 100.02 | 100 | 100 | 100.1 | 100.08 | 100
-0.2 | 99.82 | 99.98 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.88 | 99.92 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.58 | 99.82 | 100 | 100 | 100 | 99.9 | 100 | 100
0.4 | 99.38 | 99.72 | 100 | 100 | 100 | 100 | 100 | 100.07
0.5 | 99.04 | 99.62 | 100 | 100 | 100 | 100 | 100 | 100.1
0.6 | 98.44 | 99.5 | 99.92 | 100 | 100 | 100 | 100.02 | 100.1
0.7 | 97.72 | 99.32 | 99.9 | 100 | 100 | 100 | 100.1 | 100.1
0.8 | 96.76 | 99.12 | 99.9 | 100 | 100.02 | 100.1 | 100.1 | 100.17
0.9 | 95.22 | 98.92 | 99.9 | 100 | 100.08 | 100.1 | 100.2 | 100.2
1 | 93.1 | 98.46 | 99.8 | 100 | 100.1 | 100.1 | 100.2 | 100.2
1.1 | 89.88 | 97.94 | 99.72 | 100 | 100.1 | 100.1 | 100.2 | 100.27
1.2 | 84.72 | 97.1 | 99.7 | 100 | 100.1 | 100.2 | 100.3 | 100.3
1.3 | 77.02 | 96.2 | 99.7 | 100.08 | 100.18 | 100.3 | 100.3 | 100.37
1.4 | 67.54 | 94.86 | 99.6 | 100.1 | 100.2 | 100.4 | 100.4 | 100.4
1.5 | 56.64 | 92.9 | 99.52 | 100.1 | 100.2 | 100.4 | 100.42 | 100.4
1.6 | 45.28 | 90.04 | 99.42 | 100 | 100.2 | 100.4 | 100.5 | 100.47
1.7 | 34.42 | 85.52 | 99.12 | 100 | 100.28 | 100.5 | 100.5 | 100.57
1.8 | 25.08 | 78.92 | 98.92 | 100 | 100.3 | 100.5 | 100.52 | 100.6
1.9 | 17.84 | 69.78 | 98.62 | 99.9 | 100.3 | 100.5 | 100.6 | 100.67
2 | 12.94 | 59.3 | 98.24 | 99.82 | 100.3 | 100.5 | 100.62 | 100.7
2.1 | 9.86 | 48.22 | 97.74 | 99.8 | 100.3 | 100.6 | 100.72 | 100.77
2.2 | 7.64 | 37.28 | 96.98 | 99.72 | 100.2 | 100.7 | 100.8 | 100.8
2.3 | 6.18 | 27.62 | 96.18 | 99.62 | 100.2 | 100.8 | 100.82 | 100.87
2.4 | 5.1 | 19.9 | 95.04 | 99.52 | 100.2 | 100.8 | 100.9 | 100.9
2.5 | 4.24 | 14.64 | 93.4 | 99.42 | 100.2 | 100.8 | 100.9 | 100.97
2.6 | 3.64 | 11.18 | 90.98 | 99.3 | 100.12 | 100.9 | 100.92 | 101
2.7 | 3.16 | 8.9 | 87.34 | 99.12 | 100.1 | 100.9 | 101 | 101.07
2.8 | 2.76 | 7.34 | 81.74 | 98.92 | 100.1 | 100.9 | 101 | 101.1
2.9 | 2.48 | 6.12 | 74.04 | 98.54 | 100.1 | 101 | 101.1 | 101.17
3 | 2.28 | 5.28 | 64.3 | 98.14 | 100 | 101 | 101.1 | 101.2
3.1 | 2.08 | 4.56 | 53.68 | 97.66 | 99.92 | 101 | 101.2 | 101.2
3.2 | 1.9 | 4.04 | 42.9 | 97.24 | 99.9 | 101.1 | 101.2 | 101.27
3.3 | 1.78 | 3.54 | 33.04 | 96.48 | 99.8 | 101.1 | 101.3 | 101.37
3.4 | 1.68 | 3.24 | 24.48 | 95.42 | 99.7 | 101.1 | 101.4 | 101.4
3.5 | 1.5 | 2.92 | 18.16 | 94.22 | 99.62 | 101.2 | 101.4 | 101.47
3.6 | 1.48 | 2.72 | 13.96 | 92.4 | 99.42 | 101.2 | 101.5 | 101.57
3.7 | 1.38 | 2.52 | 11.24 | 89.54 | 99.14 | 101.3 | 101.5 | 101.6
3.8 | 1.3 | 2.32 | 9.36 | 85.4 | 99 | 101.3 | 101.52 | 101.67
3.9 | 1.28 | 2.2 | 8.02 | 79.02 | 98.74 | 101.3 | 101.6 | 101.7
4 | 1.2 | 2.1 | 6.92 | 70.6 | 98.44 | 101.4 | 101.6 | 101.77
4.1 | 1.18 | 1.92 | 6.08 | 60.64 | 98.06 | 101.4 | 101.7 | 101.8
4.2 | 1.1 | 1.9 | 5.46 | 50.16 | 97.56 | 101.5 | 101.7 | 101.87
4.3 | 1.1 | 1.8 | 4.96 | 39.8 | 96.96 | 101.5 | 101.8 | 101.9
4.4 | 1 | 1.7 | 4.54 | 30.4 | 96.18 | 101.6 | 101.9 | 101.97
4.5 | 1 | 1.62 | 4.24 | 22.88 | 95.12 | 101.6 | 101.92 | 102
4.6 | 1 | 1.52 | 3.92 | 17.38 | 93.56 | 101.6 | 102 | 102.07
4.7 | 0.9 | 1.5 | 3.72 | 13.9 | 91.44 | 101.6 | 102 | 102.17
4.8 | 0.9 | 1.42 | 3.5 | 11.5 | 88.36 | 101.7 | 102.1 | 102.2
4.9 | 0.9 | 1.4 | 3.22 | 9.86 | 83.56 | 101.7 | 102.1 | 102.27
5 | 0.9 | 1.32 | 3.1 | 8.62 | 76.48 | 101.8 | 102.1 | 102.3
5.1 | 0.88 | 1.3 | 2.92 | 7.6 | 67.46 | 101.8 | 102.2 | 102.3
5.2 | 0.8 | 1.3 | 2.8 | 6.88 | 57.32 | 101.8 | 102.2 | 102.3
5.3 | 0.8 | 1.2 | 2.7 | 6.26 | 46.64 | 101.8 | 102.22 | 102.37
5.4 | 0.8 | 1.2 | 2.52 | 5.76 | 36.68 | 101.8 | 102.3 | 102.47
5.5 | 0.7 | 1.12 | 2.42 | 5.34 | 28 | 101.8 | 102.3 | 102.5
5.6 | 0.7 | 1.1 | 2.4 | 5.04 | 21.2 | 101.9 | 102.32 | 102.5
5.7 | 0.7 | 1.1 | 2.22 | 4.72 | 16.6 | 101.9 | 102.4 | 102.57
5.8 | 0.7 | 1 | 2.12 | 4.42 | 13.64 | 101.9 | 102.4 | 102.6
5.9 | 0.7 | 1 | 2.1 | 4.22 | 11.58 | 101.9 | 102.5 | 102.6
6 | 0.7 | 1 | 2 | 4.02 | 10.12 | 101.8 | 102.5 | 102.67
6.1 | 0.7 | 0.92 | 1.92 | 3.82 | 9 | 101.8 | 102.5 | 102.7
6.2 | 0.6 | 0.9 | 1.9 | 3.7 | 8.1 | 101.8 | 102.5 | 102.77
6.3 | 0.6 | 0.9 | 1.8 | 3.52 | 7.46 | 101.8 | 102.5 | 102.8
6.4 | 0.6 | 0.9 | 1.8 | 3.32 | 6.94 | 101.8 | 102.6 | 102.8
6.5 | 0.6 | 0.9 | 1.7 | 3.22 | 6.44 | 101.8 | 102.6 | 102.8
6.6 |  | 0.8 | 1.6 | 3.1 | 6.04 | 101.8 | 102.6 | 102.87
6.7 |  | 0.8 | 1.6 | 3 | 5.64 | 101.8 | 102.6 | 102.9
6.8 |  | 0.8 | 1.5 | 2.9 | 5.42 | 101.8 | 102.6 | 102.9
6.9 |  | 0.8 | 1.5 | 2.72 | 5.12 | 101.7 | 102.62 | 102.9
7 |  | 0.8 | 1.4 | 2.62 | 4.92 | 101.7 | 102.7 | 102.97
7.1 |  |  | 1.4 | 2.6 | 4.64 | 101.7 | 102.7 | 103
7.2 |  |  | 1.4 | 2.5 | 4.52 | 101.6 | 102.7 | 103
7.3 |  |  | 1.3 | 2.42 | 4.3 | 101.6 | 102.7 | 103
7.4 |  |  | 1.3 | 2.3 | 4.12 | 101.5 | 102.7 | 103.07
7.5 |  |  | 1.22 | 2.3 | 4 | 101.5 | 102.8 | 103.1
7.6 |  |  | 1.2 | 2.2 | 3.82 | 101.5 | 102.7 | 103.1
7.7 |  |  | 1.2 | 2.1 | 3.7 | 101.4 | 102.72 | 103.1
7.8 |  |  | 1.2 | 2.1 | 3.6 | 101.3 | 102.8 | 103.1
7.9 |  |  | 1.12 | 2 | 3.42 | 101.3 | 102.78 | 103.17
8 |  |  | 1.1 | 1.92 | 3.3 | 101.2 | 102.7 | 103.2
8.1 |  |  | 1.1 | 1.9 | 3.2 | 101.1 | 102.8 | 103.2
8.2 |  |  |  | 1.8 | 3.1 | 101 | 102.8 | 103.2
8.3 |  |  |  | 1.8 | 3 | 100.9 | 102.8 | 103.2
8.4 |  |  |  | 1.7 | 2.9 | 100.8 | 102.8 | 103.27
8.5 |  |  |  | 1.7 | 2.82 | 100.7 | 102.8 | 103.3
8.6 |  |  |  | 1.7 | 2.72 | 100.6 | 102.8 | 103.3
8.7 |  |  |  | 1.6 | 2.7 | 100.4 | 102.78 | 103.3
8.8 |  |  |  | 1.6 | 2.6 | 100.2 | 102.72 | 103.3
8.9 |  |  |  | 1.52 | 2.5 | 99.9 | 102.7 | 103.3
9 |  |  |  | 1.5 | 2.42 | 99.7 | 102.7 | 103.3
9.1 |  |  |  | 1.5 | 2.4 | 99.5 | 102.7 | 103.3
9.2 |  |  |  |  | 2.3 | 99.1 | 102.7 | 103.3
9.3 |  |  |  |  | 2.22 | 98.8 | 102.7 | 103.3
9.4 |  |  |  |  | 2.2 | 98.3 | 102.7 | 103.37
9.5 |  |  |  |  | 2.1 | 97.8 | 102.7 | 103.4
9.6 |  |  |  |  | 2.02 | 97.2 | 102.7 | 103.4
9.7 |  |  |  |  | 2 | 96.3 | 102.7 | 103.47
9.8 |  |  |  |  | 2 | 95.2 | 102.7 | 103.5
9.9 |  |  |  |  | 1.9 | 93.6 | 102.7 | 103.5
10 |  |  |  |  | 1.9 | 91.3 | 102.7 | 103.5
10.1 |  |  |  |  | 1.82 | 88 | 102.7 | 103.5
10.2 |  |  |  |  | 1.8 | 82.7 | 102.7 | 103.57
10.3 | [5 empty] | 75 | 102.68 | 103.6
10.4 | [5 empty] | 65.8 | 102.6 | 103.6
10.5 | [5 empty] | 55.8 | 102.6 | 103.6
10.6 | [5 empty] | 45.3 | 102.6 | 103.6
10.7 | [5 empty] | 36.1 | 102.6 | 103.6
10.8 | [5 empty] | 28 | 102.5 | 103.6
10.9 | [5 empty] | 22 | 102.5 | 103.6
11 | [5 empty] | 18.1 | 102.5 | 103.6
11.1 | [5 empty] | 15.6 | 102.5 | 103.6
11.2 | [5 empty] | 13.8 | 102.5 | 103.6
11.3 | [5 empty] | 12.5 | 102.48 | 103.6
11.4 | [5 empty] | 11.5 | 102.4 | 103.6
11.5 | [5 empty] | 10.7 | 102.4 | 103.6
11.6 | [5 empty] | 10.1 | 102.4 | 103.6
11.7 | [5 empty] | 9.5 | 102.3 | 103.67
11.8 | [5 empty] | 9.1 | 102.3 | 103.7
11.9 | [5 empty] | 8.7 | 102.28 | 103.7
12 | [5 empty] | 8.4 | 102.2 | 103.7
12.1 | [5 empty] | 8.1 | 102.2 | 103.7
12.2 | [5 empty] | 7.8 | 102.1 | 103.7
12.3 | [5 empty] | 7.5 | 102.1 | 103.7
12.4 | [5 empty] | 7.3 | 102 | 103.7
12.5 | [5 empty] | 7 | 101.98 | 103.7
12.6 | [5 empty] | 6.8 | 101.9 | 103.7
12.7 | [5 empty] | 6.6 | 101.88 | 103.7
12.8 | [5 empty] | 6.4 | 101.8 | 103.7
12.9 | [5 empty] | 6.3 | 101.7 | 103.7
13 | [5 empty] | 6.1 | 101.68 | 103.7
13.1 | [5 empty] | 5.9 | 101.58 | 103.63
13.2 | [5 empty] | 5.7 | 101.5 | 103.6
13.3 | [5 empty] | 5.6 | 101.4 | 103.6
13.4 | [5 empty] | 5.5 | 101.38 | 103.6
13.5 | [5 empty] | 5.3 | 101.2 | 103.6
13.6 | [5 empty] | 5.2 | 101.18 | 103.6
13.7 | [5 empty] | 5 | 101 | 103.67
13.8 | [5 empty] | 4.9 | 100.9 | 103.7
13.9 | [5 empty] | 4.8 | 100.7 | 103.7
14 | [5 empty] | 4.7 | 100.68 | 103.7
14.1 | [5 empty] | 4.5 | 100.48 | 103.7
14.2 | [5 empty] | 4.4 | 100.28 | 103.7
14.3 | [5 empty] | 4.3 | 100.06 | 103.7
14.4 | [5 empty] | 4.2 | 99.78 | 103.7
14.5 | [5 empty] | 4.1 | 99.48 | 103.7
14.6 | [5 empty] | 4 | 99.16 | 103.7
14.7 | [5 empty] | 3.9 | 98.66 | 103.7
14.8 | [5 empty] | 3.8 | 98.12 | 103.7
14.9 | [5 empty] | 3.7 | 97.4 | 103.7
15 | [5 empty] | 3.7 | 96.5 | 103.7
15.1 | [5 empty] | 3.6 | 95.34 | 103.7
15.2 | [5 empty] | 3.5 | 93.58 | 103.7
15.3 | [5 empty] | 3.4 | 90.96 | 103.7
15.4 | [5 empty] | 3.3 | 86.88 | 103.7
15.5 | [5 empty] | 3.3 | 80.74 | 103.7
15.6 | [6 empty] | 72.66 | 103.7
15.7 | [6 empty] | 63.08 | 103.7
15.8 | [6 empty] | 52.98 | 103.7
15.9 | [6 empty] | 43.04 | 103.7
16 | [6 empty] | 34 | 103.77
16.1 | [6 empty] | 26.8 | 103.8
16.2 | [6 empty] | 21.62 | 103.8
16.3 | [6 empty] | 18.36 | 103.87
16.4 | [6 empty] | 16.32 | 103.83
16.5 | [6 empty] | 14.68 | 103.8
16.6 | [6 empty] | 13.62 | 103.8
16.7 | [6 empty] | 12.82 | 103.8
16.8 | [6 empty] | 12.14 | 103.8
16.9 | [6 empty] | 11.54 | 103.8
17 | [6 empty] | 11.06 | 103.8
17.1 | [6 empty] | 10.66 | 103.8
17.2 | [6 empty] | 10.28 | 103.8
17.3 | [6 empty] | 9.98 | 103.8
17.4 | [6 empty] | 9.68 | 103.8
17.5 | [6 empty] | 9.38 | 103.8
17.6 | [6 empty] | 9.18 | 103.8
17.7 | [6 empty] | 8.88 | 103.8
17.8 | [6 empty] | 8.68 | 103.73
17.9 | [6 empty] | 8.48 | 103.7
18 | [6 empty] | 8.28 | 103.7
18.1 | [6 empty] | 8.08 | 103.63
18.2 | [6 empty] | 7.88 | 103.6
18.3 | [6 empty] | 7.68 | 103.53
18.4 | [6 empty] | 7.48 | 103.43
18.5 | [6 empty] | 7.3 | 103.33
18.6 | [6 empty] | 7.18 | 103.3
18.7 | [6 empty] | 6.98 | 103.23
18.8 | [6 empty] | 6.88 | 103.13
18.9 | [6 empty] | 6.68 | 102.96
19 | [6 empty] | 6.5 | 102.83
19.1 | [6 empty] | 6.38 | 102.73
19.2 | [6 empty] | 6.2 | 102.56
19.3 | [6 empty] | 6.08 | 102.36
19.4 | [6 empty] | 5.98 | 102.16
19.5 | [6 empty] | 5.8 | 101.89
19.6 | [6 empty] | 5.7 | 101.66
19.7 | [6 empty] | 5.6 | 101.32
19.8 | [6 empty] | 5.48 | 100.99
19.9 | [6 empty] | 5.38 | 100.55
20 | [6 empty] | 5.2 | 99.98
20.1 | [6 empty] | 5.18 | 99.31
20.2 | [6 empty] | 5 | 98.54
20.3 | [6 empty] | 4.9 | 97.53
20.4 | [6 empty] | 4.8 | 96.15
20.5 | [6 empty] | 4.7 | 94.16
20.6 | [6 empty] | 4.6 | 91.19
20.7 | [6 empty] | 4.5 | 86.42
20.8 | [7 empty] | 79.27
20.9 | [7 empty] | 70.46
21 | [7 empty] | 60.7
21.1 | [7 empty] | 50.49
21.2 | [7 empty] | 40.75
21.3 | [7 empty] | 32.3
21.4 | [7 empty] | 25.7
21.5 | [7 empty] | 21.24
21.6 | [7 empty] | 18.42
21.7 | [7 empty] | 16.51
21.8 | [7 empty] | 15.16
21.9 | [7 empty] | 14.17
22 | [7 empty] | 13.41
22.1 | [7 empty] | 12.78
22.2 | [7 empty] | 12.25
22.3 | [7 empty] | 11.75
22.4 | [7 empty] | 11.39
22.5 | [7 empty] | 11.02
22.6 | [7 empty] | 10.69
22.7 | [7 empty] | 10.39
22.8 | [7 empty] | 10.09
22.9 | [7 empty] | 9.79
23 | [7 empty] | 9.56
23.1 | [7 empty] | 9.29
23.2 | [7 empty] | 9.06
23.3 | [7 empty] | 8.86
23.4 | [7 empty] | 8.59
23.5 | [7 empty] | 8.36
23.6 | [7 empty] | 8.16
23.7 | [7 empty] | 7.96
23.8 | [7 empty] | 7.76
23.9 | [7 empty] | 7.56
24 | [7 empty] | 7.36
24.1 | [7 empty] | 7.16
24.2 | [7 empty] | 7.03
24.3 | [7 empty] | 6.86
24.4 | [7 empty] | 6.66
24.5 | [7 empty] | 6.46
24.6 | [7 empty] | 6.33
24.7 | [7 empty] | 6.16
24.8 | [7 empty] | 6.03
24.9 | [7 empty] | 5.93
25 | [7 empty] | 5.76
25.1 | [7 empty] | 5.63
25.2 | [7 empty] | 5.53
25.3 | [7 empty] | 5.43
25.4 | [7 empty] | 5.26
25.5 | [7 empty] | 5.13
25.6 | [7 empty] | 5.03
25.7 | [7 empty] | 4.93
25.8 | [7 empty] | 4.83
25.9 | [7 empty] | 4.73

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 4.68
-28.8 | [7 empty] | 4.7
-28.7 | [7 empty] | 4.78
-28.6 | [7 empty] | 4.88
-28.5 | [7 empty] | 4.98
-28.4 | [7 empty] | 5.08
-28.3 | [7 empty] | 5.1
-28.2 | [7 empty] | 5.18
-28.1 | [7 empty] | 5.28
-28 | [7 empty] | 5.38
-27.9 | [7 empty] | 5.48
-27.8 | [7 empty] | 5.58
-27.7 | [7 empty] | 5.68
-27.6 | [7 empty] | 5.78
-27.5 | [7 empty] | 5.88
-27.4 | [7 empty] | 5.98
-27.3 | [7 empty] | 6.08
-27.2 | [7 empty] | 6.18
-27.1 | [7 empty] | 6.28
-27 | [7 empty] | 6.38
-26.9 | [7 empty] | 6.56
-26.8 | [7 empty] | 6.68
-26.7 | [7 empty] | 6.78
-26.6 | [7 empty] | 6.96
-26.5 | [7 empty] | 7.08
-26.4 | [7 empty] | 7.26
-26.3 | [7 empty] | 7.38
-26.2 | [7 empty] | 7.56
-26.1 | [7 empty] | 7.68
-26 | [7 empty] | 7.86
-25.9 | [7 empty] | 8.06
-25.8 | [7 empty] | 8.18
-25.7 | [7 empty] | 8.36
-25.6 | [7 empty] | 8.56
-25.5 | [7 empty] | 8.76
-25.4 | [7 empty] | 8.96
-25.3 | [7 empty] | 9.16
-25.2 | [7 empty] | 9.36
-25.1 | [7 empty] | 9.56
-25 | [7 empty] | 9.76
-24.9 | [7 empty] | 9.96
-24.8 | [7 empty] | 10.24
-24.7 | [7 empty] | 10.46
-24.6 | [7 empty] | 10.66
-24.5 | [7 empty] | 10.94
-24.4 | [7 empty] | 11.16
-24.3 | [7 empty] | 11.36
-24.2 | [7 empty] | 11.64
-24.1 | [7 empty] | 11.94
-24 | [7 empty] | 12.16
-23.9 | [7 empty] | 12.44
-23.8 | [7 empty] | 12.74
-23.7 | [7 empty] | 13.04
-23.6 | [7 empty] | 13.42
-23.5 | [7 empty] | 13.74
-23.4 | [7 empty] | 14.12
-23.3 | [7 empty] | 14.52
-23.2 | [7 empty] | 15
-23.1 | [7 empty] | 15.58
-23 | [7 empty] | 16.26
-22.9 | [7 empty] | 17.04
-22.8 | [7 empty] | 18
-22.7 | [7 empty] | 19.24
-22.6 | [7 empty] | 20.94
-22.5 | [7 empty] | 23.46
-22.4 | [7 empty] | 27.2
-22.3 | [7 empty] | 32.72
-22.2 | [7 empty] | 39.82
-22.1 | [7 empty] | 48.1
-22 | [7 empty] | 56.92
-21.9 | [7 empty] | 65.5
-21.8 | [7 empty] | 73.36
-21.7 | [7 empty] | 79.78
-21.6 | [7 empty] | 84.44
-21.5 | [6 empty] | 6 | 87.54
-21.4 | [6 empty] | 6.12 | 89.54
-21.3 | [6 empty] | 6.22 | 90.94
-21.2 | [6 empty] | 6.4 | 92.08
-21.1 | [6 empty] | 6.5 | 93.02
-21 | [6 empty] | 6.62 | 93.68
-20.9 | [6 empty] | 6.72 | 94.36
-20.8 | [6 empty] | 6.9 | 94.9
-20.7 | [6 empty] | 7.02 | 95.32
-20.6 | [6 empty] | 7.12 | 95.64
-20.5 | [6 empty] | 7.22 | 96.02
-20.4 | [6 empty] | 7.42 | 96.34
-20.3 | [6 empty] | 7.62 | 96.64
-20.2 | [6 empty] | 7.72 | 96.94
-20.1 | [6 empty] | 7.92 | 97.16
-20 | [6 empty] | 8.12 | 97.44
-19.9 | [6 empty] | 8.3 | 97.66
-19.8 | [6 empty] | 8.42 | 97.86
-19.7 | [6 empty] | 8.62 | 98.06
-19.6 | [6 empty] | 8.74 | 98.18
-19.5 | [6 empty] | 9.02 | 98.36
-19.4 | [6 empty] | 9.22 | 98.56
-19.3 | [6 empty] | 9.42 | 98.68
-19.2 | [6 empty] | 9.62 | 98.78
-19.1 | [6 empty] | 9.82 | 98.96
-19 | [6 empty] | 10.02 | 99.08
-18.9 | [6 empty] | 10.22 | 99.18
-18.8 | [6 empty] | 10.52 | 99.36
-18.7 | [6 empty] | 10.72 | 99.4
-18.6 | [6 empty] | 11.02 | 99.48
-18.5 | [6 empty] | 11.22 | 99.58
-18.4 | [6 empty] | 11.52 | 99.68
-18.3 | [6 empty] | 11.82 | 99.78
-18.2 | [6 empty] | 12.14 | 99.8
-18.1 | [6 empty] | 12.44 | 99.88
-18 | [6 empty] | 12.84 | 99.98
-17.9 | [6 empty] | 13.14 | 100
-17.8 | [6 empty] | 13.64 | 100.08
-17.7 | [6 empty] | 14.08 | 100.18
-17.6 | [6 empty] | 14.66 | 100.28
-17.5 | [6 empty] | 15.28 | 100.3
-17.4 | [6 empty] | 16.1 | 100.3
-17.3 | [6 empty] | 17.1 | 100.38
-17.2 | [6 empty] | 18.34 | 100.4
-17.1 | [6 empty] | 20 | 100.4
-17 | [6 empty] | 22.44 | 100.48
-16.9 | [6 empty] | 26.28 | 100.5
-16.8 | [6 empty] | 31.78 | 100.58
-16.7 | [6 empty] | 39.02 | 100.6
-16.6 | [6 empty] | 47.58 | 100.68
-16.5 | [6 empty] | 56.5 | 100.7
-16.4 | [6 empty] | 65.44 | 100.7
-16.3 | [6 empty] | 73.42 | 100.7
-16.2 | [6 empty] | 79.82 | 100.78
-16.1 | [6 empty] | 84.54 | 100.8
-16 | [5 empty] | 4.58 | 87.62 | 100.88
-15.9 | [5 empty] | 4.6 | 89.66 | 100.9
-15.8 | [5 empty] | 4.7 | 91.04 | 100.98
-15.7 | [5 empty] | 4.8 | 92.2 | 101
-15.6 | [5 empty] | 4.9 | 93.06 | 101
-15.5 | [5 empty] | 5 | 93.68 | 101.08
-15.4 | [5 empty] | 5.1 | 94.36 | 101.1
-15.3 | [5 empty] | 5.2 | 94.84 | 101.1
-15.2 | [5 empty] | 5.3 | 95.24 | 101.18
-15.1 | [5 empty] | 5.4 | 95.62 | 101.2
-15 | [5 empty] | 5.58 | 95.94 | 101.2
-14.9 | [5 empty] | 5.68 | 96.22 | 101.28
-14.8 | [5 empty] | 5.8 | 96.44 | 101.3
-14.7 | [5 empty] | 5.98 | 96.74 | 101.3
-14.6 | [5 empty] | 6.08 | 96.94 | 101.38
-14.5 | [5 empty] | 6.28 | 97.22 | 101.4
-14.4 | [5 empty] | 6.4 | 97.32 | 101.4
-14.3 | [5 empty] | 6.58 | 97.52 | 101.48
-14.2 | [5 empty] | 6.78 | 97.7 | 101.5
-14.1 | [5 empty] | 6.9 | 97.8 | 101.5
-14 | [5 empty] | 7.08 | 97.92 | 101.5
-13.9 | [5 empty] | 7.28 | 98.12 | 101.58
-13.8 | [5 empty] | 7.48 | 98.3 | 101.6
-13.7 | [5 empty] | 7.68 | 98.42 | 101.6
-13.6 | [5 empty] | 7.88 | 98.52 | 101.68
-13.5 | [5 empty] | 8.08 | 98.62 | 101.7
-13.4 | [5 empty] | 8.28 | 98.82 | 101.7
-13.3 | [5 empty] | 8.48 | 98.92 | 101.78
-13.2 | [5 empty] | 8.76 | 99.12 | 101.8
-13.1 | [5 empty] | 8.98 | 99.22 | 101.8
-13 | [5 empty] | 9.26 | 99.32 | 101.8
-12.9 | [5 empty] | 9.48 | 99.5 | 101.8
-12.8 | [5 empty] | 9.78 | 99.5 | 101.88
-12.7 | [5 empty] | 10.08 | 99.6 | 101.9
-12.6 | [5 empty] | 10.36 | 99.62 | 101.9
-12.5 | [5 empty] | 10.76 | 99.8 | 101.98
-12.4 | [5 empty] | 11.16 | 99.9 | 102
-12.3 | [5 empty] | 11.56 | 99.92 | 102
-12.2 | [5 empty] | 12.14 | 100.02 | 102
-12.1 | [5 empty] | 12.64 | 100.2 | 102
-12 | [5 empty] | 13.32 | 100.2 | 102
-11.9 | [5 empty] | 14.12 | 100.3 | 102
-11.8 | [5 empty] | 15.18 | 100.3 | 102
-11.7 | [5 empty] | 16.38 | 100.4 | 102
-11.6 | [5 empty] | 18.12 | 100.42 | 102.08
-11.5 | [5 empty] | 20.52 | 100.5 | 102.1
-11.4 | [5 empty] | 24.2 | 100.52 | 102.1
-11.3 | [5 empty] | 29.4 | 100.6 | 102.1
-11.2 | [5 empty] | 36.72 | 100.7 | 102.1
-11.1 | [5 empty] | 45.22 | 100.72 | 102.1
-11 | [5 empty] | 54.54 | 100.8 | 102.1
-10.9 | [5 empty] | 63.88 | 100.82 | 102.1
-10.8 | [5 empty] | 72.36 | 100.9 | 102.1
-10.7 | [5 empty] | 79.28 | 100.92 | 102.1
-10.6 | [5 empty] | 84.38 | 101 | 102.1
-10.5 |  |  |  |  | 2.52 | 87.8 | 101.1 | 102.1
-10.4 |  |  |  |  | 2.6 | 90 | 101.1 | 102.1
-10.3 |  |  |  |  | 2.62 | 91.66 | 101.1 | 102.18
-10.2 |  |  |  |  | 2.7 | 92.9 | 101.2 | 102.2
-10.1 |  |  |  |  | 2.8 | 93.82 | 101.2 | 102.2
-10 |  |  |  |  | 2.82 | 94.54 | 101.2 | 102.2
-9.9 |  |  |  |  | 2.9 | 95.14 | 101.3 | 102.2
-9.8 |  |  |  |  | 3 | 95.66 | 101.3 | 102.2
-9.7 |  |  |  |  | 3.02 | 96.16 | 101.4 | 102.2
-9.6 |  |  |  |  | 3.12 | 96.56 | 101.42 | 102.2
-9.5 |  |  |  |  | 3.2 | 96.88 | 101.5 | 102.2
-9.4 |  |  |  | 2.1 | 3.3 | 97.26 | 101.5 | 102.2
-9.3 |  |  |  | 2.1 | 3.4 | 97.48 | 101.52 | 102.2
-9.2 |  |  |  | 2.2 | 3.5 | 97.84 | 101.6 | 102.2
-9.1 |  |  |  | 2.2 | 3.6 | 98.08 | 101.6 | 102.2
-9 |  |  |  | 2.22 | 3.72 | 98.28 | 101.6 | 102.2
-8.9 |  |  |  | 2.3 | 3.82 | 98.4 | 101.62 | 102.2
-8.8 |  |  |  | 2.4 | 3.92 | 98.58 | 101.7 | 102.2
-8.7 |  |  |  | 2.42 | 4.02 | 98.78 | 101.7 | 102.2
-8.6 |  |  |  | 2.5 | 4.2 | 98.98 | 101.7 | 102.2
-8.5 |  |  |  | 2.52 | 4.3 | 99.1 | 101.7 | 102.2
-8.4 |  |  |  | 2.62 | 4.42 | 99.28 | 101.7 | 102.2
-8.3 |  |  | 1.5 | 2.7 | 4.6 | 99.4 | 101.72 | 102.2
-8.2 |  |  | 1.6 | 2.8 | 4.72 | 99.5 | 101.8 | 102.2
-8.1 |  |  | 1.6 | 2.82 | 4.82 | 99.68 | 101.8 | 102.2
-8 |  |  | 1.7 | 3 | 5.02 | 99.78 | 101.82 | 102.2
-7.9 |  |  | 1.7 | 3.02 | 5.22 | 99.88 | 101.9 | 102.2
-7.8 |  |  | 1.7 | 3.12 | 5.42 | 99.98 | 101.9 | 102.2
-7.7 |  |  | 1.8 | 3.22 | 5.62 | 100.08 | 101.9 | 102.2
-7.6 |  |  | 1.8 | 3.32 | 5.8 | 100.18 | 101.9 | 102.2
-7.5 |  |  | 1.9 | 3.42 | 6.02 | 100.2 | 101.9 | 102.2
-7.4 |  |  | 1.9 | 3.52 | 6.24 | 100.3 | 101.9 | 102.2
-7.3 |  |  | 2 | 3.62 | 6.44 | 100.4 | 101.9 | 102.12
-7.2 |  | 1.1 | 2 | 3.8 | 6.74 | 100.5 | 101.9 | 102.1
-7.1 |  | 1.1 | 2.1 | 3.9 | 7.04 | 100.58 | 101.9 | 102.18
-7 |  | 1.1 | 2.1 | 4.02 | 7.42 | 100.6 | 101.9 | 102.2
-6.9 |  | 1.1 | 2.2 | 4.2 | 7.74 | 100.68 | 102 | 102.2
-6.8 |  | 1.1 | 2.3 | 4.32 | 8.14 | 100.7 | 102 | 102.12
-6.7 |  | 1.2 | 2.4 | 4.52 | 8.64 | 100.78 | 102 | 102.1
-6.6 | 0.8 | 1.2 | 2.5 | 4.72 | 9.18 | 100.88 | 102 | 102.1
-6.5 | 0.9 | 1.2 | 2.5 | 4.84 | 9.88 | 100.98 | 102 | 102.1
-6.4 | 0.9 | 1.2 | 2.6 | 5.1 | 10.7 | 101 | 102 | 102.1
-6.3 | 0.9 | 1.3 | 2.7 | 5.32 | 11.72 | 101.1 | 102 | 102.1
-6.2 | 0.9 | 1.3 | 2.8 | 5.54 | 13.06 | 101.1 | 102 | 102.1
-6.1 | 0.9 | 1.4 | 2.9 | 5.82 | 14.82 | 101.1 | 102 | 102.02
-6 | 0.9 | 1.4 | 3 | 6.12 | 17.2 | 101.18 | 101.98 | 102
-5.9 | 1 | 1.4 | 3.1 | 6.44 | 20.78 | 101.2 | 101.9 | 102
-5.8 | 1 | 1.5 | 3.2 | 6.74 | 26.18 | 101.28 | 101.9 | 102
-5.7 | 1 | 1.6 | 3.4 | 7.16 | 33.56 | 101.3 | 101.9 | 102
-5.6 | 1 | 1.6 | 3.5 | 7.66 | 42.44 | 101.3 | 101.82 | 102
-5.5 | 1 | 1.6 | 3.6 | 8.24 | 52 | 101.3 | 101.8 | 101.92
-5.4 | 1.1 | 1.7 | 3.8 | 8.88 | 61.92 | 101.3 | 101.8 | 101.9
-5.3 | 1.1 | 1.8 | 4 | 9.7 | 71.04 | 101.3 | 101.8 | 101.9
-5.2 | 1.1 | 1.8 | 4.2 | 10.72 | 78.84 | 101.3 | 101.8 | 101.82
-5.1 | 1.12 | 1.9 | 4.4 | 12.06 | 84.72 | 101.3 | 101.8 | 101.8
-5 | 1.2 | 2 | 4.6 | 13.8 | 88.5 | 101.3 | 101.8 | 101.8
-4.9 | 1.3 | 2 | 4.9 | 16.2 | 91.08 | 101.3 | 101.7 | 101.72
-4.8 | 1.3 | 2.1 | 5.1 | 19.76 | 92.76 | 101.3 | 101.7 | 101.7
-4.7 | 1.32 | 2.2 | 5.5 | 25.14 | 94.12 | 101.3 | 101.7 | 101.7
-4.6 | 1.4 | 2.3 | 5.9 | 32.24 | 95.18 | 101.3 | 101.7 | 101.62
-4.5 | 1.42 | 2.4 | 6.3 | 41.06 | 95.96 | 101.3 | 101.7 | 101.6
-4.4 | 1.5 | 2.5 | 6.8 | 50.9 | 96.46 | 101.3 | 101.6 | 101.52
-4.3 | 1.6 | 2.6 | 7.4 | 60.84 | 97.06 | 101.3 | 101.6 | 101.5
-4.2 | 1.6 | 2.7 | 8.2 | 70.06 | 97.54 | 101.3 | 101.5 | 101.42
-4.1 | 1.7 | 2.9 | 9.1 | 78.06 | 97.84 | 101.2 | 101.5 | 101.4
-4 | 1.72 | 3.1 | 10.4 | 83.98 | 98.02 | 101.2 | 101.4 | 101.4
-3.9 | 1.82 | 3.2 | 12.1 | 88.3 | 98.22 | 101.2 | 101.4 | 101.4
-3.8 | 1.9 | 3.4 | 14.4 | 91 | 98.62 | 101.2 | 101.38 | 101.32
-3.7 | 2.02 | 3.7 | 17.7 | 92.86 | 98.82 | 101.2 | 101.3 | 101.3
-3.6 | 2.12 | 3.9 | 22.8 | 94.24 | 99.02 | 101.12 | 101.3 | 101.22
-3.5 | 2.3 | 4.3 | 29.9 | 95.3 | 99.12 | 101.1 | 101.3 | 101.2
-3.4 | 2.42 | 4.7 | 38.8 | 96.16 | 99.32 | 101.1 | 101.28 | 101.12
-3.3 | 2.52 | 5.1 | 48.7 | 96.76 | 99.5 | 101 | 101.2 | 101.1
-3.2 | 2.72 | 5.7 | 59.1 | 97.34 | 99.52 | 101.02 | 101.18 | 101.1
-3.1 | 2.94 | 6.5 | 68.8 | 97.74 | 99.62 | 101 | 101.1 | 101.02
-3 | 3.22 | 7.4 | 77.4 | 98.12 | 99.8 | 101 | 101.1 | 101
-2.9 | 3.54 | 8.7 | 84 | 98.34 | 99.82 | 101 | 101 | 101
-2.8 | 3.94 | 10.4 | 88.4 | 98.62 | 99.92 | 100.9 | 101 | 100.92
-2.7 | 4.46 | 12.6 | 91.4 | 98.82 | 100.02 | 100.9 | 101 | 100.9
-2.6 | 5.06 | 16 | 93.3 | 99.12 | 100.1 | 100.9 | 101 | 100.82
-2.5 | 5.88 | 21.3 | 94.8 | 99.22 | 100.2 | 100.9 | 100.98 | 100.8
-2.4 | 6.92 | 28.5 | 95.8 | 99.32 | 100.12 | 100.8 | 100.9 | 100.8
-2.3 | 8.28 | 37.8 | 96.8 | 99.5 | 100.2 | 100.8 | 100.88 | 100.72
-2.2 | 10.36 | 48.3 | 97.4 | 99.52 | 100.2 | 100.7 | 100.8 | 100.7
-2.1 | 13.28 | 59.1 | 97.9 | 99.6 | 100.2 | 100.7 | 100.8 | 100.62
-2 | 17.64 | 69.2 | 98.2 | 99.7 | 100.22 | 100.7 | 100.78 | 100.6
-1.9 | 23.9 | 78 | 98.5 | 99.8 | 100.3 | 100.7 | 100.7 | 100.52
-1.8 | 32.56 | 84.8 | 98.8 | 99.9 | 100.3 | 100.62 | 100.68 | 100.5
-1.7 | 42.76 | 89.5 | 99.1 | 100 | 100.3 | 100.6 | 100.6 | 100.42
-1.6 | 53.48 | 92.5 | 99.3 | 100 | 100.3 | 100.52 | 100.6 | 100.4
-1.5 | 64.34 | 94.4 | 99.4 | 100 | 100.3 | 100.42 | 100.5 | 100.32
-1.4 | 74.18 | 95.8 | 99.5 | 100.02 | 100.3 | 100.4 | 100.5 | 100.3
-1.3 | 82.28 | 96.9 | 99.7 | 100.1 | 100.2 | 100.4 | 100.48 | 100.3
-1.2 | 88.04 | 97.7 | 99.7 | 100.1 | 100.2 | 100.32 | 100.4 | 100.3
-1.1 | 91.76 | 98.3 | 99.8 | 100.1 | 100.2 | 100.3 | 100.3 | 100.22
-1 | 94.22 | 98.7 | 99.8 | 100.1 | 100.2 | 100.3 | 100.3 | 100.2
-0.9 | 96.02 | 99.2 | 99.8 | 100.1 | 100.2 | 100.2 | 100.3 | 100.2
-0.8 | 97.14 | 99.5 | 99.8 | 100.1 | 100.2 | 100.2 | 100.2 | 100.12
-0.7 | 98.06 | 99.6 | 99.9 | 100.1 | 100.2 | 100.2 | 100.2 | 100.1
-0.6 | 98.64 | 99.7 | 99.9 | 100.1 | 100.22 | 100.2 | 100.2 | 100.02
-0.5 | 99.04 | 99.8 | 99.9 | 100.1 | 100.28 | 100.1 | 100.2 | 100
-0.4 | 99.44 | 99.9 | 99.9 | 100.08 | 100.2 | 100.1 | 100.1 | 100
-0.3 | 99.72 | 99.9 | 100 | 100 | 100.18 | 100.1 | 100.1 | 100
-0.2 | 99.9 | 99.9 | 100 | 100 | 100.1 | 100 | 100.1 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.98 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.88 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100
0.3 | 99.76 | 99.9 | 99.9 | 100 | 100 | 100 | 100 | 100
0.4 | 99.38 | 99.9 | 99.9 | 100 | 100.1 | 100 | 100 | 100
0.5 | 99.06 | 99.8 | 99.8 | 100 | 100.1 | 100 | 100 | 100
0.6 | 98.64 | 99.6 | 99.8 | 99.98 | 100.1 | 100 | 100.02 | 100
0.7 | 98.1 | 99.5 | 99.8 | 99.9 | 100.1 | 100.1 | 100.1 | 100.07
0.8 | 97.16 | 99.3 | 99.8 | 99.9 | 100.1 | 100.1 | 100.1 | 100.1
0.9 | 95.88 | 99.1 | 99.7 | 99.9 | 100.1 | 100.1 | 100.1 | 100.17
1 | 94.08 | 98.7 | 99.7 | 99.9 | 100.1 | 100.1 | 100.1 | 100.2
1.1 | 91.6 | 98.2 | 99.6 | 99.9 | 100.1 | 100.2 | 100.12 | 100.27
1.2 | 87.9 | 97.7 | 99.6 | 99.9 | 100.1 | 100.2 | 100.2 | 100.3
1.3 | 82.26 | 96.9 | 99.5 | 99.88 | 100.1 | 100.2 | 100.2 | 100.3
1.4 | 74.18 | 95.8 | 99.3 | 99.8 | 100.1 | 100.2 | 100.22 | 100.3
1.5 | 64.24 | 94.4 | 99.1 | 99.8 | 100.1 | 100.2 | 100.3 | 100.37
1.6 | 53.52 | 92.4 | 99.1 | 99.8 | 100.1 | 100.2 | 100.3 | 100.4
1.7 | 42.84 | 89.5 | 98.9 | 99.7 | 100 | 100.28 | 100.3 | 100.47
1.8 | 32.76 | 85.1 | 98.6 | 99.7 | 100 | 100.3 | 100.4 | 100.5
1.9 | 23.96 | 78.5 | 98.3 | 99.58 | 100 | 100.38 | 100.4 | 100.5
2 | 17.62 | 69.8 | 98 | 99.48 | 99.98 | 100.4 | 100.4 | 100.57
2.1 | 13.26 | 59.9 | 97.7 | 99.4 | 99.9 | 100.48 | 100.5 | 100.67
2.2 | 10.36 | 49.1 | 97.1 | 99.28 | 99.9 | 100.5 | 100.5 | 100.7
2.3 | 8.26 | 38.6 | 96.6 | 99.2 | 99.88 | 100.5 | 100.52 | 100.7
2.4 | 6.88 | 29.1 | 95.7 | 99.08 | 99.8 | 100.5 | 100.6 | 100.77
2.5 | 5.82 | 21.7 | 94.6 | 98.88 | 99.78 | 100.5 | 100.6 | 100.8
2.6 | 5.04 | 16.3 | 93.2 | 98.68 | 99.68 | 100.5 | 100.7 | 100.8
2.7 | 4.44 | 12.7 | 91.3 | 98.5 | 99.6 | 100.6 | 100.7 | 100.87
2.8 | 3.96 | 10.4 | 88.4 | 98.38 | 99.5 | 100.6 | 100.72 | 100.9
2.9 | 3.56 | 8.8 | 84.1 | 98.16 | 99.5 | 100.6 | 100.72 | 100.97
3 | 3.18 | 7.4 | 77.7 | 97.76 | 99.38 | 100.68 | 100.8 | 101
3.1 | 2.98 | 6.5 | 69.4 | 97.46 | 99.28 | 100.7 | 100.82 | 101
3.2 | 2.78 | 5.7 | 59.6 | 97.06 | 99.1 | 100.7 | 100.9 | 101.07
3.3 | 2.58 | 5.1 | 49.3 | 96.44 | 99 | 100.78 | 100.9 | 101.1
3.4 | 2.4 | 4.7 | 39.2 | 95.82 | 98.98 | 100.8 | 101 | 101.1
3.5 | 2.28 | 4.3 | 30.4 | 94.92 | 98.78 | 100.8 | 101 | 101.17
3.6 | 2.18 | 4 | 23.1 | 93.98 | 98.68 | 100.8 | 101 | 101.2
3.7 | 2.08 | 3.7 | 17.9 | 92.62 | 98.38 | 100.8 | 101 | 101.27
3.8 | 1.98 | 3.5 | 14.5 | 90.64 | 98.18 | 100.8 | 101.1 | 101.3
3.9 | 1.88 | 3.3 | 12.1 | 87.84 | 97.98 | 100.8 | 101.1 | 101.37
4 | 1.78 | 3.1 | 10.5 | 83.68 | 97.68 | 100.8 | 101.1 | 101.4
4.1 | 1.68 | 2.9 | 9.2 | 77.56 | 97.36 | 100.9 | 101.1 | 101.4
4.2 | 1.6 | 2.8 | 8.3 | 69.6 | 97.08 | 100.9 | 101.1 | 101.4
4.3 | 1.58 | 2.6 | 7.4 | 60.22 | 96.56 | 100.9 | 101.2 | 101.47
4.4 | 1.5 | 2.5 | 6.8 | 50.38 | 96.12 | 100.9 | 101.2 | 101.5
4.5 | 1.48 | 2.4 | 6.3 | 40.58 | 95.34 | 101 | 101.22 | 101.57
4.6 | 1.4 | 2.3 | 5.9 | 31.84 | 94.6 | 100.9 | 101.3 | 101.6
4.7 | 1.38 | 2.2 | 5.5 | 24.78 | 93.7 | 100.9 | 101.3 | 101.67
4.8 | 1.3 | 2.1 | 5.2 | 19.5 | 92.44 | 100.9 | 101.32 | 101.7
4.9 | 1.3 | 2 | 4.9 | 16.02 | 90.6 | 100.9 | 101.4 | 101.77
5 | 1.2 | 2 | 4.7 | 13.62 | 88.16 | 100.9 | 101.4 | 101.8
5.1 | 1.2 | 1.9 | 4.4 | 11.94 | 84.32 | 100.9 | 101.5 | 101.8
5.2 | 1.1 | 1.8 | 4.2 | 10.58 | 78.7 | 100.9 | 101.5 | 101.87
5.3 | 1.1 | 1.8 | 4 | 9.62 | 71.04 | 100.9 | 101.5 | 101.9
5.4 | 1.1 | 1.7 | 3.9 | 8.82 | 62.12 | 100.9 | 101.5 | 101.9
5.5 | 1 | 1.6 | 3.7 | 8.14 | 52.24 | 100.9 | 101.5 | 101.97
5.6 | 1 | 1.6 | 3.6 | 7.56 | 42.7 | 100.9 | 101.5 | 102
5.7 | 1 | 1.5 | 3.4 | 7.14 | 33.82 | 100.9 | 101.52 | 102
5.8 | 1 | 1.5 | 3.3 | 6.68 | 26.56 | 100.8 | 101.6 | 102
5.9 | 1 | 1.5 | 3.1 | 6.36 | 21.08 | 100.8 | 101.6 | 102
6 | 0.9 | 1.4 | 3 | 6.08 | 17.42 | 100.8 | 101.6 | 102
6.1 | 0.9 | 1.4 | 2.9 | 5.78 | 14.9 | 100.72 | 101.6 | 102
6.2 | 0.9 | 1.3 | 2.8 | 5.56 | 13.16 | 100.7 | 101.6 | 102.07
6.3 | 0.9 | 1.3 | 2.7 | 5.28 | 11.8 | 100.6 | 101.6 | 102.1
6.4 | 0.88 | 1.3 | 2.6 | 5.08 | 10.8 | 100.6 | 101.6 | 102.1
6.5 | 0.8 | 1.2 | 2.5 | 4.88 | 10.02 | 100.5 | 101.6 | 102.1
6.6 | 0.8 | 1.2 | 2.5 | 4.68 | 9.24 | 100.42 | 101.6 | 102.1
6.7 |  | 1.2 | 2.4 | 4.58 | 8.64 | 100.4 | 101.58 | 102.1
6.8 |  | 1.2 | 2.3 | 4.38 | 8.24 | 100.4 | 101.58 | 102.1
6.9 |  | 1.1 | 2.2 | 4.18 | 7.76 | 100.3 | 101.5 | 102.17
7 |  | 1.1 | 2.2 | 4.08 | 7.38 | 100.22 | 101.5 | 102.2
7.1 |  | 1.1 | 2.1 | 3.88 | 7.06 | 100.2 | 101.5 | 102.2
7.2 |  | 1 | 2 | 3.78 | 6.78 | 100.1 | 101.5 | 102.13
7.3 |  |  | 2 | 3.68 | 6.46 | 100 | 101.5 | 102.17
7.4 |  |  | 1.9 | 3.5 | 6.18 | 99.9 | 101.5 | 102.2
7.5 |  |  | 1.9 | 3.4 | 5.98 | 99.9 | 101.5 | 102.2
7.6 |  |  | 1.8 | 3.3 | 5.78 | 99.8 | 101.5 | 102.2
7.7 |  |  | 1.8 | 3.2 | 5.48 | 99.72 | 101.5 | 102.2
7.8 |  |  | 1.7 | 3.1 | 5.38 | 99.62 | 101.5 | 102.2
7.9 |  |  | 1.7 | 3 | 5.2 | 99.5 | 101.4 | 102.2
8 |  |  | 1.6 | 2.9 | 5.08 | 99.4 | 101.4 | 102.2
8.1 |  |  | 1.6 | 2.8 | 4.88 | 99.3 | 101.4 | 102.2
8.2 |  |  | 1.6 | 2.78 | 4.68 | 99.2 | 101.4 | 102.2
8.3 |  |  | 1.5 | 2.7 | 4.5 | 99.02 | 101.4 | 102.2
8.4 |  |  |  | 2.6 | 4.4 | 98.92 | 101.4 | 102.2
8.5 |  |  |  | 2.5 | 4.28 | 98.64 | 101.4 | 102.2
8.6 |  |  |  | 2.5 | 4.18 | 98.52 | 101.3 | 102.2
8.7 |  |  |  | 2.4 | 3.98 | 98.42 | 101.3 | 102.2
8.8 |  |  |  | 2.38 | 3.9 | 98.22 | 101.3 | 102.2
8.9 |  |  |  | 2.3 | 3.78 | 98.02 | 101.2 | 102.2
9 |  |  |  | 2.2 | 3.68 | 97.82 | 101.2 | 102.2
9.1 |  |  |  | 2.2 | 3.58 | 97.62 | 101.18 | 102.2
9.2 |  |  |  | 2.1 | 3.48 | 97.42 | 101.1 | 102.2
9.3 |  |  |  | 2.1 | 3.4 | 97.14 | 101.1 | 102.2
9.4 |  |  |  |  | 3.3 | 96.82 | 101.08 | 102.2
9.5 |  |  |  |  | 3.2 | 96.52 | 101 | 102.2
9.6 |  |  |  |  | 3.1 | 96.14 | 101 | 102.2
9.7 |  |  |  |  | 3 | 95.74 | 101 | 102.2
9.8 |  |  |  |  | 2.98 | 95.34 | 100.98 | 102.2
9.9 |  |  |  |  | 2.88 | 94.86 | 100.9 | 102.2
10 |  |  |  |  | 2.8 | 94.28 | 100.9 | 102.2
10.1 |  |  |  |  | 2.7 | 93.48 | 100.9 | 102.2
10.2 |  |  |  |  | 2.7 | 92.6 | 100.8 | 102.2
10.3 |  |  |  |  | 2.6 | 91.32 | 100.8 | 102.13
10.4 |  |  |  |  | 2.58 | 89.86 | 100.78 | 102.1
10.5 | [5 empty] | 87.64 | 100.7 | 102.1
10.6 | [5 empty] | 84.38 | 100.6 | 102.1
10.7 | [5 empty] | 79.44 | 100.6 | 102.1
10.8 | [5 empty] | 72.56 | 100.5 | 102.1
10.9 | [5 empty] | 64.28 | 100.48 | 102.1
11 | [5 empty] | 55.14 | 100.4 | 102.1
11.1 | [5 empty] | 46.02 | 100.4 | 102.1
11.2 | [5 empty] | 37.42 | 100.3 | 102.1
11.3 | [5 empty] | 30.08 | 100.28 | 102.1
11.4 | [5 empty] | 24.5 | 100.2 | 102.1
11.5 | [5 empty] | 20.8 | 100.18 | 102.1
11.6 | [5 empty] | 18.22 | 100.08 | 102.03
11.7 | [5 empty] | 16.46 | 99.98 | 102
11.8 | [5 empty] | 15.22 | 99.9 | 102
11.9 | [5 empty] | 14.1 | 99.88 | 102
12 | [5 empty] | 13.28 | 99.8 | 102
12.1 | [5 empty] | 12.66 | 99.7 | 102
12.2 | [5 empty] | 12.06 | 99.68 | 102
12.3 | [5 empty] | 11.54 | 99.5 | 102
12.4 | [5 empty] | 11.14 | 99.48 | 101.93
12.5 | [5 empty] | 10.74 | 99.38 | 101.9
12.6 | [5 empty] | 10.34 | 99.3 | 101.9
12.7 | [5 empty] | 10.04 | 99.2 | 101.9
12.8 | [5 empty] | 9.74 | 99.18 | 101.9
12.9 | [5 empty] | 9.42 | 99.08 | 101.83
13 | [5 empty] | 9.22 | 98.98 | 101.8
13.1 | [5 empty] | 8.92 | 98.88 | 101.8
13.2 | [5 empty] | 8.7 | 98.7 | 101.8
13.3 | [5 empty] | 8.44 | 98.58 | 101.73
13.4 | [5 empty] | 8.22 | 98.4 | 101.7
13.5 | [5 empty] | 8.02 | 98.38 | 101.7
13.6 | [5 empty] | 7.82 | 98.28 | 101.7
13.7 | [5 empty] | 7.62 | 98.08 | 101.63
13.8 | [5 empty] | 7.42 | 97.98 | 101.6
13.9 | [5 empty] | 7.22 | 97.78 | 101.6
14 | [5 empty] | 7.02 | 97.58 | 101.53
14.1 | [5 empty] | 6.9 | 97.4 | 101.5
14.2 | [5 empty] | 6.72 | 97.38 | 101.5
14.3 | [5 empty] | 6.52 | 97.18 | 101.43
14.4 | [5 empty] | 6.4 | 96.88 | 101.4
14.5 | [5 empty] | 6.22 | 96.68 | 101.4
14.6 | [5 empty] | 6.1 | 96.58 | 101.33
14.7 | [5 empty] | 5.92 | 96.28 | 101.3
14.8 | [5 empty] | 5.82 | 96.08 | 101.3
14.9 | [5 empty] | 5.72 | 95.86 | 101.3
15 | [5 empty] | 5.52 | 95.46 | 101.23
15.1 | [5 empty] | 5.42 | 95.16 | 101.2
15.2 | [5 empty] | 5.32 | 94.76 | 101.2
15.3 | [5 empty] | 5.22 | 94.36 | 101.13
15.4 | [5 empty] | 5.1 | 93.84 | 101.1
15.5 | [5 empty] | 5 | 93.34 | 101.1
15.6 | [5 empty] | 4.82 | 92.62 | 101.03
15.7 | [5 empty] | 4.8 | 91.7 | 101
15.8 | [5 empty] | 4.7 | 90.58 | 101
15.9 | [5 empty] | 4.6 | 89.12 | 100.93
16 | [6 empty] | 87.02 | 100.9
16.1 | [6 empty] | 83.88 | 100.83
16.2 | [6 empty] | 79.1 | 100.8
16.3 | [6 empty] | 72.62 | 100.73
16.4 | [6 empty] | 64.44 | 100.7
16.5 | [6 empty] | 55.62 | 100.7
16.6 | [6 empty] | 46.74 | 100.63
16.7 | [6 empty] | 38.42 | 100.6
16.8 | [6 empty] | 31.3 | 100.6
16.9 | [6 empty] | 25.88 | 100.53
17 | [6 empty] | 22.22 | 100.5
17.1 | [6 empty] | 19.8 | 100.43
17.2 | [6 empty] | 18.06 | 100.4
17.3 | [6 empty] | 16.9 | 100.4
17.4 | [6 empty] | 15.92 | 100.33
17.5 | [6 empty] | 15.12 | 100.3
17.6 | [6 empty] | 14.52 | 100.23
17.7 | [6 empty] | 13.86 | 100.13
17.8 | [6 empty] | 13.46 | 100.1
17.9 | [6 empty] | 13.06 | 100.03
18 | [6 empty] | 12.66 | 100
18.1 | [6 empty] | 12.28 | 99.93
18.2 | [6 empty] | 11.96 | 99.83
18.3 | [6 empty] | 11.68 | 99.73
18.4 | [6 empty] | 11.38 | 99.7
18.5 | [6 empty] | 11.08 | 99.63
18.6 | [6 empty] | 10.88 | 99.53
18.7 | [6 empty] | 10.58 | 99.43
18.8 | [6 empty] | 10.38 | 99.33
18.9 | [6 empty] | 10.18 | 99.23
19 | [6 empty] | 9.9 | 99.13
19.1 | [6 empty] | 9.76 | 98.96
19.2 | [6 empty] | 9.48 | 98.83
19.3 | [6 empty] | 9.28 | 98.66
19.4 | [6 empty] | 9.08 | 98.53
19.5 | [6 empty] | 8.88 | 98.36
19.6 | [6 empty] | 8.68 | 98.23
19.7 | [6 empty] | 8.58 | 98.06
19.8 | [6 empty] | 8.38 | 97.86
19.9 | [6 empty] | 8.18 | 97.66
20 | [6 empty] | 7.98 | 97.46
20.1 | [6 empty] | 7.8 | 97.19
20.2 | [6 empty] | 7.68 | 96.89
20.3 | [6 empty] | 7.48 | 96.66
20.4 | [6 empty] | 7.38 | 96.32
20.5 | [6 empty] | 7.18 | 95.99
20.6 | [6 empty] | 7.08 | 95.69
20.7 | [6 empty] | 6.9 | 95.32
20.8 | [6 empty] | 6.8 | 94.92
20.9 | [6 empty] | 6.68 | 94.38
21 | [6 empty] | 6.5 | 93.71
21.1 | [6 empty] | 6.38 | 92.94
21.2 | [6 empty] | 6.28 | 92.07
21.3 | [6 empty] | 6.18 | 90.96
21.4 | [6 empty] | 6.08 | 89.55
21.5 | [7 empty] | 87.49
21.6 | [7 empty] | 84.35
21.7 | [7 empty] | 79.73
21.8 | [7 empty] | 73.3
21.9 | [7 empty] | 65.53
22 | [7 empty] | 56.87
22.1 | [7 empty] | 48.11
22.2 | [7 empty] | 39.83
22.3 | [7 empty] | 32.71
22.4 | [7 empty] | 27.27
22.5 | [7 empty] | 23.49
22.6 | [7 empty] | 20.96
22.7 | [7 empty] | 19.25
22.8 | [7 empty] | 18.03
22.9 | [7 empty] | 17.07
23 | [7 empty] | 16.24
23.1 | [7 empty] | 15.58
23.2 | [7 empty] | 15.05
23.3 | [7 empty] | 14.55
23.4 | [7 empty] | 14.12
23.5 | [7 empty] | 13.72
23.6 | [7 empty] | 13.39
23.7 | [7 empty] | 13.09
23.8 | [7 empty] | 12.79
23.9 | [7 empty] | 12.49
24 | [7 empty] | 12.19
24.1 | [7 empty] | 11.96
24.2 | [7 empty] | 11.69
24.3 | [7 empty] | 11.39
24.4 | [7 empty] | 11.16
24.5 | [7 empty] | 10.89
24.6 | [7 empty] | 10.66
24.7 | [7 empty] | 10.46
24.8 | [7 empty] | 10.26
24.9 | [7 empty] | 9.99
25 | [7 empty] | 9.76
25.1 | [7 empty] | 9.56
25.2 | [7 empty] | 9.36
25.3 | [7 empty] | 9.16
25.4 | [7 empty] | 8.96
25.5 | [7 empty] | 8.76
25.6 | [7 empty] | 8.56
25.7 | [7 empty] | 8.36
25.8 | [7 empty] | 8.16
25.9 | [7 empty] | 8.03
26 | [7 empty] | 7.86
26.1 | [7 empty] | 7.66
26.2 | [7 empty] | 7.53
26.3 | [7 empty] | 7.36
26.4 | [7 empty] | 7.23
26.5 | [7 empty] | 7.06
26.6 | [7 empty] | 6.93
26.7 | [7 empty] | 6.83
26.8 | [7 empty] | 6.66
26.9 | [7 empty] | 6.53
27 | [7 empty] | 6.43
27.1 | [7 empty] | 6.33
27.2 | [7 empty] | 6.23
27.3 | [7 empty] | 6.06
27.4 | [7 empty] | 5.93
27.5 | [7 empty] | 5.83
27.6 | [7 empty] | 5.73
27.7 | [7 empty] | 5.63
27.8 | [7 empty] | 5.53
27.9 | [7 empty] | 5.43
28 | [7 empty] | 5.33
28.1 | [7 empty] | 5.3
28.2 | [7 empty] | 5.23
28.3 | [7 empty] | 5.13
28.4 | [7 empty] | 5.03
28.5 | [7 empty] | 4.93
28.6 | [7 empty] | 4.9
28.7 | [7 empty] | 4.83
28.8 | [7 empty] | 4.73
28.9 | [7 empty] | 4.7

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 9.58
-28.8 | [7 empty] | 9.76
-28.7 | [7 empty] | 9.88
-28.6 | [7 empty] | 10.06
-28.5 | [7 empty] | 10.26
-28.4 | [7 empty] | 10.38
-28.3 | [7 empty] | 10.48
-28.2 | [7 empty] | 10.74
-28.1 | [7 empty] | 10.88
-28 | [7 empty] | 11.06
-27.9 | [7 empty] | 11.26
-27.8 | [7 empty] | 11.46
-27.7 | [7 empty] | 11.66
-27.6 | [7 empty] | 11.86
-27.5 | [7 empty] | 12.06
-27.4 | [7 empty] | 12.26
-27.3 | [7 empty] | 12.46
-27.2 | [7 empty] | 12.74
-27.1 | [7 empty] | 12.96
-27 | [7 empty] | 13.16
-26.9 | [7 empty] | 13.44
-26.8 | [7 empty] | 13.66
-26.7 | [7 empty] | 13.86
-26.6 | [7 empty] | 14.14
-26.5 | [7 empty] | 14.36
-26.4 | [7 empty] | 14.64
-26.3 | [7 empty] | 14.94
-26.2 | [7 empty] | 15.16
-26.1 | [7 empty] | 15.44
-26 | [7 empty] | 15.74
-25.9 | [7 empty] | 16.04
-25.8 | [7 empty] | 16.34
-25.7 | [7 empty] | 16.72
-25.6 | [7 empty] | 17.04
-25.5 | [7 empty] | 17.42
-25.4 | [7 empty] | 17.82
-25.3 | [7 empty] | 18.22
-25.2 | [7 empty] | 18.7
-25.1 | [7 empty] | 19.28
-25 | [7 empty] | 19.96
-24.9 | [7 empty] | 20.74
-24.8 | [7 empty] | 21.7
-24.7 | [7 empty] | 22.94
-24.6 | [7 empty] | 24.56
-24.5 | [7 empty] | 26.9
-24.4 | [7 empty] | 30.28
-24.3 | [7 empty] | 34.84
-24.2 | [7 empty] | 40.52
-24.1 | [7 empty] | 47.06
-24 | [7 empty] | 54.08
-23.9 | [7 empty] | 60.78
-23.8 | [7 empty] | 66.74
-23.7 | [7 empty] | 71.74
-23.6 | [7 empty] | 75.42
-23.5 | [7 empty] | 77.94
-23.4 | [7 empty] | 79.76
-23.3 | [7 empty] | 81.14
-23.2 | [7 empty] | 82.12
-23.1 | [7 empty] | 83.1
-23 | [6 empty] | 8.62 | 83.86
-22.9 | [6 empty] | 8.8 | 84.4
-22.8 | [6 empty] | 8.92 | 84.9
-22.7 | [6 empty] | 9.12 | 85.4
-22.6 | [6 empty] | 9.22 | 85.9
-22.5 | [6 empty] | 9.42 | 86.24
-22.4 | [6 empty] | 9.52 | 86.62
-22.3 | [6 empty] | 9.72 | 86.94
-22.2 | [6 empty] | 9.92 | 87.24
-22.1 | [6 empty] | 10.02 | 87.62
-22 | [6 empty] | 10.22 | 87.94
-21.9 | [6 empty] | 10.42 | 88.16
-21.8 | [6 empty] | 10.6 | 88.44
-21.7 | [6 empty] | 10.82 | 88.66
-21.6 | [6 empty] | 10.92 | 88.94
-21.5 | [6 empty] | 11.12 | 89.24
-21.4 | [6 empty] | 11.32 | 89.46
-21.3 | [6 empty] | 11.62 | 89.74
-21.2 | [6 empty] | 11.72 | 89.96
-21.1 | [6 empty] | 11.92 | 90.08
-21 | [6 empty] | 12.22 | 90.34
-20.9 | [6 empty] | 12.42 | 90.56
-20.8 | [6 empty] | 12.64 | 90.76
-20.7 | [6 empty] | 12.92 | 90.96
-20.6 | [6 empty] | 13.22 | 91.08
-20.5 | [6 empty] | 13.42 | 91.26
-20.4 | [6 empty] | 13.62 | 91.46
-20.3 | [6 empty] | 13.84 | 91.58
-20.2 | [6 empty] | 14.14 | 91.76
-20.1 | [6 empty] | 14.42 | 91.96
-20 | [6 empty] | 14.72 | 92.08
-19.9 | [6 empty] | 15.02 | 92.18
-19.8 | [6 empty] | 15.34 | 92.36
-19.7 | [6 empty] | 15.64 | 92.56
-19.6 | [6 empty] | 15.94 | 92.68
-19.5 | [6 empty] | 16.26 | 92.78
-19.4 | [6 empty] | 16.66 | 92.88
-19.3 | [6 empty] | 17.14 | 93.06
-19.2 | [6 empty] | 17.64 | 93.18
-19.1 | [6 empty] | 18.08 | 93.28
-19 | [6 empty] | 18.78 | 93.46
-18.9 | [6 empty] | 19.58 | 93.58
-18.8 | [6 empty] | 20.42 | 93.68
-18.7 | [6 empty] | 21.64 | 93.78
-18.6 | [6 empty] | 23.28 | 93.88
-18.5 | [6 empty] | 25.48 | 93.98
-18.4 | [6 empty] | 28.6 | 94
-18.3 | [6 empty] | 33.04 | 94.16
-18.2 | [6 empty] | 38.78 | 94.28
-18.1 | [6 empty] | 45.7 | 94.38
-18 | [6 empty] | 53.04 | 94.48
-17.9 | [6 empty] | 60.28 | 94.58
-17.8 | [6 empty] | 66.88 | 94.68
-17.7 | [6 empty] | 72.48 | 94.7
-17.6 | [6 empty] | 76.64 | 94.86
-17.5 | [6 empty] | 79.6 | 94.98
-17.4 | [6 empty] | 81.54 | 95.08
-17.3 | [6 empty] | 83.02 | 95.18
-17.2 | [6 empty] | 84.1 | 95.28
-17.1 | [6 empty] | 85.08 | 95.38
-17 | [5 empty] | 6.9 | 85.86 | 95.4
-16.9 | [5 empty] | 7 | 86.46 | 95.48
-16.8 | [5 empty] | 7.2 | 86.96 | 95.58
-16.7 | [5 empty] | 7.3 | 87.54 | 95.68
-16.6 | [5 empty] | 7.4 | 87.94 | 95.78
-16.5 | [5 empty] | 7.6 | 88.26 | 95.88
-16.4 | [5 empty] | 7.7 | 88.64 | 95.98
-16.3 | [5 empty] | 7.9 | 89.12 | 96.08
-16.2 | [5 empty] | 8.1 | 89.42 | 96.18
-16.1 | [5 empty] | 8.2 | 89.72 | 96.28
-16 | [5 empty] | 8.4 | 89.94 | 96.38
-15.9 | [5 empty] | 8.5 | 90.22 | 96.4
-15.8 | [5 empty] | 8.7 | 90.42 | 96.48
-15.7 | [5 empty] | 8.9 | 90.64 | 96.58
-15.6 | [5 empty] | 9 | 90.92 | 96.68
-15.5 | [5 empty] | 9.2 | 91.12 | 96.78
-15.4 | [5 empty] | 9.4 | 91.32 | 96.8
-15.3 | [5 empty] | 9.6 | 91.54 | 96.88
-15.2 | [5 empty] | 9.8 | 91.84 | 96.98
-15.1 | [5 empty] | 10 | 92.1 | 97.08
-15 | [5 empty] | 10.2 | 92.22 | 97.18
-14.9 | [5 empty] | 10.4 | 92.52 | 97.2
-14.8 | [5 empty] | 10.6 | 92.72 | 97.28
-14.7 | [5 empty] | 10.9 | 92.9 | 97.3
-14.6 | [5 empty] | 11.1 | 93.02 | 97.38
-14.5 | [5 empty] | 11.4 | 93.22 | 97.48
-14.4 | [5 empty] | 11.6 | 93.32 | 97.58
-14.3 | [5 empty] | 11.8 | 93.52 | 97.68
-14.2 | [5 empty] | 12.1 | 93.72 | 97.78
-14.1 | [5 empty] | 12.4 | 93.92 | 97.8
-14 | [5 empty] | 12.7 | 94.12 | 97.88
-13.9 | [5 empty] | 13 | 94.3 | 97.9
-13.8 | [5 empty] | 13.3 | 94.42 | 97.98
-13.7 | [5 empty] | 13.6 | 94.6 | 98.08
-13.6 | [5 empty] | 14 | 94.7 | 98.1
-13.5 | [5 empty] | 14.3 | 94.82 | 98.18
-13.4 | [5 empty] | 14.8 | 95.02 | 98.28
-13.3 | [5 empty] | 15.2 | 95.2 | 98.3
-13.2 | [5 empty] | 15.7 | 95.32 | 98.3
-13.1 | [5 empty] | 16.2 | 95.42 | 98.38
-13 | [5 empty] | 17 | 95.52 | 98.4
-12.9 | [5 empty] | 17.8 | 95.7 | 98.48
-12.8 | [5 empty] | 18.8 | 95.72 | 98.58
-12.7 | [5 empty] | 20 | 95.9 | 98.68
-12.6 | [5 empty] | 21.6 | 96.02 | 98.7
-12.5 | [5 empty] | 23.9 | 96.12 | 98.7
-12.4 | [5 empty] | 27.2 | 96.3 | 98.78
-12.3 | [5 empty] | 31.9 | 96.4 | 98.8
-12.2 | [5 empty] | 38 | 96.52 | 98.88
-12.1 | [5 empty] | 45.2 | 96.62 | 98.9
-12 | [5 empty] | 53 | 96.72 | 98.98
-11.9 | [5 empty] | 60.9 | 96.82 | 99.08
-11.8 | [5 empty] | 68.1 | 97 | 99.1
-11.7 | [5 empty] | 74.3 | 97.02 | 99.18
-11.6 | [5 empty] | 78.9 | 97.14 | 99.2
-11.5 | [5 empty] | 82.1 | 97.3 | 99.2
-11.4 | [5 empty] | 84.4 | 97.32 | 99.28
-11.3 | [5 empty] | 86 | 97.42 | 99.3
-11.2 | [5 empty] | 87.3 | 97.52 | 99.3
-11.1 | [5 empty] | 88.2 | 97.62 | 99.38
-11 |  |  |  |  | 4.1 | 89.1 | 97.72 | 99.4
-10.9 |  |  |  |  | 4.2 | 89.6 | 97.82 | 99.4
-10.8 |  |  |  |  | 4.3 | 90.3 | 97.92 | 99.48
-10.7 |  |  |  |  | 4.4 | 90.8 | 98.02 | 99.5
-10.6 |  |  |  |  | 4.5 | 91.2 | 98.12 | 99.58
-10.5 |  |  |  |  | 4.54 | 91.7 | 98.2 | 99.6
-10.4 |  |  |  |  | 4.6 | 92 | 98.3 | 99.6
-10.3 |  |  |  |  | 4.74 | 92.3 | 98.4 | 99.68
-10.2 |  |  |  |  | 4.9 | 92.7 | 98.4 | 99.7
-10.1 |  |  |  |  | 5 | 93 | 98.5 | 99.78
-10 |  |  |  |  | 5.1 | 93.3 | 98.6 | 99.8
-9.9 |  |  |  |  | 5.24 | 93.6 | 98.6 | 99.8
-9.8 |  |  |  |  | 5.34 | 93.8 | 98.7 | 99.88
-9.7 |  |  |  | 3.42 | 5.5 | 94.1 | 98.8 | 99.9
-9.6 |  |  |  | 3.52 | 5.64 | 94.3 | 98.9 | 99.9
-9.5 |  |  |  | 3.6 | 5.84 | 94.6 | 98.92 | 99.98
-9.4 |  |  |  | 3.7 | 5.94 | 94.9 | 99 | 100.08
-9.3 |  |  |  | 3.8 | 6.1 | 95.2 | 99.02 | 100.1
-9.2 |  |  |  | 3.9 | 6.24 | 95.3 | 99.12 | 100.1
-9.1 |  |  |  | 3.92 | 6.4 | 95.5 | 99.2 | 100.1
-9 |  |  |  | 4.02 | 6.54 | 95.6 | 99.3 | 100.18
-8.9 |  |  |  | 4.12 | 6.68 | 95.9 | 99.32 | 100.2
-8.8 |  |  |  | 4.22 | 6.94 | 96.1 | 99.4 | 100.2
-8.7 |  |  |  | 4.32 | 7.04 | 96.2 | 99.5 | 100.28
-8.6 |  |  |  | 4.5 | 7.3 | 96.4 | 99.5 | 100.3
-8.5 |  |  | 2.7 | 4.52 | 7.44 | 96.5 | 99.52 | 100.3
-8.4 |  |  | 2.8 | 4.7 | 7.74 | 96.8 | 99.6 | 100.38
-8.3 |  |  | 2.8 | 4.8 | 7.94 | 96.8 | 99.6 | 100.4
-8.2 |  |  | 2.9 | 4.92 | 8.14 | 97 | 99.7 | 100.4
-8.1 |  |  | 2.9 | 5.12 | 8.44 | 97.1 | 99.8 | 100.4
-8 |  |  | 3 | 5.22 | 8.68 | 97.3 | 99.8 | 100.48
-7.9 |  |  | 3.1 | 5.32 | 8.98 | 97.4 | 99.9 | 100.58
-7.8 |  |  | 3.1 | 5.52 | 9.24 | 97.6 | 100 | 100.52
-7.7 |  |  | 3.2 | 5.7 | 9.58 | 97.7 | 100 | 100.5
-7.6 |  |  | 3.3 | 5.82 | 9.88 | 97.8 | 100 | 100.5
-7.5 |  |  | 3.4 | 6.02 | 10.28 | 98 | 100.02 | 100.58
-7.4 |  |  | 3.5 | 6.2 | 10.68 | 98.1 | 100.1 | 100.6
-7.3 |  | 1.8 | 3.6 | 6.32 | 11.18 | 98.1 | 100.1 | 100.6
-7.2 |  | 1.82 | 3.7 | 6.62 | 11.62 | 98.3 | 100.12 | 100.6
-7.1 |  | 1.9 | 3.7 | 6.72 | 12.22 | 98.4 | 100.2 | 100.68
-7 |  | 1.9 | 3.9 | 6.94 | 13.02 | 98.5 | 100.2 | 100.7
-6.9 |  | 1.9 | 4 | 7.22 | 13.9 | 98.6 | 100.3 | 100.7
-6.8 |  | 2 | 4.1 | 7.42 | 14.94 | 98.7 | 100.3 | 100.7
-6.7 | 1.3 | 2 | 4.2 | 7.72 | 16.22 | 98.8 | 100.4 | 100.7
-6.6 | 1.32 | 2.1 | 4.3 | 8.02 | 18.04 | 98.9 | 100.4 | 100.7
-6.5 | 1.4 | 2.1 | 4.4 | 8.32 | 20.5 | 99 | 100.4 | 100.7
-6.4 | 1.4 | 2.12 | 4.6 | 8.62 | 23.94 | 99 | 100.5 | 100.7
-6.3 | 1.4 | 2.3 | 4.7 | 8.94 | 29 | 99.1 | 100.5 | 100.78
-6.2 | 1.5 | 2.3 | 4.9 | 9.34 | 35.68 | 99.2 | 100.52 | 100.8
-6.1 | 1.5 | 2.4 | 5.1 | 9.76 | 43.76 | 99.3 | 100.6 | 100.8
-6 | 1.5 | 2.42 | 5.2 | 10.26 | 52.5 | 99.4 | 100.6 | 100.8
-5.9 | 1.6 | 2.5 | 5.4 | 10.88 | 61.4 | 99.5 | 100.6 | 100.8
-5.8 | 1.6 | 2.52 | 5.5 | 11.58 | 69.72 | 99.5 | 100.7 | 100.8
-5.7 | 1.7 | 2.6 | 5.7 | 12.48 | 76.7 | 99.6 | 100.7 | 100.8
-5.6 | 1.7 | 2.72 | 6 | 13.52 | 81.98 | 99.6 | 100.7 | 100.8
-5.5 | 1.72 | 2.8 | 6.1 | 14.86 | 85.8 | 99.7 | 100.7 | 100.8
-5.4 | 1.8 | 2.9 | 6.4 | 16.72 | 88.44 | 99.7 | 100.7 | 100.8
-5.3 | 1.8 | 3 | 6.7 | 19.2 | 90.22 | 99.8 | 100.7 | 100.8
-5.2 | 1.9 | 3.1 | 7 | 22.66 | 91.58 | 99.8 | 100.7 | 100.72
-5.1 | 1.92 | 3.2 | 7.3 | 27.72 | 92.6 | 99.9 | 100.7 | 100.7
-5 | 2 | 3.32 | 7.6 | 34.48 | 93.46 | 99.9 | 100.7 | 100.7
-4.9 | 2.02 | 3.4 | 8.1 | 42.56 | 94.22 | 99.9 | 100.7 | 100.7
-4.8 | 2.1 | 3.52 | 8.5 | 51.62 | 94.86 | 100 | 100.7 | 100.7
-4.7 | 2.2 | 3.7 | 9.2 | 60.7 | 95.38 | 100 | 100.7 | 100.7
-4.6 | 2.3 | 3.82 | 9.8 | 69.28 | 95.78 | 100 | 100.7 | 100.7
-4.5 | 2.4 | 3.92 | 10.7 | 76.6 | 96.14 | 100.1 | 100.7 | 100.7
-4.4 | 2.42 | 4.1 | 11.7 | 82.18 | 96.42 | 100.1 | 100.7 | 100.7
-4.3 | 2.5 | 4.22 | 13.1 | 86.2 | 96.68 | 100.1 | 100.7 | 100.7
-4.2 | 2.6 | 4.42 | 14.8 | 89.02 | 96.94 | 100.1 | 100.7 | 100.7
-4.1 | 2.62 | 4.62 | 17.1 | 91.04 | 97.14 | 100.2 | 100.7 | 100.7
-4 | 2.72 | 4.84 | 20.6 | 92.5 | 97.44 | 100.2 | 100.7 | 100.62
-3.9 | 2.9 | 5.14 | 25.6 | 93.62 | 97.74 | 100.2 | 100.7 | 100.6
-3.8 | 3 | 5.44 | 32.4 | 94.46 | 97.94 | 100.2 | 100.7 | 100.52
-3.7 | 3.12 | 5.84 | 40.7 | 95.16 | 98.1 | 100.2 | 100.7 | 100.5
-3.6 | 3.32 | 6.34 | 50 | 95.74 | 98.24 | 100.2 | 100.7 | 100.5
-3.5 | 3.5 | 6.84 | 59.7 | 96.24 | 98.38 | 100.2 | 100.7 | 100.5
-3.4 | 3.62 | 7.36 | 68.7 | 96.66 | 98.54 | 100.2 | 100.7 | 100.5
-3.3 | 3.92 | 8.1 | 76.6 | 97.12 | 98.74 | 100.2 | 100.7 | 100.5
-3.2 | 4.12 | 9.22 | 82.8 | 97.34 | 98.84 | 100.3 | 100.68 | 100.5
-3.1 | 4.44 | 10.48 | 87.2 | 97.64 | 98.94 | 100.3 | 100.6 | 100.5
-3 | 4.84 | 12.32 | 90.2 | 97.92 | 99.14 | 100.4 | 100.6 | 100.5
-2.9 | 5.44 | 14.8 | 92.3 | 98.22 | 99.24 | 100.4 | 100.6 | 100.5
-2.8 | 5.96 | 18.22 | 93.8 | 98.4 | 99.4 | 100.4 | 100.6 | 100.5
-2.7 | 6.68 | 23.24 | 94.8 | 98.52 | 99.5 | 100.4 | 100.6 | 100.42
-2.6 | 7.62 | 30.34 | 95.7 | 98.64 | 99.54 | 100.3 | 100.6 | 100.4
-2.5 | 8.96 | 39.06 | 96.5 | 98.9 | 99.64 | 100.3 | 100.6 | 100.4
-2.4 | 10.72 | 48.92 | 97.1 | 98.92 | 99.7 | 100.4 | 100.6 | 100.4
-2.3 | 13.12 | 59.1 | 97.5 | 99.12 | 99.8 | 100.4 | 100.6 | 100.4
-2.2 | 16.74 | 68.68 | 98 | 99.22 | 99.84 | 100.3 | 100.6 | 100.4
-2.1 | 21.76 | 77.24 | 98.2 | 99.3 | 99.9 | 100.3 | 100.5 | 100.4
-2 | 29.02 | 83.82 | 98.5 | 99.4 | 99.9 | 100.3 | 100.5 | 100.32
-1.9 | 37.9 | 88.46 | 98.8 | 99.42 | 100 | 100.3 | 100.4 | 100.3
-1.8 | 48.16 | 91.46 | 99 | 99.62 | 100 | 100.2 | 100.4 | 100.3
-1.7 | 58.62 | 93.76 | 99.2 | 99.7 | 100 | 100.2 | 100.4 | 100.22
-1.6 | 68.42 | 95.22 | 99.3 | 99.72 | 100.1 | 100.2 | 100.3 | 100.2
-1.5 | 77.08 | 96.4 | 99.4 | 99.8 | 100.1 | 100.2 | 100.3 | 100.2
-1.4 | 84.06 | 97.16 | 99.5 | 99.82 | 100.1 | 100.2 | 100.3 | 100.12
-1.3 | 89.06 | 97.88 | 99.6 | 99.9 | 100.1 | 100.2 | 100.3 | 100.1
-1.2 | 92.16 | 98.44 | 99.7 | 99.9 | 100.1 | 100.2 | 100.2 | 100.1
-1.1 | 94.42 | 98.82 | 99.8 | 100 | 100.1 | 100.2 | 100.2 | 100.1
-1 | 96.12 | 99.12 | 99.8 | 100 | 100.1 | 100.1 | 100.2 | 100.1
-0.9 | 97.3 | 99.32 | 99.9 | 100 | 100.06 | 100.1 | 100.2 | 100.1
-0.8 | 98.16 | 99.52 | 99.9 | 100 | 100 | 100.1 | 100.2 | 100.1
-0.7 | 98.74 | 99.7 | 99.9 | 100 | 100.1 | 100 | 100.18 | 100.1
-0.6 | 99.14 | 99.8 | 99.9 | 100 | 100.1 | 100 | 100.1 | 100.02
-0.5 | 99.5 | 99.9 | 100 | 100 | 100.1 | 100 | 100.08 | 100
-0.4 | 99.62 | 99.92 | 100 | 100 | 100.1 | 100 | 100 | 100
-0.3 | 99.72 | 100 | 100 | 100 | 100.1 | 100 | 100 | 100
-0.2 | 99.92 | 100 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.98 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.78 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.4 | 99.68 | 99.88 | 99.9 | 99.9 | 100 | 100 | 100 | 100
0.5 | 99.36 | 99.8 | 99.9 | 99.9 | 99.96 | 100 | 100 | 100
0.6 | 98.96 | 99.78 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100
0.7 | 98.54 | 99.58 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100.07
0.8 | 97.84 | 99.48 | 99.7 | 99.8 | 99.9 | 100 | 100 | 100.1
0.9 | 97.18 | 99.26 | 99.7 | 99.8 | 99.94 | 100 | 100 | 100.1
1 | 95.98 | 98.98 | 99.6 | 99.8 | 100 | 100 | 100 | 100.1
1.1 | 94.38 | 98.74 | 99.6 | 99.8 | 100 | 100 | 100.02 | 100.1
1.2 | 92.12 | 98.14 | 99.5 | 99.78 | 100 | 100 | 100.1 | 100.1
1.3 | 88.96 | 97.64 | 99.4 | 99.7 | 99.96 | 100 | 100.1 | 100.1
1.4 | 84.1 | 96.94 | 99.4 | 99.7 | 99.9 | 100 | 100.1 | 100.17
1.5 | 77.28 | 96.22 | 99.3 | 99.58 | 99.9 | 100 | 100.1 | 100.2
1.6 | 68.76 | 95.08 | 99.1 | 99.48 | 99.9 | 100 | 100.12 | 100.2
1.7 | 58.86 | 93.52 | 99 | 99.38 | 99.8 | 100 | 100.2 | 100.2
1.8 | 48.34 | 91.26 | 98.7 | 99.3 | 99.76 | 100 | 100.2 | 100.27
1.9 | 38.24 | 88.2 | 98.5 | 99.2 | 99.66 | 100.1 | 100.2 | 100.3
2 | 29.22 | 83.7 | 98.3 | 99.18 | 99.6 | 100.1 | 100.2 | 100.3
2.1 | 22.02 | 77.12 | 98 | 99.08 | 99.5 | 100.1 | 100.28 | 100.37
2.2 | 16.88 | 68.76 | 97.7 | 98.98 | 99.5 | 100.1 | 100.22 | 100.4
2.3 | 13.22 | 59.06 | 97.4 | 98.78 | 99.46 | 100.1 | 100.3 | 100.4
2.4 | 10.72 | 48.88 | 96.8 | 98.68 | 99.4 | 100.1 | 100.3 | 100.4
2.5 | 8.94 | 38.98 | 96.3 | 98.58 | 99.3 | 100.1 | 100.3 | 100.4
2.6 | 7.6 | 30.22 | 95.6 | 98.48 | 99.2 | 100.1 | 100.3 | 100.4
2.7 | 6.72 | 23.12 | 94.6 | 98.18 | 99.1 | 100 | 100.3 | 100.4
2.8 | 5.94 | 18.04 | 93.4 | 98 | 99 | 100 | 100.3 | 100.47
2.9 | 5.34 | 14.72 | 91.9 | 97.88 | 98.86 | 100 | 100.3 | 100.5
3 | 4.86 | 12.22 | 89.8 | 97.68 | 98.8 | 100 | 100.3 | 100.5
3.1 | 4.46 | 10.46 | 86.7 | 97.38 | 98.7 | 100 | 100.3 | 100.5
3.2 | 4.26 | 9.18 | 82.2 | 97.06 | 98.56 | 100.1 | 100.3 | 100.5
3.3 | 3.88 | 8.12 | 76 | 96.74 | 98.36 | 100 | 100.32 | 100.5
3.4 | 3.68 | 7.32 | 68 | 96.36 | 98.26 | 100 | 100.4 | 100.5
3.5 | 3.48 | 6.76 | 58.9 | 95.96 | 98 | 100 | 100.4 | 100.5
3.6 | 3.3 | 6.26 | 49.3 | 95.46 | 97.9 | 100 | 100.4 | 100.5
3.7 | 3.18 | 5.76 | 39.9 | 94.84 | 97.76 | 99.9 | 100.4 | 100.5
3.8 | 2.98 | 5.46 | 31.8 | 94.22 | 97.56 | 99.9 | 100.4 | 100.57
3.9 | 2.88 | 5.16 | 25.1 | 93.28 | 97.36 | 99.9 | 100.4 | 100.6
4 | 2.78 | 4.88 | 20.3 | 92.14 | 97.16 | 99.9 | 100.4 | 100.67
4.1 | 2.6 | 4.66 | 16.9 | 90.6 | 96.92 | 99.9 | 100.4 | 100.7
4.2 | 2.5 | 4.4 | 14.6 | 88.56 | 96.56 | 99.8 | 100.38 | 100.7
4.3 | 2.48 | 4.28 | 12.9 | 85.86 | 96.36 | 99.8 | 100.32 | 100.7
4.4 | 2.4 | 4.06 | 11.5 | 81.7 | 96.06 | 99.8 | 100.4 | 100.7
4.5 | 2.3 | 3.88 | 10.6 | 76 | 95.76 | 99.7 | 100.4 | 100.7
4.6 | 2.2 | 3.78 | 9.7 | 68.66 | 95.36 | 99.6 | 100.4 | 100.7
4.7 | 2.1 | 3.6 | 9.1 | 60 | 94.98 | 99.6 | 100.4 | 100.7
4.8 | 2.1 | 3.48 | 8.4 | 50.88 | 94.42 | 99.6 | 100.38 | 100.7
4.9 | 2 | 3.38 | 8.1 | 41.98 | 93.78 | 99.6 | 100.3 | 100.7
5 | 2 | 3.3 | 7.6 | 33.88 | 93.14 | 99.6 | 100.3 | 100.7
5.1 | 1.98 | 3.18 | 7.3 | 27.26 | 92.3 | 99.6 | 100.3 | 100.7
5.2 | 1.9 | 3 | 7 | 22.32 | 91.22 | 99.5 | 100.3 | 100.77
5.3 | 1.8 | 2.98 | 6.6 | 18.94 | 89.78 | 99.4 | 100.3 | 100.8
5.4 | 1.8 | 2.88 | 6.3 | 16.5 | 87.9 | 99.4 | 100.3 | 100.8
5.5 | 1.7 | 2.78 | 6.1 | 14.76 | 85.26 | 99.3 | 100.3 | 100.8
5.6 | 1.7 | 2.68 | 6 | 13.48 | 81.58 | 99.2 | 100.28 | 100.8
5.7 | 1.68 | 2.58 | 5.8 | 12.34 | 76.16 | 99.2 | 100.2 | 100.8
5.8 | 1.6 | 2.5 | 5.5 | 11.54 | 69.32 | 99.1 | 100.2 | 100.8
5.9 | 1.58 | 2.48 | 5.4 | 10.84 | 61.12 | 99.1 | 100.2 | 100.8
6 | 1.5 | 2.4 | 5.2 | 10.26 | 52.34 | 99.1 | 100.2 | 100.8
6.1 | 1.5 | 2.38 | 5.1 | 9.74 | 43.78 | 99 | 100.2 | 100.8
6.2 | 1.48 | 2.3 | 4.9 | 9.36 | 35.72 | 98.9 | 100.2 | 100.8
6.3 | 1.4 | 2.28 | 4.7 | 8.88 | 28.98 | 98.8 | 100.1 | 100.73
6.4 | 1.4 | 2.2 | 4.6 | 8.56 | 23.96 | 98.7 | 100.1 | 100.7
6.5 | 1.3 | 2.1 | 4.4 | 8.26 | 20.54 | 98.6 | 100.1 | 100.7
6.6 | 1.3 | 2.08 | 4.3 | 7.98 | 18.1 | 98.5 | 100.08 | 100.7
6.7 | 1.3 | 2 | 4.2 | 7.68 | 16.28 | 98.5 | 100 | 100.7
6.8 |  | 1.98 | 4.1 | 7.38 | 14.92 | 98.3 | 100 | 100.7
6.9 |  | 1.9 | 4 | 7.18 | 13.84 | 98.2 | 100 | 100.7
7 |  | 1.88 | 3.8 | 6.98 | 12.94 | 98.2 | 100 | 100.7
7.1 |  | 1.8 | 3.7 | 6.68 | 12.18 | 98 | 99.9 | 100.63
7.2 |  | 1.78 | 3.7 | 6.58 | 11.62 | 97.9 | 99.8 | 100.6
7.3 |  | 1.7 | 3.6 | 6.38 | 11.18 | 97.8 | 99.8 | 100.6
7.4 |  |  | 3.5 | 6.18 | 10.66 | 97.6 | 99.7 | 100.6
7.5 |  |  | 3.4 | 5.98 | 10.32 | 97.6 | 99.68 | 100.6
7.6 |  |  | 3.3 | 5.8 | 9.92 | 97.5 | 99.6 | 100.53
7.7 |  |  | 3.2 | 5.68 | 9.62 | 97.4 | 99.6 | 100.5
7.8 |  |  | 3.1 | 5.48 | 9.22 | 97.2 | 99.6 | 100.5
7.9 |  |  | 3 | 5.38 | 8.96 | 97.1 | 99.58 | 100.5
8 |  |  | 2.9 | 5.18 | 8.72 | 96.9 | 99.5 | 100.5
8.1 |  |  | 2.9 | 5.08 | 8.36 | 96.7 | 99.48 | 100.43
8.2 |  |  | 2.8 | 4.98 | 8.16 | 96.6 | 99.3 | 100.4
8.3 |  |  | 2.8 | 4.78 | 7.96 | 96.6 | 99.3 | 100.4
8.4 |  |  | 2.7 | 4.68 | 7.76 | 96.4 | 99.2 | 100.33
8.5 |  |  | 2.6 | 4.6 | 7.52 | 96.3 | 99.2 | 100.3
8.6 |  |  |  | 4.48 | 7.26 | 96.1 | 99.2 | 100.3
8.7 |  |  |  | 4.28 | 7.06 | 95.9 | 99.1 | 100.3
8.8 |  |  |  | 4.2 | 6.96 | 95.7 | 99.08 | 100.23
8.9 |  |  |  | 4.1 | 6.76 | 95.4 | 99 | 100.2
9 |  |  |  | 4 | 6.6 | 95.3 | 98.9 | 100.2
9.1 |  |  |  | 3.9 | 6.46 | 95.2 | 98.8 | 100.13
9.2 |  |  |  | 3.8 | 6.26 | 95 | 98.8 | 100.1
9.3 |  |  |  | 3.78 | 6.1 | 94.7 | 98.78 | 100.1
9.4 |  |  |  | 3.68 | 5.96 | 94.5 | 98.58 | 100.1
9.5 |  |  |  | 3.6 | 5.76 | 94.3 | 98.5 | 100.03
9.6 |  |  |  | 3.5 | 5.66 | 94 | 98.4 | 99.93
9.7 |  |  |  | 3.4 | 5.56 | 93.8 | 98.4 | 99.9
9.8 |  |  |  |  | 5.4 | 93.6 | 98.3 | 99.9
9.9 |  |  |  |  | 5.3 | 93.3 | 98.28 | 99.83
10 |  |  |  |  | 5.2 | 93 | 98.2 | 99.8
10.1 |  |  |  |  | 5.06 | 92.8 | 98.2 | 99.73
10.2 |  |  |  |  | 4.9 | 92.5 | 98.18 | 99.7
10.3 |  |  |  |  | 4.8 | 92.1 | 98.08 | 99.7
10.4 |  |  |  |  | 4.7 | 91.7 | 97.98 | 99.63
10.5 |  |  |  |  | 4.6 | 91.4 | 97.8 | 99.6
10.6 |  |  |  |  | 4.5 | 91 | 97.7 | 99.53
10.7 |  |  |  |  | 4.4 | 90.6 | 97.6 | 99.5
10.8 |  |  |  |  | 4.3 | 90 | 97.58 | 99.5
10.9 |  |  |  |  | 4.26 | 89.4 | 97.48 | 99.43
11 | [5 empty] | 88.8 | 97.38 | 99.4
11.1 | [5 empty] | 88 | 97.3 | 99.33
11.2 | [5 empty] | 87 | 97.28 | 99.3
11.3 | [5 empty] | 85.9 | 97.18 | 99.3
11.4 | [5 empty] | 84.3 | 97.08 | 99.23
11.5 | [5 empty] | 82.1 | 96.98 | 99.2
11.6 | [5 empty] | 79 | 96.88 | 99.2
11.7 | [5 empty] | 74.6 | 96.7 | 99.2
11.8 | [5 empty] | 68.9 | 96.58 | 99.13
11.9 | [5 empty] | 61.7 | 96.4 | 99.03
12 | [5 empty] | 53.9 | 96.38 | 98.93
12.1 | [5 empty] | 46 | 96.28 | 98.9
12.2 | [5 empty] | 38.7 | 96.18 | 98.9
12.3 | [5 empty] | 32.5 | 96.08 | 98.83
12.4 | [5 empty] | 27.6 | 95.9 | 98.8
12.5 | [5 empty] | 24.2 | 95.8 | 98.73
12.6 | [5 empty] | 21.9 | 95.7 | 98.7
12.7 | [5 empty] | 20.2 | 95.58 | 98.63
12.8 | [5 empty] | 18.9 | 95.4 | 98.53
12.9 | [5 empty] | 17.9 | 95.3 | 98.5
13 | [5 empty] | 17 | 95.18 | 98.43
13.1 | [5 empty] | 16.4 | 95.1 | 98.4
13.2 | [5 empty] | 15.8 | 94.98 | 98.33
13.3 | [5 empty] | 15.2 | 94.88 | 98.3
13.4 | [5 empty] | 14.8 | 94.68 | 98.3
13.5 | [5 empty] | 14.3 | 94.48 | 98.23
13.6 | [5 empty] | 13.9 | 94.38 | 98.13
13.7 | [5 empty] | 13.6 | 94.18 | 98.03
13.8 | [5 empty] | 13.3 | 93.98 | 98
13.9 | [5 empty] | 13 | 93.88 | 97.93
14 | [5 empty] | 12.6 | 93.68 | 97.9
14.1 | [5 empty] | 12.4 | 93.5 | 97.83
14.2 | [5 empty] | 12.1 | 93.38 | 97.8
14.3 | [5 empty] | 11.8 | 93.28 | 97.73
14.4 | [5 empty] | 11.6 | 92.98 | 97.63
14.5 | [5 empty] | 11.3 | 92.8 | 97.46
14.6 | [5 empty] | 11.1 | 92.68 | 97.4
14.7 | [5 empty] | 10.8 | 92.48 | 97.33
14.8 | [5 empty] | 10.6 | 92.28 | 97.3
14.9 | [5 empty] | 10.4 | 92.08 | 97.23
15 | [5 empty] | 10.2 | 91.88 | 97.2
15.1 | [5 empty] | 10 | 91.68 | 97.06
15.2 | [5 empty] | 9.8 | 91.56 | 96.93
15.3 | [5 empty] | 9.5 | 91.28 | 96.9
15.4 | [5 empty] | 9.4 | 91.08 | 96.83
15.5 | [5 empty] | 9.2 | 90.86 | 96.73
15.6 | [5 empty] | 9 | 90.48 | 96.63
15.7 | [5 empty] | 8.8 | 90.38 | 96.6
15.8 | [5 empty] | 8.7 | 90.06 | 96.53
15.9 | [5 empty] | 8.5 | 89.78 | 96.43
16 | [5 empty] | 8.3 | 89.56 | 96.33
16.1 | [5 empty] | 8.2 | 89.28 | 96.3
16.2 | [5 empty] | 8 | 88.96 | 96.23
16.3 | [5 empty] | 7.9 | 88.66 | 96.13
16.4 | [5 empty] | 7.7 | 88.28 | 96.03
16.5 | [5 empty] | 7.6 | 87.96 | 95.93
16.6 | [5 empty] | 7.4 | 87.54 | 95.83
16.7 | [5 empty] | 7.3 | 87.16 | 95.73
16.8 | [5 empty] | 7.2 | 86.74 | 95.63
16.9 | [5 empty] | 7.1 | 86.16 | 95.53
17 | [5 empty] | 6.9 | 85.54 | 95.43
17.1 | [6 empty] | 84.8 | 95.33
17.2 | [6 empty] | 83.88 | 95.23
17.3 | [6 empty] | 82.78 | 95.13
17.4 | [6 empty] | 81.4 | 95.03
17.5 | [6 empty] | 79.38 | 94.93
17.6 | [6 empty] | 76.56 | 94.83
17.7 | [6 empty] | 72.48 | 94.73
17.8 | [6 empty] | 67.08 | 94.63
17.9 | [6 empty] | 60.58 | 94.53
18 | [6 empty] | 53.26 | 94.5
18.1 | [6 empty] | 46.06 | 94.43
18.2 | [6 empty] | 39.18 | 94.33
18.3 | [6 empty] | 33.38 | 94.16
18.4 | [6 empty] | 28.74 | 94.03
18.5 | [6 empty] | 25.54 | 93.93
18.6 | [6 empty] | 23.32 | 93.9
18.7 | [6 empty] | 21.68 | 93.83
18.8 | [6 empty] | 20.5 | 93.66
18.9 | [6 empty] | 19.52 | 93.53
19 | [6 empty] | 18.74 | 93.43
19.1 | [6 empty] | 18.14 | 93.33
19.2 | [6 empty] | 17.56 | 93.16
19.3 | [6 empty] | 17.06 | 93.03
19.4 | [6 empty] | 16.66 | 92.93
19.5 | [6 empty] | 16.26 | 92.83
19.6 | [6 empty] | 15.88 | 92.66
19.7 | [6 empty] | 15.56 | 92.53
19.8 | [6 empty] | 15.18 | 92.43
19.9 | [6 empty] | 14.88 | 92.26
20 | [6 empty] | 14.66 | 92.06
20.1 | [6 empty] | 14.36 | 91.93
20.2 | [6 empty] | 14.08 | 91.76
20.3 | [6 empty] | 13.78 | 91.63
20.4 | [6 empty] | 13.56 | 91.46
20.5 | [6 empty] | 13.28 | 91.26
20.6 | [6 empty] | 13.06 | 91.13
20.7 | [6 empty] | 12.78 | 90.96
20.8 | [6 empty] | 12.58 | 90.76
20.9 | [6 empty] | 12.28 | 90.56
21 | [6 empty] | 12.08 | 90.29
21.1 | [6 empty] | 11.88 | 90.06
21.2 | [6 empty] | 11.68 | 89.93
21.3 | [6 empty] | 11.48 | 89.69
21.4 | [6 empty] | 11.28 | 89.46
21.5 | [6 empty] | 11.08 | 89.26
21.6 | [6 empty] | 10.88 | 88.99
21.7 | [6 empty] | 10.68 | 88.69
21.8 | [6 empty] | 10.48 | 88.39
21.9 | [6 empty] | 10.3 | 88.16
22 | [6 empty] | 10.18 | 87.89
22.1 | [6 empty] | 9.98 | 87.59
22.2 | [6 empty] | 9.78 | 87.22
22.3 | [6 empty] | 9.6 | 86.96
22.4 | [6 empty] | 9.48 | 86.55
22.5 | [6 empty] | 9.28 | 86.26
22.6 | [6 empty] | 9.18 | 85.85
22.7 | [6 empty] | 8.98 | 85.35
22.8 | [6 empty] | 8.88 | 84.85
22.9 | [6 empty] | 8.7 | 84.42
23 | [7 empty] | 83.88
23.1 | [7 empty] | 83.07
23.2 | [7 empty] | 82.17
23.3 | [7 empty] | 81.13
23.4 | [7 empty] | 79.82
23.5 | [7 empty] | 77.93
23.6 | [7 empty] | 75.41
23.7 | [7 empty] | 71.73
23.8 | [7 empty] | 66.72
23.9 | [7 empty] | 60.76
24 | [7 empty] | 54.07
24.1 | [7 empty] | 47.1
24.2 | [7 empty] | 40.52
24.3 | [7 empty] | 34.89
24.4 | [7 empty] | 30.29
24.5 | [7 empty] | 26.9
24.6 | [7 empty] | 24.53
24.7 | [7 empty] | 22.92
24.8 | [7 empty] | 21.73
24.9 | [7 empty] | 20.7
25 | [7 empty] | 19.91
25.1 | [7 empty] | 19.28
25.2 | [7 empty] | 18.75
25.3 | [7 empty] | 18.25
25.4 | [7 empty] | 17.82
25.5 | [7 empty] | 17.42
25.6 | [7 empty] | 17.09
25.7 | [7 empty] | 16.72
25.8 | [7 empty] | 16.39
25.9 | [7 empty] | 16.09
26 | [7 empty] | 15.72
26.1 | [7 empty] | 15.46
26.2 | [7 empty] | 15.19
26.3 | [7 empty] | 14.89
26.4 | [7 empty] | 14.66
26.5 | [7 empty] | 14.39
26.6 | [7 empty] | 14.09
26.7 | [7 empty] | 13.86
26.8 | [7 empty] | 13.66
26.9 | [7 empty] | 13.39
27 | [7 empty] | 13.16
27.1 | [7 empty] | 12.89
27.2 | [7 empty] | 12.73
27.3 | [7 empty] | 12.49
27.4 | [7 empty] | 12.26
27.5 | [7 empty] | 12.06
27.6 | [7 empty] | 11.86
27.7 | [7 empty] | 11.66
27.8 | [7 empty] | 11.46
27.9 | [7 empty] | 11.26
28 | [7 empty] | 11.06
28.1 | [7 empty] | 10.93
28.2 | [7 empty] | 10.69
28.3 | [7 empty] | 10.53
28.4 | [7 empty] | 10.36
28.5 | [7 empty] | 10.23
28.6 | [7 empty] | 10.06
28.7 | [7 empty] | 9.93
28.8 | [7 empty] | 9.76
28.9 | [7 empty] | 9.63

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-30.9 | [7 empty] | 11.84
-30.8 | [7 empty] | 12.04
-30.7 | [7 empty] | 12.24
-30.6 | [7 empty] | 12.37
-30.5 | [7 empty] | 12.54
-30.4 | [7 empty] | 12.74
-30.3 | [7 empty] | 12.94
-30.2 | [7 empty] | 13.14
-30.1 | [7 empty] | 13.34
-30 | [7 empty] | 13.47
-29.9 | [7 empty] | 13.64
-29.8 | [7 empty] | 13.84
-29.7 | [7 empty] | 14.04
-29.6 | [7 empty] | 14.24
-29.5 | [7 empty] | 14.44
-29.4 | [7 empty] | 14.57
-29.3 | [7 empty] | 14.81
-29.2 | [7 empty] | 15.11
-29.1 | [7 empty] | 15.34
-29 | [7 empty] | 15.47
-28.9 | [7 empty] | 15.71
-28.8 | [7 empty] | 15.94
-28.7 | [7 empty] | 16.21
-28.6 | [7 empty] | 16.44
-28.5 | [7 empty] | 16.71
-28.4 | [7 empty] | 16.94
-28.3 | [7 empty] | 17.14
-28.2 | [7 empty] | 17.41
-28.1 | [7 empty] | 17.71
-28 | [7 empty] | 18.01
-27.9 | [7 empty] | 18.31
-27.8 | [7 empty] | 18.54
-27.7 | [7 empty] | 18.81
-27.6 | [7 empty] | 19.18
-27.5 | [7 empty] | 19.51
-27.4 | [7 empty] | 19.88
-27.3 | [7 empty] | 20.28
-27.2 | [7 empty] | 20.75
-27.1 | [7 empty] | 21.25
-27 | [7 empty] | 21.89
-26.9 | [7 empty] | 22.59
-26.8 | [7 empty] | 23.43
-26.7 | [7 empty] | 24.54
-26.6 | [7 empty] | 26.09
-26.5 | [7 empty] | 28.21
-26.4 | [7 empty] | 31
-26.3 | [7 empty] | 34.84
-26.2 | [7 empty] | 39.53
-26.1 | [7 empty] | 44.99
-26 | [7 empty] | 50.69
-25.9 | [7 empty] | 56.25
-25.8 | [7 empty] | 61.4
-25.7 | [7 empty] | 65.63
-25.6 | [7 empty] | 68.97
-25.5 | [7 empty] | 71.37
-25.4 | [7 empty] | 73.05
-25.3 | [7 empty] | 74.41
-25.2 | [7 empty] | 75.43
-25.1 | [7 empty] | 76.26
-25 | [7 empty] | 76.92
-24.9 | [7 empty] | 77.52
-24.8 | [7 empty] | 78.12
-24.7 | [7 empty] | 78.65
-24.6 | [7 empty] | 79.01
-24.5 | [6 empty] | 11 | 79.45
-24.4 | [6 empty] | 11.2 | 79.88
-24.3 | [6 empty] | 11.3 | 80.21
-24.2 | [6 empty] | 11.5 | 80.51
-24.1 | [6 empty] | 11.6 | 80.81
-24 | [6 empty] | 11.8 | 81.11
-23.9 | [6 empty] | 12 | 81.48
-23.8 | [6 empty] | 12.1 | 81.74
-23.7 | [6 empty] | 12.3 | 82.08
-23.6 | [6 empty] | 12.5 | 82.41
-23.5 | [6 empty] | 12.7 | 82.64
-23.4 | [6 empty] | 12.9 | 82.91
-23.3 | [6 empty] | 13 | 83.14
-23.2 | [6 empty] | 13.3 | 83.34
-23.1 | [6 empty] | 13.5 | 83.61
-23 | [6 empty] | 13.7 | 83.84
-22.9 | [6 empty] | 13.9 | 84.11
-22.8 | [6 empty] | 14.1 | 84.34
-22.7 | [6 empty] | 14.3 | 84.61
-22.6 | [6 empty] | 14.6 | 84.84
-22.5 | [6 empty] | 14.8 | 85.04
-22.4 | [6 empty] | 15 | 85.24
-22.3 | [6 empty] | 15.2 | 85.44
-22.2 | [6 empty] | 15.4 | 85.64
-22.1 | [6 empty] | 15.7 | 85.84
-22 | [6 empty] | 15.9 | 86.04
-21.9 | [6 empty] | 16.2 | 86.24
-21.8 | [6 empty] | 16.4 | 86.44
-21.7 | [6 empty] | 16.7 | 86.64
-21.6 | [6 empty] | 16.9 | 86.84
-21.5 | [6 empty] | 17.3 | 86.97
-21.4 | [6 empty] | 17.6 | 87.14
-21.3 | [6 empty] | 17.9 | 87.27
-21.2 | [6 empty] | 18.2 | 87.44
-21.1 | [6 empty] | 18.5 | 87.57
-21 | [6 empty] | 18.8 | 87.74
-20.9 | [6 empty] | 19.3 | 87.94
-20.8 | [6 empty] | 19.7 | 88.07
-20.7 | [6 empty] | 20.2 | 88.24
-20.6 | [6 empty] | 20.8 | 88.37
-20.5 | [6 empty] | 21.5 | 88.54
-20.4 | [6 empty] | 22.2 | 88.67
-20.3 | [6 empty] | 23.2 | 88.77
-20.2 | [6 empty] | 24.5 | 88.94
-20.1 | [6 empty] | 26.1 | 89.14
-20 | [6 empty] | 28.4 | 89.27
-19.9 | [6 empty] | 31.6 | 89.37
-19.8 | [6 empty] | 35.9 | 89.47
-19.7 | [6 empty] | 41.3 | 89.64
-19.6 | [6 empty] | 47.2 | 89.77
-19.5 | [6 empty] | 53.4 | 89.87
-19.4 | [6 empty] | 59.5 | 89.97
-19.3 | [6 empty] | 64.7 | 90.14
-19.2 | [6 empty] | 69.2 | 90.27
-19.1 | [6 empty] | 72.6 | 90.44
-19 | [6 empty] | 75.1 | 90.57
-18.9 | [6 empty] | 76.8 | 90.74
-18.8 | [6 empty] | 78.1 | 90.87
-18.7 | [6 empty] | 79 | 90.97
-18.6 | [6 empty] | 79.8 | 91.07
-18.5 | [6 empty] | 80.5 | 91.17
-18.4 | [6 empty] | 81.1 | 91.27
-18.3 | [6 empty] | 81.6 | 91.37
-18.2 | [6 empty] | 82.2 | 91.47
-18.1 | [6 empty] | 82.6 | 91.64
-18 | [5 empty] | 9 | 83 | 91.77
-17.9 | [5 empty] | 9.12 | 83.4 | 91.87
-17.8 | [5 empty] | 9.32 | 83.8 | 91.97
-17.7 | [5 empty] | 9.42 | 84 | 92.07
-17.6 | [5 empty] | 9.62 | 84.4 | 92.24
-17.5 | [5 empty] | 9.72 | 84.6 | 92.37
-17.4 | [5 empty] | 9.92 | 84.9 | 92.47
-17.3 | [5 empty] | 10.02 | 85.2 | 92.64
-17.2 | [5 empty] | 10.2 | 85.5 | 92.77
-17.1 | [5 empty] | 10.32 | 85.8 | 92.8
-17 | [5 empty] | 10.54 | 86 | 92.87
-16.9 | [5 empty] | 10.72 | 86.3 | 92.97
-16.8 | [5 empty] | 10.92 | 86.5 | 93.07
-16.7 | [5 empty] | 11.2 | 86.8 | 93.17
-16.6 | [5 empty] | 11.32 | 87.1 | 93.34
-16.5 | [5 empty] | 11.52 | 87.4 | 93.47
-16.4 | [5 empty] | 11.74 | 87.6 | 93.57
-16.3 | [5 empty] | 12.02 | 87.8 | 93.6
-16.2 | [5 empty] | 12.22 | 88 | 93.67
-16.1 | [5 empty] | 12.42 | 88.2 | 93.84
-16 | [5 empty] | 12.62 | 88.4 | 94.04
-15.9 | [5 empty] | 12.74 | 88.7 | 94.17
-15.8 | [5 empty] | 13.12 | 88.8 | 94.27
-15.7 | [5 empty] | 13.32 | 89 | 94.37
-15.6 | [5 empty] | 13.54 | 89.2 | 94.47
-15.5 | [5 empty] | 13.82 | 89.5 | 94.57
-15.4 | [5 empty] | 14.02 | 89.7 | 94.67
-15.3 | [5 empty] | 14.24 | 89.9 | 94.77
-15.2 | [5 empty] | 14.54 | 90.1 | 94.87
-15.1 | [5 empty] | 14.82 | 90.2 | 94.9
-15 | [5 empty] | 15.12 | 90.3 | 94.97
-14.9 | [5 empty] | 15.42 | 90.4 | 95.07
-14.8 | [5 empty] | 15.74 | 90.7 | 95.17
-14.7 | [5 empty] | 16.04 | 90.9 | 95.27
-14.6 | [5 empty] | 16.46 | 91.1 | 95.3
-14.5 | [5 empty] | 16.84 | 91.3 | 95.37
-14.4 | [5 empty] | 17.24 | 91.4 | 95.47
-14.3 | [5 empty] | 17.74 | 91.5 | 95.57
-14.2 | [5 empty] | 18.18 | 91.7 | 95.67
-14.1 | [5 empty] | 18.78 | 91.9 | 95.77
-14 | [5 empty] | 19.48 | 92 | 95.87
-13.9 | [5 empty] | 20.3 | 92.2 | 95.97
-13.8 | [5 empty] | 21.3 | 92.4 | 96
-13.7 | [5 empty] | 22.56 | 92.5 | 96.07
-13.6 | [5 empty] | 24.08 | 92.6 | 96.1
-13.5 | [5 empty] | 26.26 | 92.9 | 96.17
-13.4 | [5 empty] | 29.36 | 93 | 96.27
-13.3 | [5 empty] | 33.6 | 93.1 | 96.37
-13.2 | [5 empty] | 39.1 | 93.2 | 96.47
-13.1 | [5 empty] | 45.4 | 93.3 | 96.5
-13 | [5 empty] | 52.32 | 93.5 | 96.57
-12.9 | [5 empty] | 59.36 | 93.6 | 96.67
-12.8 | [5 empty] | 65.72 | 93.7 | 96.77
-12.7 | [5 empty] | 71.18 | 93.9 | 96.8
-12.6 | [5 empty] | 75.44 | 94 | 96.87
-12.5 | [5 empty] | 78.44 | 94.1 | 96.97
-12.4 | [5 empty] | 80.68 | 94.3 | 97.07
-12.3 | [5 empty] | 82.22 | 94.4 | 97.1
-12.2 | [5 empty] | 83.62 | 94.5 | 97.17
-12.1 | [5 empty] | 84.5 | 94.6 | 97.27
-12 | [5 empty] | 85.36 | 94.7 | 97.3
-11.9 | [5 empty] | 86.06 | 94.9 | 97.37
-11.8 | [5 empty] | 86.56 | 95 | 97.47
-11.7 | [5 empty] | 87.14 | 95.1 | 97.5
-11.6 | [5 empty] | 87.58 | 95.2 | 97.57
-11.5 |  |  |  |  | 5.6 | 88.14 | 95.4 | 97.67
-11.4 |  |  |  |  | 5.7 | 88.42 | 95.4 | 97.7
-11.3 |  |  |  |  | 5.8 | 88.76 | 95.5 | 97.7
-11.2 |  |  |  |  | 5.9 | 89.14 | 95.6 | 97.77
-11.1 |  |  |  |  | 6.1 | 89.52 | 95.7 | 97.8
-11 |  |  |  |  | 6.3 | 89.86 | 95.8 | 97.87
-10.9 |  |  |  |  | 6.4 | 90.22 | 95.9 | 97.9
-10.8 |  |  |  |  | 6.5 | 90.54 | 96.1 | 97.97
-10.7 |  |  |  |  | 6.6 | 90.74 | 96.2 | 98.07
-10.6 |  |  |  |  | 6.7 | 91.04 | 96.2 | 98.1
-10.5 |  |  |  |  | 6.9 | 91.3 | 96.3 | 98.17
-10.4 |  |  |  |  | 7 | 91.52 | 96.4 | 98.2
-10.3 |  |  |  |  | 7.2 | 91.64 | 96.5 | 98.27
-10.2 |  |  |  |  | 7.4 | 92.02 | 96.7 | 98.37
-10.1 |  |  |  | 4.76 | 7.5 | 92.22 | 96.7 | 98.4
-10 |  |  |  | 4.86 | 7.7 | 92.42 | 96.8 | 98.47
-9.9 |  |  |  | 4.96 | 7.9 | 92.62 | 96.9 | 98.57
-9.8 |  |  |  | 5.06 | 8 | 92.94 | 97 | 98.6
-9.7 |  |  |  | 5.2 | 8.2 | 93.14 | 97.1 | 98.67
-9.6 |  |  |  | 5.3 | 8.3 | 93.42 | 97.2 | 98.7
-9.5 |  |  |  | 5.4 | 8.6 | 93.6 | 97.3 | 98.77
-9.4 |  |  |  | 5.5 | 8.7 | 93.74 | 97.3 | 98.87
-9.3 |  |  |  | 5.66 | 8.9 | 94.02 | 97.4 | 98.9
-9.2 |  |  |  | 5.76 | 9.1 | 94.3 | 97.5 | 98.9
-9.1 |  |  |  | 5.86 | 9.4 | 94.42 | 97.6 | 98.97
-9 |  |  |  | 5.96 | 9.5 | 94.52 | 97.7 | 99
-8.9 |  |  |  | 6.06 | 9.7 | 94.82 | 97.8 | 99.07
-8.8 |  |  | 3.8 | 6.26 | 10 | 95.02 | 97.8 | 99.1
-8.7 |  |  | 3.8 | 6.36 | 10.2 | 95.12 | 97.9 | 99.17
-8.6 |  |  | 3.9 | 6.56 | 10.4 | 95.22 | 98.1 | 99.2
-8.5 |  |  | 4 | 6.66 | 10.7 | 95.4 | 98.2 | 99.2
-8.4 |  |  | 4.06 | 6.86 | 11 | 95.62 | 98.2 | 99.27
-8.3 |  |  | 4.1 | 7.06 | 11.4 | 95.72 | 98.3 | 99.37
-8.2 |  |  | 4.2 | 7.2 | 11.7 | 95.82 | 98.3 | 99.4
-8.1 |  |  | 4.26 | 7.36 | 12 | 96.02 | 98.4 | 99.47
-8 |  |  | 4.3 | 7.46 | 12.4 | 96.2 | 98.4 | 99.5
-7.9 |  |  | 4.46 | 7.66 | 12.8 | 96.32 | 98.5 | 99.57
-7.8 |  |  | 4.56 | 7.86 | 13.3 | 96.52 | 98.5 | 99.6
-7.7 |  |  | 4.7 | 8.12 | 13.8 | 96.7 | 98.6 | 99.67
-7.6 |  |  | 4.8 | 8.36 | 14.5 | 96.72 | 98.7 | 99.7
-7.5 |  | 2.4 | 4.9 | 8.52 | 15.2 | 96.82 | 98.7 | 99.7
-7.4 |  | 2.46 | 5 | 8.76 | 16.1 | 97.02 | 98.8 | 99.77
-7.3 |  | 2.5 | 5.1 | 8.96 | 17.1 | 97.1 | 98.8 | 99.8
-7.2 |  | 2.56 | 5.26 | 9.16 | 18.4 | 97.22 | 98.9 | 99.8
-7.1 |  | 2.6 | 5.4 | 9.62 | 20.3 | 97.32 | 99 | 99.87
-7 |  | 2.7 | 5.5 | 9.96 | 22.7 | 97.5 | 99 | 99.9
-6.9 | 1.66 | 2.7 | 5.66 | 10.16 | 26.1 | 97.6 | 99.1 | 99.9
-6.8 | 1.7 | 2.76 | 5.86 | 10.42 | 30.9 | 97.72 | 99.1 | 99.9
-6.7 | 1.8 | 2.8 | 5.96 | 10.78 | 37.4 | 97.82 | 99.2 | 99.9
-6.6 | 1.8 | 2.9 | 6.16 | 11.22 | 44.6 | 98 | 99.2 | 99.97
-6.5 | 1.9 | 3 | 6.3 | 11.62 | 52.8 | 98.02 | 99.2 | 100
-6.4 | 1.96 | 3.1 | 6.46 | 12.18 | 61 | 98.12 | 99.3 | 100
-6.3 | 2 | 3.1 | 6.66 | 12.82 | 68.9 | 98.3 | 99.3 | 100
-6.2 | 2.06 | 3.2 | 6.86 | 13.54 | 75.3 | 98.4 | 99.4 | 100
-6.1 | 2.1 | 3.3 | 7 | 14.44 | 80.4 | 98.42 | 99.4 | 100.07
-6 | 2.1 | 3.46 | 7.26 | 15.56 | 84 | 98.52 | 99.5 | 100.1
-5.9 | 2.16 | 3.5 | 7.46 | 16.94 | 86.8 | 98.6 | 99.5 | 100.17
-5.8 | 2.26 | 3.6 | 7.72 | 18.64 | 88.5 | 98.7 | 99.5 | 100.2
-5.7 | 2.3 | 3.66 | 7.96 | 21.14 | 89.9 | 98.7 | 99.6 | 100.2
-5.6 | 2.3 | 3.76 | 8.16 | 24.66 | 91.1 | 98.82 | 99.6 | 100.2
-5.5 | 2.4 | 3.9 | 8.46 | 29.42 | 92 | 98.9 | 99.6 | 100.2
-5.4 | 2.4 | 4 | 8.78 | 35.8 | 92.7 | 99 | 99.7 | 100.2
-5.3 | 2.4 | 4.1 | 9.16 | 43.4 | 93.3 | 99 | 99.7 | 100.27
-5.2 | 2.5 | 4.2 | 9.48 | 51.88 | 93.9 | 99.02 | 99.7 | 100.3
-5.1 | 2.6 | 4.3 | 10.08 | 60.42 | 94.4 | 99.12 | 99.7 | 100.3
-5 | 2.7 | 4.46 | 10.68 | 68.58 | 94.7 | 99.22 | 99.8 | 100.3
-4.9 | 2.76 | 4.6 | 11.44 | 75.52 | 95.1 | 99.22 | 99.8 | 100.3
-4.8 | 2.86 | 4.76 | 12.24 | 80.94 | 95.3 | 99.3 | 99.8 | 100.3
-4.7 | 2.96 | 4.86 | 13.26 | 84.88 | 95.6 | 99.4 | 99.8 | 100.3
-4.6 | 3.06 | 5.06 | 14.68 | 87.82 | 95.9 | 99.4 | 99.9 | 100.3
-4.5 | 3.16 | 5.26 | 16.46 | 89.78 | 96.1 | 99.5 | 99.9 | 100.3
-4.4 | 3.2 | 5.46 | 18.74 | 91.28 | 96.6 | 99.52 | 99.9 | 100.23
-4.3 | 3.3 | 5.56 | 22.2 | 92.5 | 96.9 | 99.6 | 99.9 | 100.27
-4.2 | 3.4 | 5.86 | 26.88 | 93.4 | 96.9 | 99.6 | 99.9 | 100.3
-4.1 | 3.5 | 6.12 | 33.14 | 94.18 | 97.1 | 99.7 | 99.9 | 100.3
-4 | 3.6 | 6.52 | 40.98 | 94.82 | 97.3 | 99.7 | 99.9 | 100.3
-3.9 | 3.82 | 6.92 | 49.8 | 95.38 | 97.5 | 99.7 | 99.9 | 100.3
-3.8 | 3.96 | 7.32 | 58.64 | 95.82 | 97.8 | 99.8 | 99.9 | 100.3
-3.7 | 4.1 | 7.82 | 67.32 | 96.18 | 98 | 99.8 | 100 | 100.3
-3.6 | 4.26 | 8.44 | 75.14 | 96.46 | 98.1 | 99.8 | 100 | 100.3
-3.5 | 4.46 | 9.4 | 81.12 | 96.76 | 98.2 | 99.8 | 100 | 100.3
-3.4 | 4.76 | 10.4 | 85.64 | 96.96 | 98.4 | 99.82 | 100 | 100.3
-3.3 | 5.12 | 11.78 | 88.94 | 97.42 | 98.5 | 99.9 | 100 | 100.37
-3.2 | 5.48 | 13.56 | 91.1 | 97.6 | 98.6 | 99.92 | 100 | 100.4
-3.1 | 5.92 | 16 | 92.66 | 97.86 | 98.7 | 100 | 100 | 100.4
-3 | 6.48 | 19.5 | 93.96 | 98.02 | 98.9 | 100 | 100 | 100.4
-2.9 | 7.18 | 24.64 | 95.04 | 98.26 | 99 | 100 | 100 | 100.33
-2.8 | 8.16 | 31.18 | 95.78 | 98.4 | 99.1 | 100.02 | 100 | 100.3
-2.7 | 9.22 | 39.5 | 96.38 | 98.5 | 99.2 | 100.1 | 100 | 100.3
-2.6 | 10.82 | 48.88 | 96.98 | 98.66 | 99.3 | 100.1 | 100 | 100.3
-2.5 | 12.98 | 58.64 | 97.42 | 98.8 | 99.3 | 100.1 | 100 | 100.3
-2.4 | 15.92 | 67.9 | 97.76 | 98.96 | 99.4 | 100.1 | 100 | 100.3
-2.3 | 20.04 | 76.08 | 97.96 | 99.06 | 99.5 | 100.18 | 100 | 100.3
-2.2 | 26.04 | 82.54 | 98.32 | 99.2 | 99.6 | 100.1 | 100 | 100.23
-2.1 | 33.78 | 87.4 | 98.52 | 99.3 | 99.6 | 100.1 | 100 | 100.2
-2 | 43.12 | 90.74 | 98.66 | 99.3 | 99.8 | 100.1 | 100.1 | 100.2
-1.9 | 52.96 | 92.9 | 98.86 | 99.36 | 99.8 | 100.1 | 100.1 | 100.2
-1.8 | 62.94 | 94.68 | 99 | 99.46 | 99.8 | 100.1 | 100.1 | 100.13
-1.7 | 72.22 | 95.96 | 99.06 | 99.56 | 99.9 | 100.1 | 100.1 | 100.1
-1.6 | 79.76 | 96.84 | 99.2 | 99.66 | 99.9 | 100.1 | 100 | 100.1
-1.5 | 85.72 | 97.48 | 99.36 | 99.8 | 100 | 100.1 | 100 | 100.1
-1.4 | 89.62 | 98.02 | 99.46 | 99.8 | 100 | 100.1 | 100 | 100.1
-1.3 | 92.42 | 98.36 | 99.56 | 99.8 | 100 | 100.1 | 100 | 100.1
-1.2 | 94.5 | 98.72 | 99.6 | 99.86 | 99.9 | 100.1 | 100 | 100.1
-1.1 | 95.86 | 99.06 | 99.76 | 99.8 | 99.9 | 100 | 99.9 | 100.1
-1 | 96.96 | 99.26 | 99.8 | 99.9 | 100 | 100 | 99.9 | 100.1
-0.9 | 97.92 | 99.46 | 99.86 | 99.9 | 100 | 100 | 99.9 | 100.1
-0.8 | 98.48 | 99.6 | 99.9 | 100 | 100 | 100 | 99.9 | 100.03
-0.7 | 98.98 | 99.76 | 99.9 | 100 | 100 | 100 | 99.9 | 100
-0.6 | 99.26 | 99.86 | 100 | 100 | 100 | 100 | 99.9 | 100
-0.5 | 99.46 | 99.9 | 100 | 100 | 100 | 100 | 99.9 | 100
-0.4 | 99.66 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100
-0.3 | 99.76 | 100 | 100.1 | 100 | 100 | 100 | 100 | 100
-0.2 | 99.9 | 100 | 100.04 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.94 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.84 | 99.94 | 100 | 100 | 100 | 99.98 | 100 | 100
0.3 | 99.7 | 99.9 | 100 | 100 | 100 | 99.9 | 100 | 100
0.4 | 99.54 | 99.84 | 100 | 100 | 100 | 99.9 | 100 | 100
0.5 | 99.34 | 99.8 | 100 | 100 | 100 | 99.9 | 99.9 | 100
0.6 | 99.14 | 99.64 | 100 | 99.94 | 99.9 | 99.9 | 99.9 | 100
0.7 | 98.88 | 99.54 | 99.94 | 99.9 | 99.9 | 99.9 | 99.9 | 100
0.8 | 98.42 | 99.5 | 99.9 | 99.8 | 99.9 | 100 | 99.9 | 100
0.9 | 97.78 | 99.4 | 99.84 | 99.8 | 99.8 | 100 | 99.9 | 100.07
1 | 97.04 | 99.14 | 99.7 | 99.8 | 99.8 | 100 | 99.9 | 100.1
1.1 | 95.72 | 98.94 | 99.6 | 99.74 | 99.8 | 100 | 99.9 | 100.1
1.2 | 94.32 | 98.68 | 99.5 | 99.7 | 99.8 | 100 | 100 | 100.1
1.3 | 92.42 | 98.28 | 99.44 | 99.7 | 99.8 | 99.98 | 100 | 100.1
1.4 | 89.64 | 97.82 | 99.34 | 99.7 | 99.8 | 99.9 | 100 | 100.1
1.5 | 85.52 | 97.28 | 99.3 | 99.6 | 99.7 | 99.9 | 100 | 100.1
1.6 | 79.72 | 96.56 | 99.2 | 99.54 | 99.7 | 99.9 | 100 | 100.1
1.7 | 72.4 | 95.6 | 99.1 | 99.5 | 99.7 | 99.9 | 100.1 | 100.1
1.8 | 63.32 | 94.32 | 99 | 99.34 | 99.6 | 99.9 | 100.1 | 100.17
1.9 | 53.44 | 92.5 | 98.84 | 99.3 | 99.6 | 99.9 | 100.1 | 100.2
2 | 43.42 | 90.48 | 98.64 | 99.14 | 99.5 | 99.9 | 100.1 | 100.2
2.1 | 34 | 87.24 | 98.28 | 99.1 | 99.4 | 99.9 | 100 | 100.2
2.2 | 26.3 | 82.48 | 98.04 | 99 | 99.3 | 99.9 | 100 | 100.2
2.3 | 20.18 | 76.04 | 97.84 | 98.84 | 99.2 | 99.9 | 100 | 100.27
2.4 | 15.88 | 67.9 | 97.54 | 98.74 | 99.1 | 99.9 | 100 | 100.3
2.5 | 12.92 | 58.66 | 97.34 | 98.6 | 99 | 99.8 | 100 | 100.3
2.6 | 10.82 | 48.82 | 96.82 | 98.38 | 99 | 99.8 | 100 | 100.3
2.7 | 9.18 | 39.4 | 96.22 | 98.14 | 98.9 | 99.8 | 100 | 100.3
2.8 | 8.1 | 31.26 | 95.72 | 98.04 | 98.8 | 99.8 | 100 | 100.3
2.9 | 7.22 | 24.5 | 94.96 | 97.84 | 98.7 | 99.8 | 100 | 100.3
3 | 6.48 | 19.44 | 93.84 | 97.7 | 98.5 | 99.7 | 100 | 100.37
3.1 | 5.86 | 15.96 | 92.64 | 97.54 | 98.3 | 99.7 | 100 | 100.4
3.2 | 5.48 | 13.5 | 91.1 | 97.34 | 98.2 | 99.7 | 100 | 100.4
3.3 | 5.08 | 11.72 | 89.02 | 97.04 | 98.2 | 99.6 | 100 | 100.4
3.4 | 4.78 | 10.34 | 85.82 | 96.78 | 98 | 99.6 | 100 | 100.33
3.5 | 4.44 | 9.3 | 81.5 | 96.48 | 97.9 | 99.6 | 100 | 100.3
3.6 | 4.24 | 8.4 | 75.42 | 96.12 | 97.7 | 99.5 | 100 | 100.3
3.7 | 4.14 | 7.78 | 68 | 95.84 | 97.6 | 99.5 | 100 | 100.3
3.8 | 3.94 | 7.24 | 59.36 | 95.52 | 97.5 | 99.4 | 99.9 | 100.3
3.9 | 3.74 | 6.88 | 50.36 | 95.02 | 97.2 | 99.4 | 99.9 | 100.3
4 | 3.64 | 6.48 | 41.58 | 94.42 | 97 | 99.3 | 99.9 | 100.3
4.1 | 3.54 | 6.08 | 33.68 | 93.82 | 96.8 | 99.3 | 99.9 | 100.3
4.2 | 3.44 | 5.84 | 27.32 | 93.1 | 96.7 | 99.3 | 99.9 | 100.3
4.3 | 3.3 | 5.64 | 22.44 | 92.14 | 96.4 | 99.2 | 99.9 | 100.3
4.4 | 3.2 | 5.44 | 19 | 90.88 | 96.1 | 99.2 | 99.9 | 100.3
4.5 | 3.14 | 5.24 | 16.6 | 89.42 | 95.9 | 99.2 | 99.9 | 100.3
4.6 | 3.04 | 4.98 | 14.76 | 87.38 | 95.7 | 99.1 | 99.9 | 100.3
4.7 | 2.94 | 4.84 | 13.34 | 84.42 | 95.4 | 99.1 | 99.8 | 100.3
4.8 | 2.9 | 4.74 | 12.3 | 80.48 | 95 | 99 | 99.8 | 100.3
4.9 | 2.8 | 4.64 | 11.46 | 75.04 | 94.6 | 98.98 | 99.8 | 100.3
5 | 2.7 | 4.4 | 10.72 | 68.12 | 94.3 | 98.88 | 99.8 | 100.3
5.1 | 2.6 | 4.24 | 10.12 | 60.08 | 93.9 | 98.8 | 99.7 | 100.3
5.2 | 2.5 | 4.2 | 9.68 | 51.46 | 93.3 | 98.78 | 99.7 | 100.3
5.3 | 2.44 | 4.1 | 9.12 | 43.04 | 92.9 | 98.7 | 99.7 | 100.23
5.4 | 2.4 | 3.94 | 8.74 | 35.5 | 92.3 | 98.6 | 99.7 | 100.2
5.5 | 2.4 | 3.8 | 8.42 | 29.18 | 91.5 | 98.5 | 99.6 | 100.2
5.6 | 2.34 | 3.74 | 8.14 | 24.44 | 90.6 | 98.4 | 99.6 | 100.2
5.7 | 2.3 | 3.6 | 7.94 | 21.06 | 89.5 | 98.4 | 99.6 | 100.2
5.8 | 2.24 | 3.5 | 7.74 | 18.64 | 88 | 98.3 | 99.5 | 100.2
5.9 | 2.2 | 3.4 | 7.48 | 16.86 | 86.1 | 98.28 | 99.5 | 100.2
6 | 2.1 | 3.34 | 7.24 | 15.44 | 83.5 | 98.18 | 99.5 | 100.13
6.1 | 2.1 | 3.3 | 6.94 | 14.46 | 79.7 | 98.1 | 99.4 | 100.1
6.2 | 2.1 | 3.24 | 6.8 | 13.56 | 74.6 | 98 | 99.4 | 100.03
6.3 | 2 | 3.1 | 6.54 | 12.82 | 68.2 | 97.9 | 99.3 | 100
6.4 | 1.94 | 3 | 6.44 | 12.18 | 60.6 | 97.78 | 99.3 | 100
6.5 | 1.9 | 2.94 | 6.3 | 11.68 | 52.4 | 97.7 | 99.2 | 99.93
6.6 | 1.84 | 2.84 | 6.14 | 11.28 | 44.3 | 97.6 | 99.2 | 99.9
6.7 | 1.8 | 2.8 | 5.94 | 10.78 | 36.9 | 97.5 | 99.2 | 99.9
6.8 | 1.8 | 2.74 | 5.74 | 10.48 | 30.7 | 97.4 | 99.1 | 99.9
6.9 | 1.7 | 2.74 | 5.64 | 10.14 | 25.9 | 97.3 | 99.1 | 99.9
7 |  | 2.7 | 5.54 | 9.84 | 22.5 | 97.18 | 99 | 99.9
7.1 |  | 2.6 | 5.44 | 9.58 | 20.2 | 97.08 | 99 | 99.9
7.2 |  | 2.6 | 5.24 | 9.24 | 18.4 | 96.98 | 98.9 | 99.83
7.3 |  | 2.5 | 5.14 | 8.98 | 17.1 | 96.9 | 98.8 | 99.8
7.4 |  | 2.44 | 5.04 | 8.8 | 16 | 96.68 | 98.8 | 99.73
7.5 |  | 2.4 | 4.94 | 8.54 | 15.2 | 96.58 | 98.7 | 99.7
7.6 |  |  | 4.8 | 8.34 | 14.4 | 96.4 | 98.7 | 99.7
7.7 |  |  | 4.7 | 8.08 | 13.8 | 96.28 | 98.6 | 99.63
7.8 |  |  | 4.54 | 7.9 | 13.3 | 96.1 | 98.5 | 99.6
7.9 |  |  | 4.44 | 7.74 | 12.8 | 96.08 | 98.5 | 99.6
8 |  |  | 4.4 | 7.54 | 12.3 | 95.9 | 98.4 | 99.53
8.1 |  |  | 4.3 | 7.44 | 11.9 | 95.88 | 98.4 | 99.43
8.2 |  |  | 4.2 | 7.24 | 11.6 | 95.58 | 98.3 | 99.4
8.3 |  |  | 4.2 | 7.04 | 11.3 | 95.48 | 98.3 | 99.33
8.4 |  |  | 4.1 | 6.9 | 11 | 95.36 | 98.2 | 99.3
8.5 |  |  | 3.94 | 6.74 | 10.7 | 95.18 | 98.2 | 99.23
8.6 |  |  | 3.84 | 6.54 | 10.4 | 94.98 | 98.1 | 99.2
8.7 |  |  | 3.74 | 6.4 | 10.2 | 94.78 | 97.9 | 99.13
8.8 |  |  | 3.7 | 6.24 | 10 | 94.66 | 97.8 | 99.1
8.9 |  |  |  | 6.1 | 9.7 | 94.4 | 97.8 | 99.1
9 |  |  |  | 6.04 | 9.5 | 94.3 | 97.7 | 99.03
9.1 |  |  |  | 5.9 | 9.3 | 94.18 | 97.6 | 99
9.2 |  |  |  | 5.74 | 9 | 93.98 | 97.5 | 98.93
9.3 |  |  |  | 5.64 | 8.8 | 93.78 | 97.4 | 98.9
9.4 |  |  |  | 5.44 | 8.6 | 93.58 | 97.3 | 98.83
9.5 |  |  |  | 5.4 | 8.5 | 93.26 | 97.3 | 98.8
9.6 |  |  |  | 5.3 | 8.3 | 93.08 | 97.2 | 98.73
9.7 |  |  |  | 5.24 | 8.1 | 92.9 | 97.1 | 98.7
9.8 |  |  |  | 5.04 | 8 | 92.76 | 97 | 98.63
9.9 |  |  |  | 4.94 | 7.7 | 92.48 | 96.9 | 98.53
10 |  |  |  | 4.84 | 7.6 | 92.28 | 96.8 | 98.5
10.1 |  |  |  | 4.8 | 7.4 | 91.98 | 96.7 | 98.43
10.2 |  |  |  |  | 7.3 | 91.78 | 96.7 | 98.33
10.3 |  |  |  |  | 7.1 | 91.58 | 96.5 | 98.23
10.4 |  |  |  |  | 7 | 91.36 | 96.4 | 98.2
10.5 |  |  |  |  | 6.8 | 91.08 | 96.3 | 98.13
10.6 |  |  |  |  | 6.7 | 90.78 | 96.2 | 98.1
10.7 |  |  |  |  | 6.6 | 90.58 | 96.2 | 98.03
10.8 |  |  |  |  | 6.5 | 90.26 | 96.1 | 98
10.9 |  |  |  |  | 6.3 | 89.86 | 95.9 | 97.93
11 |  |  |  |  | 6.2 | 89.58 | 95.8 | 97.9
11.1 |  |  |  |  | 6.1 | 89.16 | 95.7 | 97.83
11.2 |  |  |  |  | 6 | 88.86 | 95.6 | 97.8
11.3 |  |  |  |  | 5.9 | 88.56 | 95.5 | 97.73
11.4 |  |  |  |  | 5.7 | 88.18 | 95.4 | 97.7
11.5 | [5 empty] | 87.78 | 95.4 | 97.63
11.6 | [5 empty] | 87.44 | 95.2 | 97.53
11.7 | [5 empty] | 86.86 | 95.1 | 97.5
11.8 | [5 empty] | 86.36 | 95 | 97.43
11.9 | [5 empty] | 85.74 | 94.9 | 97.4
12 | [5 empty] | 85.12 | 94.7 | 97.33
12.1 | [5 empty] | 84.28 | 94.6 | 97.23
12.2 | [5 empty] | 83.28 | 94.5 | 97.2
12.3 | [5 empty] | 82.06 | 94.4 | 97.13
12.4 | [5 empty] | 80.5 | 94.3 | 97.03
12.5 | [5 empty] | 78.42 | 94.1 | 97
12.6 | [5 empty] | 75.42 | 94 | 96.93
12.7 | [5 empty] | 71.26 | 93.9 | 96.83
12.8 | [5 empty] | 66.12 | 93.7 | 96.73
12.9 | [5 empty] | 59.92 | 93.6 | 96.63
13 | [5 empty] | 53 | 93.5 | 96.6
13.1 | [5 empty] | 46.22 | 93.3 | 96.53
13.2 | [5 empty] | 39.7 | 93.2 | 96.5
13.3 | [5 empty] | 34.12 | 93.1 | 96.43
13.4 | [5 empty] | 29.74 | 93 | 96.33
13.5 | [5 empty] | 26.54 | 92.9 | 96.23
13.6 | [5 empty] | 24.24 | 92.6 | 96.13
13.7 | [5 empty] | 22.64 | 92.5 | 96.1
13.8 | [5 empty] | 21.3 | 92.4 | 96.03
13.9 | [5 empty] | 20.32 | 92.2 | 95.93
14 | [5 empty] | 19.52 | 92 | 95.83
14.1 | [5 empty] | 18.84 | 91.9 | 95.73
14.2 | [5 empty] | 18.24 | 91.7 | 95.63
14.3 | [5 empty] | 17.74 | 91.5 | 95.53
14.4 | [5 empty] | 17.26 | 91.4 | 95.5
14.5 | [5 empty] | 16.86 | 91.3 | 95.43
14.6 | [5 empty] | 16.46 | 91.1 | 95.33
14.7 | [5 empty] | 16.08 | 90.9 | 95.23
14.8 | [5 empty] | 15.76 | 90.7 | 95.13
14.9 | [5 empty] | 15.38 | 90.4 | 95.03
15 | [5 empty] | 15.16 | 90.3 | 94.93
15.1 | [5 empty] | 14.86 | 90.2 | 94.9
15.2 | [5 empty] | 14.58 | 90.1 | 94.83
15.3 | [5 empty] | 14.36 | 89.9 | 94.73
15.4 | [5 empty] | 14.06 | 89.7 | 94.7
15.5 | [5 empty] | 13.78 | 89.5 | 94.56
15.6 | [5 empty] | 13.56 | 89.2 | 94.43
15.7 | [5 empty] | 13.28 | 89 | 94.33
15.8 | [5 empty] | 13.08 | 88.8 | 94.23
15.9 | [5 empty] | 12.88 | 88.7 | 94.13
16 | [5 empty] | 12.66 | 88.4 | 94.03
16.1 | [5 empty] | 12.48 | 88.2 | 93.86
16.2 | [5 empty] | 12.16 | 88 | 93.73
16.3 | [5 empty] | 11.98 | 87.8 | 93.63
16.4 | [5 empty] | 11.78 | 87.6 | 93.53
16.5 | [5 empty] | 11.58 | 87.4 | 93.43
16.6 | [5 empty] | 11.38 | 87.1 | 93.33
16.7 | [5 empty] | 11.18 | 86.8 | 93.23
16.8 | [5 empty] | 10.98 | 86.5 | 93.13
16.9 | [5 empty] | 10.78 | 86.3 | 93.03
17 | [5 empty] | 10.58 | 86 | 92.93
17.1 | [5 empty] | 10.38 | 85.8 | 92.83
17.2 | [5 empty] | 10.28 | 85.5 | 92.73
17.3 | [5 empty] | 10.08 | 85.2 | 92.63
17.4 | [5 empty] | 9.9 | 84.9 | 92.53
17.5 | [5 empty] | 9.68 | 84.6 | 92.43
17.6 | [5 empty] | 9.58 | 84.4 | 92.26
17.7 | [5 empty] | 9.48 | 84 | 92.06
17.8 | [5 empty] | 9.36 | 83.8 | 91.93
17.9 | [5 empty] | 9.1 | 83.4 | 91.83
18 | [6 empty] | 83 | 91.73
18.1 | [6 empty] | 82.6 | 91.63
18.2 | [6 empty] | 82.2 | 91.53
18.3 | [6 empty] | 81.6 | 91.43
18.4 | [6 empty] | 81.1 | 91.33
18.5 | [6 empty] | 80.5 | 91.23
18.6 | [6 empty] | 79.8 | 91.13
18.7 | [6 empty] | 79 | 90.96
18.8 | [6 empty] | 78.1 | 90.83
18.9 | [6 empty] | 76.8 | 90.73
19 | [6 empty] | 75.1 | 90.56
19.1 | [6 empty] | 72.6 | 90.43
19.2 | [6 empty] | 69.2 | 90.33
19.3 | [6 empty] | 64.7 | 90.16
19.4 | [6 empty] | 59.5 | 90.03
19.5 | [6 empty] | 53.4 | 89.93
19.6 | [6 empty] | 47.2 | 89.83
19.7 | [6 empty] | 41.3 | 89.66
19.8 | [6 empty] | 35.9 | 89.53
19.9 | [6 empty] | 31.6 | 89.43
20 | [6 empty] | 28.4 | 89.26
20.1 | [6 empty] | 26.1 | 89.13
20.2 | [6 empty] | 24.5 | 88.96
20.3 | [6 empty] | 23.2 | 88.76
20.4 | [6 empty] | 22.2 | 88.63
20.5 | [6 empty] | 21.5 | 88.53
20.6 | [6 empty] | 20.8 | 88.43
20.7 | [6 empty] | 20.2 | 88.26
20.8 | [6 empty] | 19.7 | 88.06
20.9 | [6 empty] | 19.3 | 87.93
21 | [6 empty] | 18.8 | 87.76
21.1 | [6 empty] | 18.5 | 87.56
21.2 | [6 empty] | 18.2 | 87.43
21.3 | [6 empty] | 17.9 | 87.26
21.4 | [6 empty] | 17.6 | 87.13
21.5 | [6 empty] | 17.3 | 86.96
21.6 | [6 empty] | 16.9 | 86.83
21.7 | [6 empty] | 16.7 | 86.66
21.8 | [6 empty] | 16.4 | 86.46
21.9 | [6 empty] | 16.2 | 86.26
22 | [6 empty] | 15.9 | 86.06
22.1 | [6 empty] | 15.7 | 85.86
22.2 | [6 empty] | 15.4 | 85.66
22.3 | [6 empty] | 15.2 | 85.46
22.4 | [6 empty] | 15 | 85.26
22.5 | [6 empty] | 14.8 | 85.06
22.6 | [6 empty] | 14.6 | 84.86
22.7 | [6 empty] | 14.3 | 84.59
22.8 | [6 empty] | 14.1 | 84.36
22.9 | [6 empty] | 13.9 | 84.09
23 | [6 empty] | 13.7 | 83.86
23.1 | [6 empty] | 13.5 | 83.66
23.2 | [6 empty] | 13.3 | 83.39
23.3 | [6 empty] | 13 | 83.16
23.4 | [6 empty] | 12.9 | 82.89
23.5 | [6 empty] | 12.7 | 82.66
23.6 | [6 empty] | 12.5 | 82.39
23.7 | [6 empty] | 12.3 | 82.09
23.8 | [6 empty] | 12.1 | 81.79
23.9 | [6 empty] | 12 | 81.49
24 | [6 empty] | 11.8 | 81.12
24.1 | [6 empty] | 11.6 | 80.86
24.2 | [6 empty] | 11.5 | 80.59
24.3 | [6 empty] | 11.3 | 80.29
24.4 | [6 empty] | 11.2 | 79.92
24.5 | [6 empty] | 11 | 79.45
24.6 | [7 empty] | 79.09
24.7 | [7 empty] | 78.65
24.8 | [7 empty] | 78.15
24.9 | [7 empty] | 77.58
25 | [7 empty] | 76.98
25.1 | [7 empty] | 76.31
25.2 | [7 empty] | 75.54
25.3 | [7 empty] | 74.53
25.4 | [7 empty] | 73.22
25.5 | [7 empty] | 71.54
25.6 | [7 empty] | 69.18
25.7 | [7 empty] | 65.95
25.8 | [7 empty] | 61.75
25.9 | [7 empty] | 56.62
26 | [7 empty] | 51.08
26.1 | [7 empty] | 45.41
26.2 | [7 empty] | 39.92
26.3 | [7 empty] | 35.15
26.4 | [7 empty] | 31.28
26.5 | [7 empty] | 28.31
26.6 | [7 empty] | 26.17
26.7 | [7 empty] | 24.69
26.8 | [7 empty] | 23.53
26.9 | [7 empty] | 22.64
27 | [7 empty] | 21.91
27.1 | [7 empty] | 21.28
27.2 | [7 empty] | 20.75
27.3 | [7 empty] | 20.32
27.4 | [7 empty] | 19.92
27.5 | [7 empty] | 19.52
27.6 | [7 empty] | 19.19
27.7 | [7 empty] | 18.89
27.8 | [7 empty] | 18.59
27.9 | [7 empty] | 18.29
28 | [7 empty] | 17.99
28.1 | [7 empty] | 17.69
28.2 | [7 empty] | 17.46
28.3 | [7 empty] | 17.19
28.4 | [7 empty] | 16.96
28.5 | [7 empty] | 16.69
28.6 | [7 empty] | 16.46
28.7 | [7 empty] | 16.19
28.8 | [7 empty] | 15.96
28.9 | [7 empty] | 15.76
29 | [7 empty] | 15.56
29.1 | [7 empty] | 15.29
29.2 | [7 empty] | 15.06
29.3 | [7 empty] | 14.79
29.4 | [7 empty] | 14.56
29.5 | [7 empty] | 14.43
29.6 | [7 empty] | 14.26
29.7 | [7 empty] | 14.06
29.8 | [7 empty] | 13.86
29.9 | [7 empty] | 13.66
30 | [7 empty] | 13.46
30.1 | [7 empty] | 13.33
30.2 | [7 empty] | 13.16
30.3 | [7 empty] | 12.96
30.4 | [7 empty] | 12.76
30.5 | [7 empty] | 12.56
30.6 | [7 empty] | 12.36
30.7 | [7 empty] | 12.23
30.8 | [7 empty] | 12.06
30.9 | [7 empty] | 11.86

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2.4 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-43.7 |  |  |  |  | 4.9134
-43.6 |  |  |  |  | 5
-43.5 |  |  |  |  | 5.0012
-43.4 |  |  |  |  | 5.1
-43.3 |  |  |  |  | 5.1
-43.2 |  |  |  |  | 5.1042
-43.1 |  |  |  |  | 5.2053
-43 |  |  |  |  | 5.3
-42.9 |  |  |  |  | 5.3073
-42.8 |  |  |  |  | 5.4
-42.7 |  |  |  |  | 5.4093
-42.6 |  |  |  |  | 5.5
-42.5 |  |  |  |  | 5.5
-42.4 |  |  |  |  | 5.5124
-42.3 |  |  |  |  | 5.6134
-42.2 |  |  |  |  | 5.7
-42.1 |  |  |  |  | 5.7011
-42 |  |  |  |  | 5.8
-41.9 |  |  |  |  | 5.8031
-41.8 |  |  |  |  | 5.9
-41.7 |  |  |  |  | 5.9053
-41.6 |  |  |  |  | 6.0063
-41.5 |  |  |  |  | 6.1
-41.4 |  |  |  |  | 6.1083
-41.3 |  |  |  |  | 6.2093
-41.2 |  |  |  |  | 6.3103
-41.1 |  |  |  |  | 6.4
-41 |  |  |  |  | 6.4123
-40.9 |  |  |  | 3.8 | 6.5
-40.8 |  |  |  | 3.8858 | 6.5143
-40.7 |  |  |  | 3.9 | 6.6134
-40.6 |  |  |  | 3.9878 | 6.7021
-40.5 |  |  |  | 4 | 6.8
-40.4 |  |  |  | 4 | 6.8041
-40.3 |  |  |  | 4.0766 | 6.9052
-40.2 |  |  |  | 4.1 | 7.0062
-40.1 |  |  |  | 4.1 | 7.1
-40 |  |  |  | 4.1796 | 7.1082
-39.9 |  |  |  | 4.2 | 7.2092
-39.8 |  |  |  | 4.2816 | 7.3102
-39.7 |  |  |  | 4.3 | 7.4112
-39.6 |  |  |  | 4.3836 | 7.5122
-39.5 |  |  |  | 4.4 | 7.6132
-39.4 |  |  |  | 4.4857 | 7.7142
-39.3 |  |  |  | 4.5 | 7.8134
-39.2 |  |  |  | 4.5878 | 7.902
-39.1 |  |  |  | 4.6888 | 8
-39 |  |  |  | 4.7 | 8.004
-38.9 |  |  |  | 4.7794 | 8.1051
-38.8 |  |  |  | 4.8 | 8.2061
-38.7 |  |  |  | 4.8785 | 8.3071
-38.6 |  |  |  | 4.9795 | 8.4081
-38.5 |  |  |  | 5 | 8.5091
-38.4 |  |  |  | 5.0816 | 8.6101
-38.3 |  |  |  | 5.1 | 8.7222
-38.2 |  |  |  | 5.1836 | 8.9122
-38.1 |  |  | 2.7 | 5.2846 | 9.0132
-38 |  |  | 2.7571 | 5.3857 | 9.1142
-37.9 |  |  | 2.8 | 5.4 | 9.2133
-37.8 |  |  | 2.8 | 5.4877 | 9.3019
-37.7 |  |  | 2.8 | 5.5887 | 9.4059
-37.6 |  |  | 2.8611 | 5.6 | 9.6039
-37.5 |  |  | 2.9 | 5.6794 | 9.7049
-37.4 |  |  | 2.9553 | 5.7775 | 9.806
-37.3 |  |  | 3 | 5.8785 | 9.9141
-37.2 |  |  | 3 | 5.9795 | 10.1081
-37.1 |  |  | 3.0519 | 6 | 10.2091
-37 |  |  | 3.1 | 6.0815 | 10.3202
-36.9 |  |  | 3.1 | 6.1825 | 10.5111
-36.8 |  |  | 3.1549 | 6.2835 | 10.6242
-36.7 |  | 2.1 | 3.2 | 6.3845 | 10.8131
-36.6 |  | 2.1 | 3.257 | 6.4856 | 10.9141
-36.5 |  | 2.1437 | 3.3 | 6.5 | 11.0265
-36.4 |  | 2.2 | 3.359 | 6.5876 | 11.2019
-36.3 |  | 2.2 | 3.4 | 6.7772 | 11.3059
-36.2 |  | 2.2468 | 3.461 | 6.8 | 11.5039
-36.1 |  | 2.3 | 3.5 | 6.8793 | 11.6099
-36 | 1.9 | 2.3 | 3.5 | 6.9774 | 11.8119
-35.9 | 1.9 | 2.3 | 3.5498 | 7.0784 | 12.0139
-35.8 | 1.908 | 2.3445 | 3.6508 | 7.1794 | 12.216
-35.7 | 2 | 2.4 | 3.7 | 7.3608 | 12.409
-35.6 | 2 | 2.4 | 3.7528 | 7.4814 | 12.52
-35.5 | 2 | 2.4396 | 3.8 | 7.5824 | 12.722
-35.4 | 2.012 | 2.5 | 3.8548 | 7.6834 | 12.924
-35.3 | 2.1 | 2.5 | 3.9 | 7.7844 | 13.1261
-35.2 | 2.1 | 2.5426 | 3.9569 | 7.8855 | 13.3281
-35.1 | 2.1 | 2.6 | 4.058 | 7.9865 | 13.5263
-35 | 2.1141 | 2.6 | 4.1 | 8.1749 | 13.7036
-34.9 | 2.2 | 2.6457 | 4.16 | 8.2885 | 13.9085
-34.8 | 2.2 | 2.7 | 4.2 | 8.3896 | 14.2077
-34.7 | 2.2 | 2.7477 | 4.262 | 8.4792 | 14.4145
-34.6 | 2.2059 | 2.8 | 4.3551 | 8.6545 | 14.7117
-34.5 | 2.3 | 2.8 | 4.4 | 8.7783 | 14.9137
-34.4 | 2.3 | 2.8444 | 4.4507 | 8.8793 | 15.1236
-34.3 | 2.3089 | 2.9 | 4.5518 | 9.0606 | 15.4356
-34.2 | 2.4 | 2.9385 | 4.6528 | 9.1813 | 15.8396
-34.1 | 2.4 | 3 | 4.7538 | 9.2823 | 16.2436
-34 | 2.4119 | 3.0405 | 4.8 | 9.4668 | 16.6477
-33.9 | 2.5 | 3.1 | 4.8559 | 9.5844 | 17.0776
-33.8 | 2.514 | 3.1425 | 4.9569 | 9.7709 | 17.6977
-33.7 | 2.6 | 3.2 | 5.0579 | 9.9729 | 18.3921
-33.6 | 2.6 | 3.2445 | 5.1589 | 10.1749 | 19.0172
-33.5 | 2.6027 | 3.3 | 5.2599 | 10.377 | 20.0327
-33.4 | 2.7037 | 3.3467 | 5.3 | 10.4895 | 21.2486
-33.3 | 2.8 | 3.4477 | 5.3619 | 10.6583 | 22.576
-33.2 | 2.8 | 3.5 | 5.4551 | 10.8545 | 24.2094
-33.1 | 2.8068 | 3.5497 | 5.5497 | 11.0566 | 26.1557
-33 | 2.9 | 3.6 | 5.7014 | 11.2586 | 28.4944
-32.9 | 2.9088 | 3.6452 | 5.8 | 11.3803 | 31.0639
-32.8 | 3.0099 | 3.7384 | 5.8527 | 11.5626 | 34.1267
-32.7 | 3.1 | 3.8 | 5.9537 | 11.7646 | 37.5036
-32.6 | 3.1119 | 3.8404 | 6.1095 | 12.05 | 41.2172
-32.5 | 3.2 | 3.9415 | 6.2558 | 12.2687 | 44.7396
-32.4 | 3.2139 | 4.0425 | 6.3568 | 12.4707 | 48.16
-32.3 | 3.3149 | 4.1 | 6.4578 | 12.6727 | 51.3661
-32.2 | 3.4 | 4.1445 | 6.5588 | 12.9621 | 53.8388
-32.1 | 3.4026 | 4.2456 | 6.6598 | 13.1768 | 56.2499
-32 | 3.5037 | 4.3466 | 6.8216 | 13.4682 | 58.1523
-31.9 | 3.6 | 4.4476 | 6.9618 | 13.7374 | 59.557
-31.8 | 3.6058 | 4.5486 | 7.055 | 14.0315 | 60.7576
-31.7 | 3.7068 | 4.6496 | 7.1496 | 14.4127 | 61.7609
-31.6 | 3.8 | 4.7 | 7.3012 | 14.7376 | 62.6622
-31.5 | 3.8088 | 4.7452 | 7.4516 | 15.201 | 63.4527
-31.4 | 3.9098 | 4.8384 | 7.5526 | 15.7061 | 64.0784
-31.3 | 4.0108 | 4.9394 | 7.7073 | 16.2933 | 64.8648
-31.2 | 4.1118 | 5.0404 | 7.9093 | 16.9826 | 65.4709
-31.1 | 4.2128 | 5.1414 | 8.0557 | 17.8582 | 66.0898
-31 | 4.3138 | 5.2424 | 8.2133 | 18.8525 | 66.783
-30.9 | 4.4 | 5.3434 | 8.4154 | 20.0352 | 67.391
-30.8 | 4.4139 | 5.4444 | 8.5587 | 21.5964 | 68.0081
-30.7 | 4.5026 | 5.5455 | 8.7195 | 23.3891 | 68.5158
-30.6 | 4.6036 | 5.6465 | 8.9216 | 25.6537 | 69.1182
-30.5 | 4.7046 | 5.7949 | 9.1236 | 28.0342 | 69.6279
-30.4 | 4.8057 | 5.9485 | 9.3099 | 31.0663 | 70.2283
-30.3 | 4.9067 | 6.0495 | 9.499 | 34.6109 | 70.74
-30.2 | 5.0077 | 6.1505 | 9.6505 | 38.6427 | 71.3307
-30.1 | 5.1087 | 6.2451 | 9.803 | 43.0244 | 71.7347
-30 | 5.2097 | 6.3766 | 10.0576 | 47.4689 | 72.1388
-29.9 | 5.3214 | 6.5393 | 10.3071 | 51.7491 | 72.5428
-29.8 | 5.5117 | 6.6806 | 10.5092 | 55.7421 | 72.9469
-29.7 | 5.6127 | 6.8413 | 10.7113 | 59.1767 | 73.3509
-29.6 | 5.7137 | 6.9846 | 10.97 | 61.9991 | 73.7412
-29.5 | 5.8147 | 7.1433 | 11.2154 | 64.2112 | 74.0516
-29.4 | 5.9138 | 7.2443 | 11.4761 | 65.9836 | 74.4061
-29.3 | 6.0025 | 7.3907 | 11.7791 | 67.2594 | 74.8076
-29.2 | 6.1071 | 7.5927 | 12.1428 | 68.3822 | 75.1106
-29.1 | 6.3045 | 7.7474 | 12.4852 | 69.2108 | 75.4136
-29 | 6.4111 | 7.8969 | 12.8745 | 70.0158 | 75.7111
-28.9 | 6.6066 | 8.099 | 13.3475 | 70.9798 | 75.9197
-28.8 | 6.7152 | 8.2505 | 13.9535 | 71.7529 | 76.2227
-28.7 | 6.9086 | 8.3901 | 14.7121 | 72.5406 | 76.5343
-28.6 | 7.0096 | 8.5764 | 15.5202 | 73.3489 | 76.9192
-28.5 | 7.1212 | 8.7784 | 16.596 | 74.157 | 77.1212
-28.4 | 7.3233 | 8.9804 | 18.0182 | 74.8819 | 77.3349
-28.3 | 7.5127 | 9.2236 | 19.7 | 75.5048 | 77.6382
-28.2 | 7.6275 | 9.4844 | 21.801 | 76.1961 | 77.9275
-28.1 | 7.8147 | 9.7298 | 24.4697 | 76.9031 | 78.1258
-28 | 7.9276 | 10.033 | 27.6919 | 77.4359 | 78.3028
-27.9 | 8.1048 | 10.3361 | 31.6626 | 78.0291 | 78.5073
-27.8 | 8.3069 | 10.6855 | 36.3697 | 78.6352 | 78.8103
-27.7 | 8.5089 | 11.0895 | 41.6273 | 79.0948 | 79.1089
-27.6 | 8.7055 | 11.5419 | 47.3323 | 79.5848 | 79.3109
-27.5 | 8.8197 | 12.1458 | 52.7308 | 80.0899 | 79.5197
-27.4 | 9.1152 | 12.9536 | 58.0722 | 80.516 | 79.8227
-27.3 | 9.3258 | 13.85 | 63.2679 | 80.84 | 80.1086
-27.2 | 9.6384 | 14.9345 | 67.6445 | 81.143 | 80.2192
-27.1 | 10.0424 | 16.5055 | 71.1565 | 81.6101 | 80.4318
-27 | 10.4465 | 18.4844 | 73.7978 | 81.9491 | 80.7348
-26.9 | 10.8758 | 20.9539 | 75.7427 | 82.3362 | 81.0253
-26.8 | 11.5091 | 24.12 | 77.2905 | 82.6552 | 81.2409
-26.7 | 12.3611 | 28.1022 | 78.5322 | 82.8721 | 81.5256
-26.6 | 13.4645 | 32.9891 | 79.7018 | 83.1612 | 81.7014
-26.5 | 14.5436 | 38.7604 | 80.7355 | 83.5523 | 81.8073
-26.4 | 16.3756 | 45.2459 | 81.7656 | 83.7782 | 82.1069
-26.3 | 18.6289 | 52.1618 | 82.6921 | 84.0366 | 82.3133
-26.2 | 21.6073 | 59.0349 | 83.6259 | 84.2537 | 82.6055
-26.1 | 25.5038 | 65.4083 | 84.5009 | 84.5336 | 82.7194
-26 | 30.3261 | 70.9152 | 85.3527 | 84.7578 | 83.0149
-25.9 | 36.1685 | 75.1966 | 86.2105 | 85.0397 | 83.217
-25.8 | 43.0026 | 78.4663 | 87.0186 | 85.3427 | 83.419
-25.7 | 50.4984 | 80.9818 | 87.7733 | 85.5638 | 83.621
-25.6 | 58.1521 | 82.8416 | 88.4804 | 85.7659 | 83.823
-25.5 | 65.5392 | 84.3344 | 89.1321 | 85.9679 | 84.0251
-25.4 | 72.1309 | 85.5633 | 89.7382 | 86.2548 | 84.2406
-25.3 | 77.4109 | 86.6744 | 90.2872 | 86.5579 | 84.5255
-25.2 | 81.4085 | 87.8297 | 90.7924 | 86.7739 | 84.7013
-25.1 | 84.0465 | 88.9967 | 91.238 | 86.976 | 84.8046
-25 | 86.0533 | 90.0616 | 91.642 | 87.2672 | 85.0067
-24.9 | 87.6565 | 91.0245 | 92.0461 | 87.5364 | 85.2087
-24.8 | 88.9642 | 91.9818 | 92.4501 | 87.6768 | 85.4107
-24.7 | 90.17 | 92.8935 | 92.7111 | 87.8556 | 85.6064
-24.6 | 91.2811 | 93.7518 | 92.9004 | 88.0576 | 85.7221
-24.5 | 92.3922 | 94.5138 | 93.2051 | 88.3394 | 86.0084
-24.4 | 93.4845 | 95.1659 | 93.557 | 88.5616 | 86.1094
-24.3 | 94.404 | 95.8339 | 93.8067 | 88.7636 | 86.2208
-24.2 | 95.4141 | 96.44 | 94.063 | 88.9658 | 86.4228
-24.1 | 96.4118 | 96.964 | 94.3107 | 89.1679 | 86.6248
-24 | 97.3076 | 97.4101 | 94.5691 | 89.3699 | 86.8134
-23.9 | 98.0873 | 97.8291 | 94.8147 | 89.6579 | 86.9254
-23.8 | 98.6952 | 98.1762 | 95.0168 | 89.8739 | 87.1012
-23.7 | 99.3133 | 98.4902 | 95.2188 | 90.076 | 87.2044
-23.6 | 99.9162 | 98.7385 | 95.4812 | 90.278 | 87.4032
-23.5 | 100.417 | 99.0415 | 95.7228 | 90.4574 | 87.5085
-23.4 | 100.821 | 99.2964 | 95.9248 | 90.5767 | 87.7105
-23.3 | 101.2251 | 99.4984 | 96.1111 | 90.7556 | 87.9063
-23.2 | 101.6218 | 99.7004 | 96.3004 | 91.0364 | 88.0073
-23.1 | 101.9166 | 99.9344 | 96.5024 | 91.1798 | 88.1166
-23 | 102.1188 | 100.1758 | 96.7044 | 91.3616 | 88.3094
-22.9 | 102.3312 | 100.3778 | 96.8532 | 91.4818 | 88.4208
-22.8 | 102.6114 | 100.5799 | 97.0085 | 91.6657 | 88.6114
-22.7 | 102.7248 | 100.741 | 97.2105 | 91.8677 | 88.7124
-22.6 | 102.9134 | 100.884 | 97.4125 | 91.9848 | 88.8134
-22.5 | 103.0289 | 101.0861 | 97.6145 | 92.0859 | 88.9126
-22.4 | 103.227 | 101.244 | 97.7583 | 92.2737 | 89.0022
-22.3 | 103.4021 | 101.3901 | 97.9186 | 92.4758 | 89.2021
-22.2 | 103.5065 | 101.5461 | 98.1206 | 92.6778 | 89.3032
-22.1 | 103.7042 | 101.6941 | 98.3226 | 92.8574 | 89.4042
-22 | 103.8105 | 101.8962 | 98.4623 | 92.9767 | 89.5053
-21.9 | 104.0125 | 102.0491 | 98.6109 | 93.0777 | 89.6125
-21.8 | 104.2073 | 102.2002 | 98.8002 | 93.1787 | 89.8073
-21.7 | 104.3083 | 102.3447 | 98.9511 | 93.2797 | 89.9083
-21.6 | 104.4093 | 102.4379 | 99.1042 | 93.3807 | 90.0093
-21.5 | 104.5103 | 102.5778 | 99.2531 | 93.4817 | 90.1103
-21.4 | 104.6226 | 102.7399 | 99.3541 | 93.5827 | 90.2113
-21.3 | 104.8123 | 102.8818 | 99.5103 | 93.7675 | 90.3123
-21.2 | 104.9133 | 103.0419 | 99.7123 | 93.8847 | 90.4133
-21.1 | 105.0143 | 103.1429 | 99.8572 | 93.9858 | 90.5143
-21 | 105.1307 | 103.2439 | 99.9582 | 94.0868 | 90.6134
-20.9 | 105.3143 | 103.3449 | 100.0593 | 94.1878 | 90.7021
-20.8 | 105.4063 | 103.446 | 100.2206 | 94.2888 | 90.8031
-20.7 | 105.6041 | 103.5939 | 100.3613 | 94.3786 | 90.9041
-20.6 | 105.7052 | 103.796 | 100.4623 | 94.4766 | 91.0052
-20.5 | 105.8062 | 103.9 | 100.5 | 94.5776 | 91.1062
-20.4 | 105.9072 | 103.95 | 100.55 | 94.6786 | 91.2
-20.3 | 106.0164 | 104.0447 | 100.651 | 94.7 | 91.2082
-20.2 | 106.2 | 104.1378 | 100.752 | 94.7806 | 91.3092
-20.1 | 106.2204 | 104.2388 | 100.853 | 94.8816 | 91.4102
-20 | 106.4112 | 104.3398 | 100.9541 | 94.9 | 91.5112
-19.9 | 106.5122 | 104.4408 | 101 | 95.0673 | 91.6122
-19.8 | 106.6 | 104.5418 | 101.0562 | 95.1846 | 91.7132
-19.7 | 106.6142 | 104.6428 | 101.1572 | 95.2 | 91.8142
-19.6 | 106.7153 | 104.7 | 101.2 | 95.2868 | 91.9134
-19.5 | 106.8 | 104.7448 | 101.2592 | 95.3 | 92.002
-19.4 | 106.803 | 104.8459 | 101.3 | 95.3888 | 92.1
-19.3 | 106.904 | 104.9 | 101.3612 | 95.4898 | 92.104
-19.2 | 107 | 104.9479 | 101.4622 | 95.5794 | 92.2101
-19.1 | 107.0061 | 105 | 101.5 | 95.6 | 92.4061
-19 | 107.1 | 105 | 101.5 | 95.6785 | 92.5
-18.9 | 107.1 | 105.0446 | 101.551 | 95.7 | 92.5081
-18.8 | 107.1 | 105.1 | 101.6 | 95.7806 | 92.6091
-18.7 | 107.1 | 105.1 | 101.6 | 95.8816 | 92.7101
-18.6 | 107.1 | 105.1 | 101.6 | 95.9 | 92.8112
-18.5 | 107.1122 | 105.1 | 101.6 | 95.9836 | 92.9122
-18.4 | 107.2 | 105.1 | 101.6 | 96 | 93.0132
-18.3 | 107.2 | 105.1 | 101.6 | 96.0857 | 93.1
-18.2 | 107.2 | 105.1 | 101.6 | 96.1867 | 93.1133
-18.1 | 107.2 | 105.1 | 101.6 | 96.2 | 93.2019
-18 | 107.1971 | 105.1 | 101.6 | 96.2887 | 93.3029
-17.9 | 107.1 | 105.1 | 101.6 | 96.3 | 93.4039
-17.8 | 107.1 | 105.1 | 101.6621 | 96.3794 | 93.505
-17.7 | 107.1 | 105.0511 | 101.7 | 96.4775 | 93.6061
-17.6 | 107.1 | 105 | 101.7 | 96.5 | 93.7071
-17.5 | 107.1 | 105 | 101.7 | 96.5795 | 93.8
-17.4 | 107.0909 | 105 | 101.7 | 96.6 | 93.8091
-17.3 | 107 | 104.9613 | 101.7 | 96.6815 | 93.9101
-17.2 | 106.9889 | 104.9 | 101.7 | 96.7 | 94.0111
-17.1 | 106.9 | 104.9 | 101.7 | 96.7835 | 94.1121
-17 | 106.8869 | 104.8583 | 101.7 | 96.8845 | 94.2131
-16.9 | 106.8 | 104.8 | 101.7 | 96.9 | 94.3141
-16.8 | 106.7848 | 104.8 | 101.7 | 96.9866 | 94.4
-16.7 | 106.7 | 104.7553 | 101.759 | 97 | 94.4019
-16.6 | 106.6971 | 104.7 | 101.8 | 97.0886 | 94.5029
-16.5 | 106.5961 | 104.7 | 101.8 | 97.1 | 94.6039
-16.4 | 106.5 | 104.6522 | 101.7379 | 97.1 | 94.7
-16.3 | 106.5 | 104.6 | 101.7 | 97.1774 | 94.706
-16.2 | 106.493 | 104.5502 | 101.7 | 97.2 | 94.807
-16.1 | 106.392 | 104.5 | 101.7 | 97.2794 | 94.908
-16 | 106.3 | 104.5 | 101.7 | 97.3804 | 95.009
-15.9 | 106.29 | 104.5 | 101.7 | 97.4814 | 95.11
-15.8 | 106.2 | 104.4604 | 101.7538 | 97.5 | 95.211
-15.7 | 106.188 | 104.4 | 101.8 | 97.5834 | 95.3
-15.6 | 106.1 | 104.4 | 101.8 | 97.6 | 95.313
-15.5 | 106.086 | 104.3574 | 101.8 | 97.6855 | 95.414
-15.4 | 106 | 104.3 | 101.8 | 97.7 | 95.5132
-15.3 | 105.9859 | 104.3 | 101.8 | 97.7875 | 95.6018
-15.2 | 105.9 | 104.3 | 101.8 | 97.8 | 95.7028
-15.1 | 105.9 | 104.2533 | 101.8 | 97.8896 | 95.8
-15 | 105.8952 | 104.2 | 101.8 | 97.9 | 95.8048
-14.9 | 105.8 | 104.2 | 101.8551 | 97.9773 | 95.9
-14.8 | 105.8 | 104.2 | 101.9 | 98.0783 | 95.9069
-14.7 | 105.7921 | 104.2 | 101.9 | 98.1 | 96.0079
-14.6 | 105.7 | 104.1625 | 101.9 | 98.1803 | 96.1
-14.5 | 105.7 | 104.1 | 101.9 | 98.2813 | 96.1099
-14.4 | 105.7 | 104.1 | 101.9 | 98.3 | 96.2109
-14.3 | 105.6881 | 104.1 | 101.9 | 98.3 | 96.3119
-14.2 | 105.6 | 104.1 | 101.9 | 98.3844 | 96.4
-14.1 | 105.586 | 104.1 | 101.9 | 98.4 | 96.414
-14 | 105.5 | 104.1 | 101.9 | 98.4865 | 96.5
-13.9 | 105.486 | 104.1 | 101.9 | 98.5 | 96.5017
-13.8 | 105.4 | 104.0543 | 101.9 | 98.5885 | 96.6
-13.7 | 105.4 | 104 | 101.9609 | 98.6 | 96.6037
-13.6 | 105.3953 | 104 | 102 | 98.6792 | 96.7047
-13.5 | 105.3058 | 104 | 102 | 98.7 | 96.8058
-13.4 | 105.4 | 104 | 102 | 98.7783 | 96.9068
-13.3 | 105.3922 | 104 | 102.0507 | 98.8793 | 97
-13.2 | 105.3 | 104 | 102.1 | 98.9 | 97.0089
-13.1 | 105.3 | 104 | 102.1 | 98.9 | 97.1
-13 | 105.3 | 104 | 102.1 | 98.9823 | 97.1109
-12.9 | 105.3 | 104 | 102.1 | 99 | 97.2119
-12.8 | 105.2871 | 103.9585 | 102.1 | 99.0843 | 97.3129
-12.7 | 105.2 | 103.9 | 102.1568 | 99.1 | 97.4139
-12.6 | 105.2 | 103.9 | 102.2 | 99.1 | 97.5
-12.5 | 105.2 | 103.9 | 102.2 | 99.1874 | 97.5
-12.4 | 105.2 | 103.9 | 102.2 | 99.2 | 97.5027
-12.3 | 105.1963 | 103.9 | 102.2 | 99.2894 | 97.6037
-12.2 | 105.1 | 103.9 | 102.2 | 99.3 | 97.7
-12.1 | 105.0942 | 103.9 | 102.2 | 99.3 | 97.7058
-12 | 105 | 103.9 | 102.2 | 99.3782 | 97.8068
-11.9 | 105 | 103.9 | 102.2 | 99.4 | 97.9
-11.8 | 105 | 103.9 | 102.2 | 99.4 | 97.9088
-11.7 | 105 | 103.9 | 102.2 | 99.4812 | 98
-11.6 | 104.9892 | 103.8606 | 102.2 | 99.5822 | 98
-11.5 | 104.9 | 103.8 | 102.2 | 99.6 | 98.0118
-11.4 | 104.9 | 103.8 | 102.2 | 99.6 | 98.1
-11.3 | 104.9 | 103.8 | 102.2 | 99.6853 | 98.1138
-11.2 | 104.9 | 103.8 | 102.2 | 99.7 | 98.213
-11.1 | 104.8861 | 103.8 | 102.2 | 99.7 | 98.3
-11 | 104.8 | 103.8 | 102.2 | 99.7883 | 98.3
-10.9 | 104.8 | 103.8 | 102.2 | 99.8 | 98.3036
-10.8 | 104.8 | 103.8 | 102.2618 | 99.8 | 98.4
-10.7 | 104.7943 | 103.8 | 102.2451 | 99.8 | 98.4057
-10.6 | 104.7 | 103.7505 | 102.2 | 99.8781 | 98.5067
-10.5 | 104.7 | 103.7 | 102.2 | 99.9 | 98.6
-10.4 | 104.6913 | 103.7 | 102.2 | 99.9801 | 98.6087
-10.3 | 104.6 | 103.7 | 102.2 | 100 | 98.7097
-10.2 | 104.6 | 103.7 | 102.2 | 100 | 98.8
-10.1 | 104.6 | 103.6597 | 102.2 | 100.0831 | 98.8117
-10 | 104.5873 | 103.6 | 102.2 | 100.1 | 98.9
-9.9 | 104.5 | 103.6 | 102.2 | 100.1 | 98.9138
-9.8 | 104.4852 | 103.6 | 102.2 | 100.1 | 99
-9.7 | 104.4 | 103.5556 | 102.2 | 100.1872 | 99.0015
-9.6 | 104.4 | 103.5 | 102.2 | 100.2 | 99.1
-9.5 | 104.3965 | 103.5 | 102.2 | 100.2 | 99.1
-9.4 | 104.3 | 103.5 | 102.2 | 100.2 | 99.1046
-9.3 | 104.3 | 103.5 | 102.2 | 100.277 | 99.2
-9.2 | 104.3 | 103.4505 | 102.2 | 100.3 | 99.2066
-9.1 | 104.3 | 103.4 | 102.2 | 100.379 | 99.3
-9 | 104.2914 | 103.4 | 102.2 | 100.4 | 99.3086
-8.9 | 104.2 | 103.4 | 102.2 | 100.4 | 99.4
-8.8 | 104.1893 | 103.3608 | 102.2 | 100.4 | 99.4
-8.7 | 104.1 | 103.3 | 102.2 | 100.4831 | 99.4
-8.6 | 104.1 | 103.3 | 102.2 | 100.5 | 99.4127
-8.5 | 104.0863 | 103.3 | 102.2 | 100.5 | 99.5
-8.4 | 104 | 103.3 | 102.2 | 100.5 | 99.5129
-8.3 | 104 | 103.3 | 102.2 | 100.5872 | 99.6014
-8.2 | 104 | 103.3 | 102.2 | 100.6 | 99.7
-8.1 | 103.9965 | 103.3 | 102.2 | 100.6 | 99.7035
-8 | 103.9 | 103.3 | 102.2 | 100.6 | 99.8
-7.9 | 103.9 | 103.2516 | 102.2 | 100.6 | 99.8
-7.8 | 103.8935 | 103.2 | 102.2 | 100.6 | 99.8
-7.7 | 103.8 | 103.2 | 102.2 | 100.679 | 99.8
-7.6 | 103.8 | 103.155 | 102.2 | 100.7 | 99.8086
-7.5 | 103.7904 | 103.1 | 102.1476 | 100.7 | 99.9
-7.4 | 103.7 | 103.0608 | 102.1 | 100.7 | 99.9
-7.3 | 103.7 | 103 | 102.1 | 100.7 | 99.9
-7.2 | 103.6874 | 103 | 102.1 | 100.7 | 99.9126
-7.1 | 103.6 | 102.9578 | 102.1 | 100.7 | 100
-7 | 103.5854 | 102.9 | 102.1 | 100.7 | 100
-6.9 | 103.4863 | 102.9 | 102.1 | 100.7 | 100.0014
-6.8 | 103.4 | 102.8548 | 102.1 | 100.7 | 100.1
-6.7 | 103.3966 | 102.8 | 102.1 | 100.7891 | 100.1
-6.6 | 103.3 | 102.8 | 102.1 | 100.8 | 100.1
-6.5 | 103.3 | 102.8 | 102.1 | 100.8 | 100.1
-6.4 | 103.3 | 102.8 | 102.0444 | 100.8 | 100.1
-6.3 | 103.2925 | 102.8 | 102 | 100.8 | 100.1
-6.2 | 103.2 | 102.7551 | 102 | 100.8 | 100.1
-6.1 | 103.1905 | 102.7 | 102 | 100.8 | 100.1095
-6 | 103.1 | 102.6609 | 102 | 100.8 | 100.2
-5.9 | 103.0885 | 102.6 | 102 | 100.8 | 100.2
-5.8 | 103 | 102.6 | 101.9446 | 100.8 | 100.2
-5.7 | 103 | 102.5579 | 101.9 | 100.8 | 100.2
-5.6 | 102.9854 | 102.5 | 101.8426 | 100.8 | 100.2
-5.5 | 102.8864 | 102.4558 | 101.8 | 100.8 | 100.2013
-5.4 | 102.8 | 102.4 | 101.7405 | 100.8 | 100.3
-5.3 | 102.7967 | 102.4 | 101.7 | 100.8 | 100.3
-5.2 | 102.7 | 102.4 | 101.7 | 100.8 | 100.3
-5.1 | 102.6946 | 102.3518 | 101.7 | 100.8 | 100.3
-5 | 102.6 | 102.3 | 101.7 | 100.7222 | 100.3064
-4.9 | 102.5926 | 102.2497 | 101.6497 | 100.7 | 100.4
-4.8 | 102.5 | 102.1551 | 101.6 | 100.7 | 100.4
-4.7 | 102.4906 | 102.1 | 101.5477 | 100.7 | 100.4
-4.6 | 102.3896 | 102.061 | 101.5 | 100.7 | 100.4
-4.5 | 102.2885 | 102 | 101.4457 | 100.7 | 100.4
-4.4 | 102.1875 | 101.959 | 101.4 | 100.7 | 100.4
-4.3 | 102.1 | 101.9 | 101.3437 | 100.7 | 100.4
-4.2 | 102.0855 | 101.8569 | 101.3 | 100.7 | 100.4
-4.1 | 102 | 101.8 | 101.2416 | 100.7 | 100.4
-4 | 101.9978 | 101.7549 | 101.2 | 100.7 | 100.4
-3.9 | 101.8967 | 101.6539 | 101.2 | 100.611 | 100.4
-3.8 | 101.7957 | 101.5529 | 101.1386 | 100.6 | 100.4
-3.7 | 101.6947 | 101.5 | 101.1 | 100.6 | 100.4
-3.6 | 101.5937 | 101.4508 | 101.1 | 100.6 | 100.4
-3.5 | 101.5 | 101.4 | 101.0498 | 100.6 | 100.4
-3.4 | 101.4917 | 101.3552 | 101 | 100.5202 | 100.4
-3.3 | 101.4 | 101.3 | 101 | 100.5 | 100.4
-3.2 | 101.3896 | 101.2611 | 100.9468 | 100.5 | 100.4
-3.1 | 101.3 | 101.2 | 100.9 | 100.5 | 100.3886
-3 | 101.2876 | 101.159 | 100.8447 | 100.5 | 100.3
-2.9 | 101.2 | 101.1 | 100.8 | 100.5 | 100.3
-2.8 | 101.1856 | 101.057 | 100.8 | 100.4141 | 100.3
-2.7 | 101.0865 | 101 | 100.8 | 100.4 | 100.3
-2.6 | 100.9978 | 100.955 | 100.7407 | 100.4 | 100.3
-2.5 | 100.9 | 100.9 | 100.7 | 100.4 | 100.3
-2.4 | 100.8958 | 100.8529 | 100.6387 | 100.4 | 100.3
-2.3 | 100.8 | 100.7519 | 100.6 | 100.4 | 100.2948
-2.2 | 100.7938 | 100.7 | 100.5445 | 100.3223 | 100.2
-2.1 | 100.7 | 100.7 | 100.5 | 100.3 | 100.2
-2 | 100.6917 | 100.6553 | 100.5 | 100.3 | 100.2
-1.9 | 100.5907 | 100.6 | 100.4479 | 100.2193 | 100.2
-1.8 | 100.5 | 100.5611 | 100.4 | 100.2 | 100.2
-1.7 | 100.4887 | 100.5 | 100.4 | 100.2 | 100.2
-1.6 | 100.3877 | 100.4591 | 100.3448 | 100.2 | 100.1877
-1.5 | 100.3 | 100.4 | 100.3 | 100.2 | 100.1
-1.4 | 100.2856 | 100.3571 | 100.2428 | 100.2 | 100.1
-1.3 | 100.2 | 100.3 | 100.1418 | 100.1132 | 100.1
-1.2 | 100.2 | 100.255 | 100.1 | 100.1 | 100.1
-1.1 | 100.2 | 100.2 | 100.1 | 100.1 | 100.1
-1 | 100.1959 | 100.2 | 100.1 | 100.0214 | 100.1
-0.9 | 100.1 | 100.2 | 100.0377 | 100 | 100.1
-0.8 | 100.0938 | 100.151 | 100 | 100 | 100.1
-0.7 | 100 | 100.1 | 100 | 100 | 100.1
-0.6 | 100 | 100.1 | 100 | 100 | 100.1
-0.5 | 100 | 100.0622 | 100 | 100 | 100.0908
-0.4 | 100 | 100 | 100 | 100 | 100
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100 | 100
0.4 | 100 | 100 | 100 | 100 | 100.0041
0.5 | 100 | 100.0479 | 100 | 100 | 100.1
0.6 | 100 | 100.1 | 100 | 100 | 100.1
0.7 | 100.0071 | 100.1 | 100 | 100 | 100.1
0.8 | 100.1 | 100.1446 | 100 | 100 | 100.1
0.9 | 100.1091 | 100.2 | 100.052 | 100 | 100.1
1 | 100.2 | 100.2 | 100.1 | 100 | 100.1
1.1 | 100.2 | 100.2397 | 100.1 | 100.0826 | 100.1
1.2 | 100.2 | 100.3 | 100.1 | 100.1 | 100.1
1.3 | 100.2132 | 100.3 | 100.1561 | 100.1 | 100.1
1.4 | 100.3 | 100.3428 | 100.2 | 100.1856 | 100.1
1.5 | 100.3152 | 100.4 | 100.2581 | 100.2 | 100.1133
1.6 | 100.4 | 100.4448 | 100.3591 | 100.2 | 100.2
1.7 | 100.403 | 100.5 | 100.4 | 100.2 | 100.2
1.8 | 100.504 | 100.5468 | 100.4 | 100.2 | 100.2
1.9 | 100.6 | 100.6 | 100.4621 | 100.2794 | 100.2
2 | 100.606 | 100.6489 | 100.5 | 100.3 | 100.2
2.1 | 100.707 | 100.7 | 100.5 | 100.3 | 100.2
2.2 | 100.8 | 100.7 | 100.5 | 100.3 | 100.2081
2.3 | 100.8091 | 100.7376 | 100.5519 | 100.3805 | 100.3
2.4 | 100.9 | 100.8387 | 100.6529 | 100.4 | 100.3
2.5 | 100.9111 | 100.9 | 100.7 | 100.4 | 100.3
2.6 | 101.0121 | 100.9407 | 100.755 | 100.4 | 100.3
2.7 | 101.1 | 101 | 100.8 | 100.4 | 100.3
2.8 | 101.1141 | 101.0427 | 100.8 | 100.4 | 100.3
2.9 | 101.2152 | 101.1437 | 100.8 | 100.4866 | 100.3
3 | 101.3 | 101.2 | 100.859 | 100.5 | 100.3019
3.1 | 101.3 | 101.2458 | 100.9 | 100.5 | 100.4
3.2 | 101.3039 | 101.3 | 100.9611 | 100.5 | 100.4
3.3 | 101.4049 | 101.3 | 101 | 100.5 | 100.4
3.4 | 101.5 | 101.3488 | 101 | 100.5774 | 100.4
3.5 | 101.507 | 101.4 | 101.0498 | 100.6 | 100.4
3.6 | 101.608 | 101.4445 | 101.1 | 100.6 | 100.4
3.7 | 101.709 | 101.5 | 101.1 | 100.6 | 100.4
3.8 | 101.81 | 101.5386 | 101.1529 | 100.6 | 100.4
3.9 | 101.9 | 101.6396 | 101.2 | 100.6 | 100.4
4 | 101.912 | 101.7406 | 101.2 | 100.6835 | 100.4
4.1 | 102 | 101.8 | 101.2559 | 100.7 | 100.4
4.2 | 102.0141 | 101.8426 | 101.3 | 100.7 | 100.4
4.3 | 102.1151 | 101.9 | 101.3579 | 100.7 | 100.4
4.4 | 102.2141 | 101.9447 | 101.4 | 100.7 | 100.4
4.5 | 102.3028 | 102 | 101.46 | 100.7 | 100.4
4.6 | 102.4 | 102.0467 | 101.5 | 100.7 | 100.4
4.7 | 102.4049 | 102.1477 | 101.562 | 100.7 | 100.4
4.8 | 102.5059 | 102.2 | 101.6 | 100.7 | 100.4
4.9 | 102.6 | 102.2497 | 101.6497 | 100.7 | 100.3931
5 | 102.6079 | 102.3444 | 101.7 | 100.7 | 100.3
5.1 | 102.7 | 102.4 | 101.7 | 100.7804 | 100.3
5.2 | 102.7099 | 102.4 | 101.7 | 100.8 | 100.3
5.3 | 102.8 | 102.4 | 101.7 | 100.8 | 100.3
5.4 | 102.812 | 102.4 | 101.7548 | 100.8 | 100.288
5.5 | 102.9 | 102.4416 | 101.8 | 100.8 | 100.2
5.6 | 102.914 | 102.5 | 101.8569 | 100.8 | 100.2
5.7 | 103 | 102.5436 | 101.9 | 100.8 | 100.2
5.8 | 103 | 102.6 | 101.9 | 100.8 | 100.2
5.9 | 103.0028 | 102.6 | 101.9599 | 100.8 | 100.2
6 | 103.1038 | 102.6466 | 102 | 100.8 | 100.2
6.1 | 103.2 | 102.7 | 102 | 100.8 | 100.1952
6.2 | 103.2058 | 102.7487 | 102 | 100.8 | 100.1
6.3 | 103.3 | 102.8 | 102 | 100.8 | 100.1
6.4 | 103.3 | 102.8 | 102.0507 | 100.8 | 100.1
6.5 | 103.3 | 102.8 | 102.1 | 100.8 | 100.1
6.6 | 103.3099 | 102.8 | 102.1 | 100.7187 | 100.1
6.7 | 103.4 | 102.8 | 102.1 | 100.7823 | 100.1
6.8 | 103.4119 | 102.8405 | 102.1 | 100.7167 | 100.0881
6.9 | 103.5129 | 102.9 | 102.1 | 100.7 | 100
7 | 103.6 | 102.9425 | 102.1 | 100.7 | 100
7.1 | 103.6 | 103 | 102.1 | 100.7 | 99.9869
7.2 | 103.614 | 103 | 102.1 | 100.7 | 99.9
7.3 | 103.7 | 103 | 102.1 | 100.7 | 99.9
7.4 | 103.7037 | 103.0466 | 102.1 | 100.7 | 99.9
7.5 | 103.8 | 103.1 | 102.1619 | 100.7 | 99.8953
7.6 | 103.8 | 103.1486 | 102.2 | 100.7 | 99.8
7.7 | 103.8067 | 103.2 | 102.2 | 100.7 | 99.8
7.8 | 103.9 | 103.2 | 102.2 | 100.6208 | 99.8
7.9 | 103.9 | 103.2452 | 102.2 | 100.6 | 99.8
8 | 103.9098 | 103.3 | 102.2 | 100.6 | 99.7902
8.1 | 104 | 103.3 | 102.2 | 100.6 | 99.7
8.2 | 104 | 103.3 | 102.2 | 100.6 | 99.6882
8.3 | 104 | 103.3 | 102.2 | 100.6 | 99.6
8.4 | 104 | 103.3 | 102.2 | 100.5147 | 99.5862
8.5 | 104.0149 | 103.3 | 102.2 | 100.5 | 99.487
8.6 | 104.1 | 103.3 | 102.2 | 100.5 | 99.4
8.7 | 104.1026 | 103.3 | 102.2 | 100.5 | 99.4
8.8 | 104.2 | 103.3465 | 102.2 | 100.4107 | 99.4
8.9 | 104.2046 | 103.4 | 102.2 | 100.4 | 99.3954
9 | 104.3 | 103.4 | 102.2 | 100.4 | 99.3
9.1 | 104.3 | 103.4 | 102.2 | 100.4 | 99.3
9.2 | 104.3 | 103.4505 | 102.2 | 100.3209 | 99.2923
9.3 | 104.3 | 103.5 | 102.2 | 100.3 | 99.2
9.4 | 104.3097 | 103.5 | 102.2 | 100.2189 | 99.1903
9.5 | 104.4 | 103.5 | 102.2 | 100.2 | 99.1
9.6 | 104.4 | 103.5 | 102.2 | 100.2 | 99.0883
9.7 | 104.4128 | 103.5413 | 102.2 | 100.1158 | 99
9.8 | 104.5 | 103.6 | 102.2 | 100.1 | 98.9862
9.9 | 104.5148 | 103.6 | 102.2 | 100.1 | 98.9
10 | 104.6 | 103.6 | 102.2 | 100.1 | 98.8985
10.1 | 104.6 | 103.6454 | 102.2 | 100.1 | 98.8
10.2 | 104.6 | 103.7 | 102.2 | 100.0107 | 98.7965
10.3 | 104.6045 | 103.7 | 102.2 | 100 | 98.6955
10.4 | 104.7 | 103.7 | 102.2 | 100 | 98.6
10.5 | 104.7 | 103.7 | 102.2 | 99.9219 | 98.5934
10.6 | 104.7076 | 103.7505 | 102.2 | 99.9 | 98.4924
10.7 | 104.8 | 103.8 | 102.2515 | 99.8199 | 98.4
10.8 | 104.8 | 103.8 | 102.2475 | 99.8 | 98.3903
10.9 | 104.8 | 103.8 | 102.2 | 99.8 | 98.3
11 | 104.8117 | 103.8 | 102.2 | 99.7169 | 98.3
11.1 | 104.9 | 103.8 | 102.2 | 99.7 | 98.2873
11.2 | 104.9 | 103.8 | 102.2 | 99.7 | 98.1863
11.3 | 104.9 | 103.8 | 102.2 | 99.6138 | 98.1
11.4 | 104.9 | 103.8 | 102.2 | 99.6 | 98.0986
11.5 | 104.9024 | 103.8 | 102.2 | 99.6 | 98
11.6 | 105 | 103.8464 | 102.2 | 99.5108 | 98
11.7 | 105 | 103.9 | 102.2 | 99.5 | 97.9955
11.8 | 105 | 103.9 | 102.2 | 99.423 | 97.9
11.9 | 105 | 103.9 | 102.2 | 99.4 | 97.8934
12 | 105.0076 | 103.9 | 102.2 | 99.321 | 97.8
12.1 | 105.1 | 103.9 | 102.2 | 99.3 | 97.7914
12.2 | 105.1096 | 103.9 | 102.2 | 99.3 | 97.6904
12.3 | 105.2 | 103.9 | 102.2 | 99.218 | 97.5894
12.4 | 105.2 | 103.9 | 102.2 | 99.2 | 97.5
12.5 | 105.2 | 103.9 | 102.2 | 99.2 | 97.5
12.6 | 105.2 | 103.9 | 102.2 | 99.1149 | 97.4864
12.7 | 105.2 | 103.9 | 102.1425 | 99.1 | 97.3872
12.8 | 105.2137 | 103.9442 | 102.1 | 99.0129 | 97.2986
12.9 | 105.3 | 104 | 102.1 | 99 | 97.1976
13 | 105.3 | 104 | 102.1 | 98.9109 | 97.1
13.1 | 105.3 | 104 | 102.1 | 98.9 | 97.0956
13.2 | 105.3055 | 104 | 102.1 | 98.9 | 97
13.3 | 105.4 | 104 | 102.1 | 98.8221 | 96.9935
13.4 | 105.3925 | 104 | 102.0497 | 98.7211 | 96.8925
13.5 | 105.3085 | 104 | 102 | 98.7 | 96.7915
13.6 | 105.4 | 104 | 102 | 98.7 | 96.7
13.7 | 105.4 | 104 | 102 | 98.6181 | 96.6895
13.8 | 105.4115 | 104.0401 | 101.9457 | 98.6 | 96.6
13.9 | 105.5 | 104.1 | 101.9 | 98.5161 | 96.5875
14 | 105.5135 | 104.1 | 101.9 | 98.5 | 96.4865
14.1 | 105.6 | 104.1 | 101.9 | 98.414 | 96.4
14.2 | 105.6136 | 104.1 | 101.9 | 98.4 | 96.3987
14.3 | 105.7 | 104.1 | 101.9 | 98.3119 | 96.2977
14.4 | 105.7 | 104.1 | 101.9 | 98.3 | 96.1967
14.5 | 105.7 | 104.1 | 101.9 | 98.2212 | 96.1
14.6 | 105.7 | 104.1482 | 101.9 | 98.2 | 96.0946
14.7 | 105.7064 | 104.2 | 101.9 | 98.1222 | 95.9936
14.8 | 105.8 | 104.2 | 101.9 | 98.1 | 95.9
14.9 | 105.8 | 104.2 | 101.9 | 98.0202 | 95.8916
15 | 105.8094 | 104.2 | 101.8477 | 97.9191 | 95.8
15.1 | 105.9 | 104.239 | 101.8 | 97.9 | 95.7896
15.2 | 105.9114 | 104.3 | 101.8 | 97.8171 | 95.6886
15.3 | 106 | 104.3 | 101.8 | 97.8 | 95.5875
15.4 | 106.0135 | 104.3 | 101.8 | 97.7151 | 95.4865
15.5 | 106.1 | 104.3431 | 101.8 | 97.614 | 95.4
15.6 | 106.1136 | 104.4 | 101.8 | 97.6 | 95.3988
15.7 | 106.2 | 104.4 | 101.8 | 97.512 | 95.2978
15.8 | 106.2032 | 104.4462 | 101.7396 | 97.5 | 95.1968
15.9 | 106.3 | 104.5 | 101.7 | 97.5 | 95.0958
16 | 106.3053 | 104.5 | 101.7 | 97.4232 | 94.9947
16.1 | 106.4 | 104.5492 | 101.7 | 97.3222 | 94.8937
16.2 | 106.4074 | 104.6 | 101.7 | 97.2212 | 94.7926
16.3 | 106.5 | 104.6448 | 101.7 | 97.2 | 94.7
16.4 | 106.5094 | 104.7 | 101.7522 | 97.1192 | 94.6906
16.5 | 106.6104 | 104.7 | 101.8 | 97.1 | 94.5896
16.6 | 106.7 | 104.7 | 101.8 | 97.0172 | 94.5
16.7 | 106.7124 | 104.741 | 101.8 | 97 | 94.4876
16.8 | 106.8 | 104.8 | 101.7437 | 97 | 94.3866
16.9 | 106.8144 | 104.8 | 101.7 | 96.9141 | 94.2874
17 | 106.9 | 104.844 | 101.7 | 96.8131 | 94.1988
17.1 | 106.9022 | 104.9 | 101.7 | 96.8 | 94.0978
17.2 | 107 | 104.9 | 101.7 | 96.7111 | 93.9968
17.3 | 107.0042 | 104.9471 | 101.7 | 96.7 | 93.8958
17.4 | 107.1 | 105 | 101.7 | 96.6233 | 93.8
17.5 | 107.1 | 105 | 101.7 | 96.6 | 93.7937
17.6 | 107.1 | 105 | 101.7 | 96.5213 | 93.6927
17.7 | 107.1 | 105.0447 | 101.7 | 96.4203 | 93.5917
17.8 | 107.1 | 105.1 | 101.6479 | 96.4 | 93.4907
17.9 | 107.1103 | 105.1 | 101.6 | 96.3183 | 93.3897
18 | 107.2 | 105.1 | 101.6 | 96.3 | 93.2887
18.1 | 107.2 | 105.1 | 101.6 | 96.2163 | 93.2
18.2 | 107.2 | 105.1 | 101.6 | 96.1153 | 93.1867
18.3 | 107.2 | 105.1 | 101.6 | 96.1 | 93.0857
18.4 | 107.1846 | 105.1 | 101.6 | 96.0132 | 92.9866
18.5 | 107.1 | 105.1 | 101.6 | 95.9122 | 92.8979
18.6 | 107.1 | 105.1 | 101.6 | 95.9 | 92.7969
18.7 | 107.1 | 105.1 | 101.6 | 95.9 | 92.6959
18.8 | 107.1 | 105.1 | 101.6 | 95.8234 | 92.5948
18.9 | 107.1 | 105.051 | 101.5446 | 95.7224 | 92.5
19 | 107.0928 | 105 | 101.5 | 95.7 | 92.4928
19.1 | 107 | 105 | 101.5 | 95.6204 | 92.3918
19.2 | 106.9908 | 104.9622 | 101.4479 | 95.5194 | 92.2908
19.3 | 106.8898 | 104.9 | 101.4 | 95.5 | 92.1898
19.4 | 106.8 | 104.8602 | 101.3459 | 95.4174 | 92.0888
19.5 | 106.7878 | 104.7592 | 101.2448 | 95.3163 | 91.9878
19.6 | 106.7 | 104.7 | 101.2 | 95.3 | 91.8868
19.7 | 106.6857 | 104.6572 | 101.1428 | 95.2142 | 91.7857
19.8 | 106.5846 | 104.5562 | 101.0418 | 95.1132 | 91.6866
19.9 | 106.4857 | 104.4552 | 101 | 95.0122 | 91.598
20 | 106.397 | 104.3541 | 100.9398 | 94.9112 | 91.497
20.1 | 106.296 | 104.253 | 100.8388 | 94.8215 | 91.396
20.2 | 106.1949 | 104.152 | 100.7378 | 94.8 | 91.3
20.3 | 106.0939 | 104.051 | 100.6447 | 94.7224 | 91.2939
20.4 | 105.9859 | 103.95 | 100.55 | 94.7 | 91.1929
20.5 | 105.7919 | 103.8554 | 100.5 | 94.6204 | 91.0919
20.6 | 105.6909 | 103.7623 | 100.448 | 94.5194 | 90.9909
20.7 | 105.5898 | 103.6226 | 100.347 | 94.4184 | 90.8898
20.8 | 105.4776 | 103.4603 | 100.246 | 94.2347 | 90.7888
20.9 | 105.2878 | 103.3593 | 100.0899 | 94.1164 | 90.6878
21 | 105.1868 | 103.2582 | 99.9439 | 94.1 | 90.5868
21.1 | 105.0715 | 103.1572 | 99.8429 | 93.9287 | 90.4858
21.2 | 104.9 | 103.0123 | 99.6838 | 93.8133 | 90.3867
21.3 | 104.8716 | 102.8552 | 99.4818 | 93.7123 | 90.2981
21.4 | 104.6941 | 102.7541 | 99.3399 | 93.6113 | 90.1971
21.5 | 104.496 | 102.6063 | 99.2389 | 93.5103 | 90.096
21.6 | 104.3949 | 102.4521 | 99.0758 | 93.4206 | 89.9949
21.7 | 104.2939 | 102.3511 | 98.9447 | 93.3225 | 89.8939
21.8 | 104.1859 | 102.2002 | 98.8002 | 93.143 | 89.7929
21.9 | 103.9919 | 102.0109 | 98.6491 | 93.0205 | 89.6838
22 | 103.8818 | 101.8623 | 98.4962 | 92.9195 | 89.4909
22.1 | 103.6798 | 101.7226 | 98.2941 | 92.8185 | 89.3899
22.2 | 103.4889 | 101.5603 | 98.0921 | 92.6349 | 89.2889
22.3 | 103.3879 | 101.4186 | 97.8901 | 92.5165 | 89.1758
22.4 | 103.2737 | 101.2166 | 97.744 | 92.2464 | 88.9869
22.5 | 103.0717 | 101.0573 | 97.5861 | 92.1144 | 88.8859
22.6 | 102.8697 | 100.9125 | 97.384 | 92.0134 | 88.7867
22.7 | 102.6859 | 100.7105 | 97.241 | 91.9124 | 88.6981
22.8 | 102.5941 | 100.5085 | 97.0198 | 91.6342 | 88.5971
22.9 | 102.3882 | 100.3532 | 96.8389 | 91.5104 | 88.4921
23 | 102.0901 | 100.2044 | 96.6758 | 91.3414 | 88.2951
23.1 | 101.8821 | 99.9536 | 96.4896 | 91.2226 | 88.1881
23.2 | 101.5791 | 99.7004 | 96.3004 | 91.0432 | 87.993
23.3 | 101.2681 | 99.5111 | 96.0984 | 90.8412 | 87.884
23.4 | 100.864 | 99.3248 | 95.8964 | 90.6392 | 87.682
23.5 | 100.45 | 99.0842 | 95.6943 | 90.4372 | 87.49
23.6 | 99.9449 | 98.7812 | 95.4385 | 90.2352 | 87.389
23.7 | 99.4038 | 98.4782 | 95.1902 | 90.0331 | 87.276
23.8 | 98.6218 | 98.1752 | 94.9881 | 89.9156 | 87.0739
23.9 | 97.9877 | 97.8295 | 94.7861 | 89.6436 | 86.886
24 | 97.1645 | 97.4255 | 94.584 | 89.3403 | 86.7737
24.1 | 96.2735 | 96.9768 | 94.323 | 89.2124 | 86.5982
24.2 | 95.4717 | 96.4261 | 94.08 | 88.9342 | 86.4943
24.3 | 94.4655 | 95.82 | 93.817 | 88.8104 | 86.2923
24.4 | 93.5418 | 95.1663 | 93.5139 | 88.6415 | 86.0952
24.5 | 92.3414 | 94.4588 | 93.2345 | 88.3682 | 85.9883
24.6 | 91.3244 | 93.7016 | 92.9506 | 88.0652 | 85.7863
24.7 | 90.2055 | 92.9001 | 92.6984 | 87.8414 | 85.5842
24.8 | 88.9933 | 92.0253 | 92.3927 | 87.7196 | 85.3822
24.9 | 87.7416 | 91.0152 | 91.9887 | 87.5372 | 85.1802
25 | 86.0925 | 90.0051 | 91.5846 | 87.2527 | 84.9891
25.1 | 84.0634 | 88.9544 | 91.1806 | 86.9497 | 84.876
25.2 | 81.0918 | 87.8433 | 90.7766 | 86.7311 | 84.6739
25.3 | 76.9699 | 86.7317 | 90.3157 | 86.5291 | 84.4719
25.4 | 71.5217 | 85.5764 | 89.7527 | 86.2406 | 84.2737
25.5 | 64.9746 | 84.3196 | 89.0878 | 85.9376 | 84.0948
25.6 | 58.4927 | 82.8695 | 88.4406 | 85.8115 | 83.7973
25.7 | 50.8197 | 81.0667 | 87.7127 | 85.5315 | 83.6925
25.8 | 43.2772 | 78.5697 | 86.9666 | 85.3416 | 83.4858
25.9 | 36.46 | 75.0038 | 86.2593 | 85.0682 | 83.1885
26 | 30.5566 | 70.3158 | 85.4527 | 84.8434 | 82.9863
26.1 | 25.5927 | 64.6576 | 84.5436 | 84.6414 | 82.7842
26.2 | 21.7333 | 58.6669 | 83.6345 | 84.2788 | 82.5733
26.3 | 18.7723 | 52.3445 | 82.7255 | 84.0374 | 82.2901
26.4 | 16.4927 | 45.4538 | 81.7089 | 83.753 | 82.1673
26.5 | 14.6451 | 38.9292 | 80.7073 | 83.45 | 81.8762
26.6 | 13.3578 | 33.0997 | 79.6867 | 83.147 | 81.6741
26.7 | 12.2885 | 28.1714 | 78.5188 | 82.9293 | 81.4582
26.8 | 11.5103 | 24.1763 | 77.2489 | 82.6409 | 81.1738
26.9 | 10.9162 | 20.9973 | 75.7594 | 82.2505 | 80.9966
27 | 10.3891 | 18.4978 | 73.7844 | 81.9348 | 80.7918
27.1 | 9.9888 | 16.5084 | 71.1366 | 81.5424 | 80.4925
27.2 | 9.6858 | 14.9815 | 67.5891 | 81.1835 | 80.2905
27.3 | 9.3827 | 13.8141 | 63.5152 | 80.8685 | 80.0885
27.4 | 9.0865 | 12.8536 | 58.7218 | 80.5655 | 79.8797
27.5 | 8.8844 | 12.0964 | 53.0648 | 80.104 | 79.5844
27.6 | 8.6824 | 11.5289 | 47.2063 | 79.599 | 79.3736
27.7 | 8.4804 | 11.0465 | 41.5056 | 79.1752 | 79.0804
27.8 | 8.2784 | 10.6424 | 36.2182 | 78.6067 | 78.8676
27.9 | 8.0764 | 10.2788 | 31.5502 | 78.0006 | 78.5764
28 | 7.8872 | 9.9758 | 27.652 | 77.4788 | 78.3615
28.1 | 7.7723 | 9.7152 | 24.3983 | 76.8885 | 78.0723
28.2 | 7.5852 | 9.4697 | 21.7711 | 76.1962 | 77.861
28.3 | 7.4722 | 9.2111 | 19.683 | 75.4891 | 77.5968
28.4 | 7.2974 | 9.0091 | 18.003 | 74.8698 | 77.3947
28.5 | 7.1927 | 8.8071 | 16.6095 | 74.0848 | 77.1891
28.6 | 6.9954 | 8.6051 | 15.4818 | 73.3675 | 76.8861
28.7 | 6.8887 | 8.403 | 14.6604 | 72.6605 | 76.583
28.8 | 6.6933 | 8.2505 | 14.003 | 71.7973 | 76.28
28.9 | 6.5846 | 8.099 | 13.397 | 71.0464 | 75.977
29 | 6.3913 | 7.9098 | 12.8937 | 70.1791 | 75.6739
29.1 | 6.2903 | 7.7617 | 12.4895 | 69.27 | 75.3806
29.2 | 6.1786 | 7.6214 | 12.1391 | 68.3609 | 75.1572
29.3 | 5.9883 | 7.4194 | 11.8361 | 67.2014 | 74.7648
29.4 | 5.8873 | 7.2587 | 11.533 | 65.9048 | 74.4491
29.5 | 5.7862 | 7.1154 | 11.23 | 64.0802 | 74.045
29.6 | 5.6852 | 6.9567 | 10.9846 | 61.9022 | 73.661
29.7 | 5.5861 | 6.8557 | 10.7239 | 59.0691 | 73.3939
29.8 | 5.4949 | 6.7092 | 10.4806 | 55.6101 | 72.9899
29.9 | 5.2965 | 6.5071 | 10.2786 | 51.539 | 72.5859
30 | 5.2 | 6.3525 | 10.0766 | 47.7041 | 72.1818
30.1 | 5.1889 | 6.2515 | 9.8353 | 43.4099 | 71.7778
30.2 | 4.9934 | 6.1505 | 9.6505 | 38.9644 | 71.3606
30.3 | 4.8924 | 6.0495 | 9.499 | 34.9945 | 70.7621
30.4 | 4.7913 | 5.9549 | 9.297 | 31.3965 | 70.2479
30.5 | 4.6903 | 5.8236 | 9.0949 | 28.1856 | 69.6515
30.6 | 4.6 | 5.6608 | 8.8929 | 25.547 | 69.1358
30.7 | 4.5883 | 5.5598 | 8.7455 | 23.3542 | 68.5414
30.8 | 4.4873 | 5.4587 | 8.5889 | 21.4855 | 68.0109
30.9 | 4.3863 | 5.3577 | 8.3869 | 20.0079 | 67.3176
31 | 4.2853 | 5.2567 | 8.1848 | 18.7661 | 66.7097
31.1 | 4.1862 | 5.1557 | 8.0414 | 17.7283 | 66.0915
31.2 | 4.0976 | 5.0546 | 7.8808 | 16.8945 | 65.4854
31.3 | 4 | 4.9536 | 7.7394 | 16.2648 | 64.8717
31.4 | 3.9955 | 4.8526 | 7.5768 | 15.7263 | 64.0727
31.5 | 3.8944 | 4.7516 | 7.4374 | 15.2152 | 63.4556
31.6 | 3.7934 | 4.6506 | 7.3443 | 14.7881 | 62.6475
31.7 | 3.7 | 4.5496 | 7.1992 | 14.384 | 61.8242
31.8 | 3.6914 | 4.5 | 7.0486 | 13.98 | 60.797
31.9 | 3.5904 | 4.4618 | 6.9476 | 13.738 | 59.5657
32 | 3.4894 | 4.3608 | 6.7931 | 13.4539 | 58.1091
32.1 | 3.4 | 4.2598 | 6.6456 | 13.1509 | 56.2328
32.2 | 3.3874 | 4.1588 | 6.5445 | 12.9319 | 53.8591
32.3 | 3.2864 | 4.1 | 6.4435 | 12.6448 | 51.0636
32.4 | 3.2 | 4.0568 | 6.3425 | 12.4279 | 47.8641
32.5 | 3.1863 | 3.9558 | 6.2415 | 12.2259 | 44.8505
32.6 | 3.1 | 3.8547 | 6.1404 | 12.0238 | 41.3103
32.7 | 3.0966 | 3.8 | 6.0394 | 11.8218 | 37.5832
32.8 | 3 | 3.7527 | 5.8768 | 11.5634 | 34.1667
32.9 | 2.9945 | 3.6517 | 5.8 | 11.4231 | 31.1582
33 | 2.9 | 3.6 | 5.7443 | 11.2442 | 28.5448
33.1 | 2.8925 | 3.5497 | 5.6497 | 11.0422 | 26.158
33.2 | 2.8 | 3.4551 | 5.4974 | 10.8402 | 24.2558
33.3 | 2.7905 | 3.4 | 5.4 | 10.6382 | 22.5766
33.4 | 2.7 | 3.3609 | 5.3467 | 10.5181 | 21.2739
33.5 | 2.6885 | 3.3 | 5.2456 | 10.3341 | 20.0848
33.6 | 2.6 | 3.2589 | 5.1445 | 10.1321 | 19.1123
33.7 | 2.5865 | 3.2 | 5.0435 | 10.0151 | 18.3917
33.8 | 2.5 | 3.1569 | 4.9425 | 9.8279 | 17.6236
33.9 | 2.5 | 3.1 | 4.8415 | 9.6259 | 17.0934
34 | 2.4977 | 3.0548 | 4.8 | 9.5119 | 16.5907
34.1 | 2.4 | 3 | 4.7395 | 9.3218 | 16.19
34.2 | 2.3957 | 2.9528 | 4.6385 | 9.2212 | 15.8826
34.3 | 2.3 | 2.9 | 4.5375 | 9.0465 | 15.4839
34.4 | 2.3 | 2.8507 | 4.4444 | 8.9222 | 15.1809
34.5 | 2.2926 | 2.8 | 4.4 | 8.7424 | 14.8853
34.6 | 2.2 | 2.8 | 4.3487 | 8.6201 | 14.6748
34.7 | 2.2 | 2.762 | 4.3 | 8.5191 | 14.3812
34.8 | 2.2 | 2.7 | 4.2467 | 8.4181 | 14.1792
34.9 | 2.1885 | 2.66 | 4.1457 | 8.2341 | 13.9655
35 | 2.1 | 2.6 | 4.1 | 8.1161 | 13.6749
35.1 | 2.1 | 2.6 | 4.0436 | 8.0151 | 13.4729
35.2 | 2.1 | 2.5569 | 3.9426 | 7.914 | 13.2746
35.3 | 2.0864 | 2.5 | 3.9 | 7.813 | 13.0976
35.4 | 2 | 2.5 | 3.8406 | 7.712 | 12.8956
35.5 | 2 | 2.4538 | 3.8 | 7.611 | 12.6968
35.6 | 2 | 2.4 | 3.7386 | 7.4424 | 12.5915
35.7 | 1.9946 | 2.4 | 3.7 | 7.3232 | 12.3946
35.8 | 1.9 | 2.3508 | 3.6445 | 7.2222 | 12.2873
35.9 | 1.9 | 2.3 | 3.6 | 7.1212 | 12.0779
36 |  | 2.3 | 3.5488 | 7.0202 | 11.7832
36.1 |  | 2.3 | 3.5 | 6.9192 | 11.5906
36.2 |  | 2.261 | 3.4468 | 6.8182 | 11.4896
36.3 |  | 2.2 | 3.4 | 6.7172 | 11.3772
36.4 |  | 2.2 | 3.3447 | 6.6162 | 11.1876
36.5 |  | 2.158 | 3.3 | 6.5152 | 11.0731
36.6 |  | 2.1 | 3.2427 | 6.4141 | 10.8874
36.7 |  |  | 3.2 | 6.3131 | 10.7976
36.8 |  |  | 3.1407 | 6.3 | 10.5978
36.9 |  |  | 3.1 | 6.2111 | 10.4968
37 |  |  | 3.1 | 6.1213 | 10.3915
37.1 |  |  | 3.0376 | 6.0233 | 10.1947
37.2 |  |  | 3 | 5.9223 | 10.0875
37.3 |  |  | 3 | 5.8213 | 9.8927
37.4 |  |  | 2.9489 | 5.8 | 9.7917
37.5 |  |  | 2.9 | 5.7193 | 9.6907
37.6 |  |  | 2.8469 | 5.6183 | 9.5897
37.7 |  |  | 2.8 | 5.5173 | 9.4774
37.8 |  |  | 2.8 | 5.5 | 9.2877
37.9 |  |  | 2.8 | 5.4153 | 9.1867
38 |  |  | 2.7427 | 5.3142 | 9.0857
38.1 |  |  | 2.7 | 5.3 | 8.9866
38.2 |  |  |  | 5.2122 | 8.8958
38.3 |  |  |  | 5.1111 | 8.6969
38.4 |  |  |  | 5.1 | 8.6
38.5 |  |  |  | 5.0234 | 8.5948
38.6 |  |  |  | 4.9224 | 8.4938
38.7 |  |  |  | 4.9 | 8.3857
38.8 |  |  |  | 4.8204 | 8.2
38.9 |  |  |  | 4.8 | 8.1908
39 |  |  |  | 4.7184 | 8.0898
39.1 |  |  |  | 4.7 | 7.9888
39.2 |  |  |  | 4.6163 | 7.8878
39.3 |  |  |  | 4.5153 | 7.7868
39.4 |  |  |  | 4.5 | 7.6857
39.5 |  |  |  | 4.4132 | 7.5866
39.6 |  |  |  | 4.4 | 7.498
39.7 |  |  |  | 4.3112 | 7.397
39.8 |  |  |  | 4.3 | 7.296
39.9 |  |  |  | 4.2235 | 7.2
40 |  |  |  | 4.2 | 7.1939
40.1 |  |  |  | 4.1214 | 7.0929
40.2 |  |  |  | 4.1 | 6.9919
40.3 |  |  |  | 4.1 | 6.9
40.4 |  |  |  | 4.0184 | 6.8898
40.5 |  |  |  | 4 | 6.7888
40.6 |  |  |  | 3.9164 | 6.6878
40.7 |  |  |  | 3.9 | 6.5868
40.8 |  |  |  | 3.9 | 6.5
40.9 |  |  |  | 3.8133 | 6.4867
41 |  |  |  |  | 6.4
41.1 |  |  |  |  | 6.397
41.2 |  |  |  |  | 6.296
41.3 |  |  |  |  | 6.2
41.4 |  |  |  |  | 6.1939
41.5 |  |  |  |  | 6.0929
41.6 |  |  |  |  | 5.9919
41.7 |  |  |  |  | 5.9
41.8 |  |  |  |  | 5.8899
41.9 |  |  |  |  | 5.8
42 |  |  |  |  | 5.7879
42.1 |  |  |  |  | 5.7
42.2 |  |  |  |  | 5.6859
42.3 |  |  |  |  | 5.6
42.4 |  |  |  |  | 5.5981
42.5 |  |  |  |  | 5.5
42.6 |  |  |  |  | 5.4961
42.7 |  |  |  |  | 5.4
42.8 |  |  |  |  | 5.394
42.9 |  |  |  |  | 5.3
43 |  |  |  |  | 5.292
43.1 |  |  |  |  | 5.2
43.2 |  |  |  |  | 5.19
43.3 |  |  |  |  | 5.1
43.4 |  |  |  |  | 5.088
43.5 |  |  |  |  | 5
43.6 |  |  |  |  | 5
43.7 |  |  |  |  | 4.9868

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.883 | 0.897 | 0.904 | 0.912 | 0.917 | 0.921 | 0.924 | 0.926 | 0.926
| 4 | 0.898 | 0.914 | 0.923 | 0.933 | 0.941 | 0.946 | 0.949 | 0.952 | 0.952
| 5 | 0.907 | 0.925 | 0.935 | 0.947 | 0.956 | 0.963 | 0.966 | 0.97 | 0.97
| 7 | 0.918 | 0.939 | 0.952 | 0.967 | 0.979 | 0.987 | 0.992 | 0.997 | 0.998
| 10 | 0.928 | 0.952 | 0.966 | 0.985 | 1 | 1.011 | 1.018 | 1.024 | 1.026
| 15 | 0.936 | 0.962 | 0.979 | 1 | 1.019 | 1.033 | 1.043 | 1.051 | 1.054
| 20 | 0.94 | 0.967 | 0.984 | 1.008 | 1.028 | 1.045 | 1.056 | 1.067 | 1.069
| 30 | 0.945 | 0.973 | 0.991 | 1.016 | 1.038 | 1.058 | 1.072 | 1.085 | 1.089
| 40 | 0.947 | 0.975 | 0.994 | 1.02 | 1.043 | 1.064 | 1.079 | 1.093 | 1.097

## Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 28 | 34.2 | 45.3 | 55 | 58.6
0.1 | 31.4 | 37.34 | 48.2667 | 58.3 | 62.0857
0.2 | 37 | 43.0333 | 53.5 | 62.9 | 66.44
0.3 | 46.4 | 50.32 | 60.4333 | 69.3 | 72.7
0.4 | 54.68 | 58.45 | 67.8 | 75.4 | 78.1286
0.5 | 61.6 | 65.56 | 73.35 | 80.4 | 82.62
0.6 | 68.1667 | 71.5 | 78.8 | 84.2 | 86.6333
0.7 | 73.4 | 76.38 | 82.44 | 88.1 | 89.55
0.8 | 77.64 | 80.42 | 86.1 | 90.3667 | 91.9
0.9 | 81.6 | 83.8 | 88.6429 | 92.58 | 93.95
1 | 84.75 | 86.75 | 91 | 94.4 | 95.45
1.1 | 87.64 | 89.45 | 92.9 | 95.7 | 96.6
1.2 | 89.96 | 91.4 | 94.66 | 97 | 97.7
1.3 | 91.9333 | 93.26 | 95.9667 | 97.8 | 98.4
1.4 | 93.68 | 94.74 | 97.1 | 98.5 | 98.92
1.5 | 94.9 | 96.06 | 97.85 | 99.14 | 99.44
1.6 | 96.225 | 97.1 | 98.625 | 99.4 | 99.8
1.7 | 97.3 | 97.96 | 99.06 | 99.725 | 99.9
1.8 | 98.075 | 98.55 | 99.425 | 99.9 | 100
1.9 | 98.8 | 99.1 | 99.72 | 99.9833 | 100
2 | 99.2 | 99.44 | 99.86 | 100 | 100
2.1 | 99.55 | 99.74 | 99.98 | 100 | 100
2.2 | 99.76 | 99.9 | 100 | 99.94 | 99.825
2.3 | 99.9 | 100 | 99.9 | 99.8 | 99.7
2.4 | 100 | 100 | 99.86 | 99.6 | 99.6
2.5 | 100 | 99.95 | 99.72 | 99.475 | 99.45
2.6 | 99.98 | 99.9 | 99.62 | 99.2 | 99.125
2.7 | 99.86 | 99.8 | 99.45 | 98.96 | 98.85
2.8 | 99.7 | 99.6 | 99.25 | 98.75 | 98.6
2.9 | 99.4667 | 99.4 | 98.9 | 98.4 | 98.3
3 | 99.15 | 99.225 | 98.66 | 98.15 | 98
3.1 | 98.9 | 98.9 | 98.45 | 97.9 | 97.75
3.2 | 98.55 | 98.7 | 98.1 | 97.575 | 97.4
3.3 | 98.24 | 98.3 | 97.78 | 97.25 | 97.06
3.4 | 97.9 | 98.06 | 97.45 | 96.82 | 96.725
3.5 | 97.6 | 97.7333 | 97.1667 | 96.5 | 96.4
3.6 | 97.1 | 97.32 | 96.82 | 96.225 | 96
3.7 | 96.8 | 96.95 | 96.48 | 95.85 | 95.775
3.8 | 96.3 | 96.6 | 96 | 95.55 | 95.36
3.9 | 95.88 | 96.12 | 95.7 | 95.22 | 95.125
4 | 95.45 | 95.825 | 95.4 | 94.92 | 94.75
4.1 | 95.1 | 95.44 | 95 | 94.6 | 94.45
4.2 | 94.7 | 95 | 94.6833 | 94.25 | 94.02
4.3 | 94.35 | 94.6 | 94.36 | 93.9 | 93.74
4.4 | 93.82 | 94.26 | 94.05 | 93.48 | 93.35
4.5 | 93.38 | 93.9333 | 93.6 | 93.2 | 93.05
4.6 | 93.025 | 93.55 | 93.2333 | 92.78 | 92.75
4.7 | 92.56 | 93.1 | 92.76 | 92.5 | 92.36
4.8 | 92 | 92.85 | 92.52 | 92.25 | 92.02
4.9 | 91.6333 | 92.3333 | 92.2 | 91.82 | 91.7
5 | 91.15 | 91.925 | 91.82 | 91.48 | 91.28
5.1 | 90.7 | 91.7 | 91.38 | 91.1 | 91.025
5.2 | 90.3667 | 91.4 | 91.1333 | 90.78 | 90.7167
5.3 | 89.9857 | 90.98 | 90.8333 | 90.45 | 90.3
5.4 | 89.48 | 90.45 | 90.5 | 90.2333 | 89.9333
5.5 | 88.95 | 90.1 | 90.08 | 89.84 | 89.7
5.6 | 88.4667 | 89.75 | 89.8 | 89.4667 | 89.3333
5.7 | 88.1 | 89.42 | 89.4 | 89.1333 | 88.9667
5.8 | 87.76 | 88.9667 | 89.1333 | 88.8333 | 88.62
5.9 | 87.32 | 88.5 | 88.74 | 88.375 | 88.34
6 | 86.925 | 88.2667 | 88.4 | 88.0833 | 88.12
6.1 | 86.46 | 87.92 | 88.075 | 87.8333 | 87.8
6.2 | 85.9 | 87.5333 | 87.78 | 87.525 | 87.5
6.3 | 85.6 | 87.22 | 87.375 | 87.2 | 87
6.4 | 85.2 | 86.85 | 87.1 | 86.775 | 86.725
6.5 | 84.8 | 86.4 | 86.7333 | 86.4 | 86.4667
6.6 | 84.45 | 85.9833 | 86.34 | 86.1 | 86.16
6.7 | 84.05 | 85.65 | 86.02 | 85.8 | 85.78
6.8 | 83.68 | 85.2667 | 85.64 | 85.46 | 85.48
6.9 | 83.28 | 85.06 | 85.3333 | 85.1 | 85.2286
7 | 82.82 | 84.68 | 85.1333 | 84.84 | 84.85
7.1 | 82.45 | 84.32 | 84.7 | 84.62 | 84.5
7.2 | 82.04 | 83.96 | 84.4667 | 84.26 | 84.3
7.3 | 81.74 | 83.58 | 84.05 | 83.875 | 83.9667
7.4 | 81.26 | 83.175 | 83.675 | 83.64 | 83.6333
7.5 | 80.88 | 82.85 | 83.35 | 83.3 | 83.2
7.6 | 80.7 | 82.6 | 83.12 | 83.05 | 82.9667
7.7 | 80.16 | 82.1 | 82.7 | 82.75 | 82.7
7.8 | 79.7 | 81.7167 | 82.3667 | 82.3 | 82.36
7.9 | 79.3667 | 81.4 | 82.14 | 82.04 | 82.06
8 | 79.04 | 81.1 | 81.76 | 81.72 | 81.75
8.1 | 78.7 | 80.78 | 81.38 | 81.48 | 81.425
8.2 | 78.16 | 80.3 | 81.04 | 81 | 81.08
8.3 | 77.85 | 80.06 | 80.675 | 80.74 | 80.8
8.4 | 77.4 | 79.675 | 80.38 | 80.55 | 80.45
8.5 | 76.92 | 79.425 | 80.1 | 80.2 | 80.1
8.6 | 76.56 | 79.04 | 79.775 | 79.9 | 79.86
8.7 | 76.26 | 78.72 | 79.5 | 79.58 | 79.5
8.8 | 75.9333 | 78.35 | 79.2 | 79.25 | 79.2833
8.9 | 75.425 | 78.04 | 78.8667 | 78.9 | 78.96
9 | 75.025 | 77.72 | 78.5 | 78.5667 | 78.6286
9.1 | 74.725 | 77.275 | 78.1 | 78.2 | 78.38
9.2 | 74.42 | 77 | 77.8667 | 78.0333 | 78.025
9.3 | 74.0333 | 76.55 | 77.625 | 77.76 | 77.74
9.4 | 73.625 | 76.2 | 77.2333 | 77.375 | 77.42
9.5 | 73.3 | 75.9667 | 76.95 | 76.98 | 77.075
9.6 | 72.8333 | 75.6333 | 76.66 | 76.78 | 76.825
9.7 | 72.48 | 75.2 | 76.325 | 76.46 | 76.58
9.8 | 72.2 | 74.8667 | 76.075 | 76.25 | 76.25
9.9 | 71.725 | 74.58 | 75.7667 | 75.94 | 75.925
10 | 71.38 | 74.275 | 75.45 | 75.5 | 75.5667
10.1 | 71.05 | 73.98 | 75.12 | 75.275 | 75.3167
10.2 | 70.74 | 73.6 | 74.8 | 74.95 | 75
10.3 | 70.35 | 73.26 | 74.48 | 74.7 | 74.575
10.4 | 70.08 | 72.96 | 74.26 | 74.3 | 74.3667
10.5 | 69.68 | 72.575 | 73.9 | 74.06 | 74.08
10.6 | 69.3 | 72.2333 | 73.55 | 73.8143 | 73.86
10.7 | 69 | 71.9 | 73.225 | 73.5 | 73.58
10.8 | 68.65 | 71.6 | 72.96 | 73.2 | 73.22
10.9 | 68.18 | 71.28 | 72.66 | 73 | 72.94
11 | 67.88 | 70.975 | 72.375 | 72.6667 | 72.72
11.1 | 67.6667 | 70.65 | 72.1667 | 72.4 | 72.5
11.2 | 67.25 | 70.34 | 71.875 | 72.16 | 72.175
11.3 | 66.9333 | 70 | 71.55 | 71.85 | 71.94
11.4 | 66.6 | 69.6 | 71.25 | 71.54 | 71.5667
11.5 | 66.22 | 69.425 | 70.94 | 71.28 | 71.34
11.6 | 66 | 69.08 | 70.68 | 70.98 | 71.0667
11.7 | 65.66 | 68.775 | 70.35 | 70.64 | 70.78
11.8 | 65.2 | 68.525 | 70.0667 | 70.42 | 70.55
11.9 | 64.98 | 68.14 | 69.7667 | 70.1 | 70.28
12 | 64.62 | 67.9 | 69.5 | 69.86 | 70
12.1 | 64.325 | 67.7 | 69.225 | 69.56 | 69.72
12.2 | 63.9667 | 67.25 | 68.925 | 69.325 | 69.42
12.3 | 63.7 | 67.06 | 68.62 | 69.1167 | 69.22
12.4 | 63.5 | 66.66 | 68.3667 | 68.7857 | 68.9
12.5 | 63 | 66.32 | 68.02 | 68.5 | 68.55
12.6 | 62.76 | 66.06 | 67.6667 | 68.2 | 68.3833
12.7 | 62.4 | 65.75 | 67.4833 | 67.94 | 68.1
12.8 | 62.26 | 65.46 | 67.3 | 67.72 | 67.8667
12.9 | 61.85 | 65.1286 | 66.95 | 67.4333 | 67.5
13 | 61.5 | 64.8333 | 66.72 | 67.2 | 67.275
13.1 | 61.25 | 64.575 | 66.45 | 66.8667 | 67.025
13.2 | 60.9 | 64.34 | 66.18 | 66.7143 | 66.78
13.3 | 60.55 | 64.02 | 65.9 | 66.4 | 66.48
13.4 | 60.2 | 63.6667 | 65.48 | 66.2 | 66.1833
13.5 | 59.9333 | 63.5 | 65.3 | 65.88 | 65.9
13.6 | 59.725 | 63.24 | 65.025 | 65.6 | 65.675
13.7 | 59.4 | 62.8 | 64.78 | 65.2833 | 65.5
13.8 | 59.1667 | 62.7 | 64.46 | 65.08 | 65.26
13.9 | 58.78 | 62.4 | 64.25 | 64.9 | 65
14 | 58.475 | 62.06 | 63.95 | 64.56 | 64.65
14.1 | 58.25 | 61.78 | 63.6333 | 64.25 | 64.42
14.2 | 57.875 | 61.44 | 63.38 | 64.0333 | 64.175
14.3 | 57.65 | 61.2 | 63.15 | 63.75 | 63.875
14.4 | 57.38 | 60.95 | 62.9 | 63.46 | 63.6571
14.5 | 57.1667 | 60.62 | 62.68 | 63.25 | 63.36
14.6 | 56.7 | 60.325 | 62.38 | 63.025 | 63.06
14.7 | 56.5333 | 60.1 | 62.1 | 62.7 | 62.9
14.8 | 56.3 | 59.82 | 61.8333 | 62.46 | 62.72
14.9 | 55.96 | 59.5333 | 61.66 | 62.16 | 62.5
15 | 55.66 | 59.3143 | 61.3333 | 62 | 62.175
15.1 | 55.44 | 59.02 | 61.05 | 61.675 | 61.9333
15.2 | 55.0833 | 58.8 | 60.8 | 61.45 | 61.65
15.3 | 54.9286 | 58.5 | 60.58 | 61.24 | 61.4
15.4 | 54.72 | 58.24 | 60.225 | 60.98 | 61.2
15.5 | 54.3667 | 57.9667 | 60.0333 | 60.7 | 60.95
15.6 | 54.1 | 57.65 | 59.8 | 60.4 | 60.6667
15.7 | 53.8 | 57.34 | 59.5333 | 60.2333 | 60.4167
15.8 | 53.55 | 57.2 | 59.275 | 60.025 | 60.125
15.9 | 53.22 | 56.9 | 59.025 | 59.7 | 60
16 | 53.0333 | 56.575 | 58.86 | 59.55 | 59.7167
16.1 | 52.8 | 56.35 | 58.55 | 59.36 | 59.4
16.2 | 52.5 | 56.15 | 58.35 | 59.1 | 59.3
16.3 | 52.2667 | 55.8 | 58.0667 | 58.8 | 59
16.4 | 51.98 | 55.68 | 57.9 | 58.6667 | 58.8
16.5 | 51.78 | 55.45 | 57.5833 | 58.425 | 58.54
16.6 | 51.52 | 55.1 | 57.36 | 58.2167 | 58.325
16.7 | 51.26 | 54.9 | 57.0667 | 57.95 | 58.0333
16.8 | 51 | 54.7 | 56.8 | 57.675 | 57.8
16.9 | 50.62 | 54.42 | 56.6 | 57.4667 | 57.6333
17 | 50.35 | 54.1 | 56.42 | 57.2 | 57.42
17.1 | 50.28 | 53.8333 | 56.2 | 57 | 57.125
17.2 | 50 | 53.62 | 55.94 | 56.8 | 56.9
17.3 | 49.74 | 53.4167 | 55.8 | 56.5333 | 56.65
17.4 | 49.475 | 53.06 | 55.46 | 56.3 | 56.425
17.5 | 49.325 | 52.9 | 55.2 | 56.1 | 56.25
17.6 | 49.04 | 52.675 | 55.05 | 55.82 | 56.06
17.7 | 48.78 | 52.4333 | 54.86 | 55.6 | 55.85
17.8 | 48.6 | 52.175 | 54.56 | 55.46 | 55.625
17.9 | 48.4 | 51.975 | 54.325 | 55.15 | 55.32
18 | 48.2 | 51.74 | 54.0333 | 54.9 | 55.1
18.1 | 47.85 | 51.5 | 53.9 | 54.66 | 54.95
18.2 | 47.675 | 51.275 | 53.64 | 54.45 | 54.775
18.3 | 47.44 | 51.025 | 53.44 | 54.3 | 54.525
18.4 | 47.225 | 50.775 | 53.25 | 54.1 | 54.3
18.5 | 46.975 | 50.54 | 53.05 | 53.8667 | 54.1
18.6 | 46.6833 | 50.3 | 52.8 | 53.64 | 53.9333
18.7 | 46.4833 | 50.1167 | 52.575 | 53.52 | 53.64
18.8 | 46.3 | 49.9 | 52.34 | 53.24 | 53.42
18.9 | 46.0333 | 49.6833 | 52.04 | 52.9 | 53.2333
19 | 45.775 | 49.38 | 51.86 | 52.8 | 53.025
19.1 | 45.54 | 49.18 | 51.6667 | 52.5667 | 52.8333
19.2 | 45.34 | 48.9 | 51.475 | 52.3429 | 52.64
19.3 | 45.1167 | 48.76 | 51.2667 | 52.15 | 52.34
19.4 | 44.9 | 48.46 | 51 | 51.825 | 52.175
19.5 | 44.7 | 48.25 | 50.8 | 51.72 | 51.9
19.6 | 44.5 | 48.125 | 50.5667 | 51.5 | 51.7
19.7 | 44.3 | 47.8667 | 50.325 | 51.3 | 51.52
19.8 | 44.0333 | 47.6333 | 50.1167 | 51.1 | 51.2333
19.9 | 43.74 | 47.3 | 49.92 | 50.94 | 51.0333
20 | 43.5 | 47.1333 | 49.7 | 50.64 | 50.9
20.1 | 43.375 | 46.95 | 49.4667 | 50.425 | 50.6833
20.2 | 43.16 | 46.78 | 49.3 | 50.15 | 50.48
20.3 | 43.02 | 46.5667 | 49.1 | 50 | 50.275
20.4 | 42.75 | 46.2571 | 48.9143 | 49.78 | 50.1
20.5 | 42.625 | 46.06 | 48.68 | 49.7 | 49.875
20.6 | 42.44 | 45.8714 | 48.4429 | 49.4333 | 49.64
20.7 | 42.15 | 45.7 | 48.2 | 49.2 | 49.4833
20.8 | 42 | 45.4833 | 48 | 49.1 | 49.3
20.9 | 41.7 | 45.3 | 47.8 | 48.775 | 49
21 | 41.54 | 45.14 | 47.65 | 48.6 | 48.88
21.1 | 41.3429 | 44.94 | 47.475 | 48.475 | 48.6833
21.2 | 41.0833 | 44.64 | 47.225 | 48.25 | 48.48
21.3 | 40.9 | 44.4 | 47.0333 | 48.0667 | 48.3
21.4 | 40.68 | 44.26 | 46.86 | 47.8333 | 48.04
21.5 | 40.5 | 44.0833 | 46.66 | 47.6 | 47.8
21.6 | 40.36 | 43.9167 | 46.45 | 47.45 | 47.625
21.7 | 40.14 | 43.7 | 46.3 | 47.3 | 47.5429
21.8 | 40 | 43.5 | 46.075 | 47.1 | 47.3
21.9 | 39.7 | 43.28 | 45.8 | 46.8714 | 47.175
22 | 39.6 | 43.08 | 45.65 | 46.62 | 46.95
22.1 | 39.3833 | 42.9143 | 45.4667 | 46.5 | 46.72
22.2 | 39.1833 | 42.72 | 45.3 | 46.3 | 46.55
22.3 | 38.9833 | 42.5 | 45.1167 | 46.18 | 46.38
22.4 | 38.9 | 42.3667 | 44.9 | 45.94 | 46.16
22.5 | 38.7 | 42.1667 | 44.7333 | 45.72 | 46
22.6 | 38.525 | 41.9 | 44.5 | 45.6 | 45.85
22.7 | 38.3 | 41.7333 | 44.3667 | 45.4 | 45.6833
22.8 | 38.1333 | 41.6 | 44.1167 | 45.3 | 45.5167
22.9 | 37.9667 | 41.42 | 44 | 45.025 | 45.3167
23 | 37.8 | 41.16 | 43.8333 | 44.775 | 45.1
23.1 | 37.5 | 41.04 | 43.6 | 44.6667 | 44.88
23.2 | 37.3 | 40.8333 | 43.4 | 44.4833 | 44.7
23.3 | 37.16 | 40.58 | 43.1667 | 44.3167 | 44.575
23.4 | 37 | 40.48 | 43.0833 | 44.15 | 44.34
23.5 | 36.86 | 40.28 | 42.9167 | 44.075 | 44.15
23.6 | 36.6333 | 40.06 | 42.74 | 43.75 | 44.04
23.7 | 36.44 | 39.92 | 42.6 | 43.6 | 43.825
23.8 | 36.325 | 39.72 | 42.35 | 43.3333 | 43.7
23.9 | 36.1333 | 39.525 | 42.2 | 43.2 | 43.4833
24 | 35.925 | 39.4 | 42 | 43.0333 | 43.3167
24.1 | 35.72 | 39.2 | 41.78 | 42.86 | 43.1333
24.2 | 35.5 | 39 | 41.68 | 42.6833 | 42.98
24.3 | 35.4 | 38.82 | 41.5 | 42.4833 | 42.76
24.4 | 35.2 | 38.6333 | 41.3 | 42.3167 | 42.6
24.5 | 35.1 | 38.52 | 41.08 | 42.1833 | 42.4667
24.6 | 34.88 | 38.2 | 40.95 | 42.0167 | 42.275
24.7 | 34.6667 | 38.1 | 40.775 | 41.86 | 42.1
24.8 | 34.58 | 37.95 | 40.525 | 41.6833 | 41.9
24.9 | 34.3833 | 37.8 | 40.38 | 41.5 | 41.76
25 | 34.2 | 37.62 | 40.2 | 41.3 | 41.5667
25.1 | 34.08 | 37.44 | 40.075 | 41.15 | 41.4
25.2 | 33.95 | 37.3 | 39.9 | 41 | 41.2
25.3 | 33.725 | 37.15 | 39.8 | 40.7667 | 41.0714
25.4 | 33.58 | 36.94 | 39.525 | 40.6 | 40.9
25.5 | 33.4 | 36.775 | 39.4 | 40.42 | 40.775
25.6 | 33.3 | 36.625 | 39.28 | 40.3 | 40.55
25.7 | 33.1 | 36.45 | 39.075 | 40.2 | 40.475
25.8 | 32.9 | 36.325 | 38.86 | 39.94 | 40.225
25.9 | 32.85 | 36.14 | 38.76 | 39.85 | 40.1
26 | 32.7 | 35.95 | 38.5833 | 39.7 | 39.98
26.1 | 32.5 | 35.78 | 38.4167 | 39.5 | 39.8
26.2 | 32.34 | 35.7 | 38.22 | 39.4 | 39.675
26.3 | 32.2 | 35.5 | 38.075 | 39.24 | 39.4833
26.4 | 32.1 | 35.3 | 37.8667 | 39.04 | 39.3
26.5 | 31.8333 | 35.15 | 37.7 | 38.9 | 39.2
26.6 | 31.75 | 35.05 | 37.6 | 38.7667 | 38.9833
26.7 | 31.6 | 34.85 | 37.5 | 38.54 | 38.8
26.8 | 31.46 | 34.62 | 37.3 | 38.4 | 38.675
26.9 | 31.26 | 34.5 | 37.1833 | 38.25 | 38.5
27 | 31.125 | 34.35 | 36.975 | 38.125 | 38.375
27.1 | 31 | 34.2333 | 36.86 | 37.9 | 38.2
27.2 | 30.9 | 34.05 | 36.6667 | 37.8 | 38.075
27.3 | 30.72 | 33.8833 | 36.5 | 37.6 | 37.95
27.4 | 30.5833 | 33.7167 | 36.375 | 37.5 | 37.72
27.5 | 30.3833 | 33.62 | 36.24 | 37.3333 | 37.62
27.6 | 30.2167 | 33.5 | 36.1 | 37.2 | 37.475
27.7 | 30.1333 | 33.26 | 35.9 | 37.0333 | 37.3
27.8 | 29.9 | 33.175 | 35.72 | 36.86 | 37.2
27.9 | 29.88 | 33 | 35.6 | 36.76 | 37.0167
28 | 29.7 | 32.8667 | 35.4 | 36.5667 | 36.8
28.1 | 29.5667 | 32.7 | 35.28 | 36.4 | 36.7
28.2 | 29.375 | 32.5667 | 35.1 | 36.25 | 36.54
28.3 | 29.2 | 32.42 | 34.95 | 36.1333 | 36.4333
28.4 | 29.1 | 32.3 | 34.8 | 36 | 36.225
28.5 | 29.0333 | 32.1333 | 34.7 | 35.875 | 36.1429
28.6 | 28.9 | 31.9667 | 34.575 | 35.68 | 35.925
28.7 | 28.7333 | 31.8 | 34.46 | 35.5 | 35.82
28.8 | 28.6 | 31.72 | 34.3 | 35.4 | 35.7
28.9 | 28.4667 | 31.5 | 34.14 | 35.26 | 35.525
29 | 28.3667 | 31.375 | 34 | 35.14 | 35.34
29.1 | 28.2 | 31.225 | 33.8833 | 34.96 | 35.25
29.2 | 28.0833 | 31.1 | 33.7167 | 34.8 | 35.075
29.3 | 27.9167 | 30.9833 | 33.6 | 34.68 | 34.9667
29.4 | 27.74 | 30.8 | 33.4 | 34.58 | 34.8
29.5 | 27.6333 | 30.7 | 33.225 | 34.3667 | 34.6833
29.6 | 27.5571 | 30.525 | 33.1333 | 34.2 | 34.5167
29.7 | 27.5 | 30.4 | 33.0333 | 34.1 | 34.4
29.8 | 27.3167 | 30.3 | 32.84 | 33.94 | 34.2
29.9 | 27.2 | 30.1167 | 32.75 | 33.8 | 34.1
30 | 27.04 | 30 | 32.6 | 33.6833 | 34

## SSD80 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 29.5 | 38.1 | 54.7 | 67.9 | 72.4
0.1 | 33.45 | 42.5125 | 58.975 | 71.85 | 76.05
0.2 | 40.1 | 49.0625 | 65 | 76.9 | 80.65
0.3 | 49.75 | 57.675 | 71.825 | 82.375 | 85.64
0.4 | 58.1 | 65.575 | 78.025 | 86.75 | 89.4
0.5 | 65.375 | 71.6 | 82.5 | 90.3 | 92.45
0.6 | 71.25 | 76.8 | 86.65 | 92.95 | 94.6
0.7 | 76 | 81.2 | 89.6 | 94.95 | 96.3
0.8 | 80.4 | 84.8 | 92.125 | 96.475 | 97.6
0.9 | 83.85 | 87.96 | 94.1 | 97.6 | 98.4
1 | 87.05 | 90.45 | 95.7 | 98.54 | 99.15
1.1 | 89.7 | 92.5 | 96.9 | 99.1 | 99.5375
1.2 | 91.8 | 94.2 | 98 | 99.55 | 99.8375
1.3 | 93.6 | 95.7 | 98.64 | 99.8 | 99.9625
1.4 | 95.2 | 96.725 | 99.1571 | 100 | 100
1.5 | 96.35 | 97.725 | 99.6 | 100 | 100
1.6 | 97.425 | 98.48 | 99.825 | 100 | 99.92
1.7 | 98.175 | 99.05 | 100 | 99.95 | 99.8
1.8 | 98.85 | 99.52 | 100 | 99.8 | 99.65
1.9 | 99.275 | 99.75 | 100 | 99.55 | 99.4
2 | 99.7 | 99.9 | 99.9375 | 99.4 | 99.1667
2.1 | 99.875 | 100 | 99.8125 | 99.15 | 98.85
2.2 | 100 | 100 | 99.6 | 98.85 | 98.575
2.3 | 100 | 100 | 99.4 | 98.55 | 98.25
2.4 | 100 | 99.9 | 99.12 | 98.225 | 97.95
2.5 | 99.9 | 99.65 | 98.85 | 97.85 | 97.6
2.6 | 99.85 | 99.5 | 98.525 | 97.525 | 97.2
2.7 | 99.5667 | 99.25 | 98.2 | 97.175 | 96.8333
2.8 | 99.1667 | 98.9 | 97.8 | 96.825 | 96.5
2.9 | 98.9 | 98.6 | 97.5 | 96.45 | 96.1
3 | 98.5 | 98.275 | 97.05 | 96.04 | 95.8
3.1 | 98.25 | 97.95 | 96.75 | 95.7 | 95.35
3.2 | 97.8 | 97.625 | 96.35 | 95.3 | 95
3.3 | 97.46 | 97.125 | 95.9333 | 94.9 | 94.55
3.4 | 96.975 | 96.85 | 95.6 | 94.5 | 94.24
3.5 | 96.475 | 96.35 | 95.2 | 94.1 | 93.8333
3.6 | 96.05 | 95.975 | 94.7667 | 93.8 | 93.4
3.7 | 95.675 | 95.55 | 94.35 | 93.3 | 93
3.8 | 95.25 | 95.1 | 93.95 | 92.9 | 92.6
3.9 | 94.75 | 94.7 | 93.575 | 92.5667 | 92.275
4 | 94.25 | 94.24 | 93.15 | 92.15 | 91.925
4.1 | 93.65 | 93.85 | 92.75 | 91.7 | 91.475
4.2 | 93.225 | 93.35 | 92.2875 | 91.35 | 91.075
4.3 | 92.85 | 92.9667 | 91.9571 | 91.04 | 90.75
4.4 | 92.25 | 92.5 | 91.54 | 90.575 | 90.4
4.5 | 91.7 | 92.08 | 91.1 | 90.25 | 90.0333
4.6 | 91.28 | 91.6 | 90.775 | 89.7667 | 89.6333
4.7 | 90.8 | 91.075 | 90.25 | 89.45 | 89.12
4.8 | 90.45 | 90.76 | 89.875 | 89.025 | 88.7625
4.9 | 89.9 | 90.3 | 89.45 | 88.6667 | 88.425
5 | 89.45 | 89.8 | 89.15 | 88.1667 | 88.1125
5.1 | 88.9 | 89.42 | 88.65 | 87.9 | 87.7667
5.2 | 88.55 | 89.025 | 88.25 | 87.4667 | 87.4
5.3 | 88.05 | 88.65 | 87.85 | 87.12 | 86.9429
5.4 | 87.45 | 88.15 | 87.5 | 86.7333 | 86.54
5.5 | 87.025 | 87.775 | 87.0333 | 86.375 | 86.075
5.6 | 86.625 | 87.35 | 86.7 | 86.05 | 85.7875
5.7 | 86.08 | 86.875 | 86.275 | 85.625 | 85.4125
5.8 | 85.7 | 86.6 | 85.925 | 85.2667 | 85.0375
5.9 | 85.2 | 86 | 85.575 | 84.9 | 84.6714
6 | 84.7667 | 85.65 | 85.225 | 84.48 | 84.4
6.1 | 84.36 | 85.3333 | 84.75 | 84.1 | 83.95
6.2 | 83.85 | 84.8333 | 84.45 | 83.7667 | 83.6429
6.3 | 83.35 | 84.4 | 84.1 | 83.3857 | 83.275
6.4 | 82.88 | 84 | 83.65 | 83.05 | 82.925
6.5 | 82.6 | 83.66 | 83.36 | 82.6333 | 82.55
6.6 | 82.05 | 83.175 | 82.9 | 82.26 | 82.15
6.7 | 81.6333 | 82.85 | 82.5 | 81.9286 | 81.775
6.8 | 81.1125 | 82.4571 | 82.2 | 81.5778 | 81.45
6.9 | 80.675 | 82 | 81.8 | 81.3 | 81.15
7 | 80.36 | 81.65 | 81.3667 | 80.85 | 80.8
7.1 | 79.8 | 81.26 | 81.0667 | 80.5286 | 80.5
7.2 | 79.4 | 80.85 | 80.7 | 80.15 | 80.1
7.3 | 79 | 80.42 | 80.4 | 79.75 | 79.75
7.4 | 78.54 | 80.025 | 79.95 | 79.425 | 79.375
7.5 | 78.175 | 79.675 | 79.7 | 79.1125 | 79.025
7.6 | 77.725 | 79.3 | 79.2333 | 78.8 | 78.675
7.7 | 77.35 | 78.9 | 78.875 | 78.4 | 78.34
7.8 | 76.9 | 78.45 | 78.5375 | 78.1 | 78
7.9 | 76.5 | 78.05 | 78.15 | 77.8 | 77.7
8 | 76.1 | 77.7667 | 77.825 | 77.4 | 77.275
8.1 | 75.7 | 77.35 | 77.4 | 77.1 | 77.04
8.2 | 75.3 | 76.9667 | 77.125 | 76.75 | 76.7
8.3 | 74.875 | 76.5667 | 76.7667 | 76.375 | 76.375
8.4 | 74.45 | 76.25 | 76.4 | 76.05 | 76.025
8.5 | 74.125 | 75.85 | 76.05 | 75.775 | 75.725
8.6 | 73.64 | 75.45 | 75.7333 | 75.4 | 75.325
8.7 | 73.3 | 75.1 | 75.375 | 75.05 | 75.05
8.8 | 72.85 | 74.775 | 75.05 | 74.7 | 74.625
8.9 | 72.3 | 74.425 | 74.675 | 74.35 | 74.4
9 | 72 | 73.95 | 74.325 | 74.1 | 74
9.1 | 71.65 | 73.55 | 73.975 | 73.725 | 73.65
9.2 | 71.32 | 73.2 | 73.65 | 73.475 | 73.4
9.3 | 70.95 | 72.85 | 73.35 | 73.15 | 73.1
9.4 | 70.55 | 72.6 | 73 | 72.75 | 72.725
9.5 | 70.05 | 72.1 | 72.5667 | 72.45 | 72.4
9.6 | 69.68 | 71.825 | 72.25 | 72.1333 | 72.1
9.7 | 69.425 | 71.44 | 71.9 | 71.7667 | 71.7
9.8 | 69.06 | 71.0333 | 71.625 | 71.5 | 71.5
9.9 | 68.55 | 70.725 | 71.2667 | 71.1 | 71.1
10 | 68.3 | 70.275 | 70.925 | 70.8 | 70.7
10.1 | 67.9 | 69.9875 | 70.575 | 70.5 | 70.475
10.2 | 67.5 | 69.6125 | 70.25 | 70.2 | 70.15
10.3 | 67.0333 | 69.2667 | 69.975 | 69.8625 | 69.85
10.4 | 66.675 | 68.9 | 69.625 | 69.525 | 69.525
10.5 | 66.35 | 68.625 | 69.28 | 69.2125 | 69.3
10.6 | 65.88 | 68.3 | 69.0333 | 68.925 | 68.95
10.7 | 65.6 | 67.975 | 68.6667 | 68.55 | 68.6
10.8 | 65.4 | 67.55 | 68.3333 | 68.4 | 68.3
10.9 | 64.96 | 67.175 | 68 | 68 | 68.02
11 | 64.6 | 66.85 | 67.7 | 67.7 | 67.75
11.1 | 64.3 | 66.56 | 67.35 | 67.35 | 67.525
11.2 | 63.875 | 66.2 | 67.02 | 67.12 | 67.075
11.3 | 63.525 | 65.9 | 66.85 | 66.85 | 66.8
11.4 | 63.1857 | 65.6 | 66.4 | 66.525 | 66.5
11.5 | 62.85 | 65.2 | 66.15 | 66.2 | 66.2333
11.6 | 62.45 | 64.95 | 65.8667 | 65.95 | 65.94
11.7 | 62.18 | 64.625 | 65.5 | 65.6 | 65.7
11.8 | 61.94 | 64.275 | 65.26 | 65.35 | 65.35
11.9 | 61.5667 | 63.925 | 64.925 | 65.1 | 65.125
12 | 61.2 | 63.5875 | 64.65 | 64.76 | 64.78
12.1 | 60.8667 | 63.275 | 64.325 | 64.45 | 64.525
12.2 | 60.475 | 62.9667 | 64.075 | 64.075 | 64.3
12.3 | 60.3 | 62.7 | 63.725 | 63.8 | 63.95
12.4 | 60.05 | 62.425 | 63.375 | 63.575 | 63.6333
12.5 | 59.64 | 62.0667 | 63.15 | 63.325 | 63.3667
12.6 | 59.375 | 61.7125 | 62.9 | 63.075 | 63.12
12.7 | 58.95 | 61.425 | 62.55 | 62.725 | 62.775
12.8 | 58.55 | 61.1125 | 62.3 | 62.45 | 62.525
12.9 | 58.325 | 60.8 | 61.95 | 62.125 | 62.275
13 | 58 | 60.48 | 61.7 | 61.875 | 62.025
13.1 | 57.75 | 60.22 | 61.45 | 61.68 | 61.75
13.2 | 57.3 | 59.86 | 61.1 | 61.38 | 61.425
13.3 | 57.12 | 59.65 | 60.8 | 61.1333 | 61.175
13.4 | 56.725 | 59.3 | 60.475 | 60.8 | 60.88
13.5 | 56.475 | 59.06 | 60.225 | 60.54 | 60.68
13.6 | 56.15 | 58.75 | 59.975 | 60.3 | 60.35
13.7 | 55.8 | 58.375 | 59.74 | 59.95 | 60.1
13.8 | 55.55 | 58.125 | 59.4571 | 59.8 | 59.85
13.9 | 55.3 | 57.875 | 59.1 | 59.55 | 59.625
14 | 54.85 | 57.55 | 58.85 | 59.2 | 59.3
14.1 | 54.7 | 57.25 | 58.6 | 58.875 | 59.125
14.2 | 54.3667 | 56.925 | 58.4 | 58.625 | 58.775
14.3 | 54.1 | 56.7 | 58.08 | 58.4 | 58.5
14.4 | 53.75 | 56.425 | 57.68 | 58.2 | 58.2
14.5 | 53.575 | 56.1 | 57.4667 | 57.85 | 57.95
14.6 | 53.15 | 55.9 | 57.225 | 57.6 | 57.7
14.7 | 52.875 | 55.45 | 56.975 | 57.3667 | 57.5
14.8 | 52.625 | 55.2 | 56.65 | 57.06 | 57.2
14.9 | 52.375 | 54.92 | 56.375 | 56.8 | 56.95
15 | 52.125 | 54.625 | 56.2 | 56.56 | 56.65
15.1 | 51.775 | 54.38 | 55.875 | 56.25 | 56.375
15.2 | 51.45 | 54.15 | 55.65 | 55.9667 | 56.125
15.3 | 51.2 | 53.9 | 55.275 | 55.7833 | 55.875
15.4 | 50.9 | 53.625 | 55.025 | 55.58 | 55.7
15.5 | 50.6 | 53.3 | 54.775 | 55.35 | 55.38
15.6 | 50.35 | 53 | 54.525 | 55 | 55.15
15.7 | 50.1 | 52.825 | 54.3 | 54.8 | 54.9
15.8 | 49.875 | 52.475 | 54.15 | 54.6 | 54.75
15.9 | 49.625 | 52.25 | 53.9 | 54.3667 | 54.5
16 | 49.375 | 52.0333 | 53.65 | 54.1 | 54.1667
16.1 | 49.1 | 51.7375 | 53.3 | 53.8 | 53.9667
16.2 | 48.875 | 51.45 | 53.05 | 53.56 | 53.725
16.3 | 48.6 | 51.25 | 52.825 | 53.325 | 53.4875
16.4 | 48.4 | 51 | 52.675 | 53.15 | 53.325
16.5 | 48.14 | 50.74 | 52.425 | 52.825 | 53.1
16.6 | 47.8 | 50.4 | 52.1667 | 52.65 | 52.8333
16.7 | 47.55 | 50.25 | 51.825 | 52.375 | 52.5
16.8 | 47.25 | 50 | 51.65 | 52.1625 | 52.3
16.9 | 47 | 49.65 | 51.375 | 51.9333 | 52.06
17 | 46.75 | 49.45 | 51.125 | 51.65 | 51.825
17.1 | 46.6 | 49.25 | 50.875 | 51.4 | 51.6333
17.2 | 46.36 | 49 | 50.725 | 51.225 | 51.375
17.3 | 46.025 | 48.7 | 50.4333 | 51 | 51.125
17.4 | 45.775 | 48.5 | 50.25 | 50.75 | 50.98
17.5 | 45.625 | 48.2 | 50 | 50.55 | 50.75
17.6 | 45.375 | 48 | 49.75 | 50.2857 | 50.5
17.7 | 45.125 | 47.8 | 49.5 | 50.1 | 50.25
17.8 | 44.875 | 47.55 | 49.25 | 49.8667 | 50
17.9 | 44.625 | 47.3 | 49 | 49.7125 | 49.8
18 | 44.375 | 47.0333 | 48.8 | 49.45 | 49.6
18.1 | 44.125 | 46.825 | 48.625 | 49.24 | 49.35
18.2 | 43.875 | 46.575 | 48.375 | 49 | 49.22
18.3 | 43.7 | 46.325 | 48.125 | 48.75 | 48.9667
18.4 | 43.575 | 46.1 | 47.875 | 48.6 | 48.72
18.5 | 43.28 | 45.8333 | 47.74 | 48.3 | 48.475
18.6 | 43.08 | 45.6667 | 47.4 | 48.1 | 48.3
18.7 | 42.75 | 45.4 | 47.225 | 47.9 | 48.075
18.8 | 42.5 | 45.2 | 46.975 | 47.6667 | 47.825
18.9 | 42.35 | 44.9667 | 46.825 | 47.4 | 47.675
19 | 42.2 | 44.775 | 46.58 | 47.275 | 47.425
19.1 | 41.95 | 44.475 | 46.38 | 47.025 | 47.2
19.2 | 41.7 | 44.3 | 46.1 | 46.8 | 46.95
19.3 | 41.45 | 44.125 | 45.95 | 46.625 | 46.78
19.4 | 41.2 | 43.875 | 45.7 | 46.4 | 46.55
19.5 | 41.05 | 43.65 | 45.4667 | 46.2333 | 46.4
19.6 | 40.9 | 43.5 | 45.35 | 46 | 46.15
19.7 | 40.575 | 43.25 | 45.04 | 45.75 | 46
19.8 | 40.4 | 43 | 44.875 | 45.6 | 45.7667
19.9 | 40.175 | 42.75 | 44.625 | 45.35 | 45.5
20 | 39.9 | 42.6 | 44.425 | 45.15 | 45.275
20.1 | 39.775 | 42.375 | 44.28 | 44.95 | 45.125
20.2 | 39.525 | 42.15 | 44.025 | 44.7286 | 44.95
20.3 | 39.3333 | 41.9 | 43.775 | 44.5 | 44.74
20.4 | 39.1667 | 41.7 | 43.6333 | 44.325 | 44.5
20.5 | 38.9 | 41.5 | 43.38 | 44.075 | 44.35
20.6 | 38.7 | 41.375 | 43.2 | 43.9 | 44.075
20.7 | 38.52 | 41.025 | 43 | 43.725 | 43.9
20.8 | 38.375 | 40.875 | 42.775 | 43.48 | 43.775
20.9 | 38.14 | 40.64 | 42.55 | 43.35 | 43.525
21 | 37.9 | 40.5 | 42.4 | 43.1 | 43.375
21.1 | 37.76 | 40.3 | 42.2 | 42.95 | 43.1333
21.2 | 37.525 | 40.08 | 42 | 42.8 | 43
21.3 | 37.3375 | 39.9 | 41.8 | 42.525 | 42.8
21.4 | 37.1444 | 39.7 | 41.675 | 42.375 | 42.6
21.5 | 36.9 | 39.55 | 41.5 | 42.2 | 42.45
21.6 | 36.775 | 39.3 | 41.275 | 41.975 | 42.2
21.7 | 36.55 | 39.1 | 41.025 | 41.75 | 42
21.8 | 36.325 | 38.9 | 40.875 | 41.6 | 41.8
21.9 | 36.2 | 38.75 | 40.6833 | 41.425 | 41.65
22 | 36 | 38.5 | 40.475 | 41.3 | 41.4
22.1 | 35.825 | 38.3 | 40.3 | 41.05 | 41.26
22.2 | 35.6 | 38.1667 | 40.1 | 40.85 | 41.1
22.3 | 35.425 | 38 | 39.95 | 40.7143 | 40.9
22.4 | 35.3 | 37.775 | 39.7 | 40.5625 | 40.7
22.5 | 35.05 | 37.625 | 39.55 | 40.375 | 40.575
22.6 | 34.9 | 37.38 | 39.4 | 40.15 | 40.325
22.7 | 34.65 | 37.25 | 39.2 | 39.975 | 40.175
22.8 | 34.5 | 37 | 39.02 | 39.8 | 40.04
22.9 | 34.36 | 36.85 | 38.775 | 39.6 | 39.8
23 | 34.15 | 36.7 | 38.625 | 39.45 | 39.65
23.1 | 34 | 36.45 | 38.4 | 39.3 | 39.5
23.2 | 33.825 | 36.3333 | 38.24 | 39.1 | 39.2667
23.3 | 33.675 | 36.125 | 38.1 | 38.92 | 39.1333
23.4 | 33.55 | 35.9375 | 37.95 | 38.775 | 39.0125
23.5 | 33.325 | 35.8125 | 37.8 | 38.6 | 38.7778
23.6 | 33.15 | 35.5875 | 37.6 | 38.375 | 38.625
23.7 | 33 | 35.44 | 37.4 | 38.2 | 38.4375
23.8 | 32.9 | 35.24 | 37.25 | 38.1 | 38.25
23.9 | 32.675 | 35.14 | 37.0333 | 37.875 | 38.1
24 | 32.5 | 34.9 | 36.9 | 37.7 | 37.95
24.1 | 32.325 | 34.75 | 36.7 | 37.575 | 37.8
24.2 | 32.2 | 34.6 | 36.5 | 37.4 | 37.6
24.3 | 31.975 | 34.45 | 36.425 | 37.2 | 37.5
24.4 | 31.8556 | 34.2 | 36.2 | 37.0333 | 37.26
24.5 | 31.6333 | 34.1 | 36 | 36.8 | 37.125
24.6 | 31.5 | 33.9 | 35.8 | 36.74 | 37
24.7 | 31.375 | 33.75 | 35.65 | 36.6 | 36.775
24.8 | 31.1333 | 33.6 | 35.575 | 36.4 | 36.6
24.9 | 30.95 | 33.4 | 35.325 | 36.2 | 36.4
25 | 30.8333 | 33.2 | 35.2 | 36.0375 | 36.3
25.1 | 30.66 | 33.05 | 35 | 35.8714 | 36.1
25.2 | 30.6 | 32.875 | 34.8875 | 35.7 | 35.9
25.3 | 30.35 | 32.725 | 34.725 | 35.6 | 35.8
25.4 | 30.175 | 32.6 | 34.5286 | 35.4 | 35.6
25.5 | 30.025 | 32.46 | 34.3625 | 35.2333 | 35.5
25.6 | 29.9 | 32.3 | 34.2375 | 35.0667 | 35.4
25.7 | 29.725 | 32.1333 | 34.1125 | 34.9 | 35.2
25.8 | 29.675 | 32 | 33.9 | 34.8 | 35.0375
25.9 | 29.525 | 31.75 | 33.75 | 34.6429 | 34.9125
26 | 29.34 | 31.625 | 33.66 | 34.5125 | 34.725
26.1 | 29.2 | 31.5 | 33.475 | 34.325 | 34.6
26.2 | 29.05 | 31.36 | 33.3 | 34.2 | 34.425
26.3 | 28.9 | 31.2 | 33.15 | 34.05 | 34.275
26.4 | 28.75 | 31.0667 | 33 | 33.8667 | 34.1
26.5 | 28.6 | 30.875 | 32.85 | 33.7 | 34
26.6 | 28.475 | 30.7 | 32.7 | 33.575 | 33.8
26.7 | 28.325 | 30.6 | 32.5 | 33.425 | 33.675
26.8 | 28.175 | 30.475 | 32.325 | 33.275 | 33.55
26.9 | 28.025 | 30.325 | 32.2 | 33.125 | 33.4
27 | 27.875 | 30.175 | 32.1 | 32.975 | 33.2
27.1 | 27.74 | 30.025 | 31.9 | 32.8 | 33.1
27.2 | 27.6 | 29.9 | 31.8 | 32.68 | 32.9
27.3 | 27.475 | 29.74 | 31.65 | 32.55 | 32.8
27.4 | 27.35 | 29.6 | 31.5 | 32.4 | 32.6
27.5 | 27.225 | 29.45 | 31.4 | 32.2667 | 32.5
27.6 | 27 | 29.3 | 31.2667 | 32.125 | 32.3
27.7 | 26.875 | 29.15 | 31.075 | 32 | 32.2
27.8 | 26.75 | 29 | 30.9 | 31.8714 | 32.05
27.9 | 26.625 | 28.8333 | 30.78 | 31.66 | 31.9
28 | 26.5 | 28.7 | 30.65 | 31.5 | 31.76
28.1 | 26.375 | 28.6 | 30.5 | 31.4 | 31.625
28.2 | 26.2625 | 28.4 | 30.4 | 31.225 | 31.5
28.3 | 26.1 | 28.3667 | 30.2 | 31.1 | 31.4
28.4 | 26 | 28.2 | 30.1 | 30.975 | 31.1857
28.5 | 25.9 | 28.05 | 29.9333 | 30.8 | 31.05
28.6 | 25.725 | 27.875 | 29.8 | 30.7 | 30.9143
28.7 | 25.6 | 27.75 | 29.625 | 30.575 | 30.8
28.8 | 25.5 | 27.625 | 29.5 | 30.425 | 30.64
28.9 | 25.4 | 27.5 | 29.4 | 30.275 | 30.58
29 | 25.2 | 27.3875 | 29.3 | 30.2 | 30.45
29.1 | 25.1 | 27.2625 | 29.15 | 30 | 30.3
29.2 | 24.95 | 27.075 | 29 | 29.9 | 30.15
29.3 | 24.8 | 26.9125 | 28.875 | 29.8 | 30
29.4 | 24.7 | 26.8833 | 28.75 | 29.7 | 29.85
29.5 | 24.6 | 26.68 | 28.625 | 29.5 | 29.7
29.6 | 24.45 | 26.525 | 28.5 | 29.4 | 29.6
29.7 | 24.4 | 26.4 | 28.35 | 29.28 | 29.5
29.8 | 24.275 | 26.35 | 28.2 | 29.15 | 29.375
29.9 | 24.1 | 26.225 | 28.05 | 29 | 29.225
30 | 24 | 26.1 | 27.9333 | 28.85 | 29.1

## SSD120 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 28.3 | 34.5 | 43.6 | 52.2 | 55.4
0.1 | 32.7 | 38.35 | 47.9 | 56.2333 | 59.475
0.2 | 40 | 44.95 | 54.5 | 62 | 64.7
0.3 | 48.575 | 53.95 | 61.7 | 68.9 | 71.6
0.4 | 56.475 | 61.575 | 68.7 | 74.9 | 77.16
0.5 | 63.325 | 67.7 | 73.85 | 79.375 | 81.5
0.6 | 69.125 | 73.1 | 78.625 | 83.45 | 85.3
0.7 | 74.025 | 77.525 | 82.4 | 86.55 | 88.2
0.8 | 78.3 | 81.45 | 85.7667 | 89.35 | 90.65
0.9 | 82 | 84.575 | 88.4 | 91.5 | 92.8
1 | 84.975 | 87.1 | 90.7 | 93.375 | 94.2667
1.1 | 87.825 | 89.5 | 92.5 | 94.7 | 95.7
1.2 | 89.95 | 91.8 | 94.225 | 96 | 96.7
1.3 | 91.8 | 93.35 | 95.425 | 97 | 97.55
1.4 | 93.4 | 94.825 | 96.675 | 98 | 98.25
1.5 | 94.8 | 96.05 | 97.5 | 98.45 | 98.85
1.6 | 96.1 | 97 | 98.275 | 98.95 | 99.2
1.7 | 97.1 | 97.875 | 98.775 | 99.35 | 99.7
1.8 | 97.95 | 98.4 | 99.25 | 99.675 | 99.84
1.9 | 98.5 | 98.96 | 99.58 | 99.825 | 99.95
2 | 99.15 | 99.45 | 99.8 | 99.96 | 100
2.1 | 99.4 | 99.7 | 99.9 | 100 | 100
2.2 | 99.725 | 99.85 | 100 | 100 | 100
2.3 | 99.875 | 100 | 100 | 99.925 | 99.9
2.4 | 100 | 100 | 100 | 99.82 | 99.8
2.5 | 100 | 100 | 99.9 | 99.6286 | 99.625
2.6 | 100 | 100 | 99.9 | 99.5 | 99.46
2.7 | 99.825 | 99.9 | 99.8 | 99.2875 | 99.2667
2.8 | 99.7 | 99.7556 | 99.575 | 99.125 | 99.025
2.9 | 99.5 | 99.5571 | 99.325 | 98.7875 | 98.85
3 | 99.3 | 99.3286 | 99.175 | 98.625 | 98.525
3.1 | 99.125 | 99.15 | 98.9 | 98.4 | 98.28
3.2 | 99 | 99 | 98.625 | 98.2 | 97.95
3.3 | 98.5333 | 98.625 | 98.4 | 97.9 | 97.725
3.4 | 98.2 | 98.375 | 98.0333 | 97.5 | 97.4667
3.5 | 97.9333 | 98.125 | 97.7333 | 97.25 | 97.0667
3.6 | 97.6 | 97.8 | 97.3667 | 97 | 96.8333
3.7 | 97.2 | 97.35 | 97.1 | 96.7 | 96.6
3.8 | 96.8 | 97.0714 | 96.725 | 96.3667 | 96.3
3.9 | 96.65 | 96.8 | 96.5333 | 96.1 | 96
4 | 96.2 | 96.4333 | 96.26 | 95.7667 | 95.6
4.1 | 95.65 | 96.15 | 95.8 | 95.4 | 95.275
4.2 | 95.1 | 95.7 | 95.5 | 95 | 95.025
4.3 | 94.76 | 95.35 | 95.0333 | 94.6667 | 94.675
4.4 | 94.525 | 95.02 | 94.725 | 94.4 | 94.325
4.5 | 93.95 | 94.65 | 94.5 | 94.12 | 94.0833
4.6 | 93.65 | 94.35 | 94.15 | 93.8667 | 93.86
4.7 | 93.175 | 93.9667 | 93.9 | 93.475 | 93.46
4.8 | 92.85 | 93.56 | 93.675 | 93.225 | 93.1
4.9 | 92.55 | 93.25 | 93.425 | 92.9 | 92.9
5 | 92.15 | 92.95 | 93.08 | 92.625 | 92.4
5.1 | 91.8 | 92.525 | 92.725 | 92.3 | 92.2
5.2 | 91.2 | 92.18 | 92.3333 | 92 | 91.875
5.3 | 90.925 | 91.8857 | 92 | 91.65 | 91.675
5.4 | 90.5 | 91.3857 | 91.65 | 91.35 | 91.35
5.5 | 90.05 | 91.15 | 91.35 | 91 | 90.975
5.6 | 89.7 | 90.94 | 91 | 90.86 | 90.725
5.7 | 89.3 | 90.6 | 90.8 | 90.5 | 90.4
5.8 | 89.06 | 90.2 | 90.4333 | 90.3 | 90.225
5.9 | 88.6 | 89.8 | 90.0875 | 89.8875 | 89.85
6 | 88.275 | 89.4111 | 89.775 | 89.575 | 89.5
6.1 | 88.025 | 89.175 | 89.525 | 89.325 | 89.3
6.2 | 87.525 | 88.8375 | 89.075 | 89.06 | 88.95
6.3 | 87.05 | 88.575 | 88.8625 | 88.7375 | 88.7
6.4 | 86.625 | 88.2667 | 88.6333 | 88.4333 | 88.4
6.5 | 86.15 | 87.8 | 88.275 | 88.15 | 88.05
6.6 | 85.8 | 87.5 | 87.9 | 87.9 | 87.775
6.7 | 85.45 | 87.25 | 87.5333 | 87.45 | 87.525
6.8 | 85.2 | 86.85 | 87.225 | 87.2222 | 87.2333
6.9 | 84.8 | 86.55 | 86.975 | 86.95 | 86.875
7 | 84.3 | 86.15 | 86.6375 | 86.6167 | 86.625
7.1 | 83.925 | 85.9 | 86.3714 | 86.2667 | 86.28
7.2 | 83.475 | 85.45 | 86 | 86 | 85.95
7.3 | 83.225 | 85.1 | 85.7889 | 85.7 | 85.65
7.4 | 82.78 | 84.75 | 85.4 | 85.375 | 85.38
7.5 | 82.55 | 84.5 | 85.1 | 85.1 | 85
7.6 | 82.2 | 84.15 | 84.85 | 84.75 | 84.8
7.7 | 81.8 | 83.925 | 84.55 | 84.5556 | 84.5
7.8 | 81.3 | 83.55 | 84.25 | 84.3125 | 84.2
7.9 | 80.975 | 83.125 | 83.975 | 84.025 | 84
8 | 80.575 | 82.7667 | 83.725 | 83.775 | 83.6667
8.1 | 80.25 | 82.525 | 83.375 | 83.4875 | 83.35
8.2 | 79.8125 | 82.2667 | 82.95 | 83.175 | 82.975
8.3 | 79.4375 | 81.8 | 82.8 | 82.95 | 82.74
8.4 | 79.2 | 81.575 | 82.45 | 82.6 | 82.5
8.5 | 78.8 | 81.2 | 82.1 | 82.3 | 82.3
8.6 | 78.3 | 80.85 | 81.82 | 82 | 81.9
8.7 | 78.04 | 80.6 | 81.6 | 81.6333 | 81.64
8.8 | 77.725 | 80.275 | 81.275 | 81.275 | 81.4
8.9 | 77.3 | 80.025 | 81.025 | 81.025 | 81.2
9 | 76.95 | 79.575 | 80.675 | 80.675 | 80.8667
9.1 | 76.65 | 79.15 | 80.34 | 80.45 | 80.6
9.2 | 76.25 | 79.0167 | 80.1 | 80.1333 | 80.3333
9.3 | 75.95 | 78.78 | 79.75 | 79.9 | 80.02
9.4 | 75.6 | 78.4 | 79.4 | 79.5429 | 79.725
9.5 | 75.25 | 78 | 79.1333 | 79.4 | 79.375
9.6 | 74.82 | 77.72 | 78.9 | 79.0333 | 79.1
9.7 | 74.5 | 77.3 | 78.5333 | 78.7667 | 78.85
9.8 | 74.15 | 76.9 | 78.25 | 78.5 | 78.425
9.9 | 73.975 | 76.68 | 77.975 | 78.275 | 78.3
10 | 73.55 | 76.5 | 77.725 | 78 | 78
10.1 | 73.26 | 76.25 | 77.475 | 77.675 | 77.7
10.2 | 72.85 | 75.9 | 77.15 | 77.35 | 77.55
10.3 | 72.5 | 75.4667 | 76.86 | 77.1 | 77.22
10.4 | 72.15 | 75.2 | 76.55 | 76.94 | 76.95
10.5 | 71.7 | 74.96 | 76.3 | 76.6 | 76.625
10.6 | 71.45 | 74.55 | 76 | 76.3 | 76.375
10.7 | 71.2 | 74.275 | 75.7333 | 76 | 76.125
10.8 | 70.96 | 73.85 | 75.44 | 75.7 | 75.8
10.9 | 70.475 | 73.675 | 75.2 | 75.5667 | 75.55
11 | 70.1 | 73.4667 | 74.8857 | 75.24 | 75.3333
11.1 | 69.85 | 73.175 | 74.7111 | 74.975 | 75.1
11.2 | 69.475 | 72.85 | 74.4125 | 74.725 | 74.75
11.3 | 69.225 | 72.5 | 74.125 | 74.5 | 74.5
11.4 | 68.875 | 72.175 | 73.8625 | 74.15 | 74.2667
11.5 | 68.6 | 71.875 | 73.525 | 73.9 | 73.875
11.6 | 68.25 | 71.7 | 73.2125 | 73.6 | 73.7
11.7 | 67.85 | 71.375 | 73.0125 | 73.325 | 73.35
11.8 | 67.6 | 71.05 | 72.6778 | 73.1 | 73.1
11.9 | 67.4 | 70.8 | 72.45 | 72.95 | 72.95
12 | 66.9667 | 70.5 | 72.1667 | 72.6857 | 72.7
12.1 | 66.7 | 70.1 | 71.92 | 72.35 | 72.35
12.2 | 66.44 | 69.85 | 71.5 | 72.075 | 72.05
12.3 | 66.2 | 69.6 | 71.375 | 71.9 | 71.825
12.4 | 65.7 | 69.3 | 71.125 | 71.55 | 71.575
12.5 | 65.45 | 69 | 70.825 | 71.325 | 71.34
12.6 | 65.2 | 68.75 | 70.5 | 71 | 71.025
12.7 | 64.9 | 68.4 | 70.325 | 70.75 | 70.875
12.8 | 64.5 | 68.15 | 70 | 70.4667 | 70.625
12.9 | 64.15 | 67.8 | 69.65 | 70.2 | 70.3
13 | 63.9 | 67.575 | 69.5 | 70.025 | 70.05
13.1 | 63.65 | 67.25 | 69.25 | 69.75 | 69.9
13.2 | 63.35 | 66.975 | 69.1 | 69.5 | 69.5333
13.3 | 63.1714 | 66.725 | 68.75 | 69.3 | 69.3
13.4 | 62.9 | 66.4667 | 68.525 | 68.95 | 69.12
13.5 | 62.45 | 66.2 | 68.175 | 68.8 | 68.825
13.6 | 62.15 | 65.925 | 67.925 | 68.3875 | 68.55
13.7 | 62.05 | 65.675 | 67.6 | 68.225 | 68.3
13.8 | 61.7 | 65.425 | 67.48 | 67.9125 | 68.175
13.9 | 61.35 | 65.175 | 67.2 | 67.7125 | 67.75
14 | 61 | 64.85 | 66.85 | 67.475 | 67.475
14.1 | 60.86 | 64.58 | 66.575 | 67.18 | 67.28
14.2 | 60.5625 | 64.3 | 66.4 | 67 | 67.05
14.3 | 60.225 | 64 | 66.1 | 66.7286 | 66.9
14.4 | 59.9333 | 63.72 | 65.75 | 66.4857 | 66.5429
14.5 | 59.65 | 63.45 | 65.5 | 66.25 | 66.3
14.6 | 59.38 | 63.2 | 65.275 | 66 | 66.0333
14.7 | 59.15 | 62.9 | 65.025 | 65.7 | 65.75
14.8 | 58.85 | 62.72 | 64.78 | 65.525 | 65.6
14.9 | 58.6 | 62.5 | 64.525 | 65.175 | 65.35
15 | 58.275 | 62.2 | 64.3 | 64.925 | 65.2
15.1 | 58 | 61.8429 | 64.0286 | 64.85 | 64.95
15.2 | 57.6333 | 61.66 | 63.88 | 64.525 | 64.625
15.3 | 57.3 | 61.35 | 63.625 | 64.3 | 64.375
15.4 | 57.1 | 61.075 | 63.35 | 64.05 | 64.2
15.5 | 56.8333 | 60.9 | 63.15 | 63.9 | 63.975
15.6 | 56.6333 | 60.55 | 62.8 | 63.6 | 63.625
15.7 | 56.325 | 60.325 | 62.6 | 63.3 | 63.475
15.8 | 56.2 | 60.1333 | 62.4143 | 63.1 | 63.26
15.9 | 55.925 | 59.8 | 62.15 | 62.875 | 63
16 | 55.65 | 59.65 | 62 | 62.625 | 62.8
16.1 | 55.4 | 59.3 | 61.725 | 62.375 | 62.475
16.2 | 55.175 | 59.2 | 61.4333 | 62.2125 | 62.225
16.3 | 54.925 | 58.9667 | 61.25 | 61.975 | 62.1
16.4 | 54.775 | 58.7 | 60.9571 | 61.725 | 61.85
16.5 | 54.5 | 58.4 | 60.775 | 61.5 | 61.6
16.6 | 54.15 | 58.1 | 60.54 | 61.25 | 61.45
16.7 | 53.9 | 58 | 60.3 | 61 | 61.2
16.8 | 53.675 | 57.66 | 60.15 | 60.8 | 61
16.9 | 53.425 | 57.4 | 59.8 | 60.55 | 60.775
17 | 53.2 | 57.075 | 59.55 | 60.4 | 60.525
17.1 | 52.8 | 56.9 | 59.4 | 60.16 | 60.275
17.2 | 52.66 | 56.65 | 59.15 | 60 | 60.04
17.3 | 52.3667 | 56.5111 | 58.9 | 59.725 | 59.8
17.4 | 52.2 | 56.2875 | 58.6 | 59.4667 | 59.6667
17.5 | 51.9 | 56.125 | 58.4 | 59.3 | 59.5
17.6 | 51.7 | 55.8125 | 58.2 | 59.05 | 59.175
17.7 | 51.5 | 55.55 | 58 | 58.8 | 59.0111
17.8 | 51.2 | 55.2 | 57.7 | 58.6 | 58.7333
17.9 | 51 | 55.05 | 57.55 | 58.42 | 58.525
18 | 50.7333 | 54.8 | 57.3 | 58.2 | 58.3375
18.1 | 50.6 | 54.675 | 57.04 | 58 | 58.125
18.2 | 50.25 | 54.42 | 56.8 | 57.775 | 57.8833
18.3 | 50 | 54.25 | 56.6 | 57.625 | 57.675
18.4 | 49.8 | 53.925 | 56.42 | 57.36 | 57.4667
18.5 | 49.66 | 53.675 | 56.26 | 57.125 | 57.2667
18.6 | 49.4571 | 53.44 | 56.025 | 56.875 | 57
18.7 | 49.2 | 53.28 | 55.7667 | 56.625 | 56.775
18.8 | 48.925 | 53.06 | 55.55 | 56.45 | 56.6
18.9 | 48.675 | 52.7 | 55.3333 | 56.25 | 56.375
19 | 48.4667 | 52.6 | 55.1 | 56 | 56.225
19.1 | 48.3 | 52.35 | 54.9333 | 55.775 | 55.975
19.2 | 48.05 | 52.1 | 54.7 | 55.625 | 55.725
19.3 | 47.8 | 51.86 | 54.46 | 55.375 | 55.5
19.4 | 47.58 | 51.625 | 54.3 | 55.14 | 55.35
19.5 | 47.36 | 51.4 | 54.0556 | 54.9 | 55.2
19.6 | 47.14 | 51.2 | 53.825 | 54.7 | 54.95
19.7 | 46.9 | 51 | 53.5571 | 54.55 | 54.6
19.8 | 46.64 | 50.7333 | 53.425 | 54.3 | 54.5
19.9 | 46.5111 | 50.4857 | 53.06 | 54.1 | 54.325
20 | 46.2625 | 50.2714 | 52.96 | 54 | 54.2
20.1 | 45.925 | 50.0667 | 52.75 | 53.775 | 53.975
20.2 | 45.775 | 49.85 | 52.5 | 53.45 | 53.8
20.3 | 45.625 | 49.7 | 52.35 | 53.3 | 53.575
20.4 | 45.4 | 49.4556 | 52.2 | 53.14 | 53.425
20.5 | 45.3 | 49.225 | 51.9 | 52.875 | 53.2
20.6 | 45.05 | 49.1167 | 51.8 | 52.7 | 52.94
20.7 | 44.75 | 48.9167 | 51.625 | 52.55 | 52.8
20.8 | 44.525 | 48.6714 | 51.375 | 52.3 | 52.58
20.9 | 44.4 | 48.5 | 51.125 | 52.15 | 52.325
21 | 44.275 | 48.2 | 50.975 | 52 | 52.1
21.1 | 44.025 | 48 | 50.6333 | 51.8 | 51.95
21.2 | 43.78 | 47.8 | 50.5 | 51.5 | 51.8
21.3 | 43.65 | 47.7 | 50.4 | 51.25 | 51.64
21.4 | 43.4 | 47.475 | 50.15 | 51.1 | 51.4
21.5 | 43.2 | 47.15 | 49.9 | 51 | 51.25
21.6 | 43 | 47 | 49.7667 | 50.85 | 50.9
21.7 | 42.85 | 46.85 | 49.525 | 50.65 | 50.75
21.8 | 42.52 | 46.6 | 49.375 | 50.35 | 50.6
21.9 | 42.375 | 46.35 | 49.225 | 50.125 | 50.375
22 | 42.225 | 46.2 | 49 | 49.975 | 50.22
22.1 | 41.975 | 46.0333 | 48.825 | 49.74 | 50.075
22.2 | 41.725 | 45.8 | 48.6 | 49.6 | 49.925
22.3 | 41.6 | 45.65 | 48.44 | 49.4 | 49.675
22.4 | 41.475 | 45.4 | 48.2778 | 49.2 | 49.5
22.5 | 41.225 | 45.275 | 48.0429 | 49.05 | 49.3
22.6 | 41 | 45.125 | 47.8 | 48.9 | 49.1
22.7 | 40.825 | 44.9667 | 47.65 | 48.75 | 48.9
22.8 | 40.725 | 44.7 | 47.4 | 48.4 | 48.8
22.9 | 40.575 | 44.575 | 47.3 | 48.26 | 48.5875
23 | 40.325 | 44.3 | 47.175 | 48.125 | 48.425
23.1 | 40.1167 | 44.1 | 47 | 47.975 | 48.175
23.2 | 39.9 | 44 | 46.875 | 47.8 | 48.0125
23.3 | 39.7625 | 43.775 | 46.525 | 47.6 | 47.7889
23.4 | 39.6 | 43.6 | 46.3 | 47.44 | 47.6714
23.5 | 39.425 | 43.4 | 46.225 | 47.2 | 47.475
23.6 | 39.175 | 43.24 | 45.95 | 47.05 | 47.34
23.7 | 39.025 | 43 | 45.725 | 46.9 | 47.2
23.8 | 38.9 | 42.85 | 45.675 | 46.6667 | 46.96
23.9 | 38.7 | 42.7 | 45.475 | 46.5 | 46.825
24 | 38.6 | 42.45 | 45.225 | 46.3667 | 46.675
24.1 | 38.4 | 42.3 | 45.1167 | 46.2 | 46.5
24.2 | 38.175 | 42.075 | 44.8875 | 46.1 | 46.28
24.3 | 38.025 | 41.9 | 44.7625 | 45.8 | 46.15
24.4 | 37.8 | 41.7375 | 44.575 | 45.7 | 46
24.5 | 37.7 | 41.6125 | 44.325 | 45.5 | 45.75
24.6 | 37.575 | 41.4 | 44.2 | 45.3 | 45.6
24.7 | 37.325 | 41.125 | 44.05 | 45.16 | 45.35
24.8 | 37.175 | 41.0167 | 43.9 | 44.925 | 45.2
24.9 | 37 | 40.875 | 43.7 | 44.8 | 45.075
25 | 36.9 | 40.7 | 43.6 | 44.625 | 44.825
25.1 | 36.625 | 40.575 | 43.3667 | 44.5 | 44.675
25.2 | 36.5 | 40.425 | 43.1 | 44.3 | 44.6
25.3 | 36.3 | 40.275 | 43 | 44.1 | 44.4333
25.4 | 36.125 | 40.025 | 42.825 | 44 | 44.275
25.5 | 35.9 | 39.7857 | 42.7 | 43.875 | 44.1
25.6 | 35.75 | 39.62 | 42.5 | 43.6 | 43.88
25.7 | 35.6 | 39.48 | 42.3 | 43.4 | 43.725
25.8 | 35.5 | 39.25 | 42.2 | 43.2 | 43.6
25.9 | 35.4 | 39.1 | 41.95 | 43.1 | 43.375
26 | 35.2 | 38.95 | 41.8 | 42.925 | 43.225
26.1 | 35.0429 | 38.8 | 41.65 | 42.8 | 43
26.2 | 34.9 | 38.65 | 41.5 | 42.6625 | 42.875
26.3 | 34.8 | 38.5 | 41.3 | 42.475 | 42.7
26.4 | 34.625 | 38.375 | 41.2 | 42.3125 | 42.575
26.5 | 34.475 | 38.225 | 41 | 42.0875 | 42.35
26.6 | 34.325 | 38.075 | 40.825 | 41.9667 | 42.2
26.7 | 34.2 | 37.8 | 40.675 | 41.8 | 42.1
26.8 | 34.05 | 37.7 | 40.525 | 41.55 | 41.9
26.9 | 33.8 | 37.54 | 40.3286 | 41.4 | 41.775
27 | 33.65 | 37.3 | 40.18 | 41.275 | 41.625
27.1 | 33.5 | 37.15 | 40.1 | 41.15 | 41.48
27.2 | 33.25 | 37 | 39.9 | 41.1 | 41.3
27.3 | 33.2 | 36.95 | 39.75 | 40.875 | 41.1333
27.4 | 33.075 | 36.8333 | 39.62 | 40.7 | 41
27.5 | 32.925 | 36.6 | 39.4667 | 40.575 | 40.8
27.6 | 32.6875 | 36.4714 | 39.2333 | 40.325 | 40.7
27.7 | 32.5571 | 36.3 | 39.1 | 40.175 | 40.575
27.8 | 32.4 | 36.1875 | 39 | 40.025 | 40.35
27.9 | 32.3 | 36.025 | 38.85 | 39.9 | 40.2125
28 | 32.2 | 35.8 | 38.7 | 39.8 | 40.06
28.1 | 32.075 | 35.65 | 38.6 | 39.6 | 39.9125
28.2 | 31.9 | 35.5 | 38.4 | 39.5 | 39.8
28.3 | 31.725 | 35.35 | 38.2 | 39.4 | 39.575
28.4 | 31.6 | 35.2 | 38.05 | 39.2 | 39.45
28.5 | 31.4429 | 35.1 | 37.9 | 39 | 39.325
28.6 | 31.3 | 34.9 | 37.76 | 38.8 | 39.1
28.7 | 31.2 | 34.725 | 37.625 | 38.7 | 38.975
28.8 | 30.975 | 34.6 | 37.475 | 38.6 | 38.8556
28.9 | 30.825 | 34.425 | 37.325 | 38.4 | 38.7375
29 | 30.8 | 34.3 | 37.175 | 38.25 | 38.525
29.1 | 30.6 | 34.2 | 37.04 | 38.2 | 38.4
29.2 | 30.5 | 34.1 | 36.9 | 38.0333 | 38.3
29.3 | 30.3 | 33.9 | 36.75 | 37.88 | 38.15
29.4 | 30.2 | 33.8 | 36.6 | 37.7 | 38
29.5 | 30.0667 | 33.65 | 36.45 | 37.6 | 37.85
29.6 | 29.925 | 33.52 | 36.3 | 37.425 | 37.7
29.7 | 29.8 | 33.375 | 36.175 | 37.275 | 37.6
29.8 | 29.7 | 33.225 | 36.0125 | 37.15 | 37.425
29.9 | 29.6 | 33.1 | 35.8 | 37 | 37.26
30 | 29.425 | 32.975 | 35.725 | 36.9 | 37.125

## Wedge 15 Lateral Profiles
Lateral Profiles
Wedge 15
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:2.4 cm FieldSize:4x4 cm2 | Depth:2.4 cm FieldSize:10x10 cm2 | Depth:2.4 cm FieldSize:20x20 cm2 | Depth:2.4 cm FieldSize:30x30 cm2 | Depth:2.4 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.9 | [8 empty] | 5.0991 | 5.9 |  |  |  | 6.85 | 7.7 |  |  |  | 11.36 | 12.3 |  |  |  | 16.7 | 17.76
-22.8 | [8 empty] | 5.1991 | 6 |  |  |  | 7 | 7.775 |  |  |  | 11.5 | 12.52 |  |  |  | 17 | 17.96
-22.7 | [8 empty] | 5.242 | 6 |  |  |  | 7.05 | 7.8 |  |  |  | 11.72 | 12.7 |  |  |  | 17.2 | 18.3333
-22.6 | [8 empty] | 5.2991 | 6.1 |  |  |  | 7.2 | 8 |  |  |  | 11.9 | 12.88 |  |  |  | 17.48 | 18.6
-22.5 | [8 empty] | 5.3991 | 6.2 |  |  |  | 7.3 | 8.1 |  |  |  | 12.0833 | 13.1 |  |  |  | 17.7 | 18.8333
-22.4 | [8 empty] | 5.4824 | 6.2667 |  |  |  | 7.4 | 8.22 |  |  |  | 12.3 | 13.3 |  |  |  | 17.9333 | 19.1
-22.3 |  |  |  | 5.3009 | 6.1 |  |  |  | 5.5657 | 6.3286 |  |  |  | 7.5429 | 8.4 |  |  |  | 12.45 | 13.4667 |  |  |  | 18.2333 | 19.2667
-22.2 |  |  |  | 5.4009 | 6.1667 |  |  |  | 5.6324 | 6.4833 |  |  |  | 7.6857 | 8.5 |  |  |  | 12.6 | 13.65 |  |  |  | 18.55 | 19.6333
-22.1 |  |  |  | 5.5009 | 6.2 |  |  |  | 5.7115 | 6.5 |  |  |  | 7.8143 | 8.6 |  |  |  | 12.775 | 13.8167 |  |  |  | 18.76 | 19.94
-22 |  |  |  | 5.5259 | 6.3 |  |  |  | 5.799 | 6.6286 |  |  |  | 7.95 | 8.7667 |  |  |  | 13.02 | 14.0667 |  |  | 6.3 | 19.0333 | 20.18
-21.9 |  |  |  | 5.6009 | 6.4 |  |  |  | 5.9157 | 6.7 |  |  |  | 8.0857 | 8.9 |  |  |  | 13.2333 | 14.26 |  |  | 6.4 | 19.3667 | 20.45
-21.8 |  |  |  | 5.7009 | 6.5 |  |  |  | 5.999 | 6.8 |  |  |  | 8.2571 | 9.04 |  |  |  | 13.475 | 14.5 |  |  | 6.4833 | 19.65 | 20.7
-21.7 |  |  |  | 5.781 | 6.6 |  |  |  | 6.0847 | 6.9 |  |  |  | 8.4 | 9.25 |  |  |  | 13.6667 | 14.72 |  |  | 6.55 | 19.9 | 21.0667
-21.6 |  |  |  | 5.801 | 6.62 |  |  |  | 6.2275 | 7.05 |  |  |  | 8.4571 | 9.3 |  |  |  | 13.8333 | 14.9 |  |  | 6.6 | 20.2429 | 21.2857
-21.5 |  |  |  | 5.921 | 6.8 |  |  |  | 6.2989 | 7.2 |  |  |  | 8.5875 | 9.4333 |  |  |  | 14.05 | 15.1 |  |  | 6.65 | 20.5333 | 21.7167
-21.4 |  |  |  | 6.021 | 6.85 |  |  |  | 6.3989 | 7.25 |  |  |  | 8.8 | 9.675 |  |  |  | 14.3 | 15.375 |  |  | 6.7429 | 20.85 | 21.9667
-21.3 |  |  |  | 6.101 | 6.9 |  |  |  | 6.4989 | 7.375 |  |  |  | 8.9286 | 9.8 |  |  |  | 14.55 | 15.6 |  |  | 6.8 | 21.175 | 22.3333
-21.2 |  |  |  | 6.251 | 7.0667 |  |  |  | 6.6364 | 7.5 |  |  |  | 9 | 9.9 |  |  |  | 14.775 | 15.84 |  |  | 6.9667 | 21.6 | 22.6
-21.1 |  |  |  | 6.3511 | 7.175 |  |  |  | 6.7614 | 7.6429 |  |  |  | 9.1571 | 10.025 |  |  |  | 15 | 16.05 |  |  | 7 | 21.92 | 23.1
-21 |  |  |  | 6.4411 | 7.2 |  |  |  | 6.8989 | 7.8 |  |  |  | 9.3167 | 10.2 |  |  |  | 15.3333 | 16.3 |  |  | 7.1 | 22.35 | 23.46
-20.9 |  |  |  | 6.5511 | 7.4 |  |  |  | 7.0417 | 7.8375 |  |  |  | 9.4714 | 10.35 |  |  |  | 15.5 | 16.575 |  |  | 7.2 | 22.7667 | 23.8333
-20.8 |  |  |  | 6.7011 | 7.5 |  |  |  | 7.1845 | 7.9714 |  |  |  | 9.6333 | 10.5 |  |  |  | 15.84 | 16.9 |  |  | 7.2857 | 23.24 | 24.3
-20.7 |  |  |  | 6.8011 | 7.6 |  |  |  | 7.2788 | 8.1375 |  |  |  | 9.8 | 10.7 |  |  |  | 16.05 | 17.1 |  |  | 7.4 | 23.82 | 24.84
-20.6 |  |  |  | 6.9178 | 7.8 |  |  |  | 7.3988 | 8.2 |  |  |  | 10.0143 | 10.9 |  |  |  | 16.375 | 17.42 |  |  | 7.4167 | 24.38 | 25.5
-20.5 |  |  |  | 7.0012 | 7.9 |  |  |  | 7.5187 | 8.3667 |  |  |  | 10.18 | 11.0833 |  |  |  | 16.6 | 17.7 |  |  | 7.6 | 25.12 | 26.0833
-20.4 |  |  |  | 7.1679 | 8 |  |  |  | 7.6844 | 8.5 |  |  |  | 10.3333 | 11.225 |  |  |  | 16.925 | 18 |  |  | 7.65 | 25.98 | 26.9
-20.3 |  |  |  | 7.3012 | 8.14 |  |  |  | 7.7654 | 8.6 |  |  |  | 10.5167 | 11.4 |  |  | 5.4 | 17.18 | 18.275 |  |  | 7.8 | 27 | 28.05
-20.2 |  |  |  | 7.4012 | 8.26 |  |  |  | 7.8987 | 8.7167 |  |  |  | 10.7 | 11.6 |  |  | 5.4 | 17.48 | 18.58 |  |  | 7.9 | 28.26 | 29.35
-20.1 |  |  |  | 7.5012 | 8.36 |  |  |  | 7.9987 | 8.8833 |  |  |  | 10.88 | 11.75 |  |  | 5.5 | 17.85 | 18.925 |  |  | 8 | 30 | 31
-20 |  |  |  | 7.6513 | 8.525 |  |  |  | 8.1786 | 9 |  |  |  | 11.1143 | 12 |  |  | 5.5 | 18.1 | 19.15 |  |  | 8 | 32.3333 | 33.2
-19.9 |  |  |  | 7.8013 | 8.6 |  |  |  | 8.3486 | 9.16 |  |  |  | 11.3125 | 12.2 |  |  | 5.6 | 18.45 | 19.55 |  |  | 8.1 | 35.875 | 36.55
-19.8 |  |  |  | 7.918 | 8.75 |  |  |  | 8.4153 | 9.3 |  |  |  | 11.5 | 12.35 |  |  | 5.7 | 18.8 | 19.825 |  |  | 8.2333 | 40.2 | 40.6
-19.7 |  |  |  | 8.0013 | 8.9 |  |  |  | 8.5819 | 9.4 |  |  |  | 11.7143 | 12.56 |  |  | 5.7 | 19.1429 | 20.2 |  |  | 8.3 | 45.7 | 46.4667
-19.6 |  |  |  | 8.2014 | 9 |  |  |  | 8.7485 | 9.5714 |  |  |  | 11.8667 | 12.775 |  |  | 5.8 | 19.6 | 20.7 |  |  | 8.3667 | 52.65 | 53.0286
-19.5 |  |  |  | 8.3014 | 9.2 |  |  |  | 8.9318 | 9.7333 |  |  |  | 12.16 | 13.025 |  |  | 5.9 | 20 | 21.0667 |  |  | 8.525 | 59.8667 | 60.5
-19.4 |  |  |  | 8.4681 | 9.3 |  |  |  | 9.0842 | 9.9833 |  |  |  | 12.375 | 13.25 |  |  | 5.9 | 20.5 | 21.475 |  |  | 8.66 | 67.5333 | 67.46
-19.3 |  |  |  | 8.6014 | 9.5 |  |  |  | 9.2485 | 10.0571 |  |  |  | 12.6429 | 13.5 |  |  | 6 | 20.94 | 21.98 |  |  | 8.78 | 74.68 | 74.1333
-19.2 |  |  |  | 8.7265 | 9.58 |  |  |  | 9.3318 | 10.24 |  |  |  | 12.88 | 13.72 |  |  | 6.1 | 21.5143 | 22.62 |  |  | 8.875 | 80.5 | 80.45
-19.1 |  |  |  | 8.8615 | 9.72 |  |  |  | 9.5817 | 10.4167 |  |  |  | 13.1167 | 14.04 |  |  | 6.2 | 22.1 | 23.22 |  |  | 9 | 85.05 | 84.68
-19 |  |  |  | 9.0415 | 9.96 |  |  |  | 9.7484 | 10.6 |  |  |  | 13.3667 | 14.275 |  |  | 6.3 | 22.8333 | 23.85 |  |  | 9.1 | 88.4 | 88.16
-18.9 |  |  |  | 9.1682 | 10.1 |  |  |  | 9.9269 | 10.8 |  |  |  | 13.5857 | 14.48 |  |  | 6.3333 | 23.7 | 24.72 |  |  | 9.2 | 90.8 | 90.6
-18.8 |  |  |  | 9.4216 | 10.22 |  |  |  | 10.165 | 11.0143 |  |  |  | 13.9167 | 14.8 |  |  | 6.4 | 24.9 | 25.8833 |  |  | 9.36 | 92.4667 | 92.2
-18.7 |  |  |  | 9.5266 | 10.4 |  |  |  | 10.3583 | 11.12 |  |  |  | 14.1667 | 15.05 |  |  | 6.5 | 26.22 | 27.2 |  |  | 9.575 | 93.7333 | 93.3667
-18.6 |  |  |  | 9.6616 | 10.6 |  |  |  | 10.5149 | 11.4167 |  |  | 4.2 | 14.4667 | 15.34 |  |  | 6.6 | 28.12 | 29.125 |  |  | 9.6 | 94.7 | 94.34
-18.5 |  |  |  | 9.9017 | 10.72 |  |  |  | 10.6982 | 11.6 |  |  | 4.3 | 14.8 | 15.65 |  |  | 6.7 | 30.85 | 31.62 |  |  | 9.7333 | 95.45 | 94.95
-18.4 |  |  |  | 10.0267 | 10.925 |  |  |  | 10.9125 | 11.8667 |  |  | 4.32 | 15.0857 | 15.95 |  |  | 6.8 | 34.44 | 35.54 |  |  | 9.8167 | 95.8667 | 95.5
-18.3 |  |  |  | 10.2267 | 11.075 |  |  |  | 11.1148 | 12.04 |  |  | 4.4 | 15.4 | 16.3333 |  |  | 6.8333 | 39.85 | 40.62 |  |  | 10 | 96.5333 | 95.9333
-18.2 |  |  |  | 10.4351 | 11.3 |  |  |  | 11.3648 | 12.2333 |  |  | 4.5 | 15.7333 | 16.7 |  |  | 7 | 46.85 | 47.18 |  |  | 10.2 | 96.96 | 96.5167
-18.1 |  |  |  | 10.6418 | 11.5 |  |  |  | 11.6314 | 12.5143 |  |  | 4.5 | 16.1333 | 17.05 |  |  | 7.1 | 54.9 | 55.32 |  |  | 10.325 | 97.28 | 96.88
-18 |  |  |  | 10.7768 | 11.68 |  |  |  | 11.8266 | 12.76 |  |  | 4.575 | 16.55 | 17.425 |  |  | 7.2 | 63.6 | 64 |  |  | 10.475 | 97.66 | 97.15
-17.9 |  |  |  | 10.9818 | 11.9 |  |  |  | 12.098 | 12.9429 |  |  | 4.6 | 17.02 | 17.85 |  |  | 7.3 | 72.2667 | 72.5 |  |  | 10.6667 | 97.85 | 97.4167
-17.8 |  |  |  | 11.2019 | 12.15 |  |  | 3.675 | 12.2647 | 13.2 |  |  | 4.7 | 17.5 | 18.35 |  |  | 7.3833 | 80.22 | 80.575 |  |  | 10.8 | 98.1286 | 97.5833
-17.7 |  |  |  | 11.4269 | 12.36 |  |  | 3.7 | 12.5646 | 13.5 |  |  | 4.7667 | 18.0667 | 18.9 |  |  | 7.5 | 87.14 | 86.66 |  |  | 10.94 | 98.3667 | 97.8333
-17.6 |  |  |  | 11.677 | 12.5333 |  |  | 3.8 | 12.8693 | 13.7571 |  |  | 4.8 | 18.6667 | 19.5667 |  |  | 7.6667 | 92.1667 | 91.8 |  |  | 11.06 | 98.44 | 98
-17.5 |  |  |  | 11.882 | 12.7 |  |  | 3.8 | 13.1645 | 14.04 |  |  | 4.9 | 19.5 | 20.4 |  |  | 7.75 | 95.4143 | 95.3 |  |  | 11.15 | 98.6333 | 98.2
-17.4 |  |  |  | 12.022 | 13.02 |  |  | 3.9 | 13.4478 | 14.34 |  |  | 4.9 | 20.45 | 21.3 |  |  | 7.8667 | 97.86 | 97.46 |  |  | 11.44 | 98.82 | 98.3
-17.3 |  |  |  | 12.2771 | 13.25 |  |  | 3.9 | 13.7644 | 14.65 |  |  | 5 | 21.5875 | 22.46 |  |  | 8 | 99.3571 | 99.24 |  |  | 11.6333 | 99 | 98.5333
-17.2 |  |  | 3.9 | 12.6021 | 13.42 |  |  | 3.9 | 14.1976 | 15.1 |  |  | 5.0333 | 23.0714 | 23.92 |  |  | 8.1333 | 100.5 | 100.1833 |  |  | 11.8 | 99.15 | 98.7333
-17.1 |  |  | 4 | 12.8271 | 13.74 |  |  | 4 | 14.5976 | 15.5 |  |  | 5.1 | 25.4 | 25.9333 |  |  | 8.2 | 101.5667 | 101 |  |  | 12.08 | 99.2167 | 98.8
-17 |  |  | 4 | 13.1022 | 14 |  |  | 4 | 15.0689 | 15.9571 |  |  | 5.2 | 28.4667 | 28.78 |  |  | 8.4 | 102.2 | 101.7167 |  |  | 12.2667 | 99.3 | 98.9
-16.9 |  |  | 4.1 | 13.4022 | 14.32 |  |  | 4.1 | 15.6974 | 16.54 |  |  | 5.3 | 33 | 33.6 |  |  | 8.5 | 102.6429 | 102.25 |  |  | 12.4 | 99.4 | 99
-16.8 |  |  | 4.1 | 13.6823 | 14.7 |  |  | 4.1 | 16.2973 | 17.2 |  |  | 5.3667 | 39.98 | 40 |  |  | 8.7 | 103.2 | 102.64 |  |  | 12.58 | 99.5 | 99.1
-16.7 |  |  | 4.14 | 14.1424 | 15 |  |  | 4.18 | 17.0972 | 18 |  |  | 5.4 | 48.4667 | 48.64 |  |  | 8.8 | 103.52 | 103 |  |  | 12.78 | 99.6429 | 99.2167
-16.6 |  |  | 4.2 | 14.5024 | 15.4 |  |  | 4.2333 | 18.1803 | 19.1 |  |  | 5.5 | 59.1 | 58.1333 |  |  | 8.92 | 103.8 | 103.2667 |  |  | 12.96 | 99.72 | 99.3
-16.5 |  |  | 4.3 | 15.0358 | 16 |  |  | 4.3 | 19.4567 | 20.4667 |  |  | 5.6167 | 69.2833 | 69.2 |  |  | 9.15 | 103.9667 | 103.6 |  |  | 13.1667 | 99.8333 | 99.4
-16.4 |  |  | 4.3 | 15.6526 | 16.52 |  |  | 4.4 | 21.3465 | 22.25 |  |  | 5.7 | 80.2 | 79.35 |  |  | 9.3 | 104.15 | 103.725 |  |  | 13.3833 | 99.9 | 99.5
-16.3 |  |  | 4.4 | 16.486 | 17.32 |  |  | 4.4 | 24.1627 | 24.9167 |  |  | 5.8 | 89.375 | 89.2833 |  |  | 9.4 | 104.4 | 103.9 |  |  | 13.6333 | 100 | 99.6
-16.2 |  |  | 4.4 | 17.4029 | 18.2 |  |  | 4.5 | 28.2381 | 29.2667 |  |  | 5.9 | 97.2143 | 96.675 |  |  | 9.525 | 104.5 | 104.0167 |  |  | 13.875 | 100 | 99.7
-16.1 |  |  | 4.5 | 18.8531 | 19.5 |  |  | 4.5 | 34.4942 | 35.54 |  |  | 6 | 102.44 | 102 |  |  | 9.7333 | 104.6 | 104.1833 |  |  | 14.1 | 100.1 | 99.78
-16 |  |  | 4.5 | 20.7035 | 21.475 |  |  | 4.6 | 42.9095 | 44.72 |  |  | 6.05 | 106 | 105.3 |  |  | 9.9714 | 104.7 | 104.2833 |  |  | 14.32 | 100.2 | 99.8
-15.9 |  |  | 4.6 | 23.554 | 24.1 |  |  | 4.7 | 53.791 | 54.7667 |  |  | 6.18 | 108.0833 | 107.65 |  |  | 10.1 | 104.74 | 104.36 |  |  | 14.54 | 100.2 | 99.9
-15.8 |  |  | 4.7 | 28.1547 | 28.275 |  |  | 4.7 | 65.4224 | 66.5286 |  |  | 6.28 | 109.5714 | 109 |  |  | 10.2667 | 104.82 | 104.4 |  |  | 14.84 | 100.3 | 99.9667
-15.7 |  |  | 4.7 | 35.1059 | 35.125 |  |  | 4.8 | 77.4871 | 78.3 |  |  | 6.4 | 110.58 | 110 |  |  | 10.4667 | 104.9 | 104.48 |  |  | 15.2 | 100.4 | 100
-15.6 |  |  | 4.7833 | 45.4276 | 44.68 |  |  | 4.9 | 89.5351 | 89.6 |  |  | 6.4833 | 111.3667 | 110.7 |  |  | 10.6857 | 104.9 | 104.5 |  |  | 15.4 | 100.4 | 100.0167
-15.5 |  |  | 4.825 | 56.9095 | 56.3 |  |  | 5 | 99.2032 | 98.75 |  |  | 6.65 | 111.8667 | 111.26 |  |  | 10.85 | 105 | 104.6 |  | 3.6 | 15.62 | 100.4833 | 100.1833
-15.4 |  |  | 4.9 | 70.5368 | 69.32 |  |  | 5.05 | 105.8024 | 105.6667 |  |  | 6.72 | 112.2143 | 111.7 |  |  | 11.1 | 105 | 104.6 |  | 3.6 | 15.84 | 100.5667 | 100.2
-15.3 |  |  | 5 | 83.914 | 82.4 |  |  | 5.1167 | 110.5817 | 110.3 |  |  | 6.84 | 112.5 | 111.95 |  |  | 11.28 | 105.1 | 104.7 |  | 3.6 | 16.075 | 100.6 | 100.3
-15.2 |  |  | 5.0667 | 94.6158 | 93.6 |  |  | 5.2 | 113.3953 | 113 |  |  | 7.075 | 112.7 | 112.1 |  |  | 11.425 | 105.1 | 104.7 |  | 3.6 | 16.3 | 100.7333 | 100.32
-15.1 |  |  | 5.18 | 104.617 | 103.175 |  |  | 5.3 | 115.0143 | 114.8 |  |  | 7.1 | 112.8 | 112.2 |  |  | 11.65 | 105.15 | 104.7 |  | 3.6 | 16.75 | 100.8 | 100.5
-15 |  |  | 5.2 | 111.5584 | 110.7 |  |  | 5.4 | 116.131 | 115.7571 |  |  | 7.25 | 112.8667 | 112.3 |  |  | 11.86 | 105.2 | 104.7 |  | 3.7 | 16.88 | 100.8667 | 100.6
-14.9 |  |  | 5.35 | 115.739 | 114.8 |  |  | 5.5 | 116.9308 | 116.6 |  |  | 7.38 | 112.9 | 112.3 |  |  | 12.075 | 105.2 | 104.7 |  | 3.7 | 17.25 | 100.9667 | 100.7
-14.8 |  |  | 5.4 | 118.02 | 117.16 |  |  | 5.6 | 117.38 | 116.9667 |  |  | 7.4667 | 112.9 | 112.4 |  |  | 12.3 | 105.2 | 104.7 |  | 3.8 | 17.7 | 101 | 100.8
-14.7 |  |  | 5.5 | 119.37 | 118.5 |  |  | 5.7 | 117.6467 | 117.28 |  |  | 7.6 | 112.9 | 112.4 |  |  | 12.6 | 105.2 | 104.8 |  | 3.8 | 18 | 101.16 | 100.9
-14.6 |  |  | 5.6 | 120.17 | 119.38 |  |  | 5.8 | 117.83 | 117.44 |  |  | 7.825 | 112.8 | 112.36 |  |  | 12.825 | 105.2 | 104.8 |  | 3.9 | 18.4 | 101.2667 | 101.075
-14.5 |  |  | 5.7 | 120.595 | 119.9 |  |  | 5.9 | 117.98 | 117.5 |  |  | 8 | 112.8 | 112.3 |  |  | 13.075 | 105.2 | 104.8 |  | 3.95 | 18.76 | 101.4 | 101.1
-14.4 |  |  | 5.8 | 120.82 | 120.15 |  |  | 6 | 117.98 | 117.5 |  |  | 8.2 | 112.7 | 112.2 |  |  | 13.34 | 105.14 | 104.8 |  | 4 | 19.25 | 101.5 | 101.2
-14.3 |  |  | 5.9 | 121.02 | 120.3 |  |  | 6.1 | 117.9633 | 117.5 |  |  | 8.3 | 112.7 | 112.1 |  | 2.8 | 13.62 | 105.2 | 104.8 |  | 4 | 19.68 | 101.6 | 101.2667
-14.2 |  |  | 6 | 121.02 | 120.3 |  |  | 6.2 | 117.88 | 117.4 |  |  | 8.54 | 112.6 | 112.05 |  | 2.9 | 13.95 | 105.2 | 104.8 |  | 4.1 | 20.26 | 101.6833 | 101.35
-14.1 |  |  | 6.15 | 120.92 | 120.2 |  |  | 6.3 | 117.73 | 117.3286 |  |  | 8.6333 | 112.5 | 111.95 |  | 2.9 | 14.18 | 105.275 | 104.8 |  | 4.1 | 20.75 | 101.9 | 101.5333
-14 |  |  | 6.2 | 120.84 | 120.1 |  |  | 6.45 | 117.5925 | 117.1667 |  |  | 8.825 | 112.4 | 111.8 |  | 2.9 | 14.5 | 105.3 | 104.8 |  | 4.2 | 21.6 | 102 | 101.6
-13.9 |  |  | 6.3333 | 120.72 | 119.8833 |  |  | 6.6 | 117.38 | 117 |  |  | 9.075 | 112.2667 | 111.74 |  | 3 | 14.75 | 105.3 | 104.9 |  | 4.2 | 22.4833 | 102.1 | 101.76
-13.8 |  |  | 6.4 | 120.445 | 119.7 |  |  | 6.775 | 117.1633 | 116.82 |  |  | 9.2 | 112.1143 | 111.625 |  | 3 | 15.15 | 105.35 | 104.9 |  | 4.3 | 23.4 | 102.14 | 101.875
-13.7 |  |  | 6.525 | 120.22 | 119.5167 |  |  | 6.9 | 116.9975 | 116.64 |  |  | 9.4 | 111.9625 | 111.525 |  | 3.0667 | 15.55 | 105.4 | 105 |  | 4.4 | 24.76 | 102.24 | 101.9
-13.6 |  |  | 6.7 | 119.92 | 119.24 |  |  | 7.025 | 116.8143 | 116.48 |  |  | 9.65 | 111.8286 | 111.425 |  | 3.1 | 15.9 | 105.5 | 105 |  | 4.4 | 26.48 | 102.3 | 102
-13.5 |  |  | 6.8 | 119.72 | 118.94 |  |  | 7.225 | 116.5643 | 116.1833 |  |  | 9.84 | 111.7 | 111.3 |  | 3.1 | 16.35 | 105.5 | 105.1 |  | 4.5 | 28.58 | 102.4 | 102.1
-13.4 |  |  | 7 | 119.4033 | 118.725 |  |  | 7.4 | 116.3143 | 116 |  |  | 10.1167 | 111.6 | 111.12 |  | 3.2 | 16.6667 | 105.5 | 105.2 |  | 4.5 | 31.76 | 102.5 | 102.2
-13.3 |  |  | 7.1 | 119.1033 | 118.3 |  |  | 7.54 | 116.121 | 115.8 |  |  | 10.3 | 111.4429 | 111 |  | 3.2 | 17.2 | 105.6 | 105.225 |  | 4.55 | 36.025 | 102.525 | 102.2
-13.2 |  |  | 7.3 | 118.92 | 118.075 |  |  | 7.7 | 115.8667 | 115.5333 |  |  | 10.6 | 111.4 | 110.925 |  | 3.3 | 17.75 | 105.6 | 105.3 |  | 4.6 | 42 | 102.6 | 102.3
-13.1 |  |  | 7.42 | 118.57 | 117.8667 |  |  | 7.88 | 115.681 | 115.3333 |  | 2.1 | 10.9 | 111.2 | 110.8 |  | 3.3 | 18.4 | 105.62 | 105.3167 |  | 4.7 | 48.375 | 102.7 | 102.325
-13 |  |  | 7.54 | 118.24 | 117.56 |  |  | 8.025 | 115.481 | 115.1 |  | 2.1 | 11.16 | 111.2 | 110.7 |  | 3.4 | 19.15 | 105.7 | 105.4 |  | 4.8 | 55.7 | 102.8 | 102.4
-12.9 |  |  | 7.75 | 117.9 | 117.26 |  |  | 8.26 | 115.2143 | 114.7857 |  | 2.2 | 11.4333 | 111.1 | 110.7 |  | 3.4 | 20 | 105.7 | 105.4 |  | 4.9 | 63.6833 | 102.825 | 102.5333
-12.8 |  |  | 7.85 | 117.6533 | 116.95 |  |  | 8.5 | 115.0143 | 114.5667 |  | 2.2 | 11.68 | 111.0625 | 110.6 |  | 3.5 | 21.1 | 105.8 | 105.5 |  | 5 | 70.6 | 102.9 | 102.6
-12.7 |  |  | 8 | 117.3533 | 116.7 |  |  | 8.6 | 114.781 | 114.4 |  | 2.2 | 12.0667 | 110.9375 | 110.575 |  | 3.5 | 22.55 | 105.84 | 105.6 |  | 5.05 | 76.74 | 103 | 102.8
-12.6 |  |  | 8.28 | 117.02 | 116.34 |  |  | 8.8 | 114.531 | 114.2333 |  | 2.2 | 12.3 | 110.8857 | 110.5 |  | 3.5667 | 24.3 | 105.9333 | 105.6 |  | 5.125 | 81.8833 | 103.02 | 102.8
-12.5 |  |  | 8.46 | 116.819 | 116.2 |  | 1.8 | 9.1 | 114.431 | 114.05 |  | 2.26 | 12.78 | 110.8 | 110.48 |  | 3.6 | 26.9 | 106 | 105.7 |  | 5.2 | 85.3833 | 103.2 | 102.85
-12.4 |  |  | 8.625 | 116.369 | 115.86 |  | 1.8 | 9.3667 | 114.181 | 113.8714 |  | 2.3 | 13.16 | 110.8 | 110.4 |  | 3.6 | 30.55 | 106.1 | 105.7 |  | 5.2667 | 87.5667 | 103.2 | 102.9
-12.3 |  |  | 8.82 | 116.144 | 115.625 |  | 1.8 | 9.58 | 113.981 | 113.65 |  | 2.3 | 13.7333 | 110.7429 | 110.4 |  | 3.7 | 35.7 | 106.1 | 105.8 |  | 5.4 | 89.6 | 103.2333 | 103
-12.2 |  |  | 9.0333 | 115.919 | 115.38 |  | 1.9 | 9.85 | 113.881 | 113.5857 |  | 2.4 | 14.225 | 110.7 | 110.4 |  | 3.7667 | 41.8 | 106.2 | 105.8 |  | 5.5 | 90.9667 | 103.3167 | 103.0167
-12.1 |  | 2 | 9.25 | 115.739 | 115.05 |  | 1.9 | 10.18 | 113.7643 | 113.4333 |  | 2.4 | 14.8 | 110.7 | 110.4 |  | 3.8667 | 50.225 | 106.2 | 105.8 |  | 5.6 | 92.0667 | 103.4 | 103.1
-12 |  | 2 | 9.5 | 115.519 | 114.925 |  | 1.9 | 10.5 | 113.681 | 113.3 |  | 2.4 | 15.525 | 110.7 | 110.4 |  | 3.9 | 58.95 | 106.225 | 105.9 |  | 5.7 | 93 | 103.4667 | 103.1
-11.9 |  | 2 | 9.76 | 115.319 | 114.75 |  | 1.9 | 10.9 | 113.5435 | 113.2 |  | 2.5 | 16.5 | 110.7 | 110.4 |  | 4 | 67.5 | 106.3 | 105.9 |  | 5.8 | 93.72 | 103.5 | 103.2
-11.8 |  | 2 | 10.075 | 115.219 | 114.6 |  | 2 | 11.3 | 113.481 | 113.1 |  | 2.5 | 17.6 | 110.7 | 110.4 |  | 4 | 75.06 | 106.3 | 106 |  | 5.85 | 94.18 | 103.5571 | 103.3
-11.7 |  | 2 | 10.36 | 115.1357 | 114.46 |  | 2 | 11.85 | 113.331 | 113 |  | 2.5 | 18.92 | 110.7 | 110.3 |  | 4.1 | 81.6 | 106.4 | 106 |  | 5.9 | 94.72 | 103.6 | 103.32
-11.6 |  | 2.1 | 10.675 | 115.019 | 114.3 |  | 2 | 12.5 | 113.181 | 112.9 |  | 2.6 | 20.84 | 110.7 | 110.2 |  | 4.2 | 86.54 | 106.4 | 106 | 1.3 | 6 | 95.25 | 103.6 | 103.4
-11.5 |  | 2.1 | 11 | 114.819 | 114.3 |  | 2 | 13.1 | 113.0953 | 112.85 |  | 2.6 | 23.74 | 110.7 | 110.2 |  | 4.275 | 90.05 | 106.35 | 106 | 1.4 | 6.175 | 95.7333 | 103.675 | 103.4
-11.4 |  | 2.1 | 11.54 | 114.719 | 114.15 |  | 2.05 | 14.0667 | 113.081 | 112.8 |  | 2.6 | 27.8 | 110.6 | 110.2 |  | 4.4 | 92.5 | 106.4 | 106.1 | 1.4 | 6.3 | 96.0667 | 103.7 | 103.4
-11.3 |  | 2.1 | 12.1 | 114.719 | 114 |  | 2.1 | 15.2 | 112.981 | 112.8 |  | 2.7 | 33.5 | 110.6 | 110.2 |  | 4.4333 | 94.35 | 106.4 | 106.1 | 1.5 | 6.4 | 96.4 | 103.7 | 103.5
-11.2 |  | 2.2 | 12.8333 | 114.579 | 113.9 |  | 2.1 | 16.7667 | 112.8977 | 112.7 |  | 2.7333 | 41.6 | 110.6 | 110.2 |  | 4.5 | 95.52 | 106.4 | 106.1 | 1.5 | 6.525 | 96.62 | 103.75 | 103.5
-11.1 |  | 2.2 | 13.7 | 114.419 | 113.8 |  | 2.1 | 18.75 | 112.881 | 112.6375 |  | 2.8 | 50.9 | 110.5 | 110.2 |  | 4.6 | 96.66 | 106.5 | 106.075 | 1.5 | 6.675 | 97 | 103.8 | 103.625
-11 |  | 2.2 | 14.8 | 114.319 | 113.7 |  | 2.2 | 21.76 | 112.781 | 112.5 |  | 2.85 | 60.95 | 110.5 | 110.2 |  | 4.7 | 97.35 | 106.58 | 106.1 | 1.5 | 6.725 | 97.2667 | 103.9 | 103.7
-10.9 |  | 2.2333 | 16.325 | 114.219 | 113.62 |  | 2.2 | 26.5 | 112.681 | 112.45 |  | 2.9 | 70.6 | 110.5 | 110.1 |  | 4.8 | 97.9 | 106.6 | 106.2 | 1.5 | 6.8 | 97.55 | 104 | 103.8
-10.8 |  | 2.3 | 18.75 | 114.219 | 113.6 |  | 2.2 | 33.475 | 112.6143 | 112.3833 |  | 2.9 | 79.66 | 110.4 | 110 |  | 4.88 | 98.65 | 106.6 | 106.2 | 1.5 | 6.9 | 97.84 | 104 | 103.9
-10.7 |  | 2.3 | 22.275 | 114.079 | 113.5 |  | 2.3 | 42.56 | 112.581 | 112.3 |  | 3 | 87.26 | 110.3833 | 110 | 1.1 | 4.9333 | 99.06 | 106.6 | 106.2 | 1.5 | 7.025 | 97.98 | 104 | 103.9
-10.6 |  | 2.4 | 27.4 | 113.969 | 113.5 |  | 2.3 | 52.82 | 112.4477 | 112.2 |  | 3 | 92.6 | 110.3 | 109.92 | 1.1 | 5.0667 | 99.4 | 106.6 | 106.2 | 1.5 | 7.1 | 98.2333 | 104 | 103.9
-10.5 |  | 2.4 | 35.5 | 113.859 | 113.3 |  | 2.3667 | 64.4 | 112.381 | 112.175 |  | 3.1 | 96.2 | 110.25 | 109.9 | 1.1 | 5.2 | 99.8 | 106.6 | 106.1 | 1.6 | 7.3 | 98.46 | 104 | 103.9
-10.4 |  | 2.4 | 46.14 | 113.719 | 113.2 |  | 2.4 | 74.8 | 112.3643 | 112.1 |  | 3.1333 | 98.65 | 110.2 | 109.9 | 1.1 | 5.3 | 100.1 | 106.5 | 106.18 | 1.6 | 7.45 | 98.675 | 104.1 | 104
-10.3 |  | 2.4 | 58 | 113.639 | 113.1 |  | 2.4 | 84.4 | 112.281 | 111.9833 |  | 3.2 | 100.42 | 110.2 | 109.9 | 1.2 | 5.325 | 100.4 | 106.5 | 106.2 | 1.525 | 7.6333 | 98.9 | 104.1 | 104
-10.2 |  | 2.5 | 69.5833 | 113.619 | 113.1 |  | 2.5 | 91.8 | 112.181 | 111.9 |  | 3.3 | 101.6 | 110.1429 | 109.825 | 1.2 | 5.4 | 100.7 | 106.5 | 106.2 | 1.6 | 7.8 | 99.1167 | 104.125 | 104
-10.1 |  | 2.5 | 80.56 | 113.499 | 113 |  | 2.5 | 97.3 | 112.1685 | 111.8 |  | 3.4 | 102.54 | 110.1 | 109.8 | 1.2 | 5.6 | 100.875 | 106.5 | 106.2 | 1.6333 | 7.95 | 99.2833 | 104.2 | 104
-10 |  | 2.525 | 89.9 | 113.419 | 112.9 |  | 2.6 | 100.8 | 112.0381 | 111.7833 |  | 3.425 | 103.325 | 110 | 109.7 | 1.2 | 5.7 | 101.225 | 106.5 | 106.2 | 1.7 | 8.075 | 99.45 | 104.3 | 104
-9.9 |  | 2.6 | 97.15 | 113.319 | 112.8 |  | 2.6 | 103.1 | 111.981 | 111.7 |  | 3.5 | 103.8 | 109.9333 | 109.6167 | 1.2 | 5.8667 | 101.475 | 106.5 | 106.2 | 1.7 | 8.25 | 99.675 | 104.3 | 104
-9.8 |  | 2.6 | 101.72 | 113.219 | 112.7 |  | 2.675 | 104.52 | 111.9143 | 111.7 | 0.8 | 3.6 | 104.32 | 109.8833 | 109.5833 | 1.2 | 5.95 | 101.68 | 106.5 | 106.2 | 1.8 | 8.44 | 99.8 | 104.3 | 104.1
-9.7 |  | 2.7 | 104.6 | 113.139 | 112.6 |  | 2.7667 | 105.64 | 111.8667 | 111.6 | 0.8 | 3.6333 | 104.65 | 109.8 | 109.5 | 1.2 | 6.1 | 101.84 | 106.5 | 106.2 | 1.8 | 8.675 | 99.9 | 104.3 | 104.1
-9.6 |  | 2.7 | 106.6 | 112.999 | 112.5 |  | 2.8 | 106.6 | 111.7143 | 111.4667 | 0.8 | 3.7 | 105.06 | 109.8 | 109.4833 | 1.3 | 6.2 | 102.05 | 106.5 | 106.2 | 1.8 | 8.8 | 100.04 | 104.3 | 104.2
-9.5 |  | 2.8 | 107.76 | 112.819 | 112.38 |  | 2.8 | 107.18 | 111.581 | 111.4 | 0.9 | 3.8 | 105.28 | 109.7 | 109.4 | 1.3 | 6.35 | 102.175 | 106.4 | 106.2 | 1.8 | 9 | 100.2 | 104.3 | 104.2
-9.4 |  | 2.8 | 108.625 | 112.719 | 112.3 | 0.7 | 2.9 | 107.7 | 111.481 | 111.2286 | 0.9 | 3.9 | 105.55 | 109.7 | 109.4 | 1.3 | 6.5 | 102.3 | 106.4 | 106.2 | 1.8 | 9.1333 | 100.3 | 104.3 | 104.2
-9.3 |  | 2.9 | 109.15 | 112.659 | 112.2 | 0.7 | 2.9 | 108 | 111.381 | 111.1 | 0.9 | 4 | 105.72 | 109.62 | 109.3 | 1.3 | 6.675 | 102.4714 | 106.4 | 106.2 | 1.8 | 9.35 | 100.4833 | 104.4 | 104.2
-9.2 |  | 3 | 109.575 | 112.519 | 112 | 0.7 | 3 | 108.25 | 111.281 | 111.1 | 0.9 | 4.1 | 105.925 | 109.6 | 109.3 | 1.38 | 6.9 | 102.65 | 106.4 | 106.2 | 1.875 | 9.625 | 100.5667 | 104.4 | 104.2
-9.1 |  | 3 | 109.925 | 112.439 | 111.9167 | 0.7 | 3 | 108.4667 | 111.2477 | 111 | 0.9 | 4.2 | 106.05 | 109.5 | 109.2333 | 1.4 | 7 | 102.8 | 106.4 | 106.2 | 1.9 | 9.9333 | 100.64 | 104.4 | 104.2
-9 | 0.76667 | 3.1 | 110.25 | 112.319 | 111.8 | 0.7 | 3.1 | 108.7 | 111.181 | 110.975 | 0.9 | 4.3 | 106.2 | 109.4 | 109.2 | 1.4 | 7.1667 | 102.9 | 106.4 | 106.2 | 1.9 | 10.2 | 100.84 | 104.4 | 104.3
-8.9 | 0.75 | 3.2 | 110.32 | 112.219 | 111.7 | 0.7 | 3.2 | 108.9 | 111.081 | 110.85 | 0.9 | 4.4 | 106.35 | 109.4 | 109.1 | 1.4 | 7.35 | 103 | 106.4 | 106.2 | 1.925 | 10.35 | 100.975 | 104.4 | 104.3
-8.8 | 0.7 | 3.2 | 110.4 | 112.119 | 111.6 | 0.73333 | 3.3 | 108.95 | 110.982 | 110.8 | 0.96 | 4.525 | 106.4 | 109.4 | 109.1 | 1.4 | 7.5 | 103.1 | 106.325 | 106.2 | 2 | 10.5 | 101.2 | 104.4 | 104.2667
-8.7 | 0.7 | 3.3 | 110.5 | 112.019 | 111.5 | 0.8 | 3.38 | 109 | 110.8945 | 110.7 | 1 | 4.675 | 106.5 | 109.3 | 109.05 | 1.4 | 7.7 | 103.2 | 106.3 | 106.2 | 2 | 10.775 | 101.3833 | 104.3833 | 104.24
-8.6 | 0.7 | 3.4 | 110.5 | 111.919 | 111.4333 | 0.8 | 3.4 | 109.1 | 110.782 | 110.6143 | 1 | 4.8 | 106.64 | 109.2143 | 109 | 1.4 | 7.9 | 103.3 | 106.3 | 106.2 | 2 | 11.05 | 101.4 | 104.3833 | 104.2667
-8.5 | 0.8 | 3.4 | 110.6 | 111.819 | 111.4 | 0.8 | 3.5 | 109.1 | 110.7106 | 110.4625 | 1 | 4.9 | 106.7 | 109.1 | 108.9 | 1.4 | 8.1 | 103.4667 | 106.3 | 106.2 | 2.1 | 11.3 | 101.5 | 104.4 | 104.26
-8.4 | 0.8 | 3.5 | 110.6 | 111.719 | 111.2 | 0.8 | 3.6 | 109.2 | 110.582 | 110.4 | 1 | 5 | 106.8 | 109.1 | 108.825 | 1.5 | 8.3 | 103.5333 | 106.3 | 106.2 | 2.1 | 11.6 | 101.5833 | 104.4333 | 104.3
-8.3 | 0.8 | 3.6 | 110.6 | 111.619 | 111.14 | 0.8 | 3.7 | 109.2 | 110.532 | 110.3 | 1 | 5.2 | 106.8167 | 109.0143 | 108.74 | 1.5 | 8.6 | 103.6 | 106.3 | 106.1 | 2.2 | 11.9667 | 101.625 | 104.5 | 104.3
-8.2 | 0.8 | 3.7 | 110.6 | 111.539 | 111 | 0.8 | 3.825 | 109.275 | 110.382 | 110.2143 | 1 | 5.3 | 106.9 | 108.975 | 108.7 | 1.5 | 8.8 | 103.7 | 106.3 | 106.1 | 2.2 | 12.2333 | 101.7 | 104.5 | 104.38
-8.1 | 0.8 | 3.8 | 110.5 | 111.4023 | 111 | 0.8 | 3.925 | 109.3 | 110.382 | 110.2 | 1 | 5.5 | 106.925 | 108.9 | 108.7 | 1.6 | 8.98 | 103.7 | 106.24 | 106.1 | 2.3 | 12.65 | 101.8 | 104.5 | 104.4
-8 | 0.8 | 3.9 | 110.5 | 111.2357 | 110.9 | 0.8 | 4.0333 | 109.3 | 110.2677 | 110.1 | 1 | 5.6 | 107 | 108.8 | 108.6 | 1.6 | 9.4 | 103.8 | 106.2 | 106.1 | 2.3 | 13 | 101.9 | 104.5 | 104.4
-7.9 | 0.8 | 4 | 110.4 | 111.169 | 110.8 | 0.8 | 4.1 | 109.3 | 110.182 | 110.0286 | 1.1 | 5.8 | 107 | 108.7333 | 108.6 | 1.6 | 9.6 | 103.9 | 106.1 | 106.0167 | 2.3 | 13.4 | 102 | 104.5 | 104.4
-7.8 | 0.8 | 4.1 | 110.4 | 111.119 | 110.68 | 0.8 | 4.3 | 109.3 | 110.0677 | 109.9 | 1.1 | 6 | 107 | 108.6 | 108.5 | 1.6 | 9.875 | 104 | 106.1 | 105.98 | 2.4 | 13.9333 | 102.1 | 104.5 | 104.4
-7.7 | 0.8 | 4.2 | 110.3667 | 110.918 | 110.55 | 0.9 | 4.4 | 109.2 | 109.982 | 109.8333 | 1.1 | 6.2333 | 107 | 108.58 | 108.44 | 1.7 | 10.125 | 104 | 106.1 | 105.9 | 2.4 | 14.4 | 102.1714 | 104.5 | 104.4
-7.6 | 0.8 | 4.3 | 110.3 | 110.818 | 110.4 | 0.9 | 4.55 | 109.125 | 109.882 | 109.7625 | 1.1 | 6.375 | 107 | 108.5 | 108.32 | 1.7 | 10.5333 | 104.1 | 106.1 | 105.9 | 2.5 | 15.05 | 102.25 | 104.4 | 104.4
-7.5 | 0.8 | 4.4 | 110.25 | 110.718 | 110.3 | 0.9 | 4.7 | 109.1 | 109.7445 | 109.7 | 1.1 | 6.6333 | 107 | 108.4 | 108.28 | 1.7 | 10.95 | 104.2 | 106.1 | 105.9 | 2.5 | 15.7667 | 102.3 | 104.4 | 104.4
-7.4 | 0.9 | 4.575 | 110.2 | 110.618 | 110.2333 | 0.9 | 4.8333 | 109.06 | 109.682 | 109.6 | 1.1 | 6.9 | 107 | 108.4 | 108.16 | 1.8 | 11.325 | 104.3 | 106 | 105.9 | 2.6 | 16.55 | 102.3 | 104.4 | 104.3
-7.3 | 0.9 | 4.7 | 110.15 | 110.518 | 110.1 | 0.9 | 5.0333 | 109 | 109.5487 | 109.4857 | 1.1 | 7.12 | 107 | 108.3 | 108.1 | 1.8 | 11.7667 | 104.3 | 106 | 105.8 | 2.625 | 17.7 | 102.4 | 104.4 | 104.3
-7.2 | 0.9 | 4.8333 | 110.1 | 110.398 | 110 | 0.9 | 5.2 | 109 | 109.4249 | 109.3 | 1.2 | 7.4333 | 107 | 108.2333 | 108 | 1.9 | 12.35 | 104.3 | 106 | 105.8 | 2.7 | 19 | 102.42 | 104.4 | 104.3
-7.1 | 0.9 | 5 | 110 | 110.318 | 109.9 | 0.9 | 5.4667 | 108.9 | 109.382 | 109.2167 | 1.2 | 7.7667 | 107 | 108.1667 | 107.9 | 1.9333 | 12.95 | 104.4 | 106 | 105.8 | 2.7 | 20.8 | 102.5857 | 104.4 | 104.3
-7 | 0.9 | 5.15 | 109.9 | 110.218 | 109.8 | 0.9 | 5.6 | 108.9 | 109.282 | 109.15 | 1.2 | 8 | 107 | 108.1 | 107.8 | 2 | 13.6 | 104.4 | 105.9 | 105.7 | 2.7 | 23.3 | 102.6 | 104.36 | 104.2
-6.9 | 0.9 | 5.3 | 109.8 | 110.118 | 109.66 | 0.975 | 5.9 | 108.76 | 109.1487 | 108.98 | 1.2333 | 8.45 | 106.9 | 108 | 107.78 | 2 | 14.5333 | 104.4 | 105.9 | 105.7 | 2.8 | 26.7 | 102.6 | 104.3 | 104.2
-6.8 | 0.9 | 5.5333 | 109.7 | 109.9513 | 109.5 | 1 | 6.15 | 108.7 | 109.082 | 108.9 | 1.3 | 8.9333 | 106.9 | 107.9 | 107.65 | 2.05 | 15.6333 | 104.4 | 105.8667 | 105.6 | 2.85 | 31.4 | 102.6667 | 104.3 | 104.2
-6.7 | 0.9 | 5.8 | 109.6 | 109.818 | 109.46 | 1 | 6.5 | 108.6 | 108.9106 | 108.8 | 1.3 | 9.3667 | 106.9 | 107.8 | 107.6 | 2.1 | 17.1 | 104.4 | 105.8 | 105.6 | 3 | 37.6 | 102.74 | 104.3 | 104.2
-6.6 | 0.96667 | 6.0667 | 109.52 | 109.7347 | 109.325 | 1 | 6.8 | 108.5 | 108.832 | 108.72 | 1.3 | 9.9667 | 106.8 | 107.7 | 107.5 | 2.1 | 18.8 | 104.4 | 105.7 | 105.5 | 3 | 44.7 | 102.8 | 104.3 | 104.2
-6.5 | 1 | 6.35 | 109.4 | 109.618 | 109.2 | 1 | 7.2 | 108.5 | 108.7487 | 108.6 | 1.3333 | 10.7667 | 106.8 | 107.5833 | 107.4 | 2.2 | 21.4333 | 104.4 | 105.6167 | 105.5 | 3.0333 | 53.35 | 102.8 | 104.2167 | 104.2
-6.4 | 1 | 6.6667 | 109.3 | 109.518 | 109.1 | 1 | 7.75 | 108.4 | 108.5963 | 108.5 | 1.4 | 11.7 | 106.8 | 107.5 | 107.4 | 2.25 | 25.2 | 104.4 | 105.6 | 105.5 | 3.2 | 61.975 | 102.86 | 104.2 | 104.2
-6.3 | 1 | 7.0333 | 109.1 | 109.338 | 109 | 1.1 | 8.4 | 108.28 | 108.457 | 108.4286 | 1.4 | 12.75 | 106.8 | 107.4 | 107.2 | 2.3 | 30.1667 | 104.4 | 105.5 | 105.4333 | 3.3 | 70.325 | 102.9 | 104.2 | 104.1
-6.2 | 1 | 7.6 | 109.1 | 109.218 | 108.9 | 1.1 | 9.1 | 108.175 | 108.302 | 108.3667 | 1.5 | 14.1667 | 106.7 | 107.325 | 107.2 | 2.4 | 36.9 | 104.4 | 105.5 | 105.4 | 3.4 | 77.4 | 102.9 | 104.2 | 104.1
-6.1 | 1 | 8.1667 | 109 | 109.078 | 108.75 | 1.1 | 10 | 108.1 | 108.182 | 108.1286 | 1.5 | 16.05 | 106.7 | 107.2 | 107.1 | 2.4 | 45.275 | 104.48 | 105.475 | 105.3 | 3.4 | 82.9667 | 102.9 | 104.2 | 104
-6 | 1.1 | 9 | 108.9 | 108.9513 | 108.6 | 1.1 | 11.3333 | 108 | 108.1249 | 108.05 | 1.5667 | 18.7667 | 106.6 | 107.2 | 107 | 2.5 | 54.35 | 104.4 | 105.4 | 105.3 | 3.4 | 87.14 | 102.9 | 104.1 | 104
-5.9 | 1.1 | 10.05 | 108.86 | 108.818 | 108.5 | 1.1 | 12.925 | 107.92 | 108.082 | 107.9125 | 1.6 | 22.7333 | 106.54 | 107.1 | 107 | 2.6 | 63.2 | 104.5 | 105.3 | 105.2714 | 3.5 | 89.9 | 102.9 | 104.1 | 103.9
-5.8 | 1.1 | 11.4 | 108.7 | 108.718 | 108.4 | 1.2 | 15.225 | 107.8 | 107.9653 | 107.8 | 1.625 | 28.4 | 106.46 | 107.0333 | 106.9 | 2.675 | 72.55 | 104.46 | 105.3 | 105.2 | 3.675 | 92 | 102.9 | 104.1 | 103.9
-5.7 | 1.1 | 13.5 | 108.56 | 108.618 | 108.28 | 1.2 | 18.65 | 107.8 | 107.882 | 107.7 | 1.7 | 36.2 | 106.4 | 106.9667 | 106.7833 | 2.7 | 79.95 | 104.45 | 105.3 | 105.1 | 3.7 | 93.45 | 102.8 | 104.1 | 103.9
-5.6 | 1.1667 | 16.4 | 108.425 | 108.418 | 108.1 | 1.2 | 23.9 | 107.6833 | 107.782 | 107.6625 | 1.7 | 45.9 | 106.3 | 106.7167 | 106.7 | 2.8 | 85.15 | 104.4 | 105.2 | 105.0833 | 3.8 | 94.5667 | 102.8833 | 104.05 | 103.8
-5.5 | 1.2 | 20.9 | 108.3 | 108.3347 | 107.9667 | 1.3 | 31.1 | 107.6 | 107.632 | 107.5286 | 1.8 | 55.75 | 106.3 | 106.62 | 106.6 | 2.9 | 89.4 | 104.4 | 105.2 | 105 | 3.9 | 95.5 | 102.9 | 104 | 103.8
-5.4 | 1.2 | 28.6667 | 108.28 | 108.218 | 107.8 | 1.3 | 40.4667 | 107.5 | 107.4487 | 107.4 | 1.8 | 66.5 | 106.2 | 106.55 | 106.48 | 2.9333 | 92 | 104.3 | 105.1 | 104.975 | 4 | 96.1667 | 102.9 | 103.9 | 103.8
-5.3 | 1.2 | 37.7333 | 108.175 | 108.043 | 107.7 | 1.3 | 51.3 | 107.4 | 107.382 | 107.3 | 1.9 | 75.45 | 106.16 | 106.4667 | 106.3833 | 3.025 | 94 | 104.3 | 105.1 | 104.9 | 4.1 | 96.95 | 102.9 | 103.9 | 103.775
-5.2 | 1.3 | 48.6 | 108.06 | 107.843 | 107.6 | 1.4 | 62.5 | 107.3 | 107.182 | 107.2 | 1.9 | 83.1 | 106 | 106.4 | 106.2167 | 3.2 | 95.5667 | 104.2 | 104.9833 | 104.8 | 4.2 | 97.45 | 102.9 | 103.9 | 103.7
-5.1 | 1.3 | 60.075 | 107.85 | 107.818 | 107.4667 | 1.4 | 72.6 | 107.2 | 107.082 | 107.05 | 2 | 89.3 | 106 | 106.3 | 106.1 | 3.3 | 96.5333 | 104.2 | 104.9 | 104.775 | 4.3 | 97.85 | 102.9 | 103.7333 | 103.65
-5 | 1.4 | 71.375 | 107.7 | 107.698 | 107.35 | 1.475 | 81.9 | 107.1 | 106.8987 | 106.86 | 2.1 | 92.7 | 105.9 | 106.1286 | 106 | 3.35 | 97.4 | 104.175 | 104.8 | 104.6 | 4.4 | 98.2 | 102.9 | 103.7 | 103.6
-4.9 | 1.4 | 81.5 | 107.52 | 107.478 | 107.2333 | 1.5 | 88.6 | 106.92 | 106.832 | 106.7 | 2.1667 | 95.4 | 105.8 | 106.05 | 105.9 | 3.5 | 98.1 | 104.1 | 104.7 | 104.5333 | 4.65 | 98.4 | 102.9 | 103.62 | 103.6
-4.8 | 1.4333 | 89.45 | 107.34 | 107.3751 | 107.1 | 1.5 | 93.5667 | 106.8 | 106.6695 | 106.6333 | 2.2 | 97.1333 | 105.7 | 105.9125 | 105.8 | 3.6 | 98.575 | 104.1 | 104.6 | 104.5 | 4.8 | 98.74 | 102.9 | 103.5167 | 103.6
-4.7 | 1.5 | 95.0333 | 107.2 | 107.1847 | 106.9 | 1.6 | 96.6 | 106.66 | 106.532 | 106.5 | 2.3 | 98.275 | 105.6 | 105.8 | 105.8 | 3.7 | 99.025 | 104.1 | 104.6 | 104.5 | 4.925 | 99.05 | 102.9 | 103.5 | 103.5
-4.6 | 1.5 | 98.2 | 107.1 | 107.0513 | 106.8 | 1.7 | 98.65 | 106.54 | 106.382 | 106.34 | 2.4 | 99.3 | 105.42 | 105.7429 | 105.62 | 3.8 | 99.475 | 104 | 104.5 | 104.4 | 5.175 | 99.3 | 102.8 | 103.4 | 103.4
-4.5 | 1.6 | 100 | 107 | 106.843 | 106.7 | 1.7667 | 100 | 106.4 | 106.282 | 106.15 | 2.475 | 99.95 | 105.325 | 105.6 | 105.54 | 4 | 99.7333 | 103.9 | 104.4333 | 104.3 | 5.3 | 99.45 | 102.8167 | 103.4 | 103.4
-4.4 | 1.6 | 101.5 | 106.9 | 106.738 | 106.55 | 1.85 | 100.875 | 106.28 | 106.122 | 106.1 | 2.6 | 100.5 | 105.2333 | 105.475 | 105.4 | 4.2 | 100.0667 | 103.8 | 104.3 | 104.22 | 5.4 | 99.7 | 102.8 | 103.35 | 103.3
-4.3 | 1.7 | 102.4 | 106.74 | 106.618 | 106.4 | 1.9 | 101.65 | 106.175 | 105.982 | 105.9429 | 2.7 | 101 | 105.2 | 105.35 | 105.3 | 4.2667 | 100.2 | 103.8 | 104.2 | 104.15 | 5.6333 | 99.8667 | 102.8 | 103.2833 | 103.3
-4.2 | 1.8 | 103.2 | 106.54 | 106.458 | 106.225 | 2 | 102.15 | 106 | 105.8653 | 105.9 | 2.85 | 101.3333 | 105.1 | 105.2 | 105.2 | 4.4333 | 100.45 | 103.75 | 104.1 | 104.08 | 5.9333 | 100 | 102.7 | 103.2 | 103.3
-4.1 | 1.8 | 103.5 | 106.42 | 106.218 | 106.1 | 2.1 | 102.6667 | 105.9 | 105.6987 | 105.7429 | 3 | 101.65 | 105 | 105.12 | 105.1 | 4.725 | 100.575 | 103.7 | 104 | 103.98 | 6.1667 | 100.125 | 102.7 | 103.2 | 103.2
-4 | 1.95 | 103.8 | 106.275 | 106.118 | 105.92 | 2.2 | 103.0667 | 105.7667 | 105.532 | 105.5857 | 3.175 | 101.9 | 104.9 | 105.02 | 105.02 | 4.9333 | 100.8 | 103.6 | 103.9 | 103.9 | 6.3667 | 100.3 | 102.6571 | 103.2 | 103.1
-3.9 | 2 | 103.9333 | 106.06 | 106.018 | 105.775 | 2.3 | 103.2 | 105.6 | 105.4653 | 105.45 | 3.4 | 102.1 | 104.78 | 104.8857 | 104.8 | 5.15 | 101 | 103.5 | 103.8 | 103.8 | 6.8 | 100.3667 | 102.6 | 103.1 | 103.1
-3.8 | 2.1 | 104.0667 | 105.96 | 105.858 | 105.625 | 2.4667 | 103.3667 | 105.475 | 105.3106 | 105.2833 | 3.6 | 102.2333 | 104.6857 | 104.78 | 104.7 | 5.5 | 101.125 | 103.5 | 103.8 | 103.7 | 7.35 | 100.475 | 102.6 | 103.1 | 103
-3.7 | 2.2 | 104.2 | 105.775 | 105.658 | 105.5 | 2.7 | 103.5 | 105.3 | 105.182 | 105.1286 | 3.8 | 102.4 | 104.6 | 104.6 | 104.6 | 5.86 | 101.275 | 103.5 | 103.7 | 103.6 | 7.9 | 100.6 | 102.6 | 103 | 103
-3.6 | 2.4 | 104.3 | 105.575 | 105.558 | 105.4 | 2.9 | 103.5 | 105.175 | 105.082 | 104.9857 | 4.1 | 102.4 | 104.42 | 104.5429 | 104.4 | 6.25 | 101.3 | 103.4 | 103.6333 | 103.6 | 8.45 | 100.7 | 102.6 | 102.9 | 102.9
-3.5 | 2.5667 | 104.3 | 105.475 | 105.443 | 105.225 | 3.1 | 103.6 | 105.0667 | 104.9259 | 104.9 | 4.4333 | 102.5 | 104.3 | 104.3833 | 104.3 | 6.7667 | 101.3 | 103.4 | 103.6 | 103.5 | 9.2333 | 100.8 | 102.525 | 102.8 | 102.8
-3.4 | 2.775 | 104.3 | 105.36 | 105.218 | 105.1 | 3.45 | 103.6 | 104.9 | 104.783 | 104.7833 | 4.7667 | 102.5333 | 104.2 | 104.2375 | 104.225 | 7.4 | 101.4 | 103.3 | 103.45 | 103.4 | 10.3667 | 100.9 | 102.5 | 102.8 | 102.7167
-3.3 | 3 | 104.2 | 105.2 | 105.098 | 104.8 | 3.7 | 103.5 | 104.7 | 104.633 | 104.7 | 5.2 | 102.6 | 104.1 | 104.2 | 104.08 | 8.3 | 101.5 | 103.2 | 103.3 | 103.3 | 11.6 | 100.9 | 102.475 | 102.66 | 102.7
-3.2 | 3.3333 | 104.2 | 105 | 104.917 | 104.7167 | 4.2 | 103.5 | 104.58 | 104.383 | 104.54 | 5.95 | 102.6 | 103.9 | 104.05 | 103.96 | 9.4 | 101.6 | 103.2 | 103.2167 | 103.2 | 13.2333 | 100.9 | 102.4 | 102.5857 | 102.6
-3.1 | 3.7 | 104.2 | 104.9 | 104.817 | 104.55 | 4.85 | 103.5 | 104.475 | 104.283 | 104.35 | 6.6 | 102.6 | 103.8 | 104 | 103.9 | 10.9 | 101.6 | 103.1 | 103.1 | 103.1 | 15.65 | 101.1 | 102.3714 | 102.5 | 102.6
-3 | 4.2333 | 104.0333 | 104.7333 | 104.637 | 104.4333 | 5.6333 | 103.5 | 104.35 | 104.1497 | 104.2 | 7.5 | 102.6 | 103.7 | 103.85 | 103.75 | 12.3667 | 101.7 | 103 | 103 | 103.075 | 18.7333 | 101.1 | 102.3 | 102.5 | 102.5
-2.9 | 4.9 | 104 | 104.5667 | 104.497 | 104.25 | 6.54 | 103.5 | 104.1833 | 103.9973 | 104.1 | 8.95 | 102.7 | 103.6 | 103.6556 | 103.66 | 14.8 | 101.7 | 102.9 | 102.9 | 102.9833 | 23.6 | 101.1 | 102.3 | 102.5 | 102.4
-2.8 | 6 | 103.9 | 104.4 | 104.3003 | 104.1 | 7.9333 | 103.5 | 104 | 103.8544 | 103.95 | 10.775 | 102.6 | 103.5 | 103.5286 | 103.48 | 18.5 | 101.7333 | 102.8 | 102.85 | 102.9 | 29.65 | 101.1 | 102.25 | 102.5 | 102.3
-2.7 | 7.28 | 103.7667 | 104.2 | 104.117 | 104 | 9.6 | 103.425 | 103.8 | 103.7116 | 103.7667 | 13.0667 | 102.6 | 103.4 | 103.3714 | 103.3833 | 23.2 | 101.7 | 102.76 | 102.8 | 102.8 | 37.9 | 101.1 | 102.1833 | 102.3667 | 102.3
-2.6 | 9.2 | 103.6667 | 104.06 | 104.0337 | 103.85 | 12.325 | 103.3 | 103.7 | 103.583 | 103.5857 | 16.6 | 102.6 | 103.3 | 103.2167 | 103.3 | 30.2 | 101.7 | 102.66 | 102.7 | 102.7167 | 47.4 | 101.1 | 102.1 | 102.2 | 102.2
-2.5 | 12.3 | 103.5 | 103.9 | 103.797 | 103.65 | 15.9 | 103.2 | 103.6 | 103.4497 | 103.5 | 21.9 | 102.5 | 103.22 | 103.1875 | 103.05 | 39.1 | 101.7 | 102.55 | 102.5 | 102.5667 | 56.9 | 101.1 | 102 | 102.1333 | 102.2
-2.4 | 16.7 | 103.4 | 103.75 | 103.692 | 103.5 | 21.8 | 103.1 | 103.5 | 103.233 | 103.4 | 29.3 | 102.4 | 103.12 | 103.05 | 102.9667 | 49.2 | 101.7 | 102.42 | 102.45 | 102.5 | 66.3 | 101.2 | 101.9 | 102.04 | 102.12
-2.3 | 23.9 | 103.4 | 103.56 | 103.397 | 103.4 | 30.1 | 103.025 | 103.3 | 103.0663 | 103.22 | 39.1333 | 102.4 | 103 | 102.9125 | 102.9 | 59.36 | 101.6667 | 102.3 | 102.4 | 102.3333 | 74.85 | 101.125 | 101.82 | 102 | 102
-2.2 | 33.6 | 103.25 | 103.35 | 103.217 | 103.3 | 39.8 | 102.9 | 103.15 | 102.9116 | 103.0625 | 49.4667 | 102.4 | 102.875 | 102.775 | 102.7333 | 68.875 | 101.7 | 102.225 | 102.3 | 102.3 | 82.25 | 101.2 | 101.7 | 101.9 | 101.975
-2.1 | 44.4667 | 103.1 | 103.2 | 103.117 | 103.0667 | 51 | 102.8 | 103.1 | 102.8497 | 102.8571 | 60.1 | 102.3 | 102.7333 | 102.6429 | 102.6 | 77.8 | 101.7 | 102.12 | 102.2167 | 102.3 | 87.275 | 101.2 | 101.7 | 101.8 | 101.825
-2 | 56.3 | 102.9 | 103.1 | 103.017 | 102.9 | 61.8667 | 102.65 | 102.9 | 102.6687 | 102.76 | 70.6333 | 102.2 | 102.48 | 102.5 | 102.45 | 84.625 | 101.7 | 102 | 102.1 | 102.1167 | 91.0667 | 101.2 | 101.7 | 101.8 | 101.8
-1.9 | 67.35 | 102.8 | 103 | 102.8503 | 102.7167 | 72.3667 | 102.5 | 102.76 | 102.583 | 102.6125 | 79.375 | 102.15 | 102.35 | 102.4286 | 102.3 | 89.1 | 101.6 | 102 | 101.975 | 102 | 93.4 | 101.1 | 101.6 | 101.7 | 101.625
-1.8 | 77.3333 | 102.7 | 102.84 | 102.717 | 102.6 | 81.325 | 102.4 | 102.625 | 102.4163 | 102.5 | 86.1 | 102.1 | 102.2 | 102.2667 | 102.2 | 92.55 | 101.525 | 101.9 | 101.9 | 101.9 | 95.1 | 101.1 | 101.5 | 101.58 | 101.6
-1.7 | 85.6 | 102.5 | 102.64 | 102.537 | 102.4571 | 88.1 | 102.3 | 102.4 | 102.2663 | 102.4 | 91 | 102 | 102.0833 | 102.125 | 102.0429 | 94.7 | 101.5 | 101.9 | 101.8 | 101.7 | 96.7667 | 101.2 | 101.5 | 101.5 | 101.48
-1.6 | 91.4 | 102.4 | 102.44 | 102.417 | 102.25 | 92 | 102.2 | 102.3 | 102.0997 | 102.15 | 93.5 | 101.9 | 102 | 101.98 | 101.98 | 96.4 | 101.4 | 101.7667 | 101.7 | 101.6 | 97.5 | 101.2 | 101.42 | 101.4 | 101.4
-1.5 | 94.625 | 102.2 | 102.3333 | 102.192 | 102.1 | 95 | 102.025 | 102.125 | 101.923 | 102.0143 | 95.6 | 101.7 | 101.9 | 101.8143 | 101.8 | 97.6 | 101.3 | 101.7 | 101.6667 | 101.52 | 98.2 | 101.15 | 101.3 | 101.3 | 101.3
-1.4 | 97 | 102.1 | 102.14 | 102.067 | 101.95 | 96.6 | 101.9 | 101.9167 | 101.7973 | 101.85 | 97.2 | 101.6 | 101.8 | 101.5875 | 101.65 | 98.28 | 101.3 | 101.6 | 101.5 | 101.4429 | 98.7 | 101.1 | 101.3 | 101.2 | 101.2333
-1.3 | 98.3 | 101.9 | 101.95 | 101.917 | 101.76 | 97.9 | 101.7 | 101.8 | 101.583 | 101.725 | 98 | 101.5 | 101.62 | 101.4625 | 101.6 | 98.9 | 101.2 | 101.4 | 101.4167 | 101.3833 | 99.25 | 101 | 101.2 | 101.1833 | 101.2
-1.2 | 99.2 | 101.7 | 101.8 | 101.717 | 101.6 | 98.6 | 101.7 | 101.6 | 101.4163 | 101.65 | 98.8 | 101.4 | 101.5 | 101.4 | 101.44 | 99.35 | 101.15 | 101.34 | 101.25 | 101.2167 | 99.5 | 100.8333 | 101.1 | 101.1 | 101.1
-1.1 | 99.75 | 101.6 | 101.7 | 101.617 | 101.46 | 99.325 | 101.5 | 101.52 | 101.2705 | 101.4 | 99.2667 | 101.375 | 101.3 | 101.2667 | 101.3 | 99.7 | 101.1 | 101.16 | 101.12 | 101.1 | 99.725 | 100.8 | 101 | 100.925 | 101.08
-1 | 100.16 | 101.5 | 101.575 | 101.397 | 101.3 | 99.7333 | 101.3 | 101.4 | 101.1401 | 101.3571 | 99.7 | 101.2 | 101.2 | 101.125 | 101.2 | 100.0333 | 101.075 | 101.075 | 101 | 101 | 99.975 | 100.775 | 100.9 | 100.8 | 100.9
-0.9 | 100.3 | 101.3667 | 101.44 | 101.257 | 101.2 | 99.9 | 101.2 | 101.2667 | 100.9973 | 101.2143 | 99.95 | 101.1 | 101.075 | 101.0833 | 101.05 | 100.175 | 100.975 | 100.9833 | 100.9 | 100.9 | 100.1 | 100.7 | 100.8333 | 100.7667 | 100.8167
-0.8 | 100.4 | 101.3 | 101.125 | 101.117 | 101.02 | 100.2 | 101.1 | 101.1 | 100.8663 | 101 | 100.1333 | 100.975 | 100.9 | 100.9167 | 101 | 100.3 | 100.8333 | 100.9 | 100.8 | 100.8 | 100.2 | 100.7 | 100.7 | 100.675 | 100.8
-0.7 | 100.5 | 101 | 100.98 | 101.017 | 100.9 | 100.3 | 101 | 100.95 | 100.7205 | 100.8833 | 100.25 | 100.9 | 100.8 | 100.82 | 100.8 | 100.3 | 100.7 | 100.8 | 100.62 | 100.7 | 100.25 | 100.6 | 100.5833 | 100.6 | 100.7
-0.6 | 100.6 | 100.9 | 100.86 | 100.8837 | 100.775 | 100.3 | 100.9 | 100.7 | 100.558 | 100.7167 | 100.3 | 100.7 | 100.7 | 100.7571 | 100.66 | 100.4 | 100.7 | 100.7 | 100.55 | 100.6 | 100.2 | 100.5 | 100.5 | 100.5 | 100.6286
-0.5 | 100.5333 | 100.7 | 100.74 | 100.697 | 100.65 | 100.3 | 100.7 | 100.6 | 100.4163 | 100.54 | 100.3 | 100.6 | 100.5 | 100.6143 | 100.5 | 100.3 | 100.6 | 100.6 | 100.4833 | 100.4667 | 100.1 | 100.4667 | 100.4 | 100.4 | 100.5
-0.4 | 100.5 | 100.5 | 100.525 | 100.4837 | 100.42 | 100.2333 | 100.6 | 100.475 | 100.383 | 100.4 | 100.3 | 100.5 | 100.4 | 100.5333 | 100.42 | 100.3 | 100.5 | 100.4667 | 100.4 | 100.3 | 100.1 | 100.4 | 100.3143 | 100.3 | 100.4
-0.3 | 100.4 | 100.4 | 100.4 | 100.367 | 100.3 | 100.1 | 100.4 | 100.4 | 100.2544 | 100.3 | 100.2 | 100.425 | 100.3 | 100.3333 | 100.3 | 100.2 | 100.4 | 100.3 | 100.38 | 100.2 | 100.1 | 100.3 | 100.24 | 100.2833 | 100.2333
-0.2 | 100.3 | 100.3 | 100.3 | 100.217 | 100.2 | 100.1 | 100.2 | 100.28 | 100.183 | 100.2333 | 100.1667 | 100.225 | 100.24 | 100.1143 | 100.2 | 100.1 | 100.2667 | 100.2 | 100.2 | 100.1 | 100.1 | 100.2 | 100.2 | 100.2 | 100.1667
-0.1 | 100.2 | 100.2 | 100.15 | 100.192 | 100.175 | 100.0333 | 100.1 | 100.1 | 100.083 | 100.1 | 100.1 | 100.1 | 100.1 | 100.0857 | 100.1 | 100.0667 | 100.1 | 100.1167 | 100.1 | 100.05 | 100 | 100.1 | 100.075 | 100.1 | 100.1
0 | 100 | 100 | 100 | 100.0003 | 100 | 100 | 100 | 100 | 99.9999 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.9 | 99.875 | 99.9 | 99.9167 | 99.825 | 99.875 | 99.9 | 99.8833 | 99.8834 | 99.9 | 99.9 | 99.9 | 99.9 | 99.8714 | 99.8833 | 99.9 | 99.9 | 99.9167 | 99.9 | 99.9 | 99.9 | 99.9 | 99.8833 | 99.9 | 99.9
0.2 | 99.725 | 99.7333 | 99.8 | 99.8833 | 99.7333 | 99.7 | 99.8 | 99.7833 | 99.7334 | 99.8333 | 99.7333 | 99.8 | 99.8 | 99.7375 | 99.8 | 99.7 | 99.8333 | 99.9 | 99.8 | 99.8 | 99.8 | 99.9 | 99.8 | 99.8167 | 99.9
0.3 | 99.6 | 99.7 | 99.7333 | 99.7666 | 99.7 | 99.575 | 99.7 | 99.7 | 99.5834 | 99.7 | 99.6 | 99.725 | 99.6 | 99.6 | 99.7 | 99.6 | 99.8 | 99.8 | 99.7143 | 99.7 | 99.7 | 99.8 | 99.8 | 99.72 | 99.8
0.4 | 99.4 | 99.575 | 99.6 | 99.7166 | 99.6 | 99.275 | 99.6 | 99.5857 | 99.4834 | 99.64 | 99.4 | 99.6 | 99.5286 | 99.55 | 99.6 | 99.4333 | 99.7 | 99.7 | 99.6 | 99.6 | 99.6333 | 99.7 | 99.7 | 99.7 | 99.8
0.5 | 99.3667 | 99.4 | 99.5 | 99.6166 | 99.52 | 99.1 | 99.525 | 99.5 | 99.4234 | 99.55 | 99.2667 | 99.4333 | 99.5 | 99.425 | 99.5 | 99.3 | 99.6667 | 99.6833 | 99.575 | 99.6 | 99.475 | 99.6 | 99.6 | 99.65 | 99.675
0.6 | 99.175 | 99.3 | 99.4 | 99.4666 | 99.4 | 98.85 | 99.4 | 99.4 | 99.2835 | 99.4667 | 99.1 | 99.4 | 99.4 | 99.3667 | 99.4667 | 99.1667 | 99.5 | 99.6 | 99.5 | 99.5 | 99.3 | 99.5 | 99.5 | 99.6 | 99.54
0.7 | 98.85 | 99.2 | 99.38 | 99.3416 | 99.3167 | 98.6667 | 99.3 | 99.3 | 99.196 | 99.4 | 98.875 | 99.3 | 99.3 | 99.3 | 99.4 | 99 | 99.4333 | 99.5 | 99.4 | 99.4 | 99.1 | 99.4333 | 99.4 | 99.5 | 99.44
0.8 | 98.55 | 99.15 | 99.3 | 99.3166 | 99.3 | 98.4 | 99.2667 | 99.2333 | 99.1335 | 99.3625 | 98.55 | 99.3 | 99.2 | 99.2 | 99.3 | 98.7333 | 99.4 | 99.4 | 99.3 | 99.4 | 98.95 | 99.4 | 99.3 | 99.4 | 99.4
0.9 | 98.2 | 99 | 99.2 | 99.2165 | 99.2 | 97.9 | 99.1 | 99.125 | 99.0835 | 99.2375 | 98.15 | 99.2 | 99.1 | 99.1833 | 99.2 | 98.4333 | 99.25 | 99.35 | 99.3 | 99.2667 | 98.7 | 99.225 | 99.2333 | 99.375 | 99.36
1 | 97.575 | 98.9333 | 99.175 | 99.1165 | 99.18 | 97.35 | 99.05 | 99 | 99.0002 | 99.1125 | 97.8 | 99.1 | 99 | 99.1 | 99.1 | 98.15 | 99.2 | 99.25 | 99.25 | 99.2 | 98.45 | 99.1 | 99.2 | 99.2 | 99.25
1.1 | 97 | 98.9 | 99.0667 | 99.0415 | 99.1 | 96.7 | 98.9 | 98.98 | 98.9335 | 99.1 | 96.9 | 99 | 98.9 | 99 | 99.04 | 97.65 | 99.1 | 99.1 | 99.1 | 99.1 | 98.2 | 99.1 | 99.1 | 99.1 | 99.2
1.2 | 96.2 | 98.8 | 99 | 99.0165 | 99 | 95.85 | 98.8 | 98.9 | 98.8668 | 99 | 96.425 | 98.825 | 98.9 | 98.95 | 99 | 97.3 | 99 | 99 | 99 | 99.06 | 97.775 | 99 | 99 | 99.1 | 99.1
1.3 | 95 | 98.7 | 98.9 | 98.9332 | 99 | 94.7 | 98.8 | 98.8 | 98.7121 | 98.9143 | 95.6 | 98.7 | 98.8 | 98.9 | 98.9 | 96.7667 | 98.8 | 98.95 | 98.9 | 99 | 97.3667 | 98.9 | 99 | 99 | 99
1.4 | 93.5 | 98.6333 | 98.8 | 98.9165 | 98.9 | 93.4333 | 98.7 | 98.8 | 98.655 | 98.9 | 94.6 | 98.625 | 98.7 | 98.85 | 98.8 | 96.05 | 98.7 | 98.8 | 98.9 | 99 | 96.8 | 98.8 | 98.9 | 99 | 99
1.5 | 91.25 | 98.5 | 98.775 | 98.9165 | 98.85 | 91.1667 | 98.6 | 98.7 | 98.5836 | 98.825 | 93 | 98.5 | 98.68 | 98.8 | 98.7 | 95 | 98.575 | 98.8 | 98.8 | 98.9 | 96.35 | 98.62 | 98.82 | 99 | 98.9
1.6 | 87.6 | 98.4 | 98.7 | 98.8165 | 98.8 | 88.1 | 98.5 | 98.6 | 98.5503 | 98.8 | 91 | 98.3 | 98.6 | 98.7 | 98.7 | 93.8 | 98.425 | 98.7 | 98.7667 | 98.7667 | 95.6 | 98.5 | 98.8 | 98.9 | 98.9
1.7 | 81.75 | 98.25 | 98.6 | 98.7164 | 98.7 | 83.6 | 98.4 | 98.5167 | 98.4836 | 98.7714 | 88.075 | 98.2 | 98.5 | 98.6286 | 98.6 | 92.2 | 98.3 | 98.6 | 98.7 | 98.7 | 94.7 | 98.4 | 98.7 | 98.7833 | 98.75
1.8 | 73.1 | 98.2 | 98.6 | 98.5664 | 98.6 | 77 | 98.3 | 98.5 | 98.4169 | 98.6286 | 83.48 | 98.1 | 98.4 | 98.5 | 98.5 | 90.15 | 98.2 | 98.4833 | 98.6 | 98.6 | 93.6 | 98.225 | 98.6 | 98.6167 | 98.7
1.9 | 63.275 | 98.1 | 98.5 | 98.5164 | 98.525 | 67.9 | 98.1 | 98.4 | 98.3836 | 98.6 | 76.8 | 97.9 | 98.3 | 98.475 | 98.48 | 87.05 | 98.0667 | 98.4 | 98.45 | 98.6 | 92.025 | 98.1 | 98.5 | 98.52 | 98.6
2 | 52.1333 | 98 | 98.4 | 98.5164 | 98.5 | 57.25 | 98 | 98.3 | 98.3211 | 98.5286 | 68.4 | 97.8 | 98.3 | 98.4 | 98.4 | 82.7 | 97.9 | 98.32 | 98.4 | 98.5 | 89.7 | 97.9 | 98.4 | 98.5 | 98.5
2.1 | 41.2 | 97.9 | 98.3 | 98.4164 | 98.5 | 46.9333 | 97.9 | 98.2 | 98.2693 | 98.4625 | 59.15 | 97.7 | 98.2 | 98.3 | 98.3 | 76.5667 | 97.825 | 98.3 | 98.3167 | 98.425 | 86.725 | 97.7667 | 98.3 | 98.4 | 98.5
2.2 | 30.625 | 97.8 | 98.3 | 98.4164 | 98.4 | 36.1 | 97.8 | 98.2 | 98.1836 | 98.3375 | 47.9 | 97.6 | 98.1 | 98.2778 | 98.3 | 68.8 | 97.6667 | 98.2 | 98.2 | 98.3 | 82.6 | 97.6 | 98.18 | 98.325 | 98.4
2.3 | 21.375 | 97.8 | 98.2 | 98.3164 | 98.32 | 26.5667 | 97.7 | 98.1 | 98.1636 | 98.28 | 37.7667 | 97.5 | 98.02 | 98.15 | 98.2 | 59.45 | 97.5 | 98.1 | 98.1833 | 98.2 | 77.3 | 97.4 | 98.1 | 98.3 | 98.3
2.4 | 15.1333 | 97.6667 | 98.1 | 98.2164 | 98.3 | 19.2667 | 97.6 | 98.04 | 98.0837 | 98.2 | 28.6 | 97.3 | 97.9 | 98.0833 | 98.1 | 49.1667 | 97.3 | 98 | 98.1 | 98.2 | 69.12 | 97.3 | 98.06 | 98.2 | 98.2
2.5 | 11.3333 | 97.4 | 98 | 98.2164 | 98.2 | 14.2 | 97.4 | 97.925 | 97.9237 | 98.1 | 21.3 | 97.2 | 97.875 | 98 | 98 | 39.7 | 97.1 | 97.9 | 98.0286 | 98.1 | 61.15 | 97.2 | 97.925 | 98.1833 | 98.2
2.6 | 8.7 | 97.25 | 97.9 | 98.1164 | 98.1 | 10.95 | 97.2 | 97.84 | 97.867 | 98.08 | 16.3333 | 97 | 97.7833 | 97.9 | 97.9 | 30.5 | 97 | 97.8 | 97.9667 | 98 | 51.5 | 97.0667 | 97.85 | 98.1 | 98.1
2.7 | 6.8 | 97.1 | 97.8333 | 98.0363 | 98.0333 | 8.7 | 97.1 | 97.72 | 97.7837 | 98 | 12.725 | 96.8333 | 97.68 | 97.8714 | 97.86 | 23.3 | 96.9 | 97.7 | 97.9 | 97.9 | 41.2 | 96.9 | 97.75 | 97.9333 | 98
2.8 | 5.65 | 97.0333 | 97.725 | 98.0163 | 97.95 | 7.2 | 96.9 | 97.7 | 97.7237 | 97.9 | 10.475 | 96.65 | 97.6 | 97.8 | 97.8 | 18.3333 | 96.7 | 97.6714 | 97.7 | 97.8 | 33.05 | 96.7 | 97.7 | 97.8 | 97.9
2.9 | 4.75 | 96.9 | 97.68 | 97.9363 | 97.9 | 6 | 96.7 | 97.6 | 97.6837 | 97.8833 | 8.8 | 96.45 | 97.5 | 97.7 | 97.725 | 14.9 | 96.45 | 97.6 | 97.6 | 97.7 | 25.9 | 96.5 | 97.6 | 97.7143 | 97.8
3 | 4 | 96.7 | 97.6 | 97.9163 | 97.8 | 5.2 | 96.6 | 97.5667 | 97.5837 | 97.8 | 7.5333 | 96.2 | 97.4 | 97.6 | 97.7 | 12.48 | 96.125 | 97.45 | 97.6 | 97.66 | 20.6 | 96.275 | 97.5 | 97.65 | 97.7
3.1 | 3.6 | 96.6 | 97.5 | 97.8163 | 97.7167 | 4.4 | 96.4 | 97.4 | 97.5837 | 97.6333 | 6.575 | 96 | 97.2167 | 97.5167 | 97.6 | 10.6 | 95.925 | 97.3 | 97.5 | 97.55 | 16.9 | 96.1 | 97.3 | 97.6 | 97.55
3.2 | 3.15 | 96.4 | 97.42 | 97.8163 | 97.7 | 3.9 | 96.2 | 97.375 | 97.4171 | 97.6 | 5.8 | 95.8 | 97.2 | 97.4667 | 97.5 | 9.3 | 95.7333 | 97.2 | 97.5 | 97.5 | 14.2667 | 95.85 | 97.2 | 97.5 | 97.5
3.3 | 2.8667 | 96.2 | 97.3167 | 97.7163 | 97.6 | 3.55 | 96 | 97.25 | 97.3838 | 97.5429 | 5.1667 | 95.6 | 97.1 | 97.3143 | 97.4 | 8.1333 | 95.5 | 97.2 | 97.36 | 97.38 | 12.2667 | 95.7 | 97.1 | 97.4 | 97.4
3.4 | 2.6 | 96.1 | 97.25 | 97.6163 | 97.46 | 3.2 | 95.8 | 97.2 | 97.2838 | 97.5 | 4.7333 | 95.2667 | 97 | 97.3 | 97.3833 | 7.4 | 95.325 | 97.08 | 97.3 | 97.3 | 10.8 | 95.45 | 97 | 97.325 | 97.4
3.5 | 2.4333 | 95.9 | 97.1667 | 97.5163 | 97.4 | 2.925 | 95.575 | 97.18 | 97.2838 | 97.4167 | 4.325 | 95 | 96.9 | 97.2 | 97.3 | 6.75 | 95.1 | 96.86 | 97.1167 | 97.2 | 9.6 | 95.2 | 96.9 | 97.2667 | 97.3
3.6 | 2.2667 | 95.55 | 97.1 | 97.4162 | 97.35 | 2.75 | 95.3 | 97.05 | 97.2005 | 97.3714 | 4.05 | 94.9 | 96.8 | 97.1333 | 97.225 | 6.2 | 94.9333 | 96.8 | 97 | 97.0833 | 8.7 | 95 | 96.8 | 97.2 | 97.2
3.7 | 2.1 | 95.25 | 97 | 97.3162 | 97.3 | 2.5 | 95.1 | 97 | 97.1438 | 97.2375 | 3.775 | 94.55 | 96.7 | 97.0667 | 97.1 | 5.75 | 94.7 | 96.7 | 96.9 | 97 | 7.98 | 94.8333 | 96.6 | 97.0333 | 97.05
3.8 | 2 | 94.9 | 97 | 97.3162 | 97.2 | 2.3333 | 94.65 | 96.9 | 96.9838 | 97.2 | 3.45 | 94.325 | 96.6 | 97 | 97 | 5.375 | 94.4 | 96.625 | 96.8833 | 96.9 | 7.5 | 94.5333 | 96.5167 | 96.9 | 97
3.9 | 1.9 | 94.625 | 96.9 | 97.2162 | 97.2 | 2.2 | 94.3 | 96.7714 | 96.9339 | 97.1 | 3.24 | 94.0333 | 96.5 | 96.9333 | 96.9 | 5.05 | 94.1 | 96.44 | 96.775 | 96.8 | 7 | 94.4 | 96.4 | 96.84 | 96.9
4 | 1.8 | 94.2333 | 96.8 | 97.2162 | 97.1 | 2.1 | 93.8667 | 96.6 | 96.8089 | 97.0333 | 3.1 | 93.65 | 96.425 | 96.86 | 96.8 | 4.84 | 93.75 | 96.3 | 96.7 | 96.7 | 6.5333 | 94 | 96.3 | 96.7 | 96.8
4.1 | 1.7 | 93.7 | 96.8 | 97.1562 | 97.04 | 2 | 93.4 | 96.6 | 96.7339 | 96.975 | 2.9 | 93.075 | 96.3 | 96.78 | 96.7 | 4.6 | 93.5 | 96.1333 | 96.6 | 96.6 | 6.2 | 93.85 | 96.1 | 96.6 | 96.6667
4.2 | 1.7 | 93 | 96.7 | 97.0162 | 96.9667 | 1.9 | 92.8 | 96.5 | 96.6839 | 96.8429 | 2.8 | 92.825 | 96.2 | 96.6857 | 96.68 | 4.5 | 93.1 | 96 | 96.5 | 96.5333 | 5.9 | 93.5 | 96 | 96.5 | 96.6
4.3 | 1.6 | 92.05 | 96.55 | 96.9162 | 96.9 | 1.8 | 92.1 | 96.4667 | 96.6439 | 96.7833 | 2.65 | 92.25 | 96.1 | 96.5429 | 96.6 | 4.2 | 92.85 | 95.9 | 96.4 | 96.45 | 5.7667 | 93.275 | 95.9167 | 96.4 | 96.4333
4.4 | 1.6 | 90.9 | 96.425 | 96.8962 | 96.8 | 1.7 | 91.1 | 96.4 | 96.5672 | 96.7 | 2.5 | 91.7 | 96 | 96.4833 | 96.5 | 4.1 | 92.4 | 95.8 | 96.3 | 96.32 | 5.5 | 93 | 95.7667 | 96.4 | 96.4
4.5 | 1.5 | 89.2 | 96.325 | 96.8161 | 96.8 | 1.6 | 89.825 | 96.24 | 96.4839 | 96.6 | 2.4 | 90.8 | 95.8167 | 96.3167 | 96.4 | 3.95 | 92.0333 | 95.64 | 96.22 | 96.3 | 5.3 | 92.54 | 95.66 | 96.25 | 96.35
4.6 | 1.4 | 87.2 | 96.2 | 96.7761 | 96.7 | 1.6 | 88.5 | 96.14 | 96.4039 | 96.5571 | 2.3 | 90.04 | 95.8 | 96.2571 | 96.3 | 3.825 | 91.6 | 95.5 | 96.1333 | 96.1667 | 5.125 | 92.3 | 95.56 | 96.2 | 96.26
4.7 | 1.4 | 83.1667 | 96.2 | 96.6994 | 96.6 | 1.5 | 86.25 | 96.02 | 96.284 | 96.5 | 2.2667 | 89 | 95.6667 | 96.2 | 96.2 | 3.675 | 91.1333 | 95.4 | 96 | 96.075 | 4.95 | 92 | 95.45 | 96.1 | 96.1833
4.8 | 1.3 | 77.2333 | 96.1 | 96.6161 | 96.55 | 1.5 | 83.2 | 95.9 | 96.234 | 96.3714 | 2.15 | 87.8 | 95.5 | 96.0833 | 96.2 | 3.5 | 90.5 | 95.2 | 95.9 | 95.96 | 4.775 | 91.6 | 95.3 | 96 | 96.0167
4.9 | 1.3 | 69.1 | 96 | 96.5328 | 96.4 | 1.4 | 78.2 | 95.8 | 96.1507 | 96.3 | 2.1 | 85.7 | 95.5 | 96 | 96.1 | 3.3333 | 89.95 | 95.2 | 95.8 | 95.82 | 4.625 | 91.2 | 95.2 | 95.9 | 95.9
5 | 1.3 | 60.0667 | 95.9 | 96.5161 | 96.4 | 1.4 | 70.9667 | 95.76 | 96.084 | 96.175 | 2 | 83.175 | 95.4 | 95.9 | 96 | 3.2 | 89.225 | 95 | 95.6833 | 95.775 | 4.5 | 90.7 | 95 | 95.7857 | 95.7833
5.1 | 1.2 | 49.5 | 95.8 | 96.3661 | 96.25 | 1.3 | 62.4 | 95.6 | 95.9173 | 96.1 | 2 | 79.5 | 95.2333 | 95.7667 | 95.9 | 3.1 | 88.2 | 94.9 | 95.5 | 95.7 | 4.4 | 90.3333 | 94.88 | 95.6333 | 95.7
5.2 | 1.2 | 39.8 | 95.7 | 96.241 | 96.2 | 1.3 | 52.7 | 95.5 | 95.8507 | 96 | 1.9 | 73.6667 | 95.2 | 95.7 | 95.8 | 3 | 87.1 | 94.8 | 95.4 | 95.5833 | 4.275 | 89.8 | 94.75 | 95.5333 | 95.6
5.3 | 1.2 | 30.4 | 95.6 | 96.116 | 96.1 | 1.2 | 42.7 | 95.38 | 95.784 | 95.9 | 1.8 | 66.8 | 95 | 95.55 | 95.7 | 2.95 | 85.8 | 94.7 | 95.32 | 95.4167 | 4.1 | 89.4 | 94.6 | 95.4667 | 95.44
5.4 | 1.1 | 22.475 | 95.5 | 96.016 | 96 | 1.2 | 34.325 | 95.2833 | 95.684 | 95.8167 | 1.7 | 58.3 | 94.85 | 95.46 | 95.52 | 2.9 | 84 | 94.6 | 95.3 | 95.3 | 4 | 88.5 | 94.5 | 95.3 | 95.3667
5.5 | 1.1 | 17.0667 | 95.44 | 95.916 | 95.9 | 1.2 | 26.0667 | 95.0833 | 95.5841 | 95.74 | 1.7 | 49 | 94.78 | 95.4 | 95.5 | 2.8 | 81.4 | 94.44 | 95.175 | 95.2 | 3.9 | 87.9 | 94.4 | 95.2 | 95.3
5.6 | 1.1 | 13.7 | 95.225 | 95.836 | 95.86 | 1.1 | 19.85 | 95 | 95.4555 | 95.6 | 1.7 | 39.7 | 94.6833 | 95.3 | 95.4 | 2.7 | 78.05 | 94.3 | 95.0667 | 95.12 | 3.775 | 87.025 | 94.26 | 95.0667 | 95.14
5.7 | 1 | 11.4 | 95.1 | 95.7359 | 95.74 | 1.1 | 15.8667 | 94.82 | 95.2841 | 95.525 | 1.6 | 31.6 | 94.5 | 95.2167 | 95.3 | 2.62 | 72.8333 | 94.2 | 95 | 95.02 | 3.7 | 85.95 | 94.05 | 95 | 95.05
5.8 | 1 | 9.8 | 95 | 95.6409 | 95.6 | 1.1 | 13.2 | 94.7 | 95.1675 | 95.3833 | 1.6 | 24.9 | 94.375 | 95.0889 | 95.2 | 2.6 | 66.225 | 94 | 94.8833 | 94.8571 | 3.5667 | 84.5 | 93.9 | 94.875 | 95
5.9 | 1 | 8.7 | 94.9 | 95.5159 | 95.5 | 1.1 | 11.4 | 94.62 | 95.0842 | 95.3 | 1.5 | 19.8 | 94.15 | 94.9714 | 95.0333 | 2.5 | 58.75 | 93.86 | 94.78 | 94.7857 | 3.4 | 82.6 | 93.75 | 94.75 | 94.8833
6 | 1 | 7.8 | 94.75 | 95.4159 | 95.4 | 1 | 9.9333 | 94.52 | 94.9467 | 95.2 | 1.5 | 16.8 | 94.05 | 94.8444 | 94.9 | 2.4 | 50.175 | 93.65 | 94.68 | 94.7 | 3.4 | 80.2667 | 93.56 | 94.6833 | 94.7
6.1 | 0.93333 | 7.1667 | 94.6 | 95.2492 | 95.3 | 1 | 9 | 94.3333 | 94.8009 | 95.0667 | 1.4 | 14.45 | 93.9 | 94.725 | 94.8 | 2.3667 | 42.175 | 93.4 | 94.48 | 94.55 | 3.3 | 77.2 | 93.4667 | 94.5167 | 94.58
6.2 | 0.9 | 6.72 | 94.5 | 95.2159 | 95.2 | 1 | 8.1333 | 94.25 | 94.7413 | 94.9833 | 1.4 | 12.8 | 93.8 | 94.6 | 94.675 | 2.3 | 34.4 | 93.26 | 94.4 | 94.46 | 3.3 | 72.4 | 93.3 | 94.35 | 94.45
6.3 | 0.9 | 6.3 | 94.4 | 95.0658 | 95.05 | 1 | 7.5 | 94.1 | 94.6842 | 94.8167 | 1.4 | 11.6 | 93.6 | 94.5167 | 94.58 | 2.2 | 28.3 | 93.0667 | 94.3 | 94.325 | 3.2 | 66.6 | 93.1 | 94.2833 | 94.32
6.4 | 0.9 | 5.9 | 94.225 | 95.0158 | 94.9 | 1 | 6.9 | 93.9 | 94.5675 | 94.7 | 1.3 | 10.5667 | 93.4 | 94.3429 | 94.45 | 2.2 | 23.5 | 92.9143 | 94.2 | 94.2 | 3.1 | 59.1 | 92.9 | 94.2 | 94.2
6.5 | 0.9 | 5.6 | 94.1 | 94.9158 | 94.8 | 1 | 6.5333 | 93.78 | 94.4843 | 94.5667 | 1.3 | 9.775 | 93.25 | 94.2857 | 94.3 | 2.1 | 20.2 | 92.8 | 94.0833 | 94 | 3 | 51.4 | 92.8 | 94.05 | 94.14
6.6 | 0.9 | 5.4 | 93.9333 | 94.8158 | 94.7 | 0.9 | 6.1 | 93.6 | 94.351 | 94.5 | 1.3 | 9.2 | 93.1 | 94.12 | 94.2 | 2.1 | 17.9 | 92.6167 | 93.9167 | 93.9 | 2.9 | 43.9 | 92.58 | 93.9667 | 94.0667
6.7 | 0.9 | 5.2 | 93.82 | 94.7158 | 94.62 | 0.9 | 5.8333 | 93.46 | 94.1843 | 94.4 | 1.2 | 8.6 | 92.9 | 94 | 94.1 | 2 | 16.1 | 92.45 | 93.8 | 93.8 | 2.8 | 37.15 | 92.36 | 93.8 | 93.9
6.8 | 0.9 | 4.9667 | 93.68 | 94.5825 | 94.5 | 0.9 | 5.5 | 93.3 | 94.0557 | 94.2833 | 1.2 | 8.2 | 92.6833 | 93.9167 | 94 | 2 | 14.84 | 92.2833 | 93.6333 | 93.7 | 2.8 | 31.125 | 92.2 | 93.625 | 93.8
6.9 | 0.8 | 4.8 | 93.58 | 94.4824 | 94.38 | 0.9 | 5.275 | 93.16 | 93.9843 | 94.1167 | 1.2 | 7.8333 | 92.5833 | 93.8 | 93.9 | 1.9 | 13.7 | 92 | 93.5667 | 93.6 | 2.8 | 26.275 | 92 | 93.48 | 93.6286
7 | 0.8 | 4.6667 | 93.4 | 94.3157 | 94.2667 | 0.9 | 5.0667 | 92.95 | 93.8844 | 94 | 1.2 | 7.4 | 92.4167 | 93.7 | 93.76 | 1.9 | 12.9333 | 91.9 | 93.48 | 93.38 | 2.8 | 22.7 | 91.8 | 93.4 | 93.4333
7.1 | 0.8 | 4.5 | 93.2 | 94.2157 | 94.1 | 0.9 | 4.8667 | 92.8 | 93.7594 | 93.8857 | 1.1 | 7.15 | 92.225 | 93.6 | 93.6 | 1.8 | 12.2667 | 91.75 | 93.3 | 93.3 | 2.7 | 20.225 | 91.6167 | 93.3 | 93.22
7.2 | 0.8 | 4.3 | 93.02 | 94.0957 | 93.9833 | 0.8 | 4.7 | 92.7 | 93.6844 | 93.7333 | 1.1 | 6.8 | 92.1 | 93.52 | 93.5 | 1.8 | 11.625 | 91.5 | 93.12 | 93.125 | 2.625 | 18.4667 | 91.42 | 93.1667 | 93.1
7.3 | 0.8 | 4.2 | 92.86 | 93.9907 | 93.9 | 0.8 | 4.5333 | 92.55 | 93.5344 | 93.6667 | 1.1 | 6.55 | 91.8 | 93.36 | 93.3 | 1.8 | 11.1333 | 91.28 | 92.9833 | 92.98 | 2.575 | 17.1333 | 91.32 | 93 | 93.1
7.4 | 0.8 | 4.1 | 92.725 | 93.9157 | 93.725 | 0.8 | 4.4 | 92.34 | 93.3719 | 93.5 | 1.1 | 6.35 | 91.66 | 93.1444 | 93.2 | 1.7333 | 10.6 | 91.075 | 92.9 | 92.8 | 2.575 | 16 | 91.0167 | 92.82 | 92.9
7.5 | 0.8 | 4 | 92.6 | 93.7323 | 93.6 | 0.8 | 4.2 | 92.15 | 93.2345 | 93.4 | 1.1 | 6.1 | 91.5 | 92.9714 | 93.12 | 1.7 | 10.3 | 90.8667 | 92.8 | 92.72 | 2.5 | 15.1667 | 90.8333 | 92.7 | 92.74
7.6 | 0.8 | 3.9 | 92.48 | 93.6156 | 93.5 | 0.8 | 4.125 | 92.025 | 93.1845 | 93.3 | 1 | 5.9 | 91.275 | 92.9 | 93.02 | 1.7 | 9.9667 | 90.6 | 92.7 | 92.6 | 2.5 | 14.45 | 90.66 | 92.68 | 92.625
7.7 | 0.8 | 3.8 | 92.275 | 93.5156 | 93.4 | 0.8 | 4 | 91.82 | 93.0345 | 93.2125 | 1 | 5.7 | 91.1 | 92.8 | 92.875 | 1.6 | 9.6 | 90.4 | 92.45 | 92.5 | 2.4 | 13.8 | 90.4333 | 92.45 | 92.4571
7.8 | 0.7 | 3.7 | 92.025 | 93.3656 | 93.38 | 0.8 | 3.9 | 91.58 | 92.8845 | 93.08 | 1 | 5.5667 | 90.9 | 92.7143 | 92.725 | 1.6 | 9.3 | 90.2333 | 92.3 | 92.375 | 2.4 | 13.3 | 90.2 | 92.3 | 92.4
7.9 | 0.7 | 3.5667 | 91.88 | 93.1955 | 93.15 | 0.8 | 3.8 | 91.46 | 92.6845 | 92.9833 | 1 | 5.4 | 90.775 | 92.5571 | 92.5333 | 1.6 | 9.1 | 90.0333 | 92.26 | 92.25 | 2.35 | 12.8 | 89.9333 | 92.16 | 92.2333
8 | 0.7 | 3.5 | 91.68 | 93.1155 | 93.1 | 0.76667 | 3.7 | 91.25 | 92.5846 | 92.8167 | 1 | 5.2 | 90.525 | 92.4143 | 92.5 | 1.5 | 8.8333 | 89.88 | 92.1 | 92.1 | 2.3 | 12.4667 | 89.82 | 92.05 | 92
8.1 | 0.7 | 3.4 | 91.475 | 93.0155 | 92.9333 | 0.7 | 3.6 | 91.1 | 92.4846 | 92.72 | 1 | 5.1 | 90.2 | 92.2714 | 92.3333 | 1.5 | 8.5 | 89.675 | 91.875 | 91.95 | 2.2 | 12 | 89.5833 | 91.8667 | 91.92
8.2 | 0.7 | 3.3 | 91.3333 | 92.8555 | 92.8 | 0.7 | 3.4333 | 90.8667 | 92.3846 | 92.5667 | 1 | 4.9 | 90.0833 | 92.1167 | 92.2 | 1.5 | 8.3 | 89.44 | 91.75 | 91.8 | 2.2 | 11.7 | 89.4 | 91.8 | 91.8
8.3 | 0.7 | 3.2 | 91.0667 | 92.7154 | 92.675 | 0.7 | 3.4 | 90.54 | 92.2646 | 92.3833 | 1 | 4.7667 | 89.9 | 91.9571 | 92.05 | 1.5 | 8 | 89.1857 | 91.64 | 91.6833 | 2.2 | 11.36 | 89.1333 | 91.6 | 91.72
8.4 | 0.7 | 3.2 | 90.8 | 92.5904 | 92.525 | 0.7 | 3.2667 | 90.32 | 92.099 | 92.3 | 1 | 4.6667 | 89.575 | 91.825 | 91.9 | 1.4667 | 7.825 | 88.88 | 91.5 | 91.5167 | 2.1 | 11.1 | 88.9 | 91.4 | 91.6
8.5 | 0.7 | 3.1 | 90.7 | 92.4154 | 92.375 | 0.7 | 3.2 | 90.175 | 91.9561 | 92.1429 | 0.9 | 4.5333 | 89.425 | 91.6857 | 91.775 | 1.4 | 7.55 | 88.675 | 91.4 | 91.3 | 2.1 | 10.9 | 88.7667 | 91.26 | 91.4
8.6 | 0.7 | 3.1 | 90.38 | 92.2654 | 92.28 | 0.7 | 3.1 | 89.9 | 91.8133 | 92.0125 | 0.9 | 4.4 | 89.08 | 91.52 | 91.6 | 1.4 | 7.5 | 88.5333 | 91.2 | 91.2 | 2 | 10.5 | 88.5 | 91.125 | 91.2333
8.7 | 0.7 | 3 | 90.15 | 92.2154 | 92.2 | 0.7 | 3 | 89.64 | 91.668 | 91.9 | 0.9 | 4.3 | 88.8 | 91.4 | 91.42 | 1.4 | 7.2667 | 88.26 | 91.05 | 91.04 | 2 | 10.2667 | 88.3333 | 91 | 91.16
8.8 | 0.7 | 2.9 | 89.84 | 92.0904 | 92.1 | 0.7 | 3 | 89.325 | 91.5847 | 91.8 | 0.9 | 4.2 | 88.6 | 91.2833 | 91.3 | 1.375 | 7.075 | 87.92 | 90.9 | 90.94 | 2 | 10.025 | 88.0667 | 90.85 | 91
8.9 | 0.7 | 2.8 | 89.55 | 91.9403 | 91.94 | 0.7 | 2.9 | 89.075 | 91.4348 | 91.62 | 0.9 | 4.1 | 88.3 | 91.1 | 91.2 | 1.3 | 6.8667 | 87.6 | 90.72 | 90.72 | 1.925 | 9.7 | 87.8 | 90.7 | 90.85
9 | 0.7 | 2.8 | 89.22 | 91.9153 | 91.8 | 0.7 | 2.8 | 88.825 | 91.2848 | 91.4429 | 0.9 | 3.975 | 87.9667 | 90.95 | 91.1 | 1.3 | 6.7 | 87.42 | 90.6 | 90.6 | 1.9 | 9.6 | 87.58 | 90.6 | 90.62
9.1 |  | 2.8 | 88.9 | 91.8153 | 91.725 | 0.7 | 2.8 | 88.3 | 91.1515 | 91.3 | 0.8 | 3.9 | 87.74 | 90.8857 | 90.9 | 1.3 | 6.5 | 87.2667 | 90.4167 | 90.48 | 1.9 | 9.3667 | 87.2333 | 90.44 | 90.46
9.2 |  | 2.7 | 88.3333 | 91.7153 | 91.65 | 0.7 | 2.7 | 88 | 91.0848 | 91.2375 | 0.8 | 3.8 | 87.4 | 90.7429 | 90.8 | 1.3 | 6.4 | 86.9 | 90.25 | 90.3833 | 1.8333 | 9.2 | 87.06 | 90.3 | 90.36
9.3 |  | 2.7 | 87.94 | 91.6153 | 91.6 | 0.7 | 2.6 | 87.55 | 91.0562 | 91.2 | 0.8 | 3.7 | 87.02 | 90.5714 | 90.7 | 1.2 | 6.3 | 86.575 | 90.2 | 90.175 | 1.8 | 8.975 | 86.74 | 90.06 | 90.2
9.4 |  | 2.6 | 87.275 | 91.5753 | 91.5 | 0.7 | 2.6 | 87 | 90.9848 | 91.02 | 0.8 | 3.6 | 86.65 | 90.5 | 90.5833 | 1.2 | 6.15 | 86.325 | 89.9667 | 90.025 | 1.8 | 8.8 | 86.52 | 89.925 | 90.0667
9.5 |  | 2.5 | 86.375 | 91.5153 | 91.5 |  | 2.5 | 86.525 | 90.8182 | 91 | 0.8 | 3.5 | 86.26 | 90.35 | 90.3833 | 1.2 | 6 | 86.12 | 89.8 | 89.8 | 1.8 | 8.6333 | 86.4 | 89.775 | 89.88
9.6 |  | 2.5 | 85.1 | 91.5153 | 91.4 |  | 2.5 | 85.85 | 90.7849 | 90.9 | 0.8 | 3.4667 | 85.95 | 90.2 | 90.2667 | 1.2 | 5.9 | 85.775 | 89.72 | 89.65 | 1.7 | 8.4 | 86.1167 | 89.5571 | 89.675
9.7 |  | 2.4 | 83.3 | 91.4652 | 91.4 |  | 2.5 | 84.98 | 90.6849 | 90.9 | 0.8 | 3.4 | 85.54 | 90.12 | 90.125 | 1.2 | 5.7 | 85.42 | 89.5 | 89.5 | 1.7 | 8.25 | 85.8 | 89.5 | 89.5
9.8 |  | 2.4 | 80.7 | 91.4152 | 91.3 |  | 2.4 | 83.88 | 90.6182 | 90.8 | 0.8 | 3.3 | 85.08 | 90.0429 | 90 | 1.2 | 5.625 | 85.18 | 89.3 | 89.3 | 1.7 | 8.1 | 85.58 | 89.3 | 89.4
9.9 |  | 2.4 | 76.325 | 91.2752 | 91.2 |  | 2.4 | 82.44 | 90.5849 | 90.75 |  | 3.2333 | 84.575 | 89.9 | 89.925 | 1.2 | 5.5 | 84.85 | 89.2 | 89.1 | 1.7 | 7.9 | 85.28 | 89.12 | 89.2
10 |  | 2.3 | 70 | 91.2152 | 91.2 |  | 2.3 | 80.325 | 90.5016 | 90.7 |  | 3.2 | 84.0333 | 89.8333 | 89.9 | 1.2 | 5.3333 | 84.5167 | 89 | 89.025 | 1.7 | 7.8 | 84.975 | 89 | 89
10.1 |  | 2.3 | 62.35 | 91.2152 | 91.125 |  | 2.3 | 76.9 | 90.4563 | 90.6444 |  | 3.1 | 83.4 | 89.7 | 89.78 | 1.1 | 5.3 | 84.15 | 88.9 | 88.9 | 1.6 | 7.6 | 84.65 | 88.8833 | 88.9
10.2 |  | 2.3 | 53.3 | 91.1952 | 91.1 |  | 2.2 | 71.9 | 90.3849 | 90.4857 |  | 3 | 82.42 | 89.6 | 89.6 | 1.1 | 5.1333 | 83.82 | 88.8 | 88.7667 | 1.6 | 7.425 | 84.38 | 88.7167 | 88.7167
10.3 |  | 2.2 | 43.9333 | 91.1152 | 91 |  | 2.2 | 64.9 | 90.3474 | 90.4 |  | 3 | 81.1 | 89.6 | 89.6 | 1.1 | 5.1 | 83.4667 | 88.6 | 88.68 | 1.6 | 7.3 | 84.075 | 88.5 | 88.5833
10.4 |  | 2.2 | 34.9 | 91.0952 | 90.92 |  | 2.2 | 57.2 | 90.285 | 90.4 |  | 2.925 | 79.76 | 89.5 | 89.5 | 1.1 | 4.9333 | 83.1 | 88.45 | 88.5 | 1.6 | 7.2333 | 83.7333 | 88.38 | 88.4167
10.5 |  | 2.2 | 27.3 | 91.0152 | 90.9 |  | 2.1667 | 48.36 | 90.185 | 90.3 |  | 2.9 | 77.6 | 89.4 | 89.4667 | 1.1 | 4.85 | 82.6 | 88.36 | 88.4 | 1.6 | 7.1 | 83.4 | 88.2 | 88.275
10.6 |  | 2.1 | 21.2 | 90.9152 | 90.875 |  | 2.1 | 39.2 | 90.1421 | 90.2875 |  | 2.8 | 74.925 | 89.3429 | 89.4 | 1.1 | 4.72 | 82.2667 | 88.14 | 88.2 | 1.6 | 6.9333 | 83.1 | 88.02 | 88.1
10.7 |  | 2.1 | 17.54 | 90.8151 | 90.8 |  | 2.1 | 31.5667 | 90.085 | 90.2 |  | 2.8 | 70.25 | 89.2 | 89.25 | 1.1 | 4.7 | 81.8 | 87.95 | 88 | 1.6 | 6.8 | 82.84 | 87.875 | 87.9
10.8 |  | 2.1 | 15.2333 | 90.8151 | 90.7 |  | 2.1 | 25.1 | 89.985 | 90.1 |  | 2.8 | 64.0333 | 89.1167 | 89.1833 |  | 4.6 | 81.22 | 87.8 | 87.9 | 1.5 | 6.6 | 82.6 | 87.66 | 87.82
10.9 |  | 2.1 | 13.575 | 90.7151 | 90.7 |  | 2 | 20 | 89.985 | 90.05 |  | 2.7 | 57.2 | 89.05 | 89.1 |  | 4.5 | 80.8 | 87.8 | 87.7167 | 1.5 | 6.5 | 82.3 | 87.4857 | 87.5
11 |  | 2 | 12.4 | 90.7151 | 90.6 |  | 2 | 17.1 | 89.885 | 90 |  | 2.7 | 48.8833 | 89 | 89 |  | 4.4 | 80.14 | 87.7 | 87.62 | 1.5 | 6.4 | 81.88 | 87.3333 | 87.35
11.1 |  | 2 | 11.5 | 90.6151 | 90.5 |  | 2 | 15.2 | 89.845 | 89.9167 |  | 2.6 | 41.2286 | 88.9375 | 88.9 |  | 4.3 | 79.325 | 87.5667 | 87.4667 | 1.5 | 6.3 | 81.625 | 87.2333 | 87.16
11.2 |  | 2 | 10.9 | 90.6151 | 90.475 |  | 1.9 | 13.6 | 89.785 | 89.9 |  | 2.6 | 33.84 | 88.8 | 88.8 |  | 4.3 | 78.4 | 87.4 | 87.3 | 1.4 | 6.2 | 81.24 | 87.0429 | 87.08
11.3 |  | 2 | 10.35 | 90.5151 | 90.4 |  | 1.9 | 12.5 | 89.685 | 89.8 |  | 2.5 | 27.6 | 88.7 | 88.7 |  | 4.1333 | 77.36 | 87.3 | 87.26 | 1.4 | 6.1 | 80.8 | 86.8167 | 86.8833
11.4 |  | 2 | 9.8 | 90.4151 | 90.3 |  | 1.9 | 11.7 | 89.5851 | 89.76 |  | 2.5 | 22.86 | 88.6 | 88.6 |  | 4.1 | 75.9 | 87.1 | 87.2 | 1.4 | 5.9 | 80.4 | 86.7 | 86.7
11.5 |  | 1.9 | 9.5 | 90.4151 | 90.3 |  | 1.9 | 11.025 | 89.5851 | 89.6833 |  | 2.5 | 19.6 | 88.5429 | 88.525 |  | 4 | 73.9 | 87 | 87.075 | 1.4 | 5.8 | 80.1571 | 86.55 | 86.575
11.6 |  | 1.9 | 9.275 | 90.3151 | 90.2 |  | 1.9 | 10.48 | 89.4851 | 89.6 |  | 2.4 | 17.5 | 88.4125 | 88.5 |  | 3.9 | 71.2333 | 86.9 | 86.9 | 1.4 | 5.7 | 79.8 | 86.3833 | 86.4667
11.7 |  | 1.9 | 8.95 | 90.2951 | 90.2 |  | 1.9 | 10.025 | 89.3976 | 89.5 |  | 2.4 | 15.9667 | 88.3 | 88.38 |  | 3.9 | 67.125 | 86.8167 | 86.8 |  | 5.575 | 79.1 | 86.2 | 86.35
11.8 |  | 1.9 | 8.72 | 90.215 | 90.1 |  | 1.8 | 9.65 | 89.2851 | 89.4857 |  | 2.4 | 14.9 | 88.2143 | 88.3 |  | 3.8 | 61.7 | 86.7 | 86.6667 |  | 5.475 | 78.6 | 86.05 | 86.2
11.9 |  | 1.8 | 8.45 | 90.115 | 90.025 |  | 1.8 | 9.34 | 89.2851 | 89.3429 |  | 2.3 | 13.925 | 88.075 | 88.1 |  | 3.7 | 55.5 | 86.575 | 86.58 |  | 5.4 | 78.1 | 85.9667 | 85.975
12 |  | 1.8 | 8.24 | 90.065 | 89.92 |  | 1.8 | 9 | 89.1851 | 89.3 |  | 2.3 | 13.24 | 87.95 | 88.0667 |  | 3.7 | 48.4 | 86.44 | 86.4143 |  | 5.3 | 77.52 | 85.8 | 85.825
12.1 |  | 1.8 | 8.1 | 90.015 | 89.9 |  | 1.7 | 8.7333 | 89.0852 | 89.2 |  | 2.2 | 12.575 | 87.9 | 87.96 |  | 3.6 | 41.7 | 86.3 | 86.2 |  | 5.3 | 76.8 | 85.7 | 85.7
12.2 |  |  | 7.9 | 89.955 | 89.8 |  | 1.7 | 8.44 | 88.9852 | 89.1375 |  | 2.2 | 12.06 | 87.8333 | 87.8 |  | 3.6 | 35.2 | 86.1833 | 86.1667 |  | 5.2 | 75.95 | 85.56 | 85.5667
12.3 |  |  | 7.7 | 89.815 | 89.7 |  | 1.7 | 8.24 | 88.9685 | 89.0833 |  | 2.1667 | 11.6667 | 87.6667 | 87.7 |  | 3.5 | 30.1 | 86 | 86.1 |  | 5.1 | 74.825 | 85.425 | 85.45
12.4 |  |  | 7.6 | 89.795 | 89.64 |  | 1.7 | 8 | 88.8852 | 88.9 |  | 2.1 | 11.275 | 87.6 | 87.6 |  | 3.4 | 25.8333 | 85.95 | 85.9 |  | 5.0667 | 73.7 | 85.28 | 85.325
12.5 |  |  | 7.375 | 89.7149 | 89.6 |  | 1.7 | 7.825 | 88.7185 | 88.8 |  | 2.1 | 10.975 | 87.5 | 87.5 |  | 3.4 | 22.9 | 85.74 | 85.8 |  | 5 | 71.7 | 85.1 | 85.1333
12.6 |  |  | 7.26 | 89.6149 | 89.5 |  |  | 7.65 | 88.6452 | 88.7333 |  | 2.1 | 10.625 | 87.32 | 87.4 |  | 3.3 | 20.85 | 85.6333 | 85.6 |  | 4.9 | 69.2 | 85 | 85
12.7 |  |  | 7.125 | 89.5149 | 89.4 |  |  | 7.45 | 88.5686 | 88.6667 |  | 2.1 | 10.3143 | 87.24 | 87.3 |  | 3.3 | 19.45 | 85.5 | 85.5 |  | 4.8 | 66.0333 | 84.9 | 84.9167
12.8 |  |  | 7 | 89.4149 | 89.3 |  |  | 7.35 | 88.402 | 88.5143 |  | 2.1 | 10.0667 | 87.1667 | 87.2 |  | 3.3 | 18.1 | 85.3667 | 85.3333 |  | 4.8 | 61.2 | 84.7833 | 84.74
12.9 |  |  | 6.88 | 89.3149 | 89.1667 |  |  | 7.1333 | 88.3103 | 88.46 |  | 2 | 9.8 | 87 | 87 |  | 3.2 | 17.275 | 85.28 | 85.1667 |  | 4.7 | 55.72 | 84.6167 | 84.6
13 |  |  | 6.7 | 89.2149 | 89.1 |  |  | 7.0333 | 88.2853 | 88.3 |  | 2 | 9.575 | 86.9 | 86.9 |  | 3.1667 | 16.6 | 85.15 | 85 |  | 4.6 | 49.82 | 84.45 | 84.5
13.1 |  |  | 6.56 | 89.2149 | 88.925 |  |  | 6.9 | 88.152 | 88.2 |  | 2 | 9.34 | 86.8 | 86.725 |  | 3.1 | 15.88 | 84.92 | 84.95 |  | 4.6 | 43.5333 | 84.3833 | 84.3667
13.2 |  |  | 6.44 | 89.0148 | 88.9 |  |  | 6.7 | 87.9853 | 88.1 |  |  | 9.2 | 86.65 | 86.625 |  | 3.1 | 15.3667 | 84.75 | 84.8 |  | 4.5 | 37.96 | 84.2 | 84.2
13.3 |  |  | 6.35 | 88.9148 | 88.8 |  |  | 6.575 | 87.8854 | 88.0286 |  |  | 8.9833 | 86.4833 | 86.5 |  | 3 | 14.94 | 84.6 | 84.675 |  | 4.45 | 32.92 | 84.0333 | 83.925
13.4 |  |  | 6.2 | 88.8148 | 88.7 |  |  | 6.45 | 87.7054 | 87.9 |  |  | 8.7333 | 86.3 | 86.38 |  | 3 | 14.4 | 84.4667 | 84.45 |  | 4.4 | 28.9667 | 83.8667 | 83.8
13.5 |  |  | 6.1 | 88.7148 | 88.6 |  |  | 6.325 | 87.6854 | 87.74 |  |  | 8.55 | 86.3 | 86.2 |  | 3 | 14.06 | 84.3 | 84.34 |  | 4.3 | 25.8833 | 83.7 | 83.66
13.6 |  |  | 6 | 88.6148 | 88.5 |  |  | 6.2 | 87.5654 | 87.64 |  |  | 8.375 | 86.1167 | 86.075 |  | 2.9 | 13.8 | 84.2 | 84.12 |  | 4.2 | 23.7 | 83.5333 | 83.5
13.7 |  |  | 5.9 | 88.5148 | 88.325 |  |  | 6.1 | 87.4854 | 87.4667 |  |  | 8.18 | 85.92 | 85.8667 |  | 2.9 | 13.475 | 84.15 | 84 |  | 4.2 | 21.98 | 83.4 | 83.4
13.8 |  |  | 5.8 | 88.3747 | 88.22 |  |  | 5.92 | 87.3021 | 87.38 |  |  | 7.9833 | 85.75 | 85.78 |  | 2.9 | 13.18 | 83.98 | 83.8833 |  | 4.1 | 20.8 | 83.2 | 83.2667
13.9 |  |  | 5.7 | 88.1897 | 88 |  |  | 5.82 | 87.1055 | 87.18 |  |  | 7.9 | 85.6125 | 85.575 |  | 2.8 | 12.88 | 83.7333 | 83.7 |  | 4 | 19.8 | 83.1 | 83.05
14 |  |  | 5.6 | 87.9747 | 87.875 |  |  | 5.7 | 86.9355 | 87.06 |  |  | 7.725 | 85.48 | 85.3833 |  | 2.8 | 12.66 | 83.525 | 83.5167 |  | 4 | 19.1 | 83.0167 | 82.88
14.1 |  |  | 5.5 | 87.7396 | 87.65 |  |  | 5.6 | 86.7455 | 86.7833 |  |  | 7.5333 | 85.28 | 85.3 |  | 2.7 | 12.45 | 83.45 | 83.42 |  | 3.95 | 18.3857 | 82.75 | 82.7
14.2 |  |  | 5.46 | 87.4146 | 87.4 |  |  | 5.56 | 86.557 | 86.6 |  |  | 7.44 | 85.1 | 85.0167 |  | 2.7 | 12.1 | 83.3 | 83.2286 |  | 3.9 | 17.85 | 82.5 | 82.6
14.3 |  |  | 5.3 | 87.1645 | 86.95 |  |  | 5.44 | 86.3856 | 86.42 |  |  | 7.25 | 84.92 | 84.9 |  | 2.7 | 11.86 | 83.15 | 83 |  | 3.825 | 17.5 | 82.4167 | 82.4667
14.4 |  |  | 5.2 | 86.6944 | 86.6 |  |  | 5.4 | 86.0523 | 86.0857 |  |  | 7.15 | 84.725 | 84.7 |  |  | 11.625 | 82.95 | 82.9 |  | 3.8 | 17.02 | 82.225 | 82.2167
14.5 |  |  | 5.1167 | 86.2144 | 86.1 |  |  | 5.28 | 85.7857 | 85.82 |  |  | 7 | 84.5333 | 84.5667 |  |  | 11.4 | 82.8 | 82.66 |  | 3.7 | 16.6667 | 82.1 | 82.05
14.6 |  |  | 5.1 | 85.4142 | 85.24 |  |  | 5.2 | 85.4191 | 85.4333 |  |  | 6.8833 | 84.34 | 84.275 |  |  | 11.25 | 82.55 | 82.54 |  | 3.7 | 16.26 | 81.88 | 81.8857
14.7 |  |  | 5 | 84.3141 | 84.225 |  |  | 5.1 | 84.9858 | 85.04 |  |  | 6.7167 | 84.15 | 84.08 |  |  | 11.025 | 82.4 | 82.25 |  | 3.7 | 16.02 | 81.7 | 81.8
14.8 |  |  | 4.9 | 82.8972 | 82.7 |  |  | 5 | 84.4193 | 84.5 |  |  | 6.7 | 83.8833 | 83.76 |  |  | 10.8 | 82.2 | 82.06 |  | 3.625 | 15.64 | 81.575 | 81.6
14.9 |  |  | 4.8667 | 80.4134 | 80.8 |  |  | 4.9 | 83.7527 | 83.7667 |  |  | 6.5 | 83.6571 | 83.5 |  |  | 10.6 | 82.02 | 81.98 |  | 3.6 | 15.34 | 81.4 | 81.46
15 |  |  | 4.8 | 76.8628 | 77.04 |  |  | 4.8 | 82.8862 | 82.85 |  |  | 6.4 | 83.3333 | 83.25 |  |  | 10.42 | 81.7333 | 81.76 |  | 3.5 | 14.98 | 81.28 | 81.2
15.1 |  |  | 4.7 | 70.8318 | 71.675 |  |  | 4.775 | 81.6864 | 81.84 |  |  | 6.2833 | 83 | 83.0333 |  |  | 10.1714 | 81.5333 | 81.55 |  | 3.5 | 14.75 | 81.1 | 81
15.2 |  |  | 4.65 | 63.5106 | 64.64 |  |  | 4.7 | 80.2533 | 80.3429 |  |  | 6.1833 | 82.7 | 82.7 |  |  | 10 | 81.4 | 81.4 |  | 3.44 | 14.525 | 80.8 | 80.85
15.3 |  |  | 4.6 | 55.2843 | 56 |  |  | 4.6 | 77.867 | 77.7333 |  |  | 6.0286 | 82.25 | 82.175 |  |  | 9.875 | 81.12 | 81.14 |  | 3.4 | 14.3 | 80.7 | 80.6
15.4 |  |  | 4.55 | 46.0827 | 47.6 |  |  | 4.5 | 74.2076 | 73.95 |  |  | 5.9 | 81.8333 | 81.775 |  |  | 9.65 | 80.98 | 80.925 |  | 3.4 | 14.025 | 80.5167 | 80.5
15.5 |  |  | 4.5 | 37.8063 | 38.675 |  |  | 4.4667 | 68.8885 | 68.7 |  |  | 5.85 | 81.26 | 81.26 |  |  | 9.525 | 80.76 | 80.62 |  | 3.3 | 13.8333 | 80.32 | 80.34
15.6 |  |  | 4.4 | 29.905 | 31.075 |  |  | 4.4 | 61.6297 | 62.2 |  |  | 5.8 | 80.6286 | 80.5333 |  |  | 9.42 | 80.45 | 80.46 |  |  | 13.5667 | 80.1167 | 80.0667
15.7 |  |  | 4.35 | 24.1641 | 25.38 |  |  | 4.3333 | 53.771 | 53.7 |  |  | 5.7 | 79.7833 | 79.9 |  |  | 9.2 | 80.225 | 80.25 |  |  | 13.35 | 80 | 79.86
15.8 |  |  | 4.3 | 19.9434 | 21.0667 |  |  | 4.3 | 45.2758 | 45.9333 |  |  | 5.6 | 78.8 | 78.84 |  |  | 9.1 | 80.04 | 79.94 |  |  | 13.12 | 79.68 | 79.6
15.9 |  |  | 4.22 | 17.4529 | 18.5 |  |  | 4.225 | 37.5271 | 37.9 |  |  | 5.5 | 77.4667 | 77.46 |  |  | 8.96 | 79.8167 | 79.6667 |  |  | 12.875 | 79.5 | 79.4
16 |  |  | 4.2 | 15.9027 | 16.9 |  |  | 4.2 | 30.475 | 30.48 |  |  | 5.3833 | 75.6 | 75.9 |  |  | 8.75 | 79.575 | 79.5 |  |  | 12.7 | 79.3 | 79.25
16.1 |  |  | 4.2 | 14.7358 | 15.7 |  |  | 4.1 | 24.6102 | 25.05 |  |  | 5.3 | 72.6 | 72.9667 |  |  | 8.64 | 79.325 | 79.14 |  |  | 12.5 | 79.06 | 79.025
16.2 |  |  | 4.1833 | 13.8273 | 14.85 |  |  | 4.1 | 20.5966 | 21.34 |  |  | 5.2143 | 68.3333 | 68.72 |  |  | 8.44 | 78.96 | 78.9 |  |  | 12.28 | 78.8667 | 78.75
16.3 |  |  | 4.1 | 13.2022 | 14.25 |  |  | 4.05 | 18.177 | 18.8667 |  |  | 5.2 | 62.7 | 63.62 |  |  | 8.34 | 78.6 | 78.6333 |  |  | 12.08 | 78.7 | 78.58
16.4 |  |  | 4.1 | 12.7521 | 13.75 |  |  | 4 | 16.5172 | 17.2667 |  |  | 5.1333 | 56.3333 | 56.6 |  |  | 8.2 | 78.4 | 78.38 |  |  | 11.9333 | 78.48 | 78.38
16.5 |  |  | 4 | 12.2821 | 13.3 |  |  | 4 | 15.2974 | 16.1 |  |  | 5.05 | 49 | 49.9 |  |  | 8.1 | 78.1 | 77.975 |  |  | 11.7 | 78.3 | 78.2
16.6 |  |  | 4 | 11.962 | 13.02 |  |  | 3.9833 | 14.2976 | 15.2 |  |  | 4.9857 | 41.4667 | 41.84 |  |  | 7.92 | 77.74 | 77.68 |  |  | 11.56 | 78.05 | 78
16.7 |  |  | 3.9 | 11.752 | 12.65 |  |  | 3.9 | 13.6477 | 14.5 |  |  | 4.9 | 34.0571 | 35.25 |  |  | 7.8 | 77.42 | 77.24 |  |  | 11.4 | 77.825 | 77.7
16.8 |  |  | 3.9 | 11.4419 | 12.4667 |  |  | 3.8 | 13.0479 | 13.95 |  |  | 4.875 | 28.56 | 29.775 |  |  | 7.7 | 76.925 | 76.8333 |  |  | 11.2571 | 77.5667 | 77.475
16.9 |  |  | 3.8 | 11.2269 | 12.225 |  |  | 3.8 | 12.5979 | 13.46 |  |  | 4.8 | 24.1667 | 25 |  |  | 7.56 | 76.3667 | 76.38 |  |  | 11.1 | 77.32 | 77.24
17 |  |  | 3.8 | 11.0518 | 11.98 |  |  | 3.8 | 12.258 | 13.1 |  |  | 4.7 | 21.25 | 22.2833 |  |  | 7.45 | 75.775 | 75.82 |  |  | 10.95 | 77.2 | 76.98
17.1 |  |  | 3.7333 | 10.9018 | 11.875 |  |  | 3.7 | 11.9266 | 12.8429 |  |  | 4.65 | 19.34 | 20.3667 |  |  | 7.34 | 75.175 | 75 |  |  | 10.7 | 76.95 | 76.76
17.2 |  |  | 3.7 | 10.6618 | 11.7 |  |  | 3.7 | 11.5581 | 12.5333 |  |  | 4.6 | 17.86 | 18.8 |  |  | 7.2667 | 74.32 | 74.3 |  |  | 10.5 | 76.62 | 76.45
17.3 |  |  |  | 10.5418 | 11.55 |  |  | 3.6333 | 11.3381 | 12.2667 |  |  | 4.6 | 16.8167 | 17.8 |  |  | 7.1 | 73.4 | 73.3 |  |  | 10.4333 | 76.4 | 76.3
17.4 |  |  |  | 10.3517 | 11.325 |  |  | 3.6 | 11.1482 | 11.98 |  |  | 4.5 | 16 | 17.04 |  |  | 7 | 72.15 | 72 |  |  | 10.2 | 76.15 | 75.9667
17.5 |  |  |  | 10.3017 | 11.2 |  |  | 3.6 | 10.9315 | 11.86 |  |  | 4.4833 | 15.4143 | 16.35 |  |  | 6.9167 | 70.45 | 70.4 |  |  | 10.1 | 75.7667 | 75.7167
17.6 |  |  |  | 10.1517 | 11.075 |  |  | 3.5 | 10.7125 | 11.5667 |  |  | 4.4 | 14.8714 | 15.8 |  |  | 6.8167 | 68.12 | 67.8 |  |  | 10.0167 | 75.55 | 75.52
17.7 |  |  |  | 10.0017 | 10.96 |  |  | 3.5 | 10.5149 | 11.4 |  |  | 4.375 | 14.4 | 15.38 |  |  | 6.7 | 64.6 | 64.44 |  |  | 9.8 | 75.24 | 75.2167
17.8 |  |  |  | 9.8766 | 10.8 |  |  | 3.4333 | 10.3383 | 11.3143 |  |  | 4.3 | 13.9857 | 14.95 |  |  | 6.66 | 59.76 | 59.78 |  |  | 9.6167 | 74.9 | 74.875
17.9 |  |  |  | 9.7266 | 10.7 |  |  |  | 10.1583 | 11.14 |  |  | 4.3 | 13.6667 | 14.6667 |  |  | 6.525 | 54.22 | 54.36 |  |  | 9.52 | 74.575 | 74.525
18 |  |  |  | 9.6016 | 10.56 |  |  |  | 10.0783 | 10.95 |  |  | 4.2 | 13.3333 | 14.34 |  |  | 6.5 | 47.8833 | 48.05 |  |  | 9.3571 | 74.26 | 74.15
18.1 |  |  |  | 9.5016 | 10.425 |  |  |  | 9.8984 | 10.8 |  |  | 4.2 | 13.08 | 14 |  |  | 6.38 | 42.1 | 42.05 |  |  | 9.275 | 73.72 | 73.88
18.2 |  |  |  | 9.4016 | 10.32 |  |  |  | 9.715 | 10.7 |  |  | 4.1 | 12.8 | 13.7 |  |  | 6.3 | 36.2 | 36.54 |  |  | 9.1 | 73.4 | 73.45
18.3 |  |  |  | 9.2615 | 10.2 |  |  |  | 9.6184 | 10.55 |  |  | 4.06 | 12.5833 | 13.5167 |  |  | 6.2 | 30.9 | 31.4 |  |  | 9 | 73 | 72.96
18.4 |  |  |  | 9.1265 | 10.1143 |  |  |  | 9.4984 | 10.4 |  |  | 4 | 12.3571 | 13.3 |  |  | 6.15 | 27.1 | 27.7 |  |  | 8.9 | 72.45 | 72.54
18.5 |  |  |  | 9.0015 | 10 |  |  |  | 9.3651 | 10.28 |  |  | 4 | 12.0857 | 13.05 |  |  | 6.08 | 24.28 | 25.04 |  |  | 8.8 | 71.94 | 71.96
18.6 |  |  |  | 8.9015 | 9.9 |  |  |  | 9.1985 | 10.18 |  |  | 3.9 | 11.92 | 12.9 |  |  | 6 | 22.3 | 23.175 |  |  | 8.6667 | 71.28 | 71.3
18.7 |  |  |  | 8.8015 | 9.7 |  |  |  | 9.0842 | 10.0143 |  |  |  | 11.7143 | 12.65 |  |  | 5.9 | 20.975 | 21.9 |  |  | 8.5 | 70.46 | 70.6
18.8 |  |  |  | 8.6764 | 9.6167 |  |  |  | 8.9985 | 9.9 |  |  |  | 11.4857 | 12.5 |  |  | 5.82 | 19.9 | 20.86 |  |  | 8.4333 | 69.45 | 69.66
18.9 |  |  |  | 8.5814 | 9.5 |  |  |  | 8.8652 | 9.74 |  |  |  | 11.35 | 12.3 |  |  | 5.8 | 19.1 | 20.06 |  |  | 8.3 | 68.3 | 68.4
19 |  |  |  | 8.4814 | 9.4 |  |  |  | 8.6985 | 9.66 |  |  |  | 11.2 | 12.08 |  |  | 5.7 | 18.45 | 19.4 |  |  | 8.2333 | 66.55 | 66.775
19.1 |  |  |  | 8.3264 | 9.24 |  |  |  | 8.6129 | 9.5 |  |  |  | 11.0167 | 11.9 |  |  | 5.6 | 17.85 | 18.84 |  |  | 8.2 | 64.3 | 64.7
19.2 |  |  |  | 8.3014 | 9.16 |  |  |  | 8.5152 | 9.4 |  |  |  | 10.84 | 11.78 |  |  | 5.6 | 17.45 | 18.4 |  |  | 8.05 | 61.1 | 61.56
19.3 |  |  |  | 8.1014 | 9 |  |  |  | 8.3986 | 9.3 |  |  |  | 10.7 | 11.6 |  |  | 5.5 | 17 | 17.9 |  |  | 7.9 | 57.18 | 57.8
19.4 |  |  |  | 8.0413 | 8.94 |  |  |  | 8.232 | 9.1833 |  |  |  | 10.5333 | 11.48 |  |  | 5.44 | 16.6 | 17.6333 |  |  | 7.8 | 52.16 | 52.9429
19.5 |  |  |  | 7.9013 | 8.8 |  |  |  | 8.1653 | 9.0167 |  |  |  | 10.3667 | 11.3 |  |  | 5.4 | 16.2 | 17.18 |  |  | 7.75 | 47 | 47.8
19.6 |  |  |  | 7.8013 | 8.725 |  |  |  | 7.9987 | 8.92 |  |  |  | 10.2 | 11.15 |  |  | 5.3 | 16 | 17 |  |  | 7.62 | 41.8 | 42.6
19.7 |  |  |  | 7.7013 | 8.6 |  |  |  | 7.9415 | 8.82 |  |  |  | 10.1 | 11 |  |  | 5.225 | 15.675 | 16.6 |  |  | 7.5 | 36.375 | 37.6
19.8 |  |  |  | 7.6013 | 8.5 |  |  |  | 7.8987 | 8.7 |  |  |  | 9.9571 | 10.9 |  |  | 5.2 | 15.4333 | 16.375 |  |  | 7.4333 | 32.22 | 33.3
19.9 |  |  |  | 7.5012 | 8.4 |  |  |  | 7.732 | 8.5833 |  |  |  | 9.8 | 10.675 |  |  | 5.1 | 15.12 | 16.12 |  |  | 7.3 | 28.6 | 29.75
20 |  |  |  | 7.4012 | 8.3 |  |  |  | 7.6321 | 8.5 |  |  |  | 9.68 | 10.6 |  |  | 5.1 | 14.88 | 15.84 |  |  | 7.26 | 26.1 | 27.3
20.1 |  |  |  | 7.3012 | 8.22 |  |  |  | 7.5654 | 8.4 |  |  |  | 9.5 | 10.42 |  |  | 5 | 14.5857 | 15.6333 |  |  | 7.2 | 24.3 | 25.52
20.2 |  |  |  | 7.1345 | 8.1 |  |  |  | 7.4821 | 8.3 |  |  |  | 9.4 | 10.28 |  |  | 5 | 14.42 | 15.4 |  |  | 7.1 | 22.94 | 24.05
20.3 |  |  |  | 7.1012 | 8 |  |  |  | 7.3154 | 8.2 |  |  |  | 9.3 | 10.14 |  |  | 4.9 | 14.16 | 15.1 |  |  | 7.05 | 21.95 | 23.1
20.4 |  |  |  | 7.0012 | 7.86 |  |  |  | 7.2559 | 8.1 |  |  |  | 9.12 | 10.025 |  |  |  | 14.05 | 14.92 |  |  | 7 | 21.2 | 22.3571
20.5 |  |  |  | 6.8611 | 7.76 |  |  |  | 7.1845 | 8 |  |  |  | 9 | 9.86 |  |  |  | 13.7 | 14.76 |  |  | 6.9 | 20.52 | 21.5833
20.6 |  |  |  | 6.7261 | 7.64 |  |  |  | 7.0417 | 7.9 |  |  |  | 8.8667 | 9.725 |  |  |  | 13.56 | 14.55 |  |  | 6.8 | 19.98 | 20.95
20.7 |  |  |  | 6.6761 | 7.54 |  |  |  | 6.9822 | 7.8 |  |  |  | 8.7 | 9.7 |  |  |  | 13.45 | 14.3 |  |  | 6.7 | 19.5 | 20.62
20.8 |  |  |  | 6.5611 | 7.42 |  |  |  | 6.8989 | 7.7 |  |  |  | 8.62 | 9.58 |  |  |  | 13.26 | 14.175 |  |  | 6.65 | 19.14 | 20.24
20.9 |  |  |  | 6.5011 | 7.3 |  |  |  | 6.7989 | 7.58 |  |  |  | 8.5 | 9.4 |  |  |  | 13 | 13.925 |  |  | 6.6 | 18.725 | 19.775
21 |  |  |  | 6.4011 | 7.22 |  |  |  | 6.6703 | 7.48 |  |  |  | 8.4125 | 9.3 |  |  |  | 12.8 | 13.7 |  |  | 6.5 | 18.3667 | 19.45
21.1 |  |  |  | 6.3011 | 7.1 |  |  |  | 6.5275 | 7.4 |  |  |  | 8.3 | 9.2 |  |  |  | 12.6 | 13.6167 |  |  | 6.48 | 18 | 19.1333
21.2 |  |  |  | 6.201 | 7 |  |  |  | 6.4489 | 7.3 |  |  |  | 8.16 | 9.1 |  |  |  | 12.425 | 13.44 |  |  | 6.35 | 17.775 | 18.7667
21.3 |  |  |  | 6.101 | 6.9 |  |  |  | 6.2989 | 7.2 |  |  |  | 8.0667 | 8.98 |  |  |  | 12.28 | 13.26 |  |  | 6.3 | 17.425 | 18.5667
21.4 |  |  |  | 6.001 | 6.8 |  |  |  | 6.2156 | 7 |  |  |  | 8 | 8.86 |  |  |  | 12.1 | 13.14 |  |  | 6.2 | 17.15 | 18.375
21.5 |  |  |  | 5.901 | 6.78 |  |  |  | 6.199 | 6.95 |  |  |  | 7.8333 | 8.775 |  |  |  | 11.96 | 12.94 |  |  | 6.125 | 16.9 | 18.16
21.6 |  |  |  | 5.801 | 6.7 |  |  |  | 6.0418 | 6.8667 |  |  |  | 7.7 | 8.6 |  |  |  | 11.76 | 12.72 |  |  | 6.1 | 16.66 | 17.7667
21.7 |  |  |  | 5.7009 | 6.6 |  |  |  | 5.999 | 6.8 |  |  |  | 7.6143 | 8.56 |  |  |  | 11.625 | 12.6333 |  |  | 6 | 16.425 | 17.6
21.8 |  |  |  | 5.6843 | 6.5 |  |  |  | 5.899 | 6.7333 |  |  |  | 7.5 | 8.425 |  |  |  | 11.5 | 12.45 |  |  | 5.9 | 16.3 | 17.3
21.9 |  |  |  | 5.6009 | 6.44 |  |  |  | 5.799 | 6.6 |  |  |  | 7.425 | 8.3 |  |  |  | 11.35 | 12.3 |  |  | 5.9 | 16 | 17.1
22 |  |  |  | 5.5009 | 6.34 |  |  |  | 5.699 | 6.5 |  |  |  | 7.275 | 8.2 |  |  |  | 11.2 | 12.1 |  |  | 5.9 | 15.7 | 16.86
22.1 |  |  |  | 5.4009 | 6.3 |  |  |  | 5.5991 | 6.4333 |  |  |  | 7.2 | 8.1 |  |  |  | 11.02 | 12.025 |  |  |  | 15.5 | 16.6667
22.2 |  |  |  | 5.3759 | 6.2 |  |  |  | 5.5366 | 6.375 |  |  |  | 7.1 | 7.975 |  |  |  | 10.9 | 11.875 |  |  |  | 15.4 | 16.46
22.3 |  |  |  | 5.3009 | 6.1 |  |  |  | 5.4824 | 6.3 |  |  |  | 6.9429 | 7.8 |  |  |  | 10.775 | 11.68 |  |  |  | 15.14 | 16.2667
22.4 | [8 empty] | 5.3991 | 6.2 |  |  |  | 6.9 | 7.7333 |  |  |  | 10.64 | 11.52 |  |  |  | 14.875 | 16.0667
22.5 | [8 empty] | 5.3324 | 6.2 |  |  |  | 6.7571 | 7.6 |  |  |  | 10.42 | 11.4 |  |  |  | 14.66 | 15.9
22.6 | [8 empty] | 5.2825 | 6.1 |  |  |  | 6.6143 | 7.5 |  |  |  | 10.3333 | 11.2 |  |  |  | 14.5 | 15.76
22.7 | [8 empty] | 5.1991 | 6.0167 |  |  |  | 6.5167 | 7.475 |  |  |  | 10.2 | 11.1 |  |  |  | 14.3 | 15.56
22.8 | [8 empty] | 5.1991 | 5.9714 |  |  |  | 6.42 | 7.36 |  |  |  | 10.0667 | 11 |  |  |  | 14.15 | 15.375
22.9 | [8 empty] | 5.0991 | 5.8167 |  |  |  | 6.4 | 7.2 |  |  |  | 9.9 | 10.775 |  |  |  | 14.025 | 15.1667

## Wedge 15 Longitudinal Profiles
Longitudinal Profiles
Wedge 15
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2.4 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 8.9
-34.9 |  |  |  |  | 8.9889
-34.8 |  |  |  |  | 9
-34.7 |  |  |  |  | 9.1
-34.6 |  |  |  |  | 9.12
-34.5 |  |  |  |  | 9.22
-34.4 |  |  |  |  | 9.3222
-34.3 |  |  |  |  | 9.4375
-34.2 |  |  |  |  | 9.5556
-34.1 |  |  |  |  | 9.66
-34 |  |  |  |  | 9.76
-33.9 |  |  |  |  | 9.8545
-33.8 |  |  |  |  | 9.95
-33.7 |  |  |  |  | 10.0625
-33.6 |  |  |  |  | 10.1
-33.5 |  |  |  |  | 10.2125
-33.4 |  |  |  |  | 10.33
-33.3 |  |  |  |  | 10.43
-33.2 |  |  |  |  | 10.5
-33.1 |  |  |  |  | 10.5375
-33 |  |  |  |  | 10.6714
-32.9 |  |  |  |  | 10.7
-32.8 |  |  |  |  | 10.8143
-32.7 |  |  |  |  | 10.9667
-32.6 |  |  |  |  | 11.1
-32.5 |  |  |  |  | 11.2
-32.4 |  |  |  |  | 11.33
-32.3 |  |  |  | 8 | 11.43
-32.2 |  |  |  | 8.0889 | 11.5333
-32.1 |  |  |  | 8.1125 | 11.6571
-32 |  |  |  | 8.2375 | 11.7875
-31.9 |  |  |  | 8.3 | 11.8429
-31.8 |  |  |  | 8.3875 | 11.9857
-31.7 |  |  |  | 8.4143 | 12.1
-31.6 |  |  |  | 8.5 | 12.2125
-31.5 |  |  |  | 8.6 | 12.3375
-31.4 |  |  |  | 8.6667 | 12.4833
-31.3 |  |  |  | 8.8 | 12.6222
-31.2 |  |  |  | 8.8833 | 12.7333
-31.1 |  |  |  | 9 | 12.8444
-31 |  |  |  | 9.0667 | 13.0111
-30.9 |  |  |  | 9.1222 | 13.16
-30.8 |  |  |  | 9.2375 | 13.275
-30.7 |  |  |  | 9.3625 | 13.46
-30.6 |  |  |  | 9.4875 | 13.66
-30.5 |  |  |  | 9.5143 | 13.7727
-30.4 |  |  |  | 9.6 | 13.87
-30.3 |  |  |  | 9.75 | 14.075
-30.2 |  |  |  | 9.8143 | 14.2333
-30.1 |  |  |  | 9.9667 | 14.35
-30 |  |  |  | 10.1 | 14.475
-29.9 |  |  |  | 10.15 | 14.6778
-29.8 |  |  |  | 10.275 | 14.9125
-29.7 |  |  |  | 10.4 | 15.0429
-29.6 |  |  | 7.0125 | 10.525 | 15.22
-29.5 |  |  | 7.16 | 10.65 | 15.4091
-29.4 |  |  | 7.2 | 10.775 | 15.6111
-29.3 |  |  | 7.2429 | 10.82 | 15.7667
-29.2 |  |  | 7.3 | 11 | 15.9556
-29.1 |  |  | 7.4 | 11.15 | 16.1
-29 |  |  | 7.4 | 11.3 | 16.425
-28.9 |  |  | 7.5 | 11.4 | 16.6143
-28.8 |  |  | 7.6 | 11.5667 | 16.8
-28.7 |  |  | 7.7 | 11.7 | 16.9571
-28.6 |  |  | 7.7667 | 11.7833 | 17.1556
-28.5 |  |  | 7.8286 | 11.9 | 17.44
-28.4 |  |  | 7.9 | 12.0167 | 17.6778
-28.3 |  | 6.6 | 8 | 12.1833 | 17.85
-28.2 |  | 6.6 | 8.05 | 12.3429 | 18.0889
-28.1 |  | 6.66 | 8.1857 | 12.5 | 18.2833
-28 |  | 6.7429 | 8.2571 | 12.6333 | 18.5667
-27.9 |  | 6.8 | 8.3143 | 12.7857 | 18.8
-27.8 |  | 6.85 | 8.4 | 12.9375 | 19.0333
-27.7 |  | 6.9 | 8.5 | 13.1667 | 19.275
-27.6 |  | 6.9571 | 8.6 | 13.3333 | 19.5571
-27.5 |  | 7 | 8.6667 | 13.5 | 19.95
-27.4 | 6.7571 | 7.0833 | 8.8143 | 13.6667 | 20.2333
-27.3 | 6.8 | 7.15 | 8.9 | 13.8429 | 20.6
-27.2 | 6.84 | 7.2 | 8.9875 | 13.9857 | 21.0714
-27.1 | 6.9167 | 7.3 | 9.1 | 14.2286 | 21.48
-27 | 7 | 7.3714 | 9.2 | 14.3833 | 21.95
-26.9 | 7.0286 | 7.4 | 9.3 | 14.64 | 22.5714
-26.8 | 7.1 | 7.4625 | 9.4167 | 14.875 | 23.4286
-26.7 | 7.2 | 7.6 | 9.5 | 15.0333 | 24.4167
-26.6 | 7.25 | 7.66 | 9.66 | 15.25 | 25.7143
-26.5 | 7.3 | 7.7 | 9.8 | 15.4167 | 27.4
-26.4 | 7.4 | 7.8 | 9.85 | 15.6667 | 29.5286
-26.3 | 7.4 | 7.8833 | 10 | 15.84 | 32.225
-26.2 | 7.5 | 7.9286 | 10.16 | 16.1167 | 35.3778
-26.1 | 7.5143 | 8 | 10.25 | 16.4 | 39.28
-26 | 7.6 | 8.1 | 10.4 | 16.55 | 43.74
-25.9 | 7.7 | 8.2143 | 10.4667 | 16.8286 | 48.46
-25.8 | 7.7857 | 8.3 | 10.6375 | 17.1143 | 53.275
-25.7 | 7.8375 | 8.4 | 10.7833 | 17.4 | 57.9
-25.6 | 7.9714 | 8.5 | 10.9286 | 17.6333 | 61.6667
-25.5 | 8.0143 | 8.58 | 11.0714 | 18 | 65.0333
-25.4 | 8.1 | 8.68 | 11.2 | 18.2857 | 68.01
-25.3 | 8.2 | 8.8 | 11.3429 | 18.7429 | 70.08
-25.2 | 8.3 | 8.8857 | 11.55 | 19.15 | 71.8222
-25.1 | 8.3833 | 8.9333 | 11.7333 | 19.6 | 73.0667
-25 | 8.4429 | 9.1 | 11.8857 | 20.2 | 74.125
-24.9 | 8.5167 | 9.2 | 12.0429 | 20.9 | 75
-24.8 | 8.6833 | 9.2857 | 12.2167 | 21.8 | 75.6
-24.7 | 8.7286 | 9.4 | 12.4 | 22.78 | 76.2571
-24.6 | 8.8714 | 9.5167 | 12.5833 | 24.2 | 76.8222
-24.5 | 8.9143 | 9.6 | 12.8375 | 26.05 | 77.3222
-24.4 | 9.0571 | 9.76 | 13.025 | 28.64 | 77.8
-24.3 | 9.1 | 9.925 | 13.225 | 31.82 | 78.0429
-24.2 | 9.25 | 10.0571 | 13.4429 | 35.7 | 78.45
-24.1 | 9.3 | 10.1143 | 13.6167 | 40.6 | 78.8
-24 | 9.4571 | 10.2667 | 13.8429 | 45.825 | 79.1286
-23.9 | 9.6 | 10.4333 | 14.0857 | 51.6 | 79.5
-23.8 | 9.6667 | 10.5857 | 14.2857 | 57.2 | 79.81
-23.7 | 9.8 | 10.7 | 14.5286 | 62.5 | 80.075
-23.6 | 9.9333 | 10.8571 | 14.725 | 67 | 80.3571
-23.5 | 10.02 | 11.0333 | 15 | 70.6571 | 80.6429
-23.4 | 10.2 | 11.2 | 15.4429 | 73.7857 | 81
-23.3 | 10.3375 | 11.35 | 15.8429 | 76.1 | 81.2286
-23.2 | 10.4714 | 11.5667 | 16.2333 | 77.6857 | 81.5
-23.1 | 10.6333 | 11.725 | 16.7714 | 78.9714 | 81.8
-23 | 10.7 | 11.9333 | 17.3429 | 79.9714 | 82.0429
-22.9 | 10.8667 | 12.1167 | 18.0571 | 80.7333 | 82.25
-22.8 | 11.0167 | 12.3429 | 19 | 81.5 | 82.52
-22.7 | 11.1833 | 12.6167 | 20.1 | 82.04 | 82.71
-22.6 | 11.3571 | 12.7833 | 21.6 | 82.5429 | 82.83
-22.5 | 11.5143 | 13.05 | 23.9143 | 83.05 | 83.12
-22.4 | 11.7143 | 13.3 | 26.9714 | 83.3667 | 83.3111
-22.3 | 11.9167 | 13.65 | 30.8571 | 83.8333 | 83.45
-22.2 | 12.0833 | 14 | 35.775 | 84.1429 | 83.7
-22.1 | 12.25 | 14.4857 | 41.8714 | 84.4286 | 83.95
-22 | 12.5 | 15.0571 | 48.9714 | 84.76 | 84.2125
-21.9 | 12.7429 | 15.7333 | 56.04 | 85.12 | 84.3857
-21.8 | 13.025 | 16.7 | 63.0333 | 85.425 | 84.575
-21.7 | 13.35 | 17.8857 | 69.2 | 85.675 | 84.7
-21.6 | 13.7667 | 19.5167 | 74.7833 | 85.9667 | 84.8571
-21.5 | 14.3 | 21.9143 | 79.0857 | 86.3125 | 85.125
-21.4 | 14.9429 | 25.1167 | 82.2 | 86.52 | 85.3
-21.3 | 15.8 | 29.8286 | 84.25 | 86.725 | 85.4375
-21.2 | 16.9 | 35.66 | 85.86 | 86.975 | 85.5556
-21.1 | 18.4333 | 42.3 | 87.2857 | 87.225 | 85.7714
-21 | 20.8143 | 50.05 | 88.1333 | 87.4 | 85.9
-20.9 | 23.9667 | 58.3375 | 88.9333 | 87.65 | 86.1375
-20.8 | 29.1 | 66.65 | 89.5 | 87.8 | 86.35
-20.7 | 34.9333 | 73.5 | 90 | 88.0667 | 86.5625
-20.6 | 42.4167 | 79.5833 | 90.6 | 88.2667 | 86.6778
-20.5 | 51 | 84.02 | 90.9125 | 88.4 | 86.8125
-20.4 | 59.2 | 87.3 | 91.35 | 88.65 | 86.9375
-20.3 | 68.35 | 89.5833 | 91.7333 | 88.8667 | 87.0714
-20.2 | 76.4 | 91.0833 | 91.975 | 89 | 87.1889
-20.1 | 82.2857 | 92.26 | 92.25 | 89.1875 | 87.3
-20 | 87.1333 | 93.1333 | 92.5143 | 89.3875 | 87.5
-19.9 | 90.4 | 93.9 | 92.76 | 89.5125 | 87.6818
-19.8 | 92.6 | 94.4 | 93.06 | 89.6375 | 87.79
-19.7 | 94.12 | 94.8857 | 93.225 | 89.825 | 87.89
-19.6 | 95.2 | 95.3143 | 93.4375 | 89.9875 | 88.025
-19.5 | 96.0333 | 95.5571 | 93.625 | 90.1 | 88.26
-19.4 | 96.6 | 95.95 | 93.9 | 90.225 | 88.43
-19.3 | 96.9714 | 96.175 | 94.0375 | 90.3667 | 88.5375
-19.2 | 97.375 | 96.3714 | 94.1833 | 90.5333 | 88.6556
-19.1 | 97.75 | 96.5571 | 94.36 | 90.7 | 88.7667
-19 | 98.0167 | 96.825 | 94.4143 | 90.8429 | 88.9273
-18.9 | 98.2667 | 97 | 94.5667 | 90.9167 | 89.06
-18.8 | 98.4167 | 97.1 | 94.74 | 91.0625 | 89.2333
-18.7 | 98.5714 | 97.2125 | 94.8333 | 91.1875 | 89.3778
-18.6 | 98.75 | 97.36 | 94.9857 | 91.35 | 89.56
-18.5 | 98.9167 | 97.5 | 95.08 | 91.4125 | 89.68
-18.4 | 99 | 97.5857 | 95.2 | 91.5 | 89.7
-18.3 | 99.1 | 97.7571 | 95.3143 | 91.575 | 89.8429
-18.2 | 99.2 | 97.9 | 95.4 | 91.8143 | 89.9857
-18.1 | 99.3167 | 98 | 95.5 | 91.95 | 90.1
-18 | 99.4 | 98.08 | 95.5857 | 92.0857 | 90.2222
-17.9 | 99.5 | 98.1 | 95.7 | 92.1429 | 90.4
-17.8 | 99.5833 | 98.2 | 95.8 | 92.2857 | 90.43
-17.7 | 99.68 | 98.3 | 95.8667 | 92.4 | 90.5333
-17.6 | 99.7 | 98.3 | 96 | 92.5 | 90.6444
-17.5 | 99.7 | 98.4 | 96.1 | 92.6 | 90.8111
-17.4 | 99.8 | 98.45 | 96.1286 | 92.625 | 90.975
-17.3 | 99.8 | 98.5 | 96.2 | 92.7571 | 91.08
-17.2 | 99.85 | 98.6 | 96.3 | 92.8875 | 91.2
-17.1 | 99.9167 | 98.6 | 96.4 | 93.1 | 91.325
-17 | 100 | 98.6 | 96.5 | 93.225 | 91.44
-16.9 | 100 | 98.7 | 96.5 | 93.3 | 91.55
-16.8 | 100.1 | 98.7667 | 96.6 | 93.425 | 91.675
-16.7 | 100.1 | 98.8 | 96.7 | 93.5 | 91.8143
-16.6 | 100.12 | 98.8 | 96.7 | 93.5 | 91.95
-16.5 | 100.2 | 98.9 | 96.8286 | 93.5875 | 92
-16.4 | 100.2 | 98.9 | 96.9 | 93.7333 | 92.11
-16.3 | 100.2 | 98.95 | 96.9333 | 93.8143 | 92.21
-16.2 | 100.2 | 99 | 97 | 93.9571 | 92.31
-16.1 | 100.2 | 99 | 97.1 | 94.02 | 92.41
-16 | 100.2 | 99.06 | 97.1 | 94.1 | 92.5125
-15.9 | 100.28 | 99.1 | 97.2 | 94.2375 | 92.6
-15.8 | 100.3 | 99.1 | 97.22 | 94.3 | 92.675
-15.7 | 100.3 | 99.2 | 97.3 | 94.4 | 92.8
-15.6 | 100.3 | 99.2 | 97.4 | 94.475 | 93
-15.5 | 100.4 | 99.2 | 97.5 | 94.6 | 93.1
-15.4 | 100.4 | 99.2875 | 97.6 | 94.7 | 93.175
-15.3 | 100.4 | 99.3 | 97.6 | 94.8 | 93.3
-15.2 | 100.4 | 99.4 | 97.6875 | 94.8571 | 93.3778
-15.1 | 100.4 | 99.4 | 97.7429 | 95 | 93.4889
-15 | 100.5 | 99.4167 | 97.8 | 95.1 | 93.51
-14.9 | 100.5 | 99.5 | 97.8429 | 95.1667 | 93.61
-14.8 | 100.5 | 99.5 | 97.9 | 95.2 | 93.72
-14.7 | 100.5 | 99.6 | 97.98 | 95.3 | 93.9
-14.6 | 100.6 | 99.6 | 98.0375 | 95.4429 | 93.9444
-14.5 | 100.6 | 99.6 | 98.1 | 95.5 | 94.05
-14.4 | 100.6 | 99.62 | 98.1 | 95.55 | 94.15
-14.3 | 100.6 | 99.7 | 98.2 | 95.6 | 94.2625
-14.2 | 100.6 | 99.7 | 98.2 | 95.7667 | 94.3
-14.1 | 100.7 | 99.8 | 98.3 | 95.86 | 94.4222
-14 | 100.7 | 99.8 | 98.4 | 95.925 | 94.5
-13.9 | 100.7 | 99.8 | 98.4 | 96 | 94.5714
-13.8 | 100.8 | 99.9 | 98.5 | 96.1 | 94.6111
-13.7 | 100.8 | 99.9 | 98.5 | 96.125 | 94.725
-13.6 | 100.82 | 99.9 | 98.6 | 96.25 | 94.8444
-13.5 | 100.9 | 99.9833 | 98.6333 | 96.32 | 94.95
-13.4 | 100.9 | 100 | 98.7 | 96.4 | 95.05
-13.3 | 101 | 100.1 | 98.8 | 96.5 | 95.15
-13.2 | 101 | 100.1 | 98.8167 | 96.5857 | 95.2455
-13.1 | 101 | 100.1429 | 98.9 | 96.6 | 95.34
-13 | 101 | 100.2 | 98.9 | 96.6833 | 95.44
-12.9 | 101.0857 | 100.2714 | 99 | 96.7333 | 95.54
-12.8 | 101.1 | 100.3 | 99 | 96.84 | 95.6
-12.7 | 101.1857 | 100.32 | 99.1 | 96.9 | 95.6875
-12.6 | 101.2 | 100.4 | 99.1667 | 97 | 95.8
-12.5 | 101.275 | 100.4 | 99.2 | 97.075 | 95.82
-12.4 | 101.3 | 100.4 | 99.2286 | 97.1 | 95.9286
-12.3 | 101.3 | 100.4 | 99.3 | 97.1222 | 96
-12.2 | 101.3 | 100.4625 | 99.3 | 97.2375 | 96.0857
-12.1 | 101.4 | 100.5 | 99.4 | 97.3 | 96.19
-12 | 101.4 | 100.58 | 99.4 | 97.4 | 96.2
-11.9 | 101.4 | 100.6 | 99.4 | 97.4 | 96.2286
-11.8 | 101.4 | 100.6 | 99.5 | 97.5 | 96.3
-11.7 | 101.5 | 100.6857 | 99.5 | 97.5167 | 96.3778
-11.6 | 101.5 | 100.7 | 99.5333 | 97.6625 | 96.5
-11.5 | 101.5 | 100.7 | 99.6 | 97.7 | 96.5125
-11.4 | 101.5333 | 100.8 | 99.6 | 97.8 | 96.6333
-11.3 | 101.6 | 100.8 | 99.7 | 97.8333 | 96.7
-11.2 | 101.6 | 100.8 | 99.8 | 97.9 | 96.7556
-11.1 | 101.6 | 100.8 | 99.8 | 98 | 96.8545
-11 | 101.6 | 100.8429 | 99.8 | 98 | 96.95
-10.9 | 101.6 | 100.9 | 99.9 | 98.1 | 97.05
-10.8 | 101.64 | 100.9 | 99.9 | 98.1625 | 97.15
-10.7 | 101.7 | 100.9 | 99.9 | 98.2 | 97.2
-10.6 | 101.7 | 100.9 | 100 | 98.2 | 97.3
-10.5 | 101.7 | 101 | 100 | 98.3 | 97.3222
-10.4 | 101.8 | 101 | 100.0667 | 98.35 | 97.4
-10.3 | 101.8 | 101.0167 | 100.1 | 98.4 | 97.42
-10.2 | 101.8 | 101.1 | 100.1 | 98.4 | 97.52
-10.1 | 101.82 | 101.1 | 100.1 | 98.4143 | 97.6
-10 | 101.8 | 101.1 | 100.2 | 98.5 | 97.64
-9.9 | 101.8 | 101.1 | 100.2 | 98.6 | 97.7
-9.8 | 101.8333 | 101.1 | 100.2 | 98.625 | 97.7556
-9.7 | 101.9 | 101.1 | 100.2167 | 98.7 | 97.8
-9.6 | 101.9 | 101.1 | 100.3 | 98.7167 | 97.9
-9.5 | 101.9 | 101.2 | 100.3 | 98.8 | 97.9
-9.4 | 101.9 | 101.2 | 100.36 | 98.8875 | 98.01
-9.3 | 101.9 | 101.2 | 100.4 | 98.9125 | 98.1
-9.2 | 101.9714 | 101.2 | 100.4167 | 99 | 98.1
-9.1 | 102 | 101.2857 | 100.5 | 99 | 98.13
-9 | 102 | 101.3 | 100.5 | 99.1 | 98.2
-8.9 | 102 | 101.3 | 100.5 | 99.1 | 98.2444
-8.8 | 102 | 101.3 | 100.6 | 99.2 | 98.3
-8.7 | 102 | 101.34 | 100.5 | 99.2 | 98.3875
-8.6 | 102 | 101.3 | 100.5 | 99.2571 | 98.4
-8.5 | 102 | 101.4 | 100.6 | 99.3 | 98.5
-8.4 | 102 | 101.4 | 100.6 | 99.3 | 98.6
-8.3 | 102 | 101.4 | 100.6167 | 99.4 | 98.6111
-8.2 | 101.9571 | 101.4 | 100.7 | 99.4 | 98.7
-8.1 | 101.9167 | 101.4571 | 100.8 | 99.4 | 98.7714
-8 | 102 | 101.5 | 100.725 | 99.5 | 98.8
-7.9 | 102 | 101.5 | 100.7286 | 99.5 | 98.8111
-7.8 | 102 | 101.5 | 100.8 | 99.5375 | 98.9
-7.7 | 102 | 101.5 | 100.8 | 99.6 | 98.9429
-7.6 | 102 | 101.5 | 100.8 | 99.6889 | 99
-7.5 | 102 | 101.5 | 100.8 | 99.7429 | 99
-7.4 | 102 | 101.5 | 100.8 | 99.8 | 99.1
-7.3 | 102 | 101.5 | 100.9 | 99.8 | 99.1
-7.2 | 102 | 101.5 | 100.9 | 99.8 | 99.15
-7.1 | 102 | 101.5 | 100.9 | 99.8 | 99.2
-7 | 102 | 101.5 | 100.9 | 99.8 | 99.2
-6.9 | 102 | 101.5 | 100.9 | 99.9 | 99.3
-6.8 | 102 | 101.5 | 101 | 99.9 | 99.3
-6.7 | 102 | 101.5 | 100.9333 | 99.98 | 99.34
-6.6 | 102 | 101.5 | 101 | 100 | 99.4
-6.5 | 102 | 101.5 | 101 | 100 | 99.4
-6.4 | 102 | 101.5 | 101 | 100 | 99.4778
-6.3 | 102 | 101.5333 | 101 | 100 | 99.5
-6.2 | 102 | 101.6 | 101 | 100 | 99.5
-6.1 | 102 | 101.6 | 101 | 100 | 99.5857
-6 | 102 | 101.6 | 101 | 100.0375 | 99.6
-5.9 | 101.9 | 101.6 | 101 | 100.1 | 99.6111
-5.8 | 102 | 101.6 | 101 | 100.1 | 99.7
-5.7 | 101.9167 | 101.5 | 101 | 100.1 | 99.7
-5.6 | 101.9 | 101.5 | 101 | 100.1 | 99.7
-5.5 | 101.9 | 101.5 | 101 | 100.2 | 99.7556
-5.4 | 101.9 | 101.5 | 101 | 100.2 | 99.8
-5.3 | 101.8857 | 101.5 | 101 | 100.2 | 99.8
-5.2 | 101.8 | 101.5 | 101 | 100.2 | 99.8
-5.1 | 101.8 | 101.4 | 101 | 100.2 | 99.8
-5 | 101.7 | 101.4 | 101 | 100.2 | 99.8
-4.9 | 101.7 | 101.3714 | 101 | 100.3 | 99.8
-4.8 | 101.68 | 101.3 | 100.9 | 100.3 | 99.8
-4.7 | 101.6 | 101.3 | 100.9 | 100.3 | 99.8556
-4.6 | 101.6 | 101.3 | 100.9 | 100.2 | 99.9
-4.5 | 101.6 | 101.3 | 100.9 | 100.2 | 99.9
-4.4 | 101.55 | 101.3 | 100.8429 | 100.2 | 99.9
-4.3 | 101.5 | 101.2 | 100.8 | 100.2 | 99.9
-4.2 | 101.46 | 101.2 | 100.8 | 100.2 | 99.9
-4.1 | 101.4 | 101.2 | 100.8 | 100.2 | 99.9
-4 | 101.4 | 101.1 | 100.8 | 100.2 | 99.9556
-3.9 | 101.3 | 101.1 | 100.8 | 100.2 | 100
-3.8 | 101.3 | 101.0857 | 100.7 | 100.2 | 100
-3.7 | 101.3 | 101 | 100.7 | 100.2 | 100
-3.6 | 101.2667 | 100.925 | 100.7 | 100.2 | 100
-3.5 | 101.2 | 100.9 | 100.7 | 100.2 | 100
-3.4 | 101.1875 | 100.9 | 100.7 | 100.2 | 100
-3.3 | 101.1 | 100.8333 | 100.6 | 100.2 | 100
-3.2 | 101.0857 | 100.8 | 100.6 | 100.2 | 100
-3.1 | 100.9333 | 100.8 | 100.6 | 100.2 | 100
-3 | 100.9 | 100.8 | 100.6 | 100.2 | 100.0375
-2.9 | 100.9 | 100.7 | 100.6 | 100.1375 | 100.1
-2.8 | 100.86 | 100.7 | 100.5 | 100.2 | 100.1
-2.7 | 100.8 | 100.7 | 100.5 | 100.2 | 100.1
-2.6 | 100.8 | 100.7 | 100.5 | 100.2 | 100.1
-2.5 | 100.7 | 100.6 | 100.5 | 100.1 | 100.0875
-2.4 | 100.7 | 100.6 | 100.5 | 100.1 | 100
-2.3 | 100.6571 | 100.6 | 100.4 | 100.1 | 100
-2.2 | 100.6 | 100.5667 | 100.4 | 100.1 | 100
-2.1 | 100.6 | 100.5 | 100.3 | 100.1 | 100
-2 | 100.6 | 100.5 | 100.3 | 100.1 | 100
-1.9 | 100.5 | 100.4 | 100.3 | 100.1 | 100.0444
-1.8 | 100.5 | 100.4 | 100.3 | 100.1 | 100.1
-1.7 | 100.44 | 100.3 | 100.3 | 100.1 | 100.1
-1.6 | 100.4 | 100.3 | 100.2 | 100.1 | 100.1
-1.5 | 100.35 | 100.2 | 100.2 | 100.1 | 100.025
-1.4 | 100.2 | 100.2 | 100.2 | 100.1 | 100
-1.3 | 100.2 | 100.125 | 100.2 | 100.1 | 100
-1.2 | 100.2 | 100.1 | 100.1 | 100.1 | 100
-1.1 | 100.2 | 100.1 | 100.1 | 100.1 | 100
-1 | 100.12 | 100.1 | 100.1 | 100.0286 | 100
-0.9 | 100.1 | 100.1 | 100 | 100 | 100
-0.8 | 100.1 | 100.1 | 100 | 100 | 100
-0.7 | 100.1 | 100 | 100 | 100 | 100
-0.6 | 100 | 100 | 100 | 100 | 100
-0.5 | 100 | 100 | 100 | 100 | 100
-0.4 | 100 | 100 | 100 | 100 | 100
-0.3 | 100 | 100 | 99.9 | 100 | 100
-0.2 | 100 | 99.9 | 99.975 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 99.9571 | 100 | 100 | 100
0.2 | 100 | 99.9 | 99.975 | 100 | 100
0.3 | 100 | 100 | 99.9 | 100 | 100
0.4 | 100 | 100 | 100 | 100 | 100
0.5 | 100 | 100 | 100 | 100 | 100
0.6 | 100 | 100 | 100 | 100 | 100
0.7 | 100.1 | 100 | 100 | 100 | 100
0.8 | 100.1 | 100.0714 | 100 | 100 | 100
0.9 | 100.1 | 100.1 | 100 | 100 | 100
1 | 100.1 | 100.1 | 100.1 | 100.05 | 100
1.1 | 100.2 | 100.1 | 100.1 | 100.1 | 100
1.2 | 100.2 | 100.1 | 100.1 | 100.1 | 100
1.3 | 100.2 | 100.175 | 100.2 | 100.1 | 100
1.4 | 100.25 | 100.2 | 100.2 | 100.1 | 100
1.5 | 100.32 | 100.2571 | 100.2 | 100.1 | 100
1.6 | 100.4 | 100.3 | 100.2 | 100.1 | 100.1
1.7 | 100.4167 | 100.3 | 100.3 | 100.1 | 100.1
1.8 | 100.5 | 100.4 | 100.3 | 100.1 | 100.1
1.9 | 100.55 | 100.42 | 100.3 | 100.1 | 100
2 | 100.6 | 100.5 | 100.3 | 100.1 | 100
2.1 | 100.6 | 100.5 | 100.3 | 100.1 | 100
2.2 | 100.6 | 100.5167 | 100.4 | 100.1 | 100
2.3 | 100.6333 | 100.6 | 100.4 | 100.1 | 100
2.4 | 100.7 | 100.6 | 100.5 | 100.1 | 100.0111
2.5 | 100.7 | 100.6 | 100.5 | 100.1 | 100.1
2.6 | 100.7667 | 100.7 | 100.5 | 100.2 | 100.1
2.7 | 100.8 | 100.7 | 100.5 | 100.2 | 100.1
2.8 | 100.8333 | 100.7 | 100.5 | 100.175 | 100.1
2.9 | 100.9 | 100.7 | 100.6 | 100.1571 | 100.1
3 | 100.9 | 100.8 | 100.6 | 100.2 | 100.1
3.1 | 101 | 100.8 | 100.6 | 100.2 | 100
3.2 | 101.0667 | 100.8 | 100.6 | 100.2 | 100
3.3 | 101.1286 | 100.8833 | 100.6 | 100.2 | 100
3.4 | 101.2 | 100.9 | 100.7 | 100.2 | 100
3.5 | 101.2 | 100.9 | 100.7 | 100.2 | 100
3.6 | 101.24 | 100.975 | 100.7 | 100.2 | 100
3.7 | 101.3 | 101 | 100.7 | 100.2 | 100
3.8 | 101.3 | 101.04 | 100.7 | 100.2 | 100
3.9 | 101.3 | 101.1 | 100.8 | 100.2 | 100
4 | 101.4 | 101.1143 | 100.8 | 100.2 | 99.92
4.1 | 101.4 | 101.2 | 100.8 | 100.2 | 99.9
4.2 | 101.44 | 101.2 | 100.8 | 100.2 | 99.9
4.3 | 101.5 | 101.2 | 100.8 | 100.2 | 99.9
4.4 | 101.5333 | 101.2833 | 100.8429 | 100.2 | 99.9
4.5 | 101.6 | 101.3 | 100.9 | 100.2 | 99.9
4.6 | 101.6 | 101.3 | 100.9 | 100.2 | 99.9
4.7 | 101.6 | 101.3 | 100.9 | 100.2375 | 99.9
4.8 | 101.7 | 101.3 | 100.9 | 100.3 | 99.8
4.9 | 101.7 | 101.3167 | 101 | 100.2833 | 99.8
5 | 101.7 | 101.4 | 101 | 100.2 | 99.8
5.1 | 101.8 | 101.4571 | 101 | 100.2 | 99.8
5.2 | 101.8 | 101.5 | 101 | 100.2 | 99.8
5.3 | 101.8667 | 101.5 | 101 | 100.2 | 99.8
5.4 | 101.9 | 101.5 | 101 | 100.2 | 99.8
5.5 | 101.9 | 101.5 | 101 | 100.2 | 99.72
5.6 | 101.9 | 101.5 | 101 | 100.12 | 99.7
5.7 | 101.9 | 101.5429 | 101 | 100.1 | 99.7
5.8 | 102 | 101.6 | 101 | 100.1 | 99.6875
5.9 | 101.9 | 101.6 | 101 | 100.1 | 99.6
6 | 101.9857 | 101.6 | 101 | 100.0125 | 99.6
6.1 | 102 | 101.6 | 101 | 100 | 99.5429
6.2 | 102 | 101.6 | 101 | 100 | 99.5
6.3 | 102 | 101.5833 | 101 | 100 | 99.5
6.4 | 102 | 101.5 | 101 | 100 | 99.4444
6.5 | 102 | 101.5 | 101 | 100 | 99.4
6.6 | 102 | 101.5 | 101 | 100 | 99.4
6.7 | 102 | 101.5 | 100.9333 | 99.95 | 99.3
6.8 | 102 | 101.5 | 101 | 99.9 | 99.3
6.9 | 102 | 101.5 | 100.9 | 99.9 | 99.3
7 | 102 | 101.5 | 100.9 | 99.8714 | 99.26
7.1 | 102 | 101.5 | 100.9 | 99.8 | 99.2
7.2 | 102 | 101.5 | 100.9 | 99.8 | 99.1143
7.3 | 102 | 101.5 | 100.9 | 99.8 | 99.1
7.4 | 102 | 101.5 | 100.8 | 99.8 | 99.0875
7.5 | 102 | 101.5 | 100.8 | 99.7167 | 99
7.6 | 102 | 101.5 | 100.8 | 99.675 | 99
7.7 | 102 | 101.5 | 100.8 | 99.6 | 98.9
7.8 | 102 | 101.5 | 100.8 | 99.5125 | 98.8875
7.9 | 102 | 101.5 | 100.7286 | 99.5 | 98.8
8 | 102 | 101.5 | 100.725 | 99.4714 | 98.8
8.1 | 101.94 | 101.5 | 100.8 | 99.4 | 98.7286
8.2 | 101.9333 | 101.4 | 100.7 | 99.4 | 98.6875
8.3 | 102 | 101.4 | 100.6167 | 99.4 | 98.6
8.4 | 102 | 101.4 | 100.6 | 99.3 | 98.57
8.5 | 102 | 101.4 | 100.6 | 99.3 | 98.4667
8.6 | 102 | 101.3286 | 100.5 | 99.3 | 98.4
8.7 | 102 | 101.3 | 100.5 | 99.2 | 98.3444
8.8 | 102 | 101.3 | 100.6 | 99.2 | 98.3
8.9 | 102 | 101.3 | 100.5 | 99.1 | 98.2
9 | 102 | 101.3 | 100.5 | 99.1 | 98.2
9.1 | 102 | 101.3 | 100.5 | 99 | 98.19
9.2 | 102 | 101.2 | 100.4167 | 99 | 98.1
9.3 | 101.9 | 101.2 | 100.4 | 98.9 | 98.07
9.4 | 101.9 | 101.2 | 100.36 | 98.8625 | 97.9667
9.5 | 101.9 | 101.2 | 100.3 | 98.8 | 97.9
9.6 | 101.9 | 101.1 | 100.3 | 98.7 | 97.8875
9.7 | 101.9 | 101.1 | 100.2167 | 98.7 | 97.8
9.8 | 101.8 | 101.1 | 100.2 | 98.6 | 97.7143
9.9 | 101.8 | 101.1 | 100.2 | 98.5714 | 97.7
10 | 101.8 | 101.1 | 100.2 | 98.5 | 97.7
10.1 | 101.8 | 101.1 | 100.1 | 98.5 | 97.5889
10.2 | 101.8 | 101.1 | 100.1 | 98.4 | 97.5
10.3 | 101.8 | 101 | 100.1 | 98.4 | 97.4818
10.4 | 101.8 | 101 | 100.0667 | 98.3286 | 97.39
10.5 | 101.7333 | 101 | 100 | 98.2714 | 97.3
10.6 | 101.7 | 100.925 | 100 | 98.2 | 97.2714
10.7 | 101.7 | 100.9 | 99.9 | 98.2 | 97.2
10.8 | 101.6 | 100.9 | 99.9 | 98.1375 | 97.1111
10.9 | 101.6 | 100.9 | 99.9 | 98.1 | 97.01
11 | 101.6 | 100.8 | 99.8 | 98 | 96.91
11.1 | 101.6 | 100.8 | 99.8 | 97.98 | 96.8182
11.2 | 101.6 | 100.8 | 99.8 | 97.9 | 96.72
11.3 | 101.6 | 100.8 | 99.7 | 97.8143 | 96.7
11.4 | 101.5571 | 100.76 | 99.6 | 97.8 | 96.6
11.5 | 101.5 | 100.7 | 99.6 | 97.7 | 96.5778
11.6 | 101.5 | 100.7 | 99.5333 | 97.6375 | 96.4625
11.7 | 101.5 | 100.7 | 99.5 | 97.5 | 96.3375
11.8 | 101.4 | 100.6 | 99.5 | 97.5 | 96.3
11.9 | 101.4 | 100.6 | 99.4 | 97.4 | 96.2875
12 | 101.4 | 100.52 | 99.4 | 97.4 | 96.2
12.1 | 101.4 | 100.5 | 99.4 | 97.3571 | 96.1545
12.2 | 101.3 | 100.5 | 99.3 | 97.2143 | 96.0429
12.3 | 101.3 | 100.4 | 99.3 | 97.1889 | 96
12.4 | 101.3 | 100.4 | 99.2286 | 97.1 | 95.9
12.5 | 101.3 | 100.4 | 99.2 | 97.05 | 95.8778
12.6 | 101.2 | 100.4 | 99.1667 | 97 | 95.7667
12.7 | 101.1143 | 100.3 | 99.1 | 96.9 | 95.6429
12.8 | 101.1 | 100.3 | 99 | 96.8167 | 95.6
12.9 | 101.1 | 100.2167 | 99 | 96.7 | 95.5
13 | 101 | 100.2 | 98.9 | 96.6571 | 95.4
13.1 | 101 | 100.2 | 98.9 | 96.6 | 95.4
13.2 | 101 | 100.1 | 98.8167 | 96.5571 | 95.3091
13.3 | 101 | 100.1 | 98.8 | 96.5 | 95.21
13.4 | 100.9 | 100 | 98.7 | 96.4 | 95.11
13.5 | 100.9 | 100 | 98.6333 | 96.3833 | 95.01
13.6 | 100.84 | 99.9 | 98.6 | 96.2286 | 94.8
13.7 | 100.8 | 99.9 | 98.5 | 96.1 | 94.7
13.8 | 100.8 | 99.88 | 98.5 | 96.1 | 94.6714
13.9 | 100.7 | 99.8 | 98.4 | 96.025 | 94.5286
14 | 100.7 | 99.8 | 98.4 | 95.9 | 94.5
14.1 | 100.7 | 99.8 | 98.3 | 95.8333 | 94.3778
14.2 | 100.6167 | 99.7167 | 98.2 | 95.74 | 94.3
14.3 | 100.6 | 99.7 | 98.2 | 95.6 | 94.225
14.4 | 100.6 | 99.6 | 98.1 | 95.6 | 94.1111
14.5 | 100.6 | 99.6 | 98.1 | 95.5 | 94.01
14.6 | 100.6 | 99.6 | 98.0375 | 95.4167 | 94
14.7 | 100.5333 | 99.6 | 97.98 | 95.3 | 93.8875
14.8 | 100.5 | 99.5 | 97.9 | 95.22 | 93.77
14.9 | 100.5 | 99.5 | 97.8429 | 95.1333 | 93.67
15 | 100.5 | 99.4667 | 97.8 | 95.0667 | 93.5667
15.1 | 100.4286 | 99.4 | 97.7429 | 95 | 93.5
15.2 | 100.4 | 99.4 | 97.6875 | 94.9286 | 93.4333
15.3 | 100.4 | 99.3 | 97.6 | 94.8 | 93.3222
15.4 | 100.4 | 99.2429 | 97.6 | 94.7 | 93.2
15.5 | 100.4 | 99.2 | 97.5 | 94.6 | 93.075
15.6 | 100.3167 | 99.2 | 97.4 | 94.5 | 92.96
15.7 | 100.3 | 99.2 | 97.3 | 94.4 | 92.8111
15.8 | 100.3 | 99.1125 | 97.22 | 94.3 | 92.7
15.9 | 100.2 | 99.1 | 97.2 | 94.2143 | 92.6
16 | 100.2 | 99.1 | 97.1 | 94.1714 | 92.5
16.1 | 100.2 | 99 | 97.1 | 94.1 | 92.47
16.2 | 100.2 | 99 | 97 | 93.9333 | 92.34
16.3 | 100.2 | 98.9 | 96.9333 | 93.8833 | 92.17
16.4 | 100.2 | 98.9 | 96.9 | 93.7 | 92.1
16.5 | 100.2 | 98.9 | 96.8286 | 93.6 | 92.0375
16.6 | 100.15 | 98.8167 | 96.7 | 93.5167 | 91.9
16.7 | 100.1 | 98.8 | 96.7 | 93.5 | 91.8714
16.8 | 100.1 | 98.7167 | 96.6 | 93.4 | 91.7375
16.9 | 100 | 98.7 | 96.5 | 93.3 | 91.6
17 | 100 | 98.6 | 96.5 | 93.2 | 91.4
17.1 | 99.9 | 98.6 | 96.4 | 93.075 | 91.2875
17.2 | 99.88 | 98.54 | 96.3 | 92.9111 | 91.2
17.3 | 99.8 | 98.5 | 96.2 | 92.7286 | 91.08
17.4 | 99.8 | 98.4 | 96.1286 | 92.6 | 90.9333
17.5 | 99.7143 | 98.4 | 96.1 | 92.5667 | 90.8125
17.6 | 99.7 | 98.34 | 96 | 92.5 | 90.62
17.7 | 99.7 | 98.25 | 95.8667 | 92.4 | 90.5909
17.8 | 99.6 | 98.2 | 95.8 | 92.2667 | 90.5
17.9 | 99.5 | 98.1 | 95.7 | 92.2 | 90.3875
18 | 99.4143 | 98.02 | 95.5857 | 92.0571 | 90.24
18.1 | 99.2333 | 98 | 95.5 | 91.925 | 90.0625
18.2 | 99.2 | 97.9 | 95.4 | 91.8 | 90
18.3 | 99.1143 | 97.8 | 95.3143 | 91.65 | 89.8
18.4 | 99.02 | 97.64 | 95.2 | 91.5143 | 89.7
18.5 | 98.94 | 97.4833 | 95.08 | 91.4857 | 89.6444
18.6 | 98.7667 | 97.4 | 94.9857 | 91.32 | 89.54
18.7 | 98.5 | 97.2143 | 94.8333 | 91.225 | 89.38
18.8 | 98.4429 | 97.1 | 94.74 | 91.0375 | 89.225
18.9 | 98.3125 | 96.9571 | 94.5667 | 90.9 | 89.12
19 | 98 | 96.8143 | 94.4143 | 90.8286 | 88.9545
19.1 | 97.725 | 96.5667 | 94.36 | 90.6714 | 88.83
19.2 | 97.3875 | 96.4 | 94.1833 | 90.5 | 88.7125
19.3 | 97.0333 | 96.1714 | 94.0375 | 90.4429 | 88.5
19.4 | 96.6667 | 95.9333 | 93.9 | 90.3 | 88.39
19.5 | 96 | 95.5667 | 93.625 | 90.075 | 88.2778
19.6 | 95.2571 | 95.3286 | 93.4375 | 89.9556 | 88.075
19.7 | 94.1333 | 94.8857 | 93.225 | 89.8375 | 87.9111
19.8 | 92.58 | 94.4333 | 93.06 | 89.7 | 87.75
19.9 | 90.2 | 93.9 | 92.76 | 89.4875 | 87.6545
20 | 87.16 | 93.1333 | 92.5143 | 89.3625 | 87.5111
20.1 | 81.9286 | 92.3 | 92.25 | 89.175 | 87.3545
20.2 | 75.9167 | 91.0667 | 91.975 | 89 | 87.25
20.3 | 67.8333 | 89.56 | 91.7333 | 88.85 | 87.125
20.4 | 59.7 | 87.4 | 91.35 | 88.6 | 87
20.5 | 50.84 | 84.1333 | 90.9125 | 88.4714 | 86.7625
20.6 | 42.75 | 79.8 | 90.6 | 88.3 | 86.6375
20.7 | 35.2 | 73.96 | 90 | 88.0857 | 86.5125
20.8 | 29.2667 | 66.9 | 89.5 | 87.8833 | 86.3
20.9 | 24.2 | 58.2429 | 88.9333 | 87.6571 | 86.1
21 | 20.8 | 50.3 | 88.1333 | 87.4429 | 85.925
21.1 | 18.4571 | 42.4 | 87.2857 | 87.175 | 85.7375
21.2 | 16.86 | 35.4 | 85.86 | 86.925 | 85.5444
21.3 | 15.7 | 29.7625 | 84.25 | 86.7714 | 85.4
21.4 | 14.9143 | 25.38 | 82.2 | 86.5333 | 85.2571
21.5 | 14.3 | 21.8333 | 79.0857 | 86.275 | 85.1143
21.6 | 13.75 | 19.5 | 74.7833 | 85.95 | 84.8875
21.7 | 13.4 | 17.8429 | 69.2 | 85.7125 | 84.7625
21.8 | 13 | 16.6 | 63.0333 | 85.4 | 84.575
21.9 | 12.7429 | 15.78 | 56.04 | 85.1667 | 84.4
22 | 12.5286 | 15.0667 | 48.9714 | 84.8 | 84.1429
22.1 | 12.28 | 14.4571 | 41.8714 | 84.4714 | 83.9375
22.2 | 12.1 | 14.0333 | 35.775 | 84.15 | 83.725
22.3 | 11.9 | 13.6625 | 30.8571 | 83.85 | 83.4875
22.4 | 11.6857 | 13.3 | 26.9714 | 83.35 | 83.31
22.5 | 11.5333 | 13 | 23.9143 | 83.04 | 83.04
22.6 | 11.3571 | 12.8333 | 21.6 | 82.5667 | 82.84
22.7 | 11.2143 | 12.5667 | 20.1 | 82.1 | 82.67
22.8 | 11.0429 | 12.35 | 19 | 81.5143 | 82.525
22.9 | 10.8857 | 12.16 | 18.0571 | 80.76 | 82.2571
23 | 10.7333 | 11.9143 | 17.3429 | 80 | 82
23.1 | 10.6 | 11.7714 | 16.7714 | 78.95 | 81.8143
23.2 | 10.4875 | 11.5333 | 16.2333 | 77.7 | 81.5286
23.3 | 10.3625 | 11.375 | 15.8429 | 76.0429 | 81.25
23.4 | 10.1571 | 11.2429 | 15.4429 | 73.6143 | 80.9667
23.5 | 10.04 | 11.06 | 15 | 70.7333 | 80.675
23.6 | 9.9 | 10.9 | 14.725 | 66.8286 | 80.3714
23.7 | 9.7667 | 10.725 | 14.5286 | 62.525 | 80.0857
23.8 | 9.7 | 10.5375 | 14.2857 | 57.12 | 79.79
23.9 | 9.6 | 10.4833 | 14.0857 | 51.6 | 79.49
24 | 9.475 | 10.3167 | 13.8429 | 45.85 | 79.1
24.1 | 9.3333 | 10.1571 | 13.6167 | 40.55 | 78.8143
24.2 | 9.2667 | 10.0143 | 13.4429 | 35.86 | 78.5286
24.3 | 9.1143 | 9.8625 | 13.225 | 31.9 | 78.0714
24.4 | 9.0857 | 9.7167 | 13.025 | 28.4 | 77.7286
24.5 | 8.9429 | 9.6167 | 12.8375 | 26.1333 | 77.2889
24.6 | 8.8 | 9.4714 | 12.5833 | 24.225 | 76.8
24.7 | 8.7571 | 9.4 | 12.4 | 22.84 | 76.275
24.8 | 8.6143 | 9.3 | 12.2167 | 21.8 | 75.62
24.9 | 8.54 | 9.1857 | 12.0429 | 20.92 | 74.9333
25 | 8.4714 | 9.0429 | 11.8857 | 20.22 | 74.1
25.1 | 8.4 | 9 | 11.7333 | 19.6 | 73.2
25.2 | 8.2333 | 8.9 | 11.55 | 19.1 | 71.72
25.3 | 8.16 | 8.7857 | 11.3429 | 18.6714 | 70.03
25.4 | 8.1 | 8.7 | 11.2 | 18.275 | 67.875
25.5 | 8.0333 | 8.6 | 11.0714 | 17.9429 | 65
25.6 | 8 | 8.5 | 10.9286 | 17.6125 | 61.79
25.7 | 7.8571 | 8.3833 | 10.7833 | 17.3833 | 57.8
25.8 | 7.8 | 8.3 | 10.6375 | 17.1571 | 53.2444
25.9 | 7.7 | 8.2 | 10.4667 | 16.7833 | 48.6
26 | 7.6 | 8.1143 | 10.4 | 16.5571 | 43.72
26.1 | 7.5375 | 8.04 | 10.25 | 16.3667 | 39.4
26.2 | 7.5 | 7.975 | 10.16 | 16.0833 | 35.55
26.3 | 7.4 | 7.8333 | 10 | 15.9167 | 32.3429
26.4 | 7.4 | 7.8 | 9.85 | 15.6571 | 29.5
26.5 | 7.3 | 7.7 | 9.8 | 15.4714 | 27.375
26.6 | 7.275 | 7.6167 | 9.66 | 15.2 | 25.7143
26.7 | 7.15 | 7.5571 | 9.5 | 15 | 24.4
26.8 | 7.1 | 7.5 | 9.4167 | 14.8111 | 23.4
26.9 | 7.0571 | 7.4 | 9.3 | 14.6167 | 22.6143
27 | 7 | 7.3167 | 9.2 | 14.4 | 21.95
27.1 | 6.9333 | 7.2714 | 9.1 | 14.2125 | 21.5333
27.2 | 6.88 | 7.2 | 8.9875 | 13.9571 | 21.05
27.3 | 6.8 | 7.1 | 8.9 | 13.8143 | 20.62
27.4 |  | 7.1 | 8.8143 | 13.65 | 20.225
27.5 |  | 7 | 8.6667 | 13.4667 | 19.8857
27.6 |  | 7 | 8.6 | 13.3 | 19.5714
27.7 |  | 6.9 | 8.5 | 13.15 | 19.3
27.8 |  | 6.8 | 8.4 | 13.0125 | 19.05
27.9 |  | 6.8 | 8.3143 | 12.8 | 18.8222
28 |  | 6.7 | 8.2571 | 12.6286 | 18.5
28.1 |  | 6.7 | 8.1857 | 12.5714 | 18.35
28.2 |  | 6.6 | 8.05 | 12.4167 | 18.125
28.3 |  |  | 8 | 12.1667 | 17.8714
28.4 |  |  | 7.9 | 12.1 | 17.6
28.5 |  |  | 7.8286 | 11.9286 | 17.38
28.6 |  |  | 7.7667 | 11.75 | 17.1667
28.7 |  |  | 7.7 | 11.7 | 16.925
28.8 |  |  | 7.6 | 11.5333 | 16.7667
28.9 |  |  | 7.5 | 11.38 | 16.6143
29 |  |  | 7.4 | 11.2833 | 16.375
29.1 |  |  | 7.4 | 11.1286 | 16.125
29.2 |  |  | 7.3 | 10.9714 | 15.9375
29.3 |  |  | 7.2429 | 10.8 | 15.76
29.4 |  |  | 7.2 | 10.75 | 15.6125
29.5 |  |  | 7.16 | 10.625 | 15.4364
29.6 |  |  | 7.0125 | 10.5 | 15.24
29.7 |  |  |  | 10.375 | 15
29.8 |  |  |  | 10.25 | 14.8714
29.9 |  |  |  | 10.125 | 14.6889
30 |  |  |  | 10.0857 | 14.5333
30.1 |  |  |  | 9.9429 | 14.4125
30.2 |  |  |  | 9.8 | 14.1875
30.3 |  |  |  | 9.72 | 14.0571
30.4 |  |  |  | 9.65 | 13.9333
30.5 |  |  |  | 9.5 | 13.7727
30.6 |  |  |  | 9.4625 | 13.58
30.7 |  |  |  | 9.3375 | 13.44
30.8 |  |  |  | 9.2125 | 13.3333
30.9 |  |  |  | 9.1 | 13.1444
31 |  |  |  | 9.05 | 13.0111
31.1 |  |  |  | 8.9875 | 12.91
31.2 |  |  |  | 8.8571 | 12.7
31.3 |  |  |  | 8.7857 | 12.6
31.4 |  |  |  | 8.7 | 12.5333
31.5 |  |  |  | 8.6 | 12.4
31.6 |  |  |  | 8.5429 | 12.2778
31.7 |  |  |  | 8.4875 | 12.1333
31.8 |  |  |  | 8.3625 | 11.9429
31.9 |  |  |  | 8.3 | 11.9
32 |  |  |  | 8.2125 | 11.75
32.1 |  |  |  | 8.1 | 11.6125
32.2 |  |  |  | 8.0667 | 11.49
32.3 |  |  |  | 8 | 11.39
32.4 |  |  |  |  | 11.2875
32.5 |  |  |  |  | 11.1667
32.6 |  |  |  |  | 11.0556
32.7 |  |  |  |  | 11
32.8 |  |  |  |  | 10.875
32.9 |  |  |  |  | 10.75
33 |  |  |  |  | 10.6333
33.1 |  |  |  |  | 10.59
33.2 |  |  |  |  | 10.49
33.3 |  |  |  |  | 10.39
33.4 |  |  |  |  | 10.2889
33.5 |  |  |  |  | 10.2
33.6 |  |  |  |  | 10.1375
33.7 |  |  |  |  | 10.0125
33.8 |  |  |  |  | 9.91
33.9 |  |  |  |  | 9.81
34 |  |  |  |  | 9.7182
34.1 |  |  |  |  | 9.62
34.2 |  |  |  |  | 9.52
34.3 |  |  |  |  | 9.4
34.4 |  |  |  |  | 9.38
34.5 |  |  |  |  | 9.28
34.6 |  |  |  |  | 9.18
34.7 |  |  |  |  | 9.075
34.8 |  |  |  |  | 9
34.9 |  |  |  |  | 8.9444
35 |  |  |  |  | 8.9

## Wedge 15 Output Factors
Output factors
Wedge 15 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.88 | 0.894 | 0.901 | 0.909 | 0.915 | 0.92 | 0.924 | 0.927
| 4 | 0.895 | 0.912 | 0.921 | 0.931 | 0.939 | 0.946 | 0.95 | 0.955
| 5 | 0.904 | 0.922 | 0.933 | 0.945 | 0.955 | 0.963 | 0.968 | 0.975
| 7 | 0.916 | 0.937 | 0.949 | 0.965 | 0.978 | 0.989 | 0.996 | 1.004
| 10 | 0.926 | 0.949 | 0.965 | 0.984 | 1 | 1.015 | 1.025 | 1.037
| 15 | 0.935 | 0.961 | 0.978 | 1.001 | 1.022 | 1.042 | 1.056 | 1.072
| 20 | 0.939 | 0.967 | 0.985 | 1.011 | 1.034 | 1.057 | 1.075 | 1.095
| 30 | 0.945 | 0.974 | 0.994 | 1.022 | 1.049 | 1.078 | 1.1 | 1.126
| 40 | 0.949 | 0.979 | 0.999 | 1.029 | 1.058 | 1.09 | 1.115 | 1.144

## Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 27.1 | 33.9 | 45.4 | 55.9 | 60.2
0.1 | 31.3 | 37.9 | 49.425 | 60.175 | 64.05
0.2 | 38.76 | 44.7 | 56.425 | 65.85 | 69.62
0.3 | 47.3 | 53.6 | 63.4 | 72.625 | 75.9
0.4 | 55.55 | 61.16 | 70.14 | 78.225 | 81
0.5 | 62.6 | 67.7 | 75.9167 | 82.8333 | 85.175
0.6 | 68.725 | 73.4333 | 80.35 | 86.45 | 88.575
0.7 | 73.7 | 77.7 | 84.175 | 89.38 | 91.1667
0.8 | 78.06 | 81.55 | 87.175 | 91.9 | 93.24
0.9 | 81.75 | 84.8333 | 89.8 | 93.74 | 95
1 | 84.95 | 87.84 | 92.075 | 95.3 | 96.4
1.1 | 87.6 | 90.25 | 93.84 | 96.6 | 97.4
1.2 | 90.14 | 92 | 95.45 | 97.6 | 98.25
1.3 | 92.06 | 93.8 | 96.5 | 98.34 | 98.85
1.4 | 93.7 | 95.125 | 97.425 | 98.9 | 99.24
1.5 | 95.2 | 96.34 | 98.225 | 99.45 | 99.675
1.6 | 96.3 | 97.325 | 98.74 | 99.65 | 99.86
1.7 | 97.34 | 98.08 | 99.25 | 99.875 | 100
1.8 | 98.08 | 98.8 | 99.55 | 100 | 100
1.9 | 98.8 | 99.2 | 99.86 | 100 | 100
2 | 99.2667 | 99.6 | 99.975 | 100 | 99.95
2.1 | 99.56 | 99.84 | 100 | 99.9 | 99.8
2.2 | 99.8 | 99.96 | 100 | 99.8 | 99.675
2.3 | 99.95 | 100 | 99.9 | 99.6 | 99.5
2.4 | 100 | 100 | 99.84 | 99.48 | 99.325
2.5 | 99.95 | 99.92 | 99.64 | 99.275 | 99.1
2.6 | 99.9 | 99.8 | 99.4 | 98.95 | 98.76
2.7 | 99.775 | 99.7 | 99.2 | 98.725 | 98.54
2.8 | 99.65 | 99.6 | 99.0333 | 98.55 | 98.25
2.9 | 99.5 | 99.275 | 98.8 | 98.16 | 97.98
3 | 99.16 | 99.025 | 98.42 | 97.86 | 97.7333
3.1 | 98.9 | 98.8 | 98.22 | 97.625 | 97.35
3.2 | 98.6 | 98.54 | 97.92 | 97.3 | 96.9
3.3 | 98.2 | 98.22 | 97.6 | 96.96 | 96.7333
3.4 | 97.9 | 98 | 97.25 | 96.6 | 96.36
3.5 | 97.48 | 97.62 | 96.8 | 96.28 | 96.05
3.6 | 97.05 | 97.25 | 96.5 | 95.9667 | 95.7
3.7 | 96.74 | 96.84 | 96.16 | 95.62 | 95.4
3.8 | 96.2 | 96.425 | 95.8 | 95.3 | 95.05
3.9 | 95.82 | 96.1333 | 95.475 | 94.94 | 94.7
4 | 95.48 | 95.8 | 95.125 | 94.575 | 94.36
4.1 | 95.1333 | 95.35 | 94.72 | 94.2 | 94.08
4.2 | 94.575 | 94.94 | 94.4 | 93.9 | 93.75
4.3 | 94.175 | 94.58 | 94.1 | 93.5 | 93.3
4.4 | 93.775 | 94.24 | 93.64 | 93.175 | 93
4.5 | 93.3 | 93.7667 | 93.32 | 92.8333 | 92.66
4.6 | 92.9 | 93.4 | 93 | 92.575 | 92.32
4.7 | 92.46 | 93 | 92.65 | 92.24 | 91.925
4.8 | 92.06 | 92.75 | 92.35 | 91.9 | 91.6
4.9 | 91.675 | 92.35 | 92 | 91.55 | 91.1833
5 | 91.18 | 91.76 | 91.4571 | 91.1 | 90.9333
5.1 | 90.825 | 91.45 | 91.2 | 90.65 | 90.56
5.2 | 90.2 | 91.06 | 90.8 | 90.325 | 90.25
5.3 | 89.8 | 90.72 | 90.35 | 90.1 | 89.9
5.4 | 89.4333 | 90.34 | 90 | 89.675 | 89.52
5.5 | 89.1333 | 89.85 | 89.66 | 89.425 | 89.15
5.6 | 88.65 | 89.4667 | 89.36 | 89.0667 | 88.825
5.7 | 88.1 | 89.2 | 89.125 | 88.7 | 88.5
5.8 | 87.625 | 88.7667 | 88.66 | 88.475 | 88.16
5.9 | 87.15 | 88.5 | 88.4 | 88.05 | 87.86
6 | 86.8 | 88.08 | 87.96 | 87.7333 | 87.55
6.1 | 86.4 | 87.7833 | 87.6 | 87.35 | 87.225
6.2 | 86 | 87.3667 | 87.3 | 87.02 | 86.875
6.3 | 85.65 | 87.02 | 87 | 86.66 | 86.525
6.4 | 85.25 | 86.68 | 86.6667 | 86.325 | 86.1
6.5 | 84.7 | 86.3 | 86.325 | 86 | 85.8667
6.6 | 84.4167 | 85.95 | 86 | 85.72 | 85.6
6.7 | 84.05 | 85.525 | 85.56 | 85.425 | 85.2
6.8 | 83.56 | 85.3 | 85.2667 | 85.06 | 84.86
6.9 | 83.22 | 84.95 | 85 | 84.76 | 84.66
7 | 82.9333 | 84.56 | 84.72 | 84.38 | 84.28
7.1 | 82.5 | 84.15 | 84.3667 | 84.075 | 83.94
7.2 | 82.05 | 83.76 | 83.9667 | 83.84 | 83.7
7.3 | 81.66 | 83.46 | 83.65 | 83.5 | 83.4
7.4 | 81.26 | 83.1 | 83.35 | 83.16 | 82.9667
7.5 | 80.85 | 82.65 | 82.95 | 82.84 | 82.725
7.6 | 80.525 | 82.25 | 82.6 | 82.6 | 82.475
7.7 | 80.1 | 82.0667 | 82.275 | 82.3 | 82.225
7.8 | 79.5667 | 81.75 | 82 | 81.9667 | 81.875
7.9 | 79.35 | 81.3 | 81.7 | 81.64 | 81.5667
8 | 78.9333 | 80.94 | 81.4 | 81.24 | 81.22
8.1 | 78.55 | 80.54 | 81.12 | 80.9667 | 80.9
8.2 | 78.1 | 80.3 | 80.65 | 80.7333 | 80.575
8.3 | 77.7 | 79.9 | 80.36 | 80.34 | 80.3
8.4 | 77.4 | 79.675 | 80.16 | 80.12 | 79.9
8.5 | 77.06 | 79.38 | 79.74 | 79.74 | 79.6667
8.6 | 76.64 | 78.85 | 79.4 | 79.4667 | 79.3
8.7 | 76.3 | 78.5 | 79.05 | 79.08 | 78.975
8.8 | 76.05 | 78.25 | 78.76 | 78.8667 | 78.72
8.9 | 75.56 | 77.8 | 78.54 | 78.475 | 78.475
9 | 75.06 | 77.56 | 78.2 | 78.2 | 78.15
9.1 | 74.7667 | 77.26 | 77.9 | 77.84 | 77.8
9.2 | 74.42 | 76.82 | 77.5 | 77.55 | 77.56
9.3 | 74.06 | 76.48 | 77.3 | 77.3 | 77.24
9.4 | 73.62 | 76.175 | 76.9 | 76.9667 | 76.8667
9.5 | 73.28 | 75.88 | 76.5333 | 76.7 | 76.58
9.6 | 72.85 | 75.55 | 76.18 | 76.3 | 76.32
9.7 | 72.5 | 75.2667 | 75.95 | 76 | 76.02
9.8 | 72.125 | 74.7667 | 75.6 | 75.7667 | 75.6833
9.9 | 71.8667 | 74.5 | 75.36 | 75.45 | 75.4
10 | 71.4286 | 74.18 | 74.98 | 75.16 | 75.06
10.1 | 71.1667 | 73.825 | 74.74 | 74.8667 | 74.725
10.2 | 70.75 | 73.4857 | 74.4 | 74.48 | 74.55
10.3 | 70.38 | 73.05 | 74.1 | 74.28 | 74.16
10.4 | 70 | 72.7 | 73.725 | 73.96 | 73.88
10.5 | 69.6 | 72.4 | 73.48 | 73.7 | 73.5333
10.6 | 69.24 | 72.1 | 73.15 | 73.32 | 73.4
10.7 | 68.82 | 71.875 | 72.9 | 73.06 | 72.9667
10.8 | 68.55 | 71.5 | 72.5 | 72.85 | 72.8
10.9 | 68.14 | 71.15 | 72.32 | 72.48 | 72.38
11 | 67.84 | 70.85 | 72.0167 | 72.15 | 72.16
11.1 | 67.5 | 70.52 | 71.7 | 71.88 | 71.9
11.2 | 67.25 | 70.3 | 71.3833 | 71.7 | 71.625
11.3 | 66.9 | 69.98 | 71.12 | 71.35 | 71.35
11.4 | 66.6 | 69.7 | 70.86 | 71.0333 | 71.06
11.5 | 66.35 | 69.35 | 70.52 | 70.82 | 70.875
11.6 | 66.1 | 69.04 | 70.25 | 70.5 | 70.525
11.7 | 65.82 | 68.74 | 69.98 | 70.3 | 70.22
11.8 | 65.4 | 68.4 | 69.675 | 69.875 | 69.94
11.9 | 65 | 68.14 | 69.38 | 69.64 | 69.64
12 | 64.7 | 67.72 | 69.15 | 69.38 | 69.42
12.1 | 64.35 | 67.45 | 68.8 | 69.125 | 69.1
12.2 | 64.14 | 67.2 | 68.55 | 68.8 | 68.86
12.3 | 63.72 | 66.975 | 68.4 | 68.6 | 68.6
12.4 | 63.42 | 66.725 | 68.1167 | 68.44 | 68.28
12.5 | 62.975 | 66.34 | 67.7333 | 68.1 | 68.04
12.6 | 62.75 | 66.05 | 67.48 | 67.775 | 67.84
12.7 | 62.44 | 65.68 | 67.15 | 67.5 | 67.425
12.8 | 62.2 | 65.3333 | 66.925 | 67.12 | 67.28
12.9 | 61.86 | 65.04 | 66.7 | 66.9333 | 66.98
13 | 61.64 | 64.84 | 66.3 | 66.7333 | 66.68
13.1 | 61.3 | 64.55 | 66 | 66.45 | 66.38
13.2 | 60.975 | 64.26 | 65.8 | 66.2 | 66.1
13.3 | 60.825 | 63.95 | 65.475 | 65.9 | 65.95
13.4 | 60.4 | 63.66 | 65.2 | 65.64 | 65.6667
13.5 | 60.1 | 63.45 | 64.9667 | 65.4333 | 65.4
13.6 | 59.8167 | 63.1 | 64.68 | 65.15 | 65.16
13.7 | 59.625 | 62.8667 | 64.5 | 64.8667 | 64.85
13.8 | 59.3167 | 62.56 | 64.2 | 64.55 | 64.6
13.9 | 58.96 | 62.2 | 63.88 | 64.2667 | 64.4
14 | 58.5667 | 62 | 63.54 | 64.025 | 64.125
14.1 | 58.36 | 61.7 | 63.32 | 63.8333 | 63.9
14.2 | 58.05 | 61.45 | 63.12 | 63.5 | 63.6
14.3 | 57.84 | 61.22 | 62.775 | 63.35 | 63.4
14.4 | 57.525 | 60.925 | 62.575 | 63.05 | 63.0833
14.5 | 57.22 | 60.62 | 62.25 | 62.725 | 62.875
14.6 | 56.9667 | 60.4 | 62.075 | 62.5 | 62.6
14.7 | 56.65 | 60.1333 | 61.7667 | 62.2 | 62.36
14.8 | 56.3 | 59.8 | 61.56 | 62 | 62.15
14.9 | 56.08 | 59.45 | 61.26 | 61.78 | 61.8667
15 | 55.82 | 59.175 | 61.025 | 61.575 | 61.6
15.1 | 55.5429 | 58.925 | 60.74 | 61.28 | 61.35
15.2 | 55.2333 | 58.64 | 60.52 | 61.08 | 61.125
15.3 | 55.02 | 58.34 | 60.3 | 60.78 | 60.8833
15.4 | 54.68 | 58.1 | 59.9833 | 60.5 | 60.65
15.5 | 54.4667 | 57.86 | 59.7333 | 60.3 | 60.4
15.6 | 54.1333 | 57.62 | 59.45 | 60 | 60.1667
15.7 | 53.85 | 57.35 | 59.3 | 59.85 | 59.9
15.8 | 53.45 | 57.0333 | 59 | 59.6167 | 59.65
15.9 | 53.2333 | 56.84 | 58.7667 | 59.28 | 59.4
16 | 53.02 | 56.58 | 58.525 | 59 | 59.075
16.1 | 52.82 | 56.35 | 58.2 | 58.85 | 58.925
16.2 | 52.58 | 56.1 | 58.06 | 58.625 | 58.7167
16.3 | 52.2833 | 55.775 | 57.8 | 58.4 | 58.4333
16.4 | 52.1 | 55.55 | 57.55 | 58.15 | 58.2
16.5 | 51.8 | 55.3 | 57.225 | 57.9 | 58
16.6 | 51.5667 | 55.04 | 57 | 57.7667 | 57.775
16.7 | 51.2 | 54.8 | 56.775 | 57.4833 | 57.525
16.8 | 51.0167 | 54.55 | 56.55 | 57.2333 | 57.26
16.9 | 50.76 | 54.32 | 56.38 | 56.975 | 57.02
17 | 50.575 | 54.05 | 56.2 | 56.8 | 56.85
17.1 | 50.34 | 53.8 | 55.94 | 56.5 | 56.54
17.2 | 50.05 | 53.6 | 55.65 | 56.32 | 56.42
17.3 | 49.8 | 53.35 | 55.4 | 56.1 | 56.25
17.4 | 49.5 | 53.1 | 55.25 | 55.975 | 55.9714
17.5 | 49.3333 | 52.85 | 54.95 | 55.65 | 55.7333
17.6 | 49.1 | 52.675 | 54.8 | 55.425 | 55.575
17.7 | 48.84 | 52.38 | 54.575 | 55.14 | 55.325
17.8 | 48.62 | 52.1667 | 54.36 | 54.98 | 55.1
17.9 | 48.4 | 51.9 | 54.05 | 54.76 | 54.8833
18 | 48.175 | 51.68 | 53.9 | 54.56 | 54.68
18.1 | 47.9 | 51.5 | 53.5833 | 54.26 | 54.4
18.2 | 47.6 | 51.3 | 53.38 | 54.06 | 54.3
18.3 | 47.3 | 51.0333 | 53.18 | 53.875 | 54.02
18.4 | 47.24 | 50.8 | 52.9833 | 53.675 | 53.8
18.5 | 47.05 | 50.55 | 52.75 | 53.45 | 53.5667
18.6 | 46.8 | 50.25 | 52.46 | 53.2 | 53.375
18.7 | 46.6143 | 50.05 | 52.3 | 52.96 | 53.16
18.8 | 46.35 | 49.8333 | 52.1 | 52.7833 | 52.86
18.9 | 46.025 | 49.62 | 51.875 | 52.6 | 52.7
19 | 45.8667 | 49.425 | 51.66 | 52.34 | 52.4667
19.1 | 45.625 | 49.18 | 51.44 | 52.2 | 52.25
19.2 | 45.4333 | 48.975 | 51.2 | 51.96 | 52.02
19.3 | 45.175 | 48.76 | 51 | 51.725 | 51.85
19.4 | 44.9 | 48.46 | 50.76 | 51.6 | 51.625
19.5 | 44.8 | 48.325 | 50.5 | 51.35 | 51.42
19.6 | 44.58 | 48.1 | 50.375 | 51.02 | 51.2333
19.7 | 44.26 | 47.875 | 50.125 | 50.8333 | 51.02
19.8 | 44.0667 | 47.65 | 49.9 | 50.6667 | 50.75
19.9 | 43.875 | 47.3667 | 49.7167 | 50.45 | 50.6
20 | 43.65 | 47.2 | 49.5 | 50.25 | 50.425
20.1 | 43.38 | 47 | 49.18 | 50.04 | 50.2
20.2 | 43.2 | 46.7167 | 49.05 | 49.84 | 49.9167
20.3 | 43.06 | 46.5333 | 48.8333 | 49.625 | 49.8
20.4 | 42.8 | 46.325 | 48.625 | 49.4333 | 49.5333
20.5 | 42.6 | 46.12 | 48.38 | 49.125 | 49.3333
20.6 | 42.3 | 45.9333 | 48.2 | 48.975 | 49.2
20.7 | 42.1 | 45.76 | 47.975 | 48.84 | 49.02
20.8 | 41.8833 | 45.475 | 47.74 | 48.625 | 48.7667
20.9 | 41.7167 | 45.3 | 47.54 | 48.375 | 48.5333
21 | 41.6 | 45.1 | 47.4 | 48.225 | 48.35
21.1 | 41.4 | 44.9 | 47.25 | 48 | 48.2
21.2 | 41.18 | 44.675 | 47.02 | 47.82 | 48
21.3 | 41 | 44.5 | 46.84 | 47.6 | 47.8333
21.4 | 40.74 | 44.275 | 46.6 | 47.42 | 47.64
21.5 | 40.525 | 44.0857 | 46.42 | 47.275 | 47.45
21.6 | 40.3 | 43.9 | 46.175 | 47.04 | 47.2857
21.7 | 40.2 | 43.8 | 46.08 | 46.9 | 47.12
21.8 | 39.9333 | 43.575 | 45.875 | 46.72 | 46.9
21.9 | 39.88 | 43.325 | 45.625 | 46.55 | 46.7
22 | 39.5833 | 43.1 | 45.475 | 46.32 | 46.5167
22.1 | 39.4167 | 42.9167 | 45.3667 | 46.1333 | 46.325
22.2 | 39.22 | 42.7 | 45.2 | 45.96 | 46.2
22.3 | 39.02 | 42.5 | 44.9 | 45.78 | 45.9833
22.4 | 38.9 | 42.4 | 44.7333 | 45.6 | 45.7333
22.5 | 38.65 | 42.1333 | 44.56 | 45.4 | 45.55
22.6 | 38.54 | 41.9667 | 44.375 | 45.2333 | 45.425
22.7 | 38.34 | 41.76 | 44.14 | 45.08 | 45.2
22.8 | 38.2 | 41.4667 | 43.9667 | 44.7714 | 45
22.9 | 38.0333 | 41.3 | 43.8 | 44.6167 | 44.8333
23 | 37.8 | 41.2 | 43.65 | 44.48 | 44.65
23.1 | 37.64 | 40.9333 | 43.46 | 44.3 | 44.4833
23.2 | 37.4 | 40.84 | 43.1833 | 44.15 | 44.35
23.3 | 37.2333 | 40.62 | 43.08 | 43.92 | 44.16
23.4 | 37 | 40.5 | 42.875 | 43.74 | 43.94
23.5 | 36.9 | 40.275 | 42.65 | 43.6 | 43.72
23.6 | 36.65 | 40.06 | 42.5333 | 43.4 | 43.6
23.7 | 36.6 | 39.8667 | 42.3333 | 43.22 | 43.38
23.8 | 36.3833 | 39.76 | 42.125 | 43.0333 | 43.26
23.9 | 36.2 | 39.55 | 41.9 | 42.84 | 43.1
24 | 36 | 39.4 | 41.8 | 42.64 | 42.9
24.1 | 35.9 | 39.2 | 41.6143 | 42.5 | 42.7333
24.2 | 35.6 | 38.9833 | 41.4333 | 42.4 | 42.56
24.3 | 35.475 | 38.8167 | 41.3 | 42.2 | 42.35
24.4 | 35.34 | 38.7 | 41.12 | 42 | 42.225
24.5 | 35.175 | 38.58 | 40.9 | 41.8 | 42
24.6 | 34.96 | 38.325 | 40.7 | 41.675 | 41.78
24.7 | 34.8 | 38.1 | 40.575 | 41.5 | 41.625
24.8 | 34.6 | 37.9 | 40.34 | 41.2833 | 41.5
24.9 | 34.3333 | 37.7333 | 40.15 | 41.08 | 41.3333
25 | 34.225 | 37.6 | 40.04 | 40.96 | 41.14
25.1 | 34.12 | 37.5 | 39.825 | 40.825 | 40.95
25.2 | 33.92 | 37.28 | 39.775 | 40.6 | 40.8
25.3 | 33.8 | 37.1 | 39.525 | 40.3667 | 40.6
25.4 | 33.6 | 36.9 | 39.44 | 40.26 | 40.5
25.5 | 33.475 | 36.8 | 39.2 | 40.15 | 40.35
25.6 | 33.3 | 36.6 | 38.975 | 39.94 | 40.15
25.7 | 33.15 | 36.4 | 38.82 | 39.8 | 40.0333
25.8 | 33.02 | 36.28 | 38.7 | 39.7 | 39.9
25.9 | 32.82 | 36.08 | 38.5 | 39.46 | 39.7
26 | 32.68 | 35.95 | 38.38 | 39.3 | 39.48
26.1 | 32.56 | 35.78 | 38.1667 | 39.125 | 39.375
26.2 | 32.32 | 35.6 | 38 | 38.98 | 39.225
26.3 | 32.2333 | 35.4 | 37.9 | 38.8 | 39.1
26.4 | 32.04 | 35.3 | 37.7 | 38.68 | 38.9
26.5 | 31.925 | 35.2 | 37.5667 | 38.5 | 38.74
26.6 | 31.8 | 35.0333 | 37.4 | 38.4 | 38.6
26.7 | 31.625 | 34.8 | 37.22 | 38.2 | 38.4
26.8 | 31.52 | 34.6833 | 37.12 | 38.1 | 38.3
26.9 | 31.3 | 34.5 | 37 | 37.9 | 38.1
27 | 31.16 | 34.38 | 36.76 | 37.8 | 37.9333
27.1 | 31.04 | 34.2 | 36.6 | 37.6 | 37.8
27.2 | 30.88 | 34.1 | 36.5 | 37.48 | 37.6
27.3 | 30.7 | 33.9 | 36.38 | 37.28 | 37.5
27.4 | 30.6 | 33.725 | 36.15 | 37.16 | 37.35
27.5 | 30.425 | 33.6 | 36.02 | 37.025 | 37.2667
27.6 | 30.32 | 33.5 | 35.9 | 36.8667 | 37.1
27.7 | 30.15 | 33.3 | 35.7333 | 36.7 | 36.875
27.8 | 30 | 33.16 | 35.525 | 36.5 | 36.7
27.9 | 29.9 | 32.96 | 35.36 | 36.4167 | 36.6167
28 | 29.7167 | 32.7833 | 35.28 | 36.325 | 36.45
28.1 | 29.52 | 32.675 | 35.1 | 36.1167 | 36.3
28.2 | 29.4 | 32.5 | 34.92 | 36 | 36.14
28.3 | 29.2667 | 32.32 | 34.8 | 35.9 | 36.0333
28.4 | 29.18 | 32.2 | 34.72 | 35.7 | 35.9
28.5 | 29.0833 | 32.1 | 34.575 | 35.52 | 35.725
28.6 | 28.9 | 31.95 | 34.4 | 35.4 | 35.58
28.7 | 28.72 | 31.8 | 34.25 | 35.2 | 35.48
28.8 | 28.62 | 31.6833 | 34.1 | 35.1 | 35.3
28.9 | 28.5 | 31.6 | 33.9833 | 34.9 | 35.1833
29 | 28.4 | 31.4 | 33.7833 | 34.8 | 35
29.1 | 28.2 | 31.3 | 33.7 | 34.7 | 34.88
29.2 | 28.1 | 31.1 | 33.475 | 34.5 | 34.7
29.3 | 27.9 | 30.95 | 33.375 | 34.4 | 34.55
29.4 | 27.8 | 30.8 | 33.24 | 34.3 | 34.42
29.5 | 27.7 | 30.7 | 33.14 | 34.1 | 34.32
29.6 | 27.5 | 30.54 | 32.925 | 33.925 | 34.2
29.7 | 27.4 | 30.4 | 32.8333 | 33.8 | 34.08
29.8 | 27.3 | 30.3 | 32.6667 | 33.65 | 33.9
29.9 | 27.15 | 30.125 | 32.55 | 33.6 | 33.7333
30 | 26.95 | 29.98 | 32.44 | 33.45 | 33.6

## SSD80 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 28.4 | 38.1 | 55.5 | 69.2 | 73.8
0.1 | 32.3625 | 41.975 | 59.525 | 72.825 | 77.35
0.2 | 39.2 | 48.6 | 65.425 | 77.8667 | 81.8
0.3 | 48.28 | 57 | 72.7 | 83.3 | 86.65
0.4 | 56.5 | 64.3 | 78.525 | 87.65 | 90.4
0.5 | 63.775 | 71.15 | 82.98 | 91.05 | 93.25
0.6 | 69.9125 | 76.475 | 86.9 | 93.5714 | 95.3
0.7 | 75.2111 | 80.925 | 89.85 | 95.475 | 96.9
0.8 | 79.3571 | 84.44 | 92.4 | 96.975 | 98
0.9 | 83.25 | 87.8 | 94.35 | 98 | 98.82
1 | 86.3222 | 90.075 | 95.9 | 98.825 | 99.35
1.1 | 88.9 | 92.225 | 97.05 | 99.3 | 99.7667
1.2 | 91.2 | 94.04 | 98 | 99.75 | 99.925
1.3 | 93.05 | 95.45 | 98.75 | 99.9 | 100
1.4 | 94.75 | 96.7 | 99.3 | 100 | 100
1.5 | 96.1 | 97.6 | 99.625 | 100 | 100
1.6 | 97.175 | 98.4 | 99.9 | 100 | 99.9
1.7 | 97.9667 | 99.05 | 100 | 99.8667 | 99.7
1.8 | 98.7167 | 99.4 | 100 | 99.675 | 99.475
1.9 | 99.2333 | 99.65 | 100 | 99.425 | 99.24
2 | 99.62 | 100 | 99.9 | 99.22 | 98.95
2.1 | 99.9 | 100 | 99.725 | 98.925 | 98.625
2.2 | 100 | 100 | 99.48 | 98.66 | 98.3
2.3 | 100 | 100 | 99.34 | 98.3333 | 97.95
2.4 | 100 | 99.975 | 99.1 | 97.9667 | 97.5667
2.5 | 99.9667 | 99.7 | 98.725 | 97.575 | 97.2333
2.6 | 99.85 | 99.475 | 98.5 | 97.25 | 96.9
2.7 | 99.55 | 99.325 | 98.05 | 96.9 | 96.5
2.8 | 99.3 | 99.05 | 97.6 | 96.5 | 96.1
2.9 | 98.95 | 98.7 | 97.3 | 96.1 | 95.7333
3 | 98.7 | 98.4 | 97 | 95.675 | 95.3667
3.1 | 98.3571 | 98.0333 | 96.45 | 95.3 | 95
3.2 | 97.95 | 97.7 | 96.1 | 94.95 | 94.5667
3.3 | 97.475 | 97.15 | 95.775 | 94.4625 | 94.15
3.4 | 97.1286 | 96.8 | 95.35 | 94.05 | 93.75
3.5 | 96.65 | 96.4 | 94.95 | 93.7286 | 93.325
3.6 | 96.1875 | 95.9 | 94.55 | 93.3 | 92.95
3.7 | 95.75 | 95.5 | 94.075 | 92.8571 | 92.525
3.8 | 95.25 | 95.1333 | 93.65 | 92.44 | 92.25
3.9 | 94.85 | 94.65 | 93.25 | 92.1 | 91.78
4 | 94.35 | 94.3 | 92.75 | 91.7 | 91.4
4.1 | 93.9125 | 93.8 | 92.36 | 91.2 | 91
4.2 | 93.45 | 93.3 | 92 | 90.9 | 90.6
4.3 | 92.9375 | 92.975 | 91.6 | 90.54 | 90.25
4.4 | 92.38 | 92.4 | 91.2 | 90.15 | 89.8
4.5 | 91.85 | 92.075 | 90.72 | 89.75 | 89.4
4.6 | 91.4 | 91.64 | 90.325 | 89.325 | 89
4.7 | 90.9333 | 91.16 | 89.95 | 88.9 | 88.5667
4.8 | 90.55 | 90.775 | 89.5 | 88.5 | 88.2333
4.9 | 90.05 | 90.3 | 89.1 | 88.05 | 87.8
5 | 89.65 | 89.9 | 88.7 | 87.725 | 87.5333
5.1 | 89.1333 | 89.5 | 88.24 | 87.36 | 87.125
5.2 | 88.7 | 88.96 | 87.875 | 86.95 | 86.65
5.3 | 88.15 | 88.6 | 87.525 | 86.4625 | 86.28
5.4 | 87.8 | 88.15 | 87.06 | 86.14 | 85.9
5.5 | 87.225 | 87.7 | 86.7 | 85.8 | 85.5333
5.6 | 86.725 | 87.3 | 86.3 | 85.4 | 85.1
5.7 | 86.325 | 86.95 | 85.9 | 85.02 | 84.75
5.8 | 85.9 | 86.525 | 85.4 | 84.675 | 84.375
5.9 | 85.42 | 86.05 | 85.175 | 84.25 | 84.0333
6 | 84.8 | 85.625 | 84.75 | 83.95 | 83.6625
6.1 | 84.525 | 85.2 | 84.375 | 83.525 | 83.3
6.2 | 84.14 | 84.95 | 84.05 | 83.2 | 82.95
6.3 | 83.6 | 84.4 | 83.65 | 82.8 | 82.6
6.4 | 83.3 | 84.05 | 83.3 | 82.4 | 82.2
6.5 | 82.75 | 83.5 | 82.8 | 82 | 81.9
6.6 | 82.34 | 83.22 | 82.425 | 81.6 | 81.4667
6.7 | 81.85 | 82.85 | 82 | 81.275 | 81.1333
6.8 | 81.475 | 82.45 | 81.675 | 80.95 | 80.775
6.9 | 81.05 | 82.05 | 81.4 | 80.65 | 80.44
7 | 80.5 | 81.6 | 80.95 | 80.22 | 80.0667
7.1 | 80.15 | 81.35 | 80.6 | 79.9 | 79.7
7.2 | 79.65 | 80.9 | 80.25 | 79.5 | 79.35
7.3 | 79.25 | 80.4 | 79.825 | 79.2 | 79.1
7.4 | 78.875 | 79.9 | 79.5 | 78.8333 | 78.7
7.5 | 78.5 | 79.65 | 79.1 | 78.425 | 78.4
7.6 | 77.9 | 79.36 | 78.75 | 78.1 | 78
7.7 | 77.54 | 78.8 | 78.3667 | 77.8286 | 77.6333
7.8 | 77.18 | 78.5 | 78.025 | 77.45 | 77.275
7.9 | 76.8 | 78.16 | 77.6714 | 77.1625 | 76.925
8 | 76.5333 | 77.75 | 77.3222 | 76.78 | 76.575
8.1 | 76 | 77.35 | 76.975 | 76.375 | 76.25
8.2 | 75.7 | 77.025 | 76.6375 | 76.1 | 76
8.3 | 75.1667 | 76.64 | 76.28 | 75.75 | 75.65
8.4 | 74.7 | 76.2 | 75.925 | 75.375 | 75.3
8.5 | 74.375 | 75.82 | 75.55 | 75 | 74.95
8.6 | 73.975 | 75.525 | 75.25 | 74.6429 | 74.6
8.7 | 73.575 | 75.075 | 74.9 | 74.4 | 74.22
8.8 | 73.2 | 74.65 | 74.46 | 74.05 | 73.925
8.9 | 72.8571 | 74.3571 | 74.15 | 73.725 | 73.575
9 | 72.4 | 73.9667 | 73.9 | 73.375 | 73.25
9.1 | 72 | 73.675 | 73.5 | 73.05 | 72.95
9.2 | 71.7 | 73.25 | 73.1 | 72.7 | 72.625
9.3 | 71.4 | 72.8625 | 72.75 | 72.45 | 72.275
9.4 | 70.9 | 72.4875 | 72.5 | 72.1 | 71.95
9.5 | 70.4333 | 72.2 | 72.125 | 71.825 | 71.56
9.6 | 70.025 | 71.8 | 71.8 | 71.425 | 71.25
9.7 | 69.6571 | 71.4 | 71.4 | 71.075 | 71
9.8 | 69.3333 | 71.0333 | 71.1 | 70.725 | 70.7
9.9 | 68.975 | 70.75 | 70.75 | 70.45 | 70.3333
10 | 68.625 | 70.34 | 70.5 | 70.1 | 69.95
10.1 | 68.25 | 69.975 | 70.05 | 69.9 | 69.65
10.2 | 67.85 | 69.65 | 69.7 | 69.55 | 69.4
10.3 | 67.44 | 69.3 | 69.375 | 69.2 | 69.075
10.4 | 67.2 | 68.95 | 69.125 | 68.8 | 68.75
10.5 | 66.7 | 68.5667 | 68.8 | 68.4333 | 68.45
10.6 | 66.3 | 68.2 | 68.42 | 68.2222 | 68.1
10.7 | 66.05 | 67.86 | 68.18 | 67.9625 | 67.86
10.8 | 65.7 | 67.55 | 67.85 | 67.5875 | 67.54
10.9 | 65.3333 | 67.2 | 67.5 | 67.2889 | 67.25
11 | 64.975 | 66.9 | 67.25 | 66.975 | 66.925
11.1 | 64.65 | 66.6 | 66.9 | 66.7 | 66.6333
11.2 | 64.25 | 66.22 | 66.55 | 66.425 | 66.325
11.3 | 63.9 | 65.925 | 66.2 | 66.1111 | 65.975
11.4 | 63.6 | 65.575 | 65.95 | 65.7857 | 65.725
11.5 | 63.275 | 65.3 | 65.7 | 65.5222 | 65.475
11.6 | 62.9 | 64.875 | 65.36 | 65.2625 | 65.2
11.7 | 62.62 | 64.625 | 65.025 | 64.9 | 64.88
11.8 | 62.15 | 64.275 | 64.675 | 64.625 | 64.6
11.9 | 61.8333 | 63.95 | 64.46 | 64.4 | 64.25
12 | 61.475 | 63.7 | 64.1375 | 64.0667 | 64
12.1 | 61.2429 | 63.35 | 63.8 | 63.75 | 63.75
12.2 | 60.9125 | 62.9 | 63.55 | 63.4 | 63.42
12.3 | 60.5429 | 62.7 | 63.25 | 63.1286 | 63.16
12.4 | 60.2889 | 62.4 | 62.875 | 62.875 | 62.8
12.5 | 59.9 | 62.075 | 62.7 | 62.625 | 62.6
12.6 | 59.6 | 61.825 | 62.35 | 62.4 | 62.35
12.7 | 59.325 | 61.45 | 62.05 | 62.0333 | 62.025
12.8 | 58.95 | 61.15 | 61.725 | 61.8 | 61.75
12.9 | 58.7143 | 60.8 | 61.55 | 61.45 | 61.5
13 | 58.4 | 60.5333 | 61.225 | 61.2 | 61.2
13.1 | 57.975 | 60.1667 | 60.95 | 60.85 | 60.9
13.2 | 57.6875 | 59.9 | 60.625 | 60.62 | 60.7
13.3 | 57.375 | 59.65 | 60.38 | 60.375 | 60.3667
13.4 | 57.0375 | 59.325 | 60.15 | 60.025 | 60.1333
13.5 | 56.8571 | 59 | 59.8 | 59.8 | 59.85
13.6 | 56.44 | 58.8 | 59.5333 | 59.5714 | 59.58
13.7 | 56.2 | 58.45 | 59.175 | 59.3333 | 59.3
13.8 | 55.96 | 58.2 | 58.925 | 58.9 | 59
13.9 | 55.64 | 57.95 | 58.7 | 58.75 | 58.75
14 | 55.375 | 57.64 | 58.45 | 58.475 | 58.5
14.1 | 55.025 | 57.375 | 58.1 | 58.3 | 58.175
14.2 | 54.775 | 57.05 | 57.9 | 58.0333 | 57.925
14.3 | 54.45 | 56.75 | 57.625 | 57.6875 | 57.675
14.4 | 54.175 | 56.425 | 57.375 | 57.375 | 57.425
14.5 | 53.85 | 56.175 | 57.025 | 57.2125 | 57.175
14.6 | 53.5 | 55.925 | 56.775 | 56.9125 | 56.94
14.7 | 53.2 | 55.575 | 56.525 | 56.6571 | 56.7333
14.8 | 53 | 55.34 | 56.175 | 56.42 | 56.375
14.9 | 52.76 | 55.1 | 55.925 | 56.1444 | 56.125
15 | 52.525 | 54.8 | 55.675 | 55.9 | 55.88
15.1 | 52.075 | 54.6 | 55.38 | 55.55 | 55.6333
15.2 | 51.825 | 54.2429 | 55.1 | 55.2857 | 55.4
15.3 | 51.575 | 53.9 | 54.95 | 55.1 | 55.15
15.4 | 51.35 | 53.65 | 54.6667 | 54.8667 | 54.9
15.5 | 51.1 | 53.4 | 54.45 | 54.6 | 54.65
15.6 | 50.825 | 53.15 | 54.15 | 54.4111 | 54.4
15.7 | 50.55 | 52.9 | 53.9 | 54.0857 | 54.175
15.8 | 50.3 | 52.55 | 53.6 | 53.8625 | 53.925
15.9 | 49.98 | 52.325 | 53.4 | 53.5875 | 53.575
16 | 49.75 | 52.08 | 53.15 | 53.425 | 53.425
16.1 | 49.5 | 51.85 | 52.9 | 53.175 | 53.2167
16.2 | 49.15 | 51.7 | 52.675 | 52.925 | 53
16.3 | 49 | 51.4 | 52.425 | 52.7333 | 52.75
16.4 | 48.75 | 51.15 | 52.25 | 52.475 | 52.425
16.5 | 48.4 | 50.84 | 51.9667 | 52.2667 | 52.225
16.6 | 48.175 | 50.575 | 51.6714 | 52.02 | 52
16.7 | 47.925 | 50.325 | 51.4 | 51.775 | 51.8
16.8 | 47.7167 | 50.075 | 51.2 | 51.525 | 51.5
16.9 | 47.475 | 49.825 | 50.95 | 51.225 | 51.35
17 | 47.15 | 49.55 | 50.6 | 51.06 | 51.1
17.1 | 46.975 | 49.28 | 50.475 | 50.8 | 50.875
17.2 | 46.725 | 49.1 | 50.2 | 50.5 | 50.6
17.3 | 46.45 | 48.75 | 50 | 50.35 | 50.3333
17.4 | 46.2 | 48.52 | 49.8 | 50.1 | 50.2
17.5 | 46 | 48.35 | 49.6 | 49.9 | 50
17.6 | 45.75 | 48.05 | 49.325 | 49.7 | 49.725
17.7 | 45.5 | 47.875 | 49.1 | 49.45 | 49.5
17.8 | 45.3 | 47.625 | 48.9 | 49.2 | 49.3333
17.9 | 45 | 47.375 | 48.7 | 49 | 49.06
18 | 44.76 | 47.15 | 48.4 | 48.85 | 48.925
18.1 | 44.525 | 46.9 | 48.2 | 48.5667 | 48.675
18.2 | 44.3 | 46.65 | 47.925 | 48.3 | 48.425
18.3 | 44.05 | 46.4 | 47.8 | 48.16 | 48.2
18.4 | 43.875 | 46.25 | 47.55 | 47.925 | 48.05
18.5 | 43.625 | 46 | 47.3 | 47.6571 | 47.8
18.6 | 43.375 | 45.775 | 47 | 47.475 | 47.55
18.7 | 43.225 | 45.525 | 46.9 | 47.225 | 47.4
18.8 | 42.975 | 45.275 | 46.6333 | 46.975 | 47.15
18.9 | 42.75 | 45.025 | 46.35 | 46.825 | 46.9
19 | 42.525 | 44.825 | 46.22 | 46.58 | 46.7
19.1 | 42.3375 | 44.58 | 46 | 46.4333 | 46.5
19.2 | 42.08 | 44.4 | 45.775 | 46.2 | 46.325
19.3 | 41.8 | 44.25 | 45.5 | 46 | 46.075
19.4 | 41.65 | 44 | 45.3667 | 45.85 | 45.9
19.5 | 41.4 | 43.78 | 45.12 | 45.6 | 45.675
19.6 | 41.2 | 43.55 | 44.975 | 45.35 | 45.5
19.7 | 41 | 43.3 | 44.725 | 45.2 | 45.18
19.8 | 40.825 | 43.1 | 44.48 | 44.975 | 45
19.9 | 40.6 | 42.85 | 44.275 | 44.725 | 44.825
20 | 40.3333 | 42.7 | 44.0714 | 44.58 | 44.6
20.1 | 40.1667 | 42.46 | 43.875 | 44.35 | 44.45
20.2 | 39.9 | 42.3 | 43.625 | 44.2 | 44.3
20.3 | 39.6667 | 42.05 | 43.5 | 43.95 | 44.05
20.4 | 39.525 | 41.825 | 43.25 | 43.7 | 43.9
20.5 | 39.275 | 41.575 | 43.05 | 43.5 | 43.6
20.6 | 39.05 | 41.425 | 42.85 | 43.3667 | 43.4
20.7 | 38.8 | 41.275 | 42.6 | 43.1 | 43.275
20.8 | 38.7 | 41.025 | 42.4556 | 42.95 | 43.025
20.9 | 38.5 | 40.775 | 42.2625 | 42.7 | 42.875
21 | 38.3 | 40.625 | 42.075 | 42.55 | 42.72
21.1 | 38.1 | 40.4 | 41.9 | 42.325 | 42.475
21.2 | 37.9 | 40.2333 | 41.6 | 42.175 | 42.325
21.3 | 37.75 | 39.9 | 41.48 | 41.9833 | 42.175
21.4 | 37.6 | 39.8 | 41.3 | 41.775 | 41.925
21.5 | 37.3 | 39.62 | 41.1 | 41.625 | 41.8
21.6 | 37.12 | 39.475 | 40.8667 | 41.4 | 41.6
21.7 | 36.975 | 39.225 | 40.725 | 41.24 | 41.375
21.8 | 36.825 | 39.075 | 40.5667 | 41 | 41.25
21.9 | 36.65 | 38.9 | 40.3 | 40.85 | 41
22 | 36.375 | 38.675 | 40.1833 | 40.6 | 40.8
22.1 | 36.25 | 38.48 | 39.9571 | 40.5 | 40.6
22.2 | 35.94 | 38.22 | 39.7778 | 40.3 | 40.475
22.3 | 35.86 | 38.08 | 39.5667 | 40.15 | 40.225
22.4 | 35.6714 | 37.85 | 39.4 | 40 | 40.1
22.5 | 35.5 | 37.7 | 39.2 | 39.8 | 39.925
22.6 | 35.3 | 37.4778 | 39.05 | 39.6 | 39.775
22.7 | 35.12 | 37.3625 | 38.8 | 39.4 | 39.525
22.8 | 34.875 | 37.175 | 38.7 | 39.25 | 39.4
22.9 | 34.725 | 37.025 | 38.5 | 39.1 | 39.25
23 | 34.575 | 36.8 | 38.3 | 38.875 | 39
23.1 | 34.425 | 36.6286 | 38.15 | 38.725 | 38.85
23.2 | 34.2 | 36.4 | 38 | 38.575 | 38.7
23.3 | 34.05 | 36.26 | 37.75 | 38.425 | 38.5
23.4 | 33.875 | 36.1 | 37.62 | 38.2 | 38.3833
23.5 | 33.75 | 35.96 | 37.475 | 38.05 | 38.175
23.6 | 33.6 | 35.725 | 37.325 | 37.9 | 38.025
23.7 | 33.4 | 35.6 | 37.075 | 37.6667 | 37.8
23.8 | 33.25 | 35.4 | 36.9 | 37.525 | 37.6667
23.9 | 33.1 | 35.2 | 36.725 | 37.4 | 37.55
24 | 32.9 | 35.05 | 36.6 | 37.26 | 37.35
24.1 | 32.7 | 34.8333 | 36.4 | 37.1 | 37.1667
24.2 | 32.5 | 34.7 | 36.2 | 36.8714 | 37.025
24.3 | 32.325 | 34.5 | 36.1 | 36.6556 | 36.9
24.4 | 32.2 | 34.4 | 35.95 | 36.5375 | 36.6556
24.5 | 32.05 | 34.2 | 35.7 | 36.325 | 36.4625
24.6 | 31.9 | 34.025 | 35.5333 | 36.2 | 36.3375
24.7 | 31.75 | 33.8125 | 35.4 | 36.06 | 36.2125
24.8 | 31.5667 | 33.7 | 35.2429 | 35.9 | 36.075
24.9 | 31.45 | 33.575 | 35.05 | 35.7167 | 35.85
25 | 31.2333 | 33.4 | 34.975 | 35.5333 | 35.7
25.1 | 31.1286 | 33.2 | 34.7625 | 35.4333 | 35.5444
25.2 | 31 | 33.05 | 34.575 | 35.2 | 35.3875
25.3 | 30.8 | 32.85 | 34.4 | 35.1 | 35.2667
25.4 | 30.6 | 32.7143 | 34.2 | 34.92 | 35.0857
25.5 | 30.4 | 32.6 | 34.075 | 34.7 | 34.875
25.6 | 30.3 | 32.375 | 33.9571 | 34.6 | 34.75
25.7 | 30.15 | 32.2556 | 33.8143 | 34.5 | 34.55
25.8 | 30.025 | 32.1375 | 33.675 | 34.3 | 34.4
25.9 | 29.8 | 31.9833 | 33.5 | 34.1333 | 34.275
26 | 29.7286 | 31.8375 | 33.35 | 33.98 | 34.15
26.1 | 29.6 | 31.7 | 33.1875 | 33.84 | 34.0143
26.2 | 29.4 | 31.4889 | 33.025 | 33.7 | 33.9
26.3 | 29.3 | 31.35 | 32.9 | 33.575 | 33.7
26.4 | 29.075 | 31.2 | 32.75 | 33.425 | 33.575
26.5 | 28.95 | 31.075 | 32.6 | 33.2667 | 33.425
26.6 | 28.8 | 30.9 | 32.45 | 33.1 | 33.3
26.7 | 28.65 | 30.7 | 32.3 | 32.95 | 33.1
26.8 | 28.55 | 30.6 | 32.125 | 32.8 | 32.95
26.9 | 28.4 | 30.475 | 32 | 32.7 | 32.8
27 | 28.2 | 30.3 | 31.9 | 32.525 | 32.7
27.1 | 28.125 | 30.175 | 31.7 | 32.375 | 32.56
27.2 | 27.9 | 30.04 | 31.6 | 32.225 | 32.425
27.3 | 27.84 | 29.9 | 31.5 | 32.18 | 32.2
27.4 | 27.7 | 29.75 | 31.3 | 31.95 | 32.1
27.5 | 27.525 | 29.6 | 31.1 | 31.8 | 31.975
27.6 | 27.5 | 29.45 | 30.95 | 31.7 | 31.825
27.7 | 27.3 | 29.3 | 30.8 | 31.575 | 31.7
27.8 | 27.1 | 29.2 | 30.65 | 31.4 | 31.525
27.9 | 27 | 29 | 30.6 | 31.2 | 31.4
28 | 26.9 | 28.9 | 30.3875 | 31.1 | 31.3
28.1 | 26.8 | 28.725 | 30.2625 | 30.9 | 31.125
28.2 | 26.675 | 28.6 | 30.1375 | 30.8 | 31
28.3 | 26.525 | 28.5 | 30.0125 | 30.72 | 30.8
28.4 | 26.375 | 28.3 | 29.8875 | 30.6 | 30.7
28.5 | 26.3 | 28.2 | 29.7625 | 30.425 | 30.6
28.6 | 26.1 | 28.06 | 29.5875 | 30.3 | 30.4
28.7 | 25.9 | 27.9 | 29.425 | 30.2 | 30.3
28.8 | 25.8 | 27.775 | 29.375 | 30 | 30.2
28.9 | 25.75 | 27.625 | 29.2 | 29.9 | 30.05
29 | 25.6 | 27.5 | 29.075 | 29.78 | 29.8556
29.1 | 25.45 | 27.4 | 28.925 | 29.6 | 29.8
29.2 | 25.3 | 27.3 | 28.8 | 29.5 | 29.6222
29.3 | 25.2 | 27.15 | 28.7 | 29.35 | 29.4875
29.4 | 25.1 | 27 | 28.6 | 29.2 | 29.3625
29.5 | 24.96 | 26.85 | 28.425 | 29.1667 | 29.2375
29.6 | 24.825 | 26.8 | 28.3 | 29 | 29.1625
29.7 | 24.7 | 26.6 | 28.15 | 28.8 | 29
29.8 | 24.6 | 26.5 | 28.1 | 28.7 | 28.8571
29.9 | 24.5 | 26.4 | 27.95 | 28.6 | 28.78
30 | 24.4 | 26.3 | 27.7667 | 28.45 | 28.65

## SSD120 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 26.6 | 32.4 | 42.7 | 51.7 | 55.1
0.1 | 30.975 | 37.25 | 46.925 | 55.5 | 59.3125
0.2 | 37.75 | 43.775 | 52.875 | 61.225 | 64.875
0.3 | 46.8375 | 52.1286 | 60.775 | 68.25 | 71.275
0.4 | 54.8375 | 59.6143 | 67.1667 | 74.15 | 76.66
0.5 | 62.0125 | 66.425 | 73.275 | 79.225 | 81.4
0.6 | 67.4 | 71.725 | 77.55 | 83.18 | 85.05
0.7 | 72.4 | 76 | 81.78 | 86.3 | 88.1
0.8 | 76.85 | 80.22 | 84.94 | 89 | 90.5
0.9 | 80.7 | 83.7 | 87.8 | 91.5 | 92.7
1 | 84.08 | 86.4 | 90.15 | 93.3 | 94.15
1.1 | 86.675 | 88.95 | 92.4 | 94.85 | 95.6
1.2 | 89.075 | 91.1 | 93.8 | 96.1 | 96.7
1.3 | 91.35 | 92.75 | 95.2 | 97.15 | 97.625
1.4 | 93.1 | 94.4 | 96.35 | 97.85 | 98.275
1.5 | 94.3 | 95.5 | 97.3 | 98.6 | 98.92
1.6 | 95.7 | 96.5 | 97.975 | 99.05 | 99.275
1.7 | 96.8 | 97.5 | 98.575 | 99.5 | 99.575
1.8 | 97.65 | 98.225 | 99.18 | 99.64 | 99.82
1.9 | 98.4 | 98.7 | 99.55 | 99.875 | 99.95
2 | 98.9667 | 99.275 | 99.825 | 100 | 100
2.1 | 99.35 | 99.5 | 99.9 | 100 | 100
2.2 | 99.775 | 99.66 | 100 | 100 | 100
2.3 | 99.9 | 99.9 | 100 | 99.925 | 99.9
2.4 | 99.9778 | 100 | 100 | 99.9 | 99.8125
2.5 | 100 | 100 | 99.9333 | 99.8 | 99.5875
2.6 | 100 | 99.95 | 99.8667 | 99.68 | 99.44
2.7 | 100 | 99.8 | 99.7 | 99.425 | 99.1875
2.8 | 99.875 | 99.65 | 99.54 | 99.2 | 99.0333
2.9 | 99.7333 | 99.46 | 99.3 | 99 | 98.7857
3 | 99.5333 | 99.3 | 99.05 | 98.8 | 98.4375
3.1 | 99.18 | 99.1833 | 98.9 | 98.55 | 98.3
3.2 | 98.9 | 98.9333 | 98.6 | 98.2667 | 97.95
3.3 | 98.625 | 98.5625 | 98.3 | 97.9 | 97.65
3.4 | 98.4 | 98.2333 | 98.05 | 97.55 | 97.425
3.5 | 97.95 | 97.9857 | 97.725 | 97.3 | 97.15
3.6 | 97.7 | 97.66 | 97.5 | 96.95 | 96.8857
3.7 | 97.3 | 97.375 | 97.075 | 96.6 | 96.6286
3.8 | 96.9 | 97.025 | 96.925 | 96.36 | 96.25
3.9 | 96.55 | 96.65 | 96.65 | 96.125 | 95.98
4 | 96.22 | 96.325 | 96.125 | 95.875 | 95.6
4.1 | 95.675 | 95.98 | 95.875 | 95.625 | 95.3
4.2 | 95.425 | 95.65 | 95.5333 | 95.22 | 94.9
4.3 | 95.025 | 95.3 | 95.2333 | 94.88 | 94.6
4.4 | 94.58 | 94.9857 | 94.9 | 94.55 | 94.4333
4.5 | 94.15 | 94.75 | 94.575 | 94.275 | 94.05
4.6 | 93.85 | 94.4286 | 94.32 | 93.92 | 93.725
4.7 | 93.65 | 93.975 | 93.975 | 93.6667 | 93.38
4.8 | 93.2333 | 93.65 | 93.6 | 93.3333 | 93.15
4.9 | 92.675 | 93.4 | 93.35 | 93 | 92.85
5 | 92.18 | 92.9667 | 93 | 92.625 | 92.45
5.1 | 91.8167 | 92.6 | 92.7 | 92.35 | 92.28
5.2 | 91.375 | 92.2 | 92.3 | 92.15 | 91.8667
5.3 | 90.9 | 91.9 | 91.95 | 91.75 | 91.6
5.4 | 90.55 | 91.675 | 91.82 | 91.5 | 91.35
5.5 | 90.225 | 91.25 | 91.45 | 91.2 | 91.2
5.6 | 89.9333 | 90.925 | 91.025 | 90.825 | 90.72
5.7 | 89.46 | 90.58 | 90.7 | 90.7 | 90.325
5.8 | 89.025 | 90.3143 | 90.45 | 90.35 | 90.175
5.9 | 88.9 | 89.95 | 90.18 | 90 | 89.85
6 | 88.5 | 89.6 | 89.9333 | 89.7 | 89.5
6.1 | 88.1 | 89.25 | 89.5 | 89.35 | 89.25
6.2 | 87.8333 | 88.8667 | 89.3 | 89.04 | 89
6.3 | 87.3 | 88.575 | 88.9 | 88.775 | 88.55
6.4 | 87 | 88.225 | 88.65 | 88.45 | 88.3
6.5 | 86.5 | 88 | 88.425 | 88.2 | 88.05
6.6 | 86.05 | 87.6 | 87.975 | 87.95 | 87.7333
6.7 | 85.8 | 87.1 | 87.625 | 87.5 | 87.475
6.8 | 85.35 | 86.7333 | 87.3 | 87.18 | 87.1
6.9 | 84.825 | 86.5 | 87.075 | 86.9667 | 86.75
7 | 84.575 | 86.3 | 86.7 | 86.5833 | 86.525
7.1 | 84.25 | 85.8667 | 86.275 | 86.3125 | 86.175
7.2 | 84 | 85.6 | 86.04 | 86.0571 | 85.925
7.3 | 83.55 | 85.15 | 85.75 | 85.82 | 85.75
7.4 | 83.1 | 84.8 | 85.5 | 85.4375 | 85.325
7.5 | 82.7 | 84.425 | 85.2 | 85.1 | 85.1
7.6 | 82.425 | 84.1333 | 85 | 84.75 | 84.75
7.7 | 82.05 | 83.8 | 84.7 | 84.5 | 84.45
7.8 | 81.7 | 83.46 | 84.2333 | 84.1833 | 84.1667
7.9 | 81.375 | 83.25 | 83.95 | 83.975 | 83.825
8 | 80.9 | 82.9 | 83.8 | 83.725 | 83.58
8.1 | 80.46 | 82.55 | 83.4 | 83.475 | 83.35
8.2 | 80.15 | 82.3 | 83.05 | 83.2 | 83
8.3 | 79.875 | 82.075 | 82.75 | 82.8333 | 82.75
8.4 | 79.55 | 81.825 | 82.5 | 82.5 | 82.5
8.5 | 79.1 | 81.425 | 82.175 | 82.25 | 82.175
8.6 | 78.95 | 80.925 | 81.9 | 82 | 81.9
8.7 | 78.5 | 80.675 | 81.55 | 81.825 | 81.725
8.8 | 78 | 80.5 | 81.3 | 81.5 | 81.35
8.9 | 77.75 | 80.08 | 80.95 | 81.1 | 81.1
9 | 77.28 | 79.75 | 80.625 | 80.8 | 80.675
9.1 | 76.9 | 79.5 | 80.4 | 80.55 | 80.425
9.2 | 76.7 | 79.225 | 80 | 80.2 | 80.175
9.3 | 76.4 | 78.9 | 79.75 | 80 | 79.75
9.4 | 76 | 78.6 | 79.5 | 79.65 | 79.5
9.5 | 75.7333 | 78.3 | 79.1 | 79.5 | 79.25
9.6 | 75.3 | 77.9 | 78.825 | 79.175 | 79
9.7 | 74.975 | 77.6 | 78.575 | 79 | 78.7667
9.8 | 74.625 | 77.35 | 78.3 | 78.66 | 78.4
9.9 | 74.2333 | 76.96 | 78.1 | 78.3 | 78.2
10 | 73.9667 | 76.6 | 77.8 | 77.94 | 78
10.1 | 73.6 | 76.4 | 77.5 | 77.7 | 77.75
10.2 | 73.1667 | 76.05 | 77.2 | 77.5 | 77.325
10.3 | 72.7 | 75.725 | 76.8667 | 77.1 | 77.075
10.4 | 72.5 | 75.375 | 76.575 | 76.8333 | 76.825
10.5 | 72.26 | 75.125 | 76.325 | 76.575 | 76.5125
10.6 | 71.98 | 74.775 | 75.975 | 76.35 | 76.2444
10.7 | 71.45 | 74.525 | 75.75 | 76.15 | 75.8857
10.8 | 71.2 | 74.18 | 75.575 | 75.825 | 75.775
10.9 | 70.8875 | 73.85 | 75.3 | 75.6333 | 75.6
11 | 70.575 | 73.6 | 75.025 | 75.375 | 75.225
11.1 | 70.2667 | 73.3 | 74.7 | 75.04 | 74.95
11.2 | 70 | 73.075 | 74.45 | 74.7 | 74.675
11.3 | 69.6333 | 72.7 | 74.2 | 74.5 | 74.4
11.4 | 69.1333 | 72.4 | 73.86 | 74.15 | 74.1429
11.5 | 68.9 | 72.05 | 73.525 | 73.95 | 73.875
11.6 | 68.6667 | 71.9 | 73.2714 | 73.7333 | 73.6667
11.7 | 68.3 | 71.55 | 73.075 | 73.5 | 73.3333
11.8 | 68.1 | 71.325 | 72.775 | 73.225 | 73.1
11.9 | 67.775 | 70.975 | 72.5 | 72.9 | 72.9
12 | 67.45 | 70.675 | 72.15 | 72.625 | 72.55
12.1 | 67.075 | 70.45 | 71.85 | 72.425 | 72.4
12.2 | 66.8 | 70.15 | 71.6 | 72.1333 | 72.15
12.3 | 66.45 | 69.75 | 71.375 | 71.775 | 71.75
12.4 | 66.2 | 69.475 | 71.2 | 71.5 | 71.475
12.5 | 65.96 | 69.225 | 70.95 | 71.35 | 71.225
12.6 | 65.6 | 68.9857 | 70.7222 | 71.12 | 71
12.7 | 65.35 | 68.74 | 70.425 | 70.7 | 70.8
12.8 | 65.05 | 68.5 | 70.2 | 70.5 | 70.5667
12.9 | 64.6333 | 68.1333 | 69.95 | 70.32 | 70.2
13 | 64.375 | 67.8 | 69.6 | 70.1667 | 70.075
13.1 | 64.125 | 67.6 | 69.35 | 69.875 | 69.825
13.2 | 63.775 | 67.3 | 69.1 | 69.6 | 69.4833
13.3 | 63.525 | 67 | 68.86 | 69.3 | 69.275
13.4 | 63.26 | 66.7 | 68.625 | 69.1 | 69.1125
13.5 | 62.9 | 66.5 | 68.375 | 68.8 | 68.78
13.6 | 62.7 | 66.2 | 68.0375 | 68.5333 | 68.475
13.7 | 62.35 | 65.95 | 67.775 | 68.2 | 68.16
13.8 | 62 | 65.8 | 67.5625 | 68.075 | 68
13.9 | 61.8 | 65.45 | 67.375 | 67.825 | 67.7
14 | 61.575 | 65.125 | 67.0667 | 67.5 | 67.56
14.1 | 61.3333 | 64.875 | 66.7857 | 67.25 | 67.325
14.2 | 61.05 | 64.5667 | 66.45 | 66.9 | 67
14.3 | 60.8 | 64.2 | 66.3 | 66.8 | 66.78
14.4 | 60.35 | 63.94 | 65.9571 | 66.56 | 66.58
14.5 | 60.125 | 63.7 | 65.8 | 66.1 | 66.3
14.6 | 59.875 | 63.42 | 65.4625 | 65.9667 | 66.1
14.7 | 59.7 | 63.2 | 65.125 | 65.72 | 65.75
14.8 | 59.45 | 63.05 | 64.875 | 65.475 | 65.5
14.9 | 59.05 | 62.8 | 64.75 | 65.325 | 65.3
15 | 58.7 | 62.5 | 64.5 | 65.1 | 65.12
15.1 | 58.45 | 62.3 | 64.25 | 64.85 | 64.875
15.2 | 58.2667 | 61.9667 | 64 | 64.575 | 64.6
15.3 | 58.1 | 61.7 | 63.75 | 64.325 | 64.4
15.4 | 57.65 | 61.375 | 63.6 | 64 | 64.1
15.5 | 57.4 | 61.125 | 63.3 | 63.825 | 63.9
15.6 | 57.15 | 60.975 | 63.025 | 63.625 | 63.65
15.7 | 57 | 60.65 | 62.8667 | 63.5 | 63.4
15.8 | 56.6 | 60.3 | 62.7 | 63.3 | 63.26
15.9 | 56.3 | 60.175 | 62.35 | 63.1 | 62.95
16 | 56.05 | 59.925 | 62.05 | 62.75 | 62.675
16.1 | 55.9 | 59.7375 | 61.9 | 62.5 | 62.5
16.2 | 55.675 | 59.475 | 61.65 | 62.25 | 62.375
16.3 | 55.425 | 59.175 | 61.4 | 62.1 | 62.125
16.4 | 55.1286 | 58.925 | 61.15 | 61.85 | 61.9
16.5 | 54.9 | 58.575 | 60.925 | 61.7 | 61.65
16.6 | 54.675 | 58.425 | 60.62 | 61.3333 | 61.5
16.7 | 54.4 | 58.15 | 60.35 | 61.1 | 61.25
16.8 | 54.1833 | 57.925 | 60.1429 | 61 | 61.05
16.9 | 53.875 | 57.675 | 59.96 | 60.76 | 60.8
17 | 53.675 | 57.425 | 59.725 | 60.525 | 60.6
17.1 | 53.44 | 57.175 | 59.5625 | 60.175 | 60.32
17.2 | 53.1857 | 56.925 | 59.225 | 60 | 60.15
17.3 | 52.975 | 56.775 | 59.0375 | 59.8 | 59.825
17.4 | 52.75 | 56.44 | 58.825 | 59.6 | 59.6
17.5 | 52.3 | 56.34 | 58.575 | 59.375 | 59.4
17.6 | 52 | 56.1 | 58.325 | 59.125 | 59.225
17.7 | 51.8 | 55.85 | 58.08 | 58.95 | 59
17.8 | 51.625 | 55.625 | 57.95 | 58.7 | 58.8
17.9 | 51.375 | 55.275 | 57.7 | 58.5 | 58.5333
18 | 51.125 | 55.125 | 57.55 | 58.35 | 58.3
18.1 | 50.88 | 54.875 | 57.3 | 58.05 | 58.125
18.2 | 50.65 | 54.625 | 57.05 | 57.8 | 57.9667
18.3 | 50.5 | 54.4 | 56.8 | 57.54 | 57.75
18.4 | 50.15 | 54.25 | 56.6333 | 57.3 | 57.425
18.5 | 49.9 | 54.0333 | 56.45 | 57.15 | 57.3
18.6 | 49.8 | 53.7 | 56.2333 | 56.9 | 57.05
18.7 | 49.6 | 53.55 | 56.025 | 56.75 | 56.9
18.8 | 49.35 | 53.3 | 55.9 | 56.5 | 56.65
18.9 | 49.025 | 52.95 | 55.55 | 56.275 | 56.4
19 | 48.875 | 52.7556 | 55.3429 | 56.125 | 56.25
19.1 | 48.625 | 52.5571 | 55.125 | 55.9667 | 56
19.2 | 48.5 | 52.3222 | 54.9 | 55.725 | 55.82
19.3 | 48.2 | 52.1 | 54.7 | 55.48 | 55.625
19.4 | 47.95 | 51.94 | 54.525 | 55.3 | 55.4667
19.5 | 47.75 | 51.6 | 54.3 | 55.0667 | 55.125
19.6 | 47.62 | 51.45 | 54.1 | 54.8667 | 54.9667
19.7 | 47.4 | 51.2 | 53.85 | 54.625 | 54.7333
19.8 | 47.2 | 51.05 | 53.6 | 54.4 | 54.5
19.9 | 47 | 50.8 | 53.475 | 54.25 | 54.36
20 | 46.75 | 50.575 | 53.325 | 54 | 54.125
20.1 | 46.525 | 50.425 | 53.075 | 53.725 | 53.8333
20.2 | 46.2 | 50.1 | 52.8 | 53.62 | 53.7
20.3 | 46.05 | 49.925 | 52.6 | 53.4 | 53.55
20.4 | 45.825 | 49.8 | 52.325 | 53.25 | 53.3
20.5 | 45.6 | 49.525 | 52.1 | 53.05 | 53.2
20.6 | 45.35 | 49.35 | 52 | 52.875 | 52.9667
20.7 | 45.175 | 49.2 | 51.9 | 52.6 | 52.7
20.8 | 45.025 | 48.96 | 51.65 | 52.4 | 52.56
20.9 | 44.8375 | 48.8 | 51.4 | 52.225 | 52.4125
21 | 44.625 | 48.575 | 51.25 | 52 | 52.25
21.1 | 44.3875 | 48.4 | 51.1 | 51.9 | 52.14
21.2 | 44.225 | 48.125 | 50.75 | 51.675 | 51.9
21.3 | 44.0444 | 48 | 50.6 | 51.5 | 51.65
21.4 | 43.85 | 47.86 | 50.4 | 51.375 | 51.45
21.5 | 43.6 | 47.6444 | 50.225 | 51.08 | 51.25
21.6 | 43.45 | 47.5 | 50 | 50.9 | 51
21.7 | 43.2 | 47.25 | 49.85 | 50.75 | 50.875
21.8 | 43.025 | 47 | 49.6 | 50.5 | 50.625
21.9 | 42.8 | 46.825 | 49.45 | 50.3 | 50.375
22 | 42.65 | 46.6 | 49.2 | 50.1 | 50.225
22.1 | 42.5 | 46.2857 | 49.0333 | 50 | 50.075
22.2 | 42.15 | 46.1429 | 48.9 | 49.825 | 49.925
22.3 | 42 | 46.025 | 48.65 | 49.575 | 49.775
22.4 | 41.9 | 45.775 | 48.4 | 49.425 | 49.58
22.5 | 41.8 | 45.65 | 48.25 | 49.175 | 49.4
22.6 | 41.5 | 45.5 | 48.1 | 48.94 | 49.15
22.7 | 41.375 | 45.3 | 47.95 | 48.84 | 49
22.8 | 41.2 | 45.1 | 47.75 | 48.7 | 48.8
22.9 | 41 | 44.875 | 47.6 | 48.45 | 48.625
23 | 40.825 | 44.725 | 47.4 | 48.2333 | 48.475
23.1 | 40.6 | 44.6 | 47.22 | 48.1 | 48.325
23.2 | 40.525 | 44.325 | 47.075 | 47.95 | 48.0667
23.3 | 40.3 | 44.1 | 46.8 | 47.82 | 47.9
23.4 | 40.15 | 43.925 | 46.7 | 47.675 | 47.7
23.5 | 39.9 | 43.8 | 46.575 | 47.34 | 47.525
23.6 | 39.75 | 43.55 | 46.35 | 47.1571 | 47.375
23.7 | 39.6 | 43.4 | 46.2 | 47 | 47.225
23.8 | 39.375 | 43.25 | 46 | 46.825 | 46.975
23.9 | 39.2 | 43 | 45.8 | 46.5875 | 46.8
24 | 39 | 42.8 | 45.575 | 46.44 | 46.6
24.1 | 38.8 | 42.6 | 45.4 | 46.3 | 46.5
24.2 | 38.7 | 42.4875 | 45.15 | 46.16 | 46.275
24.3 | 38.56 | 42.325 | 44.9571 | 45.925 | 46.1
24.4 | 38.2 | 42.0333 | 44.875 | 45.875 | 45.9
24.5 | 38.0429 | 41.925 | 44.7 | 45.7 | 45.8
24.6 | 37.925 | 41.7875 | 44.5 | 45.575 | 45.7
24.7 | 37.7 | 41.625 | 44.325 | 45.35 | 45.45
24.8 | 37.525 | 41.5 | 44.1 | 45.2 | 45.3
24.9 | 37.4625 | 41.28 | 43.9 | 44.95 | 45.1
25 | 37.275 | 41.08 | 43.8 | 44.8 | 44.9333
25.1 | 37.1125 | 40.85 | 43.6667 | 44.6 | 44.775
25.2 | 36.9875 | 40.7111 | 43.4 | 44.5 | 44.6
25.3 | 36.825 | 40.575 | 43.275 | 44.3 | 44.4
25.4 | 36.6444 | 40.3625 | 43.1 | 44.1 | 44.3
25.5 | 36.5 | 40.2375 | 42.975 | 43.95 | 44.1
25.6 | 36.25 | 40.025 | 42.825 | 43.8 | 43.925
25.7 | 36.1 | 39.875 | 42.7 | 43.65 | 43.8
25.8 | 35.94 | 39.6625 | 42.5 | 43.5 | 43.625
25.9 | 35.775 | 39.5 | 42.35 | 43.2667 | 43.475
26 | 35.625 | 39.45 | 42.15 | 43.1 | 43.325
26.1 | 35.375 | 39.2 | 41.95 | 42.925 | 43.1
26.2 | 35.2 | 39.075 | 41.8 | 42.7111 | 42.95
26.3 | 35.1 | 38.9 | 41.65 | 42.6 | 42.8
26.4 | 34.95 | 38.775 | 41.5 | 42.4222 | 42.6
26.5 | 34.8 | 38.55 | 41.3 | 42.275 | 42.5
26.6 | 34.6 | 38.3 | 41.2 | 42.0571 | 42.35
26.7 | 34.4 | 38.2 | 41.1 | 41.95 | 42.1
26.8 | 34.3 | 38 | 40.9 | 41.8 | 41.975
26.9 | 34.2 | 37.9 | 40.775 | 41.7 | 41.9
27 | 34.1 | 37.75 | 40.525 | 41.55 | 41.7
27.1 | 33.975 | 37.55 | 40.275 | 41.3333 | 41.525
27.2 | 33.725 | 37.35 | 40.2 | 41.2 | 41.375
27.3 | 33.54 | 37.3 | 40.1 | 41.05 | 41.2
27.4 | 33.4 | 37.1 | 39.975 | 40.9 | 41.0333
27.5 | 33.3 | 36.925 | 39.8 | 40.76 | 40.925
27.6 | 33.175 | 36.775 | 39.6333 | 40.525 | 40.8
27.7 | 33 | 36.6 | 39.425 | 40.375 | 40.6333
27.8 | 32.9 | 36.45 | 39.2833 | 40.225 | 40.4667
27.9 | 32.8111 | 36.28 | 39.1 | 40.1 | 40.225
28 | 32.66 | 36.1 | 38.925 | 40 | 40.1
28.1 | 32.525 | 35.975 | 38.8 | 39.8 | 40
28.2 | 32.2833 | 35.8556 | 38.625 | 39.7 | 39.9
28.3 | 32.2 | 35.7375 | 38.475 | 39.5 | 39.7
28.4 | 32.1 | 35.575 | 38.325 | 39.36 | 39.5333
28.5 | 32 | 35.425 | 38.16 | 39.225 | 39.4
28.6 | 31.825 | 35.375 | 38.025 | 39 | 39.3
28.7 | 31.6 | 35.225 | 37.975 | 38.9 | 39.075
28.8 | 31.525 | 34.98 | 37.8 | 38.8 | 38.9111
28.9 | 31.4 | 34.85 | 37.6 | 38.7 | 38.8
29 | 31.225 | 34.7 | 37.44 | 38.54 | 38.7125
29.1 | 31.08 | 34.5 | 37.4 | 38.325 | 38.5
29.2 | 31 | 34.4 | 37.2 | 38.1375 | 38.4
29.3 | 30.75 | 34.3 | 37.1 | 38.0222 | 38.2
29.4 | 30.6 | 34.15 | 36.9 | 37.9 | 38.1
29.5 | 30.46 | 34 | 36.8 | 37.7333 | 37.95
29.6 | 30.3 | 33.86 | 36.625 | 37.6 | 37.8
29.7 | 30.25 | 33.725 | 36.4 | 37.45 | 37.625
29.8 | 30.1 | 33.58 | 36.3 | 37.3 | 37.5
29.9 | 30 | 33.4 | 36.1 | 37.2 | 37.3
30 | 29.775 | 33.2 | 36 | 37 | 37.2

## Wedge 30 Lateral Profiles
Lateral Profiles
Wedge 30
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:2.4 cm FieldSize:4x4 cm2 | Depth:2.4 cm FieldSize:10x10 cm2 | Depth:2.4 cm FieldSize:20x20 cm2 | Depth:2.4 cm FieldSize:30x30 cm2 | Depth:2.4 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.9 | [18 empty] | 13.4667 | 14.7286 |  |  |  | 19.4 | 20.72
-22.8 | [18 empty] | 13.65 | 14.9667 |  |  |  | 19.6667 | 20.9333
-22.7 | [18 empty] | 13.9167 | 15.1429 |  |  |  | 19.92 | 21.24
-22.6 | [18 empty] | 14.1 | 15.425 |  |  |  | 20.2333 | 21.6
-22.5 | [18 empty] | 14.28 | 15.6429 |  |  |  | 20.475 | 21.9
-22.4 | [18 empty] | 14.5286 | 15.7857 |  |  |  | 20.82 | 22.14
-22.3 | [18 empty] | 14.7429 | 15.98 |  |  |  | 21.0714 | 22.44
-22.2 | [18 empty] | 14.95 | 16.1667 |  |  |  | 21.4333 | 22.75
-22.1 | [18 empty] | 15.15 | 16.45 |  |  |  | 21.8 | 23.1
-22 | [18 empty] | 15.4 | 16.7125 |  |  |  | 22.1 | 23.35
-21.9 | [18 empty] | 15.7 | 16.92 |  |  |  | 22.3667 | 23.8
-21.8 | [18 empty] | 15.9286 | 17.2 |  |  |  | 22.7 | 24.0667
-21.7 | [18 empty] | 16.18 | 17.4857 |  |  |  | 23.1 | 24.4
-21.6 | [18 empty] | 16.4 | 17.7714 |  |  |  | 23.4667 | 24.74
-21.5 | [18 empty] | 16.7333 | 18.025 |  |  |  | 23.9 | 25.1
-21.4 | [13 empty] | 10.78 | 11.8976 |  |  |  | 16.9833 | 18.2333 |  |  |  | 24.18 | 25.54
-21.3 | [13 empty] | 10.92 | 11.9976 |  |  |  | 17.2857 | 18.5 |  |  |  | 24.58 | 26
-21.2 | [13 empty] | 11.05 | 12.1976 |  |  |  | 17.5714 | 18.8 |  |  |  | 24.98 | 26.35
-21.1 | [13 empty] | 11.28 | 12.4225 |  |  |  | 17.82 | 19.12 |  |  |  | 25.5 | 26.84
-21 | [13 empty] | 11.4 | 12.5575 |  |  |  | 18.1333 | 19.4286 |  |  |  | 25.94 | 27.25
-20.9 | [13 empty] | 11.575 | 12.7724 |  |  |  | 18.4667 | 19.7143 |  |  |  | 26.5 | 27.7
-20.8 | [13 empty] | 11.8 | 12.8974 |  |  |  | 18.76 | 20.0167 |  |  |  | 27.1 | 28.3
-20.7 | [8 empty] | 9.15 | 10.2 |  |  |  | 11.98 | 13.1574 |  |  |  | 19.0667 | 20.2875 |  |  |  | 27.56 | 28.925
-20.6 | [8 empty] | 9.3 | 10.4 |  |  |  | 12.2 | 13.364 |  |  |  | 19.4 | 20.6429 |  |  |  | 28.4 | 29.7
-20.5 | [8 empty] | 9.4 | 10.52 |  |  |  | 12.375 | 13.5473 |  |  |  | 19.7286 | 20.98 |  |  |  | 29.2 | 30.525
-20.4 | [8 empty] | 9.6 | 10.65 |  |  |  | 12.6 | 13.7972 |  |  |  | 20.0714 | 21.32 |  |  |  | 30.2 | 31.4
-20.3 |  |  |  | 9.2 | 10.22 |  |  |  | 9.725 | 10.8 |  |  |  | 12.9 | 13.9972 |  |  |  | 20.4 | 21.6667 |  |  |  | 31.3 | 32.6833
-20.2 |  |  |  | 9.3 | 10.44 |  |  |  | 9.9 | 11 |  |  |  | 13 | 14.2222 |  |  |  | 20.75 | 22 |  |  |  | 32.8 | 34.16
-20.1 |  |  |  | 9.5 | 10.56 |  |  |  | 10.025 | 11.175 |  |  |  | 13.275 | 14.4721 |  |  |  | 21.14 | 22.3857 |  |  |  | 35.06 | 36.0667
-20 |  |  |  | 9.6 | 10.7 |  |  |  | 10.275 | 11.3 |  |  |  | 13.46 | 14.6721 |  |  |  | 21.45 | 22.7333 |  |  |  | 37.6333 | 38.86
-19.9 |  |  |  | 9.725 | 10.84 |  |  |  | 10.4 | 11.5 |  |  |  | 13.7 | 14.957 |  |  |  | 21.9 | 23.1 |  |  |  | 41.7 | 42.7333
-19.8 |  |  |  | 9.9 | 11 |  |  |  | 10.575 | 11.68 |  |  |  | 13.925 | 15.172 |  |  |  | 22.2 | 23.5143 |  |  |  | 46.9 | 47.875
-19.7 |  |  |  | 10.025 | 11.14 |  |  |  | 10.7 | 11.8 |  |  |  | 14.25 | 15.4469 |  |  |  | 22.7 | 23.9 |  |  |  | 53.64 | 54.65
-19.6 |  |  |  | 10.175 | 11.3 |  |  |  | 10.9 | 12 |  |  |  | 14.525 | 15.6969 |  |  |  | 23.1286 | 24.3571 |  |  |  | 61.84 | 62.78
-19.5 |  |  |  | 10.325 | 11.425 |  |  |  | 11.1 | 12.225 |  |  |  | 14.775 | 15.9718 |  |  |  | 23.6 | 24.84 |  |  |  | 71.25 | 71.35
-19.4 |  |  |  | 10.5 | 11.7 |  |  |  | 11.3 | 12.4 |  |  |  | 14.9833 | 16.1968 |  |  |  | 24.1429 | 25.3667 |  |  |  | 79.8 | 80.6
-19.3 |  |  |  | 10.625 | 11.86 |  |  |  | 11.5 | 12.625 |  |  |  | 15.3 | 16.53 |  |  |  | 24.75 | 25.95 |  |  |  | 88.2 | 88.6667
-19.2 |  |  |  | 10.86 | 12.025 |  |  |  | 11.7 | 12.84 |  |  |  | 15.625 | 16.8166 |  |  |  | 25.4 | 26.58 |  |  |  | 95.65 | 95.24
-19.1 |  |  |  | 11.025 | 12.175 |  |  |  | 11.9 | 13.0667 |  |  |  | 15.86 | 17.0166 |  |  |  | 26.125 | 27.34 |  |  |  | 101.325 | 100.775
-19 |  |  |  | 11.2 | 12.35 |  |  |  | 12.1 | 13.25 |  |  |  | 16.22 | 17.3465 |  |  |  | 27 | 28.275 |  |  |  | 105.4 | 104.58
-18.9 |  |  |  | 11.4 | 12.5 |  |  |  | 12.35 | 13.5 |  |  |  | 16.54 | 17.7964 |  |  |  | 28.05 | 29.2571 |  |  |  | 108.08 | 107.25
-18.8 |  |  |  | 11.54 | 12.72 |  |  |  | 12.6 | 13.6667 |  |  |  | 16.875 | 17.9564 |  |  |  | 29.3571 | 30.4 |  |  |  | 109.94 | 109.08
-18.7 |  |  |  | 11.775 | 12.9 |  |  |  | 12.8 | 13.925 |  |  |  | 17.15 | 18.3163 |  |  |  | 30.95 | 31.95 |  |  |  | 111.425 | 110.5
-18.6 |  |  |  | 11.95 | 13.225 |  |  |  | 12.95 | 14.175 |  |  |  | 17.48 | 18.6763 |  |  |  | 32.94 | 34.1 |  |  |  | 112.4 | 111.5
-18.5 |  |  |  | 12.2 | 13.375 |  |  |  | 13.3 | 14.4 |  |  |  | 17.85 | 19.0712 |  |  |  | 36.0667 | 37.075 |  |  |  | 113.2 | 112.4
-18.4 |  |  |  | 12.4 | 13.6 |  |  |  | 13.5 | 14.65 |  |  |  | 18.2 | 19.4961 |  |  |  | 40.25 | 41.1714 |  |  |  | 113.9333 | 113.1
-18.3 |  |  |  | 12.575 | 13.75 |  |  |  | 13.8 | 14.95 |  |  |  | 18.6 | 19.796 |  |  |  | 46.5833 | 46.94 |  |  |  | 114.64 | 113.6333
-18.2 |  |  |  | 12.84 | 13.98 |  |  |  | 14.025 | 15.2 |  |  |  | 19 | 20.196 |  |  |  | 54.4 | 55.3 |  |  |  | 115.0667 | 114.0333
-18.1 |  |  |  | 13.04 | 14.22 |  |  |  | 14.25 | 15.475 |  |  |  | 19.54 | 20.6459 |  |  |  | 64.6143 | 65.3 |  |  |  | 115.48 | 114.4
-18 |  |  |  | 13.28 | 14.48 |  |  |  | 14.6 | 15.74 |  |  |  | 19.96 | 21.0958 |  |  |  | 75.1286 | 75.1143 |  |  |  | 115.75 | 114.7667
-17.9 |  |  |  | 13.525 | 14.7 |  |  |  | 14.85 | 16.05 |  |  |  | 20.45 | 21.6957 |  |  |  | 85.8714 | 85.625 |  |  | 11.175 | 116.1 | 115.16
-17.8 |  |  |  | 13.775 | 14.925 |  |  |  | 15.18 | 16.4 |  |  |  | 21.125 | 22.2955 |  |  |  | 96 | 94.7571 |  |  | 11.3667 | 116.3 | 115.45
-17.7 |  |  |  | 14.04 | 15.3 |  |  |  | 15.4833 | 16.6 |  |  |  | 21.7 | 22.8354 |  |  |  | 104.2 | 102.8 |  |  | 11.56 | 116.5667 | 115.6167
-17.6 |  |  |  | 14.35 | 15.4667 |  |  |  | 15.85 | 17 |  |  |  | 22.425 | 23.6953 |  |  |  | 110.0286 | 109.2 |  |  | 11.7333 | 116.75 | 115.7833
-17.5 |  |  |  | 14.575 | 15.74 |  |  |  | 16.2 | 17.28 |  |  |  | 23.375 | 24.5701 |  |  |  | 114.3 | 113.4 |  |  | 11.9 | 116.875 | 116
-17.4 |  |  |  | 14.85 | 16.08 |  |  |  | 16.55 | 17.725 |  |  |  | 24.56 | 25.5949 |  |  |  | 117.1 | 116.1 |  |  | 12.0667 | 117.1 | 116.1667
-17.3 |  |  |  | 15.125 | 16.375 |  |  |  | 16.95 | 18.075 |  |  |  | 25.88 | 26.9546 |  |  |  | 118.9 | 117.9 |  |  | 12.28 | 117.2 | 116.3
-17.2 |  |  |  | 15.475 | 16.65 |  |  |  | 17.4 | 18.58 |  |  |  | 27.625 | 28.7443 |  |  |  | 120.1667 | 119.15 |  |  | 12.5 | 117.34 | 116.46
-17.1 |  |  |  | 15.7333 | 16.9 |  |  |  | 17.9 | 19.1 |  |  |  | 29.95 | 30.9439 |  |  |  | 121.2 | 120.0857 |  |  | 12.6833 | 117.46 | 116.6
-17 |  |  |  | 16.05 | 17.2333 |  |  |  | 18.52 | 19.56 |  |  |  | 33.5 | 34.5931 |  |  |  | 122 | 120.8667 |  |  | 12.96 | 117.575 | 116.7
-16.9 |  |  |  | 16.475 | 17.6 |  |  |  | 19.08 | 20.175 |  |  |  | 38.6 | 39.2921 |  |  | 9.4 | 122.6143 | 121.55 |  |  | 13.1 | 117.72 | 116.8
-16.8 |  |  |  | 16.84 | 18.05 |  |  |  | 19.9 | 20.94 |  |  |  | 45.975 | 46.8156 |  |  | 9.6 | 123.1333 | 122.05 |  |  | 13.3667 | 117.8167 | 116.9
-16.7 |  |  |  | 17.2667 | 18.45 |  |  |  | 20.86 | 21.9 |  |  |  | 57.2 | 57.0386 |  |  | 9.72 | 123.5143 | 122.3333 |  |  | 13.5667 | 117.9 | 116.9
-16.6 |  |  |  | 17.76 | 18.9 |  |  |  | 22.1 | 23.1 |  |  |  | 68.86 | 69.5611 |  |  | 9.8 | 123.86 | 122.6667 |  |  | 13.8 | 117.9 | 117
-16.5 |  |  |  | 18.35 | 19.525 |  |  |  | 23.575 | 24.58 |  |  |  | 81.92 | 81.3837 |  |  | 10 | 124.0667 | 122.96 |  |  | 14.1 | 118 | 117.1
-16.4 |  |  |  | 19.1 | 20.25 |  |  |  | 25.6 | 26.75 |  |  |  | 95.1667 | 94.2311 |  |  | 10.2 | 124.2857 | 123.1667 |  |  | 14.3 | 118.1 | 117.2
-16.3 |  |  |  | 20 | 21.1 |  |  |  | 28.78 | 29.65 |  |  |  | 106.5 | 106.2592 |  |  | 10.3 | 124.4571 | 123.3143 |  |  | 14.4667 | 118.175 | 117.2
-16.2 |  |  |  | 21.15 | 22.2 |  |  |  | 33.16 | 34.1 |  |  |  | 116.15 | 115.1773 |  |  | 10.4833 | 124.5875 | 123.4667 |  |  | 14.75 | 118.3 | 117.3
-16.1 |  |  |  | 22.725 | 23.675 |  |  |  | 40.175 | 41.26 |  |  |  | 122.94 | 121.776 |  |  | 10.7 | 124.7 | 123.6333 |  |  | 15.05 | 118.3667 | 117.3
-16 |  |  |  | 24.825 | 25.78 |  |  |  | 51.3 | 51.6 |  |  |  | 127.1 | 125.825 |  |  | 10.92 | 124.7857 | 123.7 |  |  | 15.35 | 118.4 | 117.3
-15.9 |  |  |  | 28 | 29.1 |  |  |  | 63.95 | 64.2 |  |  | 7.2 | 129.5 | 128.574 |  |  | 11.125 | 124.9 | 123.8 |  |  | 15.56 | 118.5 | 117.38
-15.8 |  |  |  | 33.06 | 33.76 |  |  |  | 77.95 | 77.9 |  |  | 7.3833 | 131.2833 | 130.274 |  |  | 11.36 | 124.9286 | 123.9 |  |  | 15.8 | 118.5 | 117.4
-15.7 |  |  |  | 41.45 | 41.6 |  |  |  | 93.175 | 93.1 |  |  | 7.4 | 132.5 | 131.474 |  |  | 11.6 | 125 | 123.9 |  |  | 16.1167 | 118.5 | 117.4
-15.6 |  |  |  | 52.6 | 53.15 |  |  |  | 107.075 | 107.1 |  |  | 7.54 | 133.3 | 132.324 |  |  | 11.78 | 125 | 123.9 |  |  | 16.45 | 118.5 | 117.45
-15.5 |  |  |  | 67.28 | 67.8 |  |  |  | 119.12 | 118.35 |  |  | 7.7 | 134 | 132.853 |  |  | 11.98 | 125 | 123.9 |  |  | 16.66 | 118.5 | 117.5
-15.4 |  |  |  | 82.925 | 82.5333 |  |  | 6.2 | 127.66 | 126.6 |  |  | 7.8 | 134.45 | 133.248 |  |  | 12.18 | 125 | 123.9429 |  |  | 16.98 | 118.5 | 117.5
-15.3 |  |  |  | 98.8 | 97.3 |  |  | 6.3 | 132.95 | 131.95 |  |  | 8 | 134.7 | 133.598 |  |  | 12.4 | 125 | 124 |  |  | 17.325 | 118.5 | 117.52
-15.2 |  |  | 6.24 | 113.175 | 112.15 |  |  | 6.4 | 136.4 | 135.1 |  |  | 8.1 | 134.9 | 133.833 |  |  | 12.6 | 125 | 124 |  |  | 17.575 | 118.5 | 117.6
-15.1 |  |  | 6.3 | 124.85 | 124.2 |  |  | 6.44 | 138.4 | 136.88 |  |  | 8.3 | 135.1 | 133.973 |  |  | 12.9 | 125 | 124 |  |  | 17.9 | 118.5 | 117.6
-15 |  |  | 6.4 | 133.6 | 132.26 |  |  | 6.56 | 139.56 | 138.2 |  |  | 8.42 | 135.1 | 134.023 |  |  | 13.08 | 124.9167 | 124 |  |  | 18.26 | 118.5 | 117.6
-14.9 |  |  | 6.5 | 138.35 | 137.3 |  |  | 6.7 | 140.4 | 139.075 |  |  | 8.6 | 135.2 | 134.073 |  |  | 13.35 | 124.9 | 124 |  |  | 18.6333 | 118.5 | 117.6
-14.8 |  |  | 6.6 | 141.125 | 139.8 |  |  | 6.8 | 140.98 | 139.58 |  |  | 8.725 | 135.2 | 134.073 |  |  | 13.66 | 124.9 | 123.9429 |  |  | 18.9667 | 118.5 | 117.6
-14.7 |  |  | 6.7 | 142.725 | 141.5 |  |  | 6.875 | 141.36 | 140.02 |  |  | 8.875 | 135.2 | 134.023 |  |  | 13.9 | 124.9 | 123.9 |  |  | 19.4 | 118.42 | 117.6
-14.6 |  |  | 6.8 | 143.8 | 142.46 |  |  | 7 | 141.525 | 140.3 |  |  | 9.0167 | 135.1 | 133.973 |  |  | 14.2 | 124.9 | 123.9 |  |  | 19.88 | 118.4 | 117.6
-14.5 |  |  | 6.9 | 144.375 | 143.06 |  |  | 7.15 | 141.6 | 140.35 |  |  | 9.25 | 135 | 133.923 |  |  | 14.45 | 124.8167 | 123.86 |  |  | 20.2 | 118.4 | 117.5
-14.4 |  |  | 7 | 144.7 | 143.34 |  |  | 7.3 | 141.6 | 140.3 |  |  | 9.4333 | 134.8 | 133.793 |  |  | 14.78 | 124.7571 | 123.7667 |  |  | 20.7667 | 118.3 | 117.5
-14.3 |  |  | 7.1 | 144.8 | 143.4 |  |  | 7.4 | 141.5 | 140.3 |  |  | 9.625 | 134.7 | 133.673 |  |  | 15.05 | 124.6857 | 123.7 |  |  | 21.3 | 118.3 | 117.5
-14.2 |  |  | 7.24 | 144.8 | 143.4 |  |  | 7.55 | 141.46 | 140.2 |  |  | 9.84 | 134.55 | 133.5563 |  |  | 15.28 | 124.5429 | 123.7 |  |  | 21.85 | 118.2143 | 117.4
-14.1 |  |  | 7.36 | 144.75 | 143.35 |  |  | 7.66 | 141.24 | 140.02 |  |  | 10.025 | 134.4 | 133.353 |  |  | 15.6 | 124.4125 | 123.6 |  |  | 22.6 | 118.2 | 117.4
-14 |  |  | 7.5 | 144.625 | 143.2 |  |  | 7.8 | 141.025 | 139.875 |  |  | 10.2 | 134.2 | 133.223 |  |  | 15.96 | 124.4 | 123.5 |  |  | 23.32 | 118.2 | 117.3167
-13.9 |  |  | 7.65 | 144.3667 | 143.0333 |  |  | 8 | 140.86 | 139.62 |  |  | 10.3667 | 134.1 | 133.173 |  |  | 16.4 | 124.3143 | 123.4571 |  |  | 24.3 | 118.2 | 117.3
-13.8 |  |  | 7.8 | 144.1 | 142.7667 |  |  | 8.15 | 140.575 | 139.4 |  |  | 10.58 | 133.94 | 132.948 |  |  | 16.7 | 124.3 | 123.3 |  |  | 25.55 | 118.1 | 117.24
-13.7 |  |  | 7.96 | 143.85 | 142.4 |  |  | 8.3 | 140.2 | 139.16 |  |  | 10.925 | 133.7 | 132.798 |  |  | 17.125 | 124.2167 | 123.25 |  |  | 26.9 | 118 | 117.2
-13.6 |  |  | 8.075 | 143.6 | 142.175 |  |  | 8.5 | 139.95 | 138.8333 |  |  | 11.14 | 133.6 | 132.648 |  |  | 17.52 | 124.1 | 123.1833 |  |  | 28.8 | 118 | 117.2
-13.5 |  |  | 8.2 | 143.25 | 141.85 |  |  | 8.6 | 139.7 | 138.5667 |  |  | 11.34 | 133.4 | 132.434 |  |  | 18.04 | 124 | 123.1 |  |  | 31.4 | 117.9 | 117.1833
-13.4 |  |  | 8.4 | 142.8 | 141.55 |  |  | 8.775 | 139.5 | 138.225 |  |  | 11.65 | 133.2 | 132.224 |  |  | 18.46 | 123.85 | 122.9333 |  |  | 35.1833 | 117.8 | 117.1
-13.3 |  |  | 8.55 | 142.38 | 141.125 |  |  | 9 | 139.18 | 138 |  |  | 11.875 | 132.975 | 131.999 |  |  | 19 | 123.7 | 122.8833 |  |  | 39.7667 | 117.8 | 117
-13.2 |  |  | 8.775 | 142.2 | 140.82 |  |  | 9.2 | 138.875 | 137.66 |  |  | 12.1 | 132.75 | 131.854 |  |  | 19.6 | 123.6286 | 122.7286 |  |  | 46.7333 | 117.7 | 116.9333
-13.1 |  |  | 8.9 | 141.75 | 140.45 |  |  | 9.4 | 138.6 | 137.425 |  |  | 12.4 | 132.65 | 131.5407 |  |  | 20.3 | 123.5 | 122.6833 |  |  | 54.4833 | 117.7 | 116.9
-13 |  |  | 9.1 | 141.34 | 140.15 |  |  | 9.54 | 138.35 | 137.15 |  |  | 12.72 | 132.4 | 131.374 |  |  | 21.14 | 123.4 | 122.5286 |  |  | 63.5 | 117.6 | 116.8
-12.9 |  |  | 9.25 | 140.9 | 139.68 |  |  | 9.75 | 138.02 | 136.85 |  |  | 13 | 132.16 | 131.034 |  |  | 22.125 | 123.3 | 122.4 |  |  | 72.55 | 117.55 | 116.7
-12.8 |  |  | 9.5 | 140.65 | 139.325 |  |  | 10.025 | 137.64 | 136.46 |  |  | 13.4 | 131.925 | 130.874 |  |  | 23.36 | 123.2143 | 122.2333 |  |  | 80.4667 | 117.5 | 116.66
-12.7 |  |  | 9.62 | 140.15 | 138.975 |  |  | 10.24 | 137.3 | 136.225 |  |  | 13.8 | 131.55 | 130.674 |  |  | 24.86 | 123.05 | 122.16 |  |  | 87.95 | 117.4 | 116.6
-12.6 |  |  | 9.875 | 139.775 | 138.5 |  |  | 10.4667 | 137.0333 | 135.86 |  |  | 14.15 | 131.26 | 130.454 |  |  | 26.85 | 122.8833 | 122 |  |  | 93.3 | 117.3 | 116.6
-12.5 |  |  | 10 | 139.5 | 138.175 |  |  | 10.7333 | 136.6667 | 135.55 |  |  | 14.52 | 130.95 | 130.224 |  |  | 29.7 | 122.7286 | 121.9 |  |  | 97.2429 | 117.3 | 116.5
-12.4 |  |  | 10.2 | 139.08 | 137.85 |  |  | 11.1 | 136.26 | 135.38 |  |  | 15.1 | 130.825 | 130.074 |  |  | 33.9 | 122.6 | 121.8429 |  |  | 100 | 117.1833 | 116.4
-12.3 |  |  | 10.475 | 138.6 | 137.46 |  |  | 11.3667 | 135.98 | 134.9667 |  |  | 15.55 | 130.58 | 129.814 |  |  | 39.425 | 122.5 | 121.6714 |  |  | 101.85 | 117.1 | 116.3
-12.2 |  |  | 10.725 | 138.18 | 137.12 |  |  | 11.6 | 135.625 | 134.575 |  |  | 16.16 | 130.4 | 129.534 |  |  | 47.3 | 122.32 | 121.5167 |  |  | 103.2 | 117 | 116.2333
-12.1 |  |  | 10.975 | 137.85 | 136.675 |  |  | 11.94 | 135.3 | 134.3 |  |  | 16.86 | 130.15 | 129.224 |  |  | 56.7 | 122.2 | 121.4 |  |  | 104.2 | 116.9 | 116.14
-12 |  |  | 11.25 | 137.375 | 136.35 |  |  | 12.3 | 134.9 | 133.9 |  |  | 17.725 | 129.92 | 129.014 |  |  | 66.3 | 122.0857 | 121.2667 |  |  | 105.02 | 116.8 | 116.1
-11.9 |  |  | 11.5667 | 137.08 | 135.9 |  |  | 12.75 | 134.55 | 133.6 |  |  | 18.62 | 129.6 | 128.814 |  |  | 76.1 | 121.9333 | 121.1143 |  |  | 105.66 | 116.7 | 116
-11.8 |  |  | 11.925 | 136.64 | 135.475 |  |  | 13.28 | 134.3 | 133.35 |  |  | 20 | 129.3 | 128.514 |  |  | 85.6 | 121.7714 | 120.875 |  |  | 106.3 | 116.6 | 115.9
-11.7 |  |  | 12.25 | 136.3 | 135.2 |  |  | 13.8 | 133.875 | 132.98 |  |  | 21.4667 | 129 | 128.249 |  |  | 93.3 | 121.6167 | 120.7333 |  |  | 106.74 | 116.5 | 115.8
-11.6 |  |  | 12.6 | 135.88 | 134.86 |  |  | 14.45 | 133.52 | 132.62 |  |  | 23.7 | 128.7667 | 128.074 |  |  | 99.05 | 121.425 | 120.6625 |  |  | 107.25 | 116.4 | 115.7
-11.5 |  |  | 13.05 | 135.45 | 134.38 |  |  | 15.275 | 133.24 | 132.18 |  |  | 26.6 | 128.5 | 127.7407 |  |  | 102.8 | 121.25 | 120.5375 |  |  | 107.6 | 116.3 | 115.6
-11.4 |  |  | 13.5 | 135.08 | 134.1 |  |  | 16.2 | 132.9 | 131.8667 |  |  | 31.1 | 128.2 | 127.4997 |  |  | 105.175 | 121.1 | 120.325 |  | 5.6 | 107.82 | 116.1 | 115.48
-11.3 |  |  | 14.16 | 134.62 | 133.6667 |  |  | 17.5 | 132.4667 | 131.6 |  |  | 37.54 | 128.0333 | 127.2417 |  |  | 106.9333 | 120.9286 | 120.1571 |  | 5.7 | 107.95 | 116 | 115.4
-11.2 |  |  | 15 | 134.34 | 133.35 |  |  | 19 | 132.15 | 131.275 |  |  | 46.775 | 127.6667 | 126.975 |  |  | 108.35 | 120.8 | 120 |  | 5.75 | 108.175 | 115.92 | 115.3
-11.1 |  |  | 15.925 | 133.95 | 133 |  |  | 21.28 | 131.825 | 130.92 |  |  | 58 | 127.5 | 126.735 |  |  | 109.5667 | 120.6286 | 119.875 |  | 5.9 | 108.32 | 115.8 | 115.2
-11 |  |  | 17.2 | 133.525 | 132.6 |  |  | 24.5 | 131.4 | 130.625 |  |  | 68.95 | 127.15 | 126.4417 |  |  | 110.25 | 120.4667 | 119.7429 |  | 6.05 | 108.52 | 115.7 | 115.1
-10.9 |  |  | 18.9 | 133.15 | 132.225 |  |  | 29.4 | 131.15 | 130.24 |  |  | 80.775 | 126.875 | 126.235 |  | 4.5 | 110.8 | 120.3143 | 119.5833 |  | 6.2 | 108.66 | 115.6 | 114.9667
-10.8 |  |  | 21.4 | 132.675 | 131.75 |  |  | 37.3 | 130.7 | 129.86 |  |  | 90.925 | 126.625 | 125.9 |  | 4.6 | 111.2 | 120.15 | 119.4286 |  | 6.3 | 108.76 | 115.5 | 114.9
-10.7 |  |  | 25 | 132.4 | 131.46 |  |  | 47.35 | 130.45 | 129.5333 |  |  | 99.75 | 126.4 | 125.65 |  | 4.7 | 111.64 | 119.98 | 119.1667 |  | 6.475 | 109 | 115.4667 | 114.8
-10.6 |  |  | 31.1 | 131.94 | 131.06 |  |  | 59.2 | 130.06 | 129.175 |  |  | 106.3 | 126.2 | 125.535 |  | 4.7 | 111.94 | 119.78 | 119.0143 |  | 6.6 | 109.08 | 115.3 | 114.7
-10.5 |  |  | 40.2 | 131.6 | 130.78 |  |  | 72.25 | 129.7 | 128.86 |  |  | 110.26 | 125.9 | 125.2 |  | 4.8333 | 112.14 | 119.5857 | 118.8625 |  | 6.8 | 109.1 | 115.1667 | 114.6
-10.4 |  |  | 52.025 | 131.4 | 130.45 |  |  | 84.5 | 129.3333 | 128.52 |  | 3.2 | 112.775 | 125.64 | 124.875 |  | 4.95 | 112.36 | 119.4429 | 118.7167 |  | 6.8333 | 109.2 | 115 | 114.5
-10.3 |  |  | 64.34 | 130.84 | 129.9 |  |  | 95.6 | 129.12 | 128.225 |  | 3.2333 | 114.5 | 125.25 | 124.615 |  | 5.1 | 112.56 | 119.225 | 118.6 |  | 7.075 | 109.2833 | 114.9 | 114.3833
-10.2 |  |  | 78.1 | 130.35 | 129.55 |  | 2.6 | 105.2 | 128.675 | 127.875 |  | 3.3 | 115.65 | 125.02 | 124.325 |  | 5.225 | 112.66 | 119.08 | 118.4 |  | 7.2 | 109.3 | 114.7667 | 114.2833
-10.1 |  | 2.7 | 91.8 | 130.125 | 129.25 |  | 2.6333 | 111.35 | 128.35 | 127.525 |  | 3.4 | 116.6 | 124.8 | 124.1 |  | 5.3333 | 112.775 | 118.8286 | 118.2333 |  | 7.4 | 109.4 | 114.675 | 114.1
-10 |  | 2.8 | 102.66 | 129.76 | 128.85 |  | 2.7 | 115.3 | 128.1 | 127.2 |  | 3.5 | 117.2 | 124.475 | 123.895 |  | 5.4 | 112.9 | 118.6 | 117.9714 |  | 7.55 | 109.4 | 114.5333 | 114.05
-9.9 |  | 2.8 | 110.88 | 129.32 | 128.575 |  | 2.7 | 117.7 | 127.65 | 126.92 |  | 3.5333 | 117.72 | 124.225 | 123.615 |  | 5.575 | 112.96 | 118.4333 | 117.8167 |  | 7.6667 | 109.5 | 114.425 | 113.925
-9.8 |  | 2.8 | 116.4667 | 128.88 | 128.25 |  | 2.8 | 119.25 | 127.4 | 126.6 |  | 3.6 | 118.02 | 123.875 | 123.275 |  | 5.7 | 113 | 118.25 | 117.6857 |  | 7.8 | 109.5 | 114.3 | 113.775
-9.7 |  | 2.9 | 119.725 | 128.62 | 127.84 |  | 2.9 | 120.4 | 126.975 | 126.25 |  | 3.7 | 118.32 | 123.56 | 122.875 |  | 5.8 | 113.0667 | 118.1 | 117.4857 |  | 8 | 109.55 | 114.2 | 113.6
-9.6 |  | 2.9667 | 121.6 | 128.2 | 127.5 |  | 2.9 | 121.2 | 126.525 | 125.94 |  | 3.8 | 118.525 | 123.34 | 122.675 |  | 5.9667 | 113.1 | 117.95 | 117.3 |  | 8.2 | 109.6 | 114.2 | 113.6
-9.5 |  | 3 | 122.76 | 127.825 | 127.05 |  | 3 | 121.6333 | 126.3 | 125.4 |  | 3.8 | 118.6 | 123.1 | 122.426 |  | 6.125 | 113.1 | 117.6833 | 117.1286 |  | 8.3667 | 109.6 | 114.0333 | 113.48
-9.4 |  | 3 | 123.42 | 127.56 | 126.7 |  | 3 | 122 | 125.95 | 125.18 |  | 4 | 118.7 | 122.8 | 122.1427 |  | 6.22 | 113.1 | 117.6 | 116.95 |  | 8.6 | 109.6 | 113.92 | 113.3
-9.3 |  | 3.1 | 124 | 127.225 | 126.38 |  | 3.1 | 122.1667 | 125.62 | 124.85 |  | 4.1 | 118.7 | 122.5 | 121.801 |  | 6.4 | 113.1 | 117.4 | 116.78 |  | 8.8 | 109.6 | 113.7 | 113.1833
-9.2 |  | 3.2 | 124.2667 | 126.875 | 125.9667 |  | 3.125 | 122.2 | 125.18 | 124.6 |  | 4.2 | 118.7 | 122.275 | 121.676 |  | 6.6 | 113 | 117.2 | 116.6 |  | 9.0333 | 109.6 | 113.6 | 113.1
-9.1 |  | 3.2 | 124.34 | 126.45 | 125.7 |  | 3.2 | 122.3 | 124.95 | 124.275 |  | 4.3 | 118.7 | 121.98 | 121.4427 |  | 6.8 | 113 | 117.06 | 116.3571 |  | 9.25 | 109.6 | 113.5 | 112.925
-9 |  | 3.3 | 124.4 | 126 | 125.3 |  | 3.3 | 122.24 | 124.64 | 123.9 |  | 4.4 | 118.6 | 121.65 | 121.176 |  | 6.9333 | 113 | 116.8833 | 116.2 |  | 9.4667 | 109.5 | 113.3667 | 112.8
-8.9 |  | 3.3667 | 124.3 | 125.6 | 125 |  | 3.4 | 122.2 | 124.16 | 123.5 |  | 4.4333 | 118.5 | 121.38 | 120.776 |  | 7.05 | 112.9 | 116.7 | 116.1375 |  | 9.7333 | 109.5 | 113.1167 | 112.6
-8.8 |  | 3.4 | 124.1 | 125.32 | 124.575 |  | 3.5 | 122.1 | 123.85 | 123.18 |  | 4.6 | 118.4 | 121.16 | 120.576 |  | 7.2333 | 112.9 | 116.4857 | 115.925 |  | 9.9667 | 109.5 | 113 | 112.5
-8.7 |  | 3.5 | 124.02 | 124.86 | 124.225 |  | 3.55 | 122 | 123.58 | 122.925 |  | 4.7 | 118.325 | 120.86 | 120.326 |  | 7.4667 | 112.8167 | 116.1857 | 115.7 |  | 10.12 | 109.5 | 112.9 | 112.4
-8.6 |  | 3.6 | 123.9 | 124.5333 | 123.8 |  | 3.65 | 121.8667 | 123.25 | 122.65 |  | 4.8 | 118.2 | 120.56 | 120.076 |  | 7.6 | 112.8 | 115.9667 | 115.5571 |  | 10.5 | 109.4 | 112.7 | 112.2667
-8.5 |  | 3.65 | 123.75 | 124.15 | 123.55 |  | 3.7 | 121.7 | 122.94 | 122.35 |  | 4.9 | 118.04 | 120.26 | 119.696 |  | 7.9 | 112.7 | 115.8 | 115.3667 |  | 10.7 | 109.4 | 112.6 | 112.1
-8.4 |  | 3.75 | 123.45 | 123.8 | 123.175 |  | 3.8 | 121.54 | 122.5 | 122.025 |  | 5.1 | 117.9 | 120 | 119.476 |  | 8.025 | 112.6 | 115.6833 | 115.2 |  | 11.0667 | 109.4 | 112.44 | 112.1
-8.3 |  | 3.8 | 123.25 | 123.525 | 122.925 |  | 3.9 | 121.3 | 122.3 | 121.6 |  | 5.3 | 117.7 | 119.75 | 119.236 |  | 8.3 | 112.5 | 115.475 | 114.9714 |  | 11.3 | 109.325 | 112.24 | 112
-8.2 |  | 3.9 | 122.95 | 123.2 | 122.475 |  | 4.025 | 121.175 | 122.1 | 121.2667 |  | 5.4 | 117.6 | 119.48 | 118.896 |  | 8.675 | 112.4 | 115.2 | 114.7333 |  | 11.65 | 109.3 | 112.1 | 111.85
-8.1 |  | 4 | 122.6 | 122.85 | 122.075 |  | 4.1 | 120.825 | 121.58 | 121.02 |  | 5.5 | 117.48 | 119.05 | 118.616 |  | 8.8333 | 112.4 | 115.0143 | 114.5429 |  | 12.05 | 109.3 | 112 | 111.62
-8 |  | 4.1 | 122.25 | 122.5 | 121.84 |  | 4.3 | 120.6 | 121.3 | 120.68 |  | 5.7 | 117.3 | 118.75 | 118.426 |  | 9.05 | 112.3 | 114.8143 | 114.36 |  | 12.5 | 109.2 | 111.9 | 111.4
-7.9 |  | 4.2 | 122.02 | 122.2 | 121.52 |  | 4.4 | 120.3667 | 120.975 | 120.35 |  | 5.9 | 117.125 | 118.55 | 118.156 |  | 9.4 | 112.16 | 114.5625 | 114.1857 |  | 12.925 | 109.1 | 111.7333 | 111.2857
-7.8 |  | 4.3 | 121.72 | 121.8 | 121.1 |  | 4.5 | 120.1 | 120.58 | 120.025 |  | 6.05 | 116.975 | 118.34 | 117.826 |  | 9.775 | 112.05 | 114.4286 | 114.02 |  | 13.375 | 109.1 | 111.6 | 111.1333
-7.7 |  | 4.4667 | 121.5 | 121.35 | 120.8 |  | 4.6 | 119.86 | 120.3 | 119.76 |  | 6.3 | 116.74 | 117.96 | 117.6093 |  | 10 | 111.96 | 114.26 | 113.875 |  | 13.96 | 109 | 111.425 | 110.9667
-7.6 |  | 4.575 | 121.1 | 121.1 | 120.4 |  | 4.7333 | 119.5667 | 120 | 119.4 |  | 6.5 | 116.475 | 117.64 | 117.377 |  | 10.35 | 111.9 | 114.0333 | 113.7429 |  | 14.6 | 108.9 | 111.3 | 110.9
-7.5 |  | 4.7 | 120.725 | 120.76 | 120.05 |  | 4.875 | 119.2 | 119.65 | 118.95 |  | 6.7 | 116.35 | 117.5 | 117.027 |  | 10.8 | 111.74 | 113.8143 | 113.4667 | 1.3 | 15.3 | 108.9 | 111.1167 | 110.72
-7.4 |  | 4.9 | 120.6 | 120.45 | 119.775 |  | 5.1 | 118.925 | 119.25 | 118.68 |  | 6.92 | 116.06 | 117.08 | 116.777 | 1.1 | 11.15 | 111.625 | 113.6667 | 113.2 | 1.3 | 16.3 | 108.76 | 111 | 110.6
-7.3 |  | 5 | 120.32 | 120.125 | 119.4667 |  | 5.2667 | 118.6667 | 118.95 | 118.42 |  | 7.2 | 115.8333 | 116.76 | 116.502 | 1.1 | 11.7 | 111.5 | 113.5 | 113.0429 | 1.3 | 17.4 | 108.7 | 110.85 | 110.5
-7.2 |  | 5.1 | 120.04 | 119.7 | 119.1 |  | 5.4 | 118.3 | 118.625 | 118.05 | 0.7 | 7.5 | 115.6 | 116.48 | 116.277 | 1.1 | 12.25 | 111.4 | 113.3143 | 112.8 | 1.3 | 18.8333 | 108.65 | 110.725 | 110.4
-7.1 |  | 5.3 | 119.6 | 119.35 | 118.775 | 0.7 | 5.7 | 118.1 | 118.275 | 117.82 | 0.7 | 7.8667 | 115.425 | 116.25 | 115.957 | 1.1 | 12.9 | 111.2 | 113.1 | 112.65 | 1.4 | 20.7 | 108.6 | 110.68 | 110.3
-7 | 0.7 | 5.4667 | 119.175 | 118.9333 | 118.3 | 0.6 | 5.9 | 117.8 | 118.05 | 117.5667 | 0.7 | 8.25 | 115.22 | 115.88 | 115.752 | 1.1 | 13.7 | 111.1 | 112.85 | 112.3857 | 1.5 | 23.22 | 108.6 | 110.5 | 110.2
-6.9 | 0.7 | 5.625 | 118.88 | 118.675 | 118.0333 | 0.6 | 6.1667 | 117.475 | 117.64 | 117.1 | 0.8 | 8.5 | 115 | 115.625 | 115.402 | 1.2 | 14.6 | 111 | 112.5857 | 112.2 | 1.5 | 26.775 | 108.5286 | 110.4 | 110
-6.8 | 0.7 | 5.8333 | 118.5 | 118.28 | 117.72 | 0.6 | 6.4 | 117.22 | 117.36 | 116.8 | 0.8 | 9.0667 | 114.7 | 115.4 | 115.152 | 1.3 | 15.85 | 110.8833 | 112.45 | 112.0571 | 1.6 | 31.6 | 108.5 | 110.25 | 109.86
-6.7 | 0.7 | 6.1 | 118.275 | 117.9 | 117.36 | 0.6 | 6.85 | 116.9 | 117.04 | 116.575 | 0.9 | 9.55 | 114.55 | 115.08 | 114.757 | 1.4 | 17.2 | 110.775 | 112.3 | 111.8833 | 1.6 | 37.875 | 108.5 | 110.14 | 109.7
-6.6 | 0.7 | 6.4 | 117.875 | 117.58 | 116.96 | 0.65 | 7.2 | 116.6 | 116.7 | 116.18 | 0.9 | 10.15 | 114.32 | 114.82 | 114.477 | 1.4 | 19.1 | 110.625 | 112.1 | 111.675 | 1.7 | 45.925 | 108.4 | 110.02 | 109.6333
-6.5 | 0.7 | 6.65 | 117.5 | 117.15 | 116.7 | 0.7 | 7.6 | 116.3 | 116.36 | 115.8 | 0.9 | 10.9333 | 114.1 | 114.4833 | 114.252 | 1.4 | 21.8 | 110.5 | 111.85 | 111.4833 | 1.7 | 55.4 | 108.32 | 109.9 | 109.5
-6.4 | 0.7 | 7 | 117.25 | 116.95 | 116.3 | 0.7 | 8.0667 | 116.08 | 115.95 | 115.54 | 1 | 11.75 | 113.975 | 114.3 | 114.102 | 1.5 | 25.5333 | 110.4 | 111.7 | 111.3167 | 1.8 | 64.7 | 108.3 | 109.78 | 109.4
-6.3 | 0.7 | 7.3667 | 116.86 | 116.5 | 116 | 0.7 | 8.7 | 115.7667 | 115.6 | 115.2 | 1 | 12.9 | 113.725 | 114.06 | 113.837 | 1.5 | 30.8 | 110.24 | 111.54 | 111.14 | 1.9 | 73 | 108.2 | 109.65 | 109.3333
-6.2 | 0.8 | 7.9 | 116.6 | 116.15 | 115.7 | 0.7 | 9.4333 | 115.4333 | 115.4 | 114.975 | 1 | 14.5 | 113.475 | 113.8 | 113.527 | 1.5 | 38.5 | 110.2 | 111.3667 | 110.975 | 2 | 80.8 | 108.1 | 109.5 | 109.2
-6.1 | 0.8 | 8.5667 | 116.25 | 115.84 | 115.325 | 0.8 | 10.4 | 115.1667 | 115 | 114.625 | 1.1 | 16.3 | 113.24 | 113.55 | 113.227 | 1.6 | 47.0333 | 110.1 | 111.18 | 110.82 | 2 | 86.4 | 108.02 | 109.36 | 109.1
-6 | 0.8 | 9.3667 | 115.94 | 115.5 | 115.06 | 0.8 | 11.775 | 114.84 | 114.8 | 114.32 | 1.1 | 19 | 113.05 | 113.4 | 112.977 | 1.7 | 56.5 | 110 | 111 | 110.625 | 2.1 | 91 | 107.98 | 109.16 | 108.9
-5.9 | 0.8 | 10.4 | 115.62 | 115.15 | 114.775 | 0.8 | 13.45 | 114.58 | 114.45 | 114.05 | 1.2 | 23 | 112.8 | 113.075 | 112.757 | 1.8 | 66.7 | 109.9 | 110.8429 | 110.5 | 2.2 | 94.2 | 107.9 | 109.05 | 108.7833
-5.8 | 0.8 | 11.9333 | 115.175 | 114.9 | 114.5 | 0.9 | 15.75 | 114.3 | 114.125 | 113.7 | 1.2 | 28.8 | 112.62 | 112.825 | 112.5629 | 1.9 | 76.4 | 109.7667 | 110.6 | 110.3833 | 2.3 | 96.2333 | 107.8 | 108.9 | 108.7
-5.7 | 0.9 | 13.8 | 114.9 | 114.525 | 114.25 | 0.9 | 19.2 | 114.04 | 113.9 | 113.46 | 1.2 | 36.9 | 112.3 | 112.6 | 112.3447 | 1.9 | 83.8 | 109.6 | 110.4286 | 110.2 | 2.4 | 97.5333 | 107.8 | 108.8167 | 108.5167
-5.6 | 0.9 | 16.7 | 114.76 | 114.3 | 113.875 | 0.9 | 24.4333 | 113.75 | 113.7 | 113.26 | 1.3 | 46.45 | 112.1833 | 112.4 | 112.178 | 2 | 89.7 | 109.475 | 110.3 | 110.1 | 2.4 | 98.7 | 107.6857 | 108.6333 | 108.42
-5.5 | 0.9 | 21.35 | 114.5333 | 114.025 | 113.56 | 0.95 | 31.8667 | 113.5 | 113.4667 | 113 | 1.3 | 57.4 | 112 | 112.12 | 111.9113 | 2 | 93.7 | 109.36 | 110.2 | 110 | 2.6 | 99.6667 | 107.6 | 108.5 | 108.3
-5.4 | 0.9 | 28 | 114.2 | 113.68 | 113.325 | 1 | 42.125 | 113.28 | 113.1667 | 112.78 | 1.4 | 67.5333 | 111.86 | 111.975 | 111.678 | 2.2 | 96.45 | 109.25 | 110.05 | 109.8286 | 2.7 | 100.3 | 107.5 | 108.4 | 108.2
-5.3 | 1 | 38.1 | 113.95 | 113.45 | 113.05 | 1 | 53.2 | 113.06 | 112.9 | 112.56 | 1.4 | 78.2333 | 111.68 | 111.75 | 111.528 | 2.3 | 98.3333 | 109.1 | 109.88 | 109.65 | 2.8 | 101 | 107.3667 | 108.34 | 108.1333
-5.2 | 1 | 49.2 | 113.64 | 113.225 | 112.725 | 1.05 | 63.975 | 112.82 | 112.68 | 112.2667 | 1.5 | 86.5667 | 111.475 | 111.525 | 111.3447 | 2.3 | 99.8 | 108.96 | 109.7 | 109.4857 | 2.9 | 101.5 | 107.26 | 108.2 | 107.96
-5.1 | 1.1 | 61.35 | 113.4 | 113 | 112.64 | 1.1 | 75.7 | 112.6 | 112.45 | 112.12 | 1.6 | 92.3 | 111.2333 | 111.3 | 111.1613 | 2.4 | 100.9 | 108.9 | 109.5333 | 109.3429 | 3 | 101.9667 | 107.2 | 108 | 107.8
-5 | 1.1 | 73.3 | 113.1 | 112.775 | 112.4 | 1.15 | 84.8333 | 112.35 | 112.225 | 111.8 | 1.6 | 96.8 | 111.1 | 111.15 | 110.978 | 2.5 | 101.75 | 108.8 | 109.35 | 109.18 | 3.1 | 102.1667 | 107.1 | 107.825 | 107.62
-4.9 | 1.1 | 83.5667 | 112.86 | 112.475 | 112.1667 | 1.2 | 92.7 | 112.18 | 112 | 111.6 | 1.7 | 99.6 | 110.975 | 110.9 | 110.728 | 2.6 | 102.3667 | 108.6 | 109.0833 | 109.0143 | 3.2 | 102.3333 | 107 | 107.7167 | 107.5
-4.8 | 1.2 | 92.1 | 112.54 | 112.2 | 111.96 | 1.2 | 97.6 | 111.85 | 111.65 | 111.34 | 1.8 | 101.3 | 110.625 | 110.65 | 110.538 | 2.8 | 102.7667 | 108.5 | 109 | 108.8571 | 3.3 | 102.625 | 107 | 107.62 | 107.325
-4.7 | 1.2 | 98.35 | 112.3333 | 111.9333 | 111.625 | 1.3 | 101.05 | 111.6 | 111.5 | 111.075 | 1.8 | 102.7 | 110.5 | 110.45 | 110.278 | 2.9 | 103.2 | 108.4 | 108.8833 | 108.6667 | 3.5 | 102.8 | 106.88 | 107.55 | 107.3
-4.6 | 1.3 | 102.0333 | 112.1 | 111.74 | 111.4 | 1.4 | 103.15 | 111.45 | 111.16 | 110.925 | 1.9 | 103.55 | 110.35 | 110.3 | 110.078 | 3 | 103.525 | 108.32 | 108.7167 | 108.525 | 3.7 | 102.9 | 106.7 | 107.46 | 107.1
-4.5 | 1.3 | 104.4 | 111.9 | 111.4 | 111.1 | 1.4 | 104.5333 | 111.1 | 110.94 | 110.6 | 2 | 104.2 | 110.125 | 109.975 | 109.878 | 3.1 | 103.775 | 108.08 | 108.6 | 108.2714 | 3.9 | 103.075 | 106.52 | 107.2667 | 106.9333
-4.4 | 1.4 | 105.85 | 111.6 | 111.0667 | 110.94 | 1.5 | 105.4333 | 110.975 | 110.8 | 110.4 | 2.1 | 104.7 | 109.88 | 109.8 | 109.628 | 3.25 | 104 | 107.8833 | 108.4714 | 108.1167 | 4.1 | 103.1 | 106.4 | 107.1 | 106.7667
-4.3 | 1.4 | 106.9 | 111.325 | 110.8667 | 110.625 | 1.6 | 106.3 | 110.7 | 110.5 | 110.125 | 2.2 | 105.2 | 109.7 | 109.625 | 109.378 | 3.5 | 104.125 | 107.8 | 108.1857 | 107.9833 | 4.3 | 103.25 | 106.3 | 106.9333 | 106.6
-4.2 | 1.5 | 107.5667 | 111.1 | 110.62 | 110.34 | 1.7 | 106.7333 | 110.5 | 110.325 | 109.9667 | 2.4 | 105.5 | 109.4 | 109.4667 | 109.158 | 3.7 | 104.25 | 107.6667 | 108 | 107.8167 | 4.5 | 103.4 | 106.3 | 106.8 | 106.5
-4.1 | 1.6 | 108 | 110.9 | 110.4 | 110.2 | 1.8 | 107 | 110.24 | 110.04 | 109.7 | 2.5 | 105.65 | 109.3 | 109.12 | 108.978 | 3.95 | 104.3 | 107.475 | 107.8 | 107.5714 | 4.7 | 103.5 | 106.2 | 106.7 | 106.4333
-4 | 1.65 | 108.2 | 110.62 | 110.225 | 109.85 | 1.9 | 107.2667 | 110 | 109.825 | 109.5 | 2.7 | 105.8667 | 109.075 | 108.92 | 108.778 | 4.15 | 104.4 | 107.26 | 107.6333 | 107.4286 | 5 | 103.5 | 106.04 | 106.6 | 106.32
-3.9 | 1.7 | 108.3 | 110.32 | 109.95 | 109.6 | 2 | 107.4 | 109.68 | 109.5 | 109.3 | 2.8 | 105.9 | 108.775 | 108.7 | 108.578 | 4.4 | 104.475 | 107.16 | 107.44 | 107.2667 | 5.6 | 103.6 | 105.9 | 106.4 | 106.15
-3.8 | 1.9 | 108.4 | 109.95 | 109.6 | 109.35 | 2.15 | 107.5 | 109.425 | 109.34 | 109.075 | 3 | 106 | 108.525 | 108.52 | 108.278 | 4.8 | 104.5 | 106.95 | 107.2667 | 107.0833 | 5.95 | 103.6 | 105.8 | 106.3 | 105.9667
-3.7 | 2 | 108.4 | 109.68 | 109.425 | 109.1 | 2.3 | 107.575 | 109.26 | 109.06 | 108.8 | 3.3 | 106.1 | 108.3 | 108.32 | 108.038 | 5 | 104.5 | 106.82 | 107.125 | 106.875 | 6.5 | 103.6 | 105.7 | 106.1667 | 105.8
-3.6 | 2.1 | 108.4 | 109.5 | 109.1667 | 108.85 | 2.6 | 107.5 | 108.94 | 108.78 | 108.6 | 3.6 | 106 | 108.15 | 108.05 | 107.818 | 5.4 | 104.4 | 106.725 | 106.8714 | 106.625 | 7.15 | 103.6 | 105.54 | 106 | 105.7
-3.5 | 2.3 | 108.2333 | 109.2667 | 108.8 | 108.5 | 2.8 | 107.5 | 108.65 | 108.6 | 108.275 | 3.9 | 106 | 108 | 107.84 | 107.728 | 5.95 | 104.4 | 106.5 | 106.6333 | 106.475 | 7.8667 | 103.6 | 105.5 | 105.9 | 105.6
-3.4 | 2.5 | 108.1 | 108.9 | 108.54 | 108.38 | 3 | 107.4 | 108.5 | 108.375 | 108.025 | 4.25 | 106 | 107.7 | 107.6 | 107.419 | 6.65 | 104.4 | 106.375 | 106.5 | 106.3 | 9 | 103.55 | 105.3333 | 105.68 | 105.44
-3.3 | 2.7 | 108 | 108.7 | 108.4 | 108.05 | 3.3 | 107.3 | 108.175 | 108.16 | 107.775 | 4.8 | 105.9 | 107.4 | 107.4 | 107.279 | 7.45 | 104.4 | 106.3 | 106.36 | 106.2167 | 10.1333 | 103.6 | 105.2 | 105.425 | 105.25
-3.2 | 3.1 | 107.85 | 108.42 | 108.05 | 107.8 | 3.8 | 107.2 | 108.06 | 107.85 | 107.6 | 5.3 | 105.8 | 107.28 | 107.125 | 106.959 | 8.4 | 104.4 | 106.0333 | 106.1 | 106.04 | 12.35 | 103.5 | 105.1 | 105.2 | 105.0833
-3.1 | 3.35 | 107.7 | 108.22 | 107.8 | 107.62 | 4.5 | 107.1 | 107.84 | 107.56 | 107.32 | 6.1 | 105.6333 | 107.05 | 107 | 106.729 | 9.85 | 104.4 | 105.825 | 105.9571 | 105.8571 | 14.45 | 103.5 | 105 | 105.1 | 105
-3 | 3.85 | 107.55 | 107.925 | 107.5 | 107.3 | 5.15 | 106.9 | 107.5 | 107.25 | 107.025 | 7 | 105.5 | 106.82 | 106.7 | 106.599 | 11.6 | 104.4 | 105.7 | 105.8429 | 105.6833 | 17.6 | 103.5 | 104.8 | 105 | 104.8
-2.9 | 4.6 | 107.3 | 107.7 | 107.32 | 107 | 6 | 106.8 | 107.26 | 107 | 106.775 | 8.4 | 105.4 | 106.52 | 106.5167 | 106.399 | 13.7 | 104.4 | 105.4333 | 105.6 | 105.4571 | 22.4 | 103.4 | 104.7 | 104.88 | 104.7
-2.8 | 5.6 | 107.1667 | 107.3 | 107.025 | 106.825 | 7.3 | 106.6 | 107.025 | 106.8167 | 106.625 | 10 | 105.3 | 106.3 | 106.3 | 106.179 | 17.4 | 104.3 | 105.3 | 105.5 | 105.2667 | 29.2 | 103.3 | 104.5667 | 104.8 | 104.55
-2.7 | 7 | 106.9 | 107.05 | 106.775 | 106.62 | 9 | 106.4 | 106.7 | 106.55 | 106.375 | 12.4 | 105.3 | 106.075 | 106 | 105.929 | 21.9 | 104.2 | 105.1 | 105.2 | 105.1 | 37.1 | 103.3 | 104.4333 | 104.5833 | 104.35
-2.6 | 8.8 | 106.6667 | 106.8 | 106.54 | 106.45 | 11.3 | 106.2 | 106.45 | 106.25 | 106.125 | 15.85 | 105.2 | 105.8 | 105.82 | 105.729 | 28.9 | 104.1 | 104.9 | 105.0667 | 104.9333 | 47.1 | 103.3 | 104.3 | 104.4 | 104.14
-2.5 | 11.4 | 106.45 | 106.55 | 106.24 | 106.06 | 14.8 | 106.1 | 106.2 | 106.1 | 105.875 | 21 | 105.1 | 105.7 | 105.62 | 105.504 | 37.3 | 104 | 104.775 | 104.9143 | 104.7333 | 57.5 | 103.12 | 104.15 | 104.28 | 103.9667
-2.4 | 15.65 | 106.125 | 106.3 | 105.98 | 105.825 | 19.6 | 105.8667 | 106 | 105.9 | 105.6333 | 27.9 | 104.9 | 105.48 | 105.32 | 105.229 | 47.5 | 103.8667 | 104.6833 | 104.7429 | 104.525 | 67.3 | 103.1 | 103.92 | 104.1 | 103.86
-2.3 | 22.45 | 106 | 105.96 | 105.8333 | 105.58 | 27.55 | 105.7 | 105.7667 | 105.64 | 105.36 | 37.3 | 104.6333 | 105.18 | 105.12 | 104.979 | 58.7 | 103.7333 | 104.4667 | 104.5167 | 104.475 | 76.3 | 103.075 | 103.8 | 103.92 | 103.725
-2.2 | 31.7 | 105.7333 | 105.725 | 105.48 | 105.32 | 37.9 | 105.35 | 105.54 | 105.3333 | 105.2 | 48.6 | 104.5 | 104.98 | 104.84 | 104.819 | 69.3 | 103.65 | 104.3 | 104.325 | 104.3 | 83.4 | 103 | 103.6286 | 103.7 | 103.58
-2.1 | 42.3 | 105.4333 | 105.4667 | 105.25 | 105.025 | 48.1 | 105.1 | 105.26 | 105.1667 | 104.975 | 59.1 | 104.4 | 104.84 | 104.6 | 104.599 | 77.6 | 103.5 | 104.1 | 104.075 | 104.0667 | 88.55 | 102.9 | 103.44 | 103.6 | 103.5
-2 | 54.6 | 105.2 | 105.14 | 104.9 | 104.92 | 60 | 105 | 104.95 | 104.9 | 104.7 | 69.9 | 104.3 | 104.64 | 104.375 | 104.354 | 85.4 | 103.4 | 103.8833 | 103.8857 | 103.85 | 92.1 | 102.8 | 103.3 | 103.44 | 103.275
-1.9 | 66.1 | 105.025 | 104.925 | 104.675 | 104.65 | 70.3 | 104.8 | 104.64 | 104.68 | 104.45 | 79.4 | 104.0667 | 104.3 | 104.125 | 104.199 | 90.2 | 103.3333 | 103.78 | 103.7429 | 103.6667 | 94.75 | 102.7 | 103.2167 | 103.25 | 103.1
-1.8 | 76.7 | 104.7333 | 104.65 | 104.4333 | 104.4 | 80.1 | 104.5 | 104.425 | 104.4333 | 104.225 | 86.8 | 103.8 | 104.1 | 103.9167 | 103.929 | 93.5 | 103.3 | 103.6 | 103.5 | 103.5 | 96.6 | 102.5 | 103.1167 | 103.0667 | 102.95
-1.7 | 85.6 | 104.5 | 104.325 | 104.2 | 104.125 | 87.9 | 104.3 | 104.22 | 104.16 | 103.9 | 91.4 | 103.6 | 103.92 | 103.725 | 103.699 | 95.85 | 103.1 | 103.3 | 103.3571 | 103.2571 | 97.7 | 102.4 | 102.95 | 102.9 | 102.75
-1.6 | 91.5 | 104.3 | 104.075 | 104 | 103.9 | 92.4 | 104.0333 | 103.98 | 103.825 | 103.75 | 94.7 | 103.4 | 103.675 | 103.5 | 103.479 | 97.1 | 102.9 | 103.16 | 103.1 | 103 | 98.6 | 102.3 | 102.775 | 102.72 | 102.6
-1.5 | 95.3 | 103.9 | 103.84 | 103.65 | 103.6 | 95.5 | 103.72 | 103.72 | 103.475 | 103.48 | 96.9 | 103.2 | 103.45 | 103.175 | 103.279 | 98.45 | 102.625 | 102.8833 | 102.8333 | 102.8429 | 99.6 | 102.2 | 102.58 | 102.6167 | 102.5
-1.4 | 97.6 | 103.6 | 103.625 | 103.4 | 103.4 | 97.9 | 103.5 | 103.425 | 103.2 | 103.125 | 98.3 | 103 | 103.18 | 103 | 103.029 | 99.3 | 102.4667 | 102.7 | 102.6667 | 102.675 | 99.9 | 102.14 | 102.4 | 102.425 | 102.2833
-1.3 | 99.3 | 103.375 | 103.2 | 103.1333 | 103.1 | 99 | 103.2667 | 103.15 | 103.1 | 102.975 | 99.2 | 102.8 | 102.98 | 102.8 | 102.854 | 99.95 | 102.3 | 102.5 | 102.4833 | 102.4833 | 100.2 | 102 | 102.225 | 102.3 | 102.1167
-1.2 | 100.1 | 103.1333 | 102.98 | 102.95 | 102.875 | 99.6 | 103 | 102.9 | 102.8167 | 102.74 | 99.9 | 102.6333 | 102.75 | 102.6 | 102.604 | 100.2 | 102.1 | 102.3 | 102.3286 | 102.2333 | 100.4 | 101.8667 | 102 | 102.05 | 101.92
-1.1 | 100.7 | 102.9333 | 102.8 | 102.6333 | 102.58 | 100.4 | 102.8 | 102.6 | 102.62 | 102.54 | 100.2 | 102.4 | 102.5 | 102.34 | 102.38 | 100.5 | 102.025 | 102.1 | 102.1 | 102.0667 | 100.6 | 101.65 | 101.8 | 101.9143 | 101.82
-1 | 100.9 | 102.65 | 102.575 | 102.36 | 102.4 | 100.6 | 102.475 | 102.4 | 102.35 | 102.24 | 100.5 | 102.1 | 102.275 | 102.1 | 102.06 | 100.65 | 101.9 | 101.875 | 101.96 | 101.7857 | 100.8 | 101.5 | 101.6 | 101.68 | 101.6
-0.9 | 101.1 | 102.3 | 102.26 | 102.06 | 102.1 | 100.85 | 102.1 | 102.08 | 102 | 102.025 | 100.7 | 101.9667 | 102.02 | 101.85 | 101.855 | 100.7 | 101.75 | 101.725 | 101.7 | 101.6333 | 100.8 | 101.3667 | 101.5 | 101.475 | 101.475
-0.8 | 101.2 | 101.95 | 101.85 | 101.84 | 101.9 | 101 | 101.9 | 101.86 | 101.875 | 101.825 | 100.8 | 101.7 | 101.75 | 101.54 | 101.705 | 100.7 | 101.6 | 101.6 | 101.5571 | 101.44 | 100.8 | 101.2 | 101.3 | 101.36 | 101.28
-0.7 | 101.2 | 101.75 | 101.6 | 101.6 | 101.62 | 101.1 | 101.7 | 101.58 | 101.55 | 101.55 | 100.9 | 101.475 | 101.52 | 101.4 | 101.38 | 100.8 | 101.4 | 101.38 | 101.3 | 101.3 | 100.7 | 101 | 101.1667 | 101.225 | 101.06
-0.6 | 101.1 | 101.5 | 101.325 | 101.35 | 101.38 | 101 | 101.4667 | 101.28 | 101.35 | 101.3 | 100.9 | 101.2 | 101.3 | 101.2333 | 101.28 | 100.7 | 101.1 | 101.18 | 101.2167 | 101.1 | 100.7 | 100.9 | 101 | 101.1167 | 100.98
-0.5 | 101 | 101.2 | 101.175 | 101.05 | 101.1 | 100.8 | 101.3 | 101.075 | 101.2 | 101.04 | 100.8 | 101.0333 | 101.04 | 101.05 | 101.02 | 100.6 | 101 | 100.98 | 101 | 100.8714 | 100.7 | 100.8 | 100.8333 | 100.9571 | 100.8
-0.4 | 100.9 | 100.9 | 100.92 | 100.9 | 100.8667 | 100.6 | 101 | 100.84 | 100.9 | 100.8 | 100.7 | 100.7333 | 100.85 | 100.84 | 100.78 | 100.5 | 100.7 | 100.75 | 100.7667 | 100.7167 | 100.6 | 100.675 | 100.64 | 100.7 | 100.6333
-0.3 | 100.7 | 100.6 | 100.65 | 100.66 | 100.65 | 100.5 | 100.7333 | 100.625 | 100.65 | 100.5 | 100.5 | 100.5333 | 100.6 | 100.6 | 100.53 | 100.5 | 100.5 | 100.5 | 100.6 | 100.4714 | 100.45 | 100.5 | 100.425 | 100.5 | 100.5
-0.2 | 100.4 | 100.4333 | 100.44 | 100.42 | 100.4 | 100.4 | 100.5 | 100.3667 | 100.425 | 100.3 | 100.3 | 100.3 | 100.4 | 100.3 | 100.305 | 100.3 | 100.3 | 100.3 | 100.35 | 100.3375 | 100.2 | 100.3333 | 100.3 | 100.3 | 100.325
-0.1 | 100.3 | 100.2 | 100.22 | 100.2 | 100.2 | 100.2 | 100.225 | 100.2 | 100.1667 | 100.15 | 100.1 | 100.1667 | 100.25 | 100.125 | 100.1967 | 100.1 | 100.2 | 100.14 | 100.1833 | 100.2 | 100.1 | 100.2 | 100.125 | 100.25 | 100.1833
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.8 | 99.8 | 99.78 | 99.75 | 99.8 | 99.8 | 99.7 | 99.6 | 99.8333 | 99.8 | 99.8 | 99.7667 | 99.78 | 99.75 | 99.7967 | 99.8 | 99.7333 | 99.82 | 99.7556 | 99.8714 | 99.8 | 99.8 | 99.8833 | 99.85 | 99.88
0.2 | 99.6 | 99.5333 | 99.4667 | 99.6 | 99.575 | 99.6 | 99.575 | 99.45 | 99.7 | 99.62 | 99.55 | 99.5 | 99.5667 | 99.4667 | 99.5801 | 99.6 | 99.575 | 99.78 | 99.6286 | 99.7167 | 99.6 | 99.7 | 99.7167 | 99.725 | 99.75
0.3 | 99.3 | 99.35 | 99.3 | 99.425 | 99.4 | 99.4 | 99.3333 | 99.25 | 99.36 | 99.46 | 99.3 | 99.2 | 99.38 | 99.26 | 99.4301 | 99.5 | 99.4 | 99.4667 | 99.3857 | 99.5 | 99.3 | 99.5667 | 99.52 | 99.6 | 99.6
0.4 | 98.95 | 99.1 | 99.04 | 99.24 | 99.2 | 99 | 99.2 | 99.02 | 99.125 | 99.22 | 98.85 | 99 | 99.2333 | 99.125 | 99.2801 | 99.25 | 99.2 | 99.32 | 99.2429 | 99.28 | 99.15 | 99.3667 | 99.3333 | 99.45 | 99.3
0.5 | 98.6 | 98.8667 | 98.8 | 99 | 99 | 98.7 | 99 | 98.8667 | 99 | 98.975 | 98.6 | 98.8333 | 98.975 | 98.98 | 99.0552 | 99 | 99.05 | 99.1167 | 99.1 | 99.15 | 98.95 | 99.14 | 99.1667 | 99.2833 | 99.2
0.6 | 98.4 | 98.7 | 98.7 | 98.875 | 98.9 | 98.3 | 98.8 | 98.6833 | 98.8167 | 98.85 | 98.3 | 98.6667 | 98.825 | 98.75 | 98.8202 | 98.7 | 98.9 | 99 | 99 | 98.9714 | 98.7 | 99 | 99 | 99.1167 | 99.1
0.7 | 98.1 | 98.4 | 98.6 | 98.625 | 98.7 | 97.9 | 98.5333 | 98.575 | 98.7 | 98.6333 | 98 | 98.4333 | 98.58 | 98.6 | 98.7136 | 98.3 | 98.675 | 98.74 | 98.8429 | 98.8 | 98.4 | 98.825 | 98.8333 | 99.0333 | 99
0.8 | 97.5 | 98.2 | 98.2833 | 98.475 | 98.5 | 97.55 | 98.2667 | 98.325 | 98.525 | 98.5 | 97.6 | 98.15 | 98.34 | 98.5 | 98.5803 | 98 | 98.5 | 98.54 | 98.5833 | 98.65 | 98.1 | 98.6667 | 98.74 | 98.8 | 98.775
0.9 | 97.15 | 98 | 98.1 | 98.3 | 98.375 | 97 | 98 | 98.1667 | 98.35 | 98.375 | 97.2 | 98 | 98.2 | 98.2667 | 98.3803 | 97.5 | 98.3333 | 98.4 | 98.4167 | 98.4857 | 97.7 | 98.5 | 98.55 | 98.675 | 98.55
1 | 96.6 | 97.85 | 97.975 | 98.15 | 98.2 | 96.2 | 97.9 | 97.9 | 98.2 | 98.15 | 96.65 | 97.8 | 98.05 | 98.08 | 98.2403 | 97.1 | 98.1 | 98.24 | 98.28 | 98.3333 | 97.3333 | 98.3 | 98.4 | 98.4833 | 98.4667
1.1 | 95.7 | 97.5667 | 97.76 | 97.925 | 98 | 95.7 | 97.675 | 97.76 | 97.96 | 97.92 | 96 | 97.6333 | 97.8 | 97.95 | 98.1304 | 96.8 | 97.9 | 98.1 | 98.1143 | 98.1833 | 96.95 | 98.05 | 98.22 | 98.3167 | 98.3
1.2 | 95.1 | 97.3 | 97.55 | 97.775 | 97.75 | 94.6 | 97.45 | 97.625 | 97.825 | 97.8 | 95 | 97.4 | 97.65 | 97.72 | 97.9804 | 96.3 | 97.8 | 97.8833 | 97.98 | 98.0167 | 96.65 | 97.8 | 98.14 | 98.1667 | 98.15
1.3 | 94 | 97.1333 | 97.325 | 97.66 | 97.5 | 93.9 | 97.3 | 97.38 | 97.6 | 97.65 | 94.4 | 97.15 | 97.5 | 97.6 | 97.7304 | 95.4 | 97.6 | 97.8 | 97.8 | 97.85 | 96.15 | 97.7 | 97.9 | 98 | 98
1.4 | 92.4 | 96.9 | 97.2 | 97.44 | 97.475 | 92 | 97.0667 | 97.2 | 97.45 | 97.42 | 93.1 | 97.0333 | 97.28 | 97.475 | 97.5805 | 94.9 | 97.4 | 97.5667 | 97.6429 | 97.6833 | 95.45 | 97.45 | 97.7 | 97.7833 | 97.82
1.5 | 89.9 | 96.7 | 97.05 | 97.24 | 97.3 | 90.2 | 96.9333 | 97.025 | 97.3 | 97.2667 | 91.5 | 96.7667 | 97.15 | 97.2667 | 97.3805 | 93.7 | 97.1 | 97.4 | 97.55 | 97.6 | 94.7 | 97.3 | 97.6 | 97.7 | 97.6333
1.6 | 86.6 | 96.5 | 96.82 | 97.04 | 97.1 | 87.2 | 96.775 | 96.8 | 97.18 | 97.14 | 89.6 | 96.6 | 97 | 97.125 | 97.2206 | 92.4 | 96.925 | 97.2 | 97.4 | 97.3167 | 94.05 | 97.1 | 97.35 | 97.5 | 97.4333
1.7 | 81.7 | 96.3 | 96.675 | 96.9 | 97 | 83.2 | 96.5 | 96.6 | 96.975 | 97.02 | 86.4 | 96.3 | 96.75 | 96.975 | 97.1006 | 90.9 | 96.75 | 97.1 | 97.1333 | 97.1857 | 93.1 | 97 | 97.2 | 97.4 | 97.4
1.8 | 74 | 96.1 | 96.5 | 96.7333 | 96.82 | 76.9 | 96.175 | 96.42 | 96.75 | 96.88 | 82.9 | 96.1 | 96.62 | 96.84 | 96.8806 | 88.5 | 96.5 | 96.9 | 96.96 | 96.9833 | 91.6 | 96.6333 | 97.075 | 97.2 | 97.3
1.9 | 64.5 | 95.9 | 96.275 | 96.6 | 96.675 | 68.75 | 96 | 96.25 | 96.6333 | 96.7 | 76.5 | 95.9 | 96.3 | 96.5667 | 96.7007 | 85.9 | 96.4 | 96.68 | 96.8714 | 96.8167 | 90.15 | 96.475 | 96.9 | 97.05 | 97.2
2 | 53.5 | 95.7 | 96.1 | 96.45 | 96.45 | 58.55 | 95.8 | 96.025 | 96.4 | 96.46 | 68.3 | 95.6667 | 96.2 | 96.325 | 96.4807 | 81.8 | 96.2 | 96.5 | 96.7 | 96.7 | 88.2 | 96.2 | 96.8 | 96.88 | 97
2.1 | 41.9 | 95.4667 | 95.92 | 96.24 | 96.325 | 47.85 | 95.6 | 95.9 | 96.3 | 96.34 | 59.2 | 95.5 | 96 | 96.2 | 96.3607 | 75.9 | 95.8667 | 96.4 | 96.52 | 96.6143 | 84.7 | 95.975 | 96.6 | 96.8 | 96.78
2.2 | 31.5 | 95.3 | 95.7667 | 96.075 | 96.175 | 38.1 | 95.3 | 95.7 | 96.16 | 96.12 | 48.85 | 95.1667 | 95.85 | 96 | 96.1808 | 68.5 | 95.5333 | 96.2 | 96.35 | 96.48 | 80.6 | 95.8 | 96.5 | 96.6 | 96.675
2.3 | 23 | 95.05 | 95.52 | 95.9 | 95.925 | 27.9 | 95.05 | 95.5 | 95.94 | 95.98 | 38.1 | 94.875 | 95.6 | 95.85 | 96.0208 | 58.7 | 95.3 | 95.9167 | 96.2 | 96.2667 | 74.6 | 95.55 | 96.35 | 96.4333 | 96.4833
2.4 | 15.85 | 94.7 | 95.3 | 95.65 | 95.8 | 20.1 | 94.8 | 95.325 | 95.7 | 95.8333 | 29.5 | 94.7 | 95.4167 | 95.725 | 95.7808 | 49.2 | 95.0667 | 95.8 | 95.95 | 96.1143 | 67.1 | 95.4 | 96.15 | 96.2 | 96.3286
2.5 | 11.5 | 94.5 | 95.16 | 95.5 | 95.65 | 15.1 | 94.55 | 95.075 | 95.6 | 95.6 | 21.9 | 94.5 | 95.24 | 95.5 | 95.6559 | 39.8 | 94.8333 | 95.5833 | 95.8 | 95.84 | 59 | 95.275 | 95.9667 | 96.1 | 96.1
2.6 | 8.9 | 94.3 | 95 | 95.35 | 95.5 | 11.2 | 94.2667 | 94.82 | 95.38 | 95.4 | 16.5 | 94.25 | 95.1 | 95.3667 | 95.4809 | 30.4 | 94.5333 | 95.38 | 95.675 | 95.64 | 48.8 | 95.0333 | 95.68 | 95.92 | 95.9714
2.7 | 6.8 | 94 | 94.75 | 95.1333 | 95.3833 | 8.75 | 93.975 | 94.6 | 95.2 | 95.3 | 12.8 | 93.85 | 94.98 | 95.175 | 95.3309 | 23 | 94.35 | 95.18 | 95.5333 | 95.54 | 39.6 | 94.7 | 95.5143 | 95.8 | 95.8
2.8 | 5.55 | 93.8 | 94.6 | 95 | 95.2 | 7.1 | 93.7 | 94.5 | 95 | 95.05 | 10.4 | 93.6 | 94.7 | 94.94 | 95.141 | 18 | 94.1 | 95 | 95.3143 | 95.3875 | 31.6 | 94.45 | 95.4333 | 95.68 | 95.6833
2.9 | 4.6 | 93.6 | 94.36 | 94.86 | 94.9833 | 5.9 | 93.5667 | 94.4 | 94.8 | 94.92 | 8.5 | 93.3 | 94.55 | 94.8 | 94.906 | 14.3 | 93.8667 | 94.75 | 95.1143 | 95.2571 | 24.75 | 94.2333 | 95.15 | 95.5 | 95.475
3 | 3.8 | 93.3 | 94.05 | 94.65 | 94.8 | 4.9 | 93.3667 | 94.15 | 94.575 | 94.75 | 7.1 | 93.1333 | 94.325 | 94.7 | 94.6811 | 11.9 | 93.4667 | 94.5 | 94.9667 | 95.1 | 19.2 | 93.9 | 95 | 95.3 | 95.35
3.1 | 3.3 | 92.95 | 93.9 | 94.5 | 94.5667 | 4.15 | 93 | 94 | 94.4 | 94.5 | 6.1 | 92.8 | 94.18 | 94.48 | 94.5811 | 9.95 | 93.2 | 94.4 | 94.8 | 94.8667 | 15.4 | 93.6 | 94.86 | 95.2 | 95.2
3.2 | 2.85 | 92.7 | 93.825 | 94.3 | 94.375 | 3.6 | 92.7 | 93.74 | 94.3 | 94.375 | 5.25 | 92.5 | 93.9333 | 94.3 | 94.4811 | 8.5 | 92.8 | 94.2833 | 94.5667 | 94.66 | 12.8 | 93.35 | 94.55 | 95.04 | 94.925
3.3 | 2.5 | 92.4333 | 93.6 | 94.1 | 94.2333 | 3.2 | 92.475 | 93.525 | 94.125 | 94.22 | 4.8 | 92.2667 | 93.8 | 94.2 | 94.2811 | 7.3 | 92.6 | 94.0333 | 94.36 | 94.4667 | 10.85 | 93.1 | 94.4 | 94.85 | 94.8
3.4 | 2.3 | 92.1 | 93.35 | 93.975 | 94.0667 | 2.9 | 92.1 | 93.3 | 93.975 | 93.9667 | 4.2 | 92.025 | 93.675 | 94 | 94.1412 | 6.5667 | 92.225 | 93.75 | 94.1667 | 94.2875 | 9.3 | 92.75 | 94.24 | 94.6714 | 94.7
3.5 | 2.1 | 91.8333 | 93.175 | 93.8 | 93.82 | 2.6 | 91.775 | 93.14 | 93.775 | 93.72 | 3.8 | 91.7 | 93.44 | 93.75 | 93.9062 | 6 | 91.94 | 93.6 | 94.0143 | 94.1625 | 8.2 | 92.4333 | 94.08 | 94.5167 | 94.425
3.6 | 1.9 | 91.4 | 93 | 93.7 | 93.7 | 2.4 | 91.4333 | 92.98 | 93.62 | 93.625 | 3.5 | 91.2 | 93.225 | 93.5 | 93.7812 | 5.4 | 91.7 | 93.4667 | 93.9 | 94.0286 | 7.4 | 92.2 | 93.85 | 94.32 | 94.2833
3.7 | 1.8 | 90.9667 | 92.86 | 93.4 | 93.6 | 2.1 | 91.025 | 92.825 | 93.4 | 93.4167 | 3.1 | 90.7333 | 93 | 93.35 | 93.5563 | 4.95 | 91.25 | 93.2 | 93.74 | 93.88 | 6.65 | 91.9333 | 93.7 | 94.2 | 94.2
3.8 | 1.7 | 90.6 | 92.65 | 93.25 | 93.375 | 1.9 | 90.6333 | 92.65 | 93.2 | 93.275 | 2.9 | 90.4 | 92.825 | 93.25 | 93.4313 | 4.5 | 90.9 | 92.98 | 93.5 | 93.7 | 5.9 | 91.5667 | 93.5 | 93.9667 | 94
3.9 | 1.6 | 90.2 | 92.4 | 93.1 | 93.2 | 1.8 | 90.1 | 92.325 | 93 | 93.1 | 2.75 | 89.95 | 92.55 | 93.025 | 93.0814 | 4.25 | 90.5 | 92.775 | 93.3667 | 93.4333 | 5.4 | 91.35 | 93.3 | 93.8 | 93.88
4 | 1.5 | 89.6 | 92.25 | 92.84 | 93.075 | 1.7 | 89.6 | 92.1 | 92.8667 | 92.9 | 2.5 | 89.5667 | 92.32 | 92.8167 | 92.9564 | 3.95 | 90.225 | 92.64 | 93.225 | 93.2667 | 5.05 | 90.925 | 93.16 | 93.575 | 93.7
4.1 | 1.4 | 89.0667 | 92 | 92.7 | 92.825 | 1.6 | 88.95 | 91.9333 | 92.7 | 92.725 | 2.4 | 89.1 | 92.06 | 92.6 | 92.7814 | 3.7333 | 89.7 | 92.4333 | 93 | 93.1143 | 4.7 | 90.5667 | 92.96 | 93.44 | 93.5
4.2 | 1.3 | 88.2667 | 91.8 | 92.55 | 92.6 | 1.5 | 88.4 | 91.7 | 92.52 | 92.5 | 2.2667 | 88.6 | 91.96 | 92.4 | 92.6815 | 3.6 | 89.4 | 92.22 | 92.7857 | 92.9333 | 4.45 | 90.2 | 92.825 | 93.22 | 93.28
4.3 | 1.3 | 87.4 | 91.6 | 92.3 | 92.45 | 1.4 | 87.4 | 91.54 | 92.275 | 92.35 | 2.1 | 87.9 | 91.64 | 92.225 | 92.4565 | 3.3 | 88.9 | 92.04 | 92.58 | 92.7 | 4.2 | 89.9333 | 92.575 | 93.1 | 93
4.4 | 1.2 | 86.35 | 91.425 | 92.2 | 92.3 | 1.4 | 86.6 | 91.4 | 92.125 | 92.125 | 2 | 87.2333 | 91.5 | 92.1 | 92.2415 | 3.1 | 88.6 | 91.84 | 92.4 | 92.4833 | 4.05 | 89.55 | 92.35 | 92.95 | 92.92
4.5 | 1.2 | 84.7 | 91.2 | 92.1 | 92.16 | 1.3 | 85.3 | 91.16 | 91.9 | 92.0167 | 1.9 | 86.525 | 91.3 | 91.82 | 92.0816 | 3 | 88.05 | 91.65 | 92.2167 | 92.3167 | 3.8 | 89.2333 | 92.025 | 92.8 | 92.68
4.6 | 1.1 | 82.5 | 90.95 | 91.825 | 91.925 | 1.2 | 83.8333 | 90.9 | 91.775 | 91.7833 | 1.8 | 85.6 | 91.075 | 91.68 | 91.9316 | 2.8 | 87.4667 | 91.4667 | 92.05 | 92.16 | 3.55 | 88.8 | 91.9 | 92.58 | 92.45
4.7 | 1 | 79.3 | 90.8 | 91.575 | 91.775 | 1.15 | 81.8 | 90.7333 | 91.6 | 91.675 | 1.7 | 84.5 | 90.85 | 91.5 | 91.7317 | 2.7 | 87 | 91.18 | 91.8 | 91.9833 | 3.4 | 88.26 | 91.675 | 92.4667 | 92.35
4.8 | 1 | 74.6 | 90.575 | 91.4 | 91.54 | 1.1 | 78.9 | 90.6 | 91.4 | 91.4 | 1.6 | 83.2 | 90.64 | 91.4 | 91.6017 | 2.6 | 86.3333 | 90.975 | 91.6429 | 91.8167 | 3.2 | 87.7667 | 91.5333 | 92.1667 | 92.24
4.9 | 1 | 67.3 | 90.45 | 91.14 | 91.34 | 1 | 74.7 | 90.35 | 91.18 | 91.2667 | 1.5 | 81.3 | 90.5 | 91.1833 | 91.415 | 2.4 | 85.6667 | 90.725 | 91.5 | 91.65 | 3 | 87.525 | 91.275 | 91.9 | 92.1333
5 | 1 | 58.7 | 90.26 | 91 | 91.24 | 1 | 68.5 | 90.12 | 90.9333 | 91 | 1.5 | 79.3 | 90.15 | 91.0167 | 91.1818 | 2.3 | 84.9 | 90.5 | 91.38 | 91.4714 | 2.9 | 87.0667 | 91.1 | 91.8333 | 91.9667
5.1 | 0.9 | 49.4 | 89.94 | 90.86 | 90.98 | 0.95 | 60.5333 | 89.8 | 90.8 | 90.85 | 1.4 | 75.575 | 90.025 | 90.82 | 90.9985 | 2.3 | 83.9333 | 90.28 | 91.18 | 91.275 | 2.9 | 86.6 | 90.78 | 91.62 | 91.8143
5.2 | 0.9 | 40.3 | 89.775 | 90.625 | 90.86 | 0.9 | 51.75 | 89.65 | 90.55 | 90.7 | 1.3 | 70.7667 | 89.86 | 90.58 | 90.8218 | 2.2 | 82.7 | 90.15 | 91 | 91.075 | 2.7 | 85.9 | 90.58 | 91.5 | 91.54
5.3 | 0.8 | 30.9 | 89.5333 | 90.38 | 90.7 | 0.9 | 42 | 89.4 | 90.4 | 90.55 | 1.3 | 64 | 89.6 | 90.38 | 90.4819 | 2.1 | 81.3 | 89.9 | 90.8 | 90.9 | 2.6 | 85.325 | 90.4143 | 91.28 | 91.3667
5.4 | 0.8 | 23.45 | 89.32 | 90.2 | 90.45 | 0.8 | 33.6667 | 89.2 | 90.18 | 90.3 | 1.2 | 56.2333 | 89.3 | 90.2 | 90.3569 | 2 | 79.625 | 89.6833 | 90.5833 | 90.7 | 2.5 | 84.5 | 90.2 | 91.1 | 91.15
5.5 | 0.8 | 17.7667 | 89.08 | 90 | 90.3 | 0.8 | 26.3667 | 89.0167 | 90 | 90.1 | 1.1 | 47.9 | 89.12 | 90.02 | 90.182 | 1.9 | 77.3 | 89.4833 | 90.4167 | 90.5 | 2.4 | 83.7 | 89.9833 | 90.92 | 90.9833
5.6 | 0.8 | 14.2333 | 88.84 | 89.8 | 90.06 | 0.8 | 20 | 88.725 | 89.82 | 89.9 | 1.1 | 39.5 | 88.84 | 89.82 | 89.932 | 1.8 | 74.15 | 89.3 | 90.2571 | 90.2875 | 2.35 | 82.9333 | 89.8 | 90.7 | 90.8
5.7 | 0.7 | 11.7 | 88.6 | 89.6333 | 89.7333 | 0.7 | 15.9 | 88.52 | 89.5 | 89.68 | 1.1 | 31.425 | 88.64 | 89.6 | 89.807 | 1.8 | 69.5 | 89.08 | 90.05 | 90.1571 | 2.2 | 81.7 | 89.55 | 90.56 | 90.64
5.8 | 0.7 | 10.1 | 88.35 | 89.5 | 89.6 | 0.7 | 13 | 88.3667 | 89.36 | 89.48 | 1 | 24.875 | 88.425 | 89.34 | 89.6488 | 1.7 | 63.3 | 88.85 | 89.7833 | 89.9286 | 2.05 | 80.2 | 89.3833 | 90.3833 | 90.4667
5.9 | 0.7 | 8.8 | 88.125 | 89.26 | 89.375 | 0.7 | 11.2667 | 88.14 | 89.14 | 89.3 | 1 | 19.95 | 88.15 | 89.175 | 89.3821 | 1.6 | 56.3 | 88.55 | 89.6167 | 89.7333 | 2 | 78.525 | 89.175 | 90.2167 | 90.26
6 | 0.7 | 8 | 87.9 | 89 | 89.1333 | 0.6 | 10 | 87.88 | 88.975 | 89.075 | 0.9 | 16.6 | 87.94 | 88.94 | 89.1822 | 1.6 | 48.7 | 88.325 | 89.4286 | 89.58 | 1.9 | 76.35 | 89.025 | 90.02 | 90.0833
6.1 | 0.7 | 7.2 | 87.7 | 88.7 | 88.875 | 0.6 | 8.825 | 87.6 | 88.7 | 88.8 | 0.9 | 14.25 | 87.7 | 88.725 | 88.9822 | 1.5 | 41.05 | 88.12 | 89.26 | 89.3833 | 1.8 | 73.2 | 88.6667 | 89.8 | 89.9167
6.2 | 0.6 | 6.7 | 87.3333 | 88.5 | 88.7 | 0.6 | 8 | 87.3 | 88.575 | 88.575 | 0.9 | 12.45 | 87.5 | 88.58 | 88.8422 | 1.4 | 33.7 | 87.9167 | 89.1 | 89.1571 | 1.8 | 68.8 | 88.5 | 89.6 | 89.8
6.3 | 0.6 | 6.3 | 87.1 | 88.35 | 88.56 | 0.6 | 7.4 | 87.1 | 88.35 | 88.425 | 0.8 | 11.2 | 87.25 | 88.35 | 88.5623 | 1.4 | 27.3 | 87.625 | 88.9 | 89.025 | 1.7 | 63.25 | 88.2167 | 89.38 | 89.58
6.4 | 0.6 | 5.9333 | 86.85 | 88.1 | 88.275 | 0.6 | 6.85 | 86.825 | 88.1 | 88.175 | 0.8 | 10.2 | 86.9 | 88.12 | 88.3623 | 1.3 | 22.6 | 87.3667 | 88.7333 | 88.8571 | 1.6 | 56.8 | 87.92 | 89.3 | 89.35
6.5 | 0.6 | 5.65 | 86.7 | 87.94 | 88.025 | 0.6 | 6.45 | 86.6 | 87.925 | 87.94 | 0.8 | 9.4333 | 86.7 | 87.975 | 88.1824 | 1.3 | 19.3 | 87.1667 | 88.4333 | 88.6 | 1.6 | 49.8667 | 87.8 | 89.08 | 89.15
6.6 | 0.6 | 5.4 | 86.35 | 87.675 | 87.8 | 0.6 | 6.05 | 86.35 | 87.78 | 87.8 | 0.7 | 8.8 | 86.42 | 87.725 | 87.8824 | 1.2 | 16.95 | 86.9 | 88.2143 | 88.4333 | 1.5 | 42.1333 | 87.5 | 88.8286 | 89.025
6.7 | 0.6 | 5.2 | 86.1 | 87.425 | 87.625 | 0.5 | 5.7 | 86.15 | 87.475 | 87.575 | 0.7 | 8.2333 | 86.3 | 87.5 | 87.7624 | 1.2 | 15.25 | 86.6 | 88.0143 | 88.1857 | 1.5 | 35.3 | 87.3333 | 88.6429 | 88.8
6.8 | 0.6 | 5.025 | 85.88 | 87.175 | 87.4 | 0.5 | 5.4333 | 85.84 | 87.1667 | 87.35 | 0.7 | 7.7667 | 86 | 87.36 | 87.5658 | 1.1 | 14 | 86.28 | 87.78 | 88.0429 | 1.45 | 29.2 | 87.1429 | 88.575 | 88.58
6.9 | 0.6 | 4.8 | 85.65 | 87 | 87.225 | 0.5 | 5.15 | 85.65 | 87 | 87.2 | 0.7 | 7.4 | 85.725 | 87.05 | 87.2625 | 1.1 | 13 | 86.1 | 87.6 | 87.8667 | 1.4 | 25.2 | 86.8333 | 88.25 | 88.36
7 | 0.6 | 4.6667 | 85.35 | 86.78 | 86.9857 | 0.5 | 4.9 | 85.3333 | 86.8 | 86.9333 | 0.6 | 7.0667 | 85.475 | 86.9 | 87.0626 | 1 | 12.0667 | 85.82 | 87.4429 | 87.5333 | 1.3 | 21.5 | 86.6 | 87.98 | 88.15
7.1 |  | 4.5 | 85.075 | 86.625 | 86.82 |  | 4.75 | 85.06 | 86.575 | 86.6667 | 0.6 | 6.75 | 85.15 | 86.65 | 86.8326 | 1 | 11.5333 | 85.55 | 87.1571 | 87.24 | 1.3 | 19.1 | 86.35 | 87.85 | 87.96
7.2 |  | 4.4 | 84.85 | 86.4 | 86.6 |  | 4.6 | 84.85 | 86.3333 | 86.5 |  | 6.5 | 84.92 | 86.4 | 86.7027 | 1 | 10.8333 | 85.24 | 86.96 | 87.0667 | 1.25 | 17.35 | 85.98 | 87.625 | 87.72
7.3 |  | 4.2 | 84.64 | 86.2 | 86.3 |  | 4.4 | 84.5 | 86.16 | 86.25 |  | 6.225 | 84.6 | 86.2 | 86.4827 | 1 | 10.2 | 85 | 86.7143 | 86.9 | 1.2 | 16 | 85.675 | 87.36 | 87.6
7.4 |  | 4.1 | 84.325 | 85.94 | 86.15 |  | 4.3 | 84.2667 | 85.96 | 86.02 |  | 6 | 84.28 | 86.02 | 86.2228 | 0.9 | 9.8 | 84.8 | 86.5833 | 86.7 | 1.2 | 14.9 | 85.52 | 87.1833 | 87.4
7.5 |  | 4 | 84.1 | 85.8 | 85.925 |  | 4.2 | 84.02 | 85.8 | 85.8 |  | 5.8 | 84.06 | 85.7667 | 86.0028 |  | 9.4667 | 84.4667 | 86.3333 | 86.5 | 1.2 | 13.9 | 85.225 | 87 | 87.2
7.6 |  | 3.9 | 83.85 | 85.45 | 85.66 |  | 4 | 83.8 | 85.5 | 85.575 |  | 5.6 | 83.8 | 85.54 | 85.8161 |  | 9.2 | 84.15 | 86.1571 | 86.24 | 1.2 | 13.2 | 84.9667 | 86.8 | 87.025
7.7 |  | 3.8 | 83.52 | 85.3 | 85.425 |  | 3.9 | 83.48 | 85.225 | 85.44 |  | 5.4333 | 83.52 | 85.325 | 85.6829 |  | 8.8 | 83.9 | 86 | 86.06 |  | 12.6 | 84.66 | 86.5833 | 86.85
7.8 |  | 3.7 | 83.25 | 85 | 85.18 |  | 3.8 | 83.25 | 85 | 85.12 |  | 5.175 | 83.15 | 85.175 | 85.3329 |  | 8.6 | 83.62 | 85.7333 | 85.8714 |  | 12.05 | 84.32 | 86.4286 | 86.525
7.9 |  | 3.6 | 83 | 84.78 | 84.95 |  | 3.7 | 82.94 | 84.7667 | 84.9 |  | 5.0333 | 83.02 | 85 | 85.103 |  | 8.3 | 83.275 | 85.5333 | 85.6333 |  | 11.55 | 84.06 | 86.15 | 86.325
8 |  | 3.5 | 82.65 | 84.56 | 84.7 |  | 3.6 | 82.64 | 84.6 | 84.72 |  | 4.9667 | 82.675 | 84.7 | 84.8997 |  | 8 | 83.06 | 85.3 | 85.3667 |  | 11.2333 | 83.88 | 85.975 | 86.14
8.1 |  | 3.4 | 82.3 | 84.34 | 84.5 |  | 3.5 | 82.3667 | 84.4 | 84.5 |  | 4.8 | 82.36 | 84.5 | 84.783 |  | 7.825 | 82.775 | 85.12 | 85.2 |  | 10.9 | 83.66 | 85.8 | 85.92
8.2 |  | 3.4 | 82.1167 | 84.04 | 84.2 |  | 3.4 | 82.1 | 84.15 | 84.2 |  | 4.6 | 82.05 | 84.2 | 84.4081 |  | 7.55 | 82.54 | 84.8667 | 85 |  | 10.4333 | 83.34 | 85.6 | 85.7333
8.3 |  | 3.3 | 81.78 | 83.9 | 84 |  | 3.3 | 81.66 | 83.8 | 84.02 |  | 4.5 | 81.8 | 84.075 | 84.2581 |  | 7.35 | 82.2 | 84.56 | 84.8333 |  | 10.125 | 83.1 | 85.4 | 85.525
8.4 |  | 3.2 | 81.44 | 83.6 | 83.825 |  | 3.2 | 81.3333 | 83.66 | 83.7 |  | 4.4 | 81.44 | 83.85 | 84.0332 |  | 7.075 | 82 | 84.3833 | 84.6 |  | 9.9 | 82.84 | 85.14 | 85.3
8.5 |  | 3.1 | 81.1 | 83.4 | 83.6667 |  | 3.1 | 81.1 | 83.44 | 83.6 |  | 4.3 | 81.1 | 83.52 | 83.7832 |  | 6.8333 | 81.7 | 84.1 | 84.4333 |  | 9.6667 | 82.5667 | 85 | 85.1
8.6 |  | 3.1 | 80.8 | 83.2 | 83.3 |  | 3.1 | 80.85 | 83.2 | 83.3 |  | 4.2 | 80.8 | 83.4167 | 83.6333 |  | 6.7 | 81.26 | 83.9286 | 84.0667 |  | 9.3 | 82.225 | 84.7857 | 85
8.7 |  | 3 | 80.5 | 82.95 | 83.15 |  | 3 | 80.42 | 83 | 83.08 |  | 4.0667 | 80.48 | 83.1 | 83.3633 |  | 6.6 | 80.975 | 83.7714 | 83.8 |  | 9.1 | 81.8667 | 84.5667 | 84.7
8.8 |  | 2.9 | 80.1333 | 82.65 | 82.875 |  | 2.9667 | 80.08 | 82.675 | 82.775 |  | 3.9 | 80.1 | 82.925 | 83.1034 |  | 6.4 | 80.625 | 83.52 | 83.72 |  | 8.8 | 81.58 | 84.34 | 84.5
8.9 |  | 2.9 | 79.75 | 82.5 | 82.625 |  | 2.9 | 79.7 | 82.4333 | 82.625 |  | 3.8 | 79.82 | 82.675 | 82.9501 |  | 6.2 | 80.4 | 83.3667 | 83.5 |  | 8.6667 | 81.36 | 84.0667 | 84.2
9 |  | 2.8 | 79.22 | 82.175 | 82.375 |  | 2.8 | 79.26 | 82.2 | 82.475 |  | 3.7 | 79.4 | 82.45 | 82.6335 |  | 6.0667 | 80.1 | 83.1 | 83.2667 |  | 8.5 | 81.0333 | 83.9 | 83.9333
9.1 |  | 2.8 | 78.775 | 82.05 | 82.2 |  | 2.7 | 78.85 | 82.06 | 82.15 |  | 3.6 | 79.02 | 82.1 | 82.3835 |  | 5.8667 | 79.65 | 82.8333 | 83 |  | 8.3 | 80.75 | 83.675 | 83.76
9.2 |  | 2.7 | 78.35 | 81.8 | 82.04 |  | 2.65 | 78.375 | 81.84 | 81.95 |  | 3.5 | 78.6 | 81.9333 | 82.1336 |  | 5.7 | 79.275 | 82.7 | 82.8167 |  | 8.1 | 80.45 | 83.4667 | 83.56
9.3 |  | 2.7 | 77.75 | 81.64 | 81.8 |  | 2.6 | 78 | 81.62 | 81.775 |  | 3.4 | 78.25 | 81.725 | 81.9586 |  | 5.6 | 79.04 | 82.4167 | 82.5571 |  | 7.8667 | 80.15 | 83.275 | 83.35
9.4 |  | 2.6 | 76.95 | 81.4 | 81.65 |  | 2.6 | 77.45 | 81.4 | 81.625 |  | 3.4 | 77.85 | 81.5 | 81.7087 |  | 5.4333 | 78.7 | 82.225 | 82.4 |  | 7.7 | 79.8667 | 83.1 | 83.2
9.5 |  | 2.5 | 76.125 | 81.2 | 81.5 |  | 2.5 | 76.775 | 81.2 | 81.375 |  | 3.3 | 77.45 | 81.25 | 81.5637 |  | 5.3333 | 78.4 | 82.0167 | 82.1 |  | 7.5 | 79.58 | 82.86 | 82.9
9.6 |  | 2.5 | 75.1 | 81.1 | 81.3 |  | 2.5 | 76.025 | 81 | 81.16 |  | 3.2667 | 77.05 | 81.12 | 81.3437 |  | 5.2 | 77.9667 | 81.74 | 81.88 |  | 7.35 | 79.175 | 82.6333 | 82.7
9.7 |  | 2.5 | 73.625 | 80.975 | 81.125 |  | 2.4 | 75.36 | 80.82 | 80.95 |  | 3.2 | 76.54 | 80.86 | 81.0838 |  | 5.1 | 77.6 | 81.5333 | 81.7143 |  | 7.175 | 79 | 82.4 | 82.5
9.8 |  | 2.4 | 71.4 | 80.8 | 81 |  | 2.4 | 74.175 | 80.6 | 80.8 |  | 3.1 | 75.96 | 80.65 | 80.8588 |  | 5 | 77.3 | 81.2833 | 81.4333 |  | 7 | 78.64 | 82.2 | 82.275
9.9 |  | 2.4 | 68.15 | 80.6 | 80.85 |  | 2.3 | 72.85 | 80.5 | 80.7 |  | 3 | 75.45 | 80.4 | 80.6839 |  | 4.9 | 76.9333 | 81.1 | 81.2 |  | 6.8667 | 78.35 | 81.9667 | 82
10 |  | 2.4 | 63.175 | 80.52 | 80.62 |  | 2.3 | 71.15 | 80.3 | 80.46 |  | 2.9 | 74.85 | 80.24 | 80.4839 |  | 4.8 | 76.5 | 80.8833 | 80.9429 |  | 6.7 | 77.94 | 81.775 | 81.7833
10.1 |  | 2.3 | 56.6667 | 80.3667 | 80.5 |  | 2.2 | 68.5 | 80.2 | 80.325 |  | 2.8333 | 74.02 | 80.0667 | 80.2839 |  | 4.7 | 76.08 | 80.6333 | 80.68 |  | 6.6 | 77.5 | 81.55 | 81.6
10.2 |  |  | 48.7 | 80.22 | 80.375 |  | 2.2 | 64.425 | 79.98 | 80.2333 |  | 2.8 | 73.2 | 79.8 | 80.084 |  | 4.5333 | 75.76 | 80.4333 | 80.5143 |  | 6.4 | 77.2167 | 81.3 | 81.36
10.3 |  |  | 40.94 | 80.06 | 80.3 |  |  | 59.08 | 79.9 | 80.05 |  | 2.8 | 72.3 | 79.7 | 79.934 |  | 4.425 | 75.325 | 80.18 | 80.2333 |  | 6.4 | 77.02 | 81.1 | 81.125
10.4 |  |  | 33.4 | 79.8833 | 80.1 |  |  | 52.4 | 79.76 | 79.925 |  | 2.7 | 70.9 | 79.52 | 79.784 |  | 4.3 | 74.96 | 79.96 | 80.0833 |  | 6.2333 | 76.5667 | 80.9 | 80.94
10.5 |  |  | 27.0333 | 79.7 | 80 |  |  | 45.55 | 79.6 | 79.82 |  |  | 69.1 | 79.3333 | 79.5841 |  | 4.2333 | 74.575 | 79.775 | 79.9 |  | 6.0667 | 76.2333 | 80.6333 | 80.7667
10.6 |  |  | 21.4 | 79.62 | 79.775 |  |  | 37.9 | 79.36 | 79.6 |  |  | 66.54 | 79.2 | 79.4841 |  | 4.15 | 74.08 | 79.525 | 79.6667 |  | 5.9 | 75.8667 | 80.45 | 80.5
10.7 |  |  | 17.675 | 79.5 | 79.7 |  |  | 30.9 | 79.2333 | 79.475 |  |  | 63.1 | 79 | 79.2841 |  | 4.1 | 73.52 | 79.3167 | 79.4833 |  | 5.7333 | 75.44 | 80.2 | 80.36
10.8 |  |  | 15.2 | 79.375 | 79.58 |  |  | 24.8167 | 79.1 | 79.225 |  |  | 58.2 | 78.9 | 79.0842 |  | 4 | 72.95 | 79.12 | 79.2333 |  | 5.6667 | 75.14 | 80 | 80.14
10.9 |  |  | 13.7 | 79.225 | 79.4 |  |  | 20.2333 | 79 | 79.1 |  |  | 52.1 | 78.7667 | 78.9442 |  | 3.9 | 72.42 | 78.92 | 79 |  | 5.6 | 74.7 | 79.74 | 79.9
11 |  |  | 12.7 | 79.1 | 79.2 |  |  | 17 | 78.8 | 79.025 |  |  | 45.3 | 78.64 | 78.8342 |  |  | 71.78 | 78.65 | 78.7667 |  | 5.5 | 74.42 | 79.4667 | 79.725
11.1 |  |  | 11.8 | 79 | 79.04 |  |  | 15.15 | 78.65 | 78.9 |  |  | 38.22 | 78.5 | 78.6843 |  |  | 71.075 | 78.4714 | 78.6143 |  | 5.3667 | 74.08 | 79.24 | 79.5
11.2 |  |  | 11.2 | 78.75 | 78.9 |  |  | 13.725 | 78.525 | 78.72 |  |  | 31.85 | 78.3 | 78.4843 |  |  | 70.35 | 78.3 | 78.42 |  | 5.25 | 73.6 | 79.12 | 79.26
11.3 |  |  | 10.56 | 78.64 | 78.8 |  |  | 12.6 | 78.4 | 78.6 |  |  | 26.4 | 78.0833 | 78.3343 |  |  | 69.225 | 78.0667 | 78.28 |  | 5.0333 | 73.2 | 78.88 | 79.05
11.4 |  |  | 10.2 | 78.5 | 78.7 |  |  | 11.8 | 78.3 | 78.4333 |  |  | 21.95 | 78 | 78.2094 |  |  | 67.92 | 77.95 | 78.0667 |  | 4.9 | 72.85 | 78.6 | 78.84
11.5 |  |  | 9.85 | 78.3667 | 78.55 |  |  | 11.2 | 78.1 | 78.3 |  |  | 19.1 | 77.76 | 77.9844 |  |  | 66.3 | 77.6667 | 77.8833 |  |  | 72.3 | 78.4 | 78.62
11.6 |  |  | 9.575 | 78.24 | 78.4 |  |  | 10.65 | 77.975 | 78.1 |  |  | 16.95 | 77.65 | 77.7844 |  |  | 63.96 | 77.5 | 77.8 |  |  | 72.025 | 78.15 | 78.4
11.7 |  |  | 9.34 | 78.04 | 78.275 |  |  | 10.15 | 77.825 | 77.95 |  |  | 15.525 | 77.5 | 77.6845 |  |  | 60.55 | 77.3286 | 77.66 |  |  | 71.45 | 78 | 78.2
11.8 |  |  | 9.1 | 77.9 | 78.1 |  |  | 9.8 | 77.74 | 77.8 |  |  | 14.4 | 77.3 | 77.5845 |  |  | 56.1 | 77.2 | 77.4667 |  |  | 70.875 | 77.85 | 77.9667
11.9 |  |  | 8.85 | 77.75 | 77.9 |  |  | 9.5 | 77.54 | 77.6 |  |  | 13.56 | 77.1 | 77.3845 |  |  | 51 | 76.98 | 77.2286 |  |  | 70.425 | 77.725 | 77.7
12 |  |  | 8.7 | 77.62 | 77.775 |  |  | 9.175 | 77.42 | 77.5 |  |  | 12.9 | 76.94 | 77.1846 |  |  | 44.725 | 76.8 | 77 |  |  | 69.75 | 77.4 | 77.5
12.1 |  |  | 8.475 | 77.4667 | 77.6 |  |  | 8.96 | 77.2 | 77.325 |  |  | 12.3 | 76.8 | 76.9846 |  |  | 38.925 | 76.6333 | 76.9 |  |  | 68.9667 | 77.22 | 77.34
12.2 |  |  | 8.3 | 77.325 | 77.425 |  |  | 8.74 | 77 | 77.2 |  |  | 11.82 | 76.68 | 76.8513 |  |  | 33.3 | 76.4667 | 76.7 |  |  | 68.2 | 77.0333 | 77.2
12.3 |  |  | 8.175 | 77.15 | 77.3 |  |  | 8.5 | 76.9 | 77.06 |  |  | 11.425 | 76.4667 | 76.6847 |  |  | 28.275 | 76.3 | 76.5 |  |  | 67.1333 | 76.94 | 77
12.4 |  |  | 8 | 77 | 77.1 |  |  | 8.34 | 76.8 | 76.9 |  |  | 11.06 | 76.25 | 76.5597 |  |  | 24.5 | 76.2167 | 76.3167 |  |  | 65.7667 | 76.65 | 76.8
12.5 |  |  | 7.8 | 76.85 | 76.9 |  |  | 8.175 | 76.6 | 76.66 |  |  | 10.7 | 76.1 | 76.3847 |  |  | 21.55 | 76.02 | 76.1667 |  |  | 64.2 | 76.52 | 76.575
12.6 |  |  | 7.7 | 76.575 | 76.8 |  |  | 8 | 76.425 | 76.5 |  |  | 10.46 | 75.925 | 76.2447 |  |  | 19.6 | 75.8714 | 75.98 |  |  | 61.7 | 76.325 | 76.35
12.7 |  |  | 7.54 | 76.4 | 76.6333 |  |  | 7.8 | 76.28 | 76.3 |  |  | 10.18 | 75.7 | 76.0248 |  |  | 18.2667 | 75.7167 | 75.7857 |  |  | 58.825 | 76.1 | 76.12
12.8 |  |  | 7.5 | 76.3 | 76.4667 |  |  | 7.7 | 76.1 | 76.175 |  |  | 9.925 | 75.55 | 75.8048 |  |  | 17.2 | 75.4714 | 75.6 |  |  | 54.58 | 75.95 | 76.0333
12.9 |  |  | 7.3 | 76.1 | 76.275 |  |  | 7.5 | 75.975 | 76.025 |  |  | 9.7667 | 75.42 | 75.6849 |  |  | 16.225 | 75.3 | 75.46 |  |  | 50.04 | 75.7 | 75.85
13 |  |  | 7.225 | 75.96 | 76.125 |  |  | 7.4 | 75.74 | 75.875 |  |  | 9.5 | 75.26 | 75.5249 |  |  | 15.5333 | 75.1333 | 75.18 |  |  | 44.88 | 75.5571 | 75.74
13.1 |  |  | 7.1 | 75.8 | 76 |  |  | 7.275 | 75.6 | 75.7 |  |  | 9.275 | 75.1 | 75.3349 |  |  | 14.975 | 74.92 | 75 |  |  | 39.45 | 75.4 | 75.6
13.2 |  |  | 6.95 | 75.6 | 75.8 |  |  | 7.125 | 75.45 | 75.52 |  |  | 9.1 | 74.9 | 75.085 |  |  | 14.4 | 74.5857 | 74.8333 |  |  | 34.44 | 75.2 | 75.36
13.3 |  |  | 6.9 | 75.5 | 75.58 |  |  | 7 | 75.3 | 75.28 |  |  | 8.9 | 74.7333 | 74.885 |  |  | 14 | 74.42 | 74.6571 |  |  | 29.98 | 74.94 | 75.0833
13.4 |  |  | 6.8 | 75.28 | 75.4 |  |  | 6.875 | 74.975 | 75.05 |  |  | 8.78 | 74.5 | 74.765 |  |  | 13.7 | 74.3 | 74.5 |  |  | 26.4 | 74.7667 | 75
13.5 |  |  | 6.65 | 75.1 | 75.3 |  |  | 6.725 | 74.8 | 74.9 |  |  | 8.6 | 74.3 | 74.5451 |  |  | 13.375 | 74.1667 | 74.2167 |  |  | 23.75 | 74.65 | 74.75
13.6 |  |  | 6.5 | 74.9 | 74.975 |  |  | 6.68 | 74.65 | 74.725 |  |  | 8.4 | 74.15 | 74.3851 |  |  | 13.1 | 73.9167 | 74.06 |  |  | 21.66 | 74.38 | 74.52
13.7 |  |  | 6.4 | 74.75 | 74.825 |  |  | 6.5 | 74.4667 | 74.6 |  |  | 8.3 | 73.98 | 74.1852 |  |  | 12.875 | 73.7833 | 73.8 |  |  | 20.2 | 74.2333 | 74.32
13.8 |  |  | 6.325 | 74.4 | 74.65 |  |  | 6.4 | 74.275 | 74.4 |  |  | 8.1 | 73.65 | 73.9852 |  |  | 12.52 | 73.575 | 73.6167 |  |  | 19.1 | 74 | 74.12
13.9 |  |  | 6.2 | 74.2333 | 74.38 |  |  | 6.28 | 74.04 | 74.16 |  |  | 7.98 | 73.48 | 73.7852 |  |  | 12.25 | 73.28 | 73.4143 |  |  | 18.3667 | 73.8 | 74.02
14 |  |  | 6.18 | 73.98 | 74.16 |  |  | 6.16 | 73.84 | 73.95 |  |  | 7.8 | 73.2667 | 73.5353 |  |  | 12 | 73.1143 | 73.2 |  |  | 17.5333 | 73.68 | 73.7833
14.1 |  |  | 6.025 | 73.6667 | 73.84 |  |  | 6.1 | 73.6 | 73.7 |  |  | 7.7 | 73.1 | 73.2853 |  |  | 11.8 | 72.9857 | 73.0333 |  |  | 16.9 | 73.4 | 73.6
14.2 |  |  | 5.975 | 73.5 | 73.6 |  |  | 6 | 73.3 | 73.425 |  |  | 7.525 | 72.9 | 73.1854 |  |  | 11.475 | 72.8 | 72.8 |  |  | 16.4333 | 73.3 | 73.38
14.3 |  |  | 5.85 | 73.02 | 73.2 |  |  | 5.9 | 73.05 | 73.175 |  |  | 7.4 | 72.7 | 72.9854 |  |  | 11.28 | 72.5857 | 72.68 |  |  | 16.0333 | 73 | 73.15
14.4 |  |  | 5.8 | 72.675 | 72.8 |  |  | 5.8 | 72.8 | 72.9 |  |  | 7.3 | 72.45 | 72.6605 |  |  | 11.1 | 72.36 | 72.4 |  |  | 15.64 | 72.8333 | 73
14.5 |  |  | 5.7 | 72.1 | 72.2 |  |  | 5.725 | 72.36 | 72.5 |  |  | 7.1 | 72.175 | 72.4855 |  |  | 10.85 | 72.25 | 72.2429 |  |  | 15.38 | 72.68 | 72.8167
14.6 |  |  | 5.675 | 71.5 | 71.65 |  |  | 5.675 | 72 | 72.2333 |  |  | 7.06 | 71.94 | 72.1856 |  |  | 10.72 | 71.88 | 71.98 |  |  | 14.9 | 72.5143 | 72.68
14.7 |  |  | 5.54 | 70.54 | 70.8 |  |  | 5.6 | 71.6333 | 71.75 |  |  | 6.92 | 71.75 | 71.8856 |  |  | 10.52 | 71.6571 | 71.8 |  |  | 14.7 | 72.25 | 72.4833
14.8 |  |  | 5.5 | 69.425 | 69.7 |  |  | 5.5 | 71.1 | 71.25 |  |  | 6.8167 | 71.4 | 71.6357 |  |  | 10.275 | 71.4667 | 71.58 |  |  | 14.3 | 72.05 | 72.3
14.9 |  |  | 5.4 | 67.8 | 67.9 |  |  | 5.4 | 70.46 | 70.7 |  |  | 6.65 | 71.175 | 71.3607 |  |  | 10.14 | 71.2167 | 71.3833 |  |  | 14.075 | 71.9 | 72.08
15 |  |  | 5.325 | 65.05 | 65.36 |  |  | 5.3 | 69.725 | 69.85 |  |  | 6.6 | 70.8 | 71.1108 |  |  | 9.975 | 71.0333 | 71.1 |  |  | 13.84 | 71.72 | 71.8
15.1 |  |  | 5.3 | 61.2 | 61.4 |  |  | 5.2 | 68.8 | 68.98 |  |  | 6.5 | 70.55 | 70.7358 |  |  | 9.76 | 70.75 | 70.88 |  |  | 13.6333 | 71.52 | 71.5667
15.2 |  |  | 5.12 | 55.7333 | 56.08 |  |  | 5.2 | 67.6 | 67.6333 |  |  | 6.38 | 70.2 | 70.4109 |  |  | 9.625 | 70.6 | 70.6333 |  |  | 13.4333 | 71.3 | 71.4
15.3 |  |  |  | 48.95 | 49.6 |  |  | 5.1 | 65.8 | 66 |  |  | 6.3 | 69.7 | 70.0193 |  |  | 9.4667 | 70.34 | 70.3833 |  |  | 13.15 | 71.05 | 71.2
15.4 |  |  |  | 42.2 | 42.6 |  |  | 5.1 | 63.2 | 63.2333 |  |  | 6.2 | 69.3 | 69.6361 |  |  | 9.2167 | 70.0167 | 70.2 |  |  | 12.9 | 70.82 | 71.025
15.5 |  |  |  | 35.05 | 35.95 |  |  |  | 59.2 | 59.3 |  |  | 6.1 | 68.8667 | 69.0862 |  |  | 9.1 | 69.84 | 69.8833 |  |  | 12.675 | 70.6 | 70.8
15.6 |  |  |  | 29.05 | 30 |  |  |  | 53.65 | 54.2 |  |  | 6 | 68.2 | 68.553 |  |  | 8.9 | 69.6 | 69.7 |  |  | 12.38 | 70.4 | 70.55
15.7 |  |  |  | 23.7 | 24.85 |  |  |  | 47.4 | 47.875 |  |  | 5.925 | 67.6 | 67.8114 |  |  | 8.8 | 69.3667 | 69.5 |  |  | 12.2 | 70.15 | 70.25
15.8 |  |  |  | 20.1 | 21.3 |  |  |  | 41.05 | 41.875 |  |  | 5.8 | 66.65 | 66.8866 |  |  | 8.65 | 69.1 | 69.1667 |  |  | 12 | 69.9333 | 70.1
15.9 |  |  |  | 17.875 | 18.9 |  |  |  | 34.8 | 35.5 |  |  | 5.8 | 65.55 | 65.8868 |  |  | 8.5333 | 68.8571 | 68.9714 |  |  | 11.8 | 69.65 | 69.8667
16 |  |  |  | 16.1667 | 17.38 |  |  |  | 28.5667 | 29.75 |  |  |  | 64 | 64.4371 |  |  | 8.4 | 68.6143 | 68.7 |  |  | 11.6167 | 69.425 | 69.7
16.1 |  |  |  | 15.04 | 16.3 |  |  |  | 23.94 | 24.6 |  |  |  | 61.92 | 62.2875 |  |  | 8.3 | 68.4 | 68.4571 |  |  | 11.5 | 69.22 | 69.35
16.2 |  |  |  | 14.35 | 15.62 |  |  |  | 20.425 | 21.275 |  |  |  | 58.85 | 59.4548 |  |  | 8.15 | 67.95 | 68.16 |  |  | 11.4 | 69 | 69.1
16.3 |  |  |  | 13.75 | 14.9667 |  |  |  | 18 | 19.1 |  |  |  | 54.75 | 55.3689 |  |  | 8.05 | 67.72 | 67.92 |  |  | 11.1 | 68.85 | 68.975
16.4 |  |  |  | 13.25 | 14.55 |  |  |  | 16.4 | 17.5667 |  |  |  | 49.35 | 49.9234 |  |  | 7.9 | 67.42 | 67.5667 |  |  | 10.925 | 68.6 | 68.76
16.5 |  |  |  | 12.9 | 14.18 |  |  |  | 15.3 | 16.5 |  |  |  | 43.3 | 44.3911 |  |  | 7.8 | 67.1333 | 67.2 |  |  | 10.8 | 68.3 | 68.5
16.6 |  |  |  | 12.625 | 13.9 |  |  |  | 14.45 | 15.6 |  |  |  | 37.64 | 38.3923 |  |  | 7.6 | 66.8 | 66.8333 |  |  | 10.675 | 68.175 | 68.25
16.7 |  |  |  | 12.3 | 13.6 |  |  |  | 13.85 | 15.12 |  |  |  | 31.975 | 32.7934 |  |  | 7.5 | 66.34 | 66.5333 |  |  | 10.525 | 67.9 | 68
16.8 |  |  |  | 12.125 | 13.4 |  |  |  | 13.3667 | 14.55 |  |  |  | 27.05 | 27.9944 |  |  | 7.425 | 65.94 | 66.12 |  |  | 10.3 | 67.6333 | 67.7667
16.9 |  |  |  | 11.9 | 13.2 |  |  |  | 12.96 | 14.2 |  |  |  | 23.06 | 24.0951 |  |  | 7.2 | 65.4429 | 65.62 |  |  | 10.2 | 67.375 | 67.55
17 |  |  |  | 11.75 | 13 |  |  |  | 12.64 | 13.85 |  |  |  | 20.4 | 21.5757 |  |  |  | 64.98 | 65.08 |  |  | 10.025 | 67.1 | 67.25
17.1 |  |  |  | 11.575 | 12.8333 |  |  |  | 12.3 | 13.6 |  |  |  | 18.6 | 19.796 |  |  |  | 64.36 | 64.5429 |  |  | 9.8833 | 66.875 | 66.95
17.2 |  |  |  | 11.45 | 12.675 |  |  |  | 12.1 | 13.3 |  |  |  | 17.4 | 18.5963 |  |  |  | 63.68 | 63.8 |  |  | 9.7167 | 66.55 | 66.675
17.3 |  |  |  | 11.3 | 12.5 |  |  |  | 11.94 | 13.05 |  |  |  | 16.4 | 17.5465 |  |  |  | 62.8333 | 63 |  |  | 9.52 | 66.32 | 66.4
17.4 |  |  |  | 11.2 | 12.4 |  |  |  | 11.675 | 12.9 |  |  |  | 15.7 | 16.8966 |  |  |  | 61.7 | 61.96 |  |  | 9.4 | 66.1 | 66.2
17.5 |  |  |  | 11.05 | 12.3 |  |  |  | 11.525 | 12.75 |  |  |  | 15.08 | 16.2967 |  |  |  | 60.4 | 60.6167 |  |  | 9.24 | 65.8333 | 65.95
17.6 |  |  |  | 10.92 | 12.2 |  |  |  | 11.3 | 12.6 |  |  |  | 14.55 | 15.7635 |  |  |  | 58.3667 | 58.7 |  |  | 9.1 | 65.475 | 65.65
17.7 |  |  |  | 10.8 | 12.0333 |  |  |  | 11.1 | 12.3 |  |  |  | 14.2 | 15.4469 |  |  |  | 55.6143 | 56 |  |  | 9 | 65.225 | 65.34
17.8 |  |  |  | 10.76 | 11.9 |  |  |  | 11 | 12.2 |  |  |  | 13.85 | 14.997 |  |  |  | 51.9 | 52.5 |  |  | 8.9 | 64.9 | 65
17.9 |  |  |  | 10.625 | 11.9 |  |  |  | 10.9 | 12.08 |  |  |  | 13.54 | 14.817 |  |  |  | 47.45 | 48 |  |  | 8.65 | 64.5 | 64.775
18 |  |  |  | 10.575 | 11.76 |  |  |  | 10.75 | 11.98 |  |  |  | 13.3 | 14.5221 |  |  |  | 42.2667 | 43.1667 |  |  |  | 64.22 | 64.4
18.1 |  |  |  | 10.5 | 11.7 |  |  |  | 10.6 | 11.8 |  |  |  | 13.125 | 14.2971 |  |  |  | 37.4 | 38.14 |  |  |  | 63.8333 | 64.05
18.2 |  |  |  | 10.35 | 11.54 |  |  |  | 10.5 | 11.7 |  |  |  | 12.8 | 14.0972 |  |  |  | 32.7 | 33.5 |  |  |  | 63.52 | 63.725
18.3 |  |  |  | 10.22 | 11.5 |  |  |  | 10.4 | 11.6 |  |  |  | 12.7 | 13.9472 |  |  |  | 28.4 | 29.35 |  |  |  | 63.1 | 63.25
18.4 |  |  |  | 10.2 | 11.4 |  |  |  | 10.3 | 11.5 |  |  |  | 12.4 | 13.6973 |  |  |  | 25.1167 | 26.175 |  |  |  | 62.6667 | 62.8
18.5 |  |  |  | 10.1 | 11.3 |  |  |  | 10.2 | 11.36 |  |  |  | 12.22 | 13.4473 |  |  |  | 22.64 | 23.8 |  |  |  | 62.14 | 62.2
18.6 |  |  |  | 10.0667 | 11.2 |  |  |  | 10.12 | 11.3 |  |  |  | 12 | 13.3723 |  |  |  | 21 | 22.05 |  |  |  | 61.5333 | 61.6
18.7 |  |  |  | 9.975 | 11.1 |  |  |  | 10 | 11.15 |  |  |  | 11.9 | 13.1224 |  |  |  | 19.6167 | 20.9 |  |  |  | 60.7 | 60.875
18.8 |  |  |  | 9.8 | 11.05 |  |  |  | 9.9 | 11.1 |  |  |  | 11.75 | 12.9974 |  |  |  | 18.72 | 19.9 |  |  |  | 59.975 | 60
18.9 |  |  |  | 9.75 | 11 |  |  |  | 9.8 | 11 |  |  |  | 11.6 | 12.7974 |  |  |  | 18.04 | 19.2333 |  |  |  | 58.9 | 59.1
19 |  |  |  | 9.7 | 10.8333 |  |  |  | 9.7 | 10.875 |  |  |  | 11.425 | 12.6975 |  |  |  | 17.42 | 18.6 |  |  |  | 57.5 | 57.7
19.1 |  |  |  | 9.6 | 10.775 |  |  |  | 9.7 | 10.8 |  |  |  | 11.3 | 12.5475 |  |  |  | 16.86 | 18.15 |  |  |  | 55.6 | 55.74
19.2 |  |  |  | 9.5 | 10.7 |  |  |  | 9.575 | 10.7 |  |  |  | 11.15 | 12.3975 |  |  |  | 16.5286 | 17.64 |  |  |  | 52.925 | 53.325
19.3 |  |  |  | 9.4 | 10.6 |  |  |  | 9.5 | 10.6 |  |  |  | 11.02 | 12.2975 |  |  |  | 16.1333 | 17.3333 |  |  |  | 49.675 | 50.1
19.4 |  |  |  | 9.3 | 10.5 |  |  |  | 9.4 | 10.55 |  |  |  | 10.9 | 12.0976 |  |  |  | 15.7833 | 16.98 |  |  |  | 45.9 | 46.125
19.5 |  |  |  | 9.25 | 10.4 |  |  |  | 9.3 | 10.5 |  |  |  | 10.8 | 11.9976 |  |  |  | 15.5333 | 16.7286 |  |  |  | 41.25 | 41.95
19.6 |  |  |  | 9.1 | 10.3 |  |  |  | 9.2 | 10.4 |  |  |  | 10.7 | 11.8976 |  |  |  | 15.2667 | 16.48 |  |  |  | 36.7 | 37.86
19.7 |  |  |  | 9 | 10.26 |  |  |  | 9.1 | 10.28 |  |  |  | 10.525 | 11.7976 |  |  |  | 15 | 16.2333 |  |  |  | 32.7667 | 33.6
19.8 |  |  |  | 9 | 10.14 |  |  |  | 9 | 10.2 |  |  |  | 10.4 | 11.6477 |  |  |  | 14.8 | 16 |  |  |  | 29.1 | 30.1667
19.9 |  |  |  | 8.9 | 10 |  |  |  | 8.9333 | 10.1 |  |  |  | 10.34 | 11.4977 |  |  |  | 14.5167 | 15.7833 |  |  |  | 26.375 | 27.32
20 |  |  |  | 8.8 | 9.925 |  |  |  | 8.9 | 10 |  |  |  | 10.22 | 11.3977 |  |  |  | 14.3333 | 15.6 |  |  |  | 24.08 | 25.25
20.1 |  |  |  | 8.68 | 9.9 |  |  |  | 8.8 | 9.9 |  |  |  | 10.1 | 11.2977 |  |  |  | 14.12 | 15.42 |  |  |  | 22.4 | 23.6
20.2 |  |  |  | 8.55 | 9.725 |  |  |  | 8.7 | 9.8 |  |  |  | 10.02 | 11.2178 |  |  |  | 13.9 | 15.2 |  |  |  | 21.2 | 22.425
20.3 |  |  |  | 8.5 | 9.6 |  |  |  | 8.6 | 9.75 |  |  |  | 9.9 | 11.0978 |  |  |  | 13.74 | 15 |  |  |  | 20.35 | 21.54
20.4 | [8 empty] | 8.54 | 9.6 |  |  |  | 9.8 | 10.9978 |  |  |  | 13.6 | 14.8167 |  |  |  | 19.85 | 20.88
20.5 | [8 empty] | 8.4 | 9.6 |  |  |  | 9.7 | 10.8978 |  |  |  | 13.4 | 14.6333 |  |  |  | 19.1333 | 20.35
20.6 | [8 empty] | 8.3 | 9.5 |  |  |  | 9.6 | 10.7978 |  |  |  | 13.2 | 14.5 |  |  |  | 18.575 | 19.9
20.7 | [8 empty] | 8.24 | 9.325 |  |  |  | 9.5 | 10.6979 |  |  |  | 13.1 | 14.2667 |  |  |  | 18.15 | 19.525
20.8 | [13 empty] | 9.5 | 10.5979 |  |  |  | 12.8333 | 14.1 |  |  |  | 17.8 | 19.14
20.9 | [13 empty] | 9.3 | 10.4979 |  |  |  | 12.74 | 13.9 |  |  |  | 17.45 | 18.86
21 | [13 empty] | 9.2 | 10.3979 |  |  |  | 12.5667 | 13.8 |  |  |  | 17.1 | 18.475
21.1 | [13 empty] | 9.1333 | 10.2979 |  |  |  | 12.4 | 13.6667 |  |  |  | 16.825 | 18.2
21.2 | [13 empty] | 9 | 10.198 |  |  |  | 12.2286 | 13.4833 |  |  |  | 16.7 | 17.88
21.3 | [13 empty] | 8.9 | 10.098 |  |  |  | 12.16 | 13.3 |  |  |  | 16.475 | 17.66
21.4 | [13 empty] | 8.85 | 9.998 |  |  |  | 11.96 | 13.22 |  |  |  | 16.2 | 17.45
21.5 | [18 empty] | 11.8 | 13 |  |  |  | 15.95 | 17.275
21.6 | [18 empty] | 11.68 | 12.9167 |  |  |  | 15.72 | 17
21.7 | [18 empty] | 11.6 | 12.8 |  |  |  | 15.5 | 16.85
21.8 | [18 empty] | 11.42 | 12.66 |  |  |  | 15.325 | 16.7
21.9 | [18 empty] | 11.3429 | 12.5 |  |  |  | 15.2 | 16.4667
22 | [18 empty] | 11.2 | 12.4167 |  |  |  | 14.98 | 16.25
22.1 | [18 empty] | 11.06 | 12.3 |  |  |  | 14.8 | 16.1
22.2 | [18 empty] | 10.96 | 12.15 |  |  |  | 14.7 | 15.85
22.3 | [18 empty] | 10.86 | 12.0667 |  |  |  | 14.475 | 15.7
22.4 | [18 empty] | 10.78 | 11.9 |  |  |  | 14.35 | 15.525
22.5 | [18 empty] | 10.5667 | 11.8 |  |  |  | 14.175 | 15.4
22.6 | [18 empty] | 10.4833 | 11.7 |  |  |  | 13.9667 | 15.2167
22.7 | [18 empty] | 10.4 | 11.54 |  |  |  | 13.86 | 15
22.8 | [18 empty] | 10.2333 | 11.4 |  |  |  | 13.7 | 14.85
22.9 | [18 empty] | 10.1 | 11.3 |  |  |  | 13.52 | 14.7

## Wedge 30 Longitudinal Profiles
Longitudinal Profiles
Wedge 30
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2.4 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 14.2
-30.8 |  |  |  |  | 14.3286
-30.7 |  |  |  |  | 14.4714
-30.6 |  |  |  |  | 14.6
-30.5 |  |  |  |  | 14.725
-30.4 |  |  |  |  | 14.85
-30.3 |  |  |  |  | 15.05
-30.2 |  |  |  |  | 15.2
-30.1 |  |  |  |  | 15.325
-30 |  |  |  |  | 15.5
-29.9 |  |  |  |  | 15.675
-29.8 |  |  |  |  | 15.8
-29.7 |  |  |  |  | 15.9375
-29.6 |  |  |  |  | 16.125
-29.5 |  |  |  |  | 16.2875
-29.4 |  |  |  |  | 16.425
-29.3 |  |  |  |  | 16.6375
-29.2 |  |  |  |  | 16.8111
-29.1 |  |  |  |  | 17.05
-29 |  |  |  |  | 17.2
-28.9 |  |  |  | 12.45 | 17.3286
-28.8 |  |  |  | 12.6 | 17.6
-28.7 |  |  |  | 12.7 | 17.7333
-28.6 |  |  |  | 12.86 | 18
-28.5 |  |  |  | 12.94 | 18.1571
-28.4 |  |  |  | 13.1 | 18.375
-28.3 |  |  |  | 13.2 | 18.68
-28.2 |  |  |  | 13.3429 | 18.8429
-28.1 |  |  |  | 13.4857 | 19.05
-28 |  |  |  | 13.625 | 19.3
-27.9 |  |  |  | 13.8375 | 19.5667
-27.8 |  |  |  | 13.9714 | 19.825
-27.7 |  |  |  | 14.06 | 20.075
-27.6 |  |  |  | 14.2333 | 20.3429
-27.5 |  |  |  | 14.375 | 20.6333
-27.4 |  |  |  | 14.5167 | 20.9429
-27.3 |  |  |  | 14.7 | 21.3
-27.2 |  |  |  | 14.8833 | 21.7571
-27.1 |  |  |  | 15.0375 | 22.1667
-27 |  |  |  | 15.225 | 22.75
-26.9 |  |  | 10.7 | 15.475 | 23.3875
-26.8 |  |  | 10.7667 | 15.7125 | 24.1167
-26.7 |  |  | 10.92 | 15.875 | 25.1
-26.6 |  |  | 11 | 16.0714 | 26.26
-26.5 |  |  | 11.1 | 16.2333 | 27.8
-26.4 |  |  | 11.2 | 16.45 | 29.9
-26.3 |  |  | 11.3667 | 16.7 | 32.5
-26.2 |  |  | 11.4143 | 16.8857 | 35.8
-26.1 |  |  | 11.5667 | 17.1 | 39.3
-26 |  |  | 11.6125 | 17.38 | 43.5
-25.9 |  | 9.8 | 11.75 | 17.5571 | 47.8125
-25.8 |  | 9.9 | 11.9 | 17.825 | 52.4375
-25.7 |  | 9.9857 | 11.9857 | 18.1625 | 56.8
-25.6 |  | 10.1 | 12.1 | 18.44 | 60.65
-25.5 |  | 10.1571 | 12.2167 | 18.725 | 63.9
-25.4 | 9.9 | 10.2875 | 12.3714 | 19.0125 | 66.45
-25.3 | 9.9429 | 10.35 | 12.5 | 19.45 | 68.8
-25.2 | 10 | 10.4 | 12.7 | 19.7286 | 70.325
-25.1 | 10.1286 | 10.5429 | 12.7857 | 20.2 | 71.75
-25 | 10.2 | 10.6 | 12.9286 | 20.8333 | 72.8
-24.9 | 10.3 | 10.7 | 13.1429 | 21.46 | 73.5333
-24.8 | 10.4 | 10.7857 | 13.3286 | 22.2833 | 74.4
-24.7 | 10.5 | 10.9 | 13.4625 | 23.3 | 74.9667
-24.6 | 10.52 | 11.0167 | 13.675 | 24.6333 | 75.6333
-24.5 | 10.7 | 11.1 | 13.8375 | 26.5 | 76.05
-24.4 | 10.7 | 11.15 | 14.0667 | 28.84 | 76.4375
-24.3 | 10.8125 | 11.3 | 14.2375 | 32.0333 | 76.8
-24.2 | 10.9 | 11.4 | 14.3714 | 35.5 | 77.1714
-24.1 | 11.0167 | 11.5 | 14.6333 | 39.9429 | 77.6111
-24 | 11.1 | 11.65 | 14.8714 | 45.48 | 77.9
-23.9 | 11.1333 | 11.8 | 15.05 | 50.6125 | 78.25
-23.8 | 11.275 | 11.875 | 15.2667 | 55.85 | 78.56
-23.7 | 11.3429 | 12.0167 | 15.5286 | 60.7111 | 78.825
-23.6 | 11.4 | 12.1833 | 15.8333 | 65.0571 | 79.05
-23.5 | 11.5667 | 12.2571 | 16.1286 | 68.6889 | 79.35
-23.4 | 11.7 | 12.4167 | 16.4333 | 71.525 | 79.5778
-23.3 | 11.7667 | 12.6 | 16.7333 | 73.85 | 79.8375
-23.2 | 11.9 | 12.775 | 17.15 | 75.7 | 80.15
-23.1 | 12.02 | 12.9714 | 17.7 | 77 | 80.425
-23 | 12.2 | 13.1286 | 18.2 | 77.9333 | 80.6375
-22.9 | 12.3429 | 13.3571 | 18.85 | 78.7571 | 80.8429
-22.8 | 12.4857 | 13.5 | 19.8 | 79.3857 | 81.12
-22.7 | 12.6333 | 13.7571 | 20.8556 | 80.0375 | 81.35
-22.6 | 12.8 | 13.9 | 22.4 | 80.55 | 81.6
-22.5 | 12.8667 | 14.2286 | 24.2333 | 81.0286 | 81.825
-22.4 | 13.125 | 14.4429 | 27.1 | 81.3714 | 82
-22.3 | 13.2667 | 14.8167 | 30.8 | 81.9125 | 82.2714
-22.2 | 13.44 | 15.15 | 35.65 | 82.1625 | 82.4333
-22.1 | 13.6286 | 15.6286 | 41.325 | 82.4625 | 82.6167
-22 | 13.8667 | 16.1 | 47.5 | 82.8 | 82.7833
-21.9 | 14.1 | 16.775 | 54.1571 | 83.15 | 83.0333
-21.8 | 14.3333 | 17.6 | 60.8167 | 83.4 | 83.2
-21.7 | 14.6667 | 18.7429 | 66.7333 | 83.65 | 83.42
-21.6 | 15.08 | 20.5 | 72.05 | 83.9667 | 83.6286
-21.5 | 15.55 | 22.7286 | 76.025 | 84.2714 | 83.7833
-21.4 | 16.18 | 25.7 | 79.05 | 84.5571 | 84.05
-21.3 | 17 | 29.85 | 81.3 | 84.7625 | 84.2222
-21.2 | 17.98 | 35.2 | 82.9 | 85 | 84.4286
-21.1 | 19.5 | 41.2571 | 84.1 | 85.2429 | 84.5
-21 | 21.65 | 48.2857 | 85 | 85.4714 | 84.6571
-20.9 | 24.8667 | 56.2 | 85.8143 | 85.6 | 84.8714
-20.8 | 29.2 | 63.44 | 86.5 | 85.85 | 85
-20.7 | 35 | 70.4 | 87.0667 | 86.1143 | 85.225
-20.6 | 41.3286 | 76.0667 | 87.5429 | 86.3333 | 85.3444
-20.5 | 48.9667 | 80.3714 | 87.9429 | 86.5 | 85.4833
-20.4 | 56.88 | 83.55 | 88.3333 | 86.76 | 85.7286
-20.3 | 65.3 | 85.7 | 88.7125 | 86.9333 | 85.8833
-20.2 | 72.1143 | 87.25 | 88.95 | 87.1167 | 86.0125
-20.1 | 77.9714 | 88.5 | 89.3 | 87.2 | 86.1857
-20 | 82.8125 | 89.4 | 89.4857 | 87.4167 | 86.4
-19.9 | 85.96 | 90.14 | 89.7 | 87.6 | 86.4333
-19.8 | 88.2714 | 90.66 | 90 | 87.7429 | 86.5
-19.7 | 89.8333 | 91.1 | 90.3 | 87.8857 | 86.66
-19.6 | 90.96 | 91.46 | 90.3889 | 88.1125 | 86.9111
-19.5 | 91.6667 | 91.8333 | 90.6429 | 88.2375 | 87.05
-19.4 | 92.2857 | 92.1571 | 90.84 | 88.425 | 87.28
-19.3 | 92.7333 | 92.4333 | 90.9875 | 88.5875 | 87.325
-19.2 | 93.2 | 92.7 | 91.2167 | 88.7286 | 87.4571
-19.1 | 93.5667 | 92.85 | 91.3625 | 88.8714 | 87.6
-19 | 93.82 | 93.1143 | 91.575 | 89 | 87.7
-18.9 | 94.0571 | 93.3 | 91.7375 | 89.0333 | 87.8571
-18.8 | 94.2833 | 93.5 | 91.8625 | 89.2 | 87.9875
-18.7 | 94.46 | 93.65 | 92.0143 | 89.35 | 88.1125
-18.6 | 94.7143 | 93.8167 | 92.1667 | 89.475 | 88.2375
-18.5 | 94.88 | 93.9714 | 92.3 | 89.6429 | 88.3625
-18.4 | 95 | 94.1286 | 92.3571 | 89.775 | 88.4778
-18.3 | 95.12 | 94.2714 | 92.5167 | 89.925 | 88.6
-18.2 | 95.2857 | 94.32 | 92.6833 | 90.0571 | 88.725
-18.1 | 95.4 | 94.4 | 92.75 | 90.1875 | 88.8571
-18 | 95.58 | 94.625 | 92.8167 | 90.3375 | 89.075
-17.9 | 95.6429 | 94.75 | 92.9714 | 90.4 | 89.225
-17.8 | 95.7 | 94.8 | 93.1125 | 90.4875 | 89.35
-17.7 | 95.8 | 94.9 | 93.2 | 90.6 | 89.5
-17.6 | 95.9 | 95.0286 | 93.3 | 90.7 | 89.5571
-17.5 | 96.0571 | 95.1 | 93.3429 | 90.85 | 89.7286
-17.4 | 96.1 | 95.125 | 93.4857 | 91.0167 | 89.8625
-17.3 | 96.1 | 95.2571 | 93.5 | 91.1 | 89.9875
-17.2 | 96.2 | 95.3167 | 93.5857 | 91.2571 | 90.1429
-17.1 | 96.2286 | 95.4 | 93.7667 | 91.4 | 90.3
-17 | 96.3 | 95.5 | 93.85 | 91.5333 | 90.425
-16.9 | 96.4 | 95.54 | 93.9167 | 91.6857 | 90.55
-16.8 | 96.5 | 95.6 | 94 | 91.7429 | 90.6
-16.7 | 96.5 | 95.7 | 94.1 | 91.9 | 90.7
-16.6 | 96.6 | 95.7 | 94.1857 | 92.0714 | 90.8
-16.5 | 96.7 | 95.8 | 94.32 | 92.1333 | 90.8222
-16.4 | 96.78 | 95.8429 | 94.5 | 92.3 | 90.9857
-16.3 | 96.8 | 95.9 | 94.5 | 92.4 | 91.1857
-16.2 | 96.9 | 96 | 94.6 | 92.5 | 91.4429
-16.1 | 96.9 | 96 | 94.6833 | 92.6 | 91.5
-16 | 97 | 96.1 | 94.8429 | 92.6625 | 91.6
-15.9 | 97 | 96.1833 | 94.9 | 92.7875 | 91.6714
-15.8 | 97.0125 | 96.2 | 95 | 92.9 | 91.8
-15.7 | 97.1 | 96.2714 | 95.1 | 93.0167 | 91.8857
-15.6 | 97.2 | 96.34 | 95.1429 | 93.1 | 92
-15.5 | 97.2714 | 96.4286 | 95.2 | 93.2 | 92.125
-15.4 | 97.3 | 96.5 | 95.3 | 93.3125 | 92.25
-15.3 | 97.3 | 96.6 | 95.3 | 93.4429 | 92.375
-15.2 | 97.4 | 96.7 | 95.45 | 93.5857 | 92.5333
-15.1 | 97.4143 | 96.7 | 95.5 | 93.6125 | 92.6125
-15 | 97.5 | 96.8 | 95.6 | 93.7 | 92.7
-14.9 | 97.6 | 96.8 | 95.7 | 93.875 | 92.7625
-14.8 | 97.7 | 96.9 | 95.7667 | 94 | 92.9
-14.7 | 97.7667 | 96.975 | 95.8429 | 94.1 | 93.0375
-14.6 | 97.8 | 97 | 95.9 | 94.125 | 93.1625
-14.5 | 97.8 | 97.0857 | 96 | 94.25 | 93.2875
-14.4 | 97.9 | 97.2 | 96.0857 | 94.375 | 93.4
-14.3 | 97.9 | 97.2 | 96.15 | 94.5 | 93.425
-14.2 | 97.9167 | 97.2875 | 96.2875 | 94.5286 | 93.5
-14.1 | 98 | 97.4 | 96.3167 | 94.6 | 93.64
-14 | 98 | 97.4 | 96.4 | 94.7 | 93.7167
-13.9 | 98.1 | 97.4875 | 96.5 | 94.8 | 93.8833
-13.8 | 98.1 | 97.5 | 96.6 | 94.8833 | 93.9375
-13.7 | 98.2 | 97.6 | 96.7 | 95 | 94.0714
-13.6 | 98.3125 | 97.6667 | 96.8 | 95.0833 | 94.2
-13.5 | 98.4 | 97.725 | 96.8 | 95.2 | 94.2222
-13.4 | 98.4286 | 97.8 | 96.8 | 95.3167 | 94.3375
-13.3 | 98.5 | 97.8 | 96.9 | 95.4 | 94.4833
-13.2 | 98.56 | 97.9 | 96.9875 | 95.4 | 94.5
-13.1 | 98.6 | 97.9 | 97.05 | 95.5 | 94.6
-13 | 98.7 | 98 | 97.1167 | 95.5714 | 94.7167
-12.9 | 98.7 | 98.1333 | 97.2 | 95.6286 | 94.8625
-12.8 | 98.7857 | 98.2 | 97.3 | 95.7 | 94.9
-12.7 | 98.8571 | 98.3 | 97.3571 | 95.7375 | 94.9625
-12.6 | 99 | 98.3 | 97.42 | 95.8625 | 95
-12.5 | 99 | 98.4 | 97.5 | 95.9875 | 95.1
-12.4 | 99.1 | 98.4 | 97.6 | 96.1 | 95.2
-12.3 | 99.1286 | 98.45 | 97.6143 | 96.2 | 95.3429
-12.2 | 99.2 | 98.5 | 97.7 | 96.25 | 95.4857
-12.1 | 99.2333 | 98.5714 | 97.7125 | 96.3 | 95.55
-12 | 99.3 | 98.625 | 97.8 | 96.4 | 95.6
-11.9 | 99.35 | 98.7 | 97.8833 | 96.5 | 95.6571
-11.8 | 99.4 | 98.8 | 97.9375 | 96.5667 | 95.7
-11.7 | 99.4 | 98.8 | 98 | 96.6429 | 95.8
-11.6 | 99.5 | 98.9 | 98.1 | 96.7167 | 95.825
-11.5 | 99.5 | 98.9 | 98.1 | 96.8 | 95.9571
-11.4 | 99.6 | 99 | 98.1 | 96.84 | 96.02
-11.3 | 99.6 | 99.0143 | 98.2286 | 96.9 | 96.1857
-11.2 | 99.675 | 99.1 | 98.3 | 96.9571 | 96.2
-11.1 | 99.7 | 99.1143 | 98.3429 | 97.0875 | 96.3125
-11 | 99.7833 | 99.2 | 98.4 | 97.1333 | 96.4375
-10.9 | 99.8 | 99.2333 | 98.4833 | 97.2 | 96.5
-10.8 | 99.9 | 99.3167 | 98.5 | 97.3 | 96.6
-10.7 | 99.9 | 99.4 | 98.5857 | 97.4 | 96.6
-10.6 | 99.9333 | 99.4333 | 98.7 | 97.4667 | 96.74
-10.5 | 100 | 99.5 | 98.75 | 97.5 | 96.8
-10.4 | 100.05 | 99.5 | 98.8 | 97.5 | 96.9143
-10.3 | 100.1 | 99.5 | 98.85 | 97.6 | 97
-10.2 | 100.1375 | 99.5333 | 98.9 | 97.7 | 97.1
-10.1 | 100.2 | 99.6 | 98.9 | 97.7 | 97.1
-10 | 100.2571 | 99.6857 | 99 | 97.8429 | 97.2
-9.9 | 100.3 | 99.7 | 99 | 97.9 | 97.2429
-9.8 | 100.4 | 99.7167 | 99.05 | 97.9667 | 97.3
-9.7 | 100.4 | 99.8 | 99.1 | 98.0333 | 97.4
-9.6 | 100.4 | 99.8667 | 99.2 | 98.1 | 97.4833
-9.5 | 100.5 | 99.9 | 99.2 | 98.1333 | 97.6
-9.4 | 100.5 | 99.9 | 99.2333 | 98.2 | 97.7
-9.3 | 100.5667 | 100 | 99.3 | 98.3 | 97.7429
-9.2 | 100.54 | 100 | 99.3667 | 98.4 | 97.8
-9.1 | 100.6 | 100.0286 | 99.4 | 98.44 | 97.8
-9 | 100.6 | 100.1 | 99.5 | 98.5 | 97.8
-8.9 | 100.6 | 100.1167 | 99.5 | 98.5 | 97.9
-8.8 | 100.6429 | 100.2 | 99.6 | 98.6 | 97.9143
-8.7 | 100.7 | 100.2 | 99.6 | 98.6 | 98
-8.6 | 100.7375 | 100.275 | 99.6 | 98.6667 | 98.075
-8.5 | 100.8 | 100.3 | 99.7 | 98.7286 | 98.1286
-8.4 | 100.8 | 100.3 | 99.7 | 98.8 | 98.2714
-8.3 | 100.8 | 100.3 | 99.7 | 98.8167 | 98.3222
-8.2 | 100.8714 | 100.3833 | 99.78 | 98.9625 | 98.4
-8.1 | 100.9333 | 100.46 | 99.8 | 99 | 98.4
-8 | 101 | 100.5 | 99.8 | 99.1 | 98.5
-7.9 | 101 | 100.5 | 99.8571 | 99.1 | 98.5333
-7.8 | 101 | 100.5 | 99.9 | 99.2 | 98.6
-7.7 | 101 | 100.56 | 99.9375 | 99.2 | 98.7
-7.6 | 101 | 100.6 | 100 | 99.2625 | 98.8
-7.5 | 101 | 100.6 | 100 | 99.3 | 98.8
-7.4 | 101 | 100.6 | 100.1 | 99.3 | 98.9
-7.3 | 101 | 100.6 | 100.1 | 99.4 | 98.9
-7.2 | 101.1 | 100.675 | 100.2 | 99.4 | 98.9
-7.1 | 101.1 | 100.7 | 100.2 | 99.4333 | 99
-7 | 101.1571 | 100.7 | 100.2 | 99.5 | 99
-6.9 | 101.2 | 100.725 | 100.26 | 99.5 | 99.0667
-6.8 | 101.2 | 100.8 | 100.3 | 99.6 | 99.1
-6.7 | 101.2 | 100.8 | 100.3 | 99.6 | 99.1
-6.6 | 101.2 | 100.85 | 100.3 | 99.6 | 99.1167
-6.5 | 101.2 | 100.9 | 100.3 | 99.7 | 99.2
-6.4 | 101.2 | 100.9 | 100.4 | 99.75 | 99.2333
-6.3 | 101.2 | 100.9 | 100.4 | 99.8 | 99.3
-6.2 | 101.2 | 100.9 | 100.4 | 99.8 | 99.35
-6.1 | 101.2286 | 100.9 | 100.5 | 99.8 | 99.4
-6 | 101.2286 | 100.9 | 100.5 | 99.8 | 99.4
-5.9 | 101.2333 | 101 | 100.5 | 99.8 | 99.4571
-5.8 | 101.3 | 101 | 100.5 | 99.8875 | 99.5
-5.7 | 101.3 | 101 | 100.5 | 99.9333 | 99.5
-5.6 | 101.3 | 101 | 100.4 | 100 | 99.5
-5.5 | 101.3 | 101 | 100.4 | 100 | 99.5
-5.4 | 101.3 | 101 | 100.4571 | 100 | 99.5
-5.3 | 101.3 | 101 | 100.5 | 100.0571 | 99.5
-5.2 | 101.3 | 101 | 100.5 | 100 | 99.52
-5.1 | 101.3 | 101 | 100.5 | 100.0625 | 99.6
-5 | 101.3 | 101 | 100.5 | 100.1 | 99.7
-4.9 | 101.2833 | 101 | 100.5 | 100.1 | 99.7
-4.8 | 101.2 | 101 | 100.5 | 100.1 | 99.7
-4.7 | 101.2 | 100.9 | 100.5 | 100.1 | 99.7
-4.6 | 101.2 | 100.9 | 100.5 | 100.1 | 99.7714
-4.5 | 101.2 | 100.9 | 100.5 | 100.1 | 99.8
-4.4 | 101.16 | 100.9 | 100.5 | 100.1 | 99.8
-4.3 | 101.1 | 100.9 | 100.5 | 100.1 | 99.8
-4.2 | 101.1 | 100.9 | 100.5 | 100.1 | 99.86
-4.1 | 101.1 | 100.8167 | 100.5 | 100.1 | 99.9
-4 | 101.08 | 100.8 | 100.4 | 100.1 | 99.9
-3.9 | 101 | 100.8 | 100.4 | 100.1 | 99.9
-3.8 | 101 | 100.8 | 100.4 | 100.2 | 99.9
-3.7 | 100.92 | 100.8 | 100.4 | 100.2 | 99.9
-3.6 | 100.9 | 100.8 | 100.4 | 100.1 | 99.8571
-3.5 | 100.9 | 100.7 | 100.4 | 100.1 | 99.875
-3.4 | 100.9 | 100.7 | 100.4 | 100.1 | 99.9
-3.3 | 100.84 | 100.7 | 100.3667 | 100.1 | 99.9
-3.2 | 100.8 | 100.7 | 100.3 | 100.1 | 99.95
-3.1 | 100.7 | 100.7 | 100.3 | 100.1429 | 100
-3 | 100.7 | 100.7 | 100.3 | 100.2 | 100
-2.9 | 100.7 | 100.6 | 100.3 | 100.2 | 100
-2.8 | 100.7 | 100.6 | 100.3 | 100.2 | 100
-2.7 | 100.6 | 100.6 | 100.3 | 100.2 | 100
-2.6 | 100.6 | 100.6 | 100.3 | 100.2 | 100
-2.5 | 100.6 | 100.6 | 100.3 | 100.2 | 100
-2.4 | 100.5 | 100.55 | 100.3 | 100.2 | 100
-2.3 | 100.5 | 100.5 | 100.2 | 100.2 | 100
-2.2 | 100.5 | 100.4 | 100.2 | 100.2 | 100
-2.1 | 100.4667 | 100.4 | 100.2 | 100.2 | 100
-2 | 100.4 | 100.4 | 100.2 | 100.2 | 100
-1.9 | 100.4 | 100.3 | 100.2 | 100.2 | 100
-1.8 | 100.4 | 100.3 | 100.2 | 100.1 | 100
-1.7 | 100.4 | 100.3 | 100.2 | 100.1 | 100
-1.6 | 100.3333 | 100.2 | 100.2 | 100.1 | 100
-1.5 | 100.3 | 100.2 | 100.2 | 100.0571 | 100
-1.4 | 100.2833 | 100.2 | 100.1 | 100 | 100
-1.3 | 100.2 | 100.2 | 100 | 100 | 100
-1.2 | 100.2 | 100.1 | 100 | 100 | 100
-1.1 | 100.2 | 100.1 | 100 | 100 | 100
-1 | 100.2 | 100.1 | 100 | 100 | 100
-0.9 | 100.2 | 100.1 | 100 | 99.9875 | 100
-0.8 | 100.2 | 100.0286 | 100 | 99.9 | 100
-0.7 | 100.2 | 100 | 100 | 100 | 100
-0.6 | 100.1714 | 100 | 100 | 99.9286 | 100
-0.5 | 100.1 | 100 | 100 | 99.8875 | 100
-0.4 | 100 | 100 | 100 | 99.8 | 99.9429
-0.3 | 100 | 99.9833 | 99.9286 | 99.8857 | 99.9
-0.2 | 100 | 100 | 99.9143 | 99.9429 | 99.9
-0.1 | 100 | 100 | 100 | 100 | 99.9167
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 99.9167
0.2 | 100 | 100 | 99.9857 | 99.9429 | 99.9
0.3 | 100 | 99.9833 | 99.9571 | 99.8857 | 99.9
0.4 | 100 | 100 | 100 | 99.8 | 99.9429
0.5 | 100.1 | 100 | 100 | 99.8875 | 100
0.6 | 100.1714 | 100 | 100 | 99.9286 | 100
0.7 | 100.2 | 100 | 100 | 100 | 100
0.8 | 100.2 | 100.0286 | 100 | 99.9 | 100
0.9 | 100.2 | 100.1 | 100 | 99.9875 | 100
1 | 100.2 | 100.1 | 100 | 100 | 100
1.1 | 100.2 | 100.1 | 100 | 100 | 100
1.2 | 100.2 | 100.1 | 100 | 100 | 100
1.3 | 100.2 | 100.2 | 100 | 100 | 100
1.4 | 100.2833 | 100.2 | 100.1 | 100 | 100
1.5 | 100.3 | 100.2 | 100.2 | 100.0571 | 100
1.6 | 100.3333 | 100.2 | 100.2 | 100.1 | 100
1.7 | 100.4 | 100.3 | 100.2 | 100.1 | 100
1.8 | 100.4 | 100.3 | 100.2 | 100.1 | 100
1.9 | 100.4 | 100.3 | 100.2 | 100.2 | 100
2 | 100.4 | 100.4 | 100.2 | 100.2 | 100
2.1 | 100.4667 | 100.4 | 100.2 | 100.2 | 100
2.2 | 100.5 | 100.4 | 100.2 | 100.2 | 100
2.3 | 100.5 | 100.5 | 100.2 | 100.2 | 100
2.4 | 100.5 | 100.55 | 100.3 | 100.2 | 100
2.5 | 100.6 | 100.6 | 100.3 | 100.2 | 100
2.6 | 100.6 | 100.6 | 100.3 | 100.2 | 100
2.7 | 100.6 | 100.6 | 100.3 | 100.2 | 100
2.8 | 100.7 | 100.6 | 100.3 | 100.2 | 100
2.9 | 100.7 | 100.6 | 100.3 | 100.2 | 100
3 | 100.7 | 100.7 | 100.3 | 100.2 | 100
3.1 | 100.7 | 100.7 | 100.3 | 100.1429 | 100
3.2 | 100.8 | 100.7 | 100.3 | 100.1 | 99.95
3.3 | 100.84 | 100.7 | 100.3 | 100.1 | 99.9
3.4 | 100.9 | 100.7 | 100.4 | 100.1 | 99.9
3.5 | 100.9 | 100.7 | 100.4 | 100.1 | 99.875
3.6 | 100.9 | 100.8 | 100.4 | 100.1 | 99.8571
3.7 | 100.92 | 100.8 | 100.4 | 100.2 | 99.9
3.8 | 101 | 100.8 | 100.4 | 100.2 | 99.9
3.9 | 101 | 100.8 | 100.4 | 100.1 | 99.9
4 | 101.08 | 100.8 | 100.4167 | 100.1 | 99.9
4.1 | 101.1 | 100.8167 | 100.5 | 100.1 | 99.9
4.2 | 101.1 | 100.9 | 100.5 | 100.1 | 99.86
4.3 | 101.1 | 100.9 | 100.5 | 100.1 | 99.8
4.4 | 101.16 | 100.9 | 100.5 | 100.1 | 99.8
4.5 | 101.2 | 100.9 | 100.5 | 100.1 | 99.8
4.6 | 101.2 | 100.9 | 100.5 | 100.1 | 99.7714
4.7 | 101.2 | 100.9 | 100.5 | 100.1 | 99.7
4.8 | 101.2 | 101 | 100.5 | 100.1 | 99.7
4.9 | 101.2833 | 101 | 100.5 | 100.1 | 99.7
5 | 101.3 | 101 | 100.5 | 100.1 | 99.7
5.1 | 101.3 | 101 | 100.5 | 100.0625 | 99.6
5.2 | 101.3 | 101 | 100.5 | 100 | 99.52
5.3 | 101.3 | 101 | 100.5 | 100.0571 | 99.5
5.4 | 101.3 | 101 | 100.4286 | 100 | 99.5
5.5 | 101.3 | 101 | 100.4 | 100 | 99.5
5.6 | 101.3 | 101 | 100.4 | 100 | 99.5
5.7 | 101.3 | 101 | 100.5 | 99.9333 | 99.5
5.8 | 101.3 | 101 | 100.5 | 99.8875 | 99.5
5.9 | 101.2333 | 101 | 100.5 | 99.8 | 99.4571
6 | 101.2286 | 100.9 | 100.5 | 99.8 | 99.4
6.1 | 101.2286 | 100.9 | 100.5 | 99.8 | 99.4
6.2 | 101.2 | 100.9 | 100.4333 | 99.8 | 99.35
6.3 | 101.2 | 100.9 | 100.4 | 99.8 | 99.3
6.4 | 101.2 | 100.9 | 100.375 | 99.75 | 99.2333
6.5 | 101.2 | 100.9 | 100.3 | 99.7 | 99.2
6.6 | 101.2 | 100.85 | 100.3 | 99.6 | 99.1167
6.7 | 101.2 | 100.8 | 100.3 | 99.6 | 99.1
6.8 | 101.2 | 100.8 | 100.3 | 99.6 | 99.1
6.9 | 101.2 | 100.725 | 100.24 | 99.5 | 99.0667
7 | 101.1571 | 100.7 | 100.2 | 99.5 | 99
7.1 | 101.1 | 100.7 | 100.2 | 99.4333 | 99
7.2 | 101.1 | 100.675 | 100.2 | 99.4 | 98.9
7.3 | 101 | 100.6 | 100.1 | 99.4 | 98.9
7.4 | 101 | 100.6 | 100.1 | 99.3 | 98.9
7.5 | 101 | 100.6 | 100 | 99.3 | 98.8
7.6 | 101 | 100.6 | 100 | 99.2625 | 98.8
7.7 | 101 | 100.56 | 99.9125 | 99.2 | 98.7
7.8 | 101 | 100.5 | 99.9 | 99.2 | 98.6
7.9 | 101 | 100.5 | 99.9 | 99.1 | 98.5333
8 | 101 | 100.5 | 99.8 | 99.1 | 98.5
8.1 | 100.9333 | 100.46 | 99.8 | 99 | 98.4
8.2 | 100.8714 | 100.3833 | 99.75 | 98.9625 | 98.4
8.3 | 100.8 | 100.3 | 99.7 | 98.8167 | 98.3222
8.4 | 100.8 | 100.3 | 99.7 | 98.8 | 98.2714
8.5 | 100.8 | 100.3 | 99.7 | 98.7286 | 98.1286
8.6 | 100.7375 | 100.275 | 99.6 | 98.6667 | 98.075
8.7 | 100.7 | 100.2 | 99.6 | 98.6 | 98
8.8 | 100.6429 | 100.2 | 99.6 | 98.6 | 97.9143
8.9 | 100.6 | 100.1167 | 99.5 | 98.5 | 97.9
9 | 100.6 | 100.1 | 99.4833 | 98.5 | 97.8
9.1 | 100.6 | 100.0286 | 99.4 | 98.44 | 97.8
9.2 | 100.54 | 100 | 99.3333 | 98.4 | 97.8
9.3 | 100.5667 | 100 | 99.3 | 98.3 | 97.7429
9.4 | 100.5 | 99.9 | 99.3 | 98.2 | 97.7
9.5 | 100.5 | 99.9 | 99.2 | 98.1333 | 97.6
9.6 | 100.4 | 99.8667 | 99.2 | 98.1 | 97.4833
9.7 | 100.4 | 99.8 | 99.1 | 98.0333 | 97.4
9.8 | 100.4 | 99.7167 | 99.02 | 97.9667 | 97.3
9.9 | 100.3 | 99.7 | 99 | 97.9 | 97.2429
10 | 100.2571 | 99.6857 | 99 | 97.8429 | 97.2
10.1 | 100.2 | 99.6 | 98.9 | 97.7 | 97.1
10.2 | 100.1375 | 99.5333 | 98.9 | 97.7 | 97.1
10.3 | 100.1 | 99.5 | 98.8286 | 97.6 | 97
10.4 | 100.05 | 99.5 | 98.8 | 97.5 | 96.9143
10.5 | 100 | 99.5 | 98.7286 | 97.5 | 96.8
10.6 | 99.9333 | 99.4333 | 98.7 | 97.4667 | 96.74
10.7 | 99.9 | 99.4 | 98.5571 | 97.4 | 96.6
10.8 | 99.9 | 99.3167 | 98.5 | 97.3 | 96.6
10.9 | 99.8 | 99.2333 | 98.46 | 97.2 | 96.5
11 | 99.7833 | 99.2 | 98.4 | 97.1333 | 96.4375
11.1 | 99.7 | 99.1143 | 98.3167 | 97.0875 | 96.3125
11.2 | 99.675 | 99.1 | 98.3 | 96.9571 | 96.2
11.3 | 99.6 | 99.0143 | 98.2 | 96.9 | 96.1857
11.4 | 99.6 | 99 | 98.15 | 96.84 | 96.02
11.5 | 99.5 | 98.9 | 98.1 | 96.8 | 95.9571
11.6 | 99.5 | 98.9 | 98.1 | 96.7167 | 95.825
11.7 | 99.4 | 98.8 | 98 | 96.6429 | 95.8
11.8 | 99.4 | 98.8 | 97.9143 | 96.5667 | 95.7
11.9 | 99.35 | 98.7 | 97.86 | 96.5 | 95.6571
12 | 99.3 | 98.625 | 97.8 | 96.4 | 95.6
12.1 | 99.2333 | 98.5714 | 97.8 | 96.3 | 95.55
12.2 | 99.2 | 98.5 | 97.7 | 96.25 | 95.4857
12.3 | 99.1286 | 98.45 | 97.6 | 96.2 | 95.3429
12.4 | 99.1 | 98.4 | 97.6 | 96.1 | 95.2
12.5 | 99 | 98.4 | 97.5 | 95.9875 | 95.1
12.6 | 99 | 98.3 | 97.4 | 95.8625 | 95
12.7 | 98.8571 | 98.3 | 97.3286 | 95.7375 | 94.9625
12.8 | 98.7857 | 98.2 | 97.2875 | 95.7 | 94.9
12.9 | 98.7 | 98.1333 | 97.2 | 95.6286 | 94.8625
13 | 98.7 | 98 | 97.1 | 95.5714 | 94.7167
13.1 | 98.6 | 97.9 | 97.1 | 95.5 | 94.6
13.2 | 98.56 | 97.9 | 96.9625 | 95.4 | 94.5
13.3 | 98.5 | 97.8 | 96.9 | 95.4 | 94.4833
13.4 | 98.4286 | 97.8 | 96.8 | 95.3167 | 94.3375
13.5 | 98.4 | 97.725 | 96.8 | 95.2 | 94.2222
13.6 | 98.3125 | 97.6667 | 96.7714 | 95.0833 | 94.2
13.7 | 98.2 | 97.6 | 96.7 | 95 | 94.0714
13.8 | 98.1 | 97.5 | 96.575 | 94.8833 | 93.9375
13.9 | 98.1 | 97.4875 | 96.5 | 94.8 | 93.8833
14 | 98 | 97.4 | 96.4571 | 94.7 | 93.7167
14.1 | 98 | 97.4 | 96.3 | 94.6 | 93.64
14.2 | 97.9167 | 97.2875 | 96.2625 | 94.5286 | 93.5
14.3 | 97.9 | 97.2 | 96.2 | 94.5 | 93.425
14.4 | 97.9 | 97.2 | 96.1 | 94.375 | 93.4
14.5 | 97.8 | 97.0857 | 96.0143 | 94.25 | 93.2875
14.6 | 97.8 | 97 | 95.9714 | 94.125 | 93.1625
14.7 | 97.7667 | 96.975 | 95.8167 | 94.1 | 93.0375
14.8 | 97.7 | 96.9 | 95.8 | 94 | 92.9
14.9 | 97.6 | 96.8 | 95.7 | 93.875 | 92.7625
15 | 97.5 | 96.8 | 95.6286 | 93.7 | 92.7
15.1 | 97.4143 | 96.7 | 95.5 | 93.6125 | 92.6125
15.2 | 97.4 | 96.7 | 95.4286 | 93.5857 | 92.5333
15.3 | 97.3 | 96.6 | 95.3667 | 93.4429 | 92.375
15.4 | 97.3 | 96.5 | 95.3 | 93.3125 | 92.25
15.5 | 97.2714 | 96.4286 | 95.2667 | 93.2 | 92.125
15.6 | 97.2 | 96.34 | 95.1143 | 93.1 | 92
15.7 | 97.1 | 96.2714 | 95.075 | 93.0167 | 91.8857
15.8 | 97.0125 | 96.2 | 95 | 92.9 | 91.8
15.9 | 97 | 96.1833 | 94.9 | 92.7875 | 91.6714
16 | 97 | 96.1 | 94.8143 | 92.6625 | 91.6
16.1 | 96.9 | 96 | 94.66 | 92.6 | 91.5
16.2 | 96.9 | 96 | 94.5833 | 92.5 | 91.4429
16.3 | 96.8 | 95.9 | 94.5 | 92.4 | 91.1857
16.4 | 96.78 | 95.8429 | 94.4833 | 92.3 | 90.9857
16.5 | 96.7 | 95.8 | 94.3 | 92.1333 | 90.8222
16.6 | 96.6 | 95.7 | 94.2571 | 92.0714 | 90.8
16.7 | 96.5 | 95.7 | 94.1 | 91.9 | 90.7
16.8 | 96.5 | 95.6 | 94 | 91.7429 | 90.6
16.9 | 96.4 | 95.54 | 93.9 | 91.6857 | 90.55
17 | 96.3 | 95.5 | 93.8333 | 91.5333 | 90.425
17.1 | 96.2286 | 95.4 | 93.75 | 91.4 | 90.3
17.2 | 96.2 | 95.3167 | 93.6 | 91.2571 | 90.1429
17.3 | 96.1 | 95.2571 | 93.5143 | 91.1 | 89.9875
17.4 | 96.1 | 95.125 | 93.4571 | 91.0167 | 89.8625
17.5 | 96.0571 | 95.1 | 93.4 | 90.85 | 89.7286
17.6 | 95.9 | 95.0286 | 93.2857 | 90.7 | 89.5571
17.7 | 95.8 | 94.9 | 93.2 | 90.6 | 89.5
17.8 | 95.7 | 94.8 | 93.0857 | 90.4875 | 89.35
17.9 | 95.6429 | 94.75 | 92.9429 | 90.4 | 89.225
18 | 95.58 | 94.625 | 92.8 | 90.3375 | 89.075
18.1 | 95.4 | 94.4 | 92.7286 | 90.1875 | 88.8571
18.2 | 95.2857 | 94.32 | 92.66 | 90.0571 | 88.725
18.3 | 95.12 | 94.2714 | 92.5 | 89.925 | 88.6
18.4 | 95 | 94.1286 | 92.3571 | 89.775 | 88.4778
18.5 | 94.88 | 93.9714 | 92.2857 | 89.6429 | 88.3625
18.6 | 94.7143 | 93.8167 | 92.1429 | 89.475 | 88.2375
18.7 | 94.46 | 93.65 | 91.9857 | 89.35 | 88.1125
18.8 | 94.2833 | 93.5 | 91.8429 | 89.2 | 87.9875
18.9 | 94.0571 | 93.3 | 91.7125 | 89.0333 | 87.8571
19 | 93.82 | 93.1143 | 91.5714 | 89 | 87.7
19.1 | 93.5667 | 92.85 | 91.375 | 88.8714 | 87.6
19.2 | 93.2 | 92.7 | 91.2 | 88.7286 | 87.4571
19.3 | 92.7333 | 92.4333 | 91.025 | 88.5875 | 87.325
19.4 | 92.2857 | 92.1571 | 90.8167 | 88.425 | 87.28
19.5 | 91.6667 | 91.8333 | 90.6429 | 88.2375 | 87.05
19.6 | 90.96 | 91.46 | 90.4444 | 88.1125 | 86.9111
19.7 | 89.8333 | 91.1 | 90.2667 | 87.8857 | 86.66
19.8 | 88.2714 | 90.66 | 90.05 | 87.7429 | 86.5
19.9 | 85.96 | 90.14 | 89.7333 | 87.6 | 86.4333
20 | 82.8125 | 89.4 | 89.525 | 87.4167 | 86.4
20.1 | 77.9714 | 88.5 | 89.2429 | 87.2 | 86.1857
20.2 | 72.1143 | 87.25 | 88.9333 | 87.1167 | 86.0125
20.3 | 65.3 | 85.7 | 88.6714 | 86.9333 | 85.8833
20.4 | 56.88 | 83.55 | 88.3667 | 86.76 | 85.7286
20.5 | 48.9667 | 80.3714 | 87.9571 | 86.5 | 85.4833
20.6 | 41.3286 | 76.0667 | 87.5286 | 86.3333 | 85.3444
20.7 | 35 | 70.4 | 87.0333 | 86.1143 | 85.225
20.8 | 29.2 | 63.44 | 86.45 | 85.85 | 85
20.9 | 24.8667 | 56.2 | 85.8333 | 85.6 | 84.8714
21 | 21.65 | 48.2857 | 85.05 | 85.4714 | 84.6571
21.1 | 19.5 | 41.2571 | 84.1 | 85.2429 | 84.5
21.2 | 17.98 | 35.2 | 82.875 | 85 | 84.4286
21.3 | 17 | 29.85 | 81.2143 | 84.7625 | 84.2222
21.4 | 16.18 | 25.7 | 79.0125 | 84.5571 | 84.05
21.5 | 15.55 | 22.7286 | 75.95 | 84.2714 | 83.7833
21.6 | 15.08 | 20.5 | 71.8667 | 83.9667 | 83.6286
21.7 | 14.6667 | 18.7429 | 67.0333 | 83.65 | 83.42
21.8 | 14.3333 | 17.6 | 60.6 | 83.4 | 83.2
21.9 | 14.1 | 16.775 | 54.2 | 83.15 | 83.0333
22 | 13.8667 | 16.1 | 47.325 | 82.8 | 82.7833
22.1 | 13.6286 | 15.6286 | 41.4 | 82.4625 | 82.6167
22.2 | 13.44 | 15.15 | 35.625 | 82.1625 | 82.4333
22.3 | 13.2667 | 14.8167 | 30.8 | 81.9125 | 82.2714
22.4 | 13.125 | 14.4429 | 27.1444 | 81.3714 | 82
22.5 | 12.8667 | 14.2286 | 24.4 | 81.0286 | 81.825
22.6 | 12.8 | 13.9 | 22.4333 | 80.55 | 81.6
22.7 | 12.6333 | 13.7571 | 20.8143 | 80.0375 | 81.35
22.8 | 12.4857 | 13.5 | 19.7167 | 79.3857 | 81.12
22.9 | 12.3429 | 13.3571 | 18.8833 | 78.7571 | 80.8429
23 | 12.2 | 13.1286 | 18.1857 | 77.9333 | 80.6375
23.1 | 12.02 | 12.9714 | 17.65 | 77 | 80.425
23.2 | 11.9 | 12.775 | 17.15 | 75.7 | 80.15
23.3 | 11.7667 | 12.6 | 16.7714 | 73.85 | 79.8375
23.4 | 11.7 | 12.4167 | 16.4 | 71.525 | 79.5778
23.5 | 11.5667 | 12.2571 | 16.1 | 68.6889 | 79.35
23.6 | 11.4 | 12.1833 | 15.7667 | 65.0571 | 79.05
23.7 | 11.3429 | 12.0167 | 15.5 | 60.7111 | 78.825
23.8 | 11.275 | 11.875 | 15.2667 | 55.85 | 78.56
23.9 | 11.1333 | 11.8 | 15.1 | 50.6125 | 78.25
24 | 11.1 | 11.65 | 14.825 | 45.48 | 77.9
24.1 | 11.0167 | 11.5 | 14.6167 | 39.9429 | 77.6111
24.2 | 10.9 | 11.4 | 14.4 | 35.5 | 77.1714
24.3 | 10.8125 | 11.3 | 14.2125 | 32.0333 | 76.8
24.4 | 10.7 | 11.15 | 14.0667 | 28.84 | 76.4375
24.5 | 10.7 | 11.1 | 13.8286 | 26.5 | 76.05
24.6 | 10.52 | 11.0167 | 13.6714 | 24.6333 | 75.6333
24.7 | 10.5 | 10.9 | 13.475 | 23.3 | 74.9667
24.8 | 10.4 | 10.7857 | 13.3125 | 22.2833 | 74.4
24.9 | 10.3 | 10.7 | 13.1429 | 21.46 | 73.5333
25 | 10.2 | 10.6 | 12.9 | 20.8333 | 72.8
25.1 | 10.1286 | 10.5429 | 12.8571 | 20.2 | 71.75
25.2 | 10 | 10.4 | 12.7 | 19.7286 | 70.325
25.3 | 9.9429 | 10.35 | 12.575 | 19.45 | 68.8
25.4 | 9.9 | 10.2875 | 12.4429 | 19.0125 | 66.45
25.5 |  | 10.1571 | 12.2857 | 18.725 | 63.9
25.6 |  | 10.1 | 12.15 | 18.44 | 60.65
25.7 |  | 9.9857 | 12.025 | 18.1625 | 56.8
25.8 |  | 9.9 | 11.8833 | 17.825 | 52.4375
25.9 |  | 9.8 | 11.7286 | 17.5571 | 47.8125
26 |  |  | 11.6857 | 17.38 | 43.5
26.1 |  |  | 11.5429 | 17.1 | 39.3
26.2 |  |  | 11.4 | 16.8857 | 35.8
26.3 |  |  | 11.35 | 16.7 | 32.5
26.4 |  |  | 11.2 | 16.45 | 29.9
26.5 |  |  | 11.1 | 16.2333 | 27.8
26.6 |  |  | 11 | 16.0714 | 26.26
26.7 |  |  | 10.9 | 15.875 | 25.1
26.8 |  |  | 10.8 | 15.7125 | 24.1167
26.9 |  |  | 10.6833 | 15.475 | 23.3875
27 |  |  |  | 15.225 | 22.75
27.1 |  |  |  | 15.0375 | 22.1667
27.2 |  |  |  | 14.8833 | 21.7571
27.3 |  |  |  | 14.7 | 21.3
27.4 |  |  |  | 14.5167 | 20.9429
27.5 |  |  |  | 14.375 | 20.6333
27.6 |  |  |  | 14.2333 | 20.3429
27.7 |  |  |  | 14.06 | 20.075
27.8 |  |  |  | 13.9714 | 19.825
27.9 |  |  |  | 13.8375 | 19.5667
28 |  |  |  | 13.625 | 19.3
28.1 |  |  |  | 13.4857 | 19.05
28.2 |  |  |  | 13.3429 | 18.8429
28.3 |  |  |  | 13.2 | 18.68
28.4 |  |  |  | 13.1 | 18.375
28.5 |  |  |  | 12.94 | 18.1571
28.6 |  |  |  | 12.86 | 18
28.7 |  |  |  | 12.7 | 17.7333
28.8 |  |  |  | 12.6 | 17.6
28.9 |  |  |  | 12.45 | 17.3286
29 |  |  |  |  | 17.2
29.1 |  |  |  |  | 17.05
29.2 |  |  |  |  | 16.8111
29.3 |  |  |  |  | 16.6375
29.4 |  |  |  |  | 16.425
29.5 |  |  |  |  | 16.2875
29.6 |  |  |  |  | 16.125
29.7 |  |  |  |  | 15.9375
29.8 |  |  |  |  | 15.8
29.9 |  |  |  |  | 15.675
30 |  |  |  |  | 15.5
30.1 |  |  |  |  | 15.325
30.2 |  |  |  |  | 15.2
30.3 |  |  |  |  | 15.05
30.4 |  |  |  |  | 14.85
30.5 |  |  |  |  | 14.725
30.6 |  |  |  |  | 14.6
30.7 |  |  |  |  | 14.4714
30.8 |  |  |  |  | 14.3286
30.9 |  |  |  |  | 14.2

## Wedge 30 Output Factors
Output factors
Wedge 30 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.876 | 0.891 | 0.898 | 0.906 | 0.913 | 0.919 | 0.923 | 0.928
| 4 | 0.892 | 0.909 | 0.918 | 0.928 | 0.937 | 0.945 | 0.95 | 0.957
| 5 | 0.901 | 0.919 | 0.93 | 0.943 | 0.953 | 0.963 | 0.969 | 0.977
| 7 | 0.912 | 0.933 | 0.946 | 0.963 | 0.977 | 0.99 | 0.999 | 1.009
| 10 | 0.923 | 0.947 | 0.962 | 0.982 | 1.001 | 1.018 | 1.03 | 1.045
| 15 | 0.932 | 0.959 | 0.977 | 1.001 | 1.024 | 1.048 | 1.066 | 1.087
| 20 | 0.938 | 0.966 | 0.985 | 1.012 | 1.039 | 1.067 | 1.089 | 1.116
| 30 | 0.944 | 0.974 | 0.996 | 1.026 | 1.057 | 1.093 | 1.12 | 1.156
| 40 | 0.949 | 0.981 | 1.002 | 1.035 | 1.069 | 1.109 | 1.141 | 1.179

## Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 27.1 | 33.5 | 44.1 | 51.4
0.1 | 31.16 | 37.32 | 48.175 | 55.075
0.2 | 38.4 | 44.3 | 54.44 | 61
0.3 | 47.12 | 53.0667 | 62 | 67.9
0.4 | 55.6333 | 60.6667 | 68.6 | 73.9
0.5 | 63.0667 | 67.6667 | 74.6 | 79.35
0.6 | 69.24 | 72.94 | 79.35 | 83.44
0.7 | 74.25 | 77.7 | 83.38 | 86.725
0.8 | 78.75 | 82 | 86.6 | 89.56
0.9 | 82.6 | 85.18 | 89.34 | 91.9333
1 | 85.725 | 87.7571 | 91.6 | 93.7
1.1 | 88.3167 | 90.225 | 93.3 | 95.225
1.2 | 90.975 | 92.24 | 95.2 | 96.45
1.3 | 92.68 | 93.92 | 96.1 | 97.4
1.4 | 94.25 | 95.225 | 97.175 | 98.2667
1.5 | 95.72 | 96.4 | 97.94 | 98.8
1.6 | 96.65 | 97.3 | 98.475 | 99.2333
1.7 | 97.5 | 98.175 | 99.05 | 99.6
1.8 | 98.28 | 98.7 | 99.4 | 99.775
1.9 | 98.98 | 99.06 | 99.7 | 99.9
2 | 99.35 | 99.44 | 99.8833 | 100
2.1 | 99.7 | 99.78 | 100 | 100
2.2 | 99.84 | 99.9 | 100 | 100
2.3 | 100 | 100 | 100 | 99.9
2.4 | 100 | 100 | 99.8833 | 99.8
2.5 | 100 | 99.8833 | 99.8 | 99.6167
2.6 | 99.9333 | 99.8 | 99.65 | 99.375
2.7 | 99.7333 | 99.66 | 99.4667 | 99.1
2.8 | 99.56 | 99.5 | 99.2 | 98.9
2.9 | 99.36 | 99.3 | 98.9 | 98.68
3 | 99.125 | 98.95 | 98.6333 | 98.425
3.1 | 98.85 | 98.7333 | 98.325 | 98.1
3.2 | 98.42 | 98.55 | 98.06 | 97.7667
3.3 | 98.12 | 98.16 | 97.8 | 97.475
3.4 | 97.7667 | 97.82 | 97.4 | 97.05
3.5 | 97.46 | 97.42 | 97.16 | 96.775
3.6 | 97.1 | 97.075 | 96.775 | 96.4
3.7 | 96.7 | 96.7 | 96.45 | 96.1
3.8 | 96.3 | 96.32 | 95.95 | 95.7
3.9 | 95.7 | 95.95 | 95.7167 | 95.35
4 | 95.3333 | 95.55 | 95.4 | 95.1
4.1 | 94.94 | 95.1 | 95 | 94.66
4.2 | 94.45 | 94.8333 | 94.7 | 94.35
4.3 | 94.0333 | 94.5 | 94.4 | 94
4.4 | 93.55 | 94.2 | 93.875 | 93.62
4.5 | 93.125 | 93.76 | 93.525 | 93.35
4.6 | 92.82 | 93.35 | 93.24 | 92.95
4.7 | 92.275 | 93 | 92.9 | 92.6
4.8 | 91.9333 | 92.55 | 92.475 | 92.3
4.9 | 91.5667 | 92.2 | 92.2 | 91.8
5 | 91.075 | 91.7 | 91.92 | 91.575
5.1 | 90.725 | 91.44 | 91.575 | 91.28
5.2 | 90.38 | 90.95 | 91.22 | 91
5.3 | 89.94 | 90.625 | 90.8333 | 90.5333
5.4 | 89.4 | 90.275 | 90.4 | 90.2
5.5 | 89 | 89.9 | 90.12 | 89.9333
5.6 | 88.65 | 89.45 | 89.76 | 89.5
5.7 | 88.1 | 89.15 | 89.475 | 89.32
5.8 | 87.68 | 88.8667 | 89.125 | 88.84
5.9 | 87.4 | 88.62 | 88.8 | 88.56
6 | 86.8 | 88.22 | 88.4 | 88.2333
6.1 | 86.35 | 87.84 | 88.05 | 87.85
6.2 | 86.06 | 87.425 | 87.74 | 87.525
6.3 | 85.65 | 87 | 87.38 | 87.22
6.4 | 85.3 | 86.5667 | 87.04 | 86.9
6.5 | 84.8 | 86.2 | 86.7 | 86.6167
6.6 | 84.4 | 85.9 | 86.2333 | 86.26
6.7 | 83.9 | 85.45 | 85.95 | 85.95
6.8 | 83.46 | 85.2 | 85.5833 | 85.5
6.9 | 83.15 | 84.82 | 85.3333 | 85.28
7 | 82.84 | 84.5 | 85.05 | 85.025
7.1 | 82.475 | 84.24 | 84.66 | 84.6
7.2 | 82.125 | 83.8333 | 84.35 | 84.32
7.3 | 81.58 | 83.56 | 84.0667 | 83.94
7.4 | 81.16 | 83.2 | 83.78 | 83.625
7.5 | 80.72 | 82.82 | 83.42 | 83.3333
7.6 | 80.35 | 82.3 | 83.05 | 83.04
7.7 | 80.025 | 82.02 | 82.7667 | 82.5
7.8 | 79.64 | 81.725 | 82.3333 | 82.24
7.9 | 79.38 | 81.275 | 81.98 | 82
8 | 78.95 | 80.92 | 81.8 | 81.6
8.1 | 78.62 | 80.55 | 81.4333 | 81.3
8.2 | 78.06 | 80.3333 | 81.125 | 81
8.3 | 77.74 | 79.85 | 80.725 | 80.74
8.4 | 77.3 | 79.58 | 80.4667 | 80.26
8.5 | 76.78 | 79.3 | 80.16 | 80.02
8.6 | 76.55 | 78.9 | 79.8 | 79.74
8.7 | 76.14 | 78.5 | 79.5 | 79.48
8.8 | 75.6667 | 78.1 | 79.1 | 79.18
8.9 | 75.4 | 77.7 | 78.9 | 78.825
9 | 74.95 | 77.4 | 78.5333 | 78.575
9.1 | 74.62 | 77.075 | 78.3 | 78.2
9.2 | 74.3 | 76.825 | 77.96 | 77.86
9.3 | 73.85 | 76.325 | 77.6 | 77.625
9.4 | 73.5 | 76.06 | 77.34 | 77.3
9.5 | 73.1333 | 75.8 | 76.975 | 77
9.6 | 72.9 | 75.45 | 76.66 | 76.7333
9.7 | 72.4833 | 75.05 | 76.45 | 76.375
9.8 | 72.2 | 74.7667 | 76.1 | 76.075
9.9 | 71.9 | 74.52 | 75.82 | 75.8
10 | 71.46 | 74.1 | 75.45 | 75.48
10.1 | 70.9333 | 73.8333 | 75.15 | 75.15
10.2 | 70.6833 | 73.48 | 74.8 | 74.8
10.3 | 70.4333 | 73.15 | 74.5 | 74.45
10.4 | 70 | 72.85 | 74.26 | 74.3
10.5 | 69.72 | 72.625 | 74.04 | 74.075
10.6 | 69.54 | 72.2 | 73.65 | 73.82
10.7 | 69.05 | 71.8 | 73.24 | 73.525
10.8 | 68.625 | 71.6 | 73 | 73.175
10.9 | 68.3167 | 71.32 | 72.68 | 72.78
11 | 68.0333 | 70.98 | 72.42 | 72.525
11.1 | 67.6 | 70.6 | 72.06 | 72.25
11.2 | 67.35 | 70.4167 | 71.8 | 71.92
11.3 | 66.92 | 70.075 | 71.52 | 71.62
11.4 | 66.6 | 69.675 | 71.24 | 71.42
11.5 | 66.3 | 69.35 | 70.9333 | 71.15
11.6 | 65.9 | 69 | 70.65 | 70.8
11.7 | 65.6 | 68.76 | 70.42 | 70.56
11.8 | 65.35 | 68.46 | 70.22 | 70.25
11.9 | 64.98 | 68.225 | 69.85 | 70
12 | 64.7 | 67.8667 | 69.6667 | 69.64
12.1 | 64.34 | 67.525 | 69.2667 | 69.3667
12.2 | 64.1 | 67.175 | 68.95 | 69.2
12.3 | 63.75 | 66.85 | 68.6 | 68.9333
12.4 | 63.44 | 66.52 | 68.4 | 68.6333
12.5 | 63.08 | 66.25 | 68.1333 | 68.34
12.6 | 62.82 | 65.9667 | 67.85 | 68.12
12.7 | 62.3833 | 65.7167 | 67.65 | 67.74
12.8 | 62.2 | 65.5 | 67.26 | 67.4833
12.9 | 61.84 | 65.08 | 67.02 | 67.3
13 | 61.5 | 64.9 | 66.7333 | 66.9667
13.1 | 61.16 | 64.66 | 66.5 | 66.7
13.2 | 60.875 | 64.3 | 66.275 | 66.45
13.3 | 60.5667 | 64.04 | 65.92 | 66.2
13.4 | 60.4167 | 63.7 | 65.7143 | 65.95
13.5 | 60.12 | 63.42 | 65.46 | 65.64
13.6 | 59.9 | 63 | 65.24 | 65.42
13.7 | 59.4667 | 62.86 | 64.9 | 65.1
13.8 | 59.2 | 62.6333 | 64.6 | 64.86
13.9 | 58.9333 | 62.2667 | 64.2667 | 64.475
14 | 58.5 | 61.9333 | 64 | 64.28
14.1 | 58.1667 | 61.74 | 63.7667 | 64.075
14.2 | 57.9667 | 61.52 | 63.5 | 63.75
14.3 | 57.8 | 61.16 | 63.26 | 63.5571
14.4 | 57.5167 | 60.86 | 62.9667 | 63.28
14.5 | 57.2 | 60.55 | 62.75 | 63
14.6 | 56.8167 | 60.26 | 62.5 | 62.8
14.7 | 56.5 | 59.9833 | 62.2 | 62.52
14.8 | 56.18 | 59.8 | 62 | 62.25
14.9 | 55.86 | 59.6 | 61.8 | 62.025
15 | 55.6 | 59.24 | 61.5 | 61.775
15.1 | 55.3667 | 59.04 | 61.2 | 61.54
15.2 | 55.0667 | 58.7714 | 60.925 | 61.24
15.3 | 54.825 | 58.5333 | 60.76 | 61
15.4 | 54.6 | 58.15 | 60.5 | 60.75
15.5 | 54.4333 | 57.84 | 60.2 | 60.4667
15.6 | 54.15 | 57.6 | 60 | 60.3
15.7 | 53.92 | 57.4 | 59.7 | 60
15.8 | 53.6 | 57.1333 | 59.4 | 59.9
15.9 | 53.4167 | 56.85 | 59.15 | 59.58
16 | 53.08 | 56.7 | 58.92 | 59.25
16.1 | 52.84 | 56.475 | 58.675 | 59.04
16.2 | 52.52 | 56.16 | 58.45 | 58.82
16.3 | 52.16 | 55.98 | 58.24 | 58.66
16.4 | 51.9333 | 55.675 | 58.025 | 58.36
16.5 | 51.7 | 55.36 | 57.82 | 58.025
16.6 | 51.475 | 55.06 | 57.58 | 57.82
16.7 | 51.25 | 54.8 | 57.4 | 57.6
16.8 | 50.9833 | 54.66 | 57.05 | 57.4333
16.9 | 50.8 | 54.4 | 56.825 | 57.2
17 | 50.52 | 54.2 | 56.7 | 56.9333
17.1 | 50.2667 | 53.9 | 56.35 | 56.7
17.2 | 49.94 | 53.6 | 56.14 | 56.5
17.3 | 49.7 | 53.35 | 55.8667 | 56.32
17.4 | 49.475 | 53.1333 | 55.625 | 56.06
17.5 | 49.225 | 52.98 | 55.4 | 55.78
17.6 | 49.0333 | 52.7286 | 55.22 | 55.5667
17.7 | 48.775 | 52.5 | 55 | 55.4
17.8 | 48.5667 | 52.2 | 54.7667 | 55.1333
17.9 | 48.4 | 52 | 54.54 | 54.9
18 | 48.1667 | 51.78 | 54.3 | 54.75
18.1 | 47.9 | 51.6 | 54.05 | 54.475
18.2 | 47.66 | 51.3 | 53.8 | 54.275
18.3 | 47.3333 | 51 | 53.62 | 54.0667
18.4 | 47.075 | 50.78 | 53.3333 | 53.78
18.5 | 46.84 | 50.52 | 53.175 | 53.5
18.6 | 46.725 | 50.4333 | 52.96 | 53.36
18.7 | 46.4 | 50.16 | 52.76 | 53.08
18.8 | 46.1 | 49.88 | 52.46 | 52.8
18.9 | 45.94 | 49.65 | 52.225 | 52.7
19 | 45.75 | 49.3 | 51.9833 | 52.42
19.1 | 45.5 | 49.1333 | 51.9 | 52.2667
19.2 | 45.3 | 48.94 | 51.6333 | 52.06
19.3 | 45.1 | 48.74 | 51.45 | 51.84
19.4 | 44.8333 | 48.52 | 51.18 | 51.625
19.5 | 44.65 | 48.3167 | 50.9 | 51.4333
19.6 | 44.38 | 48.0667 | 50.725 | 51.16
19.7 | 44.125 | 47.84 | 50.5 | 51
19.8 | 43.98 | 47.6 | 50.24 | 50.8
19.9 | 43.775 | 47.35 | 50.0333 | 50.5333
20 | 43.525 | 47.08 | 49.9 | 50.3
20.1 | 43.4167 | 46.85 | 49.7 | 50.175
20.2 | 43.1333 | 46.7333 | 49.4 | 49.96
20.3 | 42.9 | 46.55 | 49.2167 | 49.7333
20.4 | 42.76 | 46.325 | 49 | 49.5
20.5 | 42.55 | 46.22 | 48.9167 | 49.35
20.6 | 42.4 | 45.95 | 48.68 | 49.2
20.7 | 42.16 | 45.75 | 48.46 | 48.9667
20.8 | 41.94 | 45.6 | 48.2 | 48.76
20.9 | 41.725 | 45.4333 | 48.1 | 48.575
21 | 41.5667 | 45.14 | 47.85 | 48.4
21.1 | 41.4 | 44.84 | 47.6 | 48.15
21.2 | 41.1333 | 44.7 | 47.4333 | 47.92
21.3 | 40.8667 | 44.5 | 47.22 | 47.8
21.4 | 40.725 | 44.275 | 47.1 | 47.52
21.5 | 40.525 | 44.025 | 46.8667 | 47.4
21.6 | 40.275 | 43.9286 | 46.62 | 47.2
21.7 | 40.1 | 43.64 | 46.5 | 47.025
21.8 | 39.8333 | 43.52 | 46.32 | 46.8
21.9 | 39.6667 | 43.32 | 46.12 | 46.62
22 | 39.5 | 43.12 | 45.9 | 46.4
22.1 | 39.2333 | 42.95 | 45.725 | 46.2
22.2 | 39.1 | 42.74 | 45.45 | 46.06
22.3 | 38.96 | 42.55 | 45.2667 | 45.775
22.4 | 38.775 | 42.3667 | 45.1 | 45.56
22.5 | 38.65 | 42.2 | 44.9 | 45.46
22.6 | 38.4667 | 42 | 44.74 | 45.28
22.7 | 38.3 | 41.7667 | 44.525 | 45.06
22.8 | 38.0667 | 41.5667 | 44.28 | 44.84
22.9 | 37.9 | 41.45 | 44.1 | 44.7
23 | 37.7 | 41.16 | 43.975 | 44.54
23.1 | 37.5 | 40.975 | 43.84 | 44.4
23.2 | 37.3 | 40.8 | 43.64 | 44.2
23.3 | 37.18 | 40.65 | 43.45 | 44
23.4 | 37 | 40.45 | 43.2833 | 43.825
23.5 | 36.86 | 40.325 | 43.1 | 43.6429
23.6 | 36.675 | 40.1 | 42.9 | 43.45
23.7 | 36.4 | 39.95 | 42.625 | 43.3
23.8 | 36.2833 | 39.8 | 42.4 | 43.15
23.9 | 36.1167 | 39.6 | 42.275 | 42.9
24 | 35.925 | 39.325 | 42.14 | 42.7167
24.1 | 35.72 | 39.18 | 41.94 | 42.5
24.2 | 35.55 | 38.9667 | 41.82 | 42.35
24.3 | 35.3 | 38.78 | 41.7 | 42.2
24.4 | 35.2 | 38.6 | 41.46 | 42.05
24.5 | 35.0667 | 38.42 | 41.25 | 41.875
24.6 | 34.7833 | 38.3 | 41.15 | 41.66
24.7 | 34.6833 | 38.1167 | 40.94 | 41.5
24.8 | 34.5 | 38 | 40.7667 | 41.375
24.9 | 34.4 | 37.8167 | 40.58 | 41.15
25 | 34.2 | 37.64 | 40.3833 | 41.02
25.1 | 34.075 | 37.4 | 40.2 | 40.9
25.2 | 33.9 | 37.3 | 40.075 | 40.66
25.3 | 33.75 | 37.2 | 39.95 | 40.4667
25.4 | 33.7 | 37.04 | 39.7333 | 40.38
25.5 | 33.48 | 36.84 | 39.5667 | 40.18
25.6 | 33.26 | 36.65 | 39.4 | 39.98
25.7 | 33.1 | 36.4833 | 39.3 | 39.825
25.8 | 33 | 36.3167 | 39.1 | 39.7
25.9 | 32.8 | 36.2 | 38.9667 | 39.5
26 | 32.6 | 36 | 38.72 | 39.4
26.1 | 32.5 | 35.9 | 38.66 | 39.22
26.2 | 32.3 | 35.72 | 38.4667 | 39.04
26.3 | 32.1333 | 35.5 | 38.3 | 38.9
26.4 | 32.02 | 35.3667 | 38.175 | 38.7167
26.5 | 31.84 | 35.1833 | 37.95 | 38.54
26.6 | 31.74 | 35 | 37.7667 | 38.4
26.7 | 31.5333 | 34.9 | 37.65 | 38.24
26.8 | 31.34 | 34.7 | 37.5 | 38.125
26.9 | 31.3 | 34.5667 | 37.3 | 37.9
27 | 31.1333 | 34.4 | 37.1167 | 37.8
27.1 | 30.9833 | 34.3 | 37 | 37.6
27.2 | 30.9 | 34.18 | 36.85 | 37.4167
27.3 | 30.7 | 33.9 | 36.725 | 37.24
27.4 | 30.5 | 33.8 | 36.6 | 37.125
27.5 | 30.3667 | 33.7 | 36.4 | 37
27.6 | 30.3 | 33.56 | 36.2833 | 36.86
27.7 | 30.1 | 33.35 | 36.1 | 36.7
27.8 | 30.04 | 33.14 | 35.9333 | 36.58
27.9 | 29.8 | 32.96 | 35.8 | 36.4143
28 | 29.65 | 32.8333 | 35.6833 | 36.24
28.1 | 29.5 | 32.7167 | 35.5 | 36.0667
28.2 | 29.375 | 32.6 | 35.38 | 35.9
28.3 | 29.25 | 32.4833 | 35.26 | 35.7
28.4 | 29.08 | 32.3 | 35.1 | 35.66
28.5 | 28.98 | 32.2 | 34.9 | 35.52
28.6 | 28.8 | 32.0571 | 34.7333 | 35.3
28.7 | 28.7 | 31.8833 | 34.65 | 35.22
28.8 | 28.6167 | 31.7 | 34.46 | 35.1
28.9 | 28.4167 | 31.6 | 34.36 | 34.95
29 | 28.2 | 31.4167 | 34.18 | 34.74
29.1 | 28.1 | 31.3 | 34 | 34.6
29.2 | 28.02 | 31.2 | 33.9 | 34.475
29.3 | 27.95 | 31.05 | 33.8 | 34.35
29.4 | 27.88 | 30.9 | 33.6 | 34.225
29.5 | 27.675 | 30.8 | 33.46 | 34.1
29.6 | 27.54 | 30.6333 | 33.2 | 33.9
29.7 | 27.4 | 30.5 | 33.1833 | 33.8
29.8 | 27.2167 | 30.38 | 33.1 | 33.7
29.9 | 27.1 | 30.2 | 32.9 | 33.5
30 | 27.06 | 30.05 | 32.775 | 33.36

## SSD80 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 27.3 | 36.2 | 51.9 | 61.6
0.1 | 32.35 | 40.55 | 56.1 | 65.45
0.2 | 39.225 | 47.3 | 62.1 | 70.7
0.3 | 48.6 | 55.7 | 69.2857 | 77.125
0.4 | 57.6 | 63.55 | 75.76 | 82.225
0.5 | 64.8 | 70.5 | 80.975 | 86.475
0.6 | 70.6333 | 75.8 | 84.9375 | 89.6
0.7 | 75.8 | 80.6 | 88.2375 | 92.38
0.8 | 80.4 | 84.3 | 91.15 | 94.15
0.9 | 83.9 | 87.4 | 93.325 | 95.9
1 | 87 | 89.82 | 95 | 97.15
1.1 | 89.7 | 92.05 | 96.3 | 98.1
1.2 | 91.8333 | 93.825 | 97.5 | 98.8
1.3 | 93.64 | 95.45 | 98.3 | 99.3
1.4 | 95.15 | 96.7 | 98.88 | 99.64
1.5 | 96.45 | 97.6 | 99.35 | 99.88
1.6 | 97.66 | 98.35 | 99.72 | 100
1.7 | 98.35 | 98.95 | 99.9 | 100
1.8 | 98.975 | 99.375 | 100 | 100
1.9 | 99.475 | 99.7333 | 100 | 99.9
2 | 99.7143 | 99.92 | 100 | 99.8
2.1 | 99.925 | 100 | 99.9 | 99.575
2.2 | 100 | 100 | 99.7571 | 99.35
2.3 | 100 | 100 | 99.55 | 99.1
2.4 | 99.9556 | 99.8 | 99.3 | 98.825
2.5 | 99.88 | 99.7 | 98.98 | 98.475
2.6 | 99.8 | 99.55 | 98.7222 | 98.14
2.7 | 99.5625 | 99.3 | 98.375 | 97.9
2.8 | 99.225 | 98.95 | 98.0875 | 97.5
2.9 | 99 | 98.7 | 97.7125 | 97.1
3 | 98.58 | 98.375 | 97.3857 | 96.75
3.1 | 98.175 | 98 | 96.9125 | 96.3
3.2 | 97.8875 | 97.65 | 96.5667 | 95.95
3.3 | 97.4857 | 97.25 | 96.2 | 95.6
3.4 | 96.9667 | 96.875 | 95.725 | 95.22
3.5 | 96.5333 | 96.525 | 95.42 | 94.75
3.6 | 96.1 | 96.05 | 95.05 | 94.35
3.7 | 95.65 | 95.625 | 94.6 | 94.025
3.8 | 95.225 | 95.15 | 94.25 | 93.65
3.9 | 94.75 | 94.75 | 93.85 | 93.15
4 | 94.175 | 94.35 | 93.2857 | 92.7625
4.1 | 93.85 | 93.95 | 92.9 | 92.425
4.2 | 93.25 | 93.5 | 92.55 | 91.95
4.3 | 92.92 | 93 | 92.05 | 91.5667
4.4 | 92.3 | 92.55 | 91.65 | 91.2
4.5 | 91.8 | 92.1 | 91.325 | 90.75
4.6 | 91.3 | 91.7286 | 90.8857 | 90.35
4.7 | 90.88 | 91.25 | 90.5333 | 89.8667
4.8 | 90.35 | 90.9 | 90.075 | 89.6
4.9 | 89.84 | 90.4 | 89.65 | 89.2
5 | 89.35 | 89.98 | 89.175 | 88.8
5.1 | 89 | 89.55 | 88.9 | 88.35
5.2 | 88.65 | 89.075 | 88.475 | 88.04
5.3 | 88.1 | 88.6667 | 88.1667 | 87.64
5.4 | 87.48 | 88.3 | 87.8 | 87.175
5.5 | 87.025 | 88 | 87.25 | 86.88
5.6 | 86.66 | 87.45 | 86.88 | 86.35
5.7 | 86.2 | 87 | 86.45 | 85.95
5.8 | 85.7 | 86.62 | 86.1 | 85.65
5.9 | 85.3333 | 86.2 | 85.72 | 85.2
6 | 84.85 | 85.75 | 85.3 | 84.9
6.1 | 84.3 | 85.325 | 84.9667 | 84.62
6.2 | 83.85 | 84.95 | 84.575 | 84.15
6.3 | 83.525 | 84.5625 | 84.2 | 83.75
6.4 | 83.1 | 84.15 | 83.8 | 83.375
6.5 | 82.6 | 83.7125 | 83.4 | 83.04
6.6 | 82.1667 | 83.3375 | 83.06 | 82.7
6.7 | 81.7 | 82.9 | 82.7286 | 82.3
6.8 | 81.4 | 82.45 | 82.35 | 81.95
6.9 | 80.9 | 82.14 | 81.95 | 81.54
7 | 80.35 | 81.775 | 81.5 | 81.175
7.1 | 80 | 81.35 | 81.2 | 80.85
7.2 | 79.6 | 80.95 | 80.9 | 80.5
7.3 | 79.1 | 80.45 | 80.4 | 80.1
7.4 | 78.725 | 80.1 | 80.125 | 79.6857
7.5 | 78.3333 | 79.8 | 79.75 | 79.4444
7.6 | 77.75 | 79.4 | 79.325 | 79.06
7.7 | 77.44 | 79.05 | 79.06 | 78.7
7.8 | 77.14 | 78.66 | 78.75 | 78.3667
7.9 | 76.65 | 78.24 | 78.3 | 78
8 | 76.2 | 77.85 | 78 | 77.675
8.1 | 75.775 | 77.4 | 77.55 | 77.325
8.2 | 75.425 | 77.05 | 77.175 | 76.95
8.3 | 75.0222 | 76.7 | 76.9222 | 76.6
8.4 | 74.6125 | 76.35 | 76.55 | 76.36
8.5 | 74.2375 | 75.95 | 76.1667 | 76
8.6 | 73.6625 | 75.55 | 75.875 | 75.6
8.7 | 73.3 | 75.15 | 75.5667 | 75.3
8.8 | 72.95 | 74.76 | 75.2 | 74.95
8.9 | 72.575 | 74.42 | 74.8 | 74.6
9 | 72.2 | 74.125 | 74.45 | 74.2
9.1 | 71.825 | 73.7333 | 74.12 | 73.9
9.2 | 71.4222 | 73.2875 | 73.75 | 73.6
9.3 | 71.0125 | 72.9125 | 73.3857 | 73.25
9.4 | 70.6 | 72.6 | 73.025 | 72.95
9.5 | 70.3375 | 72.2 | 72.7333 | 72.55
9.6 | 69.7714 | 71.85 | 72.38 | 72.26
9.7 | 69.3333 | 71.525 | 72.1 | 71.88
9.8 | 69.0714 | 71.15 | 71.7 | 71.58
9.9 | 68.72 | 70.65 | 71.35 | 71.25
10 | 68.3 | 70.375 | 71 | 70.9333
10.1 | 67.9 | 70.125 | 70.7 | 70.6
10.2 | 67.55 | 69.76 | 70.3 | 70.275
10.3 | 67.2 | 69.35 | 70.02 | 70.025
10.4 | 66.85 | 69 | 69.65 | 69.6125
10.5 | 66.52 | 68.6667 | 69.375 | 69.325
10.6 | 66.05 | 68.34 | 69.125 | 68.9778
10.7 | 65.825 | 67.9857 | 68.68 | 68.7429
10.8 | 65.4 | 67.7143 | 68.45 | 68.5
10.9 | 65.05 | 67.2667 | 68.1 | 68.125
11 | 64.6 | 66.9125 | 67.8 | 67.75
11.1 | 64.4 | 66.625 | 67.5333 | 67.45
11.2 | 64.075 | 66.3625 | 67.15 | 67.1333
11.3 | 63.775 | 65.9875 | 66.875 | 66.8667
11.4 | 63.425 | 65.6333 | 66.525 | 66.575
11.5 | 63.15 | 65.35 | 66.3 | 66.225
11.6 | 62.65 | 65 | 65.8667 | 65.8833
11.7 | 62.2 | 64.725 | 65.7 | 65.6333
11.8 | 61.95 | 64.4 | 65.3 | 65.35
11.9 | 61.6 | 64 | 64.95 | 65.1
12 | 61.25 | 63.64 | 64.7 | 64.725
12.1 | 61 | 63.3429 | 64.3857 | 64.45
12.2 | 60.7 | 63.0333 | 64.15 | 64.15
12.3 | 60.3 | 62.6833 | 63.8286 | 63.875
12.4 | 60.05 | 62.45 | 63.475 | 63.625
12.5 | 59.7 | 62.025 | 63.225 | 63.375
12.6 | 59.35 | 61.8 | 62.88 | 63.05
12.7 | 59 | 61.5 | 62.65 | 62.68
12.8 | 58.8 | 61.2 | 62.4 | 62.45
12.9 | 58.46 | 60.92 | 62.05 | 62.2
13 | 58.05 | 60.65 | 61.8 | 61.85
13.1 | 57.775 | 60.325 | 61.5 | 61.5
13.2 | 57.525 | 60 | 61.2 | 61.375
13.3 | 57.2333 | 59.75 | 60.95 | 61.025
13.4 | 56.86 | 59.4 | 60.72 | 60.775
13.5 | 56.6333 | 59.05 | 60.375 | 60.4875
13.6 | 56.3667 | 58.8 | 60.125 | 60.175
13.7 | 55.9 | 58.5333 | 59.8667 | 59.925
13.8 | 55.6333 | 58.15 | 59.45 | 59.6
13.9 | 55.34 | 57.9 | 59.26 | 59.4429
14 | 55.075 | 57.65 | 59.05 | 59.075
14.1 | 54.74 | 57.375 | 58.725 | 58.825
14.2 | 54.5 | 57.1 | 58.475 | 58.58
14.3 | 54.25 | 56.8 | 58.125 | 58.35
14.4 | 53.8 | 56.45 | 57.875 | 57.98
14.5 | 53.575 | 56.175 | 57.65 | 57.8
14.6 | 53.225 | 55.9 | 57.35 | 57.525
14.7 | 53.05 | 55.65 | 57.0333 | 57.275
14.8 | 52.725 | 55.34 | 56.8 | 56.9375
14.9 | 52.5 | 55.1 | 56.55 | 56.675
15 | 52.25 | 54.85 | 56.3 | 56.3875
15.1 | 51.9 | 54.5 | 56 | 56.1375
15.2 | 51.65 | 54.2 | 55.7 | 55.925
15.3 | 51.3 | 53.9 | 55.55 | 55.75
15.4 | 51.05 | 53.675 | 55.15 | 55.4
15.5 | 50.8 | 53.3667 | 54.875 | 55.15
15.6 | 50.55 | 53.2 | 54.7 | 54.925
15.7 | 50.2 | 52.9 | 54.4333 | 54.65
15.8 | 49.9333 | 52.5667 | 54.1 | 54.425
15.9 | 49.6333 | 52.3 | 53.9333 | 54.175
16 | 49.5 | 52.12 | 53.6 | 53.85
16.1 | 49.25 | 51.85 | 53.45 | 53.675
16.2 | 48.9667 | 51.6 | 53.1667 | 53.44
16.3 | 48.725 | 51.3 | 52.92 | 53.22
16.4 | 48.44 | 51.1 | 52.74 | 53
16.5 | 48.2 | 50.8 | 52.5 | 52.75
16.6 | 47.85 | 50.6 | 52.2 | 52.5
16.7 | 47.6 | 50.3 | 51.925 | 52.16
16.8 | 47.4 | 50.1 | 51.75 | 52
16.9 | 47.2 | 49.75 | 51.4 | 51.775
17 | 46.9 | 49.5 | 51.275 | 51.525
17.1 | 46.6 | 49.275 | 50.95 | 51.275
17.2 | 46.425 | 49.1 | 50.7 | 51.025
17.3 | 46.1 | 48.85 | 50.525 | 50.775
17.4 | 45.95 | 48.6 | 50.275 | 50.6
17.5 | 45.7 | 48.35 | 50.025 | 50.4
17.6 | 45.45 | 48.12 | 49.875 | 50.15
17.7 | 45.2 | 47.875 | 49.625 | 49.9
17.8 | 45 | 47.6 | 49.4 | 49.6667
17.9 | 44.7 | 47.45 | 49.125 | 49.4833
18 | 44.52 | 47.2 | 48.875 | 49.2333
18.1 | 44.3 | 46.975 | 48.64 | 49.025
18.2 | 44.075 | 46.625 | 48.5 | 48.8
18.3 | 43.825 | 46.4667 | 48.2 | 48.55
18.4 | 43.575 | 46.175 | 48 | 48.325
18.5 | 43.325 | 45.925 | 47.85 | 48.1444
18.6 | 43.075 | 45.775 | 47.6 | 47.95
18.7 | 42.9 | 45.525 | 47.35 | 47.7
18.8 | 42.7 | 45.225 | 47.12 | 47.45
18.9 | 42.45 | 45.075 | 46.875 | 47.25
19 | 42.3 | 44.9 | 46.625 | 47.05
19.1 | 41.96 | 44.575 | 46.475 | 46.8
19.2 | 41.8 | 44.325 | 46.225 | 46.56
19.3 | 41.575 | 44.175 | 45.975 | 46.425
19.4 | 41.325 | 43.94 | 45.825 | 46.175
19.5 | 41.075 | 43.775 | 45.6 | 45.98
19.6 | 40.925 | 43.5 | 45.4333 | 45.8
19.7 | 40.7 | 43.25 | 45.1667 | 45.6
19.8 | 40.55 | 43.1 | 44.875 | 45.35
19.9 | 40.3 | 42.8333 | 44.7 | 45.1
20 | 40.1 | 42.6 | 44.5 | 44.975
20.1 | 39.85 | 42.4 | 44.35 | 44.8
20.2 | 39.6 | 42.225 | 44.1 | 44.575
20.3 | 39.45 | 41.975 | 43.9 | 44.325
20.4 | 39.2 | 41.725 | 43.725 | 44.18
20.5 | 39.025 | 41.6 | 43.5 | 43.95
20.6 | 38.86 | 41.35 | 43.35 | 43.7429
20.7 | 38.625 | 41.2 | 43.125 | 43.525
20.8 | 38.475 | 40.9 | 42.9 | 43.275
20.9 | 38.225 | 40.75 | 42.75 | 43.1222
21 | 38 | 40.55 | 42.52 | 42.9
21.1 | 37.825 | 40.35 | 42.275 | 42.76
21.2 | 37.6286 | 40.1571 | 42.1 | 42.55
21.3 | 37.44 | 39.9875 | 41.9 | 42.325
21.4 | 37.3 | 39.825 | 41.75 | 42.175
21.5 | 37.1 | 39.575 | 41.5 | 41.975
21.6 | 36.9 | 39.4125 | 41.35 | 41.8
21.7 | 36.6 | 39.15 | 41.1 | 41.5857
21.8 | 36.44 | 39 | 40.975 | 41.35
21.9 | 36.3 | 38.86 | 40.8 | 41.2
22 | 36.1375 | 38.625 | 40.5667 | 41.05
22.1 | 35.925 | 38.4 | 40.3 | 40.775
22.2 | 35.75 | 38.225 | 40.175 | 40.65
22.3 | 35.6 | 38 | 40 | 40.4
22.4 | 35.4 | 37.84 | 39.775 | 40.25
22.5 | 35.2 | 37.7 | 39.625 | 40.1
22.6 | 35 | 37.45 | 39.475 | 39.95
22.7 | 34.8 | 37.3 | 39.35 | 39.7
22.8 | 34.675 | 37.05 | 39.1 | 39.55
22.9 | 34.5 | 36.9 | 38.85 | 39.4
23 | 34.275 | 36.76 | 38.7 | 39.25
23.1 | 34.14 | 36.6 | 38.55 | 39.075
23.2 | 34 | 36.35 | 38.4 | 38.8571
23.3 | 33.75 | 36.2 | 38.175 | 38.675
23.4 | 33.5 | 36.05 | 38 | 38.5
23.5 | 33.4 | 35.9 | 37.875 | 38.3
23.6 | 33.3 | 35.65 | 37.625 | 38.125
23.7 | 33.1 | 35.5 | 37.5 | 38
23.8 | 32.9 | 35.36 | 37.325 | 37.8333
23.9 | 32.8 | 35.2 | 37.2 | 37.6
24 | 32.62 | 34.975 | 36.975 | 37.45
24.1 | 32.375 | 34.825 | 36.85 | 37.3
24.2 | 32.225 | 34.675 | 36.65 | 37.1
24.3 | 32.075 | 34.425 | 36.5 | 36.95
24.4 | 31.925 | 34.3 | 36.2875 | 36.7
24.5 | 31.7 | 34.15 | 36.1625 | 36.65
24.6 | 31.6 | 34 | 36.0375 | 36.52
24.7 | 31.5 | 33.8 | 35.825 | 36.36
24.8 | 31.25 | 33.6 | 35.6 | 36.125
24.9 | 31.1 | 33.475 | 35.4 | 35.975
25 | 30.9 | 33.2833 | 35.28 | 35.8
25.1 | 30.8 | 33.08 | 35.125 | 35.58
25.2 | 30.625 | 32.9 | 34.95 | 35.45
25.3 | 30.475 | 32.8 | 34.775 | 35.3
25.4 | 30.38 | 32.65 | 34.6 | 35.2
25.5 | 30.15 | 32.5 | 34.475 | 35
25.6 | 30 | 32.375 | 34.3 | 34.86
25.7 | 29.85 | 32.2 | 34.1833 | 34.7
25.8 | 29.7 | 32 | 34 | 34.525
25.9 | 29.56 | 31.84 | 33.9 | 34.375
26 | 29.4 | 31.7 | 33.7 | 34.1778
26.1 | 29.275 | 31.55 | 33.6 | 34.0571
26.2 | 29.15 | 31.4 | 33.4 | 33.875
26.3 | 29.025 | 31.2667 | 33.2 | 33.725
26.4 | 28.8 | 31.1 | 33.1 | 33.6375
26.5 | 28.675 | 30.95 | 32.9 | 33.4625
26.6 | 28.55 | 30.8 | 32.7 | 33.3167
26.7 | 28.4333 | 30.65 | 32.575 | 33.1
26.8 | 28.3 | 30.5 | 32.5 | 33.0333
26.9 | 28.2 | 30.36 | 32.3 | 32.9
27 | 27.95 | 30.225 | 32.125 | 32.7
27.1 | 27.8 | 30.075 | 32.0375 | 32.6
27.2 | 27.7 | 29.925 | 31.9 | 32.45
27.3 | 27.6 | 29.8 | 31.75 | 32.3
27.4 | 27.3667 | 29.625 | 31.6 | 32.1
27.5 | 27.3 | 29.4 | 31.48 | 32
27.6 | 27.175 | 29.3 | 31.3 | 31.9
27.7 | 27.025 | 29.15 | 31.2 | 31.7
27.8 | 26.9 | 29 | 31.05 | 31.5
27.9 | 26.8 | 28.9 | 30.9 | 31.4
28 | 26.6 | 28.8 | 30.75 | 31.225
28.1 | 26.45 | 28.68 | 30.6 | 31.1
28.2 | 26.375 | 28.5 | 30.45 | 31
28.3 | 26.24 | 28.35 | 30.4 | 30.875
28.4 | 26.075 | 28.225 | 30.175 | 30.7
28.5 | 25.9 | 28.1 | 30.025 | 30.6
28.6 | 25.8 | 27.975 | 29.9 | 30.45
28.7 | 25.65 | 27.8 | 29.8 | 30.3
28.8 | 25.5 | 27.675 | 29.6 | 30.175
28.9 | 25.4 | 27.6 | 29.4 | 30
29 | 25.3 | 27.4 | 29.3 | 29.9
29.1 | 25.2 | 27.3 | 29.2 | 29.75
29.2 | 25 | 27.125 | 29.1 | 29.6
29.3 | 24.9 | 27 | 28.925 | 29.5
29.4 | 24.8 | 26.9 | 28.8 | 29.4
29.5 | 24.675 | 26.75 | 28.7 | 29.275
29.6 | 24.525 | 26.7 | 28.5 | 29.1571
29.7 | 24.475 | 26.5 | 28.4 | 29.025
29.8 | 24.325 | 26.4 | 28.3 | 28.9
29.9 | 24.2 | 26.3 | 28.15 | 28.75
30 | 24 | 26.15 | 28.1 | 28.6

## SSD120 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 26.5 | 32.9 | 41.8 | 48
0.1 | 31.15 | 36.9 | 45.9 | 52.0667
0.2 | 37.925 | 43.6 | 52.2 | 57.6667
0.3 | 47.4 | 52.5 | 60.3 | 65.1
0.4 | 55.4 | 60.525 | 67 | 71.7
0.5 | 62.6333 | 66.7 | 72.6333 | 77.1
0.6 | 68.7 | 72.3 | 77.7 | 81.3
0.7 | 73.8333 | 76.9 | 81.7667 | 84.65
0.8 | 78.0667 | 81 | 85.1 | 87.54
0.9 | 81.8667 | 84.2 | 88 | 89.9
1 | 84.8333 | 87.3 | 90.15 | 92.15
1.1 | 87.8 | 89.6 | 92.1875 | 93.725
1.2 | 90.15 | 91.75 | 93.9125 | 95.05
1.3 | 92.0333 | 93.275 | 95.225 | 96.3
1.4 | 93.8 | 94.775 | 96.3 | 97.3667
1.5 | 95.1 | 95.925 | 97.38 | 98.025
1.6 | 96.3 | 96.925 | 98.05 | 98.6333
1.7 | 97.4 | 97.75 | 98.7 | 99.15
1.8 | 97.9 | 98.5 | 99 | 99.4667
1.9 | 98.65 | 99.05 | 99.4667 | 99.7333
2 | 99.2 | 99.4667 | 99.7 | 99.9
2.1 | 99.525 | 99.75 | 99.9 | 100
2.2 | 99.7 | 99.9 | 100 | 100
2.3 | 99.925 | 100 | 100 | 100
2.4 | 100 | 100 | 100 | 99.9
2.5 | 100 | 100 | 99.9 | 99.85
2.6 | 100 | 99.9 | 99.8 | 99.62
2.7 | 99.9 | 99.8 | 99.7 | 99.48
2.8 | 99.8 | 99.7333 | 99.7 | 99.36
2.9 | 99.6 | 99.6 | 99.52 | 99.24
3 | 99.35 | 99.5 | 99.2 | 98.96
3.1 | 99.2 | 99.3333 | 98.925 | 98.7429
3.2 | 98.95 | 99.08 | 98.5833 | 98.32
3.3 | 98.64 | 98.725 | 98.475 | 98.075
3.4 | 98.36 | 98.3 | 98.2 | 97.7333
3.5 | 98.025 | 98.15 | 97.75 | 97.4
3.6 | 97.75 | 97.7667 | 97.4625 | 97.2
3.7 | 97.25 | 97.5 | 97.1 | 96.8667
3.8 | 96.925 | 97.1 | 96.9 | 96.5333
3.9 | 96.525 | 96.8 | 96.6 | 96.2
4 | 96.075 | 96.5333 | 96.26 | 95.95
4.1 | 95.7333 | 96.25 | 95.95 | 95.7
4.2 | 95.35 | 95.7 | 95.6 | 95.3
4.3 | 94.9 | 95.4 | 95.25 | 95
4.4 | 94.5 | 95.1 | 94.9 | 94.65
4.5 | 94.1 | 94.6 | 94.62 | 94.4
4.6 | 93.8 | 94.4 | 94.22 | 94.05
4.7 | 93.4333 | 94.125 | 93.98 | 93.775
4.8 | 92.98 | 93.7333 | 93.675 | 93.45
4.9 | 92.4 | 93.4 | 93.4 | 93.1
5 | 92.05 | 93 | 92.975 | 92.85
5.1 | 91.75 | 92.56 | 92.625 | 92.4
5.2 | 91.375 | 92.325 | 92.45 | 92.2
5.3 | 91.05 | 92.05 | 92.025 | 91.8
5.4 | 90.66 | 91.675 | 91.8 | 91.475
5.5 | 90.1 | 91.175 | 91.55 | 91.4
5.6 | 89.7 | 90.9 | 91.1 | 90.925
5.7 | 89.45 | 90.6 | 90.8 | 90.5667
5.8 | 89 | 90.2 | 90.4 | 90.275
5.9 | 88.45 | 89.8 | 90.15 | 89.98
6 | 88.1 | 89.4 | 89.9 | 89.66
6.1 | 87.775 | 89 | 89.525 | 89.425
6.2 | 87.4625 | 88.8 | 89.175 | 89.175
6.3 | 87.05 | 88.6 | 88.925 | 88.94
6.4 | 86.75 | 88.25 | 88.6 | 88.7
6.5 | 86.375 | 87.775 | 88.2429 | 88.2
6.6 | 85.875 | 87.4 | 87.9 | 87.96
6.7 | 85.35 | 87.2 | 87.65 | 87.725
6.8 | 85.14 | 86.85 | 87.4 | 87.375
6.9 | 84.8 | 86.425 | 87.075 | 87.125
7 | 84.5 | 86.175 | 86.8222 | 86.625
7.1 | 84.1 | 85.8667 | 86.4625 | 86.4
7.2 | 83.75 | 85.48 | 86.125 | 86.1
7.3 | 83.1 | 85.05 | 85.875 | 85.8
7.4 | 82.8 | 84.9 | 85.6444 | 85.5
7.5 | 82.52 | 84.6 | 85.3 | 85.2
7.6 | 82.1 | 84.2 | 84.9667 | 85
7.7 | 81.575 | 83.86 | 84.55 | 84.725
7.8 | 81.375 | 83.6 | 84.2 | 84.275
7.9 | 80.975 | 83.275 | 84.075 | 84.025
8 | 80.6 | 82.925 | 83.825 | 83.8
8.1 | 80.3 | 82.58 | 83.56 | 83.55
8.2 | 80.1 | 82.3 | 83.1 | 83.2
8.3 | 79.85 | 81.9 | 82.9 | 82.875
8.4 | 79.325 | 81.65 | 82.6 | 82.625
8.5 | 78.75 | 81.3667 | 82.2 | 82.275
8.6 | 78.5 | 81.1 | 81.875 | 81.94
8.7 | 78.2 | 80.575 | 81.625 | 81.7
8.8 | 77.9 | 80.2 | 81.3625 | 81.35
8.9 | 77.65 | 79.975 | 81.05 | 81.1
9 | 77.3 | 79.725 | 80.85 | 80.8
9.1 | 77.05 | 79.4333 | 80.58 | 80.5
9.2 | 76.5 | 79.075 | 80.3 | 80.1667
9.3 | 76.175 | 78.875 | 80.04 | 80
9.4 | 75.85 | 78.525 | 79.7 | 79.6667
9.5 | 75.5333 | 78.175 | 79.4 | 79.4
9.6 | 75.05 | 77.85 | 79 | 79.1
9.7 | 74.7 | 77.475 | 78.85 | 78.85
9.8 | 74.4 | 77.25 | 78.5 | 78.45
9.9 | 74.05 | 76.9667 | 78.15 | 78.15
10 | 73.7 | 76.6 | 77.8 | 77.925
10.1 | 73.26 | 76.35 | 77.675 | 77.6
10.2 | 73 | 76 | 77.44 | 77.4
10.3 | 72.65 | 75.65 | 77.15 | 77.1
10.4 | 72.24 | 75.4 | 76.84 | 76.74
10.5 | 71.975 | 75 | 76.5 | 76.5
10.6 | 71.7 | 74.725 | 76.25 | 76.2667
10.7 | 71.275 | 74.45 | 75.85 | 76.1
10.8 | 71 | 74.15 | 75.625 | 75.8
10.9 | 70.75 | 73.85 | 75.4143 | 75.5
11 | 70.375 | 73.525 | 75.15 | 75.2
11.1 | 70 | 73.275 | 74.9 | 74.975
11.2 | 69.75 | 73.04 | 74.625 | 74.7
11.3 | 69.5 | 72.7 | 74.2667 | 74.4667
11.4 | 69.175 | 72.35 | 73.9125 | 74.225
11.5 | 68.9 | 72.08 | 73.7125 | 73.9333
11.6 | 68.4 | 71.825 | 73.5667 | 73.6
11.7 | 68.125 | 71.5 | 73.28 | 73.26
11.8 | 67.825 | 71.25 | 73.05 | 73.05
11.9 | 67.5 | 71.1 | 72.7 | 72.875
12 | 67.15 | 70.775 | 72.45 | 72.55
12.1 | 67 | 70.525 | 72.1 | 72.28
12.2 | 66.65 | 70.175 | 71.775 | 72.15
12.3 | 66.15 | 69.9 | 71.525 | 71.8
12.4 | 65.86 | 69.625 | 71.3 | 71.5
12.5 | 65.625 | 69.44 | 71.125 | 71.2
12.6 | 65.275 | 69.0667 | 70.8 | 71.0667
12.7 | 65.025 | 68.75 | 70.54 | 70.9
12.8 | 64.88 | 68.5 | 70.3 | 70.6
12.9 | 64.65 | 68.0667 | 69.975 | 70.3333
13 | 64.3 | 67.8 | 69.7 | 69.9667
13.1 | 64.05 | 67.66 | 69.45 | 69.66
13.2 | 63.7 | 67.3 | 69.2 | 69.425
13.3 | 63.35 | 67.025 | 68.9667 | 69.175
13.4 | 63 | 66.775 | 68.6 | 68.95
13.5 | 62.75 | 66.525 | 68.3833 | 68.725
13.6 | 62.425 | 66.2667 | 68.1833 | 68.5
13.7 | 62.125 | 66 | 67.975 | 68.25
13.8 | 61.775 | 65.66 | 67.725 | 68.1
13.9 | 61.6 | 65.3333 | 67.4667 | 67.65
14 | 61.3167 | 65.15 | 67.1667 | 67.42
14.1 | 61.02 | 64.825 | 66.8 | 67.275
14.2 | 60.66 | 64.575 | 66.55 | 66.85
14.3 | 60.45 | 64.325 | 66.4 | 66.6
14.4 | 60.125 | 64.05 | 66.05 | 66.45
14.5 | 59.9 | 63.675 | 65.775 | 66.2
14.6 | 59.55 | 63.44 | 65.55 | 65.875
14.7 | 59.3 | 63.2 | 65.3 | 65.75
14.8 | 59.075 | 62.9 | 65.15 | 65.5
14.9 | 58.725 | 62.7 | 64.9 | 65.15
15 | 58.375 | 62.45 | 64.65 | 65
15.1 | 58.05 | 62.125 | 64.5 | 64.7
15.2 | 57.65 | 61.875 | 64.15 | 64.5
15.3 | 57.4 | 61.625 | 63.9 | 64.25
15.4 | 57.2 | 61.4 | 63.7 | 64.1
15.5 | 57.025 | 61.15 | 63.4 | 63.85
15.6 | 56.8 | 60.85 | 63 | 63.6
15.7 | 56.55 | 60.7 | 62.8 | 63.35
15.8 | 56.3 | 60.5 | 62.6 | 63.12
15.9 | 56.0667 | 60.3 | 62.4333 | 62.875
16 | 55.8 | 59.875 | 62.1667 | 62.625
16.1 | 55.5 | 59.625 | 62 | 62.475
16.2 | 55.1667 | 59.4 | 61.7 | 62.25
16.3 | 54.9 | 59.1 | 61.5 | 61.9667
16.4 | 54.65 | 58.8429 | 61.35 | 61.75
16.5 | 54.4 | 58.575 | 60.975 | 61.5
16.6 | 54.3 | 58.4125 | 60.8 | 61.25
16.7 | 54.1 | 58.1 | 60.54 | 61
16.8 | 53.8 | 57.95 | 60.4 | 60.75
16.9 | 53.5 | 57.725 | 60.1 | 60.6
17 | 53.26 | 57.475 | 59.9 | 60.45
17.1 | 53.025 | 57.225 | 59.6889 | 60.05
17.2 | 52.8 | 56.975 | 59.45 | 59.9
17.3 | 52.625 | 56.725 | 59.2 | 59.68
17.4 | 52.3 | 56.475 | 59.025 | 59.375
17.5 | 52.125 | 56.2 | 58.6857 | 59.125
17.6 | 52 | 56 | 58.4714 | 58.9625
17.7 | 51.64 | 55.82 | 58.3 | 58.8
17.8 | 51.4 | 55.675 | 58.15 | 58.575
17.9 | 51.1333 | 55.35 | 57.825 | 58.4
18 | 50.9 | 55.0667 | 57.64 | 58.15
18.1 | 50.65 | 54.825 | 57.48 | 57.9
18.2 | 50.3333 | 54.675 | 57.15 | 57.7429
18.3 | 50.2 | 54.425 | 57 | 57.5
18.4 | 50.0667 | 54.1 | 56.75 | 57.25
18.5 | 49.8 | 53.925 | 56.5 | 57.05
18.6 | 49.5 | 53.7 | 56.25 | 56.725
18.7 | 49.4 | 53.5 | 56 | 56.6
18.8 | 49.275 | 53.275 | 55.8 | 56.45
18.9 | 48.94 | 52.9667 | 55.5 | 56.2
19 | 48.6 | 52.78 | 55.4 | 55.9
19.1 | 48.425 | 52.55 | 55.2 | 55.65
19.2 | 48.08 | 52.3 | 55.05 | 55.525
19.3 | 47.8833 | 52.1 | 54.7 | 55.3
19.4 | 47.7167 | 51.85 | 54.58 | 55.1
19.5 | 47.5 | 51.7 | 54.3333 | 54.9
19.6 | 47.3 | 51.4 | 54.1 | 54.72
19.7 | 47 | 51.175 | 53.95 | 54.5
19.8 | 46.72 | 51 | 53.7 | 54.3
19.9 | 46.65 | 50.775 | 53.55 | 54.05
20 | 46.4 | 50.575 | 53.32 | 53.85
20.1 | 46.2 | 50.325 | 53.1 | 53.65
20.2 | 45.95 | 50.075 | 52.8 | 53.4
20.3 | 45.72 | 49.825 | 52.65 | 53.26
20.4 | 45.6 | 49.7167 | 52.5 | 53.025
20.5 | 45.5 | 49.58 | 52.25 | 52.7667
20.6 | 45.2 | 49.4 | 52.0333 | 52.675
20.7 | 44.84 | 49.15 | 51.8 | 52.425
20.8 | 44.7 | 48.84 | 51.7 | 52.2
20.9 | 44.5778 | 48.7 | 51.5 | 52.05
21 | 44.4 | 48.55 | 51.375 | 51.8
21.1 | 44.3 | 48.3 | 51.125 | 51.675
21.2 | 44.0333 | 48.1571 | 50.875 | 51.5
21.3 | 43.8 | 47.9 | 50.6625 | 51.4
21.4 | 43.58 | 47.675 | 50.475 | 51.1333
21.5 | 43.2 | 47.4 | 50.3125 | 50.9
21.6 | 43.0778 | 47.1 | 50.0875 | 50.65
21.7 | 42.95 | 46.95 | 49.9625 | 50.5
21.8 | 42.8556 | 46.7333 | 49.8 | 50.325
21.9 | 42.625 | 46.6 | 49.625 | 50.1
22 | 42.3333 | 46.475 | 49.38 | 49.875
22.1 | 42.225 | 46.3 | 49.125 | 49.7111
22.2 | 41.975 | 46.075 | 48.875 | 49.5375
22.3 | 41.75 | 45.825 | 48.725 | 49.4125
22.4 | 41.6 | 45.6 | 48.4667 | 49.2
22.5 | 41.44 | 45.35 | 48.3 | 49
22.6 | 41.18 | 45.2 | 48.125 | 48.825
22.7 | 41 | 45.05 | 47.9 | 48.65
22.8 | 40.8 | 44.9 | 47.8 | 48.4
22.9 | 40.65 | 44.7 | 47.55 | 48.2
23 | 40.4 | 44.5 | 47.4 | 48.02
23.1 | 40.24 | 44.375 | 47.2 | 47.8667
23.2 | 40.1 | 44.2 | 47 | 47.7
23.3 | 40 | 44.075 | 46.9 | 47.5
23.4 | 39.8 | 43.9 | 46.7 | 47.275
23.5 | 39.55 | 43.675 | 46.6 | 47.1875
23.6 | 39.4 | 43.425 | 46.375 | 47.025
23.7 | 39.2 | 43.2 | 46.2 | 46.875
23.8 | 39.025 | 43.05 | 46 | 46.725
23.9 | 38.875 | 42.9 | 45.875 | 46.5
24 | 38.7 | 42.7 | 45.6625 | 46.25
24.1 | 38.5 | 42.5 | 45.4889 | 46
24.2 | 38.325 | 42.3 | 45.325 | 45.9333
24.3 | 38.1 | 42.2 | 45.1 | 45.76
24.4 | 37.9667 | 42.1 | 44.8778 | 45.525
24.5 | 37.775 | 41.92 | 44.7625 | 45.4
24.6 | 37.6 | 41.6667 | 44.575 | 45.225
24.7 | 37.5 | 41.52 | 44.4125 | 45.075
24.8 | 37.4 | 41.275 | 44.1875 | 44.825
24.9 | 37.175 | 41.1 | 44.025 | 44.6
25 | 37.025 | 41 | 43.8 | 44.55
25.1 | 36.8 | 40.7 | 43.625 | 44.4
25.2 | 36.58 | 40.6 | 43.575 | 44.2
25.3 | 36.4143 | 40.425 | 43.4 | 44.1
25.4 | 36.2 | 40.2 | 43.2 | 43.8333
25.5 | 36.1 | 40.1 | 43 | 43.6333
25.6 | 35.95 | 39.875 | 42.775 | 43.5
25.7 | 35.825 | 39.7 | 42.7 | 43.35
25.8 | 35.6 | 39.4667 | 42.575 | 43.2
25.9 | 35.475 | 39.3 | 42.325 | 42.96
26 | 35.3333 | 39.2 | 42.2 | 42.825
26.1 | 35.12 | 39.1 | 42.1 | 42.675
26.2 | 35 | 38.9 | 41.95 | 42.525
26.3 | 34.9 | 38.7 | 41.75 | 42.375
26.4 | 34.725 | 38.525 | 41.55 | 42.24
26.5 | 34.5 | 38.375 | 41.3667 | 42.1
26.6 | 34.4 | 38.225 | 41.2125 | 41.9
26.7 | 34.2 | 38.06 | 41.1 | 41.7
26.8 | 34.1 | 37.85 | 40.925 | 41.55
26.9 | 33.925 | 37.75 | 40.7833 | 41.3857
27 | 33.775 | 37.625 | 40.6167 | 41.1714
27.1 | 33.6 | 37.5 | 40.48 | 41.0375
27.2 | 33.4 | 37.275 | 40.3 | 40.9125
27.3 | 33.3 | 37.1333 | 40.1333 | 40.8
27.4 | 33.2 | 36.9667 | 40 | 40.625
27.5 | 33.04 | 36.85 | 39.8 | 40.5
27.6 | 32.8 | 36.6 | 39.625 | 40.3333
27.7 | 32.65 | 36.5 | 39.5 | 40.175
27.8 | 32.5 | 36.4 | 39.325 | 39.9833
27.9 | 32.45 | 36.2 | 39.2 | 39.9
28 | 32.3 | 36 | 39.1 | 39.8
28.1 | 32.15 | 35.85 | 38.85 | 39.7
28.2 | 32 | 35.8 | 38.7 | 39.45
28.3 | 31.9 | 35.65 | 38.55 | 39.3
28.4 | 31.8 | 35.5 | 38.5 | 39.15
28.5 | 31.575 | 35.3667 | 38.3 | 39
28.6 | 31.4143 | 35.12 | 38.125 | 38.8
28.7 | 31.225 | 34.975 | 37.975 | 38.72
28.8 | 31.1 | 34.9 | 37.8 | 38.5
28.9 | 31 | 34.675 | 37.6 | 38.325
29 | 30.9 | 34.525 | 37.5 | 38.18
29.1 | 30.78 | 34.475 | 37.3 | 38.0714
29.2 | 30.65 | 34.3333 | 37.2 | 37.9
29.3 | 30.5 | 34.2 | 37.1 | 37.7667
29.4 | 30.35 | 34.0667 | 36.85 | 37.6
29.5 | 30.2 | 33.9 | 36.7 | 37.4
29.6 | 30.05 | 33.8 | 36.6 | 37.3
29.7 | 30 | 33.64 | 36.5 | 37.2
29.8 | 29.8 | 33.5 | 36.3 | 37.05
29.9 | 29.6 | 33.35 | 36.2 | 37
30 | 29.575 | 33.12 | 36.025 | 36.76

## Wedge 45 Lateral Profiles
Lateral Profiles
Wedge 45
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:2.4 cm FieldSize:4x4 cm2 | Depth:2.4 cm FieldSize:10x10 cm2 | Depth:2.4 cm FieldSize:20x20 cm2 | Depth:2.4 cm FieldSize:40x20 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:40x20 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:40x20 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:40x20 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:40x20 cm2
-17.9 | [18 empty] | 13.36 | 16.2694
-17.8 | [18 empty] | 13.44 | 16.5028
-17.7 | [18 empty] | 13.6667 | 16.6861
-17.6 | [18 empty] | 13.8286 | 16.8778
-17.5 | [18 empty] | 14.02 | 17.1029
-17.4 | [18 empty] | 14.225 | 17.4029
-17.3 | [18 empty] | 14.46 | 17.6029
-17.2 | [18 empty] | 14.6 | 17.8363
-17.1 | [18 empty] | 14.9333 | 18.043
-17 | [18 empty] | 15.2 | 18.243
-16.9 | [14 empty] | 10.68 | 13.62 |  |  | 15.3667 | 18.5364
-16.8 | [14 empty] | 10.9 | 13.8 |  |  | 15.6667 | 18.8698
-16.7 | [14 empty] | 11.1 | 14.0167 |  |  | 15.925 | 19.0632
-16.6 | [14 empty] | 11.3 | 14.2333 |  |  | 16.2 | 19.3782
-16.5 | [14 empty] | 11.5 | 14.42 |  |  | 16.4 | 19.6533
-16.4 | [14 empty] | 11.65 | 14.6 |  |  | 16.66 | 19.8366
-16.3 | [14 empty] | 11.88 | 14.7667 |  |  | 16.95 | 20.17
-16.2 | [14 empty] | 12.0667 | 15 |  |  | 17.28 | 20.4034
-16.1 | [14 empty] | 12.25 | 15.2 |  |  | 17.6143 | 20.7035
-16 | [14 empty] | 12.4333 | 15.45 |  |  | 17.86 | 20.9702
-15.9 | [10 empty] | 8.0667 | 10.5 |  |  | 12.725 | 15.7 |  |  | 18.1333 | 21.3035
-15.8 | [10 empty] | 8.2 | 10.625 |  |  | 12.95 | 15.8833 |  |  | 18.4667 | 21.6036
-15.7 | [10 empty] | 8.3 | 10.74 |  |  | 13.16 | 16.18 |  |  | 18.8 | 21.9037
-15.6 | [10 empty] | 8.5 | 10.96 |  |  | 13.475 | 16.46 |  |  | 19.1333 | 22.1704
-15.5 | [10 empty] | 8.6 | 11.1 |  |  | 13.66 | 16.7333 |  |  | 19.38 | 22.6538
-15.4 | [6 empty] | 6.8 | 9.022 |  |  | 8.75 | 11.24 |  |  | 13.9 | 17 |  |  | 19.7 | 22.9238
-15.3 | [6 empty] | 6.9 | 9.1969 |  |  | 8.88 | 11.52 |  |  | 14.1833 | 17.22 |  |  | 20 | 23.3372
-15.2 |  |  | 6.715 | 9.2 |  |  | 7 | 9.3469 |  |  | 9.025 | 11.675 |  |  | 14.45 | 17.5 |  |  | 20.35 | 23.6707
-15.1 |  |  | 6.7983 | 9.3 |  |  | 7.1 | 9.4768 |  |  | 9.275 | 11.85 |  |  | 14.6857 | 17.725 |  |  | 20.85 | 24.104
-15 |  |  | 6.9583 | 9.4 |  |  | 7.2 | 9.5968 |  |  | 9.3833 | 12 |  |  | 15.025 | 17.94 |  |  | 21.18 | 24.4241
-14.9 |  |  | 7.0982 | 9.6 |  |  | 7.3 | 9.6968 |  |  | 9.625 | 12.125 |  |  | 15.2667 | 18.15 |  |  | 21.64 | 24.7708
-14.8 |  |  | 7.1982 | 9.65 |  |  | 7.5 | 9.8301 |  |  | 9.775 | 12.325 |  |  | 15.5 | 18.5667 |  |  | 22.08 | 25.2042
-14.7 |  |  | 7.2982 | 9.82 |  |  | 7.6 | 9.9967 |  |  | 9.9 | 12.525 |  |  | 15.82 | 18.9 |  |  | 22.5 | 25.6293
-14.6 |  |  | 7.4481 | 9.9 |  |  | 7.7 | 10.1716 |  |  | 10.1 | 12.7667 |  |  | 16.15 | 19.1333 |  |  | 22.94 | 26.1377
-14.5 |  |  | 7.5731 | 10.1 |  |  | 7.9 | 10.2966 |  |  | 10.3667 | 12.98 |  |  | 16.52 | 19.4667 |  |  | 23.45 | 26.6044
-14.4 |  |  | 7.6981 | 10.2 |  |  | 8.05 | 10.4965 |  |  | 10.5667 | 13.2 |  |  | 16.8 | 19.8667 |  |  | 23.98 | 27.2245
-14.3 |  |  | 7.798 | 10.4 |  |  | 8.2 | 10.5965 |  |  | 10.8 | 13.4333 |  |  | 17.075 | 20.1333 |  |  | 24.7 | 27.6846
-14.2 |  |  | 7.938 | 10.5167 |  |  | 8.35 | 10.7964 |  |  | 11 | 13.64 |  |  | 17.5333 | 20.54 |  |  | 25.32 | 28.4847
-14.1 |  |  | 8.073 | 10.6833 |  |  | 8.5 | 10.963 |  |  | 11.2 | 13.84 |  |  | 17.925 | 20.9 |  |  | 26.08 | 29.1799
-14 |  |  | 8.2979 | 10.8167 |  |  | 8.7 | 11.1963 |  |  | 11.4 | 14.14 |  |  | 18.18 | 21.2 |  |  | 27 | 30.105
-13.9 |  |  | 8.3979 | 11 |  |  | 8.8 | 11.3462 |  |  | 11.64 | 14.3167 |  |  | 18.72 | 21.7 |  |  | 28.05 | 31.1802
-13.8 |  |  | 8.5579 | 11.14 |  |  | 9 | 11.4962 |  |  | 11.92 | 14.5667 |  |  | 19.075 | 22.1 |  |  | 29.4667 | 32.4304
-13.7 |  |  | 8.7478 | 11.36 |  |  | 9.2 | 11.6961 |  |  | 12.175 | 14.88 |  |  | 19.56 | 22.4714 |  |  | 31.15 | 33.9856
-13.6 |  |  | 8.9178 | 11.5 |  |  | 9.4 | 11.7961 |  |  | 12.5 | 15.175 |  |  | 20 | 23 |  |  | 33.3667 | 36.2727
-13.5 |  |  | 9.0177 | 11.6833 |  |  | 9.6 | 12.096 |  |  | 12.7 | 15.4 |  |  | 20.64 | 23.45 |  |  | 36.3667 | 38.8865
-13.4 |  |  | 9.1977 | 11.9 |  |  | 9.775 | 12.2959 |  |  | 13.025 | 15.7 |  |  | 21 | 24.05 |  |  | 40.44 | 42.7071
-13.3 |  |  | 9.3727 | 12 |  |  | 10 | 12.4958 |  |  | 13.36 | 16.02 |  |  | 21.775 | 24.62 |  |  | 46.3667 | 48.5081
-13.2 |  |  | 9.5976 | 12.2 |  |  | 10.2 | 12.6958 |  |  | 13.62 | 16.35 |  |  | 22.4 | 25.34 |  |  | 53.8 | 55.5093
-13.1 |  |  | 9.6976 | 12.4 |  |  | 10.4 | 12.8957 |  |  | 14.02 | 16.7333 |  |  | 23.25 | 26.2 |  |  | 63.26 | 64.3107
-13 |  |  | 9.9975 | 12.6 |  |  | 10.6 | 13.1956 |  |  | 14.35 | 17.02 |  |  | 24.14 | 27 |  |  | 74.6 | 74.1874
-12.9 |  |  | 10.1225 | 12.8 |  |  | 10.8667 | 13.3955 |  |  | 14.7 | 17.4333 |  |  | 25.325 | 28.14 |  |  | 84.45 | 84.6141
-12.8 |  |  | 10.2724 | 12.9667 |  |  | 11.1 | 13.6455 |  |  | 15.05 | 17.82 |  |  | 26.75 | 29.5 |  |  | 95.2 | 93.5905
-12.7 |  |  | 10.4974 | 13.2 |  |  | 11.3 | 13.8621 |  |  | 15.55 | 18.28 |  |  | 28.6 | 31.2333 |  |  | 103.0667 | 102.217
-12.6 |  |  | 10.7373 | 13.4 |  |  | 11.625 | 14.1286 |  |  | 16.0333 | 18.65 |  |  | 30.8571 | 33.42 |  |  | 109.55 | 108.238
-12.5 |  |  | 10.9473 | 13.7167 |  |  | 12 | 14.4952 |  |  | 16.45 | 19.15 |  |  | 33.8667 | 36.72 |  |  | 114.6667 | 112.719
-12.4 |  |  | 11.1472 | 13.8833 |  |  | 12.25 | 14.7951 |  |  | 16.96 | 19.7 |  |  | 38.775 | 40.8667 |  |  | 117.9167 | 115.7857
-12.3 |  |  | 11.4471 | 14.2 |  |  | 12.4667 | 15.095 |  |  | 17.65 | 20.3 |  |  | 45.275 | 47.875 |  |  | 120.06 | 117.945
-12.2 |  |  | 11.6804 | 14.4 |  |  | 12.95 | 15.5281 |  |  | 18.2333 | 20.94 |  |  | 54.4333 | 55.8 |  |  | 121.85 | 119.32
-12.1 |  |  | 11.977 | 14.68 |  |  | 13.35 | 15.8947 |  |  | 19.1167 | 21.7 |  |  | 65.4 | 66.3 |  |  | 122.8 | 120.42
-12 |  |  | 12.2969 | 15.05 |  |  | 13.8 | 16.2946 |  |  | 20.1 | 22.7 |  |  | 77.2 | 77.6333 |  |  | 123.7 | 121.2
-11.9 |  |  | 12.5969 | 15.3667 |  |  | 14.2667 | 16.7944 |  |  | 21.22 | 23.8 |  |  | 89.3 | 88.9 |  |  | 124.3 | 121.92
-11.8 |  |  | 12.9968 | 15.7 |  |  | 14.775 | 17.3276 |  |  | 22.6 | 25.2 |  |  | 99.68 | 98.925 |  |  | 124.9333 | 122.4533
-11.7 |  |  | 13.3467 | 16.125 |  |  | 15.4 | 18.044 |  |  | 24.5333 | 27.1667 |  |  | 109.4 | 108.24 |  |  | 125.2 | 122.7867
-11.6 |  |  | 13.7765 | 16.52 |  |  | 16.2 | 18.6438 |  |  | 27.28 | 29.6 |  |  | 116.4333 | 115.1167 |  |  | 125.4667 | 123.121
-11.5 |  |  | 14.2764 | 17.02 |  |  | 17.1 | 19.6268 |  |  | 30.5 | 32.86 |  |  | 121.5 | 119.8667 |  |  | 125.74 | 123.321
-11.4 |  |  | 14.8963 | 17.6 |  |  | 18.2667 | 20.6931 |  |  | 36.1333 | 37.9 |  |  | 124.6 | 122.7429 |  | 7 | 125.875 | 123.4877
-11.3 |  |  | 15.5761 | 18.3 |  |  | 19.7 | 22.126 |  |  | 43.75 | 45.9 |  |  | 126.7667 | 124.68 |  | 7.175 | 125.95 | 123.621
-11.2 |  |  | 16.4209 | 19.125 |  |  | 21.6 | 23.992 |  |  | 54.5333 | 55.8833 |  |  | 128.1833 | 126.15 |  | 7.3 | 126.1 | 123.721
-11.1 |  |  | 17.6456 | 20.3 |  |  | 24.2 | 26.4912 |  |  | 67.65 | 69.25 |  |  | 129.3 | 127.2333 |  | 7.45 | 126.1 | 123.821
-11 |  |  | 19.1952 | 21.725 |  |  | 27.775 | 30.3899 |  |  | 81.1 | 82.65 |  |  | 129.8 | 127.9 |  | 7.6 | 126.1 | 123.821
-10.9 |  |  | 21.1947 | 23.8 |  |  | 33.6 | 35.963 |  |  | 96.025 | 95.95 |  | 5.3653 | 130.3333 | 128.4 |  | 7.7 | 126.175 | 123.821
-10.8 |  |  | 24.194 | 26.5 |  |  | 43.025 | 44.3352 |  |  | 108.8333 | 107.56 |  | 5.4986 | 130.7 | 128.66 |  | 7.8 | 126.1 | 123.821
-10.7 |  |  | 28.8428 | 31.24 |  |  | 54.7 | 56.6061 |  |  | 118.9 | 117.9 |  | 5.5986 | 131.14 | 128.95 |  | 7.92 | 126.1 | 123.821
-10.6 |  |  | 36.091 | 38.2 |  |  | 68.9 | 71.0763 |  |  | 126.58 | 125 |  | 5.6986 | 131.3 | 129.125 |  | 8.075 | 126 | 123.821
-10.5 |  |  | 46.8383 | 48.1 |  |  | 85.15 | 85.3382 |  |  | 131.36 | 129.62 |  | 5.7985 | 131.4 | 129.25 |  | 8.3 | 126 | 123.761
-10.4 |  |  | 61.7346 | 62.05 |  |  | 100.5 | 100.566 |  | 3.72 | 134.36 | 132.425 |  | 5.8985 | 131.5 | 129.3 |  | 8.4 | 125.9 | 123.721
-10.3 |  |  | 76.4809 | 77.3167 |  |  | 113.95 | 113.2625 |  | 3.8 | 136.225 | 134.15 |  | 5.9985 | 131.475 | 129.3167 |  | 8.5 | 125.8 | 123.721
-10.2 |  |  | 93.5766 | 93.15 |  | 3 | 125.6 | 123.959 |  | 3.9 | 137.44 | 135.5 |  | 6.1985 | 131.4 | 129.4 |  | 8.6 | 125.8 | 123.6877
-10.1 |  | 3.025 | 109.5722 | 107.92 |  | 3.1 | 133.3 | 131.556 |  | 3.9 | 138.3 | 136.2 |  | 6.2984 | 131.3 | 129.3 |  | 8.8667 | 125.7 | 123.521
-10 |  | 3.1 | 122.5696 | 120.9 |  | 3.1 | 138 | 135.655 |  | 4 | 138.7333 | 136.7 |  | 6.4651 | 131.2 | 129.2 |  | 9.05 | 125.6 | 123.341
-9.9 |  | 3.1 | 132.7423 | 131.45 |  | 3.2 | 140.8 | 138.404 |  | 4.1 | 139.02 | 137.06 |  | 6.5983 | 131.1 | 129.1 |  | 9.2333 | 125.44 | 123.321
-9.8 |  | 3.2 | 139.69 | 137.3 |  | 3.2 | 142.3333 | 140.1863 |  | 4.2 | 139.2667 | 137.2 |  | 6.7483 | 131 | 128.92 |  | 9.3667 | 125.3 | 123.221
-9.7 |  | 3.2 | 143.014 | 141.02 |  | 3.3 | 143.35 | 141.253 |  | 4.275 | 139.3 | 137.275 |  | 6.8583 | 130.875 | 128.8 |  | 9.6 | 125.125 | 123.121
-9.6 |  | 3.3 | 145.184 | 142.76 |  | 3.3 | 144.15 | 141.903 |  | 4.325 | 139.3 | 137.3 |  | 6.9982 | 130.7333 | 128.7 |  | 9.8 | 125 | 123.021
-9.5 |  | 3.3 | 146.463 | 144.2 |  | 3.4 | 144.5333 | 142.2197 |  | 4.4333 | 139.2 | 137.2 |  | 7.1482 | 130.58 | 128.6 |  | 10 | 124.84 | 122.82
-9.4 |  | 3.4 | 146.963 | 144.75 |  | 3.5 | 144.6 | 142.453 |  | 4.5 | 139.08 | 137.08 |  | 7.3982 | 130.3 | 128.4833 |  | 10.2333 | 124.7 | 122.72
-9.3 |  | 3.4 | 147.2463 | 144.9 |  | 3.5 | 144.6 | 142.453 |  | 4.6 | 138.86 | 136.9 |  | 7.4981 | 130.1 | 128.375 |  | 10.4 | 124.55 | 122.62
-9.2 |  | 3.5 | 147.263 | 144.9 |  | 3.6 | 144.6 | 142.278 |  | 4.725 | 138.58 | 136.68 |  | 7.6731 | 129.85 | 128.08 |  | 10.7333 | 124.4667 | 122.52
-9.1 |  | 3.6 | 147.263 | 144.9 |  | 3.7 | 144.4333 | 142.128 |  | 4.8333 | 138.38 | 136.42 |  | 7.798 | 129.7833 | 127.875 |  | 10.95 | 124.3 | 122.42
-9 |  | 3.6 | 147.063 | 144.68 |  | 3.7 | 144.0333 | 141.878 |  | 5 | 138.16 | 136.125 |  | 7.998 | 129.6167 | 127.66 |  | 11.1 | 124.0667 | 122.22
-8.9 |  | 3.7 | 146.813 | 144.42 |  | 3.8 | 143.8 | 141.4863 |  | 5.05 | 137.8 | 135.9 |  | 8.098 | 129.34 | 127.46 |  | 11.4 | 123.86 | 122.06
-8.8 |  | 3.8 | 146.363 | 144.05 |  | 3.9 | 143.4333 | 141.153 |  | 5.15 | 137.5 | 135.55 |  | 8.3313 | 129.14 | 127.225 |  | 11.65 | 123.75 | 121.86
-8.7 |  | 3.875 | 145.814 | 143.5 |  | 4 | 142.95 | 140.903 |  | 5.3 | 137.16 | 135.25 |  | 8.5479 | 128.98 | 127.05 |  | 11.9667 | 123.4667 | 121.6771
-8.6 |  | 4 | 145.339 | 143.1 |  | 4.1 | 142.5 | 140.453 |  | 5.5 | 136.75 | 135 |  | 8.8228 | 128.7 | 126.8 |  | 12.3 | 123.3 | 121.52
-8.5 |  | 4 | 145.014 | 142.6333 |  | 4.2 | 142.2 | 139.953 |  | 5.6 | 136.325 | 134.7 |  | 9.0311 | 128.3667 | 126.52 |  | 12.65 | 123 | 121.42
-8.4 |  | 4.1 | 144.164 | 142.1 |  | 4.275 | 141.7 | 139.553 |  | 5.7333 | 135.9333 | 134.275 |  | 9.2977 | 128.1667 | 126.2333 |  | 12.825 | 122.775 | 121.14
-8.3 |  | 4.2 | 143.864 | 141.62 |  | 4.4 | 141.1667 | 139.054 |  | 5.875 | 135.72 | 133.8667 |  | 9.4976 | 127.8 | 125.98 |  | 13.1333 | 122.54 | 120.895
-8.2 |  | 4.3 | 143.189 | 141.025 |  | 4.5 | 140.8 | 138.674 |  | 6.025 | 135.16 | 133.65 |  | 9.7976 | 127.6 | 125.725 |  | 13.5 | 122.42 | 120.6867
-8.1 |  | 4.4 | 142.724 | 140.45 |  | 4.6 | 140.1 | 138.279 |  | 6.2 | 134.75 | 133.18 |  | 10.0975 | 127.2 | 125.52 |  | 13.8667 | 122.2 | 120.4867
-8 |  | 4.5 | 142.0248 | 139.95 |  | 4.8 | 139.6 | 137.604 |  | 6.475 | 134.28 | 132.7 |  | 10.2974 | 126.86 | 125.3 |  | 14.2 | 121.95 | 120.1533
-7.9 |  | 4.575 | 141.415 | 139.34 |  | 4.9 | 139.05 | 137.054 |  | 6.625 | 133.9 | 132.2 |  | 10.5724 | 126.525 | 124.92 |  | 14.725 | 121.8 | 119.97
-7.8 |  | 4.7 | 140.765 | 138.82 |  | 5 | 138.65 | 136.5043 |  | 6.9 | 133.5 | 131.96 |  | 10.9473 | 126.275 | 124.66 |  | 15.1667 | 121.55 | 119.8033
-7.7 |  | 4.8 | 140.325 | 138.24 |  | 5.2 | 138 | 136.0217 |  | 7.025 | 133.0333 | 131.475 |  | 11.3472 | 125.9 | 124.38 |  | 15.7 | 121.2667 | 119.6367
-7.6 |  | 4.975 | 139.585 | 137.68 |  | 5.3 | 137.45 | 135.5883 |  | 7.2333 | 132.6 | 131.15 |  | 11.6638 | 125.56 | 124.0833 | 2.5013 | 16.4333 | 120.9333 | 119.4533
-7.5 |  | 5.1 | 139.145 | 137.12 |  | 5.5 | 136.9 | 135.005 |  | 7.4667 | 132.175 | 130.62 |  | 11.997 | 125.3714 | 123.8 | 2.6013 | 17.2333 | 120.66 | 119.2
-7.4 |  | 5.2 | 138.365 | 136.6 |  | 5.7 | 136.25 | 134.455 |  | 7.7 | 131.7 | 130.1 |  | 12.4969 | 125 | 123.4667 | 2.7014 | 18.15 | 120.4 | 118.9867
-7.3 |  | 5.3 | 137.626 | 135.85 |  | 5.8667 | 135.7 | 134.0217 |  | 8.0333 | 131.2 | 129.72 | 2 | 12.9568 | 124.55 | 123.2 | 2.7014 | 19.375 | 120.2 | 118.72
-7.2 |  | 5.5 | 137.066 | 135.25 |  | 6.1 | 135.2 | 133.406 |  | 8.2667 | 130.74 | 129.44 | 2.1 | 13.4966 | 124.2333 | 122.875 | 2.8014 | 20.8 | 120.0333 | 118.52
-7.1 |  | 5.6667 | 136.516 | 134.75 |  | 6.3 | 134.5 | 132.906 | 1.4 | 8.6667 | 130.22 | 128.88 | 2.1 | 14.2964 | 123.9 | 122.55 | 2.8014 | 22.5 | 119.7 | 118.42
-7 | 1.1 | 5.86 | 135.706 | 134.12 | 1.0995 | 6.5 | 133.9 | 132.3227 | 1.4 | 9.05 | 129.8 | 128.4667 | 2.1 | 15.0962 | 123.65 | 122.38 | 2.9015 | 25.15 | 119.44 | 118.27
-6.9 | 1.1 | 6.0333 | 135.166 | 133.54 | 1.0995 | 6.8 | 133.35 | 131.856 | 1.4 | 9.5 | 129.4 | 128 | 2.2 | 15.946 | 123.3333 | 121.975 | 3.0015 | 28.8 | 119.225 | 117.97
-6.8 | 1.1 | 6.2667 | 134.666 | 132.875 | 1.0995 | 7.1 | 132.9667 | 131.306 | 1.4 | 9.9333 | 128.92 | 127.56 | 2.3 | 17.2457 | 123.05 | 121.6 | 3.0015 | 33.8 | 119.1 | 117.62
-6.7 | 1.1 | 6.5 | 134.0165 | 132.375 | 1.0995 | 7.4 | 132.2333 | 130.656 | 1.45 | 10.4 | 128.46 | 127.1333 | 2.3 | 18.7953 | 122.6 | 121.3 | 3.1016 | 40.8 | 118.875 | 117.32
-6.6 | 1.1 | 6.8 | 133.367 | 131.7 | 1.1994 | 7.8 | 131.725 | 130.157 | 1.5 | 11.05 | 127.92 | 126.7833 | 2.3 | 20.6948 | 122.2 | 120.975 | 3.1016 | 48.35 | 118.58 | 117.18
-6.5 | 1.2 | 7.05 | 132.667 | 131.1 | 1.1994 | 8.2667 | 131.1 | 129.657 | 1.5 | 11.8 | 127.48 | 126.34 | 2.4 | 23.3692 | 121.8 | 120.75 | 3.2016 | 58 | 118.275 | 116.9695
-6.4 | 1.2 | 7.4667 | 132.027 | 130.5667 | 1.1994 | 8.7333 | 130.4333 | 128.9903 | 1.6 | 12.8 | 127 | 125.85 | 2.4 | 27.3932 | 121.46 | 120.425 | 3.3016 | 68.6 | 117.9333 | 116.6857
-6.3 | 1.2 | 7.875 | 131.6003 | 130.02 | 1.1994 | 9.4333 | 130.025 | 128.5903 | 1.6 | 14 | 126.6 | 125.175 | 2.5 | 32.9918 | 121.1333 | 120.15 | 3.4017 | 78.0333 | 117.7333 | 116.419
-6.2 | 1.2 | 8.3667 | 130.9837 | 129.375 | 1.1994 | 10.3333 | 129.6333 | 128.082 | 1.6 | 15.5667 | 126 | 124.9167 | 2.5 | 41.0897 | 120.9 | 119.74 | 3.5017 | 86.225 | 117.4 | 116.169
-6.1 | 1.2 | 9.05 | 130.2337 | 128.85 | 1.2494 | 11.3 | 129 | 127.4243 | 1.7 | 17.45 | 125.45 | 124.44 | 2.6 | 50.5874 | 120.475 | 119.5 | 3.6018 | 92.4 | 117.1 | 115.859
-6 | 1.2 | 9.9 | 129.768 | 128.22 | 1.2994 | 12.5 | 128.2667 | 126.9247 | 1.7 | 20.25 | 125.05 | 124 | 2.7 | 61.1847 | 120.225 | 118.98 | 3.7018 | 97.625 | 116.7667 | 115.599
-5.9 | 1.2 | 10.9 | 129.168 | 127.75 | 1.2994 | 14.4 | 127.7 | 126.458 | 1.7 | 24.475 | 124.6 | 123.44 | 2.7 | 71.0489 | 119.8 | 118.7 | 3.8019 | 101.1667 | 116.5333 | 115.379
-5.8 | 1.3 | 12.4333 | 128.468 | 127.12 | 1.2994 | 16.9667 | 127.2 | 125.958 | 1.8 | 30.5333 | 123.9667 | 123 | 2.8 | 81.2797 | 119.4667 | 118.34 | 3.8019 | 103.3 | 116.28 | 115.219
-5.7 | 1.3 | 14.6333 | 127.818 | 126.575 | 1.3993 | 20.6 | 126.6 | 125.258 | 1.8 | 39.15 | 123.64 | 122.5 | 2.8 | 90.2774 | 119.1 | 118.12 | 3.8019 | 105 | 116.05 | 114.979
-5.6 | 1.3 | 17.7 | 127.248 | 125.96 | 1.3993 | 26.4333 | 125.9 | 124.6913 | 1.9 | 49.2333 | 123.2 | 122.04 | 2.9 | 96.3759 | 118.7667 | 117.7 | 3.9019 | 106.05 | 115.7 | 114.679
-5.5 | 1.3 | 22.2 | 126.568 | 125.45 | 1.3993 | 34.35 | 125.5 | 124.159 | 1.9 | 60.9333 | 122.65 | 121.6333 | 3.1 | 101.075 | 118.325 | 117.36 | 4.002 | 107 | 115.55 | 114.439
-5.4 | 1.4 | 30.1 | 126.2015 | 124.75 | 1.4992 | 45.45 | 124.8667 | 123.584 | 2 | 73.3 | 122.14 | 121.16 | 3.2 | 103.8407 | 118.02 | 117.1 | 4.1021 | 107.4667 | 115.2 | 114.219
-5.3 | 1.4 | 39.9 | 125.429 | 124.26 | 1.4992 | 57 | 124.3 | 123.034 | 2 | 83.7 | 121.8 | 120.8 | 3.3 | 105.9735 | 117.8 | 116.7333 | 4.2021 | 108 | 114.925 | 113.919
-5.2 | 1.4 | 53.3 | 124.969 | 123.725 | 1.4992 | 69.2667 | 123.85 | 122.684 | 2.1 | 92.9 | 121.3333 | 120.4 | 3.35 | 107.548 | 117.36 | 116.5 | 4.3022 | 108.38 | 114.7333 | 113.699
-5.1 | 1.5 | 66.3 | 124.269 | 123.1 | 1.5992 | 81.6333 | 123.1 | 122.159 | 2.1 | 99.9 | 120.72 | 119.82 | 3.4 | 108.373 | 116.9833 | 116.32 | 4.5023 | 108.6 | 114.5 | 113.519
-5 | 1.5 | 79.1 | 123.669 | 122.5 | 1.5992 | 92.8 | 122.55 | 121.559 | 2.2 | 104.6333 | 120.3 | 119.35 | 3.5 | 109.1397 | 116.6667 | 115.8667 | 4.7024 | 108.78 | 114.25 | 113.2857
-4.9 | 1.5 | 90.475 | 123.219 | 121.9 | 1.6991 | 100.5 | 122.0667 | 120.96 | 2.3 | 107.55 | 119.8333 | 118.95 | 3.6 | 109.573 | 116.275 | 115.55 | 4.8024 | 108.8 | 113.94 | 112.9857
-4.8 | 1.6 | 100.1 | 122.669 | 121.5 | 1.6991 | 106.275 | 121.6333 | 120.46 | 2.4 | 109.5 | 119.4 | 118.5 | 3.8 | 109.873 | 116.0333 | 115.16 | 4.9024 | 108.8667 | 113.55 | 112.719
-4.7 | 1.6 | 106.7 | 122.1025 | 120.9 | 1.6991 | 109.6 | 121.0667 | 120.16 | 2.4333 | 110.55 | 119.02 | 118.08 | 3.95 | 110.122 | 115.7 | 114.9 | 5.1025 | 109.0667 | 113.275 | 112.4523
-4.6 | 1.7 | 111 | 121.47 | 120.45 | 1.8491 | 111.2 | 120.6 | 119.5933 | 2.6 | 111.3333 | 118.45 | 117.76 | 4.1 | 110.4053 | 115.3 | 114.55 | 5.3026 | 109.1 | 112.98 | 112.1857
-4.5 | 1.7 | 113 | 121.03 | 120.05 | 1.899 | 112.8333 | 120 | 118.9933 | 2.7 | 111.9667 | 118.1 | 117.4 | 4.2 | 110.497 | 115 | 114.3 | 5.5027 | 109.1 | 112.82 | 111.9857
-4.4 | 1.8 | 114.3 | 120.4033 | 119.45 | 1.999 | 113.6 | 119.7 | 118.56 | 2.75 | 112.4333 | 117.65 | 116.95 | 4.4 | 110.572 | 114.76 | 114.02 | 5.7028 | 109.1 | 112.6 | 111.719
-4.3 | 1.9 | 115.4 | 119.87 | 118.9 | 1.999 | 114.25 | 119.2 | 118.161 | 2.9 | 112.6 | 117.24 | 116.5 | 4.6 | 110.497 | 114.35 | 113.74 | 5.9029 | 109.1 | 112.3667 | 111.519
-4.2 | 1.9 | 115.75 | 119.395 | 118.5333 | 2.1989 | 114.6 | 118.7 | 117.7277 | 3 | 112.7 | 116.8 | 116.06 | 4.7 | 110.472 | 114 | 113.5 | 6.103 | 108.9667 | 112.15 | 111.1857
-4.1 | 2 | 115.9 | 118.85 | 117.9 | 2.2988 | 114.7 | 118.15 | 117.286 | 3.2 | 112.7 | 116.34 | 115.65 | 4.9667 | 110.372 | 113.6333 | 113.05 | 6.3031 | 108.8 | 111.825 | 110.893
-4 | 2.1 | 115.9 | 118.37 | 117.4 | 2.3988 | 114.75 | 117.7 | 116.761 | 3.3 | 112.7 | 116 | 115.3 | 5.2 | 110.2387 | 113.2 | 112.7667 | 6.8034 | 108.7667 | 111.44 | 110.6751
-3.9 | 2.1 | 115.8333 | 117.671 | 117.1 | 2.3988 | 114.6333 | 117.2 | 116.211 | 3.5 | 112.6 | 115.46 | 114.82 | 5.45 | 110.072 | 112.9333 | 112.34 | 7.1036 | 108.7 | 111.05 | 110.498
-3.8 | 2.3 | 115.725 | 117.231 | 116.58 | 2.6653 | 114.475 | 116.65 | 115.811 | 3.7 | 112.4667 | 115 | 114.5333 | 5.75 | 109.923 | 112.7 | 112 | 7.4037 | 108.6 | 110.72 | 110.118
-3.7 | 2.4 | 115.4333 | 116.751 | 115.95 | 2.7986 | 114.3 | 116.2 | 115.337 | 4 | 112.2667 | 114.6333 | 114.18 | 6.05 | 109.873 | 112.35 | 111.78 | 8.004 | 108.4 | 110.4667 | 109.8847
-3.6 | 2.5 | 115.125 | 116.311 | 115.35 | 2.9985 | 114.0333 | 115.6 | 115.062 | 4.2 | 112.1 | 114.08 | 113.625 | 6.5333 | 109.673 | 111.9333 | 111.56 | 8.7043 | 108.3 | 110.22 | 109.6847
-3.5 | 2.7 | 114.9 | 115.746 | 114.9 | 3.2984 | 113.75 | 115.1667 | 114.4953 | 4.55 | 111.9 | 113.65 | 113.1833 | 7.05 | 109.5063 | 111.75 | 111.1 | 9.4047 | 108.0667 | 109.94 | 109.438
-3.4 | 2.9 | 114.6667 | 115.311 | 114.54 | 3.5982 | 113.3667 | 114.65 | 113.962 | 4.9 | 111.7333 | 113.25 | 112.9 | 7.65 | 109.373 | 111.35 | 110.75 | 10.4052 | 107.975 | 109.78 | 109.158
-3.3 | 3.1 | 114.1667 | 114.831 | 113.94 | 3.8981 | 113.225 | 114.2333 | 113.562 | 5.4 | 111.4 | 112.9 | 112.45 | 8.4 | 109.1397 | 110.9571 | 110.46 | 11.7059 | 107.725 | 109.5333 | 108.8513
-3.2 | 3.4333 | 113.7333 | 114.231 | 113.475 | 4.3478 | 112.85 | 113.8 | 113.062 | 5.95 | 111.1 | 112.575 | 112.1 | 9.5 | 108.923 | 110.48 | 110.08 | 13.4067 | 107.5333 | 109.275 | 108.568
-3.1 | 3.9 | 113.3 | 113.822 | 113.1 | 4.8975 | 112.5333 | 113.25 | 112.562 | 6.7 | 110.8 | 112.1 | 111.65 | 10.7 | 108.623 | 110.16 | 109.8 | 15.8079 | 107.4 | 108.925 | 108.238
-3 | 4.3 | 113.0667 | 113.222 | 112.68 | 5.5972 | 112.0333 | 112.7667 | 112.263 | 7.6 | 110.6 | 111.65 | 111.1667 | 12.5 | 108.5063 | 109.82 | 109.4 | 19.1596 | 107.14 | 108.55 | 107.8513
-2.9 | 5.1 | 112.65 | 112.672 | 112.08 | 6.5967 | 111.6 | 112.2333 | 111.6963 | 9.1 | 110.2333 | 111.2333 | 110.82 | 14.7 | 108.173 | 109.5 | 109 | 23.6118 | 106.9 | 108.24 | 107.618
-2.8 | 6 | 112.1333 | 112.192 | 111.74 | 7.7961 | 111.25 | 111.85 | 111.2297 | 10.8 | 110 | 110.825 | 110.48 | 18.15 | 107.823 | 109.2 | 108.7333 | 30.065 | 106.8 | 107.88 | 107.358
-2.7 | 7.45 | 111.7333 | 111.872 | 111.34 | 9.5952 | 110.8667 | 111.45 | 110.763 | 13 | 109.6333 | 110.3333 | 110.05 | 23.15 | 107.673 | 108.74 | 108.45 | 38.5193 | 106.6333 | 107.625 | 107.168
-2.6 | 9.2 | 111.2333 | 111.247 | 110.75 | 11.994 | 110.45 | 110.8 | 110.363 | 16.7 | 109.375 | 109.925 | 109.6 | 30.05 | 107.548 | 108.4 | 108.0833 | 47.924 | 106.45 | 107.42 | 106.9013
-2.5 | 12.1 | 110.575 | 110.812 | 110.24 | 15.1924 | 110.15 | 110.325 | 109.988 | 21.8 | 109.1333 | 109.55 | 109.16 | 39.2 | 107.298 | 108.06 | 107.8 | 58.1291 | 106.3 | 107.2167 | 106.6513
-2.4 | 16.8 | 110.1333 | 110.347 | 109.86 | 20.8396 | 109.7 | 110.0333 | 109.4307 | 29.1 | 108.6667 | 109.1 | 108.7 | 48.9 | 107.073 | 107.7 | 107.475 | 68.4342 | 106.1 | 107 | 106.468
-2.3 | 23 | 109.75 | 109.823 | 109.44 | 29.2854 | 109.3333 | 109.4667 | 108.9973 | 39.3 | 108.35 | 108.72 | 108.3333 | 59.7 | 106.773 | 107.36 | 107.125 | 77.1386 | 106 | 106.7 | 106.138
-2.2 | 32.8 | 109.4 | 109.273 | 108.88 | 38.5307 | 108.9333 | 109.0333 | 108.5973 | 50.2 | 108 | 108.35 | 108.0333 | 70.1 | 106.5063 | 107.06 | 106.8667 | 84.8424 | 105.625 | 106.4 | 105.858
-2.1 | 45.3 | 108.95 | 108.873 | 108.4 | 50.3415 | 108.45 | 108.65 | 108.264 | 61.8 | 107.6 | 107.9 | 107.56 | 79.5 | 106.273 | 106.6667 | 106.54 | 90.6453 | 105.3667 | 106.1 | 105.568
-2 | 57 | 108.4 | 108.413 | 108.12 | 62.3688 | 108.1 | 108.1667 | 107.764 | 72.5 | 107.1 | 107.48 | 107.2 | 87.1 | 106.093 | 106.3333 | 106.2 | 94.8474 | 105.1 | 105.7 | 105.358
-1.9 | 68.9 | 108.05 | 107.873 | 107.62 | 73.4633 | 107.7333 | 107.75 | 107.364 | 82.1 | 106.9 | 107.1 | 106.86 | 92.4 | 105.874 | 106.1167 | 105.9286 | 97.1486 | 104.9 | 105.6 | 105.118
-1.8 | 79.5 | 107.65 | 107.448 | 107.1 | 83.0585 | 107.2333 | 107.5 | 106.914 | 89.05 | 106.5667 | 106.58 | 106.45 | 95.8 | 105.374 | 105.8 | 105.6 | 98.8494 | 104.8 | 105.3333 | 104.7503
-1.7 | 88.2 | 107.1 | 107.073 | 106.8 | 90.4548 | 106.8 | 106.9333 | 106.5142 | 94.1 | 106.1 | 106.3 | 106.1667 | 98 | 105.074 | 105.44 | 105.3 | 100 | 104.6 | 105.02 | 104.417
-1.6 | 94.8 | 106.8 | 106.673 | 106.42 | 94.8526 | 106.525 | 106.3667 | 106.0983 | 96.7 | 105.75 | 105.8667 | 105.8 | 99.3 | 104.849 | 105.1333 | 105.06 | 100.8 | 104.3333 | 104.6333 | 104.267
-1.5 | 98.3 | 106.3333 | 106.1332 | 105.9 | 97.951 | 106 | 105.9667 | 105.765 | 99.2 | 105.38 | 105.45 | 105.375 | 100.65 | 104.474 | 104.7 | 104.74 | 101.301 | 104.05 | 104.3667 | 104.077
-1.4 | 100 | 105.9667 | 105.694 | 105.4 | 99.5502 | 105.6667 | 105.5667 | 105.2983 | 100.25 | 105.1 | 105.1333 | 105.0333 | 101.25 | 104.224 | 104.38 | 104.275 | 101.551 | 103.7 | 104.05 | 103.717
-1.3 | 101.4 | 105.7 | 105.314 | 105.02 | 100.75 | 105.15 | 105.1 | 105.025 | 101.1 | 104.8333 | 104.74 | 104.6 | 101.6 | 103.974 | 104.15 | 104.02 | 101.851 | 103.35 | 103.82 | 103.4503
-1.2 | 102.2 | 105.1 | 104.924 | 104.68 | 101.399 | 104.8 | 104.825 | 104.6317 | 101.6 | 104.4 | 104.3 | 104.25 | 101.9 | 103.674 | 103.85 | 103.6333 | 101.951 | 103.175 | 103.4 | 103.217
-1.1 | 102.6 | 104.7333 | 104.594 | 104.25 | 101.8157 | 104.3 | 104.425 | 104.14 | 101.9 | 104.1 | 103.875 | 103.85 | 102 | 103.399 | 103.54 | 103.3 | 102.051 | 102.9 | 103.15 | 103.0337
-1 | 102.7 | 104.2667 | 103.954 | 103.9333 | 101.999 | 103.875 | 103.85 | 103.7317 | 102.1 | 103.65 | 103.5333 | 103.5667 | 102 | 103.049 | 103.08 | 103.125 | 102.051 | 102.7 | 102.8333 | 102.7741
-0.9 | 102.7 | 103.8333 | 103.554 | 103.52 | 102.049 | 103.45 | 103.45 | 103.3993 | 102.1 | 103.2 | 103.2 | 103.25 | 102 | 102.8073 | 102.76 | 102.8 | 101.951 | 102.5 | 102.5 | 102.317
-0.8 | 102.7 | 103.4333 | 103.124 | 103.02 | 101.949 | 103.1333 | 103.08 | 103.0327 | 102.1 | 102.8667 | 102.7 | 102.85 | 101.9 | 102.5073 | 102.4 | 102.4667 | 101.851 | 102.225 | 102.28 | 101.9837
-0.7 | 102.5667 | 102.875 | 102.754 | 102.66 | 101.849 | 102.925 | 102.8 | 102.666 | 102.1 | 102.4 | 102.4 | 102.5333 | 101.7667 | 102.1073 | 102.1 | 102.08 | 101.651 | 101.96 | 102.0667 | 101.777
-0.6 | 102.3 | 102.55 | 102.399 | 102.3 | 101.749 | 102.35 | 102.4 | 102.266 | 101.9 | 102.1 | 102 | 102.08 | 101.5 | 101.825 | 101.85 | 101.775 | 101.451 | 101.7667 | 101.8 | 101.517
-0.5 | 102 | 102.2333 | 102.0944 | 101.95 | 101.549 | 101.85 | 101.9 | 101.7993 | 101.6 | 101.7333 | 101.65 | 101.7 | 101.3 | 101.475 | 101.5 | 101.52 | 101.251 | 101.5667 | 101.5 | 101.337
-0.4 | 101.5 | 101.7 | 101.535 | 101.55 | 101.299 | 101.55 | 101.6 | 101.366 | 101.4 | 101.4 | 101.24 | 101.425 | 101.1 | 101.3 | 101.15 | 101.26 | 100.85 | 101.2 | 101.18 | 101.0837
-0.3 | 101.2 | 101.1 | 101.175 | 101.2 | 101.049 | 101.1333 | 101.2333 | 101.0327 | 101.15 | 100.9667 | 100.9 | 101.05 | 100.9 | 101.05 | 100.8667 | 100.96 | 100.75 | 100.8 | 100.8333 | 100.7503
-0.2 | 100.8 | 100.8333 | 100.8 | 100.72 | 100.65 | 100.7 | 100.6667 | 100.6327 | 100.8 | 100.6333 | 100.575 | 100.6 | 100.75 | 100.7 | 100.6 | 100.675 | 100.55 | 100.55 | 100.5 | 100.417
-0.1 | 100.5 | 100.5 | 100.395 | 100.4 | 100.25 | 100.4 | 100.4 | 100.3337 | 100.55 | 100.2667 | 100.28 | 100.275 | 100.4 | 100.35 | 100.3 | 100.32 | 100.35 | 100.25 | 100.3 | 100.217
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0001 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0003
0.1 | 99.6 | 99.6 | 99.6351 | 99.62 | 99.5502 | 99.675 | 99.725 | 99.6668 | 99.7 | 99.55 | 99.58 | 99.66 | 99.8 | 99.6751 | 99.5857 | 99.6833 | 99.7499 | 99.7667 | 99.72 | 99.6916
0.2 | 99.3 | 99.25 | 99.2502 | 99.25 | 99.0505 | 99.15 | 99.4333 | 99.3419 | 99.4 | 99.3 | 99.22 | 99.25 | 99.25 | 99.4751 | 99.34 | 99.4 | 99.5498 | 99.52 | 99.44 | 99.4416
0.3 | 98.95 | 98.9 | 98.9003 | 99 | 98.7506 | 98.8 | 99 | 99.067 | 99.1 | 99.1 | 98.88 | 98.9 | 98.95 | 99.0752 | 99.02 | 99.16 | 99.2496 | 99.15 | 99.2667 | 99.2165
0.4 | 98.65 | 98.5 | 98.5954 | 98.7 | 98.5007 | 98.5 | 98.8 | 98.6671 | 98.6 | 98.575 | 98.5 | 98.65 | 98.7 | 98.7087 | 98.7 | 98.88 | 98.8494 | 98.85 | 99.0167 | 99.0165
0.5 | 98.1 | 98.0667 | 98.3754 | 98.3 | 97.8511 | 98.1333 | 98.25 | 98.2173 | 98.0333 | 98.2667 | 98.1 | 98.4 | 98.5 | 98.4087 | 98.35 | 98.56 | 98.4492 | 98.6 | 98.7 | 98.6831
0.6 | 97.6 | 97.7667 | 97.9088 | 97.98 | 97.4513 | 97.6667 | 98 | 97.8674 | 97.8 | 97.9667 | 97.82 | 98.16 | 97.9 | 98.1755 | 98.2 | 98.3333 | 97.949 | 98.35 | 98.46 | 98.4664
0.7 | 97.15 | 97.4333 | 97.6156 | 97.6 | 96.9515 | 97.4 | 97.65 | 97.5342 | 97.25 | 97.6 | 97.6 | 97.8 | 97.5 | 97.8255 | 97.8833 | 97.95 | 97.6488 | 98.0333 | 98.2 | 98.2164
0.8 | 96.55 | 97.1 | 97.3757 | 97.2 | 96.3518 | 97.3 | 97.3 | 97.2676 | 96.6 | 97.2 | 97.25 | 97.55 | 97.05 | 97.5756 | 97.6 | 97.6714 | 97.3487 | 97.775 | 97.9333 | 97.9163
0.9 | 95.9 | 96.75 | 96.8958 | 96.9833 | 95.5023 | 96.775 | 97.025 | 97.0926 | 96.05 | 96.86 | 96.975 | 97.18 | 96.6 | 97.2157 | 97.26 | 97.4333 | 97.0986 | 97.4333 | 97.58 | 97.6763
1 | 95.3 | 96.4 | 96.5959 | 96.75 | 94.9025 | 96.35 | 96.8 | 96.8344 | 95.65 | 96.6 | 96.68 | 96.88 | 96.1 | 96.7425 | 97.06 | 97.14 | 96.7484 | 97.05 | 97.18 | 97.4412
1.1 | 94.4 | 96.025 | 96.2759 | 96.3833 | 94.053 | 96 | 96.4 | 96.5678 | 94.7 | 96.3667 | 96.3 | 96.6 | 95.5 | 96.5009 | 96.82 | 96.825 | 95.948 | 96.8 | 96.9 | 97.1162
1.2 | 93.5 | 95.6 | 95.936 | 96.1333 | 93.2534 | 95.725 | 96 | 96.118 | 93.8 | 95.9667 | 96.175 | 96.34 | 94.9 | 96.3009 | 96.54 | 96.62 | 95.6478 | 96.575 | 96.8167 | 96.9412
1.3 | 92.1 | 95.3 | 95.5561 | 95.775 | 92.054 | 95.45 | 95.6667 | 95.7931 | 92.9333 | 95.55 | 95.84 | 96.0333 | 94.2 | 95.9427 | 96.22 | 96.2 | 95.0475 | 96.3 | 96.62 | 96.5561
1.4 | 90.4 | 95 | 95.2762 | 95.38 | 90.2549 | 95.1 | 95.3667 | 95.4682 | 91.3 | 95.3 | 95.5 | 95.6667 | 93.15 | 95.6761 | 95.9143 | 96.05 | 94.3972 | 96 | 96.38 | 96.3161
1.5 | 88.2 | 94.62 | 95.0762 | 95.16 | 88.1559 | 94.7667 | 95 | 95.2182 | 90 | 95.1 | 94.9833 | 95.46 | 92.2 | 95.3762 | 95.6 | 95.8 | 93.6468 | 95.75 | 96.04 | 96.056
1.6 | 84.6 | 94.3 | 94.6513 | 94.8 | 85.0575 | 94.4333 | 94.6667 | 94.8684 | 88 | 94.5 | 94.8167 | 95.22 | 90.9 | 95.0762 | 95.3 | 95.6 | 92.3462 | 95.5 | 95.7333 | 95.816
1.7 | 79.2 | 94 | 94.4164 | 94.55 | 81.1594 | 94.075 | 94.475 | 94.5018 | 84.9 | 94.1667 | 94.5333 | 94.82 | 89.4 | 94.7763 | 95.1 | 95.3 | 91.6959 | 95.16 | 95.475 | 95.6159
1.8 | 71.5 | 93.6 | 94.0965 | 94.225 | 74.8626 | 93.6667 | 94.2 | 94.2686 | 81 | 93.9 | 94.2143 | 94.55 | 87.1 | 94.3764 | 94.775 | 94.9 | 90.5453 | 94.825 | 95.28 | 95.4159
1.9 | 61.9333 | 93.325 | 93.7433 | 93.88 | 66.117 | 93.3333 | 93.7667 | 93.9687 | 74.3 | 93.5667 | 94.02 | 94.26 | 84.2 | 94.0265 | 94.4833 | 94.6 | 88.7944 | 94.5 | 95.02 | 94.9158
2 | 51.4 | 92.9667 | 93.4567 | 93.65 | 56.4218 | 93 | 93.5333 | 93.6188 | 66.9 | 93.225 | 93.68 | 93.96 | 80.3 | 93.7766 | 94.2333 | 94.3714 | 86.8434 | 94.25 | 94.72 | 94.7158
2.1 | 40.4 | 92.5333 | 93.1367 | 93.3333 | 46.3768 | 92.7333 | 93.2667 | 93.3689 | 57.1 | 92.9667 | 93.3 | 93.7 | 74.5 | 93.3767 | 93.9 | 94.1 | 83.6418 | 93.9 | 94.375 | 94.4824
2.2 | 30.6 | 92.2667 | 92.7268 | 93.025 | 36.082 | 92.25 | 92.9 | 93.1356 | 47.25 | 92.5 | 93.025 | 93.34 | 66.7 | 93.0267 | 93.6 | 93.74 | 79.7399 | 93.5667 | 94.22 | 94.3157
2.3 | 21.3 | 91.9 | 92.5435 | 92.8 | 26.3868 | 92 | 92.55 | 92.7691 | 37.25 | 92.2 | 92.7 | 93.04 | 58 | 92.6768 | 93.24 | 93.5429 | 74.5373 | 93.3 | 93.94 | 94.149
2.4 | 15.5 | 91.6333 | 92.2436 | 92.425 | 19.6902 | 91.65 | 92.3 | 92.4359 | 28.65 | 91.8333 | 92.5 | 92.62 | 48.3 | 92.3436 | 93 | 93.25 | 66.8334 | 93.05 | 93.6333 | 93.7156
2.5 | 11.3 | 91.1667 | 91.7271 | 92.0833 | 14.3928 | 91.3 | 92.05 | 92.0693 | 21.6 | 91.55 | 92.0333 | 92.4333 | 38.9 | 91.877 | 92.75 | 92.9167 | 58.5293 | 92.6 | 93.375 | 93.4956
2.6 | 8.7 | 90.8 | 91.5104 | 91.82 | 10.9945 | 90.8667 | 91.6333 | 91.8194 | 15.8 | 91.075 | 91.775 | 92.14 | 30.6 | 91.6771 | 92.5 | 92.65 | 49.3247 | 92.2667 | 93.1 | 93.3156
2.7 | 6.8 | 90.4 | 91.1772 | 91.4714 | 8.6957 | 90.5 | 91.3 | 91.4945 | 12.7 | 90.65 | 91.525 | 91.8667 | 23.7 | 91.2522 | 92.175 | 92.3667 | 40.7204 | 91.9333 | 92.85 | 92.9955
2.8 | 5.5 | 90 | 90.744 | 91.14 | 7.0964 | 90.05 | 91 | 91.1946 | 10.35 | 90.3 | 91.2 | 91.6 | 18.5 | 90.8273 | 91.94 | 92.1167 | 32.016 | 91.5333 | 92.58 | 92.7905
2.9 | 4.7 | 89.6 | 90.5274 | 90.84 | 5.8971 | 89.7333 | 90.6 | 90.8947 | 8.65 | 89.8333 | 90.95 | 91.24 | 14.9 | 90.5024 | 91.6 | 91.74 | 25.5628 | 91.175 | 92.22 | 92.5654
3 | 4 | 89.25 | 90.2774 | 90.5667 | 5.0475 | 89.45 | 90.425 | 90.6698 | 7.4 | 89.5 | 90.56 | 90.85 | 12.15 | 90.1442 | 91.2 | 91.5 | 19.91 | 90.825 | 91.94 | 92.2154
3.1 | 3.45 | 88.875 | 89.9775 | 90.3 | 4.2979 | 88.9667 | 90.0333 | 90.2699 | 6.5 | 89.1 | 90.3 | 90.65 | 10.5 | 89.6776 | 90.96 | 91.2 | 16.3081 | 90.55 | 91.725 | 91.8986
3.2 | 3.05 | 88.55 | 89.6776 | 89.8667 | 3.7981 | 88.45 | 89.7 | 89.97 | 5.7333 | 88.7 | 89.9667 | 90.34 | 9.3 | 89.1777 | 90.7 | 90.88 | 13.7569 | 90.15 | 91.5 | 91.6486
3.3 | 2.8 | 88.1667 | 89.3377 | 89.66 | 3.4983 | 88.1 | 89.475 | 89.6701 | 5.2 | 88.425 | 89.66 | 90 | 8.1 | 88.7278 | 90.44 | 90.7 | 12.1061 | 89.75 | 91 | 91.3402
3.4 | 2.5 | 87.575 | 89.0277 | 89.38 | 3.1984 | 87.7 | 89.24 | 89.3702 | 4.6 | 87.8 | 89.3 | 89.65 | 7.3 | 88.4779 | 90.1 | 90.4 | 10.3052 | 89.3333 | 90.76 | 91.1152
3.5 | 2.4 | 87.25 | 88.7178 | 89.125 | 2.8986 | 87.3333 | 88.9 | 89.0703 | 4.2 | 87.3 | 89 | 89.38 | 6.6 | 88.178 | 89.8 | 90.1333 | 9.4547 | 89.0667 | 90.575 | 90.8351
3.6 | 2.25 | 86.9333 | 88.4112 | 88.8333 | 2.6987 | 86.7 | 88.6 | 88.8204 | 3.9 | 86.8 | 88.6667 | 89.2 | 6.1 | 87.5781 | 89.5333 | 89.8 | 8.5043 | 88.6667 | 90.325 | 90.6151
3.7 | 2.1 | 86.3667 | 88.178 | 88.44 | 2.4987 | 86.1 | 88.3333 | 88.3705 | 3.7 | 86.5 | 88.5 | 88.85 | 5.8 | 87.2282 | 89.32 | 89.55 | 7.954 | 88.2333 | 90.04 | 90.215
3.8 | 2 | 85.8667 | 87.903 | 88.2 | 2.2988 | 85.8 | 87.9 | 88.2206 | 3.4 | 86 | 88.1333 | 88.525 | 5.4 | 86.8533 | 88.86 | 89.2667 | 7.4037 | 87.9333 | 89.775 | 90.015
3.9 | 1.9 | 85.35 | 87.5281 | 87.8333 | 2.1989 | 85.2 | 87.6333 | 87.9207 | 3.3 | 85.5 | 87.85 | 88.275 | 5.1 | 86.4284 | 88.54 | 88.95 | 6.8034 | 87.5 | 89.4667 | 89.6816
4 | 1.8 | 84.85 | 87.1032 | 87.5 | 2.0989 | 84.5 | 87.3667 | 87.6708 | 3.1 | 84.9 | 87.6 | 88.025 | 4.9 | 85.8286 | 88.2667 | 88.65 | 6.5033 | 87.1 | 89.15 | 89.3816
4.1 | 1.7 | 84 | 86.845 | 87.175 | 1.999 | 83.8667 | 87 | 87.3042 | 2.9 | 84.4667 | 87.275 | 87.6571 | 4.7 | 85.4786 | 87.9667 | 88.4667 | 6.2031 | 86.7 | 88.9 | 89.2149
4.2 | 1.6 | 83.2 | 86.6383 | 86.925 | 1.899 | 83.1 | 86.7333 | 86.9043 | 2.8 | 83.7 | 86.975 | 87.36 | 4.5 | 85.0454 | 87.76 | 88.1333 | 5.8529 | 86.3 | 88.64 | 88.8648
4.3 | 1.6 | 82.3 | 86.3118 | 86.62 | 1.7991 | 82.3 | 86.4 | 86.6378 | 2.6 | 83.1 | 86.64 | 87.1 | 4.2 | 84.4789 | 87.4 | 87.775 | 5.6028 | 85.8 | 88.4 | 88.5398
4.4 | 1.5 | 81.0333 | 85.9785 | 86.34 | 1.6991 | 81.4 | 86.15 | 86.2712 | 2.5 | 82.4 | 86.38 | 86.7 | 4.1 | 83.9457 | 87.1 | 87.6 | 5.3693 | 85.3333 | 88.1833 | 88.2147
4.5 | 1.5 | 79.6333 | 85.6786 | 86.08 | 1.6991 | 80.2 | 85.8333 | 86.0713 | 2.4 | 81.7 | 86.15 | 86.4 | 3.9 | 83.3458 | 86.7833 | 87.2571 | 5.1025 | 84.9667 | 87.85 | 88.0347
4.6 | 1.4 | 77.6 | 85.3986 | 85.8 | 1.5992 | 78.7667 | 85.5667 | 85.7381 | 2.3 | 80.7 | 85.8 | 86.16 | 3.8 | 82.7793 | 86.6 | 86.9143 | 5.0025 | 84.7333 | 87.52 | 87.6813
4.7 | 1.4 | 74.3667 | 85.0787 | 85.45 | 1.4992 | 76.5 | 85.25 | 85.3715 | 2.2 | 79.5333 | 85.475 | 85.85 | 3.6 | 82.1128 | 86.275 | 86.65 | 4.8524 | 84.2333 | 87.28 | 87.4346
4.8 | 1.4 | 69.5333 | 84.7455 | 85.075 | 1.4992 | 73.875 | 84.8333 | 85.1049 | 2.1 | 78.2 | 85.1333 | 85.48 | 3.5 | 81.5796 | 86 | 86.36 | 4.7024 | 83.6 | 86.95 | 87.1395
4.9 | 1.3 | 62.5667 | 84.4589 | 84.82 | 1.3993 | 69.5 | 84.5 | 84.8467 | 2 | 76.6 | 84.75 | 85.26 | 3.4 | 80.9131 | 85.68 | 86.1333 | 4.6023 | 83.05 | 86.72 | 86.9395
5 | 1.3 | 54.6 | 84.104 | 84.55 | 1.3993 | 63.5 | 84.2 | 84.5218 | 2 | 74.1 | 84.6 | 84.92 | 3.2 | 80.03 | 85.4 | 85.8 | 4.4022 | 82.5 | 86.42 | 86.6644
5.1 | 1.3 | 45.6 | 83.7791 | 84.25 | 1.2994 | 55.85 | 83.9 | 84.1386 | 1.9 | 71.1 | 84.1667 | 84.625 | 3.15 | 79.0053 | 85.1 | 85.5 | 4.3022 | 81.9 | 86.125 | 86.2477
5.2 | 1.2 | 37.0667 | 83.4291 | 83.86 | 1.2994 | 48.0333 | 83.65 | 83.847 | 1.9 | 66.6 | 83.85 | 84.375 | 3.05 | 78.1055 | 84.7 | 85.1 | 4.2021 | 81.35 | 85.8 | 86.0143
5.3 | 1.2 | 28.3 | 83.0793 | 83.575 | 1.2994 | 38.9 | 83.3667 | 83.5471 | 1.8 | 59.9 | 83.5667 | 83.95 | 2.95 | 76.6308 | 84.425 | 84.9 | 4.052 | 80.675 | 85.5 | 85.7476
5.4 | 1.1 | 21.3 | 82.7993 | 83.25 | 1.1994 | 31.2667 | 83.06 | 83.2972 | 1.75 | 53.2 | 83.4 | 83.6 | 2.9 | 74.9813 | 84.1 | 84.55 | 4.002 | 80 | 85.26 | 85.4476
5.5 | 1.1 | 16.5 | 82.5044 | 83 | 1.1994 | 23.8 | 82.725 | 83.0223 | 1.7 | 45.1 | 82.94 | 83.42 | 2.8 | 72.7818 | 83.85 | 84.225 | 3.9686 | 79.4 | 84.875 | 85.1142
5.6 | 1.1 | 12.9 | 82.1795 | 82.66 | 1.1994 | 18.1 | 82.4 | 82.6224 | 1.7 | 36.7667 | 82.7167 | 83.075 | 2.7 | 69.6326 | 83.5 | 84 | 3.9019 | 78.3 | 84.5 | 84.8141
5.7 | 1.1 | 10.8 | 81.9462 | 82.325 | 1.0995 | 14.6667 | 82.0333 | 82.3226 | 1.6 | 29.5 | 82.3333 | 82.6 | 2.6 | 65.3836 | 83.26 | 83.7 | 3.7018 | 77.225 | 84.3833 | 84.5941
5.8 | 1 | 9.375 | 81.6796 | 81.95 | 1.0995 | 12.3 | 81.8 | 81.9477 | 1.5 | 23.1 | 81.95 | 82.3 | 2.6 | 59.1852 | 83 | 83.425 | 3.6018 | 75.95 | 84.1333 | 84.4341
5.9 | 1 | 8.275 | 81.2797 | 81.58 | 1.0995 | 10.6 | 81.45 | 81.6728 | 1.5 | 18.8333 | 81.68 | 82.075 | 2.5 | 52.8868 | 82.64 | 83 | 3.5017 | 74.3 | 83.7 | 84.154
6 | 1 | 7.4667 | 80.8998 | 81.24 | 1.0495 | 9.4 | 81.0333 | 81.3729 | 1.45 | 15.6333 | 81.25 | 81.85 | 2.4 | 45.5886 | 82.24 | 82.75 | 3.4017 | 72.2 | 83.4714 | 83.814
6.1 | 1 | 6.8667 | 80.6298 | 81.125 | 0.9995 | 8.35 | 80.7 | 80.973 | 1.4 | 13.5 | 80.9667 | 81.4 | 2.4 | 38.657 | 81.98 | 82.3333 | 3.3016 | 69.25 | 83.14 | 83.4139
6.2 | 1 | 6.4 | 80.28 | 80.72 | 0.9995 | 7.7333 | 80.4 | 80.6231 | 1.4 | 12 | 80.7333 | 81.15 | 2.3 | 31.8254 | 81.7 | 82.02 | 3.2016 | 65.4 | 82.9143 | 83.0138
6.3 | 1 | 6.0333 | 79.9 | 80.4333 | 0.9995 | 7.15 | 80.05 | 80.3065 | 1.4 | 10.9 | 80.32 | 80.825 | 2.3 | 25.8602 | 81.45 | 81.7333 | 3.2016 | 60.4 | 82.575 | 82.8338
6.4 | 1 | 5.7 | 79.5601 | 80.05 | 0.9995 | 6.6667 | 79.65 | 80.04 | 1.3 | 10.0667 | 80.04 | 80.46 | 2.2 | 21.7446 | 81.0571 | 81.4 | 3.1016 | 54.3 | 82.36 | 82.5138
6.5 | 0.9 | 5.45 | 79.2602 | 79.74 | 0.9995 | 6.2333 | 79.4 | 79.7984 | 1.3 | 9.25 | 79.75 | 80.14 | 2.1 | 18.8953 | 80.65 | 81.2 | 3.0015 | 47.475 | 82.025 | 82.2137
6.6 | 0.9 | 5.2 | 78.8603 | 79.35 | 0.9995 | 5.95 | 79 | 79.3485 | 1.3 | 8.6333 | 79.35 | 79.85 | 2.1 | 16.7292 | 80.3 | 80.82 | 3.0015 | 40.9 | 81.7 | 81.9337
6.7 | 0.9 | 5 | 78.5604 | 79.02 | 0.89955 | 5.6 | 78.72 | 79.0486 | 1.3 | 8.2 | 79.075 | 79.5 | 2 | 15.0213 | 80 | 80.4833 | 2.9015 | 34.3 | 81.3 | 81.6136
6.8 | 0.9 | 4.85 | 78.2805 | 78.6667 | 0.89955 | 5.3667 | 78.4333 | 78.7238 | 1.2 | 7.7 | 78.825 | 79.2 | 2 | 13.7299 | 79.7333 | 80.2 | 2.8014 | 28.8 | 81.02 | 81.2635
6.9 | 0.9 | 4.7 | 77.9055 | 78.425 | 0.89955 | 5.1667 | 78.15 | 78.4072 | 1.2 | 7.4333 | 78.28 | 78.84 | 2 | 12.8635 | 79.375 | 79.86 | 2.8014 | 24.5 | 80.8 | 80.9735
7 | 0.9 | 4.6 | 77.6206 | 78.0333 | 0.89955 | 4.9667 | 77.7667 | 77.974 | 1.2 | 7.125 | 78.04 | 78.4667 | 1.9 | 12.147 | 79.02 | 79.6333 | 2.7014 | 21.5 | 80.475 | 80.7635
7.1 |  | 4.4 | 77.1807 | 77.66 |  | 4.8 | 77.3667 | 77.7741 | 1.1 | 6.875 | 77.6 | 78.16 | 1.9 | 11.4305 | 78.8 | 79.2 | 2.6013 | 19.2 | 80.1333 | 80.3384
7.2 |  | 4.275 | 76.9558 | 77.36 |  | 4.6667 | 77.1 | 77.2742 | 1.1 | 6.6 | 77.325 | 77.82 | 1.9 | 10.9973 | 78.375 | 78.85 | 2.5013 | 17.3 | 79.7667 | 80.0133
7.3 |  | 4.1667 | 76.5475 | 77.0286 |  | 4.5 | 76.7667 | 76.9743 |  | 6.35 | 77 | 77.34 | 1.8 | 10.4974 | 78.06 | 78.625 | 2.5013 | 16.2 | 79.44 | 79.7133
7.4 |  | 4 | 76.156 | 76.65 |  | 4.3 | 76.4 | 76.6744 |  | 6.2 | 76.68 | 77.125 | 1.7 | 10.1975 | 77.76 | 78.2667 | 2.4012 | 15.3 | 79.1 | 79.4883
7.5 |  | 4 | 75.881 | 76.3 |  | 4.2 | 76.1 | 76.2746 |  | 5.9 | 76.3 | 76.8 |  | 9.8475 | 77.36 | 77.8667 | 2.4012 | 14.4 | 78.8667 | 79.1799
7.6 |  | 3.9 | 75.6411 | 76 |  | 4.1 | 75.6 | 76.008 |  | 5.7667 | 75.8667 | 76.4667 |  | 9.4726 | 77.06 | 77.65 | 2.4012 | 13.7 | 78.45 | 78.8732
7.7 |  | 3.75 | 75.2612 | 75.7 |  | 4 | 75.3333 | 75.6415 |  | 5.5333 | 75.5429 | 76.04 |  | 9.0977 | 76.76 | 77.2 |  | 13.1 | 78.1333 | 78.5131
7.8 |  | 3.7 | 74.8313 | 75.4 |  | 3.9 | 74.975 | 75.3749 |  | 5.4 | 75.3167 | 75.7 |  | 8.8978 | 76.3714 | 76.875 |  | 12.7 | 77.7 | 78.253
7.9 |  | 3.6 | 74.5314 | 74.88 |  | 3.7 | 74.7 | 74.975 |  | 5.2 | 75.02 | 75.45 |  | 8.6312 | 76.1333 | 76.58 |  | 12.2667 | 77.4667 | 77.963
8 |  | 3.5 | 74.2314 | 74.65 |  | 3.7 | 74.3 | 74.5418 |  | 5.1 | 74.56 | 75.08 |  | 8.2979 | 75.825 | 76.22 |  | 11.8333 | 77.12 | 77.5129
8.1 |  | 3.4 | 73.7616 | 74.15 |  | 3.6 | 74 | 74.2253 |  | 4.9 | 74.22 | 74.8 |  | 8.1646 | 75.275 | 75.9 |  | 11.6 | 76.875 | 77.1415
8.2 |  | 3.4 | 73.4817 | 73.9 |  | 3.5 | 73.7 | 73.8754 |  | 4.8333 | 73.875 | 74.42 |  | 7.998 | 74.95 | 75.4333 |  | 11.3 | 76.625 | 76.9128
8.3 |  | 3.3 | 73.0218 | 73.55 |  | 3.4 | 73.24 | 73.4088 |  | 4.7 | 73.58 | 74.06 |  | 7.7731 | 74.64 | 75.0667 |  | 11.025 | 76.4 | 76.5961
8.4 |  | 3.2 | 72.6818 | 73.225 |  | 3.3 | 72.7667 | 73.1423 |  | 4.5 | 73.16 | 73.65 |  | 7.4981 | 74.3 | 74.7333 |  | 10.7 | 75.85 | 76.2627
8.5 |  | 3.2 | 72.3219 | 72.75 |  | 3.3 | 72.4 | 72.7424 |  | 4.5 | 72.72 | 73.2333 |  | 7.3648 | 73.8 | 74.4 |  | 10.425 | 75.5 | 75.8926
8.6 |  | 3.1 | 71.782 | 72.35 |  | 3.2 | 72.05 | 72.3259 |  | 4.4 | 72.3 | 72.84 |  | 7.1982 | 73.6 | 74.1 |  | 10.175 | 75.15 | 75.5626
8.7 |  | 3.1 | 71.5071 | 72 |  | 3.1 | 71.6667 | 71.976 |  | 4.3 | 71.9333 | 72.54 |  | 7.0982 | 73.24 | 73.7667 |  | 10 | 74.7667 | 75.2726
8.8 |  | 3 | 71.1222 | 71.58 |  | 3.0333 | 71.1667 | 71.5761 |  | 4.2 | 71.65 | 72.2 |  | 6.8983 | 72.8 | 73.36 |  | 9.7333 | 74.4333 | 74.9125
8.9 |  | 2.9 | 70.6823 | 71.15 |  | 3 | 70.7 | 71.0763 |  | 4.0667 | 71.2 | 71.7333 |  | 6.6983 | 72.52 | 73.025 |  | 9.425 | 74.1 | 74.6124
9 |  | 2.9 | 70.1025 | 70.725 |  | 2.9 | 70.4 | 70.6431 |  | 4 | 70.8 | 71.3 |  | 6.5983 | 72.1 | 72.74 |  | 9.3 | 73.7714 | 74.2791
9.1 |  | 2.8 | 69.6626 | 70.22 |  | 2.9 | 70 | 70.3265 |  | 3.825 | 70.4 | 70.95 |  | 6.3984 | 71.7667 | 72.3 |  | 9.1 | 73.44 | 73.9623
9.2 |  | 2.8 | 69.2027 | 69.65 |  | 2.8 | 69.4 | 69.7267 |  | 3.8 | 70.04 | 70.55 |  | 6.1985 | 71.3667 | 72 |  | 8.9 | 73.06 | 73.579
9.3 |  | 2.7 | 68.4629 | 68.975 |  | 2.8 | 68.8 | 69.2436 |  | 3.7 | 69.55 | 70.12 |  | 6.0985 | 70.98 | 71.575 |  | 8.8 | 72.7667 | 73.2622
9.4 |  | 2.7 | 67.7831 | 68.2 |  | 2.7 | 68.25 | 68.6771 |  | 3.65 | 69.1286 | 69.55 |  | 5.9985 | 70.74 | 71.175 |  | 8.62 | 72.46 | 73.0122
9.5 |  | 2.6 | 66.8833 | 67.325 |  | 2.7 | 67.75 | 68.1106 |  | 3.6 | 68.74 | 69.14 |  | 5.7985 | 70.28 | 70.775 |  | 8.4 | 72.1333 | 72.6621
9.6 |  | 2.6 | 65.6836 | 66.3 |  | 2.6 | 67.1 | 67.2775 |  | 3.5 | 68.12 | 68.66 |  | 5.6986 | 69.8333 | 70.35 |  | 8.225 | 71.7667 | 72.112
9.7 |  | 2.525 | 64.2639 | 64.9 |  | 2.6 | 66.05 | 66.4028 |  | 3.5 | 67.75 | 68.2 |  | 5.5986 | 69.4667 | 69.975 |  | 8 | 71.5 | 71.8787
9.8 |  | 2.5 | 62.1845 | 62.9 |  | 2.5 | 65.1 | 65.4781 |  | 3.4 | 67.05 | 67.65 |  | 5.4986 | 69.15 | 69.65 |  | 7.9 | 71.2 | 71.6286
9.9 |  | 2.5 | 58.8853 | 59.8 |  | 2.5 | 63.8333 | 64.3785 |  | 3.3 | 66.5 | 67.1 |  | 5.3986 | 68.76 | 69.3333 |  | 7.8 | 70.75 | 71.1618
10 |  | 2.5 | 54.3364 | 55.4 |  | 2.5 | 62.1 | 62.5792 |  | 3.3 | 66.0333 | 66.4 |  | 5.2487 | 68.3 | 68.82 |  | 7.7 | 70.3333 | 70.7785
10.1 |  | 2.4 | 48.3679 | 49.75 |  | 2.4 | 59.65 | 60.3132 |  | 3.2 | 65.14 | 65.725 |  | 5.0987 | 67.9 | 68.42 |  | 7.55 | 70 | 70.5118
10.2 |  |  | 41.6896 | 43.55 |  | 2.4 | 56.3 | 56.5811 |  | 3.1 | 64.375 | 64.825 |  | 5.0987 | 67.625 | 68.1 |  | 7.4 | 69.68 | 70.0974
10.3 |  |  | 34.8163 | 36.6 |  |  | 50.9 | 52.0493 |  | 3.1 | 63.36 | 63.9 |  | 4.9988 | 67.1 | 67.7 |  | 7.2 | 69.32 | 69.6449
10.4 |  |  | 28.4929 | 30.275 |  |  | 45.2 | 46.3179 |  | 3 | 62.1 | 62.65 |  | 4.8988 | 66.65 | 67.1 |  | 7 | 68.875 | 69.3916
10.5 |  |  | 22.6143 | 24.95 |  |  | 38.9 | 39.8867 |  |  | 60.4667 | 61 |  | 4.7988 | 66.05 | 66.7 |  | 6.9 | 68.42 | 69.0615
10.6 |  |  | 18.1955 | 20.45 |  |  | 32.6 | 33.7887 |  |  | 58 | 58.8 |  | 4.6988 | 65.6 | 66.22 |  | 6.8 | 67.98 | 68.6314
10.7 |  |  | 15.1962 | 17.75 |  |  | 26.7667 | 28.0906 |  |  | 54.5 | 55.8 |  | 4.4989 | 65.075 | 65.68 |  | 6.7 | 67.72 | 68.2781
10.8 |  |  | 13.3967 | 15.9 |  |  | 21.4333 | 23.2589 |  |  | 50.375 | 51.075 |  | 4.4989 | 64.65 | 65.14 |  | 6.5 | 67.2667 | 67.8113
10.9 |  |  | 12.1303 | 14.7333 |  |  | 17.4 | 19.6268 |  |  | 44.7 | 46.4 |  | 4.3989 | 64 | 64.62 |  | 6.425 | 66.9 | 67.3912
11 |  |  | 11.2972 | 13.84 |  |  | 15.125 | 17.2692 |  |  | 39.02 | 40.6 |  |  | 63.3333 | 64 |  | 6.3 | 66.54 | 67.0445
11.1 |  |  | 10.5974 | 13.18 |  |  | 13.4 | 15.6614 |  |  | 33.05 | 34.825 |  |  | 62.7 | 63.3 |  | 6.2 | 66.1 | 66.5911
11.2 |  |  | 10.1475 | 12.7 |  |  | 12.3 | 14.5951 |  |  | 27.4 | 29.56 |  |  | 61.8 | 62.4 |  | 6.1 | 65.6429 | 66.211
11.3 |  |  | 9.7476 | 12.28 |  |  | 11.4 | 13.8287 |  |  | 22.94 | 25.1 |  |  | 60.9 | 61.5667 |  | 6.0333 | 65.35 | 65.811
11.4 |  |  | 9.4776 | 12.025 |  |  | 10.8333 | 13.1623 |  |  | 19.36 | 21.64 |  |  | 59.8 | 60.3 |  | 5.9333 | 64.8667 | 65.3709
11.5 |  |  | 9.1577 | 11.7667 |  |  | 10.2667 | 12.6625 |  |  | 16.8667 | 19.2 |  |  | 58.3 | 58.95 |  |  | 64.5 | 64.9508
11.6 |  |  | 9.0227 | 11.54 |  |  | 9.9 | 12.2626 |  |  | 15.2 | 17.7 |  |  | 56.06 | 57 |  |  | 63.95 | 64.5108
11.7 |  |  | 8.8978 | 11.38 |  |  | 9.6 | 12.0293 |  |  | 14 | 16.5 |  |  | 53.225 | 54.1 |  |  | 63.46 | 64.0707
11.8 |  |  | 8.7478 | 11.275 |  |  | 9.2333 | 11.6628 |  |  | 13.2 | 15.72 |  |  | 49.3333 | 50.3 |  |  | 63.04 | 63.4906
11.9 |  |  | 8.6228 | 11.1667 |  |  | 9.05 | 11.4462 |  |  | 12.54 | 15.08 |  |  | 44.7 | 46.175 |  |  | 62.24 | 62.9605
12 |  |  | 8.4979 | 11.06 |  |  | 8.8333 | 11.2296 |  |  | 11.95 | 14.46 |  |  | 39.375 | 41.1 |  |  | 61.7 | 62.3104
12.1 |  |  | 8.4979 | 10.96 |  |  | 8.7 | 11.0963 |  |  | 11.48 | 14.05 |  |  | 34.5 | 36.12 |  |  | 61 | 61.8246
12.2 |  |  | 8.3979 | 10.9 |  |  | 8.5 | 10.8964 |  |  | 11.1 | 13.7 |  |  | 29.5 | 31.86 |  |  | 60.1571 | 60.953
12.3 |  |  | 8.2979 | 10.8 |  |  | 8.4 | 10.7964 |  |  | 10.775 | 13.36 |  |  | 25.35 | 27.8 |  |  | 59.3833 | 60.01
12.4 |  |  | 8.2979 | 10.8 |  |  | 8.3 | 10.7631 |  |  | 10.55 | 13.08 |  |  | 22.2 | 24.6 |  |  | 58.24 | 58.9098
12.5 |  |  | 8.2646 | 10.7 |  |  | 8.3 | 10.5965 |  |  | 10.3 | 12.875 |  |  | 19.95 | 22.34 |  |  | 56.65 | 57.5096
12.6 |  |  | 8.198 | 10.65 |  |  | 8.14 | 10.4965 |  |  | 10.08 | 12.625 |  |  | 18.3 | 20.82 |  |  | 54.78 | 55.6426
12.7 |  |  | 8.098 | 10.6 |  |  | 8.1 | 10.4632 |  |  | 9.975 | 12.42 |  |  | 16.975 | 19.7 |  |  | 52.2 | 53.2089
12.8 |  |  | 8.098 | 10.5 |  |  | 8 | 10.3965 |  |  | 9.7 | 12.3 |  |  | 16.16 | 18.8 |  |  | 48.72 | 49.8083
12.9 |  |  | 7.998 | 10.48 |  |  | 8 | 10.2966 |  |  | 9.6 | 12.1 |  |  | 15.4 | 18.05 |  |  | 44.8 | 45.9577
13 |  |  | 7.898 | 10.4 |  |  | 7.9 | 10.1966 |  |  | 9.4 | 11.98 |  |  | 14.8667 | 17.56 |  |  | 40.34 | 41.907
13.1 |  |  | 7.873 | 10.3167 |  |  | 7.8 | 10.1966 |  |  | 9.3 | 11.8 |  |  | 14.36 | 17.1 |  |  | 35.7333 | 37.4563
13.2 |  |  | 7.798 | 10.275 |  |  | 7.7333 | 10.0966 |  |  | 9.2 | 11.675 |  |  | 13.9 | 16.675 |  |  | 31.4 | 33.4722
13.3 |  |  | 7.6981 | 10.1667 |  |  | 7.7 | 9.9967 |  |  | 9.1 | 11.54 |  |  | 13.575 | 16.35 |  |  | 27.3333 | 29.8383
13.4 |  |  | 7.5981 | 10.1 |  |  | 7.6 | 9.8967 |  |  | 9 | 11.4 |  |  | 13.22 | 16.05 |  |  | 24.18 | 26.8045
13.5 |  |  | 7.5648 | 9.9833 |  |  | 7.5 | 9.7967 |  |  | 8.9 | 11.3 |  |  | 12.975 | 15.72 |  |  | 22.0571 | 24.8041
13.6 |  |  | 7.4981 | 9.9 |  |  | 7.5 | 9.7218 |  |  | 8.8 | 11.2 |  |  | 12.75 | 15.475 |  |  | 20.4 | 23.2372
13.7 |  |  | 7.3982 | 9.8 |  |  | 7.4 | 9.6968 |  |  | 8.7 | 11.18 |  |  | 12.425 | 15.2333 |  |  | 19.2333 | 22.1637
13.8 |  |  | 7.2982 | 9.7 |  |  | 7.3 | 9.5968 |  |  | 8.6 | 11.025 |  |  | 12.22 | 14.98 |  |  | 18.4 | 21.2835
13.9 |  |  | 7.2232 | 9.6 |  |  | 7.2 | 9.4968 |  |  | 8.5 | 10.9 |  |  | 12.05 | 14.76 |  |  | 17.6 | 20.5701
14 |  |  | 7.1482 | 9.5 |  |  | 7.1 | 9.3969 |  |  | 8.4 | 10.8 |  |  | 11.94 | 14.58 |  |  | 17.1 | 19.9533
14.1 |  |  | 7.0816 | 9.4 |  |  | 7.1 | 9.2969 |  |  | 8.3 | 10.7 |  |  | 11.7 | 14.375 |  |  | 16.55 | 19.4699
14.2 |  |  | 6.9149 | 9.3 |  |  | 7 | 9.1969 |  |  | 8.2 | 10.6 |  |  | 11.475 | 14.225 |  |  | 16.1 | 19.1199
14.3 |  |  | 6.8483 | 9.2 |  |  | 6.85 | 9.097 |  |  | 8.06 | 10.5 |  |  | 11.3 | 14.05 |  |  | 15.675 | 18.7031
14.4 |  |  | 6.7555 | 9.1 |  |  | 6.8 | 8.997 |  |  | 8 | 10.4 |  |  | 11.1667 | 13.9 |  |  | 15.3714 | 18.4364
14.5 |  |  | 6.6583 | 9 |  |  | 6.7 | 8.897 |  |  | 7.84 | 10.3 |  |  | 11.1 | 13.7 |  |  | 15.0667 | 18.078
14.6 |  |  | 6.5983 | 8.92 |  |  | 6.6 | 8.7971 |  |  | 7.8 | 10.15 |  |  | 10.8833 | 13.575 |  |  | 14.82 | 17.843
14.7 |  |  | 6.4984 | 8.8 |  |  | 6.5 | 8.6971 |  |  | 7.7 | 10.1 |  |  | 10.7 | 13.4 |  |  | 14.5 | 17.5029
14.8 |  |  | 6.3841 | 8.7167 |  |  | 6.425 | 8.5971 |  |  | 7.58 | 9.9667 |  |  | 10.6167 | 13.28 |  |  | 14.24 | 17.2529
14.9 |  |  | 6.2984 | 8.6 |  |  | 6.4 | 8.4972 |  |  | 7.5 | 9.8 |  |  | 10.5 | 13.15 |  |  | 14.1 | 17.0028
15 |  |  | 6.1985 | 8.5 |  |  | 6.3 | 8.3972 |  |  | 7.4 | 9.725 |  |  | 10.325 | 13.0333 |  |  | 13.8833 | 16.8628
15.1 |  |  | 6.0985 | 8.4 |  |  | 6.14 | 8.3472 |  |  | 7.3 | 9.65 |  |  | 10.2 | 12.9 |  |  | 13.7 | 16.7278
15.2 |  |  | 5.9985 | 8.3 |  |  | 6.04 | 8.2972 |  |  | 7.2 | 9.5 |  |  | 10.06 | 12.72 |  |  | 13.4 | 16.5028
15.3 | [6 empty] | 5.9 | 8.1973 |  |  | 7.1 | 9.4167 |  |  | 9.95 | 12.6 |  |  | 13.275 | 16.3027
15.4 | [6 empty] | 5.9 | 8.0723 |  |  | 7.05 | 9.3 |  |  | 9.84 | 12.5 |  |  | 13.1 | 16.1027
15.5 | [10 empty] | 6.95 | 9.2167 |  |  | 9.72 | 12.36 |  |  | 12.96 | 15.9027
15.6 | [10 empty] | 6.9 | 9.1 |  |  | 9.5333 | 12.18 |  |  | 12.7833 | 15.7026
15.7 | [10 empty] | 6.7 | 9 |  |  | 9.4 | 12.0833 |  |  | 12.6 | 15.5626
15.8 | [10 empty] | 6.6 | 8.9 |  |  | 9.3 | 11.975 |  |  | 12.48 | 15.3426
15.9 | [10 empty] | 6.5 | 8.8 |  |  | 9.2 | 11.8 |  |  | 12.3 | 15.2275
16 | [14 empty] | 9.1 | 11.7 |  |  | 12.1 | 15.1025
16.1 | [14 empty] | 8.96 | 11.54 |  |  | 11.95 | 14.9025
16.2 | [14 empty] | 8.8 | 11.4 |  |  | 11.8 | 14.8025
16.3 | [14 empty] | 8.7 | 11.3 |  |  | 11.7 | 14.6024
16.4 | [14 empty] | 8.6 | 11.22 |  |  | 11.6167 | 14.5024
16.5 | [14 empty] | 8.5167 | 11.0286 |  |  | 11.42 | 14.3024
16.6 | [14 empty] | 8.4 | 10.925 |  |  | 11.3429 | 14.1274
16.7 | [14 empty] | 8.2167 | 10.8 |  |  | 11.16 | 14.0774
16.8 | [14 empty] | 8.1333 | 10.7 |  |  | 11.05 | 13.9773
16.9 | [14 empty] | 8.06 | 10.62 |  |  | 10.95 | 13.8023
17 | [18 empty] | 10.75 | 13.5773
17.1 | [18 empty] | 10.6667 | 13.4022
17.2 | [18 empty] | 10.56 | 13.3355
17.3 | [18 empty] | 10.3667 | 13.1022
17.4 | [18 empty] | 10.2 | 13.0022
17.5 | [18 empty] | 10.1 | 12.8307
17.6 | [18 empty] | 10 | 12.7021
17.7 | [18 empty] | 9.9 | 12.7021
17.8 | [18 empty] | 9.78 | 12.5421
17.9 | [18 empty] | 9.6 | 12.3688

## Wedge 45 Longitudinal Profiles
Longitudinal Profiles
Wedge 45
SSD=100cm
Field Size=40x20cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2.4 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 12
-30.8 |  |  |  |  | 12.125
-30.7 |  |  |  |  | 12.25
-30.6 |  |  |  |  | 12.375
-30.5 |  |  |  |  | 12.5167
-30.4 |  |  |  |  | 12.6
-30.3 |  |  |  |  | 12.7714
-30.2 |  |  |  |  | 12.9
-30.1 |  |  |  |  | 13.05
-30 |  |  |  |  | 13.3
-29.9 |  |  |  |  | 13.475
-29.8 |  |  |  |  | 13.6
-29.7 |  |  |  |  | 13.725
-29.6 |  |  |  |  | 13.9333
-29.5 |  |  |  |  | 14.1286
-29.4 |  |  |  |  | 14.3
-29.3 |  |  |  |  | 14.4667
-29.2 |  |  |  |  | 14.6333
-29.1 |  |  |  |  | 14.775
-29 |  |  |  |  | 15.0286
-28.9 |  |  |  | 10.4 | 15.1
-28.8 |  |  |  | 10.4429 | 15.325
-28.7 |  |  |  | 10.6 | 15.5
-28.6 |  |  |  | 10.7 | 15.75
-28.5 |  |  |  | 10.7857 | 15.9
-28.4 |  |  |  | 10.9429 | 16.0889
-28.3 |  |  |  | 11.075 | 16.3429
-28.2 |  |  |  | 11.2 | 16.7
-28.1 |  |  |  | 11.25 | 16.825
-28 |  |  |  | 11.375 | 17.075
-27.9 |  |  |  | 11.6333 | 17.3111
-27.8 |  |  |  | 11.7857 | 17.625
-27.7 |  |  |  | 11.8375 | 17.8
-27.6 |  |  |  | 12.025 | 18.175
-27.5 |  |  |  | 12.1875 | 18.5
-27.4 |  |  |  | 12.4125 | 18.7714
-27.3 |  |  |  | 12.5333 | 19.0857
-27.2 |  |  |  | 12.65 | 19.4571
-27.1 |  |  |  | 12.8714 | 19.9
-27 |  |  |  | 13.0571 | 20.5
-26.9 |  |  | 8.6 | 13.2833 | 21.075
-26.8 |  |  | 8.6571 | 13.4571 | 21.95
-26.7 |  |  | 8.8 | 13.7 | 22.975
-26.6 |  |  | 8.9 | 13.85 | 24.3
-26.5 |  |  | 8.9 | 14.05 | 25.9778
-26.4 |  |  | 9 | 14.2 | 28.05
-26.3 |  |  | 9.1 | 14.4333 | 30.925
-26.2 |  |  | 9.1857 | 14.6833 | 34.2625
-26.1 |  |  | 9.32 | 14.925 | 38.3
-26 |  |  | 9.4 | 15.1429 | 42.45
-25.9 |  | 7.8 | 9.5 | 15.4 | 47.3
-25.8 |  | 7.8667 | 9.6 | 15.65 | 52.15
-25.7 |  | 7.9 | 9.76 | 15.8714 | 56.7
-25.6 |  | 8 | 9.8333 | 16.1333 | 60.85
-25.5 |  | 8.1 | 10 | 16.4 | 64.4143
-25.4 | 7.8714 | 8.1667 | 10.05 | 16.7667 | 67.2
-25.3 | 7.9 | 8.2 | 10.1667 | 17.15 | 69.4857
-25.2 | 8 | 8.25 | 10.3 | 17.6143 | 71.3125
-25.1 | 8 | 8.3167 | 10.4667 | 18.05 | 72.65
-25 | 8.1 | 8.4 | 10.6143 | 18.6625 | 73.7
-24.9 | 8.1167 | 8.5 | 10.75 | 19.35 | 74.65
-24.8 | 8.2 | 8.5625 | 10.8857 | 20.1 | 75.46
-24.7 | 8.2714 | 8.6 | 11.15 | 21.26 | 76
-24.6 | 8.3143 | 8.6571 | 11.2 | 22.6714 | 76.6333
-24.5 | 8.4 | 8.7875 | 11.4429 | 24.5 | 77.06
-24.4 | 8.48 | 8.8167 | 11.6 | 26.9857 | 77.5571
-24.3 | 8.5 | 8.9714 | 11.7714 | 30.3667 | 77.9
-24.2 | 8.6 | 9 | 12.0286 | 34.45 | 78.275
-24.1 | 8.6833 | 9.1 | 12.1833 | 38.9625 | 78.6714
-24 | 8.75 | 9.2286 | 12.425 | 44.6286 | 78.975
-23.9 | 8.8 | 9.3833 | 12.6 | 50.2 | 79.3429
-23.8 | 8.9 | 9.4667 | 12.85 | 55.8 | 79.6571
-23.7 | 8.9 | 9.6143 | 13.125 | 60.8 | 80.05
-23.6 | 9 | 9.7571 | 13.3857 | 65.48 | 80.2333
-23.5 | 9.08 | 9.9286 | 13.5857 | 69.3375 | 80.5
-23.4 | 9.15 | 10.1 | 13.9333 | 72.5 | 80.8875
-23.3 | 9.3286 | 10.15 | 14.2875 | 74.9 | 81.175
-23.2 | 9.4625 | 10.3167 | 14.7 | 76.6667 | 81.5222
-23.1 | 9.6 | 10.5429 | 15.1857 | 78.1 | 81.725
-23 | 9.6667 | 10.7 | 15.8 | 79.0667 | 81.85
-22.9 | 9.82 | 10.86 | 16.5 | 79.85 | 82.05
-22.8 | 9.9 | 11.0333 | 17.325 | 80.5 | 82.4
-22.7 | 10.0333 | 11.3 | 18.5 | 81.2 | 82.6
-22.6 | 10.1857 | 11.45 | 20.075 | 81.5 | 82.775
-22.5 | 10.38 | 11.7 | 22.0875 | 82.1 | 83
-22.4 | 10.46 | 11.975 | 25.025 | 82.6 | 83.275
-22.3 | 10.64 | 12.3 | 28.8 | 83.0333 | 83.4625
-22.2 | 10.8143 | 12.6667 | 33.9857 | 83.325 | 83.675
-22.1 | 11.0143 | 13.0143 | 39.6143 | 83.6 | 83.825
-22 | 11.2 | 13.5857 | 46 | 83.9857 | 84.0429
-21.9 | 11.4333 | 14.2667 | 53.44 | 84.3571 | 84.2167
-21.8 | 11.74 | 15.1 | 60.2 | 84.6333 | 84.4667
-21.7 | 11.9857 | 16.36 | 66.5667 | 84.9429 | 84.6429
-21.6 | 12.4286 | 18 | 72.125 | 85.225 | 84.7857
-21.5 | 12.9 | 20.26 | 76.3 | 85.4667 | 85.025
-21.4 | 13.4667 | 23.4286 | 79.6 | 85.7 | 85.18
-21.3 | 14.3 | 27.95 | 81.7333 | 85.95 | 85.3286
-21.2 | 15.4143 | 33.5 | 83.48 | 86.1333 | 85.4714
-21.1 | 16.9286 | 40.0571 | 84.85 | 86.4 | 85.64
-21 | 19.2333 | 46.9 | 85.725 | 86.6143 | 85.8
-20.9 | 22.6 | 55.25 | 86.4333 | 86.9625 | 85.96
-20.8 | 26.94 | 63.0286 | 87.3 | 87.2167 | 86.14
-20.7 | 33 | 70.1143 | 87.8 | 87.3833 | 86.3286
-20.6 | 39.5667 | 75.96 | 88.2333 | 87.58 | 86.5667
-20.5 | 47.9857 | 80.45 | 88.7 | 87.7375 | 86.7
-20.4 | 56.1 | 83.5 | 89.125 | 87.8625 | 86.7857
-20.3 | 64.4 | 86 | 89.5333 | 88.075 | 87
-20.2 | 72.0667 | 87.66 | 89.8429 | 88.3 | 87.125
-20.1 | 78.1 | 88.8 | 90 | 88.4 | 87.25
-20 | 82.9333 | 89.7143 | 90.2857 | 88.5667 | 87.375
-19.9 | 86.2 | 90.4143 | 90.6 | 88.7125 | 87.6
-19.8 | 88.6 | 91.1167 | 90.7833 | 88.875 | 87.725
-19.7 | 90.2 | 91.5 | 91.0286 | 89.0714 | 87.8667
-19.6 | 91.4 | 91.8625 | 91.2571 | 89.3 | 88
-19.5 | 92.32 | 92.2667 | 91.4 | 89.4429 | 88.1286
-19.4 | 92.9333 | 92.6571 | 91.7143 | 89.5857 | 88.2714
-19.3 | 93.4 | 92.9 | 91.9143 | 89.7286 | 88.4
-19.2 | 93.7667 | 93.14 | 92.0571 | 89.8714 | 88.5
-19.1 | 94.1125 | 93.3857 | 92.2 | 90 | 88.6429
-19 | 94.3714 | 93.6286 | 92.3667 | 90.1143 | 88.8
-18.9 | 94.6667 | 93.9 | 92.6 | 90.25 | 88.95
-18.8 | 94.9 | 94.0286 | 92.6714 | 90.3857 | 89.1
-18.7 | 95 | 94.1 | 92.8167 | 90.5222 | 89.1667
-18.6 | 95.24 | 94.2571 | 92.9625 | 90.65 | 89.375
-18.5 | 95.4 | 94.4714 | 93.1 | 90.8286 | 89.5167
-18.4 | 95.5 | 94.6 | 93.2 | 90.9625 | 89.6833
-18.3 | 95.7571 | 94.7571 | 93.3167 | 91 | 89.75
-18.2 | 95.9 | 94.8875 | 93.4 | 91.1 | 89.9
-18.1 | 95.9571 | 94.95 | 93.6125 | 91.225 | 89.9375
-18 | 96.1333 | 95.1125 | 93.7375 | 91.3286 | 90.0833
-17.9 | 96.2857 | 95.2 | 93.8625 | 91.525 | 90.1571
-17.8 | 96.35 | 95.2714 | 94 | 91.6 | 90.2875
-17.7 | 96.4 | 95.4 | 94.0429 | 91.7 | 90.46
-17.6 | 96.5 | 95.45 | 94.175 | 91.86 | 90.5167
-17.5 | 96.6 | 95.5 | 94.3 | 91.925 | 90.6714
-17.4 | 96.7 | 95.6 | 94.35 | 92.05 | 90.8
-17.3 | 96.72 | 95.7 | 94.4 | 92.1857 | 90.9125
-17.2 | 96.8 | 95.7875 | 94.4833 | 92.3125 | 91.0429
-17.1 | 96.9 | 95.9 | 94.6 | 92.4375 | 91.1857
-17 | 97 | 95.9625 | 94.7 | 92.5714 | 91.35
-16.9 | 97.05 | 96.1 | 94.7 | 92.7125 | 91.4875
-16.8 | 97.1 | 96.1375 | 94.78 | 92.8 | 91.6125
-16.7 | 97.1571 | 96.2 | 94.9 | 92.8625 | 91.7375
-16.6 | 97.2 | 96.3 | 94.9833 | 92.9143 | 91.8
-16.5 | 97.2 | 96.3 | 95.0667 | 93.0571 | 91.9
-16.4 | 97.3 | 96.3 | 95.2 | 93.2 | 91.975
-16.3 | 97.3857 | 96.4 | 95.2571 | 93.25 | 92.2
-16.2 | 97.4286 | 96.4 | 95.4 | 93.375 | 92.28
-16.1 | 97.5 | 96.5 | 95.4571 | 93.4889 | 92.34
-16 | 97.6 | 96.6 | 95.5 | 93.6286 | 92.5167
-15.9 | 97.6 | 96.6 | 95.625 | 93.7625 | 92.6
-15.8 | 97.6 | 96.66 | 95.7 | 93.8875 | 92.7571
-15.7 | 97.7 | 96.7333 | 95.7857 | 94 | 92.8875
-15.6 | 97.72 | 96.8 | 95.8333 | 94.0833 | 92.9125
-15.5 | 97.8 | 96.9 | 95.975 | 94.1167 | 93
-15.4 | 97.8667 | 96.9 | 96 | 94.2625 | 93.1167
-15.3 | 97.9 | 96.9875 | 96.1 | 94.3778 | 93.2714
-15.2 | 97.9857 | 97.1 | 96.1 | 94.5 | 93.4
-15.1 | 98.1 | 97.1 | 96.2143 | 94.525 | 93.5
-15 | 98.1857 | 97.26 | 96.3 | 94.65 | 93.55
-14.9 | 98.2 | 97.3333 | 96.4 | 94.7857 | 93.7
-14.8 | 98.2 | 97.4143 | 96.4875 | 94.8429 | 93.8714
-14.7 | 98.2667 | 97.5 | 96.5429 | 94.9 | 93.92
-14.6 | 98.3286 | 97.5 | 96.6 | 95.0222 | 94
-14.5 | 98.4 | 97.6 | 96.7 | 95.1375 | 94.1
-14.4 | 98.4167 | 97.6 | 96.775 | 95.2625 | 94.2125
-14.3 | 98.5 | 97.7 | 96.825 | 95.3 | 94.3
-14.2 | 98.6 | 97.7167 | 96.9 | 95.4286 | 94.3625
-14.1 | 98.6714 | 97.8 | 97 | 95.5 | 94.4875
-14 | 98.7 | 97.9 | 97.0833 | 95.6125 | 94.6286
-13.9 | 98.7286 | 97.9 | 97.1571 | 95.7375 | 94.7625
-13.8 | 98.8 | 97.98 | 97.2125 | 95.8 | 94.8
-13.7 | 98.9 | 98.0286 | 97.3 | 95.9 | 94.9
-13.6 | 98.9286 | 98.1 | 97.4286 | 95.9667 | 94.9
-13.5 | 99 | 98.1333 | 97.5 | 96.04 | 95
-13.4 | 99 | 98.2 | 97.6 | 96.2 | 95
-13.3 | 99.1 | 98.3 | 97.6 | 96.225 | 95.1143
-13.2 | 99.2 | 98.4 | 97.7 | 96.3 | 95.2571
-13.1 | 99.2 | 98.48 | 97.7 | 96.3 | 95.3333
-13 | 99.3 | 98.5 | 97.8375 | 96.3444 | 95.4167
-12.9 | 99.3 | 98.5833 | 97.9 | 96.4714 | 95.5833
-12.8 | 99.3714 | 98.6286 | 97.9125 | 96.5 | 95.6375
-12.7 | 99.4 | 98.7 | 98.0429 | 96.6 | 95.7
-12.6 | 99.4 | 98.8 | 98.1 | 96.6714 | 95.825
-12.5 | 99.5 | 98.8 | 98.1429 | 96.8 | 95.9571
-12.4 | 99.6 | 98.9 | 98.2 | 96.9 | 96
-12.3 | 99.6833 | 98.9 | 98.2667 | 97 | 96.06
-12.2 | 99.7429 | 99 | 98.35 | 97 | 96.2
-12.1 | 99.8 | 99.0571 | 98.4143 | 97.1 | 96.25
-12 | 99.8667 | 99.1 | 98.5 | 97.2 | 96.3
-11.9 | 99.9 | 99.15 | 98.6 | 97.2 | 96.4
-11.8 | 99.9 | 99.2 | 98.6286 | 97.2625 | 96.4
-11.7 | 99.9 | 99.3 | 98.7 | 97.3 | 96.4
-11.6 | 99.95 | 99.3 | 98.8 | 97.38 | 96.5125
-11.5 | 100 | 99.4 | 98.8 | 97.4 | 96.6375
-11.4 | 100.1 | 99.4 | 98.8333 | 97.4625 | 96.7714
-11.3 | 100.1833 | 99.4 | 98.9 | 97.5286 | 96.9
-11.2 | 100.2 | 99.5 | 98.9 | 97.6 | 96.9
-11.1 | 100.2875 | 99.5 | 99.0167 | 97.7 | 96.925
-11 | 100.3 | 99.6 | 99.1 | 97.7 | 97
-10.9 | 100.4 | 99.6571 | 99.12 | 97.8 | 97.0857
-10.8 | 100.4 | 99.7 | 99.2 | 97.9 | 97.1429
-10.7 | 100.4 | 99.8 | 99.3 | 97.9857 | 97.2857
-10.6 | 100.5 | 99.86 | 99.3 | 98.0286 | 97.3429
-10.5 | 100.5667 | 99.9 | 99.3167 | 98.1 | 97.4
-10.4 | 100.6 | 99.9667 | 99.4 | 98.2 | 97.4167
-10.3 | 100.6667 | 100 | 99.4 | 98.2 | 97.5714
-10.2 | 100.7 | 100 | 99.4333 | 98.2857 | 97.7
-10.1 | 100.7 | 100.1 | 99.5 | 98.4 | 97.7833
-10 | 100.8 | 100.1 | 99.5 | 98.4625 | 97.8
-9.9 | 100.8 | 100.1 | 99.6 | 98.5 | 97.8875
-9.8 | 100.8571 | 100.1333 | 99.6833 | 98.54 | 98.0125
-9.7 | 100.9 | 100.2 | 99.7 | 98.6143 | 98.1
-9.6 | 100.9 | 100.3 | 99.7 | 98.7 | 98.1714
-9.5 | 100.9 | 100.4 | 99.8 | 98.7 | 98.2
-9.4 | 100.9667 | 100.4 | 99.9 | 98.7167 | 98.2
-9.3 | 100.9833 | 100.4 | 99.9 | 98.8 | 98.2
-9.2 | 100.9833 | 100.4 | 99.94 | 98.9 | 98.3
-9.1 | 101 | 100.4 | 100 | 98.9857 | 98.3
-9 | 101 | 100.4429 | 100 | 99.0429 | 98.3375
-8.9 | 101.1 | 100.5 | 100.1 | 99.1 | 98.4
-8.8 | 101.1571 | 100.5333 | 100.1 | 99.1875 | 98.4143
-8.7 | 101.2 | 100.6 | 100.1 | 99.2 | 98.5
-8.6 | 101.2 | 100.6 | 100.1 | 99.2 | 98.6
-8.5 | 101.2 | 100.7 | 100.1 | 99.3 | 98.7
-8.4 | 101.2 | 100.7 | 100.2 | 99.3125 | 98.7143
-8.3 | 101.2667 | 100.7 | 100.2 | 99.4375 | 98.8
-8.2 | 101.3 | 100.8 | 100.3125 | 99.5 | 98.8
-8.1 | 101.3 | 100.8 | 100.4 | 99.5 | 98.9
-8 | 101.3 | 100.8 | 100.4 | 99.5143 | 98.96
-7.9 | 101.4 | 100.8143 | 100.4 | 99.6 | 99
-7.8 | 101.4 | 100.9 | 100.4667 | 99.6 | 99.0571
-7.7 | 101.4 | 100.9 | 100.4 | 99.6286 | 99.1
-7.6 | 101.4167 | 101 | 100.5 | 99.7 | 99.1375
-7.5 | 101.5 | 101 | 100.5 | 99.7 | 99.2
-7.4 | 101.5 | 101 | 100.5833 | 99.7 | 99.2
-7.3 | 101.5 | 101 | 100.6 | 99.8 | 99.3
-7.2 | 101.5 | 100.9 | 100.6 | 99.8 | 99.3
-7.1 | 101.5 | 101 | 100.6 | 99.82 | 99.4
-7 | 101.5 | 101 | 100.6 | 99.9 | 99.4
-6.9 | 101.5 | 101 | 100.6 | 99.9 | 99.4
-6.8 | 101.5 | 101.0857 | 100.7 | 99.9 | 99.4
-6.7 | 101.5 | 101.1 | 100.7 | 99.9 | 99.4
-6.6 | 101.6 | 101.0833 | 100.75 | 99.975 | 99.4667
-6.5 | 101.6 | 101.0833 | 100.8 | 100 | 99.5125
-6.4 | 101.6 | 101.1 | 100.8 | 100 | 99.6
-6.3 | 101.6 | 101.1 | 100.8 | 100 | 99.6
-6.2 | 101.6 | 101.1 | 100.8 | 100.1 | 99.65
-6.1 | 101.6 | 101.2 | 100.9 | 100.1 | 99.7
-6 | 101.6 | 101.2 | 100.9 | 100.15 | 99.7
-5.9 | 101.6 | 101.1 | 100.8 | 100.2 | 99.7
-5.8 | 101.6 | 101.1833 | 100.8 | 100.2 | 99.8
-5.7 | 101.6 | 101.2 | 100.85 | 100.2 | 99.7
-5.6 | 101.6 | 101.2 | 100.8 | 100.2 | 99.8
-5.5 | 101.5 | 101.1667 | 100.8 | 100.2 | 99.8
-5.4 | 101.5 | 101.2 | 100.8 | 100.2 | 99.8
-5.3 | 101.5 | 101.2 | 100.8 | 100.2 | 99.8
-5.2 | 101.5 | 101.18 | 100.8 | 100.2625 | 99.9
-5.1 | 101.5 | 101.1 | 100.8 | 100.2125 | 99.9
-5 | 101.5 | 101.1 | 100.8 | 100.24 | 99.9
-4.9 | 101.5 | 101.1 | 100.8 | 100.3 | 99.9
-4.8 | 101.5 | 101.1 | 100.7833 | 100.3 | 99.9
-4.7 | 101.5 | 101.1 | 100.7 | 100.3 | 99.9
-4.6 | 101.4 | 101.1 | 100.7 | 100.3 | 99.9
-4.5 | 101.4 | 101.0571 | 100.7 | 100.3 | 99.9714
-4.4 | 101.35 | 101 | 100.7 | 100.25 | 100
-4.3 | 101.2 | 101 | 100.7 | 100.3 | 100
-4.2 | 101.2 | 101 | 100.7 | 100.3 | 100
-4.1 | 101.2 | 101 | 100.7 | 100.3 | 100
-4 | 101.1125 | 100.95 | 100.7 | 100.3 | 100
-3.9 | 101.1 | 100.9 | 100.65 | 100.3 | 100
-3.8 | 101.1 | 100.8286 | 100.6 | 100.3 | 100
-3.7 | 101.0143 | 100.8 | 100.6 | 100.3 | 100
-3.6 | 101 | 100.8 | 100.6 | 100.3 | 100
-3.5 | 100.9 | 100.8 | 100.6 | 100.2 | 100
-3.4 | 100.9 | 100.7333 | 100.5 | 100.2 | 100
-3.3 | 100.9 | 100.7 | 100.5 | 100.2 | 100.0714
-3.2 | 100.9 | 100.78 | 100.5 | 100.2 | 100.0833
-3.1 | 100.8 | 100.7 | 100.5 | 100.2 | 100
-3 | 100.8 | 100.7 | 100.5 | 100.2 | 100
-2.9 | 100.8 | 100.6 | 100.5 | 100.2 | 100
-2.8 | 100.8 | 100.6 | 100.4143 | 100.2 | 100
-2.7 | 100.8 | 100.5125 | 100.4 | 100.2 | 100.1
-2.6 | 100.8 | 100.5 | 100.4 | 100.2 | 100.1
-2.5 | 100.8 | 100.5 | 100.4 | 100.2 | 100.1
-2.4 | 100.7 | 100.4 | 100.4 | 100.2375 | 100.1
-2.3 | 100.7 | 100.4 | 100.4 | 100.3 | 100.1
-2.2 | 100.6 | 100.4 | 100.3625 | 100.2 | 100.1
-2.1 | 100.6 | 100.4 | 100.3 | 100.1167 | 100.2
-2 | 100.5167 | 100.4 | 100.3 | 100.1 | 100.2
-1.9 | 100.5 | 100.3 | 100.3 | 100.1 | 100.2
-1.8 | 100.38 | 100.3 | 100.2667 | 100.1 | 100.2
-1.7 | 100.3 | 100.3 | 100.2 | 100.1 | 100.2
-1.6 | 100.3 | 100.2 | 100.2 | 100.1 | 100.1667
-1.5 | 100.3 | 100.1857 | 100.2 | 100.1 | 100.1
-1.4 | 100.3 | 100.1 | 100.2 | 100.06 | 100.1
-1.3 | 100.2286 | 100.0667 | 100.2 | 100 | 100.1
-1.2 | 100.2 | 100 | 100.1286 | 100 | 100.1
-1.1 | 100.1 | 100 | 100.1 | 100 | 100
-1 | 100.1 | 99.96 | 100.1 | 100 | 100
-0.9 | 100 | 99.9 | 100 | 100 | 100
-0.8 | 100 | 99.9 | 100 | 100 | 99.9
-0.7 | 99.9833 | 99.9 | 100 | 100 | 99.9
-0.6 | 99.9833 | 99.9 | 100 | 100 | 99.9
-0.5 | 100 | 99.9 | 100 | 100 | 99.9
-0.4 | 100 | 99.9 | 100 | 100 | 99.9222
-0.3 | 100 | 99.9 | 100 | 100 | 100
-0.2 | 100 | 99.9 | 100 | 100 | 100
-0.1 | 100 | 99.96 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 99.9 | 100 | 100 | 100
0.3 | 100 | 99.9 | 100 | 100 | 100
0.4 | 100 | 99.9 | 100 | 100 | 99.9
0.5 | 100 | 99.9 | 100 | 100 | 99.9
0.6 | 99.9833 | 99.9 | 100 | 100 | 99.9
0.7 | 99.9833 | 99.9 | 99.9167 | 100 | 99.9
0.8 | 100 | 99.9 | 100 | 100 | 99.9167
0.9 | 100 | 99.9 | 100 | 100 | 100
1 | 100.1 | 99.94 | 100.1 | 100 | 100
1.1 | 100.1 | 100 | 100.1 | 100 | 100
1.2 | 100.2 | 100 | 100.1143 | 100 | 100.1
1.3 | 100.2286 | 100.1 | 100.2 | 100 | 100.1
1.4 | 100.3 | 100.1 | 100.2 | 100 | 100.1
1.5 | 100.3 | 100.1667 | 100.2 | 100.1 | 100.1
1.6 | 100.3 | 100.2 | 100.2 | 100.1 | 100.2
1.7 | 100.3 | 100.3 | 100.2 | 100.1 | 100.2
1.8 | 100.38 | 100.3 | 100.2333 | 100.1 | 100.2
1.9 | 100.5 | 100.3 | 100.3 | 100.1 | 100.2
2 | 100.5167 | 100.4 | 100.3 | 100.1 | 100.2
2.1 | 100.6 | 100.4 | 100.3 | 100.1667 | 100.1714
2.2 | 100.6 | 100.4 | 100.3429 | 100.2 | 100.1
2.3 | 100.7 | 100.4 | 100.4 | 100.2857 | 100.1
2.4 | 100.7 | 100.4 | 100.4 | 100.2 | 100.1
2.5 | 100.8 | 100.5 | 100.4 | 100.2 | 100.1
2.6 | 100.8 | 100.5 | 100.4 | 100.2 | 100.1
2.7 | 100.8 | 100.5 | 100.4 | 100.2 | 100.0714
2.8 | 100.8 | 100.6 | 100.4 | 100.2 | 100
2.9 | 100.8 | 100.6 | 100.5 | 100.2 | 100
3 | 100.8 | 100.7 | 100.5 | 100.2 | 100.0167
3.1 | 100.8 | 100.7 | 100.5 | 100.2 | 100.0286
3.2 | 100.9 | 100.75 | 100.5 | 100.2 | 100.1
3.3 | 100.9 | 100.7 | 100.5 | 100.2 | 100.1
3.4 | 100.9 | 100.7167 | 100.5 | 100.2 | 100
3.5 | 100.9 | 100.8 | 100.6 | 100.2 | 100
3.6 | 101 | 100.8 | 100.6 | 100.2857 | 100
3.7 | 101.0143 | 100.8 | 100.6 | 100.3 | 100
3.8 | 101.1 | 100.9 | 100.6 | 100.3 | 100
3.9 | 101.1 | 100.9 | 100.6167 | 100.3 | 100
4 | 101.1125 | 101 | 100.7 | 100.3 | 100
4.1 | 101.2 | 101 | 100.7 | 100.3 | 100
4.2 | 101.2 | 101 | 100.7 | 100.3 | 100
4.3 | 101.2 | 101 | 100.7 | 100.3 | 100
4.4 | 101.35 | 101 | 100.7 | 100.2 | 100
4.5 | 101.4 | 101.0286 | 100.7 | 100.3 | 100
4.6 | 101.4 | 101.1 | 100.7 | 100.3 | 99.9
4.7 | 101.5 | 101.1 | 100.7 | 100.3 | 99.9
4.8 | 101.5 | 101.1 | 100.7667 | 100.3 | 99.9
4.9 | 101.5 | 101.1 | 100.8 | 100.3 | 99.9
5 | 101.5 | 101.1 | 100.8 | 100.2 | 99.9
5.1 | 101.5 | 101.1 | 100.8 | 100.2 | 99.9
5.2 | 101.5 | 101.16 | 100.8 | 100.2 | 99.9
5.3 | 101.5 | 101.2 | 100.84 | 100.2 | 99.8286
5.4 | 101.5 | 101.2 | 100.8 | 100.2 | 99.82
5.5 | 101.5 | 101.15 | 100.8 | 100.2 | 99.8
5.6 | 101.6 | 101.2 | 100.8 | 100.2 | 99.775
5.7 | 101.6 | 101.2 | 100.8714 | 100.2 | 99.7
5.8 | 101.6 | 101.2 | 100.8 | 100.2 | 99.7714
5.9 | 101.6 | 101.1333 | 100.8 | 100.2 | 99.7
6 | 101.6 | 101.2 | 100.9 | 100.2 | 99.7
6.1 | 101.6 | 101.2 | 100.9 | 100.1 | 99.7
6.2 | 101.6 | 101.1 | 100.8167 | 100.1 | 99.7
6.3 | 101.6 | 101.1 | 100.8 | 100 | 99.6
6.4 | 101.6 | 101.1 | 100.8 | 100 | 99.5875
6.5 | 101.6 | 101.1 | 100.8 | 100 | 99.5
6.6 | 101.6 | 101.06 | 100.775 | 99.9333 | 99.4167
6.7 | 101.5 | 101.1 | 100.7 | 99.9 | 99.4
6.8 | 101.5 | 101.1 | 100.7 | 99.9 | 99.4
6.9 | 101.5 | 101 | 100.6 | 99.9 | 99.4
7 | 101.5 | 101 | 100.6 | 99.9 | 99.4
7.1 | 101.5 | 100.9333 | 100.6 | 99.8667 | 99.3667
7.2 | 101.5 | 100.9 | 100.6 | 99.8 | 99.3
7.3 | 101.5 | 101 | 100.6 | 99.8 | 99.26
7.4 | 101.5 | 101 | 100.6 | 99.7 | 99.2
7.5 | 101.5 | 100.9833 | 100.5 | 99.6429 | 99.2
7.6 | 101.4167 | 100.9714 | 100.5 | 99.6875 | 99.1
7.7 | 101.4 | 100.9 | 100.4 | 99.6 | 99.1
7.8 | 101.4 | 100.9 | 100.5 | 99.6 | 99.1
7.9 | 101.4 | 100.8429 | 100.4 | 99.6 | 99
8 | 101.3 | 100.8 | 100.4 | 99.5714 | 99
8.1 | 101.3 | 100.8 | 100.4 | 99.5 | 98.85
8.2 | 101.3 | 100.8 | 100.3375 | 99.5 | 98.8
8.3 | 101.2667 | 100.7222 | 100.2 | 99.3875 | 98.8
8.4 | 101.2 | 100.7 | 100.2 | 99.3 | 98.7
8.5 | 101.2 | 100.7 | 100.1 | 99.3 | 98.7
8.6 | 101.2 | 100.6 | 100.1 | 99.24 | 98.55
8.7 | 101.2 | 100.6 | 100.1 | 99.2 | 98.5
8.8 | 101.1571 | 100.5 | 100.1 | 99.2 | 98.4
8.9 | 101.1 | 100.5 | 100.1 | 99.1222 | 98.4
9 | 101 | 100.4 | 100 | 99.1 | 98.3
9.1 | 101 | 100.4 | 100 | 99 | 98.3
9.2 | 100.9833 | 100.4 | 99.9 | 98.9 | 98.3
9.3 | 100.9833 | 100.4 | 99.9 | 98.8286 | 98.2143
9.4 | 100.9667 | 100.4 | 99.85 | 98.76 | 98.2
9.5 | 100.9 | 100.4 | 99.8 | 98.7 | 98.2
9.6 | 100.9 | 100.3143 | 99.7333 | 98.7 | 98.125
9.7 | 100.9 | 100.22 | 99.7 | 98.6571 | 98.0875
9.8 | 100.8571 | 100.1667 | 99.6125 | 98.5 | 97.9625
9.9 | 100.8 | 100.1 | 99.6 | 98.5 | 97.9
10 | 100.8 | 100.1 | 99.5 | 98.4125 | 97.8125
10.1 | 100.7 | 100.1 | 99.5 | 98.3714 | 97.8
10.2 | 100.7 | 100.025 | 99.46 | 98.3 | 97.6833
10.3 | 100.6667 | 100 | 99.4 | 98.2 | 97.5286
10.4 | 100.6 | 99.9833 | 99.4 | 98.15 | 97.4667
10.5 | 100.5667 | 99.9 | 99.34 | 98.1 | 97.4
10.6 | 100.5 | 99.8 | 99.3 | 98 | 97.4
10.7 | 100.4 | 99.8 | 99.2714 | 97.9333 | 97.3
10.8 | 100.4 | 99.7 | 99.2 | 97.9 | 97.2
10.9 | 100.4 | 99.6857 | 99.14 | 97.8143 | 97.1
11 | 100.3 | 99.6 | 99.1 | 97.7571 | 97
11.1 | 100.2875 | 99.5 | 99.0333 | 97.7 | 96.975
11.2 | 100.2 | 99.5 | 98.9 | 97.6333 | 96.9
11.3 | 100.1833 | 99.4 | 98.9 | 97.5 | 96.8833
11.4 | 100.1 | 99.4 | 98.8 | 97.5 | 96.7167
11.5 | 100 | 99.3333 | 98.8 | 97.4 | 96.6
11.6 | 99.95 | 99.3 | 98.8 | 97.3286 | 96.5667
11.7 | 99.9 | 99.3 | 98.7 | 97.3 | 96.45
11.8 | 99.9 | 99.2 | 98.6429 | 97.2125 | 96.4
11.9 | 99.9 | 99.1714 | 98.6 | 97.2 | 96.4
12 | 99.8667 | 99.1 | 98.5 | 97.15 | 96.3286
12.1 | 99.8 | 99.0833 | 98.4333 | 97.1 | 96.2
12.2 | 99.7429 | 99 | 98.38 | 97 | 96.1571
12.3 | 99.6833 | 98.9 | 98.3 | 96.9778 | 96.1
12.4 | 99.6 | 98.9 | 98.2 | 96.9 | 96
12.5 | 99.5 | 98.8167 | 98.1714 | 96.8 | 95.9125
12.6 | 99.4 | 98.8 | 98.1 | 96.7167 | 95.8
12.7 | 99.4 | 98.72 | 98.0625 | 96.5571 | 95.7286
12.8 | 99.3714 | 98.6571 | 97.9375 | 96.5 | 95.6857
12.9 | 99.3 | 98.6 | 97.9 | 96.4286 | 95.6
13 | 99.3 | 98.5 | 97.8 | 96.3875 | 95.46
13.1 | 99.2 | 98.4 | 97.7 | 96.3 | 95.38
13.2 | 99.2 | 98.3333 | 97.7 | 96.3 | 95.2
13.3 | 99.1 | 98.2667 | 97.6 | 96.2 | 95.15
13.4 | 99 | 98.2 | 97.6 | 96.15 | 95
13.5 | 99 | 98.16 | 97.5 | 96.1 | 95
13.6 | 98.9286 | 98.1 | 97.4571 | 95.9167 | 94.9167
13.7 | 98.9 | 98.0571 | 97.3 | 95.85 | 94.9
13.8 | 98.8 | 98 | 97.2375 | 95.8 | 94.8
13.9 | 98.7286 | 97.9 | 97.1833 | 95.6875 | 94.7143
14 | 98.7 | 97.9 | 97.1 | 95.6 | 94.6
14.1 | 98.6714 | 97.8 | 97 | 95.5143 | 94.5286
14.2 | 98.6 | 97.7429 | 96.9 | 95.4 | 94.4111
14.3 | 98.5 | 97.7 | 96.8429 | 95.3375 | 94.2875
14.4 | 98.4167 | 97.6286 | 96.8 | 95.2125 | 94.2
14.5 | 98.4 | 97.6 | 96.6714 | 95.0889 | 94.1286
14.6 | 98.3286 | 97.5286 | 96.6 | 95 | 94
14.7 | 98.2667 | 97.5 | 96.5 | 94.9429 | 93.9
14.8 | 98.2 | 97.4333 | 96.5 | 94.9 | 93.825
14.9 | 98.2 | 97.3667 | 96.4 | 94.7429 | 93.76
15 | 98.1857 | 97.28 | 96.3 | 94.7 | 93.6
15.1 | 98.1 | 97.1143 | 96.2 | 94.575 | 93.475
15.2 | 97.9857 | 97.08 | 96.1 | 94.45 | 93.35
15.3 | 97.9 | 97 | 96.1 | 94.325 | 93.3
15.4 | 97.8667 | 96.9 | 96.0333 | 94.2111 | 93.1667
15.5 | 97.8 | 96.8429 | 95.9 | 94.1 | 93
15.6 | 97.72 | 96.8 | 95.8667 | 94.0286 | 92.9625
15.7 | 97.7 | 96.7667 | 95.8 | 93.9714 | 92.8375
15.8 | 97.6 | 96.68 | 95.7 | 93.9 | 92.7125
15.9 | 97.6 | 96.6 | 95.65 | 93.8222 | 92.65
16 | 97.6 | 96.6 | 95.525 | 93.6833 | 92.5667
16.1 | 97.5 | 96.5 | 95.4 | 93.5375 | 92.3857
16.2 | 97.4286 | 96.4333 | 95.3429 | 93.4222 | 92.3
16.3 | 97.3857 | 96.3667 | 95.2714 | 93.3 | 92.1571
16.4 | 97.3 | 96.3 | 95.2 | 93.1571 | 92.025
16.5 | 97.2 | 96.3 | 95.0833 | 93.0143 | 91.9
16.6 | 97.2 | 96.3 | 95 | 92.9571 | 91.8
16.7 | 97.1571 | 96.2 | 94.9 | 92.9 | 91.6875
16.8 | 97.1 | 96.1625 | 94.8 | 92.8 | 91.5625
16.9 | 97.05 | 96.1 | 94.7 | 92.7625 | 91.4375
17 | 97 | 95.9875 | 94.6222 | 92.6167 | 91.4
17.1 | 96.9 | 95.8571 | 94.58 | 92.475 | 91.2
17.2 | 96.8 | 95.8 | 94.5 | 92.2625 | 90.9875
17.3 | 96.72 | 95.7 | 94.4 | 92.2 | 90.8625
17.4 | 96.7 | 95.5833 | 94.3 | 92.1 | 90.8
17.5 | 96.6 | 95.5 | 94.2286 | 91.9714 | 90.7143
17.6 | 96.5 | 95.4 | 94.2 | 91.8167 | 90.5667
17.7 | 96.4 | 95.35 | 94.0714 | 91.72 | 90.4
17.8 | 96.35 | 95.2875 | 93.9286 | 91.65 | 90.3429
17.9 | 96.2857 | 95.1625 | 93.8875 | 91.5125 | 90.2
18 | 96.1333 | 95.1 | 93.7625 | 91.3833 | 90.0333
18.1 | 95.9571 | 94.9667 | 93.575 | 91.2167 | 89.9
18.2 | 95.9 | 94.8143 | 93.4 | 91.0833 | 89.8429
18.3 | 95.7571 | 94.7714 | 93.3 | 91 | 89.8
18.4 | 95.5 | 94.5333 | 93.2 | 90.9125 | 89.64
18.5 | 95.4 | 94.5 | 93.1286 | 90.78 | 89.4714
18.6 | 95.24 | 94.25 | 92.9875 | 90.6 | 89.3375
18.7 | 95 | 94.1 | 92.85 | 90.4714 | 89.2125
18.8 | 94.9 | 94.05 | 92.7 | 90.3286 | 89.1
18.9 | 94.6667 | 93.8286 | 92.5333 | 90.2 | 88.95
19 | 94.3714 | 93.6143 | 92.4333 | 90.1 | 88.8
19.1 | 94.1125 | 93.4143 | 92.225 | 89.9833 | 88.7
19.2 | 93.7667 | 93.15 | 92.075 | 89.9 | 88.5333
19.3 | 93.4 | 92.96 | 91.9429 | 89.7714 | 88.4
19.4 | 92.9333 | 92.6286 | 91.6714 | 89.6 | 88.3143
19.5 | 92.32 | 92.2143 | 91.42 | 89.5 | 88.1833
19.6 | 91.4 | 91.9286 | 91.2667 | 89.2571 | 88
19.7 | 90.2 | 91.56 | 91.0375 | 89.0286 | 87.9
19.8 | 88.6 | 91.1 | 90.825 | 88.875 | 87.775
19.9 | 86.2 | 90.3857 | 90.54 | 88.6571 | 87.6
20 | 82.9333 | 89.6 | 90.2429 | 88.5143 | 87.35
20.1 | 78.1 | 88.8667 | 90.0286 | 88.4 | 87.2
20.2 | 72.0667 | 87.6 | 89.8 | 88.2625 | 87.075
20.3 | 64.4 | 85.94 | 89.4667 | 88.075 | 86.9556
20.4 | 56.1 | 83.66 | 89.2 | 87.9125 | 86.8375
20.5 | 47.9857 | 80.2 | 88.6714 | 87.6857 | 86.6833
20.6 | 39.5667 | 75.8 | 88.3 | 87.52 | 86.5167
20.7 | 33 | 69.8 | 87.84 | 87.38 | 86.3571
20.8 | 26.94 | 62.9 | 87.2333 | 87.1714 | 86.2
20.9 | 22.6 | 54.9 | 86.5333 | 86.9125 | 86
21 | 19.2333 | 47.2333 | 85.7714 | 86.625 | 85.85
21.1 | 16.9286 | 40.1 | 84.8333 | 86.4 | 85.7
21.2 | 15.4143 | 33.3 | 83.55 | 86.2 | 85.5286
21.3 | 14.3 | 27.875 | 81.8 | 85.9571 | 85.275
21.4 | 13.4667 | 23.6143 | 79.45 | 85.7 | 85.12
21.5 | 12.9 | 20.18 | 76.325 | 85.4556 | 85.05
21.6 | 12.4286 | 17.9167 | 72.1 | 85.225 | 84.8333
21.7 | 11.9857 | 16.3 | 66.9 | 84.9333 | 84.7
21.8 | 11.74 | 15.1714 | 60.6333 | 84.6429 | 84.4333
21.9 | 11.4333 | 14.3333 | 53.5 | 84.2857 | 84.2667
22 | 11.2 | 13.6 | 46.35 | 84.05 | 84.075
22.1 | 11.0143 | 13.0714 | 39.7667 | 83.6143 | 83.8625
22.2 | 10.8143 | 12.6 | 33.875 | 83.3286 | 83.675
22.3 | 10.64 | 12.24 | 28.9167 | 83.025 | 83.5125
22.4 | 10.46 | 11.9875 | 24.9875 | 82.55 | 83.275
22.5 | 10.38 | 11.7286 | 22.0875 | 82.1286 | 83.0143
22.6 | 10.1857 | 11.4667 | 20.0429 | 81.5714 | 82.75
22.7 | 10.0333 | 11.2167 | 18.45 | 81.175 | 82.6
22.8 | 9.9 | 11.06 | 17.35 | 80.4667 | 82.3571
22.9 | 9.82 | 10.8667 | 16.4857 | 79.825 | 82.075
23 | 9.6667 | 10.6286 | 15.78 | 79.025 | 81.8
23.1 | 9.6 | 10.5 | 15.2286 | 78.1 | 81.6778
23.2 | 9.4625 | 10.3333 | 14.6667 | 76.6167 | 81.4875
23.3 | 9.3286 | 10.1714 | 14.3571 | 74.95 | 81.175
23.4 | 9.15 | 10.0167 | 13.9286 | 72.46 | 80.925
23.5 | 9.08 | 9.95 | 13.625 | 69.3143 | 80.52
23.6 | 9 | 9.7857 | 13.3429 | 65.32 | 80.2429
23.7 | 8.9 | 9.6429 | 13.0571 | 60.9 | 79.975
23.8 | 8.9 | 9.5 | 12.9 | 55.8167 | 79.6375
23.9 | 8.8 | 9.3167 | 12.65 | 50.0429 | 79.3875
24 | 8.75 | 9.2571 | 12.4 | 44.3571 | 79.0333
24.1 | 8.6833 | 9.1143 | 12.2 | 39.0625 | 78.7286
24.2 | 8.6 | 9.0333 | 12.0143 | 34.2 | 78.25
24.3 | 8.5 | 9 | 11.8 | 30.22 | 77.95
24.4 | 8.48 | 8.8429 | 11.62 | 26.975 | 77.5333
24.5 | 8.4 | 8.7125 | 11.4429 | 24.4667 | 77.15
24.6 | 8.3143 | 8.6857 | 11.2167 | 22.7 | 76.6143
24.7 | 8.2714 | 8.6 | 11.0667 | 21.24 | 75.98
24.8 | 8.2 | 8.5 | 10.9143 | 20.1 | 75.4667
24.9 | 8.1167 | 8.4571 | 10.775 | 19.3 | 74.7
25 | 8.1 | 8.4 | 10.6429 | 18.6286 | 73.7
25.1 | 8 | 8.35 | 10.4667 | 18.1 | 72.6
25.2 | 8 | 8.2714 | 10.3 | 17.55 | 71.2143
25.3 | 7.9 | 8.2 | 10.1889 | 17.1143 | 69.4
25.4 | 7.8714 | 8.1 | 10.075 | 16.7333 | 67.15
25.5 |  | 8.0333 | 9.9333 | 16.45 | 64.4
25.6 |  | 8 | 9.8571 | 16.1714 | 60.75
25.7 |  | 7.9 | 9.7833 | 15.8667 | 56.7
25.8 |  | 7.8857 | 9.6286 | 15.7 | 52.2
25.9 |  | 7.8 | 9.4667 | 15.4 | 47.3
26 |  |  | 9.4 | 15.175 | 42.5333
26.1 |  |  | 9.3 | 14.925 | 38.3
26.2 |  |  | 9.2125 | 14.6571 | 34.3
26.3 |  |  | 9.1 | 14.4429 | 30.9125
26.4 |  |  | 9.0143 | 14.1889 | 27.9875
26.5 |  |  | 8.92 | 14.025 | 25.9667
26.6 |  |  | 8.8667 | 13.8 | 24.3
26.7 |  |  | 8.8 | 13.6571 | 22.95
26.8 |  |  | 8.6857 | 13.4286 | 22
26.9 |  |  |  | 13.24 | 21.075
27 |  |  |  | 13.0714 | 20.4714
27.1 |  |  |  | 12.8333 | 19.9
27.2 |  |  |  | 12.7 | 19.4857
27.3 |  |  |  | 12.4875 | 19.1
27.4 |  |  |  | 12.3625 | 18.7667
27.5 |  |  |  | 12.2375 | 18.5
27.6 |  |  |  | 12.025 | 18.125
27.7 |  |  |  | 11.8857 | 17.85
27.8 |  |  |  | 11.7429 | 17.575
27.9 |  |  |  | 11.5833 | 17.3222
28 |  |  |  | 11.4286 | 17.075
28.1 |  |  |  | 11.3 | 16.825
28.2 |  |  |  | 11.1667 | 16.5889
28.3 |  |  |  | 11.025 | 16.3286
28.4 |  |  |  | 10.9 | 16.1
28.5 |  |  |  | 10.8333 | 15.9143
28.6 |  |  |  | 10.7 | 15.725
28.7 |  |  |  | 10.6 | 15.5
28.8 |  |  |  | 10.5 | 15.35
28.9 |  |  |  | 10.4 | 15.1429
29 |  |  |  |  | 14.9714
29.1 |  |  |  |  | 14.8286
29.2 |  |  |  |  | 14.5833
29.3 |  |  |  |  | 14.4
29.4 |  |  |  |  | 14.3
29.5 |  |  |  |  | 14.1667
29.6 |  |  |  |  | 13.9167
29.7 |  |  |  |  | 13.775
29.8 |  |  |  |  | 13.6
29.9 |  |  |  |  | 13.425
30 |  |  |  |  | 13.3
30.1 |  |  |  |  | 13.075
30.2 |  |  |  |  | 12.9
30.3 |  |  |  |  | 12.725
30.4 |  |  |  |  | 12.6
30.5 |  |  |  |  | 12.5
30.6 |  |  |  |  | 12.425
30.7 |  |  |  |  | 12.3
30.8 |  |  |  |  | 12.175
30.9 |  |  |  |  | 12.0111

## Wedge 45 Output Factors
Output factors
Wedge 45 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20
| 3 | 0.878 | 0.891 | 0.899 | 0.907 | 0.914 | 0.919 | 0.923
| 4 | 0.893 | 0.91 | 0.918 | 0.929 | 0.937 | 0.945 | 0.95
| 5 | 0.902 | 0.92 | 0.931 | 0.943 | 0.953 | 0.962 | 0.968
| 7 | 0.913 | 0.934 | 0.947 | 0.964 | 0.976 | 0.989 | 0.997
| 10 | 0.923 | 0.947 | 0.963 | 0.983 | 1 | 1.017 | 1.028
| 15 | 0.932 | 0.959 | 0.976 | 1 | 1.023 | 1.045 | 1.061
| 20 | 0.938 | 0.965 | 0.984 | 1.011 | 1.036 | 1.062 | 1.083
| 30 | 0.945 | 0.975 | 0.995 | 1.025 | 1.054 | 1.087 | 1.112
| 40 | 0.949 | 0.979 | 1.001 | 1.033 | 1.064 | 1.101 | 1.128

## Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 27.2 | 33.8 | 39.8 | 49
0.1 | 31.3 | 38.1333 | 44.15 | 53.1333
0.2 | 38.5 | 44.7 | 50.925 | 58.9667
0.3 | 48.3 | 53.6 | 59.3 | 66.55
0.4 | 56.3 | 61.5 | 65.9 | 72.8
0.5 | 64 | 68.2 | 72.3667 | 78.2
0.6 | 69.6 | 73.5667 | 77.0333 | 82.2
0.7 | 74.9667 | 78.6 | 81.3333 | 85.7
0.8 | 79.0333 | 82.175 | 84.8 | 88.55
0.9 | 82.725 | 85.4667 | 87.9 | 91.1
1 | 85.95 | 88.1 | 90.3 | 93.1333
1.1 | 88.8 | 90.5333 | 92.32 | 94.7
1.2 | 90.7 | 92.5333 | 94.0667 | 95.9
1.3 | 92.9 | 94.15 | 95.3333 | 96.975
1.4 | 94.35 | 95.3667 | 96.5 | 98
1.5 | 95.7 | 96.66 | 97.5 | 98.6
1.6 | 96.8 | 97.35 | 98.325 | 99
1.7 | 97.5 | 98.275 | 98.72 | 99.4667
1.8 | 98.3 | 98.775 | 99.15 | 99.8
1.9 | 98.8 | 99.25 | 99.5333 | 99.9
2 | 99.15 | 99.675 | 99.8 | 100
2.1 | 99.475 | 99.85 | 99.9 | 100
2.2 | 99.725 | 100 | 100 | 100
2.3 | 99.975 | 100 | 100 | 99.9
2.4 | 100 | 100 | 99.9 | 99.8
2.5 | 100 | 99.9 | 99.825 | 99.7
2.6 | 99.9 | 99.7667 | 99.75 | 99.5
2.7 | 99.7 | 99.575 | 99.525 | 99.175
2.8 | 99.5 | 99.3667 | 99.3 | 99
2.9 | 99.3 | 99.175 | 99.05 | 98.675
3 | 99.0333 | 99.025 | 98.7 | 98.35
3.1 | 98.75 | 98.68 | 98.45 | 98.1667
3.2 | 98.45 | 98.45 | 98.25 | 97.8
3.3 | 98.05 | 98.2 | 97.9333 | 97.5333
3.4 | 97.75 | 97.75 | 97.5667 | 97.2333
3.5 | 97.375 | 97.45 | 97.2667 | 96.925
3.6 | 96.9 | 97.1 | 96.875 | 96.575
3.7 | 96.4667 | 96.7 | 96.625 | 96.2
3.8 | 96.05 | 96.42 | 96.2 | 95.85
3.9 | 95.7 | 96.15 | 95.9 | 95.5333
4 | 95.2 | 95.6 | 95.625 | 95.2
4.1 | 94.8 | 95.225 | 95.1 | 94.7
4.2 | 94.4333 | 94.85 | 94.9 | 94.475
4.3 | 94.175 | 94.4 | 94.4 | 94.15
4.4 | 93.65 | 94.1 | 93.9 | 93.7
4.5 | 93.1 | 93.75 | 93.65 | 93.28
4.6 | 92.7 | 93.2667 | 93.375 | 93
4.7 | 92.3 | 92.9 | 92.9667 | 92.75
4.8 | 91.8 | 92.5 | 92.6 | 92.44
4.9 | 91.4667 | 92.2 | 92.1 | 92.025
5 | 91.225 | 91.74 | 91.74 | 91.65
5.1 | 90.85 | 91.375 | 91.38 | 91.3
5.2 | 90.28 | 91.05 | 91.05 | 91.02
5.3 | 89.8 | 90.78 | 90.8 | 90.55
5.4 | 89.3 | 90.5 | 90.5 | 90.275
5.5 | 88.9333 | 90 | 90.1333 | 89.95
5.6 | 88.4667 | 89.6667 | 89.675 | 89.6
5.7 | 88.0667 | 89.2 | 89.4 | 89.26
5.8 | 87.5667 | 88.84 | 89.04 | 88.95
5.9 | 87.275 | 88.45 | 88.7 | 88.62
6 | 86.775 | 88.125 | 88.2 | 88.275
6.1 | 86.2667 | 87.8 | 88 | 87.95
6.2 | 86 | 87.4 | 87.72 | 87.6
6.3 | 85.65 | 87.1333 | 87.425 | 87.25
6.4 | 85.2 | 86.7 | 87.1 | 86.875
6.5 | 84.7 | 86.3 | 86.7 | 86.625
6.6 | 84.22 | 85.975 | 86.35 | 86.2
6.7 | 83.85 | 85.575 | 86.1 | 85.8667
6.8 | 83.45 | 85.2 | 85.62 | 85.525
6.9 | 83.0667 | 84.9 | 85.3 | 85.2
7 | 82.75 | 84.5667 | 84.9333 | 84.7667
7.1 | 82.425 | 84.1667 | 84.55 | 84.4
7.2 | 82 | 83.74 | 84.225 | 84.1667
7.3 | 81.4333 | 83.4 | 83.8667 | 83.925
7.4 | 81 | 83.1333 | 83.5 | 83.675
7.5 | 80.7 | 82.8667 | 83.25 | 83.42
7.6 | 80.375 | 82.4333 | 82.9 | 82.96
7.7 | 80.025 | 82.2 | 82.65 | 82.7
7.8 | 79.55 | 81.75 | 82.28 | 82.3
7.9 | 79.16 | 81.25 | 81.825 | 82
8 | 78.8 | 80.98 | 81.5 | 81.6333
8.1 | 78.5 | 80.6 | 81.25 | 81.25
8.2 | 77.9 | 80.3 | 80.9667 | 81
8.3 | 77.5 | 80 | 80.7 | 80.775
8.4 | 77.3 | 79.675 | 80.3 | 80.4
8.5 | 76.95 | 79.3 | 79.9 | 80.025
8.6 | 76.5 | 78.94 | 79.62 | 79.7
8.7 | 76.12 | 78.6 | 79.325 | 79.5
8.8 | 75.725 | 78.1333 | 78.96 | 79.15
8.9 | 75.25 | 77.9333 | 78.6 | 78.75
9 | 74.825 | 77.525 | 78.35 | 78.425
9.1 | 74.6 | 77.16 | 78 | 77.975
9.2 | 74.25 | 76.75 | 77.6 | 77.725
9.3 | 73.925 | 76.5 | 77.3 | 77.48
9.4 | 73.5 | 76.15 | 77 | 77.2
9.5 | 73.1333 | 75.8 | 76.775 | 76.9
9.6 | 72.9 | 75.4667 | 76.425 | 76.65
9.7 | 72.3 | 75.1 | 76.05 | 76.325
9.8 | 71.9667 | 74.7333 | 75.65 | 76.025
9.9 | 71.725 | 74.46 | 75.4 | 75.675
10 | 71.35 | 74.2 | 75.15 | 75.25
10.1 | 71.0333 | 74 | 74.9 | 75
10.2 | 70.6667 | 73.65 | 74.3667 | 74.675
10.3 | 70.3 | 73.2 | 74.1 | 74.425
10.4 | 70.05 | 72.775 | 73.9 | 74.2
10.5 | 69.65 | 72.5333 | 73.5667 | 73.85
10.6 | 69.475 | 72.375 | 73.2 | 73.5667
10.7 | 69.075 | 72.025 | 72.95 | 73.35
10.8 | 68.675 | 71.7 | 72.6667 | 73
10.9 | 68.35 | 71.4 | 72.42 | 72.7
11 | 68 | 71 | 72.075 | 72.425
11.1 | 67.6 | 70.7 | 71.75 | 72.1
11.2 | 67.3333 | 70.4333 | 71.5 | 71.9
11.3 | 67 | 70 | 71.1333 | 71.5
11.4 | 66.75 | 69.76 | 70.95 | 71.2
11.5 | 66.35 | 69.45 | 70.6 | 70.95
11.6 | 66.0667 | 69.15 | 70.2 | 70.7
11.7 | 65.68 | 68.78 | 69.9 | 70.475
11.8 | 65.35 | 68.6 | 69.7 | 70.15
11.9 | 64.8667 | 68.3 | 69.55 | 69.8
12 | 64.7 | 68 | 69.225 | 69.5333
12.1 | 64.35 | 67.8 | 68.88 | 69.3
12.2 | 64 | 67.4 | 68.65 | 68.9667
12.3 | 63.7 | 67.075 | 68.3667 | 68.8
12.4 | 63.4667 | 66.825 | 68 | 68.4
12.5 | 63.1 | 66.46 | 67.8 | 68.125
12.6 | 62.825 | 66.2 | 67.45 | 67.8667
12.7 | 62.5 | 65.8667 | 67.2 | 67.6
12.8 | 62.3 | 65.6 | 66.875 | 67.35
12.9 | 61.94 | 65.3 | 66.5 | 67.1
13 | 61.58 | 65.1 | 66.2 | 66.7667
13.1 | 61.3 | 64.6333 | 66 | 66.6
13.2 | 60.8667 | 64.375 | 65.7667 | 66.2
13.3 | 60.6 | 64.1 | 65.3667 | 65.9667
13.4 | 60.3667 | 63.95 | 65.2 | 65.725
13.5 | 60.1667 | 63.6 | 64.9 | 65.45
13.6 | 59.925 | 63.26 | 64.65 | 65.1
13.7 | 59.5333 | 62.95 | 64.4 | 64.85
13.8 | 59.3 | 62.675 | 64.125 | 64.7
13.9 | 59 | 62.425 | 63.875 | 64.35
14 | 58.6 | 62.075 | 63.55 | 64.1
14.1 | 58.25 | 61.825 | 63.3 | 63.86
14.2 | 58 | 61.55 | 63.05 | 63.6
14.3 | 57.675 | 61.35 | 62.8 | 63.3
14.4 | 57.325 | 61 | 62.5 | 63.025
14.5 | 57.125 | 60.8 | 62.2 | 62.9
14.6 | 56.95 | 60.5667 | 61.8 | 62.5
14.7 | 56.7 | 60.2 | 61.6333 | 62.3
14.8 | 56.3667 | 59.925 | 61.4 | 62.05
14.9 | 56.1 | 59.6 | 61.1 | 61.72
15 | 55.8 | 59.325 | 60.8 | 61.45
15.1 | 55.5 | 59.1 | 60.56 | 61.275
15.2 | 55.175 | 58.85 | 60.25 | 61.05
15.3 | 55 | 58.6 | 60 | 60.65
15.4 | 54.8 | 58.4 | 59.875 | 60.5
15.5 | 54.5 | 58.1 | 59.625 | 60.3
15.6 | 54.2 | 57.7667 | 59.4 | 60.06
15.7 | 53.875 | 57.5667 | 59.1 | 59.825
15.8 | 53.55 | 57.275 | 58.75 | 59.5667
15.9 | 53.375 | 57.125 | 58.625 | 59.3
16 | 53.125 | 56.9 | 58.275 | 59.1
16.1 | 52.8 | 56.65 | 58.0333 | 58.8667
16.2 | 52.65 | 56.4 | 57.8 | 58.625
16.3 | 52.4 | 56.075 | 57.65 | 58.3333
16.4 | 52.1333 | 55.825 | 57.4 | 57.975
16.5 | 51.875 | 55.6 | 57.25 | 57.8
16.6 | 51.6 | 55.4 | 56.8 | 57.575
16.7 | 51.375 | 55 | 56.6 | 57.275
16.8 | 51.05 | 54.75 | 56.3667 | 57.125
16.9 | 50.88 | 54.5 | 56.1 | 56.8667
17 | 50.65 | 54.25 | 55.825 | 56.6
17.1 | 50.3667 | 54 | 55.7 | 56.4
17.2 | 50.2 | 53.8 | 55.5 | 56.2
17.3 | 49.875 | 53.5667 | 55.2 | 56
17.4 | 49.6 | 53.325 | 55 | 55.7333
17.5 | 49.3 | 53.1 | 54.75 | 55.5
17.6 | 49.08 | 52.85 | 54.5 | 55.3
17.7 | 48.8 | 52.7 | 54.25 | 55.08
17.8 | 48.65 | 52.45 | 54 | 54.9333
17.9 | 48.4 | 52.2 | 53.8333 | 54.6333
18 | 48.15 | 51.95 | 53.6333 | 54.3333
18.1 | 47.9 | 51.7 | 53.36 | 54.1333
18.2 | 47.7 | 51.4667 | 53.1 | 53.975
18.3 | 47.4667 | 51.2 | 52.9667 | 53.675
18.4 | 47.225 | 50.96 | 52.7 | 53.525
18.5 | 47 | 50.7 | 52.425 | 53.2
18.6 | 46.8 | 50.475 | 52.2 | 53
18.7 | 46.6 | 50.225 | 52.05 | 52.8
18.8 | 46.325 | 49.975 | 51.7667 | 52.6
18.9 | 46 | 49.725 | 51.5 | 52.36
19 | 45.675 | 49.6 | 51.25 | 52.125
19.1 | 45.575 | 49.4 | 51 | 51.875
19.2 | 45.325 | 49.2 | 50.8 | 51.625
19.3 | 45.18 | 48.85 | 50.65 | 51.48
19.4 | 44.95 | 48.6 | 50.4 | 51.25
19.5 | 44.7 | 48.4 | 50.2 | 51
19.6 | 44.5 | 48.2333 | 49.9 | 50.825
19.7 | 44.3 | 48.025 | 49.8 | 50.675
19.8 | 44.1 | 47.725 | 49.56 | 50.425
19.9 | 43.9 | 47.5 | 49.4 | 50.25
20 | 43.65 | 47.3 | 49.1667 | 49.95
20.1 | 43.4 | 47.1 | 48.925 | 49.7667
20.2 | 43.2 | 46.9333 | 48.725 | 49.6
20.3 | 43 | 46.7667 | 48.55 | 49.4
20.4 | 42.9 | 46.4 | 48.4 | 49.1
20.5 | 42.625 | 46.28 | 48.05 | 49
20.6 | 42.4333 | 46.1 | 47.9 | 48.85
20.7 | 42.2333 | 45.9 | 47.66 | 48.6
20.8 | 42.1 | 45.675 | 47.525 | 48.36
20.9 | 41.8 | 45.475 | 47.275 | 48.125
21 | 41.6 | 45.275 | 47.1 | 47.9667
21.1 | 41.5 | 45.0333 | 46.9 | 47.74
21.2 | 41.3 | 44.85 | 46.7 | 47.6
21.3 | 41 | 44.625 | 46.5 | 47.45
21.4 | 40.8 | 44.475 | 46.3 | 47.14
21.5 | 40.6 | 44.3 | 46.2 | 47
21.6 | 40.4333 | 44.15 | 45.975 | 46.75
21.7 | 40.2 | 43.875 | 45.625 | 46.625
21.8 | 40 | 43.6333 | 45.475 | 46.3667
21.9 | 39.82 | 43.4 | 45.24 | 46.2
22 | 39.675 | 43.225 | 45.05 | 46.0333
22.1 | 39.5667 | 43.075 | 44.9333 | 45.875
22.2 | 39.275 | 42.825 | 44.7 | 45.6
22.3 | 39.0667 | 42.7 | 44.5 | 45.4
22.4 | 38.8333 | 42.55 | 44.3 | 45.2
22.5 | 38.7 | 42.4 | 44.1333 | 45.1
22.6 | 38.5 | 42.1 | 44 | 44.95
22.7 | 38.375 | 41.925 | 43.75 | 44.725
22.8 | 38.125 | 41.7 | 43.5333 | 44.6
22.9 | 37.9 | 41.6 | 43.3 | 44.4
23 | 37.8 | 41.375 | 43.2 | 44.2
23.1 | 37.7 | 41.2 | 43.1 | 44
23.2 | 37.4 | 41 | 42.9 | 43.76
23.3 | 37.26 | 40.8667 | 42.6 | 43.6
23.4 | 37.125 | 40.7 | 42.5 | 43.375
23.5 | 36.8667 | 40.4333 | 42.2333 | 43.225
23.6 | 36.75 | 40.275 | 42.1 | 43.0667
23.7 | 36.6 | 40.025 | 41.975 | 42.875
23.8 | 36.4 | 39.875 | 41.725 | 42.7
23.9 | 36.3 | 39.75 | 41.6 | 42.525
24 | 36 | 39.6 | 41.4 | 42.3667
24.1 | 35.8 | 39.4 | 41.25 | 42.125
24.2 | 35.6 | 39.25 | 41.0333 | 41.9
24.3 | 35.4667 | 39.12 | 40.9 | 41.7333
24.4 | 35.3 | 38.875 | 40.6 | 41.6
24.5 | 35.1 | 38.6 | 40.525 | 41.45
24.6 | 34.8333 | 38.425 | 40.325 | 41.3
24.7 | 34.7333 | 38.3 | 40.1 | 41.2
24.8 | 34.6 | 38.15 | 39.975 | 40.975
24.9 | 34.475 | 38 | 39.8 | 40.8
25 | 34.3 | 37.8 | 39.6667 | 40.575
25.1 | 34.1333 | 37.7 | 39.5 | 40.425
25.2 | 34 | 37.525 | 39.325 | 40.275
25.3 | 33.8 | 37.375 | 39.08 | 40.1
25.4 | 33.7 | 37.2 | 38.95 | 39.95
25.5 | 33.5 | 36.9667 | 38.8 | 39.8
25.6 | 33.3 | 36.8333 | 38.6 | 39.675
25.7 | 33.25 | 36.65 | 38.45 | 39.425
25.8 | 33.0333 | 36.5 | 38.3 | 39.3
25.9 | 32.9 | 36.3 | 38.15 | 39.175
26 | 32.775 | 36.2 | 37.9 | 39
26.1 | 32.6 | 36 | 37.725 | 38.8
26.2 | 32.5 | 35.8 | 37.6 | 38.65
26.3 | 32.3 | 35.7 | 37.5 | 38.5
26.4 | 32.1 | 35.575 | 37.3 | 38.25
26.5 | 32 | 35.3 | 37.1667 | 38.15
26.6 | 31.9 | 35.1 | 37 | 38
26.7 | 31.7 | 35 | 36.8 | 37.8667
26.8 | 31.5 | 34.8333 | 36.7 | 37.675
26.9 | 31.3 | 34.725 | 36.475 | 37.45
27 | 31.2 | 34.6 | 36.3 | 37.3
27.1 | 31.1 | 34.4 | 36.15 | 37.2
27.2 | 30.9 | 34.2 | 36.075 | 37.1
27.3 | 30.825 | 34.1 | 35.9 | 36.9
27.4 | 30.675 | 33.9333 | 35.7 | 36.8
27.5 | 30.44 | 33.75 | 35.65 | 36.6
27.6 | 30.3 | 33.6 | 35.4 | 36.45
27.7 | 30.2 | 33.4 | 35.25 | 36.3333
27.8 | 30.05 | 33.3 | 35.1 | 36.2
27.9 | 29.9 | 33.125 | 35.025 | 35.975
28 | 29.725 | 33 | 34.875 | 35.8
28.1 | 29.6 | 32.8667 | 34.7 | 35.675
28.2 | 29.5 | 32.8 | 34.5 | 35.525
28.3 | 29.3 | 32.6 | 34.4 | 35.4
28.4 | 29.2 | 32.5 | 34.225 | 35.3
28.5 | 29.1 | 32.3 | 34.075 | 35.15
28.6 | 28.975 | 32.1 | 33.94 | 35
28.7 | 28.8 | 32 | 33.8 | 34.875
28.8 | 28.675 | 31.875 | 33.7 | 34.675
28.9 | 28.5 | 31.7 | 33.5 | 34.525
29 | 28.375 | 31.6 | 33.4 | 34.4
29.1 | 28.25 | 31.5 | 33.225 | 34.3
29.2 | 28.1 | 31.3 | 33.075 | 34.0333
29.3 | 28 | 31.2 | 33 | 33.95
29.4 | 27.85 | 31 | 32.8333 | 33.75
29.5 | 27.7 | 30.9 | 32.6 | 33.7
29.6 | 27.56 | 30.775 | 32.5 | 33.525
29.7 | 27.5 | 30.6 | 32.4 | 33.35
29.8 | 27.375 | 30.5 | 32.275 | 33.2
29.9 | 27.2 | 30.3 | 32.125 | 33.1
30 | 27.1333 | 30.2 | 32 | 33

## SSD80 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 27.7 | 36.4 | 45 | 57.8
0.1 | 31.625 | 40.35 | 49.26 | 61.64
0.2 | 38.475 | 47.3667 | 55.65 | 67.1
0.3 | 48.15 | 55.9571 | 63.65 | 73.775
0.4 | 56.9 | 63.675 | 71 | 79.6
0.5 | 64.4 | 70.6 | 76.225 | 83.96
0.6 | 70.7 | 75.9 | 81.15 | 87.74
0.7 | 75.875 | 80.675 | 84.975 | 90.4
0.8 | 80.2 | 84.25 | 88.3 | 92.925
0.9 | 83.8 | 87.5 | 90.8 | 94.75
1 | 87.025 | 90.225 | 92.9 | 96.1
1.1 | 89.7 | 92.125 | 94.6 | 97.25
1.2 | 91.8 | 93.9 | 96.1 | 98.2333
1.3 | 93.6 | 95.6 | 97.12 | 98.8
1.4 | 95.1667 | 96.54 | 98 | 99.3444
1.5 | 96.5 | 97.6 | 98.72 | 99.625
1.6 | 97.5 | 98.4 | 99.35 | 99.8375
1.7 | 98.4667 | 99.025 | 99.74 | 99.9625
1.8 | 98.8 | 99.525 | 99.875 | 100
1.9 | 99.3571 | 99.775 | 100 | 100
2 | 99.7 | 99.9 | 100 | 99.9625
2.1 | 99.9 | 100 | 100 | 99.775
2.2 | 99.9375 | 100 | 99.9 | 99.6222
2.3 | 100 | 100 | 99.85 | 99.2889
2.4 | 100 | 99.9 | 99.6 | 99.125
2.5 | 99.8 | 99.7 | 99.4 | 98.8
2.6 | 99.7 | 99.5333 | 99.1 | 98.5
2.7 | 99.5 | 99.175 | 98.7714 | 98.16
2.8 | 99.2 | 98.95 | 98.52 | 97.82
2.9 | 98.96 | 98.6 | 98.1571 | 97.425
3 | 98.5 | 98.275 | 97.775 | 96.975
3.1 | 98.125 | 97.95 | 97.45 | 96.625
3.2 | 97.65 | 97.625 | 97.06 | 96.3
3.3 | 97.375 | 97.175 | 96.6 | 95.85
3.4 | 96.95 | 96.725 | 96.2 | 95.5
3.5 | 96.5 | 96.36 | 95.85 | 95.0857
3.6 | 96.1 | 96 | 95.5 | 94.75
3.7 | 95.5 | 95.56 | 95 | 94.2875
3.8 | 95.075 | 95 | 94.64 | 93.9125
3.9 | 94.5667 | 94.7 | 94.22 | 93.45
4 | 94.15 | 94.25 | 93.75 | 93.0625
4.1 | 93.6333 | 93.8 | 93.4333 | 92.6875
4.2 | 93.15 | 93.4 | 92.9 | 92.3125
4.3 | 92.6875 | 92.95 | 92.5333 | 91.825
4.4 | 92.25 | 92.5 | 92.125 | 91.4286
4.5 | 91.75 | 92.1 | 91.7333 | 90.92
4.6 | 91.32 | 91.58 | 91.25 | 90.65
4.7 | 90.9 | 91.15 | 90.8 | 90.3
4.8 | 90.425 | 90.775 | 90.44 | 89.8333
4.9 | 89.75 | 90.44 | 90.0667 | 89.4
5 | 89.3429 | 90 | 89.65 | 89
5.1 | 88.9 | 89.6 | 89.3 | 88.65
5.2 | 88.4 | 89.1 | 88.9 | 88.24
5.3 | 87.8 | 88.6 | 88.3 | 87.85
5.4 | 87.45 | 88.2 | 88 | 87.425
5.5 | 87 | 87.7 | 87.5 | 86.975
5.6 | 86.55 | 87.35 | 87.25 | 86.64
5.7 | 86.1 | 86.95 | 86.75 | 86.1667
5.8 | 85.6 | 86.45 | 86.275 | 85.8
5.9 | 85.2 | 86.05 | 85.925 | 85.45
6 | 84.74 | 85.7 | 85.55 | 85.1857
6.1 | 84.25 | 85.3333 | 85.1 | 84.8625
6.2 | 83.725 | 84.9667 | 84.7333 | 84.4125
6.3 | 83.35 | 84.45 | 84.4 | 83.85
6.4 | 83.025 | 84.02 | 84.05 | 83.4375
6.5 | 82.6 | 83.6 | 83.65 | 83.14
6.6 | 82.1 | 83.3 | 83.25 | 82.8
6.7 | 81.7 | 82.9 | 82.825 | 82.45
6.8 | 81.15 | 82.55 | 82.525 | 82.05
6.9 | 80.725 | 82.1 | 82.025 | 81.675
7 | 80.2875 | 81.75 | 81.66 | 81.325
7.1 | 79.85 | 81.225 | 81.3 | 80.875
7.2 | 79.35 | 80.85 | 80.9 | 80.55
7.3 | 78.8778 | 80.425 | 80.5 | 80.18
7.4 | 78.625 | 79.975 | 80.1333 | 79.9333
7.5 | 78.2 | 79.65 | 79.8 | 79.45
7.6 | 77.775 | 79.4 | 79.4 | 79.1
7.7 | 77.4 | 79 | 79.05 | 78.775
7.8 | 77.05 | 78.6 | 78.74 | 78.45
7.9 | 76.7 | 78.1667 | 78.35 | 78.0333
8 | 76.2625 | 77.8 | 78 | 77.7
8.1 | 75.85 | 77.35 | 77.675 | 77.375
8.2 | 75.36 | 77.02 | 77.25 | 77.025
8.3 | 74.9 | 76.55 | 76.925 | 76.65
8.4 | 74.6 | 76.225 | 76.55 | 76.28
8.5 | 74.2 | 75.85 | 76.225 | 75.9
8.6 | 73.84 | 75.48 | 75.85 | 75.65
8.7 | 73.475 | 75.1 | 75.425 | 75.2
8.8 | 73.05 | 74.76 | 75.1 | 74.8667
8.9 | 72.575 | 74.375 | 74.7 | 74.6
9 | 72.25 | 74.05 | 74.4 | 74.35
9.1 | 71.8 | 73.675 | 74.05 | 73.85
9.2 | 71.55 | 73.38 | 73.7 | 73.5667
9.3 | 71.1 | 73.025 | 73.42 | 73.2
9.4 | 70.7 | 72.55 | 73 | 72.85
9.5 | 70.2 | 72.3 | 72.675 | 72.6
9.6 | 69.775 | 71.8 | 72.25 | 72.25
9.7 | 69.6 | 71.5 | 71.925 | 71.84
9.8 | 69.05 | 71.225 | 71.675 | 71.575
9.9 | 68.65 | 70.675 | 71.3375 | 71.25
10 | 68.3333 | 70.325 | 70.9667 | 70.85
10.1 | 67.9625 | 70.075 | 70.6333 | 70.54
10.2 | 67.64 | 69.75 | 70.1667 | 70.2
10.3 | 67.3 | 69.375 | 69.9333 | 69.95
10.4 | 66.925 | 68.95 | 69.6333 | 69.6
10.5 | 66.6 | 68.65 | 69.2571 | 69.2
10.6 | 66.125 | 68.2333 | 68.8857 | 68.85
10.7 | 65.7 | 67.9 | 68.5667 | 68.54
10.8 | 65.3429 | 67.65 | 68.2333 | 68.22
10.9 | 65.1167 | 67.325 | 67.95 | 67.85
11 | 64.6667 | 67 | 67.625 | 67.575
11.1 | 64.275 | 66.65 | 67.2625 | 67.35
11.2 | 64.1 | 66.26 | 66.94 | 67
11.3 | 63.7 | 66 | 66.6667 | 66.65
11.4 | 63.35 | 65.6 | 66.3 | 66.4
11.5 | 63 | 65.2333 | 66 | 66.05
11.6 | 62.72 | 65 | 65.7 | 65.8
11.7 | 62.2 | 64.725 | 65.45 | 65.45
11.8 | 61.95 | 64.275 | 65.08 | 65.2
11.9 | 61.6333 | 64.025 | 64.75 | 64.95
12 | 61.375 | 63.675 | 64.4 | 64.68
12.1 | 61.0333 | 63.325 | 64.15 | 64.25
12.2 | 60.6333 | 63.075 | 63.9 | 63.975
12.3 | 60.25 | 62.8 | 63.52 | 63.725
12.4 | 60 | 62.4 | 63.3 | 63.425
12.5 | 59.7 | 62.15 | 62.975 | 63.1
12.6 | 59.45 | 61.8 | 62.625 | 62.875
12.7 | 59.2 | 61.5 | 62.35 | 62.55
12.8 | 58.8 | 61.2 | 62.025 | 62.325
12.9 | 58.4 | 60.9 | 61.85 | 61.98
13 | 58.175 | 60.55 | 61.54 | 61.75
13.1 | 57.8 | 60.3 | 61.16 | 61.4667
13.2 | 57.475 | 59.95 | 60.85 | 61.125
13.3 | 57.125 | 59.6667 | 60.6 | 60.85
13.4 | 56.9667 | 59.4 | 60.3 | 60.575
13.5 | 56.65 | 59.175 | 60.02 | 60.25
13.6 | 56.325 | 58.75 | 59.75 | 60
13.7 | 56.075 | 58.55 | 59.5 | 59.65
13.8 | 55.84 | 58.225 | 59.175 | 59.4667
13.9 | 55.5 | 57.98 | 58.9 | 59.2
14 | 55.1333 | 57.68 | 58.575 | 58.92
14.1 | 54.85 | 57.3 | 58.3 | 58.625
14.2 | 54.5 | 57.05 | 58 | 58.2875
14.3 | 54.275 | 56.7 | 57.7 | 58.125
14.4 | 54.025 | 56.45 | 57.42 | 57.8125
14.5 | 53.675 | 56.2 | 57.175 | 57.525
14.6 | 53.4 | 55.96 | 56.925 | 57.275
14.7 | 53.1 | 55.6 | 56.675 | 57
14.8 | 52.75 | 55.36 | 56.325 | 56.7
14.9 | 52.5 | 55.05 | 56.1 | 56.4286
15 | 52.1 | 54.775 | 55.8333 | 56.15
15.1 | 51.9 | 54.425 | 55.5 | 55.925
15.2 | 51.6667 | 54.275 | 55.3 | 55.6667
15.3 | 51.325 | 54 | 55 | 55.375
15.4 | 51 | 53.7 | 54.7667 | 55.125
15.5 | 50.75 | 53.45 | 54.525 | 54.875
15.6 | 50.5 | 53.2 | 54.1875 | 54.625
15.7 | 50.3 | 52.85 | 54.025 | 54.3889
15.8 | 50.05 | 52.7 | 53.75 | 54.15
15.9 | 49.75 | 52.35 | 53.425 | 53.9
16 | 49.4 | 52.1 | 53.175 | 53.55
16.1 | 49.2167 | 51.9 | 53 | 53.35
16.2 | 49.075 | 51.625 | 52.7 | 53.125
16.3 | 48.725 | 51.375 | 52.45 | 52.875
16.4 | 48.475 | 51.125 | 52.25 | 52.625
16.5 | 48.175 | 50.875 | 52 | 52.375
16.6 | 47.925 | 50.55 | 51.7 | 52.125
16.7 | 47.675 | 50.3 | 51.475 | 51.88
16.8 | 47.4444 | 50.075 | 51.1 | 51.625
16.9 | 47.2 | 49.725 | 50.9714 | 51.4
17 | 46.94 | 49.6 | 50.6833 | 51.1667
17.1 | 46.7 | 49.3 | 50.475 | 51
17.2 | 46.45 | 49.025 | 50.225 | 50.7
17.3 | 46.2 | 48.78 | 49.975 | 50.475
17.4 | 46 | 48.6 | 49.725 | 50.2
17.5 | 45.7 | 48.3889 | 49.5333 | 50
17.6 | 45.5 | 48.1 | 49.3 | 49.825
17.7 | 45.2333 | 47.8125 | 49.1 | 49.58
17.8 | 45 | 47.6222 | 48.9 | 49.35
17.9 | 44.75 | 47.4 | 48.65 | 49.1
18 | 44.6 | 47.175 | 48.4 | 48.95
18.1 | 44.35 | 47 | 48.175 | 48.7
18.2 | 44.1 | 46.7 | 48 | 48.35
18.3 | 43.95 | 46.4 | 47.725 | 48.2
18.4 | 43.6 | 46.1778 | 47.375 | 47.96
18.5 | 43.35 | 46.025 | 47.1667 | 47.725
18.6 | 43.1 | 45.7333 | 47 | 47.575
18.7 | 43 | 45.5 | 46.76 | 47.325
18.8 | 42.7 | 45.3 | 46.625 | 47.0625
18.9 | 42.5 | 45.1 | 46.35 | 46.875
19 | 42.3 | 44.85 | 46.125 | 46.6444
19.1 | 42.075 | 44.6 | 45.95 | 46.4
19.2 | 41.9 | 44.375 | 45.7 | 46.25
19.3 | 41.6 | 44.125 | 45.5 | 46.0333
19.4 | 41.4 | 43.975 | 45.3 | 45.8333
19.5 | 41.175 | 43.725 | 45 | 45.625
19.6 | 41 | 43.5 | 44.8 | 45.375
19.7 | 40.68 | 43.3 | 44.65 | 45.2143
19.8 | 40.44 | 43.1 | 44.4 | 44.975
19.9 | 40.3 | 42.8667 | 44.1714 | 44.725
20 | 40.1 | 42.62 | 43.92 | 44.5
20.1 | 39.85 | 42.48 | 43.775 | 44.4
20.2 | 39.6 | 42.25 | 43.525 | 44.1333
20.3 | 39.4 | 42.025 | 43.375 | 43.9
20.4 | 39.3 | 41.82 | 43.125 | 43.7333
20.5 | 39.05 | 41.575 | 42.875 | 43.6
20.6 | 38.7667 | 41.4 | 42.7 | 43.275
20.7 | 38.6 | 41.2333 | 42.5 | 43.125
20.8 | 38.45 | 40.9667 | 42.35 | 42.975
20.9 | 38.3 | 40.8 | 42.1 | 42.725
21 | 38.1 | 40.625 | 41.95 | 42.575
21.1 | 37.8667 | 40.425 | 41.7 | 42.45
21.2 | 37.7 | 40.275 | 41.56 | 42.2
21.3 | 37.4 | 40.1 | 41.36 | 41.95
21.4 | 37.3 | 39.8 | 41.125 | 41.8
21.5 | 37.2 | 39.6 | 40.9 | 41.55
21.6 | 36.975 | 39.4 | 40.725 | 41.35
21.7 | 36.75 | 39.2 | 40.575 | 41.175
21.8 | 36.525 | 39.1 | 40.325 | 40.9625
21.9 | 36.3 | 38.85 | 40.175 | 40.8375
22 | 36.18 | 38.72 | 40 | 40.65
22.1 | 35.98 | 38.475 | 39.875 | 40.425
22.2 | 35.775 | 38.25 | 39.625 | 40.2
22.3 | 35.6125 | 38.1 | 39.375 | 40.05
22.4 | 35.3857 | 37.85 | 39.2 | 39.9
22.5 | 35.2 | 37.7 | 39.1 | 39.7
22.6 | 35.025 | 37.6 | 38.9 | 39.55
22.7 | 34.875 | 37.35 | 38.7 | 39.3
22.8 | 34.725 | 37.2 | 38.56 | 39.2
22.9 | 34.575 | 37.06 | 38.325 | 38.925
23 | 34.3 | 36.825 | 38.1 | 38.825
23.1 | 34.225 | 36.575 | 37.925 | 38.6
23.2 | 34 | 36.425 | 37.78 | 38.45
23.3 | 33.9 | 36.2375 | 37.65 | 38.3
23.4 | 33.7 | 36.1 | 37.4 | 38.05
23.5 | 33.5 | 35.875 | 37.275 | 37.9
23.6 | 33.2667 | 35.725 | 37.0571 | 37.75
23.7 | 33.08 | 35.575 | 36.9 | 37.525
23.8 | 32.95 | 35.34 | 36.75 | 37.4
23.9 | 32.72 | 35.1667 | 36.5 | 37.3
24 | 32.6 | 35 | 36.35 | 37.075
24.1 | 32.425 | 34.9 | 36.2 | 36.825
24.2 | 32.375 | 34.7 | 36 | 36.675
24.3 | 32.14 | 34.5 | 35.8 | 36.54
24.4 | 31.95 | 34.35 | 35.6875 | 36.4
24.5 | 31.7833 | 34.2 | 35.5625 | 36.2
24.6 | 31.675 | 34.025 | 35.375 | 36.1
24.7 | 31.475 | 33.875 | 35.1444 | 35.9
24.8 | 31.25 | 33.625 | 34.98 | 35.7
24.9 | 31.08 | 33.5 | 34.825 | 35.5
25 | 30.925 | 33.35 | 34.675 | 35.425
25.1 | 30.82 | 33.2 | 34.5 | 35.275
25.2 | 30.62 | 33.05 | 34.32 | 35.1
25.3 | 30.48 | 32.8 | 34.2 | 34.9
25.4 | 30.3833 | 32.7 | 34.075 | 34.75
25.5 | 30.2 | 32.55 | 33.8333 | 34.6
25.6 | 30.075 | 32.375 | 33.68 | 34.45
25.7 | 29.925 | 32.15 | 33.5714 | 34.2
25.8 | 29.8 | 32 | 33.4 | 34.075
25.9 | 29.65 | 31.9 | 33.2 | 33.925
26 | 29.6 | 31.8 | 33.1 | 33.8
26.1 | 29.35 | 31.65 | 32.95 | 33.7
26.2 | 29.2 | 31.5 | 32.8 | 33.45
26.3 | 29.075 | 31.3 | 32.575 | 33.32
26.4 | 28.825 | 31.2 | 32.4 | 33.175
26.5 | 28.7333 | 30.975 | 32.275 | 33
26.6 | 28.675 | 30.825 | 32.1 | 32.9
26.7 | 28.425 | 30.7 | 31.975 | 32.725
26.8 | 28.3 | 30.6 | 31.825 | 32.575
26.9 | 28.225 | 30.425 | 31.575 | 32.4333
27 | 28 | 30.275 | 31.425 | 32.34
27.1 | 27.85 | 30.1 | 31.3 | 32.18
27.2 | 27.7 | 30 | 31.2 | 32
27.3 | 27.6 | 29.8 | 31 | 31.9
27.4 | 27.5 | 29.7 | 30.8333 | 31.65
27.5 | 27.35 | 29.5 | 30.7 | 31.5
27.6 | 27.2 | 29.4 | 30.6 | 31.4
27.7 | 27.05 | 29.3 | 30.5 | 31.325
27.8 | 26.9 | 29.1 | 30.325 | 31.1
27.9 | 26.8 | 29 | 30.15 | 30.925
28 | 26.6667 | 28.85 | 30.1 | 30.72
28.1 | 26.5 | 28.7 | 29.85 | 30.7
28.2 | 26.4 | 28.6 | 29.7 | 30.55
28.3 | 26.25 | 28.4 | 29.575 | 30.4333
28.4 | 26.1 | 28.3 | 29.45 | 30.275
28.5 | 25.975 | 28.2 | 29.325 | 30.2
28.6 | 25.825 | 28.05 | 29.2 | 30
28.7 | 25.7 | 27.9 | 29.075 | 29.9
28.8 | 25.5 | 27.7 | 28.9333 | 29.75
28.9 | 25.4 | 27.6 | 28.7 | 29.6
29 | 25.3 | 27.475 | 28.6 | 29.5
29.1 | 25.2 | 27.3 | 28.52 | 29.4
29.2 | 25.1 | 27.2 | 28.375 | 29.25
29.3 | 24.975 | 27.06 | 28.225 | 29.1
29.4 | 24.825 | 26.925 | 28.1 | 28.9667
29.5 | 24.675 | 26.8 | 28.025 | 28.8
29.6 | 24.6 | 26.625 | 27.875 | 28.7
29.7 | 24.475 | 26.5 | 27.7 | 28.56
29.8 | 24.325 | 26.4 | 27.6167 | 28.4
29.9 | 24.2 | 26.325 | 27.475 | 28.3
30 | 24.1 | 26.2 | 27.325 | 28.2

## SSD120 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 26.6 | 32.4 | 38 | 45.7
0.1 | 30.875 | 36.375 | 42.025 | 49.66
0.2 | 36.9 | 42.75 | 47.7667 | 55.1
0.3 | 46.525 | 51.85 | 56.5 | 63.1
0.4 | 54.925 | 59.4143 | 63.7 | 69.4
0.5 | 62.1 | 66.5286 | 69.8667 | 75.2
0.6 | 68.2 | 72.15 | 75.3667 | 79.6
0.7 | 73.425 | 76.95 | 79.4333 | 83.42
0.8 | 77.5 | 80.52 | 83.05 | 86.52
0.9 | 81.55 | 83.8 | 86.2 | 89.2
1 | 84.6 | 87 | 88.8 | 91.225
1.1 | 87.45 | 89.5333 | 91.125 | 93.1857
1.2 | 89.7 | 91.4 | 92.8 | 94.55
1.3 | 91.8 | 93.3 | 94.275 | 95.875
1.4 | 93.4667 | 94.55 | 95.7 | 96.95
1.5 | 94.9333 | 96 | 96.7 | 97.775
1.6 | 96.1333 | 97.02 | 97.5 | 98.2
1.7 | 97.2 | 97.575 | 98.4 | 98.95
1.8 | 97.8 | 98.5 | 98.78 | 99.3
1.9 | 98.625 | 99.15 | 99.2 | 99.6
2 | 99.125 | 99.375 | 99.475 | 99.9
2.1 | 99.375 | 99.625 | 99.7 | 99.95
2.2 | 99.55 | 99.8 | 99.88 | 100
2.3 | 99.875 | 99.9333 | 99.9167 | 100
2.4 | 100 | 100 | 100 | 100
2.5 | 100 | 100 | 100 | 99.9
2.6 | 99.9 | 99.9 | 99.9 | 99.74
2.7 | 99.8 | 99.9 | 99.775 | 99.575
2.8 | 99.7 | 99.8 | 99.6 | 99.45
2.9 | 99.5 | 99.65 | 99.4833 | 99.3
3 | 99.4 | 99.5 | 99.275 | 99.1
3.1 | 99.05 | 99.25 | 99.05 | 98.75
3.2 | 98.8 | 98.95 | 98.84 | 98.5143
3.3 | 98.5 | 98.575 | 98.7 | 98.2
3.4 | 98.26 | 98.3222 | 98.35 | 97.95
3.5 | 97.95 | 98.0625 | 98 | 97.65
3.6 | 97.55 | 97.6875 | 97.65 | 97.325
3.7 | 97.15 | 97.375 | 97.3 | 97.175
3.8 | 96.95 | 97.125 | 96.95 | 96.85
3.9 | 96.55 | 96.6875 | 96.6 | 96.5
4 | 96.05 | 96.4875 | 96.2 | 96.2
4.1 | 95.625 | 96.1125 | 95.9 | 95.86
4.2 | 95.15 | 95.825 | 95.575 | 95.46
4.3 | 94.8 | 95.26 | 95.3 | 95.225
4.4 | 94.5 | 94.9375 | 94.9 | 94.8625
4.5 | 94.15 | 94.6 | 94.7 | 94.5
4.6 | 93.7 | 94.3333 | 94.2 | 94.2
4.7 | 93.4 | 93.86 | 94.025 | 94
4.8 | 93.05 | 93.6 | 93.725 | 93.675
4.9 | 92.5 | 93.375 | 93.28 | 93.325
5 | 92.0333 | 93 | 93.05 | 92.98
5.1 | 91.7 | 92.6625 | 92.8 | 92.7
5.2 | 91.325 | 92.325 | 92.5 | 92.2
5.3 | 90.775 | 91.9222 | 92 | 91.95
5.4 | 90.45 | 91.5 | 91.6857 | 91.7
5.5 | 90.075 | 91.2 | 91.4 | 91.375
5.6 | 89.825 | 90.85 | 91.0625 | 91.125
5.7 | 89.475 | 90.5333 | 90.725 | 90.8
5.8 | 89.2 | 90.275 | 90.475 | 90.64
5.9 | 88.7333 | 89.95 | 90.275 | 90.3
6 | 88.5 | 89.45 | 89.95 | 89.9
6.1 | 87.95 | 89.125 | 89.54 | 89.52
6.2 | 87.6 | 88.875 | 89.2 | 89.1333
6.3 | 87 | 88.625 | 88.9 | 88.9
6.4 | 86.575 | 88.2667 | 88.65 | 88.74
6.5 | 86.225 | 87.8 | 88.3 | 88.275
6.6 | 86.075 | 87.62 | 88.025 | 88.025
6.7 | 85.475 | 87.375 | 87.7 | 87.7
6.8 | 85.1 | 86.95 | 87.35 | 87.4
6.9 | 84.84 | 86.675 | 87 | 87
7 | 84.4 | 86.425 | 86.75 | 86.775
7.1 | 84.0333 | 86.075 | 86.5 | 86.45
7.2 | 83.6 | 85.8 | 86.15 | 86.075
7.3 | 83.4 | 85.3 | 85.75 | 85.925
7.4 | 83.0625 | 85 | 85.475 | 85.6
7.5 | 82.7143 | 84.6 | 85.225 | 85.2
7.6 | 82.3333 | 84.35 | 84.775 | 84.9
7.7 | 81.9375 | 84.1 | 84.48 | 84.7
7.8 | 81.6 | 83.78 | 84.2 | 84.3
7.9 | 81.2375 | 83.4 | 83.95 | 84.2
8 | 80.94 | 83.0333 | 83.5 | 83.8
8.1 | 80.64 | 82.675 | 83.2 | 83.6
8.2 | 80.2 | 82.45 | 82.8667 | 83.2
8.3 | 79.76 | 82.125 | 82.6 | 82.9
8.4 | 79.4 | 81.875 | 82.175 | 82.7
8.5 | 79.075 | 81.55 | 81.925 | 82.325
8.6 | 78.725 | 81.16 | 81.775 | 82
8.7 | 78.5 | 80.85 | 81.6 | 81.75
8.8 | 77.95 | 80.4 | 81.3 | 81.5
8.9 | 77.6 | 80.05 | 80.85 | 81.12
9 | 77.25 | 79.82 | 80.6 | 80.75
9.1 | 76.9 | 79.45 | 80.2 | 80.575
9.2 | 76.55 | 79.15 | 79.9 | 80.25
9.3 | 76.2 | 78.75 | 79.65 | 79.98
9.4 | 75.85 | 78.48 | 79.42 | 79.725
9.5 | 75.5 | 78.15 | 79.15 | 79.4
9.6 | 75.275 | 77.9 | 78.725 | 79.225
9.7 | 74.95 | 77.54 | 78.5167 | 79
9.8 | 74.575 | 77.1667 | 78.15 | 78.65
9.9 | 74.34 | 76.85 | 77.875 | 78.3
10 | 73.7 | 76.7 | 77.625 | 78.05
10.1 | 73.35 | 76.3 | 77.3 | 77.8
10.2 | 73.05 | 76.05 | 76.9 | 77.5
10.3 | 72.9 | 75.8 | 76.6 | 77.1
10.4 | 72.5 | 75.4 | 76.35 | 76.85
10.5 | 72.2 | 75.04 | 76.1 | 76.6
10.6 | 71.75 | 74.8 | 75.7667 | 76.36
10.7 | 71.4875 | 74.625 | 75.475 | 76.125
10.8 | 71.175 | 74.275 | 75.275 | 75.775
10.9 | 70.875 | 73.92 | 75.025 | 75.525
11 | 70.55 | 73.575 | 74.58 | 75.275
11.1 | 70.1333 | 73.325 | 74.4 | 74.95
11.2 | 69.775 | 73.05 | 74.1 | 74.7
11.3 | 69.525 | 72.725 | 73.8 | 74.45
11.4 | 69.375 | 72.4 | 73.625 | 74.2
11.5 | 69.0667 | 72.2 | 73.375 | 74
11.6 | 68.65 | 71.9 | 73.125 | 73.7
11.7 | 68.2333 | 71.55 | 72.775 | 73.325
11.8 | 68.025 | 71.3 | 72.6 | 73.2
11.9 | 67.76 | 71.1 | 72.375 | 72.85
12 | 67.25 | 70.9 | 72.0333 | 72.4
12.1 | 67 | 70.4 | 71.6 | 72.2
12.2 | 66.55 | 70.1 | 71.35 | 71.9
12.3 | 66.3 | 69.95 | 71 | 71.675
12.4 | 66.15 | 69.7 | 70.75 | 71.5
12.5 | 65.9667 | 69.45 | 70.525 | 71.175
12.6 | 65.75 | 69.18 | 70.2 | 70.925
12.7 | 65.25 | 68.7 | 69.9667 | 70.575
12.8 | 64.95 | 68.45 | 69.8 | 70.35
12.9 | 64.65 | 68.3 | 69.575 | 70.08
13 | 64.325 | 67.95 | 69.325 | 69.8333
13.1 | 64.075 | 67.675 | 69.075 | 69.6
13.2 | 63.825 | 67.325 | 68.875 | 69.375
13.3 | 63.5714 | 66.95 | 68.575 | 69.05
13.4 | 63.3 | 66.6667 | 68.325 | 68.8
13.5 | 62.95 | 66.4333 | 68 | 68.66
13.6 | 62.6 | 66.2667 | 67.675 | 68.525
13.7 | 62.35 | 66 | 67.45 | 68.15
13.8 | 62 | 65.725 | 67.175 | 67.825
13.9 | 61.75 | 65.425 | 66.9333 | 67.6
14 | 61.46 | 65.125 | 66.7 | 67.425
14.1 | 61.1 | 64.95 | 66.4 | 67.025
14.2 | 60.7 | 64.6 | 66.1333 | 66.8
14.3 | 60.4 | 64.35 | 65.9 | 66.6
14.4 | 60.22 | 64 | 65.6 | 66.35
14.5 | 59.975 | 63.85 | 65.25 | 66.1
14.6 | 59.7625 | 63.6 | 65.125 | 65.95
14.7 | 59.4625 | 63.35 | 64.875 | 65.62
14.8 | 59.2125 | 63.1 | 64.725 | 65.32
14.9 | 58.8375 | 62.7 | 64.3667 | 65.125
15 | 58.675 | 62.45 | 64.1 | 64.875
15.1 | 58.4333 | 62.3 | 63.8 | 64.62
15.2 | 58.175 | 62.05 | 63.625 | 64.2714
15.3 | 57.925 | 61.8 | 63.475 | 64.05
15.4 | 57.5 | 61.52 | 63.15 | 63.85
15.5 | 57.2333 | 61.3 | 62.7778 | 63.7
15.6 | 57.1 | 61.1 | 62.625 | 63.4429
15.7 | 56.8 | 60.76 | 62.3667 | 63.2
15.8 | 56.5 | 60.525 | 62.18 | 62.775
15.9 | 56.25 | 60.225 | 61.95 | 62.54
16 | 56.1 | 59.9875 | 61.6 | 62.3
16.1 | 55.75 | 59.7286 | 61.45 | 62.125
16.2 | 55.5 | 59.4571 | 61.2 | 61.9
16.3 | 55.25 | 59.35 | 60.9875 | 61.65
16.4 | 55 | 59.1 | 60.825 | 61.5
16.5 | 54.775 | 58.8 | 60.575 | 61.35
16.6 | 54.425 | 58.6 | 60.4 | 61
16.7 | 54.2 | 58.25 | 60.05 | 60.85
16.8 | 54 | 58 | 59.8 | 60.5
16.9 | 53.675 | 57.875 | 59.45 | 60.3
17 | 53.425 | 57.625 | 59.125 | 60.075
17.1 | 53.2 | 57.35 | 58.975 | 59.825
17.2 | 53.04 | 57.1 | 58.825 | 59.6
17.3 | 52.8 | 56.9 | 58.575 | 59.45
17.4 | 52.65 | 56.75 | 58.25 | 59.2
17.5 | 52.425 | 56.4 | 58.1 | 58.95
17.6 | 52.1333 | 56.15 | 57.95 | 58.72
17.7 | 51.875 | 56 | 57.6 | 58.4571
17.8 | 51.6333 | 55.62 | 57.3 | 58.3
17.9 | 51.3 | 55.425 | 57.12 | 58.15
18 | 51.1667 | 55.175 | 56.9333 | 57.8857
18.1 | 50.95 | 54.925 | 56.6625 | 57.6875
18.2 | 50.6 | 54.675 | 56.475 | 57.4875
18.3 | 50.45 | 54.54 | 56.225 | 57.2
18.4 | 50.225 | 54.3 | 56.0833 | 56.925
18.5 | 49.94 | 54.05 | 55.9167 | 56.82
18.6 | 49.8111 | 53.8714 | 55.675 | 56.575
18.7 | 49.675 | 53.6 | 55.35 | 56.425
18.8 | 49.3875 | 53.4 | 55.1 | 56.15
18.9 | 49.0125 | 53.2 | 54.95 | 55.925
19 | 48.725 | 53 | 54.7 | 55.775
19.1 | 48.5 | 52.7714 | 54.55 | 55.525
19.2 | 48.4 | 52.45 | 54.3 | 55.375
19.3 | 48.1 | 52.2 | 54.075 | 55.14
19.4 | 48 | 52.0143 | 53.825 | 54.9
19.5 | 47.66 | 51.7857 | 53.66 | 54.65
19.6 | 47.52 | 51.62 | 53.4 | 54.4
19.7 | 47.25 | 51.3875 | 53.15 | 54.1333
19.8 | 47.025 | 51.2667 | 52.9 | 54.075
19.9 | 46.875 | 51.1 | 52.8 | 53.825
20 | 46.725 | 50.825 | 52.6 | 53.6667
20.1 | 46.48 | 50.5333 | 52.375 | 53.4
20.2 | 46.25 | 50.4 | 52.2 | 53.1857
20.3 | 46 | 50.05 | 52.05 | 53
20.4 | 45.8333 | 50 | 51.725 | 52.75
20.5 | 45.5 | 49.65 | 51.575 | 52.5
20.6 | 45.32 | 49.425 | 51.3625 | 52.4
20.7 | 45.1 | 49.28 | 51.2286 | 52.175
20.8 | 44.925 | 49.05 | 50.96 | 52
20.9 | 44.8167 | 48.9 | 50.8 | 51.75
21 | 44.5889 | 48.75 | 50.6 | 51.6
21.1 | 44.45 | 48.6 | 50.35 | 51.3857
21.2 | 44.1286 | 48.3 | 50.1 | 51.22
21.3 | 43.85 | 48.12 | 49.95 | 51.0875
21.4 | 43.6 | 47.875 | 49.8 | 50.85
21.5 | 43.45 | 47.725 | 49.5 | 50.6
21.6 | 43.3 | 47.575 | 49.3 | 50.425
21.7 | 43.15 | 47.325 | 49.2 | 50.2
21.8 | 43.02 | 47 | 49 | 50
21.9 | 42.775 | 46.825 | 48.85 | 49.9
22 | 42.525 | 46.68 | 48.65 | 49.7
22.1 | 42.375 | 46.34 | 48.45 | 49.4
22.2 | 42.08 | 46.2 | 48.125 | 49.2
22.3 | 42 | 46.05 | 47.9167 | 49.1
22.4 | 41.85 | 45.8 | 47.775 | 49
22.5 | 41.6 | 45.6 | 47.625 | 48.75
22.6 | 41.4 | 45.45 | 47.5 | 48.5
22.7 | 41.225 | 45.25 | 47.325 | 48.3
22.8 | 40.975 | 45.1 | 47.0333 | 48.2
22.9 | 40.8 | 45 | 46.8 | 47.975
23 | 40.675 | 44.85 | 46.7 | 47.75
23.1 | 40.5 | 44.7 | 46.5 | 47.5
23.2 | 40.25 | 44.45 | 46.4 | 47.35
23.3 | 40.1 | 44.225 | 46.2 | 47.2
23.4 | 39.8333 | 43.975 | 45.9 | 47.1
23.5 | 39.7 | 43.9 | 45.725 | 46.9
23.6 | 39.55 | 43.675 | 45.6 | 46.65
23.7 | 39.4 | 43.425 | 45.425 | 46.5
23.8 | 39.05 | 43.3 | 45.3 | 46.35
23.9 | 38.9 | 43.1 | 45.025 | 46.125
24 | 38.8 | 42.925 | 44.8 | 45.9
24.1 | 38.62 | 42.78 | 44.65 | 45.725
24.2 | 38.4 | 42.55 | 44.5 | 45.5222
24.3 | 38.2 | 42.4 | 44.4 | 45.36
24.4 | 38.1 | 42.26 | 44.2 | 45.1833
24.5 | 38 | 42.1 | 44.05 | 45.1
24.6 | 37.775 | 41.975 | 43.8 | 44.95
24.7 | 37.525 | 41.725 | 43.6 | 44.8
24.8 | 37.4 | 41.475 | 43.4 | 44.675
24.9 | 37.225 | 41.325 | 43.275 | 44.5
25 | 37.075 | 41.2 | 43.1 | 44.3
25.1 | 36.925 | 41.02 | 43 | 44.12
25.2 | 36.78 | 40.775 | 42.825 | 43.9
25.3 | 36.55 | 40.7 | 42.675 | 43.6
25.4 | 36.3 | 40.575 | 42.425 | 43.5
25.5 | 36.2 | 40.325 | 42.275 | 43.325
25.6 | 36 | 40.2 | 42.025 | 43.2
25.7 | 35.85 | 40.0333 | 41.875 | 43.025
25.8 | 35.7 | 39.75 | 41.7 | 42.9
25.9 | 35.6 | 39.6 | 41.6 | 42.76
26 | 35.5 | 39.45 | 41.4 | 42.5
26.1 | 35.3 | 39.325 | 41.2 | 42.375
26.2 | 35.125 | 39.1 | 41 | 42.2
26.3 | 34.9875 | 39 | 40.8667 | 42
26.4 | 34.85 | 38.775 | 40.7 | 41.925
26.5 | 34.675 | 38.625 | 40.525 | 41.775
26.6 | 34.4571 | 38.475 | 40.4 | 41.625
26.7 | 34.25 | 38.225 | 40.3 | 41.475
26.8 | 34.1 | 38.075 | 40.1 | 41.325
26.9 | 34 | 37.9 | 39.94 | 41.1
27 | 33.9571 | 37.7 | 39.7333 | 40.875
27.1 | 33.625 | 37.6 | 39.5333 | 40.725
27.2 | 33.5 | 37.475 | 39.4 | 40.675
27.3 | 33.45 | 37.325 | 39.25 | 40.325
27.4 | 33.2 | 37.2 | 39.1 | 40.2
27.5 | 33.05 | 36.95 | 38.95 | 40.1
27.6 | 32.9 | 36.8 | 38.8 | 40
27.7 | 32.8 | 36.65 | 38.65 | 39.8
27.8 | 32.575 | 36.6 | 38.5 | 39.6
27.9 | 32.4222 | 36.4 | 38.4 | 39.5
28 | 32.3 | 36.25 | 38.2 | 39.4
28.1 | 32.2 | 36.1 | 37.95 | 39.2
28.2 | 32.1 | 35.95 | 37.8 | 39
28.3 | 32 | 35.8 | 37.725 | 38.8667
28.4 | 31.9 | 35.65 | 37.6 | 38.7
28.5 | 31.575 | 35.5 | 37.45 | 38.5
28.6 | 31.45 | 35.3667 | 37.3 | 38.325
28.7 | 31.325 | 35.3 | 37.15 | 38.2
28.8 | 31.2 | 35.1 | 36.9667 | 38.1
28.9 | 31.075 | 34.825 | 36.8 | 38
29 | 30.9 | 34.7 | 36.7 | 37.8
29.1 | 30.85 | 34.6 | 36.55 | 37.7
29.2 | 30.7 | 34.4 | 36.4 | 37.575
29.3 | 30.6 | 34.25 | 36.25 | 37.4625
29.4 | 30.4 | 34.1 | 36.2 | 37.275
29.5 | 30.25 | 34 | 36.05 | 37.1125
29.6 | 30.1 | 33.82 | 35.8 | 36.9875
29.7 | 30 | 33.6875 | 35.7 | 36.8667
29.8 | 29.9 | 33.5625 | 35.7 | 36.75
29.9 | 29.8 | 33.4167 | 35.5 | 36.5714
30 | 29.7 | 33.325 | 35.2 | 36.4375

## Wedge 60 Lateral Profiles
Lateral Profiles
Wedge 60
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:2.4 cm FieldSize:4x4 cm2 | Depth:2.4 cm FieldSize:10x10 cm2 | Depth:2.4 cm FieldSize:15x15 cm2 | Depth:2.4 cm FieldSize:40x15 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:15x15 cm2 | Depth:5 cm FieldSize:40x15 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:15x15 cm2 | Depth:10 cm FieldSize:40x15 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:15x15 cm2 | Depth:20 cm FieldSize:40x15 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:15x15 cm2 | Depth:30 cm FieldSize:40x15 cm2
-14.7 | [18 empty] | 12.4 | 17.1667
-14.6 | [18 empty] | 12.5333 | 17.35
-14.5 | [18 empty] | 12.775 | 17.4333
-14.4 | [18 empty] | 13.025 | 17.7
-14.3 | [18 empty] | 13.3 | 17.9667
-14.2 | [18 empty] | 13.4 | 18.2
-14.1 | [18 empty] | 13.625 | 18.4
-14 | [18 empty] | 13.7667 | 18.7333
-13.9 | [14 empty] | 10 | 14.2047 |  |  | 14 | 18.9667
-13.8 | [14 empty] | 10.2 | 14.3298 |  |  | 14.3 | 19.2
-13.7 | [14 empty] | 10.38 | 14.5799 |  |  | 14.5 | 19.5
-13.6 | [14 empty] | 10.5 | 14.8049 |  |  | 14.72 | 19.8
-13.5 | [14 empty] | 10.8 | 15.005 |  |  | 15.1 | 20
-13.4 | [14 empty] | 11.05 | 15.2051 |  |  | 15.3 | 20.2
-13.3 | [14 empty] | 11.18 | 15.3384 |  |  | 15.6 | 20.4333
-13.2 | [10 empty] | 7.3333 | 11.1 |  |  | 11.4 | 15.7052 |  |  | 15.8 | 20.75
-13.1 | [10 empty] | 7.475 | 11.225 |  |  | 11.575 | 15.9453 |  |  | 16.1 | 21.025
-13 | [10 empty] | 7.66 | 11.4 |  |  | 11.7333 | 16.2387 |  |  | 16.3667 | 21.35
-12.9 | [10 empty] | 7.8 | 11.575 |  |  | 12 | 16.4055 |  |  | 16.6 | 21.725
-12.8 | [6 empty] | 6.25 | 9.7 |  |  | 7.9 | 11.75 |  |  | 12.275 | 16.7056 |  |  | 17.0667 | 21.975
-12.7 | [6 empty] | 6.4 | 9.8667 |  |  | 8.1 | 11.9 |  |  | 12.45 | 17.0057 |  |  | 17.3 | 22.3
-12.6 |  |  | 6.35 | 10.025 |  |  | 6.5 | 10 |  |  | 8.2 | 12.075 |  |  | 12.7 | 17.1307 |  |  | 17.7 | 22.5667
-12.5 |  |  | 6.4 | 10.2 |  |  | 6.6 | 10.1 |  |  | 8.34 | 12.3 |  |  | 12.96 | 17.4308 |  |  | 17.9 | 23.0333
-12.4 |  |  | 6.55 | 10.36 |  |  | 6.7 | 10.2 |  |  | 8.55 | 12.5 |  |  | 13.2 | 17.6309 |  |  | 18.2333 | 23.3
-12.3 |  |  | 6.6333 | 10.5 |  |  | 6.8 | 10.4 |  |  | 8.725 | 12.625 |  |  | 13.5 | 17.881 |  |  | 18.55 | 23.7
-12.2 |  |  | 6.8 | 10.6 |  |  | 6.9 | 10.54 |  |  | 8.9 | 12.8 |  |  | 13.75 | 18.2561 |  |  | 18.95 | 23.95
-12.1 |  |  | 6.8 | 10.7 |  |  | 7.1 | 10.7 |  |  | 9 | 13 |  |  | 14.1 | 18.5062 |  |  | 19.3 | 24.35
-12 |  |  | 6.9667 | 10.8667 |  |  | 7.24 | 10.85 |  |  | 9.3 | 13.2 |  |  | 14.4 | 18.7663 |  |  | 19.7 | 24.825
-11.9 |  |  | 7.125 | 11 |  |  | 7.4 | 10.975 |  |  | 9.4 | 13.42 |  |  | 14.6 | 19.1064 |  |  | 20.1 | 25.25
-11.8 |  |  | 7.2 | 11.2 |  |  | 7.5 | 11.125 |  |  | 9.6667 | 13.6 |  |  | 14.875 | 19.4065 |  |  | 20.5 | 25.6
-11.7 |  |  | 7.325 | 11.3 |  |  | 7.6 | 11.375 |  |  | 9.82 | 13.8 |  |  | 15.25 | 19.7066 |  |  | 20.9 | 26.05
-11.6 |  |  | 7.475 | 11.425 |  |  | 7.75 | 11.5 |  |  | 10.1 | 14.1 |  |  | 15.62 | 20.0067 |  |  | 21.35 | 26.4
-11.5 |  |  | 7.6 | 11.675 |  |  | 7.9 | 11.675 |  |  | 10.3 | 14.35 |  |  | 16 | 20.4068 |  |  | 21.8 | 26.95
-11.4 |  |  | 7.75 | 11.8 |  |  | 8.0667 | 11.8667 |  |  | 10.5 | 14.6 |  |  | 16.3 | 20.8069 |  | 8.1459 | 22.3667 | 27.4
-11.3 |  |  | 7.9 | 12 |  |  | 8.2 | 11.975 |  |  | 10.7333 | 14.7667 |  |  | 16.65 | 21.1821 |  | 8.2958 | 23 | 28.15
-11.2 |  |  | 8.1 | 12.1667 |  |  | 8.4333 | 12.3 |  |  | 10.975 | 15.075 |  |  | 17 | 21.5572 |  | 8.4458 | 23.5333 | 28.6
-11.1 |  |  | 8.2 | 12.3333 |  |  | 8.7 | 12.4 |  |  | 11.2 | 15.325 |  |  | 17.4 | 21.9573 |  | 8.5957 | 24.175 | 29.25
-11 |  |  | 8.45 | 12.525 |  |  | 8.8 | 12.6 |  |  | 11.4333 | 15.6 |  |  | 17.85 | 22.4075 |  | 8.6957 | 25 | 30.02
-10.9 |  |  | 8.58 | 12.7 |  |  | 8.975 | 12.85 |  |  | 11.7333 | 15.875 |  | 6.4 | 18.3 | 22.8409 |  | 8.8706 | 25.8 | 30.8
-10.8 |  |  | 8.75 | 12.9 |  |  | 9.2333 | 12.98 |  |  | 12 | 16.24 |  | 6.5 | 18.72 | 23.3078 |  | 9.0954 | 26.8 | 31.8
-10.7 |  |  | 8.9 | 13.1 |  |  | 9.4333 | 13.25 |  |  | 12.35 | 16.475 |  | 6.6 | 19.225 | 23.7079 |  | 9.1954 | 28 | 33
-10.6 |  |  | 9.15 | 13.2 |  |  | 9.7 | 13.5 |  |  | 12.7 | 16.8333 |  | 6.7 | 19.75 | 24.3081 |  | 9.3453 | 29.5 | 34.375
-10.5 |  |  | 9.375 | 13.45 |  |  | 9.875 | 13.625 |  |  | 13.05 | 17.05 |  | 6.8 | 20.325 | 24.9583 |  | 9.4953 | 31.3 | 36.125
-10.4 |  |  | 9.5 | 13.6 |  |  | 10.1 | 13.95 |  | 4.5982 | 13.3 | 17.45 |  | 7 | 20.95 | 25.6085 |  | 9.6702 | 33.45 | 38.275
-10.3 |  |  | 9.675 | 13.85 |  |  | 10.35 | 14.2 |  | 4.6315 | 13.74 | 17.92 |  | 7.0333 | 21.725 | 26.1687 |  | 9.8151 | 36.52 | 41.26
-10.2 |  |  | 9.9 | 14.1 |  | 3.8 | 10.6 | 14.45 |  | 4.6981 | 14.1 | 18.325 |  | 7.2 | 22.475 | 27.009 |  | 10.045 | 40.9 | 45.2
-10.1 |  | 3.8 | 10.05 | 14.375 |  | 3.8 | 10.85 | 14.825 |  | 4.7981 | 14.45 | 18.65 |  | 7.3 | 23.475 | 27.8092 |  | 10.1949 | 46.8 | 50.7
-10 |  | 3.825 | 10.325 | 14.65 |  | 3.9 | 11.1 | 15.05 |  | 4.898 | 14.85 | 19.1 |  | 7.425 | 24.525 | 28.9597 |  | 10.3449 | 55.25 | 58.45
-9.9 |  | 3.9 | 10.6 | 14.8333 |  | 3.9 | 11.45 | 15.4 |  | 4.998 | 15.35 | 19.6 |  | 7.6 | 25.85 | 30.16 |  | 10.5947 | 65.875 | 68.7
-9.8 |  | 3.975 | 10.825 | 15.1 |  | 4 | 11.7333 | 15.7 |  | 5.098 | 15.85 | 20.1 |  | 7.75 | 27.6333 | 31.6606 |  | 10.8346 | 78.12 | 79.9
-9.7 |  | 4 | 11.1 | 15.4 |  | 4.0667 | 12.1333 | 16.1 |  | 5.1479 | 16.4667 | 20.62 |  | 7.9 | 29.6 | 33.7112 |  | 10.9945 | 91.75 | 92.5
-9.6 |  | 4.0667 | 11.3333 | 15.625 |  | 4.1 | 12.5 | 16.4 |  | 5.2729 | 17.0333 | 21.25 |  | 8.075 | 32.2 | 36.012 |  | 11.1944 | 103.4 | 104.4
-9.5 |  | 4.1 | 11.6 | 16 |  | 4.2 | 12.925 | 16.8333 |  | 5.4178 | 17.8667 | 21.975 |  | 8.2667 | 36.3 | 39.5632 |  | 11.4193 | 115.4 | 114
-9.4 |  | 4.2 | 12.025 | 16.3333 |  | 4.3 | 13.35 | 17.25 |  | 5.4978 | 18.6 | 22.825 |  | 8.4667 | 41.9 | 44.6149 |  | 11.6442 | 125.2 | 123
-9.3 |  | 4.3 | 12.3667 | 16.675 |  | 4.375 | 13.875 | 17.825 |  | 5.5978 | 19.7 | 23.8 |  | 8.64 | 49.9 | 52.2674 |  | 11.7941 | 132.35 | 129.45
-9.2 |  | 4.35 | 12.7333 | 17.1333 |  | 4.4 | 14.525 | 18.375 |  | 5.7227 | 20.8 | 24.975 |  | 8.875 | 60.325 | 61.5538 |  | 12.094 | 137.2 | 134
-9.1 |  | 4.4 | 13.1667 | 17.55 |  | 4.5 | 15.2 | 19.125 |  | 5.8477 | 22.4333 | 26.475 |  | 9.025 | 73.275 | 74.5249 |  | 12.3938 | 140.35 | 136.575
-9 |  | 4.5 | 13.7 | 17.9667 |  | 4.6 | 16.025 | 19.9 |  | 5.9976 | 24.25 | 28.3 |  | 9.275 | 88.3 | 88.2294 |  | 12.6437 | 142.25 | 138.65
-8.9 |  | 4.6 | 14.275 | 18.6 |  | 4.7 | 17 | 20.8 |  | 6.0976 | 26.6 | 30.6667 |  | 9.45 | 102.9 | 101.784 |  | 12.8936 | 143.65 | 139.7
-8.8 |  | 4.7 | 14.9 | 19.3 |  | 4.8 | 18.2 | 22 |  | 6.1975 | 30.1 | 34.2 |  | 9.7 | 116.35 | 113.7882 |  | 13.1934 | 144.425 | 140.44
-8.7 |  | 4.775 | 15.825 | 20.3 |  | 4.9 | 19.8 | 23.3 |  | 6.3974 | 34.9 | 38.8 |  | 9.85 | 127.925 | 124.5915 |  | 13.46 | 144.925 | 140.8
-8.6 |  | 4.8667 | 16.8 | 21.3333 |  | 5 | 21.85 | 25.6 |  | 6.5474 | 41.65 | 46 |  | 10.1 | 136.3 | 132.9192 |  | 13.8431 | 145.25 | 141.0333
-8.5 |  | 4.9 | 18.3 | 22.65 |  | 5.1 | 24.6 | 28 |  | 6.6973 | 53.2 | 56.2 |  | 10.45 | 141.75 | 138.321 |  | 14.118 | 145.4 | 141.2
-8.4 |  | 5 | 20.2 | 24.4 |  | 5.2 | 28.55 | 31.7 |  | 6.8306 | 66.5 | 70.75 |  | 10.7 | 145.8 | 141.8972 |  | 14.3678 | 145.4 | 141.2
-8.3 |  | 5.1 | 22.7 | 27.4 |  | 5.4 | 34.2 | 37.2 |  | 7.0222 | 82.45 | 86.1 |  | 10.95 | 147.65 | 144.123 |  | 14.726 | 145.3 | 141.15
-8.2 |  | 5.25 | 26.5 | 31.1 |  | 5.475 | 43.6 | 45.775 |  | 7.1971 | 99.925 | 101.15 |  | 11.2 | 149.1 | 145.3817 |  | 15.1175 | 145.0333 | 141
-8.1 |  | 5.3667 | 32.3 | 36.7 |  | 5.6 | 56.14 | 58.1333 |  | 7.397 | 116.775 | 117.4 |  | 11.48 | 149.8 | 146.1823 |  | 15.6588 | 144.6667 | 140.7
-8 |  | 5.5 | 41.6 | 47.05 |  | 5.7 | 73.0667 | 73.3 |  | 7.617 | 131.025 | 130.325 |  | 11.88 | 150.4 | 146.609 |  | 16.052 | 144.5 | 140.475
-7.9 |  | 5.6 | 55 | 59.7 |  | 5.9 | 91.025 | 91.6 |  | 7.8469 | 142.625 | 140.675 |  | 12.175 | 150.54 | 146.749 |  | 16.5918 | 144.1667 | 140.225
-7.8 |  | 5.7 | 74.05 | 77.625 |  | 6.025 | 109.675 | 108.65 |  | 8.0968 | 149.6 | 147.1 |  | 12.55 | 150.6 | 146.749 |  | 17.1914 | 143.75 | 139.88
-7.7 |  | 5.85 | 92.1 | 95.6 |  | 6.2 | 127.175 | 125.3 |  | 8.2967 | 153.9 | 151.06 |  | 12.95 | 150.475 | 146.649 |  | 17.8411 | 143.3 | 139.65
-7.6 |  | 6 | 110.8 | 113.5 |  | 6.4 | 141.725 | 139.125 |  | 8.4966 | 156.45 | 153.3 |  | 13.32 | 150.225 | 146.449 |  | 18.4908 | 142.9333 | 139.1
-7.5 |  | 6.1 | 129.85 | 130.15 |  | 6.575 | 152.1 | 149 |  | 8.7965 | 158.1 | 154.7 |  | 13.6667 | 149.85 | 146.049 | 2.9985 | 19.6402 | 142.4333 | 138.6333
-7.4 |  | 6.3 | 144.7667 | 145 |  | 6.7333 | 158.2 | 154.85 |  | 9.0964 | 158.8 | 155.3 |  | 14.3 | 149.45 | 145.669 | 3.1984 | 20.6231 | 141.8 | 138.3
-7.3 |  | 6.425 | 156.125 | 154 |  | 6.92 | 161.8 | 158.3 |  | 9.4462 | 159.1 | 155.5 | 2.2 | 14.8333 | 149 | 145.423 | 3.2984 | 22.069 | 141.2667 | 137.9
-7.2 |  | 6.675 | 162.525 | 159.8 |  | 7.175 | 163.85 | 160.25 |  | 9.7628 | 159.1 | 155.5 | 2.3 | 15.4 | 148.46 | 144.823 | 3.2984 | 23.6632 | 140.82 | 137.4333
-7.1 |  | 6.9 | 165.7 | 162.7667 | 1.4 | 7.4 | 164.7 | 161.15 | 1.6 | 10.121 | 159 | 155.4 | 2.4 | 16.3 | 147.95 | 144.273 | 3.3983 | 25.8121 | 140.45 | 136.825
-7 | 1.4 | 7.05 | 167.2 | 164.3 | 1.4 | 7.75 | 164.84 | 161.2 | 1.6 | 10.5458 | 158.575 | 155.05 | 2.4 | 17.1 | 147.16 | 143.773 | 3.3983 | 29.1604 | 140 | 136.5
-6.9 | 1.4 | 7.3 | 167.6 | 164.5 | 1.4 | 8.02 | 164.9 | 161.3 | 1.6 | 11.0289 | 158 | 154.55 | 2.4 | 18.3 | 146.7 | 143.3147 | 3.4983 | 33.4833 | 139.1 | 135.95
-6.8 | 1.4 | 7.5 | 167.6 | 164.5 | 1.4 | 8.3 | 164.725 | 161.1 | 1.7 | 11.5621 | 157.5 | 153.95 | 2.4 | 19.575 | 146 | 142.6478 | 3.4983 | 39.0804 | 138.675 | 135.45
-6.7 | 1.4 | 7.8 | 167.6 | 164.35 | 1.4 | 8.78 | 164.325 | 160.825 | 1.7 | 12.1452 | 156.875 | 153.375 | 2.45 | 21.4 | 145.4 | 142.127 | 3.5982 | 47.5762 | 138.35 | 135.06
-6.6 | 1.4 | 8.1667 | 167.05 | 163.65 | 1.4 | 9.2333 | 163.6 | 159.95 | 1.7 | 12.9615 | 155.9333 | 152.725 | 2.5667 | 23.675 | 144.6 | 141.547 | 3.5982 | 57.2048 | 137.7 | 134.2667
-6.5 | 1.4 | 8.4667 | 166.2 | 163.1333 | 1.4 | 9.6667 | 162.4 | 159.1 | 1.7 | 13.7945 | 155.1333 | 151.975 | 2.7 | 26.85 | 143.9 | 140.947 | 3.6982 | 68.8156 | 137.1 | 133.6
-6.4 | 1.4 | 8.8667 | 165.2333 | 162.0333 | 1.4 | 10.3 | 161.7 | 158.5333 | 1.8 | 14.894 | 154.6 | 151.22 | 2.7 | 31.3 | 143.2 | 140.047 | 3.7981 | 80.6597 | 136.5 | 133.3
-6.3 | 1.4 | 9.3 | 164.0667 | 160.9667 | 1.5 | 11.0333 | 160.8 | 157.5 | 1.8 | 16.2935 | 153.7 | 150.3 | 2.8 | 37.8333 | 142.6 | 139.5465 | 3.7981 | 91.8541 | 135.8667 | 133
-6.2 | 1.4 | 10 | 163.25 | 159.925 | 1.5 | 12.1 | 159.65 | 156.5333 | 1.8 | 18.1427 | 152.7 | 149.2 | 2.8 | 47.2 | 141.8 | 139.021 | 3.8981 | 101.3827 | 135.3333 | 132.35
-6.1 | 1.5 | 10.725 | 161.925 | 158.9 | 1.5 | 13.25 | 158.6 | 155.7 | 1.9 | 20.2669 | 151.5 | 148.35 | 2.9 | 58.3 | 141.05 | 137.996 | 3.998 | 108.4955 | 134.65 | 131.6667
-6 | 1.5 | 11.5333 | 160.75 | 157.75 | 1.5 | 14.65 | 157.66 | 154.5 | 1.9 | 23.6239 | 150.65 | 147.5667 | 3 | 70.3 | 140.5 | 137.696 | 4.0979 | 114.043 | 133.8333 | 131.1333
-5.9 | 1.5 | 12.95 | 159.5 | 156.65 | 1.6 | 16.85 | 156.2 | 153.425 | 1.9 | 28.3886 | 149.9 | 146.7 | 3 | 83.175 | 139.5 | 136.896 | 4.1979 | 117.4415 | 133.3 | 130.7333
-5.8 | 1.5 | 14.6 | 158.2 | 155.4 | 1.6 | 19.5 | 155.2 | 152.375 | 2 | 35.819 | 148.75 | 145.75 | 3.1 | 95.225 | 138.8667 | 136.245 | 4.2979 | 119.7733 | 133.05 | 130.1667
-5.7 | 1.55 | 17.1 | 157 | 154.125 | 1.65 | 23.5 | 154 | 151.32 | 2 | 46.1815 | 147.6 | 144.925 | 3.1 | 105.025 | 138 | 135.445 | 4.3978 | 121.5057 | 132.3 | 129.8
-5.6 | 1.6 | 20.6 | 155.9 | 152.925 | 1.7 | 30 | 153.0333 | 150.35 | 2.1 | 58.2017 | 146.84 | 143.925 | 3.2 | 112.4 | 137.2333 | 134.895 | 4.4977 | 122.389 | 131.45 | 129.1667
-5.5 | 1.6 | 26 | 154.3 | 151.6 | 1.7 | 39.1 | 151.725 | 149.1 | 2.1 | 72.3461 | 145.88 | 143.025 | 3.4 | 117.7 | 136.6 | 134.045 | 4.5977 | 122.9782 | 130.8 | 128.6
-5.4 | 1.6 | 34.5333 | 153.3 | 150.35 | 1.7 | 52.4 | 150.575 | 148 | 2.2 | 85.6658 | 144.9 | 141.9 | 3.4 | 120.8667 | 135.8 | 133.344 | 4.6977 | 123.263 | 130.1 | 128
-5.3 | 1.6 | 47.45 | 151.7667 | 149.0333 | 1.8 | 66.5 | 149.5 | 146.9 | 2.3 | 98.5606 | 143.8 | 141.16 | 3.5 | 122.9333 | 134.9 | 132.619 | 4.7976 | 123.338 | 129.475 | 127.2
-5.2 | 1.7 | 62.35 | 150.7 | 148.2 | 1.8 | 81.24 | 148 | 145.6 | 2.3 | 109.406 | 142.95 | 140.1 | 3.6 | 124.4 | 134.25 | 131.894 | 4.8975 | 123.338 | 128.9 | 126.8
-5.1 | 1.7 | 76.775 | 149.35 | 146.9 | 1.8 | 95.05 | 146.8 | 144.6 | 2.4 | 117.653 | 141.675 | 139.2 | 3.7 | 124.94 | 133.5 | 131.119 | 5.0975 | 123.238 | 128.3 | 126.2667
-5 | 1.7 | 92.375 | 148 | 145.8 | 1.9 | 107.975 | 145.6667 | 143.45 | 2.5 | 122.801 | 140.4 | 138.2 | 3.8 | 125.3 | 132.7 | 130.5638 | 5.1974 | 123.113 | 127.5 | 125.5
-4.9 | 1.8 | 107.2333 | 146.8667 | 144.5 | 1.9 | 117.825 | 144.4 | 142.4 | 2.6 | 125.75 | 139.6333 | 137.375 | 3.9 | 125.4 | 131.7 | 129.8763 | 5.2973 | 122.839 | 126.8 | 125.15
-4.8 | 1.9 | 117.5667 | 145.5 | 143.35 | 2 | 124.4 | 143.4 | 141.4 | 2.6 | 127.449 | 138.6 | 136.3 | 4 | 125.4 | 130.95 | 129.343 | 5.3973 | 122.589 | 126.3 | 124.7
-4.7 | 1.9 | 125.8 | 144.175 | 142.05 | 2 | 128.175 | 142.1 | 140.2 | 2.7 | 128.549 | 137.6667 | 135.6 | 4.2 | 125.3 | 130.1 | 128.343 | 5.5972 | 122.239 | 125.5 | 123.9
-4.6 | 1.9 | 130 | 142.9 | 141 | 2.1 | 130.375 | 140.8667 | 139 | 2.8 | 128.8813 | 136.425 | 134.5 | 4.3 | 125.2 | 129.75 | 127.5763 | 5.7971 | 121.839 | 125 | 123.3
-4.5 | 2 | 132.5 | 141.85 | 139.8 | 2.2 | 131.2333 | 139.7667 | 137.9 | 2.9 | 129.048 | 135.575 | 133.6 | 4.5 | 124.8333 | 128.825 | 126.9753 | 5.997 | 121.539 | 124.3 | 122.7
-4.4 | 2.1 | 133.3333 | 140.35 | 138.6 | 2.25 | 131.6 | 138.6 | 136.8 | 3 | 129.0147 | 134.6 | 132.9 | 4.6 | 124.4 | 128.05 | 126.217 | 6.1969 | 121.239 | 123.6667 | 122.15
-4.3 | 2.1 | 133.7 | 139.2 | 137.3667 | 2.3 | 131.75 | 137.5 | 135.7 | 3.15 | 128.815 | 133.7667 | 131.825 | 4.8 | 123.8 | 127.2 | 125.642 | 6.4967 | 120.59 | 123.2 | 121.5
-4.2 | 2.2 | 133.7 | 137.96 | 136.3333 | 2.4 | 131.8 | 136.5 | 134.5 | 3.3 | 128.449 | 132.6 | 130.75 | 5 | 123.4 | 126.475 | 124.9087 | 6.6966 | 120.14 | 122.7 | 121.075
-4.1 | 2.2 | 133.55 | 137.1 | 135.075 | 2.5 | 131.55 | 135.5 | 133.52 | 3.4 | 127.949 | 131.7 | 129.775 | 5.2 | 122.95 | 125.7 | 124.3743 | 6.8966 | 119.94 | 122 | 120.55
-4 | 2.3 | 133.1 | 135.8 | 133.9667 | 2.7 | 131 | 134 | 132.6 | 3.6 | 127.449 | 130.8333 | 129.275 | 5.5 | 122.3 | 125.1 | 123.5743 | 7.2964 | 119.44 | 121.55 | 120.05
-3.9 | 2.4 | 132.3 | 134.5 | 132.8 | 2.8 | 130.4 | 133.1 | 131.4 | 3.8 | 126.849 | 129.7 | 128.075 | 5.8 | 121.8 | 124.4667 | 122.941 | 7.5629 | 119.14 | 121.05 | 119.725
-3.8 | 2.5 | 131.5 | 133.3 | 131.7 | 2.9 | 129.5 | 131.825 | 130.4 | 3.9 | 126.3492 | 128.825 | 127.2 | 6.2 | 121.5 | 123.7333 | 122.291 | 7.7961 | 118.741 | 120.35 | 119.3333
-3.7 | 2.6 | 130.7333 | 132 | 130.625 | 3.1 | 128.8 | 131 | 129.25 | 4.15 | 125.425 | 127.95 | 126.4 | 6.45 | 121.1 | 123 | 121.541 | 8.5457 | 118.291 | 119.925 | 118.6
-3.6 | 2.7333 | 129.8 | 130.8 | 129.5 | 3.3 | 128.1333 | 129.6333 | 128.36 | 4.6 | 124.9 | 126.975 | 125.35 | 6.8 | 120.5 | 122.2 | 121.14 | 9.1954 | 117.866 | 119.5 | 118.0667
-3.5 | 2.9 | 128.825 | 129.85 | 128.6 | 3.6 | 127.3 | 128.8 | 127.6 | 4.8 | 124.3 | 126.1 | 124.6 | 7.5 | 120.125 | 121.7 | 120.4067 | 10.095 | 117.416 | 118.8 | 117.55
-3.4 | 3.1 | 127.8667 | 128.8 | 127.7 | 3.8 | 126.8333 | 127.7 | 126.5 | 5.15 | 123.551 | 125.4 | 123.975 | 8.1 | 119.625 | 121 | 119.8067 | 11.0945 | 117.066 | 118.4 | 117.1
-3.3 | 3.3 | 127.1 | 127.7 | 126.4 | 4.1 | 125.9333 | 127 | 125.55 | 5.7 | 122.901 | 124.5 | 123.1333 | 9 | 119.025 | 120.35 | 119.29 | 12.4938 | 116.542 | 117.95 | 116.6333
-3.2 | 3.6 | 126.275 | 126.8 | 125.5 | 4.6 | 124.975 | 125.75 | 124.7 | 6.3 | 122.251 | 123.6333 | 122.4333 | 10.0667 | 118.4667 | 119.7 | 118.74 | 14.1929 | 116.082 | 117.4333 | 116.3
-3.1 | 4.1 | 125.3 | 125.6333 | 124.6 | 5.25 | 124.025 | 124.8 | 123.7 | 7.0667 | 121.4513 | 122.8 | 121.525 | 11.5 | 117.9667 | 119.15 | 117.989 | 16.6917 | 115.642 | 116.85 | 115.7
-3 | 4.6 | 124.45 | 124.7 | 123.74 | 6 | 123.35 | 124 | 122.9 | 8.1 | 120.827 | 121.9 | 120.7 | 13.3 | 117.425 | 118.375 | 117.514 | 19.8901 | 115.242 | 116.3 | 115.2
-2.9 | 5.4 | 123.7 | 123.95 | 122.85 | 6.95 | 122.7 | 123.3 | 122.1 | 9.3 | 120.3187 | 121.1 | 120.025 | 15.8 | 116.7 | 117.95 | 116.814 | 25.0875 | 114.793 | 115.65 | 114.75
-2.8 | 6.2 | 122.85 | 123 | 122.1 | 8.25 | 121.8 | 122.2 | 121.25 | 11.15 | 119.6187 | 120.4333 | 119.3 | 19.2 | 116.3 | 117.175 | 116.339 | 31.8841 | 114.343 | 115.175 | 114.325
-2.7 | 7.6 | 122 | 122.3 | 121.3 | 10.1 | 121.1 | 121.3 | 120.65 | 13.5 | 118.8025 | 119.7 | 118.5 | 24.4 | 115.85 | 116.65 | 115.5383 | 40.1799 | 113.818 | 114.55 | 113.625
-2.6 | 9.5 | 121.2 | 121.3 | 120.2 | 12.7333 | 120.15 | 120.5 | 119.625 | 17.0333 | 118.153 | 118.9 | 117.825 | 31.5 | 115.4333 | 115.9 | 115.0047 | 50.6747 | 113.493 | 114.08 | 113.25
-2.5 | 12.3 | 120.55 | 120.2667 | 119.6333 | 16 | 119.4 | 119.7 | 118.85 | 22.1 | 117.528 | 118.1 | 117.2 | 41.1 | 114.9 | 115.3 | 114.438 | 61.0695 | 112.8773 | 113.5 | 112.7
-2.4 | 16.7 | 119.5 | 119.3 | 118.6 | 21.85 | 118.6333 | 118.9667 | 117.9 | 29.6 | 116.953 | 117.4 | 116.4 | 52.3 | 114.5 | 114.6 | 113.938 | 72.1639 | 112.3107 | 112.9333 | 112.2
-2.3 | 23.3 | 118.8 | 118.3 | 117.9 | 29.6 | 118 | 118.1 | 117.2 | 39.8 | 116.1287 | 116.625 | 115.75 | 63.9 | 113.8667 | 114 | 113.338 | 81.5093 | 111.744 | 112.4 | 111.75
-2.2 | 33.7 | 117.7 | 117.65 | 116.8 | 40.8 | 117.4 | 117.2 | 116.5 | 51.5 | 115.404 | 115.825 | 115.1 | 74.7 | 113.1333 | 113.3333 | 112.838 | 89.4553 | 111.369 | 111.7 | 111.35
-2.1 | 44.8 | 116.9667 | 116.95 | 116.2 | 52.5 | 116.3 | 116.275 | 115.7 | 63.8 | 114.654 | 114.9 | 114.3 | 84.9 | 112.5333 | 112.7 | 112.1703 | 94.7526 | 111.0445 | 111.25 | 110.6
-2 | 57.5 | 115.875 | 115.9 | 115.35 | 65.3 | 115.35 | 115.5 | 114.8667 | 75.5 | 113.904 | 114.3 | 113.5 | 91.9 | 111.7667 | 112.1 | 111.6037 | 98.7506 | 110.505 | 110.75 | 110.18
-1.9 | 70.9 | 115.0333 | 115.1 | 114.6 | 77.6 | 114.58 | 114.7 | 114.05 | 84.6 | 113.255 | 113.6 | 112.725 | 97.2 | 111.2333 | 111.45 | 110.837 | 101.349 | 109.895 | 110.2 | 109.6
-1.8 | 82.35 | 114.3 | 114.2 | 113.6 | 87.6 | 113.85 | 113.9 | 113.3 | 92.75 | 112.4217 | 112.8 | 111.975 | 100.6 | 110.5333 | 110.7 | 110.237 | 102.649 | 109.445 | 109.6 | 109.1667
-1.7 | 91.2 | 113.5 | 113.3333 | 112.975 | 95.2 | 112.9667 | 113 | 112.5 | 98.15 | 111.7883 | 112 | 111.25 | 102.9 | 110.0667 | 110.1 | 109.537 | 103.648 | 108.846 | 109.15 | 108.4667
-1.6 | 98.3 | 112.5667 | 112.375 | 112.0333 | 99.85 | 112.475 | 112.2 | 111.7 | 101.25 | 111.256 | 111.25 | 110.575 | 104.15 | 109.575 | 109.55 | 109.086 | 104.398 | 108.4793 | 108.5 | 108.05
-1.5 | 102.1 | 111.8333 | 111.6 | 111.3 | 102.6 | 111.5667 | 111.425 | 110.9667 | 103.1667 | 110.306 | 110.5 | 110 | 104.85 | 109.05 | 108.88 | 108.5693 | 104.648 | 107.946 | 107.9 | 107.5667
-1.4 | 104.1 | 110.9 | 110.8 | 110.4 | 104.2 | 110.8 | 110.525 | 110.1667 | 104.3 | 109.756 | 109.8667 | 109.2 | 105.2 | 108.375 | 108.425 | 107.9027 | 104.748 | 107.396 | 107.4333 | 107.1
-1.3 | 105.3 | 110.2 | 110.1 | 109.625 | 105.2 | 109.95 | 109.7 | 109.45 | 105 | 109.056 | 109.075 | 108.6 | 105.35 | 107.5333 | 107.7 | 107.336 | 104.848 | 106.9465 | 106.9 | 106.65
-1.2 | 105.7 | 109.3 | 109.2 | 108.9667 | 105.4 | 109.3 | 109.05 | 108.7 | 105.3 | 108.257 | 108.3 | 107.85 | 105.4 | 107.1667 | 107.05 | 106.736 | 104.848 | 106.447 | 106.45 | 106.1333
-1.1 | 105.85 | 108.6 | 108.4 | 108.2 | 105.5 | 108.45 | 108.2 | 108 | 105.3 | 107.607 | 107.65 | 107.2 | 105.3 | 106.4 | 106.4 | 106.035 | 104.648 | 105.847 | 105.7333 | 105.5667
-1 | 105.8 | 107.875 | 107.6 | 107.375 | 105.5 | 107.6667 | 107.5 | 107.1667 | 105.2 | 106.8903 | 106.8 | 106.45 | 105 | 105.8 | 105.825 | 105.535 | 104.448 | 105.347 | 105.2 | 105.1
-0.9 | 105.6 | 106.9667 | 106.725 | 106.5 | 105.3 | 106.8 | 106.6 | 106.4 | 105 | 106.2823 | 106.25 | 105.9 | 104.8 | 105.2 | 105.3 | 105.0683 | 104.098 | 104.698 | 104.7 | 104.5
-0.8 | 105.15 | 105.95 | 105.925 | 105.7 | 104.95 | 106.1 | 106.0667 | 105.7 | 104.7 | 105.458 | 105.5667 | 105.1 | 104.4 | 104.575 | 104.75 | 104.335 | 103.548 | 104.123 | 104.3 | 104.05
-0.7 | 104.7 | 105.375 | 105.125 | 105.1 | 104.5 | 105.2 | 105.025 | 104.9 | 104.1 | 104.708 | 104.8667 | 104.5333 | 103.8 | 104.05 | 104.1 | 104.035 | 103.048 | 103.673 | 103.675 | 103.5
-0.6 | 104.2 | 104.425 | 104.275 | 104.35 | 103.9 | 104.3667 | 104.2 | 104.1 | 103.6 | 104.058 | 104.1 | 103.75 | 103.4 | 103.55 | 103.4 | 103.3673 | 102.749 | 102.9988 | 103.125 | 103.08
-0.5 | 103.5 | 103.65 | 103.6 | 103.4 | 103.6333 | 103.7 | 103.6 | 103.45 | 103.1 | 103.399 | 103.2 | 103.1 | 102.9 | 102.825 | 102.675 | 102.834 | 102.449 | 102.4823 | 102.45 | 102.54
-0.4 | 102.7 | 103.02 | 102.7 | 102.825 | 102.9 | 103.0667 | 102.8 | 102.85 | 102.3667 | 102.759 | 102.75 | 102.45 | 102.45 | 102.175 | 102.25 | 102.134 | 101.849 | 101.949 | 102.075 | 102.175
-0.3 | 102 | 102.3333 | 102.2 | 102.15 | 102.1 | 102.24 | 102.1 | 102.2 | 101.75 | 101.959 | 102.0667 | 101.7 | 101.95 | 101.625 | 101.65 | 101.584 | 101.349 | 101.399 | 101.6 | 101.675
-0.2 | 101.2 | 101.5 | 101.3 | 101.425 | 101.4 | 101.575 | 101.4 | 101.45 | 101.2 | 101.359 | 101.4 | 101.125 | 101.25 | 101.15 | 101.05 | 100.984 | 100.85 | 100.85 | 100.95 | 101.05
-0.1 | 100.6333 | 100.8 | 100.75 | 100.75 | 100.7333 | 100.8 | 100.625 | 100.9 | 100.7 | 100.6267 | 100.7667 | 100.55 | 100.6333 | 100.8 | 100.55 | 100.559 | 100.2833 | 100.55 | 100.5 | 100.575
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9998 | 100 | 100 | 100 | 100
0.1 | 99.4 | 99.3 | 99.05 | 99.4 | 99.5 | 99.52 | 99.4 | 99.3 | 99.4 | 99.327 | 99.2 | 99.32 | 99.4 | 99.25 | 99.35 | 99.5732 | 99.2004 | 99.4303 | 99.6 | 99.6
0.2 | 98.3 | 98.475 | 98.4 | 98.76 | 98.85 | 98.6 | 98.6 | 98.7 | 98.9 | 98.6605 | 98.65 | 98.8 | 98.8 | 98.9 | 98.7 | 98.833 | 98.7006 | 98.8506 | 99 | 99.2
0.3 | 97.6 | 97.825 | 97.8 | 98.2 | 98.05 | 97.9 | 97.9 | 98.2 | 98.1 | 98.0941 | 98.1 | 98.1 | 98.2 | 98.2667 | 98.15 | 98.1327 | 98.1509 | 98.5007 | 98.3667 | 98.55
0.4 | 97.1 | 97.25 | 97.05 | 97.45 | 97.3 | 97.4 | 97.35 | 97.4667 | 97.3 | 97.4111 | 97.4333 | 97.6 | 97.6 | 97.6 | 97.6 | 97.7326 | 97.6012 | 98.001 | 98.025 | 98.1
0.5 | 96.35 | 96.3 | 96.4 | 96.9 | 96.4 | 96.7333 | 96.7 | 96.75 | 96.7 | 96.9013 | 96.95 | 96.85 | 97 | 96.9333 | 97 | 97.2324 | 97.1514 | 97.5013 | 97.55 | 97.8333
0.6 | 95.65 | 95.8333 | 95.825 | 96.2 | 95.4 | 95.8 | 96.05 | 96.225 | 95.95 | 96.1865 | 96.3333 | 96.1667 | 96.2 | 96.4667 | 96.5 | 96.6822 | 96.5517 | 96.9515 | 96.875 | 97.2667
0.7 | 94.6 | 95.1 | 95.15 | 95.5667 | 94.7 | 95.2667 | 95.525 | 95.675 | 95 | 95.5368 | 95.6 | 95.6 | 95.4 | 96 | 96.05 | 96.182 | 95.5522 | 96.1852 | 96.4333 | 96.85
0.8 | 94.1 | 94.3333 | 94.525 | 94.95 | 94 | 94.82 | 94.8333 | 94.925 | 94.2 | 94.8371 | 95 | 95.1 | 95 | 95.4 | 95.45 | 95.7819 | 95.2524 | 95.8771 | 95.9 | 96.2667
0.9 | 92.8 | 93.9 | 93.85 | 94.3 | 93 | 94.125 | 94.1 | 94.425 | 93.3 | 94.2623 | 94.4 | 94.6333 | 94.45 | 94.7667 | 94.9667 | 95.1317 | 94.8026 | 95.5522 | 95.275 | 95.8333
1 | 91.9 | 93.0667 | 93.325 | 93.825 | 92.3 | 93.45 | 93.45 | 93.75 | 92.75 | 93.8625 | 93.9 | 94 | 93.5333 | 94.4 | 94.3333 | 94.6065 | 93.9031 | 95.0025 | 94.8 | 95.4333
1.1 | 90.7 | 92.475 | 92.6 | 93.1 | 91 | 92.825 | 93 | 93.3 | 91.6 | 92.9628 | 93.2 | 93.4 | 92.7 | 93.65 | 93.7667 | 94.2314 | 93.4532 | 94.5528 | 94.475 | 94.8333
1.2 | 89.5 | 91.9333 | 92 | 92.45 | 89.7 | 92 | 92.3 | 92.6 | 90.35 | 92.563 | 92.6 | 92.85 | 91.9 | 93.15 | 93.2 | 93.5312 | 92.5537 | 93.8531 | 93.9667 | 94.425
1.3 | 88.1 | 91.3 | 91.45 | 91.95 | 88.2 | 91.6667 | 91.65 | 92.0667 | 89.4 | 91.8832 | 92.1 | 92.3333 | 91 | 92.55 | 92.7 | 93.2644 | 91.5542 | 93.4532 | 93.3667 | 94.05
1.4 | 86.1 | 90.5 | 90.7 | 91.3333 | 86.7 | 90.9667 | 90.875 | 91.45 | 87.8333 | 91.2635 | 91.7 | 91.7333 | 89.8 | 92.05 | 92.0667 | 92.6642 | 90.7546 | 93.0535 | 93 | 93.7
1.5 | 83.7 | 89.94 | 90.1 | 90.7667 | 84.5 | 90.2667 | 90.5 | 90.8667 | 86 | 90.7137 | 91.05 | 91.2 | 88.6 | 91.5 | 91.55 | 92.1307 | 89.8551 | 92.4288 | 92.6 | 93.1
1.6 | 80.2 | 89.25 | 89.6 | 90.1667 | 81.2 | 89.6 | 89.85 | 90.4 | 83.6 | 89.939 | 90.45 | 90.7 | 87 | 91 | 91.2 | 91.4805 | 88.6057 | 91.9041 | 92.1 | 92.6
1.7 | 75.4 | 88.775 | 89.05 | 89.5 | 76.95 | 89 | 89.3 | 89.8333 | 80.7 | 89.4393 | 89.875 | 90.05 | 85.25 | 90.6 | 90.5667 | 91.1304 | 87.3563 | 91.3877 | 91.65 | 92.04
1.8 | 67.9 | 88.05 | 88.4 | 89.1 | 71.2 | 88.3 | 88.7 | 89.15 | 76.9 | 88.9644 | 89.25 | 89.5 | 83 | 89.9 | 90 | 90.5969 | 86.1569 | 90.8046 | 91 | 91.575
1.9 | 59.1 | 87.375 | 87.85 | 88.5 | 63.0667 | 87.65 | 88.2 | 88.55 | 71.4 | 88.2647 | 88.75 | 89 | 80.15 | 89.3 | 89.6 | 90.03 | 84.9575 | 90.1549 | 90.6333 | 91.25
2 | 49.2 | 86.75 | 87.3333 | 87.95 | 54.5 | 87.2 | 87.55 | 87.925 | 64.2 | 87.7316 | 88.1333 | 88.5 | 76 | 88.65 | 89.0333 | 89.5299 | 82.5587 | 89.6218 | 90.25 | 90.85
2.1 | 39.7 | 86.15 | 86.65 | 87.26 | 44.6 | 86.4667 | 86.95 | 87.425 | 55.2 | 87.0985 | 87.5333 | 87.9 | 70.3 | 88.15 | 88.46 | 89.1297 | 79.3603 | 89.1304 | 89.725 | 90.275
2.2 | 29.7 | 85.7 | 85.98 | 86.725 | 35.2 | 85.925 | 86.3667 | 86.9667 | 46 | 86.5154 | 86.8333 | 87.44 | 62.9 | 87.55 | 87.925 | 88.6629 | 75.6622 | 88.7306 | 89.175 | 89.95
2.3 | 21.9 | 84.98 | 85.45 | 86.25 | 26.4 | 85.35 | 85.775 | 86.3 | 36.3667 | 85.9156 | 86.35 | 86.85 | 54.4 | 87.075 | 87.45 | 88.0627 | 70.3149 | 88.031 | 88.575 | 89.42
2.4 | 15.6 | 84.4333 | 84.95 | 85.75 | 18.8 | 84.75 | 85.3 | 85.8 | 28.4 | 85.3659 | 85.9 | 86.35 | 46.05 | 86.5 | 87.1333 | 87.5792 | 63.1684 | 87.6362 | 88.08 | 88.9667
2.5 | 11.4 | 83.85 | 84.5 | 85.18 | 14.25 | 84.15 | 84.65 | 85.15 | 21.4 | 84.7161 | 85.4333 | 85.725 | 37.3 | 85.9667 | 86.5 | 87.0623 | 55.5722 | 87.0232 | 87.7333 | 88.3667
2.6 | 8.8 | 83.25 | 83.9667 | 84.6 | 10.7 | 83.4667 | 84.25 | 84.6667 | 16.1 | 84.1663 | 84.8 | 85.3 | 29.3 | 85.4 | 85.925 | 86.5789 | 47.7761 | 86.6567 | 87.2667 | 88.0333
2.7 | 6.8 | 82.7 | 83.25 | 84.05 | 8.6 | 83 | 83.6333 | 84.1667 | 12.6 | 83.4666 | 84.325 | 84.8 | 22.55 | 84.8 | 85.5667 | 86.1287 | 38.3808 | 86.2569 | 86.675 | 87.6
2.8 | 5.5 | 82.2 | 82.8 | 83.5 | 6.95 | 82.5 | 83.2 | 83.8667 | 10.1 | 82.9668 | 83.7667 | 84.25 | 17.6 | 84.2333 | 85 | 85.5285 | 30.8346 | 85.6072 | 86.2 | 87.0667
2.9 | 4.65 | 81.6333 | 82.4 | 83.1667 | 5.85 | 81.8 | 82.7333 | 83.3 | 8.6 | 82.347 | 83.2333 | 83.85 | 14.4 | 83.7333 | 84.5 | 85.1283 | 24.4877 | 84.9575 | 85.8 | 86.6
3 | 3.9 | 81 | 81.9 | 82.65 | 5.1333 | 81.4 | 82.075 | 82.75 | 7.3 | 81.7673 | 82.9 | 83.3 | 12 | 83.025 | 84 | 84.6282 | 19.6901 | 84.5911 | 85.3 | 86.3
3.1 | 3.5333 | 80.525 | 81.3 | 82.1333 | 4.4 | 80.7 | 81.75 | 82.3 | 6.3 | 81.3175 | 82.2 | 82.8333 | 10.4 | 82.5 | 83.45 | 84.3281 | 16.1919 | 84.0246 | 84.8 | 85.8
3.2 | 3.15 | 79.875 | 80.94 | 81.6333 | 3.9 | 80.2333 | 81.2 | 81.8333 | 5.6333 | 80.7427 | 81.85 | 82.3333 | 9.1 | 82.0333 | 83.05 | 83.8279 | 13.6932 | 83.7082 | 84.3 | 85.2
3.3 | 2.85 | 79.4333 | 80.475 | 81.3 | 3.45 | 79.6 | 80.8 | 81.4333 | 5.05 | 80.2179 | 81.3667 | 81.9 | 7.9 | 81.4 | 82.4667 | 83.2944 | 11.8941 | 83.1334 | 83.8667 | 84.85
3.4 | 2.55 | 78.9 | 80 | 80.8 | 3.15 | 79 | 80.4 | 80.85 | 4.65 | 79.6432 | 81 | 81.375 | 7.25 | 80.95 | 82.1 | 82.8026 | 10.6947 | 82.6837 | 83.5333 | 84.45
3.5 | 2.3 | 78.28 | 79.4333 | 80.3667 | 2.9 | 78.4667 | 79.8 | 80.45 | 4.2 | 79.0434 | 80.45 | 81.05 | 6.6 | 80.5 | 81.7 | 82.4675 | 9.4453 | 82.059 | 83.0667 | 84.075
3.6 | 2.2 | 77.775 | 79.0667 | 79.85 | 2.7 | 77.8333 | 79.4 | 80.0333 | 3.9 | 78.4487 | 79.9 | 80.65 | 6.1 | 79.95 | 81.2 | 81.9773 | 8.5457 | 81.4593 | 82.6 | 83.65
3.7 | 2.1 | 77.1 | 78.6 | 79.55 | 2.5 | 77.35 | 78.85 | 79.6 | 3.65 | 77.9355 | 79.6 | 80.1 | 5.8 | 79.4 | 80.85 | 81.5772 | 8.096 | 81.0594 | 82.24 | 83.15
3.8 | 2 | 76.5 | 78.15 | 79.05 | 2.3333 | 76.7333 | 78.475 | 79.1333 | 3.4 | 77.3691 | 79.12 | 79.6 | 5.5 | 78.75 | 80.35 | 81.1604 | 7.2964 | 80.5597 | 81.85 | 82.7
3.9 | 1.9 | 75.9 | 77.75 | 78.625 | 2.2 | 76.125 | 78.05 | 78.65 | 3.2 | 76.7026 | 78.7333 | 79.15 | 5.1 | 78.25 | 79.8667 | 80.5935 | 6.9465 | 80.0599 | 81.4333 | 82.25
4 | 1.8 | 75.2 | 77.15 | 78.075 | 2.1 | 75.3 | 77.45 | 78.225 | 3.1 | 76.1495 | 78.25 | 78.825 | 4.9 | 77.8 | 79.65 | 80.2267 | 6.6467 | 79.4603 | 81.05 | 81.9
4.1 | 1.7 | 74.46 | 76.8333 | 77.7 | 2 | 74.7 | 77.05 | 77.825 | 2.9 | 75.4698 | 77.8667 | 78.35 | 4.6 | 77.2 | 79.15 | 79.8766 | 6.3468 | 78.8805 | 80.6 | 81.5333
4.2 | 1.7 | 73.6333 | 76.3 | 77.3 | 1.9 | 73.95 | 76.6333 | 77.25 | 2.8 | 74.7201 | 77.325 | 77.95 | 4.4 | 76.6 | 78.5 | 79.4765 | 6.0969 | 78.4008 | 80.2333 | 81.1667
4.3 | 1.6 | 72.7 | 76 | 76.8 | 1.8 | 73.1 | 76.15 | 76.8 | 2.6 | 74.1703 | 76.9 | 77.5 | 4.3 | 75.9 | 78.2667 | 78.993 | 5.8471 | 77.8278 | 79.75 | 80.8667
4.4 | 1.5 | 71.54 | 75.55 | 76.4 | 1.7667 | 72 | 75.7 | 76.4 | 2.5 | 73.3207 | 76.3 | 76.9 | 4.1333 | 75.42 | 77.8333 | 78.5262 | 5.6638 | 77.5612 | 79.25 | 80.3
4.5 | 1.5 | 69.8 | 75 | 75.9 | 1.7 | 70.9 | 75.3 | 76 | 2.4 | 72.3377 | 75.8333 | 76.6 | 4.05 | 74.925 | 77.3 | 78.0593 | 5.3973 | 76.9615 | 78.85 | 80.1
4.6 | 1.45 | 68 | 74.45 | 75.44 | 1.6 | 69.4 | 74.8333 | 75.55 | 2.35 | 71.4714 | 75.4333 | 76.1 | 3.85 | 74.075 | 76.9 | 77.7759 | 5.1974 | 76.2619 | 78.5 | 79.65
4.7 | 1.4 | 65.2 | 74.125 | 75.02 | 1.5 | 67.4667 | 74.4667 | 75 | 2.2 | 70.3719 | 75 | 75.7 | 3.7 | 73.5 | 76.4667 | 77.3258 | 4.9975 | 76.012 | 78.25 | 79.1333
4.8 | 1.4 | 61.4 | 73.55 | 74.55 | 1.5 | 64.9 | 74.1 | 74.775 | 2.2 | 69.0723 | 74.5 | 75.2 | 3.6 | 72.8 | 76.0667 | 76.8256 | 4.8975 | 75.1874 | 77.7667 | 78.725
4.9 | 1.3 | 55.6333 | 73.025 | 74.075 | 1.5 | 61.4 | 73.5 | 74.275 | 2.1 | 67.5397 | 74.05 | 74.8 | 3.4 | 72.1 | 75.5667 | 76.4255 | 4.7976 | 74.7126 | 77.4 | 78.25
5 | 1.3 | 48.625 | 72.6 | 73.65 | 1.4 | 56.325 | 73.05 | 73.675 | 2 | 65.6737 | 73.65 | 74.35 | 3.3 | 71.3 | 75.1 | 76.0504 | 4.6977 | 74.0297 | 76.95 | 77.85
5.1 | 1.3 | 40.8333 | 72.2333 | 73.175 | 1.4 | 50.125 | 72.6 | 73.3 | 2 | 62.575 | 73.15 | 73.9333 | 3.2 | 70.45 | 74.575 | 75.5502 | 4.4977 | 73.5299 | 76.375 | 77.425
5.2 | 1.3 | 32.925 | 71.75 | 72.7667 | 1.4 | 42.9333 | 72.15 | 72.8667 | 1.9 | 58.4766 | 72.6333 | 73.46 | 3.2 | 69.2 | 74.225 | 75.075 | 4.3978 | 72.8136 | 76 | 77.0333
5.3 | 1.2 | 26.3 | 71.3333 | 72.3333 | 1.3 | 35.5 | 71.65 | 72.4 | 1.8667 | 53.1288 | 72.1667 | 73.125 | 3.1 | 67.7 | 73.75 | 74.6499 | 4.2979 | 72.2639 | 75.6 | 76.6
5.4 | 1.2 | 19.7 | 70.9 | 71.8333 | 1.3 | 28.34 | 71.05 | 71.95 | 1.8 | 46.7146 | 71.8 | 72.525 | 3 | 66.4 | 73.35 | 74.2748 | 4.1979 | 71.4643 | 75.1 | 76.25
5.5 | 1.2 | 15.3 | 70.4 | 71.3667 | 1.3 | 22.5 | 70.725 | 71.6 | 1.7 | 39.984 | 71.375 | 72 | 2.9 | 64.375 | 72.95 | 73.7746 | 4.0979 | 70.5147 | 74.6 | 75.8
5.6 | 1.1 | 12.3 | 69.95 | 71 | 1.3 | 17.35 | 70.15 | 70.95 | 1.7 | 32.7869 | 70.8667 | 71.65 | 2.8 | 61.725 | 72.42 | 73.4245 | 3.998 | 69.6652 | 74.2333 | 75.4
5.7 | 1.1 | 10.3 | 69.3667 | 70.48 | 1.2 | 14 | 69.7 | 70.5667 | 1.7 | 26.6643 | 70.5 | 71.1667 | 2.7333 | 57.95 | 72 | 73.0743 | 3.8981 | 68.6407 | 73.8 | 75.1333
5.8 | 1.1 | 8.9 | 68.9 | 69.9667 | 1.2 | 11.65 | 69.2667 | 70 | 1.6 | 21.5164 | 69.95 | 70.6 | 2.65 | 53 | 71.5 | 72.5242 | 3.7981 | 67.4663 | 73.4333 | 74.5667
5.9 | 1.1 | 7.94 | 68.5333 | 69.4333 | 1.2 | 10.1 | 68.7 | 69.6667 | 1.5 | 17.3931 | 69.45 | 70.25 | 2.55 | 47.225 | 71.1 | 71.974 | 3.6982 | 65.6672 | 73.05 | 74.2
6 | 1.1 | 7.26 | 68 | 68.9 | 1.1 | 8.9 | 68.15 | 69.1 | 1.5 | 14.5609 | 68.9667 | 69.7667 | 2.4 | 40.95 | 70.75 | 71.5238 | 3.5482 | 63.968 | 72.55 | 73.9333
6.1 | 1.0333 | 6.75 | 67.45 | 68.4667 | 1.1 | 8.1 | 67.7333 | 68.7 | 1.5 | 12.7199 | 68.5 | 69.3 | 2.4 | 34.675 | 70.15 | 71.2237 | 3.4983 | 61.2028 | 72.2 | 73.5
6.2 | 1 | 6.2 | 66.95 | 67.9333 | 1.1 | 7.45 | 67.25 | 67.9667 | 1.5 | 11.3205 | 68 | 68.8 | 2.4 | 29 | 69.675 | 70.9236 | 3.4983 | 57.7711 | 71.6667 | 72.8667
6.3 | 1 | 5.875 | 66.4 | 67.45 | 1.1 | 6.9 | 66.7667 | 67.5667 | 1.4 | 10.3458 | 67.4 | 68.5 | 2.4 | 23.825 | 69.225 | 70.2234 | 3.3983 | 53.1734 | 71.3 | 72.575
6.4 | 1 | 5.625 | 65.8 | 66.9 | 1 | 6.4667 | 66.1 | 67.1 | 1.4 | 9.4962 | 66.8 | 67.8 | 2.3 | 20.0667 | 68.85 | 69.7899 | 3.3983 | 47.8095 | 70.85 | 72.1
6.5 | 1 | 5.4 | 65.3 | 66.275 | 1 | 6.2 | 65.6 | 66.525 | 1.4 | 8.8964 | 66.3667 | 67.22 | 2.2 | 17.6 | 68.35 | 69.4231 | 3.3483 | 42.3788 | 70.45 | 71.7
6.6 | 1 | 5.175 | 64.7333 | 65.7667 | 1 | 5.8333 | 65.1 | 65.875 | 1.3 | 8.3467 | 65.85 | 66.74 | 2.2 | 15.6 | 67.85 | 68.8229 | 3.2484 | 36.5318 | 69.9 | 71.2
6.7 | 1 | 5.0333 | 63.975 | 65.04 | 1 | 5.6 | 64.4667 | 65.325 | 1.3 | 7.9968 | 65.35 | 66.2333 | 2.1 | 14.15 | 67.4333 | 68.4478 | 3.1484 | 30.9845 | 69.55 | 70.75
6.8 | 1 | 4.9 | 63.25 | 64.275 | 1 | 5.325 | 63.65 | 64.65 | 1.3 | 7.5636 | 64.75 | 65.725 | 2.1 | 13 | 66.8333 | 67.9727 | 3.0985 | 26.3869 | 69.2 | 70.4
6.9 | 0.9 | 4.7 | 62.3 | 63.5 | 1 | 5.175 | 63.04 | 63.925 | 1.3 | 7.2721 | 64.2 | 65.05 | 2 | 12.1 | 66.4 | 67.4725 | 2.9985 | 22.3888 | 68.55 | 69.9
7 | 0.9 | 4.5333 | 61.275 | 62.35 | 0.9 | 5 | 62.175 | 63 | 1.2 | 6.9472 | 63.625 | 64.625 | 2 | 11.5 | 65.9333 | 66.9023 | 2.8986 | 19.7901 | 68.15 | 69.4
7.1 |  | 4.45 | 60.0333 | 61.1 | 0.9 | 4.8 | 61.2 | 62.275 | 1.2 | 6.6723 | 62.875 | 63.92 | 1.95 | 11.0333 | 65.3 | 66.5722 | 2.8986 | 17.8911 | 67.75 | 69.075
7.2 |  | 4.3 | 58 | 59.2 |  | 4.6 | 60.2 | 61.2 |  | 6.4224 | 62.225 | 63.2 | 1.9 | 10.5333 | 64.8 | 65.922 | 2.8986 | 16.4584 | 67.25 | 68.65
7.3 |  | 4.3 | 55.6 | 56.7 |  | 4.5 | 58.5667 | 59.65 |  | 6.2308 | 61.55 | 62.45 | 1.9 | 10.15 | 64.3 | 65.4218 | 2.8319 | 15.3724 | 66.875 | 68.25
7.4 |  | 4.2 | 52 | 52.75 |  | 4.45 | 56.8 | 58 |  | 6.0726 | 60.6 | 61.6 |  | 9.775 | 63.9 | 64.9216 | 2.7986 | 14.4178 | 66.48 | 67.75
7.5 |  | 4.0667 | 47 | 47.9 |  | 4.3 | 54.12 | 55.5667 |  | 5.7977 | 59.75 | 60.6 |  | 9.525 | 63.26 | 64.5215 | 2.7986 | 13.7681 | 66.1 | 67.3
7.6 |  | 3.975 | 41.2 | 42.3 |  | 4.2 | 50.6667 | 51.9 |  | 5.6777 | 58.6 | 59.5 |  | 9.175 | 62.75 | 63.8546 | 2.5987 | 13.1601 | 65.58 | 66.7667
7.7 |  | 3.9 | 34.8 | 36.5 |  | 4.1 | 45.8 | 47.1333 |  | 5.5978 | 57.1 | 58.1667 |  | 8.9333 | 62.175 | 63.4211 |  | 12.727 | 65.1 | 66.3
7.8 |  | 3.8 | 28.7 | 30.775 |  | 4.1 | 40.2 | 41.8333 |  | 5.3978 | 55.2 | 56.3 |  | 8.58 | 61.6333 | 62.7876 |  | 12.2439 | 64.55 | 65.9
7.9 |  | 3.8 | 22.8667 | 25.225 |  | 3.9 | 34 | 36.2 |  | 5.2229 | 52.5 | 53.5333 |  | 8.45 | 60.9 | 62.0207 |  | 11.9274 | 64.05 | 65.4
8 |  | 3.7 | 18.425 | 20.85 |  | 3.9 | 28.05 | 30.875 |  | 5.098 | 48.9 | 49.9667 |  | 8.3 | 60.225 | 61.4705 |  | 11.6275 | 63.6667 | 65
8.1 |  | 3.7 | 14.8667 | 17.875 |  | 3.8 | 22.7 | 25.625 |  | 4.998 | 44.2667 | 45.35 |  | 7.95 | 59.475 | 60.6202 |  | 11.2611 | 63.1 | 64.45
8.2 |  | 3.65 | 12.7 | 16 |  | 3.7 | 18.54 | 21.3 |  | 4.8314 | 38.7 | 40.1 |  | 7.8 | 58.7 | 59.7449 |  | 11.0945 | 62.475 | 64.05
8.3 |  | 3.6 | 11.2 | 14.625 |  | 3.7 | 15.15 | 18.3 |  | 4.6981 | 33.2 | 34.98 |  | 7.66 | 57.6 | 58.9196 |  | 10.7946 | 62.125 | 63.55
8.4 |  | 3.6 | 10.2 | 13.8 |  | 3.6 | 13 | 16.3 |  | 4.5982 | 27.7333 | 30 |  | 7.45 | 56.3 | 57.5992 |  | 10.4948 | 61.525 | 62.875
8.5 |  | 3.6 | 9.5 | 13.2 |  | 3.6 | 11.525 | 15 |  | 4.5982 | 23.0667 | 25.35 |  | 7.32 | 54.8 | 56.2188 |  | 10.2949 | 60.975 | 62.35
8.6 |  | 3.6 | 9.0667 | 12.6667 |  | 3.6 | 10.625 | 14.1 |  | 4.4982 | 18.975 | 21.7 |  | 7.175 | 52.45 | 54.3181 |  | 9.995 | 60.35 | 61.8333
8.7 |  | 3.5 | 8.6333 | 12.25 |  | 3.5 | 9.8667 | 13.4 |  | 4.3982 | 16.4 | 19.3 |  | 6.925 | 49.6 | 51.4171 |  | 9.7951 | 59.65 | 61.1
8.8 |  | 3.5 | 8.35 | 12 |  | 3.5 | 9.4 | 12.8 |  | 4.2983 | 14.45 | 17.5 |  | 6.8 | 46 | 47.7159 |  | 9.5952 | 59.075 | 60.3667
8.9 |  | 3.5 | 8.1 | 11.7667 |  | 3.5 | 8.9 | 12.35 |  | 4.2983 | 13.1 | 16.475 |  | 6.6667 | 41.475 | 43.8546 |  | 9.4203 | 58.2 | 59.7
9 |  | 3.4 | 7.925 | 11.6 |  | 3.5 | 8.5667 | 12.0333 |  | 4.2483 | 12.15 | 15.6 |  | 6.575 | 36.8 | 39.2631 |  | 9.2454 | 57.35 | 58.8
9.1 |  | 3.4 | 7.725 | 11.4 |  | 3.4 | 8.2333 | 11.7 |  | 4.1983 | 11.4 | 14.875 |  | 6.34 | 31.6 | 34.5865 |  | 8.9955 | 56.2 | 57.7
9.2 |  | 3.4 | 7.6 | 11.3 |  | 3.4 | 8.0667 | 11.45 |  | 4.1233 | 10.7 | 14.35 |  | 6.2 | 26.9 | 30.3101 |  | 8.7956 | 54.8 | 56.35
9.3 |  | 3.3 | 7.6 | 11.3 |  | 3.4 | 7.8 | 11.3 |  | 4.0984 | 10.4 | 13.875 |  | 6.1 | 23.15 | 26.6339 |  | 8.6207 | 52.9 | 54.7
9.4 |  | 3.3 | 7.5 | 11.175 |  | 3.3667 | 7.7 | 11.1 |  | 4.0984 | 9.975 | 13.525 |  | 6 | 20.1333 | 23.6829 |  | 8.4708 | 50.6 | 52.5
9.5 |  | 3.3 | 7.5 | 11.1 |  | 3.3 | 7.6 | 11 |  | 4.0984 | 9.65 | 13.225 |  | 5.9 | 18 | 21.6822 |  | 8.2958 | 48 | 49.35
9.6 |  | 3.2 | 7.4 | 11.1 |  | 3.3 | 7.5 | 10.9 |  | 3.9984 | 9.4 | 12.975 |  | 5.9 | 16.38 | 20.1818 |  | 8.1959 | 44.1 | 46.225
9.7 |  | 3.2 | 7.4 | 11 |  | 3.2 | 7.4 | 10.8 |  | 3.9234 | 9.1 | 12.7 |  | 5.8 | 15.32 | 19.2064 |  | 7.996 | 39.7333 | 42.15
9.8 |  | 3.1 | 7.4 | 11 |  | 3.2 | 7.3 | 10.7 |  | 3.8984 | 8.9 | 12.5 |  | 5.7 | 14.4 | 18.4061 |  | 7.896 | 35.3667 | 37.9
9.9 |  | 3.1 | 7.4 | 11 |  | 3.1 | 7.3 | 10.7 |  | 3.8235 | 8.7 | 12.35 |  | 5.6 | 13.8 | 17.7559 |  | 7.7961 | 31.125 | 33.775
10 |  | 3.1 | 7.4 | 10.9 |  | 3.1 | 7.2 | 10.6 |  | 3.7985 | 8.625 | 12.2 |  | 5.5 | 13.3 | 17.2057 |  | 7.6962 | 27.3 | 30.3
10.1 |  | 3.0333 | 7.325 | 10.9 |  | 3.1 | 7.2 | 10.6 |  | 3.6985 | 8.5 | 12 |  | 5.4 | 12.8333 | 16.8056 |  | 7.5962 | 23.8 | 27.3
10.2 |  |  | 7.3 | 10.9 |  | 3 | 7.1 | 10.5 |  | 3.6985 | 8.4 | 11.9 |  | 5.4 | 12.45 | 16.4055 |  | 7.3963 | 21.25 | 25.2
10.3 |  |  | 7.25 | 10.8 |  |  | 7.1 | 10.5 |  | 3.5986 | 8.3 | 11.8 |  | 5.3 | 12.1 | 16.072 |  | 7.2964 | 19.4667 | 23.5
10.4 |  |  | 7.2 | 10.8 |  |  | 7.1 | 10.4 |  | 3.5986 | 8.2 | 11.7 |  | 5.3 | 11.775 | 15.7803 |  | 7.1964 | 18.1 | 22.4
10.5 |  |  | 7.1 | 10.7 |  |  | 7 | 10.4 |  |  | 8.1 | 11.625 |  | 5.2 | 11.6 | 15.5302 |  | 7.1564 | 17.1 | 21.2
10.6 |  |  | 7.1 | 10.7 |  |  | 6.975 | 10.3 |  |  | 8 | 11.5 |  | 5.1 | 11.225 | 15.2051 |  | 6.9965 | 16.4 | 20.65
10.7 |  |  | 7 | 10.6 |  |  | 6.9 | 10.3 |  |  | 7.975 | 11.5 |  | 5.1 | 11.0667 | 15.105 |  | 6.9465 | 15.7333 | 20.1
10.8 |  |  | 6.9333 | 10.5 |  |  | 6.9 | 10.18 |  |  | 7.9 | 11.4 |  | 5 | 10.825 | 14.88 |  | 6.8966 | 15.2333 | 19.6667
10.9 |  |  | 6.9 | 10.4 |  |  | 6.8 | 10.1 |  |  | 7.85 | 11.3 |  | 4.8 | 10.675 | 14.6799 |  | 6.8216 | 14.775 | 19.3
11 |  |  | 6.8 | 10.3 |  |  | 6.7 | 10 |  |  | 7.7 | 11.2 |  |  | 10.5 | 14.5048 |  | 6.7966 | 14.35 | 18.95
11.1 |  |  | 6.8 | 10.3 |  |  | 6.65 | 10 |  |  | 7.6 | 11.1 |  |  | 10.38 | 14.3048 |  | 6.6217 | 14.1 | 18.45
11.2 |  |  | 6.7 | 10.12 |  |  | 6.6 | 9.9 |  |  | 7.6 | 11 |  |  | 10.25 | 14.138 |  | 6.5967 | 13.775 | 18.175
11.3 |  |  | 6.6 | 10.1 |  |  | 6.55 | 9.7333 |  |  | 7.5 | 10.9333 |  |  | 10.1 | 13.9296 |  | 6.4967 | 13.5 | 17.925
11.4 |  |  | 6.5 | 9.9333 |  |  | 6.4 | 9.7 |  |  | 7.4 | 10.8 |  |  | 9.9667 | 13.8046 |  | 6.3635 | 13.3 | 17.7667
11.5 |  |  | 6.4 | 9.9 |  |  | 6.35 | 9.6 |  |  | 7.3 | 10.725 |  |  | 9.9 | 13.7046 |  |  | 13.075 | 17.5
11.6 |  |  | 6.3 | 9.7333 |  |  | 6.2 | 9.5 |  |  | 7.2 | 10.6 |  |  | 9.75 | 13.6045 |  |  | 12.825 | 17.3
11.7 |  |  | 6.2 | 9.6 |  |  | 6.125 | 9.4333 |  |  | 7.1 | 10.5 |  |  | 9.6 | 13.4795 |  |  | 12.5 | 16.95
11.8 |  |  | 6.1 | 9.5 |  |  | 6.1 | 9.3 |  |  | 7 | 10.4 |  |  | 9.55 | 13.3044 |  |  | 12.3667 | 16.7
11.9 |  |  | 6.04 | 9.425 |  |  | 6 | 9.2 |  |  | 6.9333 | 10.3 |  |  | 9.4 | 13.2044 |  |  | 12.2 | 16.5
12 |  |  | 5.9 | 9.3 |  |  | 5.9333 | 9.1 |  |  | 6.9 | 10.2 |  |  | 9.3 | 13.071 |  |  | 12 | 16.4
12.1 |  |  | 5.9 | 9.2 |  |  | 5.8 | 9 |  |  | 6.8 | 10.1 |  |  | 9.1667 | 13.0043 |  |  | 11.8 | 16.15
12.2 |  |  | 5.8 | 9.15 |  |  | 5.8 | 8.9 |  |  | 6.775 | 10 |  |  | 9 | 12.8043 |  |  | 11.7 | 16
12.3 |  |  | 5.7 | 9 |  |  | 5.7 | 8.8 |  |  | 6.7 | 9.85 |  |  | 8.9 | 12.7042 |  |  | 11.5667 | 15.9
12.4 |  |  | 5.6 | 8.9 |  |  | 5.625 | 8.725 |  |  | 6.5 | 9.8 |  |  | 8.8 | 12.6292 |  |  | 11.4 | 15.7
12.5 |  |  | 5.5333 | 8.825 |  |  | 5.5 | 8.6 |  |  | 6.4 | 9.7 |  |  | 8.725 | 12.4708 |  |  | 11.3 | 15.6
12.6 |  |  | 5.5 | 8.7 |  |  | 5.45 | 8.55 |  |  | 6.3 | 9.6 |  |  | 8.7 | 12.4041 |  |  | 11.2 | 15.5
12.7 | [6 empty] | 5.4 | 8.4 |  |  | 6.2 | 9.5 |  |  | 8.6 | 12.2791 |  |  | 11.0333 | 15.4
12.8 | [6 empty] | 5.3 | 8.3 |  |  | 6.1 | 9.36 |  |  | 8.44 | 12.1791 |  |  | 10.9 | 15.125
12.9 | [10 empty] | 6.1 | 9.3 |  |  | 8.3667 | 12.004 |  |  | 10.725 | 15
13 | [10 empty] | 6 | 9.2 |  |  | 8.3 | 11.904 |  |  | 10.6 | 14.8
13.1 | [10 empty] | 5.9 | 9.1 |  |  | 8.1 | 11.8039 |  |  | 10.54 | 14.8
13.2 | [10 empty] | 5.8 | 8.975 |  |  | 8 | 11.7039 |  |  | 10.3333 | 14.5667
13.3 | [14 empty] | 7.8 | 11.5705 |  |  | 10.2 | 14.3667
13.4 | [14 empty] | 7.7 | 11.4238 |  |  | 10.2 | 14.3
13.5 | [14 empty] | 7.6 | 11.3038 |  |  | 10 | 14.2
13.6 | [14 empty] | 7.5 | 11.2037 |  |  | 9.9 | 14.1
13.7 | [14 empty] | 7.4333 | 11.0037 |  |  | 9.75 | 14
13.8 | [14 empty] | 7.325 | 10.9036 |  |  | 9.7 | 13.825
13.9 | [14 empty] | 7.275 | 10.8036 |  |  | 9.55 | 13.7
14 | [18 empty] | 9.5 | 13.6
14.1 | [18 empty] | 9.35 | 13.525
14.2 | [18 empty] | 9.225 | 13.3
14.3 | [18 empty] | 9.1 | 13.2
14.4 | [18 empty] | 9 | 13
14.5 | [18 empty] | 8.85 | 12.95
14.6 | [18 empty] | 8.7 | 12.8
14.7 | [18 empty] | 8.62 | 12.6

## Wedge 60 Longitudinal Profiles
Longitudinal Profiles
Wedge 60
SSD=100cm
Field Size=40x15cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2.4 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.8 |  |  |  |  | 11.1
-30.7 |  |  |  |  | 11.18
-30.6 |  |  |  |  | 11.3
-30.5 |  |  |  |  | 11.4125
-30.4 |  |  |  |  | 11.55
-30.3 |  |  |  |  | 11.7125
-30.2 |  |  |  |  | 11.8429
-30.1 |  |  |  |  | 12
-30 |  |  |  |  | 12.1667
-29.9 |  |  |  |  | 12.3125
-29.8 |  |  |  |  | 12.4429
-29.7 |  |  |  |  | 12.6
-29.6 |  |  |  |  | 12.7143
-29.5 |  |  |  |  | 12.9
-29.4 |  |  |  |  | 13.05
-29.3 |  |  |  |  | 13.2
-29.2 |  |  |  |  | 13.375
-29.1 |  |  |  |  | 13.5556
-29 |  |  |  |  | 13.675
-28.9 |  |  |  | 9.6 | 13.9
-28.8 |  |  |  | 9.7 | 14.125
-28.7 |  |  |  | 9.8 | 14.25
-28.6 |  |  |  | 9.9125 | 14.45
-28.5 |  |  |  | 10 | 14.6
-28.4 |  |  |  | 10.0625 | 14.8375
-28.3 |  |  |  | 10.275 | 15.1375
-28.2 |  |  |  | 10.3333 | 15.3429
-28.1 |  |  |  | 10.475 | 15.525
-28 |  |  |  | 10.6 | 15.75
-27.9 |  |  |  | 10.725 | 15.975
-27.8 |  |  |  | 10.8667 | 16.325
-27.7 |  |  |  | 11 | 16.5375
-27.6 |  |  |  | 11.1286 | 16.7429
-27.5 |  |  |  | 11.2625 | 17.0571
-27.4 |  |  |  | 11.3875 | 17.3875
-27.3 |  |  |  | 11.5125 | 17.7
-27.2 |  |  |  | 11.675 | 18.2
-27.1 |  |  |  | 11.925 | 18.66
-27 |  |  |  | 12.0875 | 19.1571
-26.9 |  |  | 8.1625 | 12.3125 | 19.9
-26.8 |  |  | 8.3 | 12.4429 | 20.6
-26.7 |  |  | 8.4 | 12.65 | 21.7
-26.6 |  |  | 8.4571 | 12.8 | 23.2429
-26.5 |  |  | 8.5167 | 13.025 | 24.7286
-26.4 |  |  | 8.6 | 13.15 | 27.0875
-26.3 |  |  | 8.65 | 13.4 | 29.9667
-26.2 |  |  | 8.7 | 13.5429 | 33.25
-26.1 |  |  | 8.7667 | 13.7714 | 37.425
-26 |  |  | 8.85 | 14.04 | 41.8
-25.9 |  | 7.6 | 8.9143 | 14.3 | 46.6667
-25.8 |  | 7.6 | 9 | 14.5375 | 51.6714
-25.7 |  | 7.6667 | 9.1 | 14.7875 | 56.24
-25.6 |  | 7.725 | 9.2 | 15.075 | 60.8
-25.5 |  | 7.8 | 9.2375 | 15.4 | 64.1
-25.4 | 7.6 | 7.8 | 9.3625 | 15.7667 | 67.1857
-25.3 | 7.6286 | 7.8625 | 9.5 | 16.0714 | 69.5667
-25.2 | 7.7 | 7.9 | 9.6 | 16.5143 | 71.3667
-25.1 | 7.7375 | 7.9625 | 9.7375 | 17.1 | 72.7
-25 | 7.8 | 8 | 9.8714 | 17.6143 | 73.7
-24.9 | 7.8 | 8.0571 | 10.0333 | 18.2833 | 74.5571
-24.8 | 7.9 | 8.1 | 10.2 | 19.1167 | 75.35
-24.7 | 7.9571 | 8.1429 | 10.3571 | 20.28 | 76.1889
-24.6 | 8 | 8.2 | 10.5 | 21.9 | 76.6714
-24.5 | 8 | 8.3 | 10.66 | 23.8 | 77.1571
-24.4 | 8.1 | 8.3571 | 10.8143 | 26.3667 | 77.5375
-24.3 | 8.1 | 8.4167 | 10.98 | 29.8667 | 77.9286
-24.2 | 8.2 | 8.5 | 11.15 | 33.675 | 78.32
-24.1 | 8.2 | 8.58 | 11.4143 | 38.4 | 78.7
-24 | 8.2333 | 8.625 | 11.6143 | 43.975 | 79.1375
-23.9 | 8.3 | 8.7571 | 11.9 | 49.7 | 79.3667
-23.8 | 8.3 | 8.9 | 12.0333 | 55.125 | 79.7
-23.7 | 8.3143 | 9 | 12.3 | 60.6857 | 79.95
-23.6 | 8.4 | 9.1 | 12.5 | 65.2 | 80.2
-23.5 | 8.5 | 9.2 | 12.8 | 68.8333 | 80.525
-23.4 | 8.6 | 9.45 | 13.1667 | 72.3 | 80.8
-23.3 | 8.6571 | 9.5167 | 13.4714 | 74.72 | 81.075
-23.2 | 8.7125 | 9.6833 | 13.9 | 76.5 | 81.4143
-23.1 | 8.85 | 9.8429 | 14.35 | 77.9 | 81.675
-23 | 9 | 10.05 | 14.8875 | 78.6875 | 82.0143
-22.9 | 9.0833 | 10.2333 | 15.6375 | 79.5625 | 82.1667
-22.8 | 9.2 | 10.375 | 16.5 | 80.3125 | 82.4125
-22.7 | 9.3143 | 10.6333 | 17.6857 | 80.9375 | 82.5857
-22.6 | 9.5143 | 10.8 | 19.3143 | 81.5 | 82.9
-22.5 | 9.6875 | 11.06 | 21.4571 | 82.0286 | 83.1667
-22.4 | 9.8 | 11.2667 | 24.3 | 82.4333 | 83.3143
-22.3 | 9.9143 | 11.6 | 28.075 | 82.7 | 83.5143
-22.2 | 10.08 | 11.9714 | 33.2571 | 83.12 | 83.6778
-22.1 | 10.25 | 12.4 | 39 | 83.6333 | 83.9
-22 | 10.5 | 12.9875 | 45.8571 | 83.9429 | 84.1222
-21.9 | 10.76 | 13.6333 | 52.6 | 84.1333 | 84.2375
-21.8 | 10.95 | 14.5 | 59.3 | 84.5143 | 84.3833
-21.7 | 11.2625 | 15.75 | 66.0333 | 84.825 | 84.575
-21.6 | 11.7 | 17.5 | 71.425 | 85.075 | 84.7625
-21.5 | 12.1 | 19.9 | 75.8333 | 85.3875 | 84.975
-21.4 | 12.7857 | 23.4 | 78.9286 | 85.6 | 85.1429
-21.3 | 13.6 | 27.7571 | 81.3714 | 85.8667 | 85.275
-21.2 | 14.6833 | 33.24 | 83.2 | 86.1 | 85.4667
-21.1 | 16.3429 | 40.1857 | 84.375 | 86.3571 | 85.6857
-21 | 18.5 | 46.9857 | 85.4286 | 86.6429 | 85.96
-20.9 | 21.9333 | 54.7 | 86.1714 | 86.8 | 86.1
-20.8 | 26.3 | 62.65 | 86.8667 | 86.95 | 86.22
-20.7 | 32.3 | 69.6667 | 87.3857 | 87.2 | 86.4167
-20.6 | 39.2 | 75.3571 | 87.8143 | 87.3375 | 86.5556
-20.5 | 47.2 | 79.7 | 88.2667 | 87.5429 | 86.75
-20.4 | 54.9333 | 83.05 | 88.5714 | 87.7333 | 87
-20.3 | 63.4 | 85.4714 | 89.0714 | 87.8143 | 87.1571
-20.2 | 70.7 | 87.1333 | 89.3 | 88 | 87.225
-20.1 | 77.3 | 88.2857 | 89.6 | 88.2714 | 87.4429
-20 | 82.04 | 89.1667 | 89.9571 | 88.4286 | 87.5167
-19.9 | 85.64 | 90 | 90.175 | 88.6429 | 87.7
-19.8 | 88.0714 | 90.5 | 90.325 | 88.825 | 87.8667
-19.7 | 89.5571 | 91.0429 | 90.55 | 88.9667 | 87.9222
-19.6 | 90.6857 | 91.4333 | 90.8 | 89.1143 | 88.1286
-19.5 | 91.5429 | 91.825 | 90.9667 | 89.2667 | 88.2556
-19.4 | 92.2143 | 92.0857 | 91.2286 | 89.4286 | 88.375
-19.3 | 92.75 | 92.4 | 91.3833 | 89.5 | 88.6
-19.2 | 93.1667 | 92.6 | 91.6143 | 89.6429 | 88.8
-19.1 | 93.5667 | 92.825 | 91.7875 | 89.7857 | 88.8667
-19 | 93.9 | 93.0333 | 91.9429 | 89.9125 | 89
-18.9 | 94.16 | 93.2333 | 92.1143 | 90.075 | 89.1
-18.8 | 94.3333 | 93.5429 | 92.2 | 90.2833 | 89.2
-18.7 | 94.5714 | 93.7286 | 92.3286 | 90.4 | 89.3444
-18.6 | 94.74 | 93.8714 | 92.4833 | 90.5 | 89.4625
-18.5 | 94.9 | 94 | 92.6333 | 90.55 | 89.675
-18.4 | 95.0667 | 94.0667 | 92.7857 | 90.675 | 89.8125
-18.3 | 95.1167 | 94.2429 | 92.9 | 90.84 | 89.9375
-18.2 | 95.2833 | 94.3857 | 93 | 91 | 90.0625
-18.1 | 95.3667 | 94.4429 | 93.1 | 91.1 | 90.1
-18 | 95.5 | 94.5857 | 93.2 | 91.2429 | 90.2
-17.9 | 95.68 | 94.6167 | 93.3571 | 91.4714 | 90.2571
-17.8 | 95.7333 | 94.7 | 93.4875 | 91.6 | 90.4
-17.7 | 95.9 | 94.85 | 93.6 | 91.6 | 90.4429
-17.6 | 96 | 94.9875 | 93.7143 | 91.7286 | 90.6
-17.5 | 96.1167 | 95.06 | 93.8 | 91.8714 | 90.75
-17.4 | 96.2 | 95.1125 | 93.9143 | 92 | 90.875
-17.3 | 96.3 | 95.25 | 94 | 92.0375 | 91
-17.2 | 96.4 | 95.3 | 94.0857 | 92.1714 | 91.0375
-17.1 | 96.45 | 95.4333 | 94.2333 | 92.2333 | 91.1714
-17 | 96.52 | 95.5 | 94.3 | 92.4 | 91.3571
-16.9 | 96.6 | 95.5 | 94.35 | 92.45 | 91.5
-16.8 | 96.7 | 95.6 | 94.4857 | 92.65 | 91.7125
-16.7 | 96.7 | 95.7 | 94.68 | 92.8 | 91.8429
-16.6 | 96.8 | 95.8 | 94.7333 | 92.9 | 91.9
-16.5 | 96.8 | 95.9 | 94.8857 | 93.0286 | 92.0125
-16.4 | 96.9 | 95.9 | 94.9 | 93.1 | 92.1333
-16.3 | 96.9 | 96 | 94.9 | 93.1889 | 92.2667
-16.2 | 96.96 | 96 | 95 | 93.3143 | 92.4167
-16.1 | 97 | 96.0286 | 95.0375 | 93.4 | 92.5
-16 | 97 | 96.1 | 95.1833 | 93.5333 | 92.6
-15.9 | 97.14 | 96.1714 | 95.3167 | 93.6 | 92.7
-15.8 | 97.2167 | 96.24 | 95.4714 | 93.7 | 92.7
-15.7 | 97.3 | 96.4 | 95.525 | 93.8 | 92.8375
-15.6 | 97.4 | 96.5 | 95.6 | 93.9 | 92.9625
-15.5 | 97.4 | 96.6 | 95.7 | 94.0333 | 93
-15.4 | 97.5 | 96.7 | 95.7 | 94.2 | 93.1375
-15.3 | 97.5 | 96.7 | 95.8 | 94.2 | 93.2625
-15.2 | 97.5714 | 96.8 | 95.9167 | 94.3 | 93.3875
-15.1 | 97.6 | 96.9 | 96 | 94.4 | 93.6
-15 | 97.7 | 96.9857 | 96.0571 | 94.46 | 93.7
-14.9 | 97.76 | 97 | 96.1 | 94.6286 | 93.7571
-14.8 | 97.8167 | 97.0857 | 96.2 | 94.7833 | 93.8
-14.7 | 97.9625 | 97.1 | 96.2857 | 94.9 | 93.9222
-14.6 | 98 | 97.175 | 96.4 | 94.9875 | 94.0375
-14.5 | 98.0667 | 97.2 | 96.4833 | 95.0125 | 94.1625
-14.4 | 98.1 | 97.3 | 96.6 | 95.16 | 94.2
-14.3 | 98.12 | 97.3 | 96.6 | 95.3286 | 94.3
-14.2 | 98.2 | 97.35 | 96.675 | 95.4 | 94.3714
-14.1 | 98.25 | 97.5 | 96.8333 | 95.5 | 94.4889
-14 | 98.3 | 97.5714 | 96.9 | 95.5857 | 94.6
-13.9 | 98.4 | 97.62 | 97 | 95.6 | 94.7857
-13.8 | 98.4 | 97.7 | 97.0286 | 95.6 | 94.8429
-13.7 | 98.5333 | 97.7 | 97.1 | 95.7 | 94.9
-13.6 | 98.6 | 97.8 | 97.1 | 95.8 | 95.0571
-13.5 | 98.6333 | 97.8 | 97.2 | 95.85 | 95.1
-13.4 | 98.7 | 97.8 | 97.3333 | 95.9 | 95.1667
-13.3 | 98.8 | 97.9667 | 97.4 | 96.0222 | 95.3
-13.2 | 98.8 | 98.025 | 97.56 | 96.1 | 95.3714
-13.1 | 98.9 | 98.1571 | 97.6167 | 96.1875 | 95.4
-13 | 99 | 98.3 | 97.7 | 96.3125 | 95.4833
-12.9 | 99.0625 | 98.3625 | 97.7875 | 96.4 | 95.65
-12.8 | 99.1125 | 98.4 | 97.8 | 96.5 | 95.7
-12.7 | 99.25 | 98.4 | 97.9143 | 96.5 | 95.8
-12.6 | 99.3 | 98.4833 | 98 | 96.6333 | 95.9375
-12.5 | 99.3 | 98.6 | 98 | 96.7 | 96
-12.4 | 99.42 | 98.6286 | 98 | 96.825 | 96.1
-12.3 | 99.5 | 98.7625 | 98.1 | 96.9 | 96.1
-12.2 | 99.5667 | 98.8 | 98.1 | 97 | 96.2
-12.1 | 99.6 | 98.8714 | 98.2 | 97 | 96.2
-12 | 99.6 | 98.9 | 98.2833 | 97.0714 | 96.2
-11.9 | 99.6286 | 99 | 98.3 | 97.1 | 96.3143
-11.8 | 99.7 | 99.0833 | 98.4 | 97.1625 | 96.4571
-11.7 | 99.8429 | 99.1 | 98.46 | 97.2167 | 96.5
-11.6 | 99.9 | 99.1 | 98.55 | 97.3 | 96.5571
-11.5 | 99.9 | 99.2 | 98.7 | 97.3571 | 96.7
-11.4 | 99.9333 | 99.2 | 98.7 | 97.42 | 96.7714
-11.3 | 100 | 99.3 | 98.7429 | 97.5 | 96.84
-11.2 | 100 | 99.3857 | 98.8 | 97.6 | 96.9125
-11.1 | 100 | 99.4 | 98.85 | 97.6333 | 97
-11 | 100.1 | 99.4 | 98.9 | 97.7 | 97.02
-10.9 | 100.1571 | 99.45 | 98.9 | 97.7333 | 97.1
-10.8 | 100.2 | 99.5 | 99 | 97.9 | 97.2
-10.7 | 100.3 | 99.5833 | 99 | 97.9667 | 97.325
-10.6 | 100.3 | 99.68 | 99.0375 | 98 | 97.45
-10.5 | 100.3 | 99.7 | 99.1 | 98.0571 | 97.5857
-10.4 | 100.35 | 99.8143 | 99.2 | 98.1 | 97.6
-10.3 | 100.4 | 99.9 | 99.2286 | 98.1286 | 97.6375
-10.2 | 100.4 | 99.9 | 99.3 | 98.2 | 97.7833
-10.1 | 100.5429 | 99.9 | 99.4 | 98.3 | 97.8
-10 | 100.6 | 100 | 99.4 | 98.3 | 97.8
-9.9 | 100.6 | 100.0333 | 99.5 | 98.3 | 97.9
-9.8 | 100.7 | 100 | 99.5 | 98.4 | 97.98
-9.7 | 100.8 | 100.1 | 99.5667 | 98.5286 | 98
-9.6 | 100.8 | 100.1 | 99.6286 | 98.6625 | 98.05
-9.5 | 100.8 | 100.1 | 99.7 | 98.7 | 98.1
-9.4 | 100.8 | 100.1667 | 99.7125 | 98.8 | 98.1
-9.3 | 100.8 | 100.2167 | 99.8 | 98.9 | 98.2
-9.2 | 100.8 | 100.3 | 99.8 | 98.9 | 98.225
-9.1 | 100.86 | 100.4 | 99.8571 | 98.9 | 98.35
-9 | 100.9 | 100.4 | 99.9 | 98.9 | 98.4
-8.9 | 100.9143 | 100.4 | 99.9 | 99 | 98.5
-8.8 | 101 | 100.4 | 99.9 | 99.0833 | 98.5
-8.7 | 101 | 100.4875 | 99.9714 | 99.1 | 98.6
-8.6 | 101 | 100.5 | 100 | 99.1714 | 98.6833
-8.5 | 101.06 | 100.5 | 100 | 99.2 | 98.7143
-8.4 | 101.1167 | 100.6 | 100 | 99.2 | 98.8
-8.3 | 101.2714 | 100.6 | 100.0125 | 99.2 | 98.9
-8.2 | 101.3 | 100.7 | 100.1 | 99.3571 | 98.9
-8.1 | 101.3 | 100.7 | 100.1625 | 99.4 | 99
-8 | 101.3 | 100.7833 | 100.2 | 99.5 | 98.9
-7.9 | 101.4 | 100.8 | 100.2833 | 99.5 | 99
-7.8 | 101.4 | 100.8 | 100.3 | 99.5 | 99
-7.7 | 101.4 | 100.8 | 100.4 | 99.5667 | 99.0167
-7.6 | 101.4 | 100.8571 | 100.4 | 99.6 | 99.1
-7.5 | 101.4 | 100.9 | 100.4143 | 99.725 | 99.2
-7.4 | 101.4 | 101 | 100.5 | 99.8 | 99.2
-7.3 | 101.5 | 101 | 100.5 | 99.8 | 99.2857
-7.2 | 101.5 | 101 | 100.5 | 99.9 | 99.3
-7.1 | 101.5 | 101 | 100.5 | 99.9 | 99.3375
-7 | 101.5 | 101.0714 | 100.6 | 99.9625 | 99.4
-6.9 | 101.5 | 101.05 | 100.6 | 100 | 99.5286
-6.8 | 101.5 | 101 | 100.6 | 100 | 99.6
-6.7 | 101.5 | 101.02 | 100.7 | 100 | 99.7
-6.6 | 101.5 | 101.1 | 100.7 | 100 | 99.7
-6.5 | 101.5 | 101.1 | 100.75 | 100 | 99.7333
-6.4 | 101.5 | 101.16 | 100.8 | 100 | 99.8
-6.3 | 101.5 | 101.2 | 100.8 | 100.025 | 99.8
-6.2 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
-6.1 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
-6 | 101.5 | 101.2 | 100.8 | 100.125 | 99.82
-5.9 | 101.5 | 101.2 | 100.8 | 100.2 | 99.8143
-5.8 | 101.5 | 101.2 | 100.8 | 100.2 | 99.8
-5.7 | 101.5 | 101.2 | 100.8571 | 100.25 | 99.8857
-5.6 | 101.5 | 101.2 | 100.9 | 100.3 | 99.9
-5.5 | 101.5 | 101.2 | 100.8714 | 100.3 | 99.9625
-5.4 | 101.5667 | 101.3 | 100.8 | 100.3 | 100
-5.3 | 101.6 | 101.3 | 100.8 | 100.3 | 100
-5.2 | 101.6 | 101.3 | 100.8 | 100.3 | 100
-5.1 | 101.6 | 101.225 | 100.8 | 100.3 | 100
-5 | 101.5167 | 101.2 | 100.8 | 100.3 | 100
-4.9 | 101.5 | 101.2 | 100.8 | 100.3 | 100
-4.8 | 101.4571 | 101.2 | 100.7667 | 100.3 | 100.1
-4.7 | 101.4 | 101.2 | 100.7 | 100.3 | 100.1
-4.6 | 101.4429 | 101.2 | 100.7 | 100.3 | 100.1
-4.5 | 101.4143 | 101.2 | 100.7 | 100.3 | 100.1
-4.4 | 101.4 | 101.1333 | 100.75 | 100.3 | 100.2
-4.3 | 101.4 | 101 | 100.8 | 100.3 | 100.2
-4.2 | 101.35 | 101 | 100.7 | 100.3 | 100.2
-4.1 | 101.3 | 101 | 100.7 | 100.3 | 100.2
-4 | 101.3 | 101 | 100.7 | 100.3857 | 100.2
-3.9 | 101.2 | 101 | 100.7 | 100.4 | 100.2
-3.8 | 101.2 | 101 | 100.7 | 100.4 | 100.2
-3.7 | 101.1 | 101 | 100.7 | 100.3333 | 100.2
-3.6 | 101.1 | 101 | 100.6143 | 100.3 | 100.2
-3.5 | 101 | 101 | 100.6 | 100.3 | 100.2
-3.4 | 101 | 100.9 | 100.6 | 100.2 | 100.2
-3.3 | 101 | 100.9 | 100.6 | 100.2 | 100.2
-3.2 | 101 | 100.9 | 100.6 | 100.3 | 100.2
-3.1 | 101 | 100.9 | 100.6 | 100.3 | 100.2
-3 | 100.9 | 100.8167 | 100.5333 | 100.3 | 100.2857
-2.9 | 100.9 | 100.8 | 100.5 | 100.2 | 100.2875
-2.8 | 100.9 | 100.8 | 100.5 | 100.3 | 100.2
-2.7 | 100.85 | 100.7 | 100.4 | 100.2571 | 100.2
-2.6 | 100.8 | 100.7 | 100.4 | 100.2 | 100.1
-2.5 | 100.8 | 100.6143 | 100.4 | 100.2 | 100.0571
-2.4 | 100.7571 | 100.6 | 100.3 | 100.2 | 100
-2.3 | 100.7 | 100.6 | 100.3 | 100.1143 | 100.1
-2.2 | 100.7 | 100.6 | 100.3 | 100.1 | 100.1
-2.1 | 100.7 | 100.6 | 100.3 | 100.1 | 100.1
-2 | 100.7 | 100.5 | 100.3 | 100.125 | 100.1
-1.9 | 100.7 | 100.5 | 100.3 | 100.2 | 100.1
-1.8 | 100.6 | 100.4 | 100.2429 | 100.2 | 100.1
-1.7 | 100.5 | 100.4 | 100.2 | 100.2 | 100.1
-1.6 | 100.4833 | 100.4 | 100.2 | 100.1 | 100.1
-1.5 | 100.4 | 100.3 | 100.1143 | 100.1 | 100.1167
-1.4 | 100.3667 | 100.3 | 100.1 | 100.1 | 100.2
-1.3 | 100.3 | 100.3 | 100.1 | 100.1 | 100.2
-1.2 | 100.24 | 100.3 | 100.1 | 100.1 | 100.1286
-1.1 | 100.2 | 100.2 | 100.0667 | 100.0333 | 100.1
-1 | 100.2 | 100.2 | 100 | 100.1 | 100.1
-0.9 | 100.1857 | 100.2 | 100 | 100.1 | 100.1
-0.8 | 100.1 | 100.1667 | 100 | 100.1 | 100.1
-0.7 | 100.1 | 100.1 | 100 | 100.0571 | 100.1
-0.6 | 100.2 | 100.1 | 100 | 100 | 100.1
-0.5 | 100.1 | 100 | 100 | 100 | 100.08
-0.4 | 100.1 | 100 | 100 | 100 | 100
-0.3 | 100.1 | 100 | 100 | 100 | 100
-0.2 | 100.05 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100.0714 | 100 | 100 | 100 | 100
0.3 | 100.1 | 100 | 100 | 100 | 100
0.4 | 100.1 | 100 | 100 | 100 | 100
0.5 | 100.1286 | 100 | 100 | 100 | 100.1
0.6 | 100.2 | 100.0667 | 100 | 100 | 100.1
0.7 | 100.1 | 100.1 | 100 | 100.0714 | 100.1
0.8 | 100.1 | 100.1167 | 100 | 100.1 | 100.1
0.9 | 100.2 | 100.2 | 100 | 100.1 | 100.1
1 | 100.2 | 100.2 | 100 | 100.1 | 100.1
1.1 | 100.2 | 100.2 | 100.0143 | 100.1 | 100.1667
1.2 | 100.26 | 100.25 | 100.1 | 100.1 | 100.1
1.3 | 100.3 | 100.3 | 100.1 | 100.1 | 100.2
1.4 | 100.3 | 100.3 | 100.1 | 100.1 | 100.2
1.5 | 100.4 | 100.3 | 100.1 | 100.1 | 100.1429
1.6 | 100.5 | 100.4 | 100.2 | 100.1 | 100.1
1.7 | 100.5 | 100.4 | 100.25 | 100.2 | 100.1
1.8 | 100.6 | 100.4167 | 100.3 | 100.2 | 100.1
1.9 | 100.64 | 100.5 | 100.28 | 100.2 | 100.1
2 | 100.7 | 100.5 | 100.3 | 100.1 | 100.1
2.1 | 100.7 | 100.6 | 100.3 | 100.1 | 100.1
2.2 | 100.7 | 100.6 | 100.3 | 100.1 | 100.1
2.3 | 100.7 | 100.6 | 100.3 | 100.1429 | 100.1
2.4 | 100.7833 | 100.6 | 100.3 | 100.2 | 100
2.5 | 100.8 | 100.6 | 100.4 | 100.2 | 100.1
2.6 | 100.8 | 100.7 | 100.4 | 100.2 | 100.1
2.7 | 100.875 | 100.7 | 100.4571 | 100.2 | 100.2
2.8 | 100.9 | 100.8 | 100.5 | 100.2875 | 100.2
2.9 | 100.9 | 100.8 | 100.5 | 100.2 | 100.2625
3 | 100.9 | 100.8667 | 100.5 | 100.3 | 100.3
3.1 | 101 | 100.9 | 100.6 | 100.3 | 100.2
3.2 | 101 | 100.9 | 100.6 | 100.3 | 100.2
3.3 | 101 | 100.9 | 100.6 | 100.2 | 100.2
3.4 | 101 | 100.9 | 100.6 | 100.2 | 100.2
3.5 | 101 | 100.9833 | 100.6 | 100.3 | 100.2
3.6 | 101.1 | 101 | 100.6 | 100.3 | 100.2
3.7 | 101.1 | 101 | 100.7 | 100.35 | 100.2
3.8 | 101.2 | 101 | 100.7 | 100.4 | 100.2
3.9 | 101.2167 | 101 | 100.7 | 100.4 | 100.2
4 | 101.3 | 101 | 100.7 | 100.4 | 100.2
4.1 | 101.3 | 101 | 100.7 | 100.3167 | 100.2
4.2 | 101.3667 | 101 | 100.7 | 100.3 | 100.2
4.3 | 101.4 | 101 | 100.775 | 100.3 | 100.2
4.4 | 101.4 | 101.1 | 100.7 | 100.3 | 100.2
4.5 | 101.4286 | 101.2 | 100.7 | 100.3 | 100.1
4.6 | 101.4286 | 101.2 | 100.7 | 100.3 | 100.1
4.7 | 101.4 | 101.2 | 100.7571 | 100.3 | 100.1
4.8 | 101.4857 | 101.2 | 100.72 | 100.3 | 100.1
4.9 | 101.5 | 101.2 | 100.8 | 100.3 | 100
5 | 101.54 | 101.2 | 100.8 | 100.3 | 100
5.1 | 101.6 | 101.2 | 100.8 | 100.3 | 100
5.2 | 101.6 | 101.3 | 100.8 | 100.3 | 100
5.3 | 101.6 | 101.3 | 100.8 | 100.3 | 100
5.4 | 101.5429 | 101.3 | 100.8 | 100.3 | 100
5.5 | 101.5 | 101.2 | 100.825 | 100.3 | 99.9875
5.6 | 101.5 | 101.2 | 100.9 | 100.2833 | 99.9
5.7 | 101.5 | 101.2 | 100.9 | 100.2714 | 99.8167
5.8 | 101.5 | 101.2 | 100.8 | 100.2667 | 99.8
5.9 | 101.5 | 101.2 | 100.8 | 100.2 | 99.8
6 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
6.1 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
6.2 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
6.3 | 101.5 | 101.2 | 100.8 | 100 | 99.8
6.4 | 101.5 | 101.2 | 100.8 | 100 | 99.8
6.5 | 101.5 | 101.1 | 100.8 | 100 | 99.7571
6.6 | 101.5 | 101.1 | 100.7 | 100 | 99.7
6.7 | 101.5 | 101 | 100.7 | 100 | 99.7
6.8 | 101.5 | 101 | 100.6167 | 100 | 99.6286
6.9 | 101.5 | 101 | 100.6 | 100 | 99.55
7 | 101.5 | 101.1 | 100.6 | 99.9429 | 99.4
7.1 | 101.5 | 101 | 100.5 | 99.9 | 99.3625
7.2 | 101.5 | 101 | 100.5 | 99.9 | 99.3
7.3 | 101.5 | 101 | 100.5 | 99.8625 | 99.3
7.4 | 101.4 | 100.9857 | 100.5 | 99.8 | 99.2
7.5 | 101.4 | 100.9 | 100.4 | 99.7 | 99.2
7.6 | 101.4 | 100.9 | 100.4 | 99.675 | 99.1167
7.7 | 101.4 | 100.8 | 100.4 | 99.6 | 99.0333
7.8 | 101.4 | 100.8 | 100.3 | 99.5 | 99
7.9 | 101.3667 | 100.8 | 100.2333 | 99.5 | 99
8 | 101.3 | 100.7286 | 100.2 | 99.5 | 98.9
8.1 | 101.3 | 100.7 | 100.2 | 99.4 | 98.9286
8.2 | 101.3 | 100.7 | 100.1 | 99.3286 | 98.9
8.3 | 101.2571 | 100.6429 | 100.0571 | 99.2833 | 98.9
8.4 | 101.1 | 100.5833 | 100 | 99.2 | 98.8
8.5 | 101.0333 | 100.5 | 100 | 99.2 | 98.7375
8.6 | 101 | 100.5 | 100 | 99.1429 | 98.7
8.7 | 101 | 100.5 | 100 | 99.1 | 98.6
8.8 | 101 | 100.4 | 99.9 | 99.0571 | 98.5
8.9 | 100.9 | 100.4 | 99.9 | 99 | 98.5
9 | 100.9 | 100.4 | 99.9 | 98.9 | 98.4
9.1 | 100.84 | 100.4 | 99.8143 | 98.9 | 98.375
9.2 | 100.8 | 100.325 | 99.8 | 98.9 | 98.25
9.3 | 100.8 | 100.2714 | 99.7875 | 98.9 | 98.2
9.4 | 100.8 | 100.2 | 99.7 | 98.8 | 98.1111
9.5 | 100.8 | 100.1 | 99.7 | 98.7714 | 98.1
9.6 | 100.8 | 100.1 | 99.6 | 98.6375 | 98.08
9.7 | 100.78 | 100.1 | 99.6 | 98.5125 | 98
9.8 | 100.7 | 100.05 | 99.5 | 98.4 | 97.9
9.9 | 100.6 | 100 | 99.46 | 98.3 | 97.9
10 | 100.6 | 100 | 99.4 | 98.3 | 97.8
10.1 | 100.6 | 99.9 | 99.4 | 98.3 | 97.8
10.2 | 100.4 | 99.9 | 99.3 | 98.2429 | 97.8
10.3 | 100.4 | 99.9 | 99.2714 | 98.2 | 97.6625
10.4 | 100.4 | 99.8 | 99.2 | 98.1 | 97.6
10.5 | 100.3 | 99.7 | 99.14 | 98.0333 | 97.6
10.6 | 100.3 | 99.62 | 99.0875 | 98 | 97.45
10.7 | 100.3 | 99.6 | 99 | 98 | 97.3
10.8 | 100.2 | 99.5 | 99 | 97.9 | 97.225
10.9 | 100.1286 | 99.4 | 98.9333 | 97.8 | 97.1
11 | 100.1 | 99.4 | 98.9 | 97.7 | 97.05
11.1 | 100 | 99.4 | 98.9 | 97.6 | 97
11.2 | 100 | 99.4 | 98.8 | 97.6 | 96.9286
11.3 | 100 | 99.3 | 98.7857 | 97.5 | 96.8667
11.4 | 99.9 | 99.2333 | 98.7 | 97.4 | 96.8
11.5 | 99.9 | 99.2 | 98.6571 | 97.3375 | 96.7
11.6 | 99.9 | 99.1167 | 98.6 | 97.3 | 96.5857
11.7 | 99.8143 | 99.1 | 98.5 | 97.2 | 96.5
11.8 | 99.7 | 99.1 | 98.3429 | 97.1429 | 96.4857
11.9 | 99.6 | 99 | 98.3 | 97.1 | 96.3429
12 | 99.6 | 98.95 | 98.3 | 97.05 | 96.2
12.1 | 99.6 | 98.825 | 98.1875 | 97 | 96.2
12.2 | 99.5333 | 98.8 | 98.1 | 97 | 96.1333
12.3 | 99.5 | 98.8 | 98.1 | 96.9 | 96.1
12.4 | 99.4 | 98.6833 | 98.0286 | 96.8 | 96.0333
12.5 | 99.3333 | 98.6 | 98 | 96.7 | 96
12.6 | 99.3 | 98.5 | 98 | 96.6143 | 95.9625
12.7 | 99.2167 | 98.4 | 97.9 | 96.5 | 95.8167
12.8 | 99.1857 | 98.4 | 97.8 | 96.4833 | 95.7
12.9 | 99.0429 | 98.4 | 97.7429 | 96.4 | 95.6714
13 | 99 | 98.2667 | 97.7 | 96.3 | 95.5
13.1 | 98.9 | 98.2 | 97.6 | 96.2625 | 95.4
13.2 | 98.8333 | 98.0714 | 97.5 | 96.1286 | 95.3
13.3 | 98.8 | 98 | 97.4429 | 96 | 95.3
13.4 | 98.7 | 97.86 | 97.3714 | 95.9571 | 95.1857
13.5 | 98.7 | 97.8 | 97.2 | 95.825 | 95.1
13.6 | 98.6 | 97.7833 | 97.1429 | 95.8 | 95
13.7 | 98.5143 | 97.7 | 97.1 | 95.7 | 94.9167
13.8 | 98.4333 | 97.7 | 97 | 95.65 | 94.8571
13.9 | 98.3833 | 97.68 | 97 | 95.6 | 94.7143
14 | 98.3 | 97.5143 | 96.95 | 95.5667 | 94.5833
14.1 | 98.22 | 97.48 | 96.8833 | 95.5 | 94.5
14.2 | 98.2 | 97.4 | 96.7286 | 95.4 | 94.3
14.3 | 98.1 | 97.3 | 96.6 | 95.3143 | 94.3
14.4 | 98.1 | 97.25 | 96.5625 | 95.1667 | 94.225
14.5 | 98.1 | 97.2 | 96.4286 | 95 | 94.1875
14.6 | 98 | 97.2 | 96.3857 | 95 | 94.0625
14.7 | 97.9375 | 97.1 | 96.2429 | 94.9167 | 93.9444
14.8 | 97.8 | 97.0333 | 96.2 | 94.7667 | 93.825
14.9 | 97.7333 | 97 | 96.1429 | 94.6143 | 93.7857
15 | 97.7 | 97 | 96.1 | 94.5 | 93.6556
15.1 | 97.6167 | 96.9 | 96 | 94.4 | 93.6
15.2 | 97.55 | 96.8 | 95.9 | 94.3 | 93.4222
15.3 | 97.5 | 96.7429 | 95.8 | 94.2286 | 93.275
15.4 | 97.5 | 96.7 | 95.7429 | 94.1667 | 93.1
15.5 | 97.4 | 96.6 | 95.6833 | 94.1 | 93.0167
15.6 | 97.3833 | 96.5143 | 95.6 | 93.94 | 92.975
15.7 | 97.3 | 96.3571 | 95.5 | 93.7667 | 92.8
15.8 | 97.2 | 96.2833 | 95.42 | 93.7 | 92.7167
15.9 | 97.1167 | 96.2 | 95.2714 | 93.6 | 92.64
16 | 97.05 | 96.1 | 95.1286 | 93.5143 | 92.5857
16.1 | 97 | 96.075 | 95.0875 | 93.4 | 92.5
16.2 | 96.94 | 96 | 95 | 93.2875 | 92.45
16.3 | 96.9 | 95.9857 | 94.9286 | 93.2 | 92.2875
16.4 | 96.9 | 95.9 | 94.9 | 93.1429 | 92.1625
16.5 | 96.8 | 95.86 | 94.8286 | 93 | 92.0375
16.6 | 96.8 | 95.7833 | 94.7 | 92.9 | 91.9125
16.7 | 96.7 | 95.7 | 94.6333 | 92.78 | 91.8625
16.8 | 96.7 | 95.6 | 94.5 | 92.625 | 91.675
16.9 | 96.6 | 95.5 | 94.4 | 92.5 | 91.5
17 | 96.5 | 95.5 | 94.3 | 92.3714 | 91.3571
17.1 | 96.4333 | 95.4 | 94.2 | 92.2 | 91.2
17.2 | 96.3714 | 95.3429 | 94.1333 | 92.15 | 91.0625
17.3 | 96.3 | 95.3 | 94 | 92.1 | 91
17.4 | 96.26 | 95.1571 | 93.9714 | 91.9857 | 90.9
17.5 | 96.1 | 95 | 93.8167 | 91.8429 | 90.775
17.6 | 96.0333 | 94.9429 | 93.6571 | 91.7125 | 90.62
17.7 | 95.8833 | 94.8 | 93.6 | 91.6 | 90.4667
17.8 | 95.8 | 94.7167 | 93.5 | 91.6 | 90.325
17.9 | 95.65 | 94.6 | 93.4 | 91.4571 | 90.2857
18 | 95.5714 | 94.5286 | 93.2 | 91.3143 | 90.2
18.1 | 95.4167 | 94.4 | 93.06 | 91.1714 | 90.1
18.2 | 95.2571 | 94.3333 | 92.98 | 91 | 89.9875
18.3 | 95.1 | 94.1857 | 92.9 | 90.82 | 89.8625
18.4 | 95.0429 | 94.1 | 92.7429 | 90.6571 | 89.8
18.5 | 94.8833 | 93.9571 | 92.5833 | 90.6 | 89.625
18.6 | 94.7167 | 93.825 | 92.4286 | 90.5 | 89.4875
18.7 | 94.5625 | 93.6714 | 92.38 | 90.4 | 89.3667
18.8 | 94.3333 | 93.5167 | 92.2143 | 90.26 | 89.2333
18.9 | 94.14 | 93.2333 | 92.0571 | 90.1125 | 89.0714
19 | 93.9 | 93.0143 | 91.9 | 89.975 | 89
19.1 | 93.5 | 92.85 | 91.7429 | 89.7571 | 88.8857
19.2 | 93.2 | 92.6333 | 91.6125 | 89.625 | 88.75
19.3 | 92.8 | 92.42 | 91.44 | 89.55 | 88.625
19.4 | 92.2 | 92.075 | 91.2429 | 89.4 | 88.4
19.5 | 91.5667 | 91.7875 | 90.9333 | 89.2333 | 88.2778
19.6 | 90.7 | 91.4333 | 90.8 | 89.0857 | 88.1143
19.7 | 89.5667 | 90.9875 | 90.6 | 88.9429 | 87.9333
19.8 | 88.0571 | 90.55 | 90.275 | 88.8 | 87.8857
19.9 | 85.6667 | 89.98 | 90.1429 | 88.5857 | 87.72
20 | 82.18 | 89.2 | 90 | 88.4125 | 87.54
20.1 | 76.9714 | 88.2 | 89.62 | 88.2625 | 87.4429
20.2 | 70.8286 | 87.1571 | 89.32 | 88.0571 | 87.2375
20.3 | 63.2 | 85.4333 | 89.025 | 87.8 | 87.15
20.4 | 55.2 | 83 | 88.62 | 87.7 | 86.9333
20.5 | 47 | 79.8 | 88.25 | 87.5 | 86.7
20.6 | 38.9167 | 75.425 | 87.8571 | 87.3143 | 86.575
20.7 | 32.1 | 69.8429 | 87.3571 | 87.175 | 86.4333
20.8 | 26 | 62.4857 | 86.86 | 87.02 | 86.24
20.9 | 21.75 | 55.2667 | 86.175 | 86.775 | 86.1
21 | 18.62 | 47.2857 | 85.3833 | 86.65 | 85.9
21.1 | 16.3 | 39.9714 | 84.475 | 86.3 | 85.7286
21.2 | 14.7143 | 33.5 | 83.06 | 86.15 | 85.4571
21.3 | 13.5167 | 27.9 | 81.2714 | 85.8667 | 85.3
21.4 | 12.7429 | 23.3286 | 78.8667 | 85.6429 | 85.0714
21.5 | 12.1 | 19.9 | 75.6333 | 85.375 | 85
21.6 | 11.7 | 17.5286 | 71.4 | 85.0375 | 84.775
21.7 | 11.2875 | 15.7 | 66.0667 | 84.775 | 84.5625
21.8 | 10.9571 | 14.52 | 59.7571 | 84.5143 | 84.4
21.9 | 10.76 | 13.6667 | 52.5857 | 84.2 | 84.2667
22 | 10.4833 | 12.95 | 45.3 | 83.8857 | 84.15
22.1 | 10.2571 | 12.45 | 39 | 83.5667 | 83.95
22.2 | 10.16 | 11.95 | 33.1375 | 83.15 | 83.7
22.3 | 9.9714 | 11.6 | 28.2 | 82.725 | 83.4857
22.4 | 9.8 | 11.2714 | 24.1 | 82.4 | 83.3429
22.5 | 9.6714 | 11.0333 | 21.425 | 82 | 83.1667
22.6 | 9.5286 | 10.8 | 19.3625 | 81.5 | 82.8571
22.7 | 9.3857 | 10.5833 | 17.675 | 80.9857 | 82.6429
22.8 | 9.2333 | 10.3571 | 16.5857 | 80.325 | 82.375
22.9 | 9.0571 | 10.1833 | 15.6143 | 79.575 | 82.1857
23 | 8.9833 | 10.0286 | 14.9375 | 78.7125 | 81.9857
23.1 | 8.8286 | 9.9 | 14.3286 | 77.7667 | 81.6375
23.2 | 8.7 | 9.6667 | 13.875 | 76.55 | 81.45
23.3 | 8.6429 | 9.5667 | 13.5 | 74.8 | 81.15
23.4 | 8.58 | 9.4 | 13.0667 | 72.25 | 80.85
23.5 | 8.5 | 9.2333 | 12.8333 | 68.9714 | 80.5
23.6 | 8.4 | 9.05 | 12.5 | 65.1 | 80.15
23.7 | 8.3857 | 8.9857 | 12.25 | 60.6 | 79.9
23.8 | 8.3 | 8.8429 | 12.0714 | 55.3429 | 79.6667
23.9 | 8.3 | 8.7125 | 11.85 | 49.7 | 79.4
24 | 8.2167 | 8.675 | 11.6286 | 44 | 79.1143
24.1 | 8.2 | 8.5333 | 11.3571 | 38.425 | 78.7
24.2 | 8.2 | 8.5 | 11.1 | 33.75 | 78.28
24.3 | 8.1 | 8.4 | 10.92 | 29.9333 | 77.9143
24.4 | 8.1 | 8.3 | 10.8625 | 26.5 | 77.5125
24.5 | 8.0333 | 8.2571 | 10.6 | 23.825 | 77.1375
24.6 | 8 | 8.2 | 10.45 | 21.9143 | 76.6857
24.7 | 7.9286 | 8.1875 | 10.3 | 20.18 | 76.1222
24.8 | 7.875 | 8.1 | 10.2 | 19.16 | 75.4
24.9 | 7.8 | 8.1 | 9.9833 | 18.25 | 74.5571
25 | 7.8 | 8 | 9.9 | 17.6429 | 73.7
25.1 | 7.7143 | 8 | 9.775 | 17.0667 | 72.7333
25.2 | 7.7 | 7.9 | 9.6 | 16.5 | 71.2667
25.3 | 7.6 | 7.8143 | 9.5 | 16.0875 | 69.4
25.4 | 7.6 | 7.8 | 9.4143 | 15.8 | 67.2714
25.5 |  | 7.8 | 9.2875 | 15.42 | 63.8833
25.6 |  | 7.7 | 9.15 | 15.125 | 60.55
25.7 |  | 7.7 | 9.0857 | 14.8125 | 56.25
25.8 |  | 7.6 | 9 | 14.525 | 51.7833
25.9 |  | 7.5833 | 8.9 | 14.2714 | 46.8667
26 |  |  | 8.8 | 14.1 | 41.9333
26.1 |  |  | 8.7167 | 13.825 | 37.4
26.2 |  |  | 8.7 | 13.5333 | 33.3111
26.3 |  |  | 8.6 | 13.3667 | 30
26.4 |  |  | 8.6 | 13.1286 | 27.125
26.5 |  |  | 8.5 | 13 | 24.7571
26.6 |  |  | 8.4 | 12.775 | 23.2333
26.7 |  |  | 8.4 | 12.65 | 21.7
26.8 |  |  | 8.3 | 12.5143 | 20.6333
26.9 |  |  | 8.2143 | 12.275 | 19.85
27 |  |  |  | 12.0625 | 19.1833
27.1 |  |  |  | 11.9375 | 18.62
27.2 |  |  |  | 11.725 | 18.2143
27.3 |  |  |  | 11.4875 | 17.7143
27.4 |  |  |  | 11.3625 | 17.375
27.5 |  |  |  | 11.2375 | 17.05
27.6 |  |  |  | 11.1 | 16.8
27.7 |  |  |  | 10.975 | 16.5111
27.8 |  |  |  | 10.8429 | 16.275
27.9 |  |  |  | 10.7 | 15.9375
28 |  |  |  | 10.575 | 15.7429
28.1 |  |  |  | 10.45 | 15.5375
28.2 |  |  |  | 10.4 | 15.3
28.3 |  |  |  | 10.2625 | 15.0625
28.4 |  |  |  | 10.1375 | 14.8125
28.5 |  |  |  | 10.0125 | 14.6222
28.6 |  |  |  | 9.9 | 14.4
28.7 |  |  |  | 9.8429 | 14.275
28.8 |  |  |  | 9.7222 | 14.1
28.9 |  |  |  | 9.6 | 13.85
29 |  |  |  |  | 13.7
29.1 |  |  |  |  | 13.5556
29.2 |  |  |  |  | 13.325
29.3 |  |  |  |  | 13.2
29.4 |  |  |  |  | 13.0429
29.5 |  |  |  |  | 12.8286
29.6 |  |  |  |  | 12.6833
29.7 |  |  |  |  | 12.5167
29.8 |  |  |  |  | 12.3714
29.9 |  |  |  |  | 12.3
30 |  |  |  |  | 12.2
30.1 |  |  |  |  | 12.0286
30.2 |  |  |  |  | 11.8625
30.3 |  |  |  |  | 11.675
30.4 |  |  |  |  | 11.5833
30.5 |  |  |  |  | 11.4286
30.6 |  |  |  |  | 11.3
30.7 |  |  |  |  | 11.2
30.8 |  |  |  |  | 11.0167
30.8 |  |  |  |  | 12.7
30.9 |  |  |  |  | 12.6

## Wedge 60 Output Factors
Output factors
Wedge 60 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15
| 3 | 0.873 | 0.888 | 0.895 | 0.904 | 0.91 | 0.917
| 4 | 0.889 | 0.906 | 0.915 | 0.926 | 0.935 | 0.943
| 5 | 0.898 | 0.916 | 0.927 | 0.94 | 0.951 | 0.961
| 7 | 0.91 | 0.931 | 0.944 | 0.961 | 0.976 | 0.989
| 10 | 0.921 | 0.945 | 0.961 | 0.982 | 1 | 1.019
| 15 | 0.931 | 0.958 | 0.976 | 1.001 | 1.025 | 1.05
| 20 | 0.936 | 0.965 | 0.985 | 1.013 | 1.041 | 1.07
| 30 | 0.944 | 0.975 | 0.997 | 1.029 | 1.061 | 1.098
| 40 | 0.948 | 0.98 | 1.003 | 1.036 | 1.072 | 1.113