## Title
Varian TrueBeam
15MV Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables |  | 2. Plots
| Open field |  | Open Field
| Open Field Depth Doses |  | Wedge 15
| Open Field Profiles at 3cm |  | Wedge 30 Plots
| Open Field Profiles at 5cm |  | Wedge 45 Plots
| Open Field Profiles at 10cm |  | Wedge 60 Plots
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
Wedge 15
Wedge 15 Depth Doses
SSD80 Wedge 15 Depth Doses
SSD120 Wedge 15 Depth Doses
Wedge 15 Lateral Profiles
Wedge 15 Longitudinal Profiles
Wedge 15 Output Factors
Wedge 30
Wedge 30 Depth Doses
SSD80 Wedge 30 Depth Doses
SSD120 Wedge 30 Depth Doses
Wedge 30 Lateral Profiles
Wedge 30 Longitudinal Profiles
Wedge 30 Output Factors
Wedge 45
Wedge 45 Depth Doses
SSD80 Wedge 45 Depth Doses
Wedge 45 Longitudinal Profiles
Wedge 45 Lateral Profiles
Wedge 45 Longitudinal Profiles
Wedge 45 Output Factors
Wedge 60
Wedge 60 Depth Doses
SSD80 Wedge 60 Depth Doses
SSD120 Wedge 60 Depth Doses
Wedge 60 Lateral Profiles
Wedge 60 Longitudinal Profiles
Wedge 60 Output Factors

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 25.3 | 26.2 | 28.7 | 31.6 | 34.4 | 48 | 57.1 | 62.2
0.1 | 29.1 | 29.8 | 32.15 | 35 | 37.8667 | 52.08 | 60.825 | 65.8
0.2 | 35.725 | 36.825 | 38.6 | 41.6 | 44.7667 | 57.9333 | 66.5 | 71.12
0.3 | 44.7 | 45.3 | 47.25 | 49.8 | 52.5 | 65 | 72.72 | 76.5
0.4 | 52.675 | 52.5 | 54.8 | 57.25 | 59.8 | 71.1 | 78.1 | 81.3667
0.5 | 59.475 | 59.35 | 61.4 | 63.4667 | 66.1333 | 76.5 | 82.2 | 85
0.6 | 65.1 | 64.9 | 67 | 69.1 | 71.2 | 80.35 | 85.875 | 88.3
0.7 | 70.38 | 70.1 | 71.2667 | 73.5 | 75.36 | 84.2 | 88.85 | 90.7
0.8 | 74.3 | 74.4 | 75.4 | 77.1333 | 79.05 | 87 | 91.14 | 92.8
0.9 | 78.1 | 77.8 | 78.94 | 80.7 | 82.1 | 89.7 | 93.175 | 94.4
1 | 81.45 | 80.9333 | 81.9333 | 83.4333 | 85.025 | 91.52 | 94.75 | 95.875
1.1 | 84.4 | 83.85 | 84.6 | 86.1667 | 87.4 | 93.4 | 95.9 | 96.8
1.2 | 86.82 | 86.1 | 86.9 | 88.3 | 89.5 | 94.8 | 96.95 | 97.76
1.3 | 88.75 | 88.3 | 89.05 | 90.3 | 91.2 | 96 | 97.9 | 98.52
1.4 | 90.85 | 90.2 | 90.9667 | 91.95 | 92.82 | 96.975 | 98.5 | 99.02
1.5 | 92.35 | 91.9 | 92.4 | 93.4 | 94.15 | 97.8 | 99.04 | 99.45
1.6 | 93.675 | 93.3333 | 93.6333 | 94.6 | 95.3 | 98.375 | 99.4 | 99.7
1.7 | 94.925 | 94.7 | 94.9 | 95.575 | 96.25 | 99 | 99.725 | 99.9
1.8 | 96.125 | 95.7 | 95.96 | 96.5 | 97.1 | 99.2333 | 99.875 | 100
1.9 | 96.88 | 96.6 | 96.725 | 97.4 | 97.74 | 99.6667 | 100 | 100
2 | 97.56 | 97.375 | 97.4667 | 97.95 | 98.35 | 99.82 | 100 | 100
2.1 | 98.25 | 97.95 | 98.0333 | 98.525 | 98.8333 | 99.9667 | 100 | 99.94
2.2 | 98.9 | 98.55 | 98.65 | 98.925 | 99.225 | 100 | 99.9 | 99.8
2.3 | 99.3333 | 98.975 | 99.0333 | 99.26 | 99.475 | 100 | 99.8 | 99.65
2.4 | 99.7333 | 99.2 | 99.325 | 99.5 | 99.675 | 100 | 99.7 | 99.54
2.5 | 99.9 | 99.62 | 99.5667 | 99.8 | 99.9 | 99.92 | 99.5 | 99.35
2.6 | 100 | 99.8 | 99.8 | 99.85 | 99.9 | 99.85 | 99.3 | 99.12
2.7 | 100 | 99.9 | 99.9 | 99.9 | 100 | 99.7 | 99.05 | 98.9
2.8 | 100 | 100 | 99.95 | 100 | 100 | 99.525 | 98.8 | 98.58
2.9 | 99.9 | 100 | 100 | 100 | 99.9 | 99.3 | 98.65 | 98.375
3 | 99.8333 | 100 | 100 | 100 | 99.8 | 99.075 | 98.3 | 98.1
3.1 | 99.8 | 100 | 99.9 | 99.9 | 99.6333 | 98.825 | 97.95 | 97.82
3.2 | 99.7 | 99.9 | 99.8 | 99.775 | 99.5 | 98.5 | 97.7 | 97.45
3.3 | 99.5 | 99.7 | 99.7 | 99.625 | 99.42 | 98.35 | 97.4 | 97.175
3.4 | 99.3667 | 99.5667 | 99.5 | 99.375 | 99.275 | 98.1 | 97.1333 | 96.94
3.5 | 99.1 | 99.35 | 99.3 | 99.225 | 99 | 97.75 | 96.85 | 96.58
3.6 | 98.74 | 99.1 | 99.125 | 98.95 | 98.7 | 97.4 | 96.45 | 96.2333
3.7 | 98.55 | 98.9 | 98.7667 | 98.75 | 98.4 | 97.12 | 96.125 | 95.9
3.8 | 98.15 | 98.6 | 98.6 | 98.5 | 98.15 | 96.825 | 95.8333 | 95.6333
3.9 | 97.875 | 98.32 | 98.4 | 98.2 | 97.9 | 96.55 | 95.45 | 95.3
4 | 97.625 | 98.025 | 98.05 | 97.9 | 97.6 | 96.15 | 95.1667 | 94.98
4.1 | 97.275 | 97.675 | 97.8 | 97.55 | 97.25 | 95.875 | 94.85 | 94.68
4.2 | 96.88 | 97.36 | 97.45 | 97.175 | 96.9 | 95.525 | 94.6 | 94.3
4.3 | 96.5 | 97.025 | 97.05 | 96.825 | 96.65 | 95.25 | 94.2333 | 94.05
4.4 | 96 | 96.55 | 96.65 | 96.675 | 96.3 | 94.86 | 93.84 | 93.75
4.5 | 95.6 | 96.125 | 96.25 | 96.325 | 95.9333 | 94.425 | 93.4333 | 93.4
4.6 | 95.2 | 95.85 | 95.925 | 95.925 | 95.45 | 94.2 | 93.2333 | 93.1
4.7 | 94.8 | 95.4 | 95.6 | 95.5 | 95.15 | 93.85 | 92.925 | 92.8
4.8 | 94.42 | 95.1 | 95.1667 | 95.25 | 94.825 | 93.525 | 92.6 | 92.4667
4.9 | 94.075 | 94.7 | 94.725 | 94.8 | 94.5 | 93.1 | 92.15 | 92.1333
5 | 93.6333 | 94.32 | 94.3667 | 94.3 | 94.1 | 92.7333 | 91.8 | 91.7
5.1 | 93.2667 | 93.95 | 94 | 94.05 | 93.82 | 92.4 | 91.6 | 91.45
5.2 | 92.8 | 93.45 | 93.65 | 93.675 | 93.3 | 92.15 | 91.3 | 91.18
5.3 | 92.4333 | 93.15 | 93.3 | 93.34 | 93 | 91.8 | 90.9 | 90.78
5.4 | 91.9667 | 92.8 | 93.0333 | 93 | 92.6333 | 91.45 | 90.55 | 90.5333
5.5 | 91.575 | 92.34 | 92.55 | 92.6 | 92.325 | 91.125 | 90.2 | 90.3
5.6 | 91.18 | 91.875 | 92.1 | 92.1 | 91.88 | 90.775 | 90 | 89.9333
5.7 | 90.64 | 91.425 | 91.76 | 91.75 | 91.6333 | 90.425 | 89.6 | 89.64
5.8 | 90.24 | 91.075 | 91.4 | 91.44 | 91.2 | 90.075 | 89.3333 | 89.42
5.9 | 89.9 | 90.66 | 90.975 | 90.975 | 90.8333 | 89.68 | 89.02 | 89.02
6 | 89.52 | 90.1 | 90.6 | 90.65 | 90.4667 | 89.4 | 88.6667 | 88.6333
6.1 | 89.05 | 89.72 | 90.2 | 90.3 | 90.1 | 89.2333 | 88.35 | 88.4
6.2 | 88.65 | 89.45 | 89.95 | 90.04 | 89.8 | 88.76 | 88.0333 | 87.96
6.3 | 88.25 | 88.95 | 89.375 | 89.6 | 89.3667 | 88.425 | 87.8 | 87.6833
6.4 | 87.875 | 88.6 | 89.15 | 89.15 | 89.1 | 88.15 | 87.475 | 87.4
6.5 | 87.45 | 88.18 | 88.7333 | 88.8 | 88.6667 | 87.825 | 87.15 | 87.05
6.6 | 87.0333 | 87.88 | 88.28 | 88.475 | 88.3 | 87.525 | 86.825 | 86.8143
6.7 | 86.5 | 87.4 | 88 | 88.1 | 87.9333 | 87.1333 | 86.525 | 86.5
6.8 | 86.1667 | 87 | 87.6 | 87.775 | 87.65 | 86.7 | 86.175 | 86.22
6.9 | 85.75 | 86.7 | 87.1 | 87.38 | 87.24 | 86.5333 | 85.94 | 85.88
7 | 85.475 | 86.3333 | 86.775 | 87 | 87 | 86.25 | 85.6 | 85.5667
7.1 | 85.15 | 85.8 | 86.525 | 86.6 | 86.65 | 85.9 | 85.3 | 85.3167
7.2 | 84.78 | 85.3333 | 85.96 | 86.32 | 86.2 | 85.5 | 84.9333 | 85.04
7.3 | 84.4 | 84.925 | 85.65 | 85.975 | 85.9 | 85.25 | 84.66 | 84.7714
7.4 | 84 | 84.675 | 85.25 | 85.6 | 85.55 | 85 | 84.35 | 84.4
7.5 | 83.5 | 84.425 | 84.875 | 85.25 | 85.24 | 84.56 | 84.05 | 84.1667
7.6 | 83.12 | 83.8 | 84.55 | 84.9 | 84.775 | 84.22 | 83.7667 | 83.84
7.7 | 82.6333 | 83.4 | 84.1333 | 84.5 | 84.525 | 83.975 | 83.3667 | 83.4667
7.8 | 82.275 | 83.1 | 83.8 | 84.1 | 84.2 | 83.64 | 83.15 | 83.22
7.9 | 81.96 | 82.7333 | 83.4333 | 83.85 | 83.68 | 83.2667 | 82.9 | 82.9667
8 | 81.55 | 82.4 | 83.075 | 83.425 | 83.45 | 83.025 | 82.55 | 82.7286
8.1 | 81.2 | 82 | 82.7 | 82.975 | 83.1 | 82.725 | 82.225 | 82.35
8.2 | 80.8 | 81.5 | 82.3 | 82.7 | 82.82 | 82.3 | 81.875 | 82.1
8.3 | 80.45 | 81.1667 | 81.9 | 82.4 | 82.45 | 82 | 81.6 | 81.84
8.4 | 79.975 | 80.8 | 81.5 | 82.02 | 82.15 | 81.725 | 81.3 | 81.5
8.5 | 79.55 | 80.425 | 81.2 | 81.7 | 81.75 | 81.5 | 81.05 | 81.22
8.6 | 79.18 | 80.05 | 81 | 81.35 | 81.38 | 81.1 | 80.7 | 80.9667
8.7 | 78.85 | 79.775 | 80.55 | 81.025 | 81.05 | 80.675 | 80.35 | 80.7
8.8 | 78.5 | 79.325 | 80.25 | 80.55 | 80.8 | 80.425 | 80.1 | 80.3167
8.9 | 78 | 79 | 79.825 | 80.225 | 80.35 | 80.1667 | 79.875 | 80.0333
9 | 77.7 | 78.7 | 79.5 | 79.9 | 80.02 | 79.9 | 79.525 | 79.74
9.1 | 77.275 | 78.3 | 79.15 | 79.65 | 79.675 | 79.6 | 79.2 | 79.52
9.2 | 76.95 | 77.75 | 78.7 | 79.2 | 79.425 | 79.2 | 78.9 | 79.16
9.3 | 76.575 | 77.525 | 78.35 | 78.95 | 79.025 | 78.96 | 78.65 | 78.82
9.4 | 76.14 | 77.15 | 78.025 | 78.64 | 78.65 | 78.6667 | 78.3 | 78.62
9.5 | 75.8 | 76.625 | 77.65 | 78.2333 | 78.4 | 78.3667 | 78.1 | 78.3
9.6 | 75.4 | 76.3 | 77.325 | 77.825 | 78 | 78.0333 | 77.84 | 77.9857
9.7 | 75 | 76 | 76.875 | 77.475 | 77.7 | 77.7 | 77.55 | 77.7833
9.8 | 74.65 | 75.6 | 76.55 | 77.075 | 77.4 | 77.44 | 77.15 | 77.45
9.9 | 74.35 | 75.3 | 76.3 | 76.8 | 77 | 77.1 | 76.875 | 77.18
10 | 73.85 | 75.025 | 76 | 76.5 | 76.675 | 76.8 | 76.6 | 76.925
10.1 | 73.7 | 74.6 | 75.6 | 76.15 | 76.425 | 76.5 | 76.35 | 76.6
10.2 | 73.4 | 74.18 | 75.15 | 75.9 | 75.975 | 76.1667 | 76 | 76.4
10.3 | 73 | 73.88 | 74.9 | 75.5 | 75.725 | 75.925 | 75.75 | 76.0667
10.4 | 72.6667 | 73.5 | 74.55 | 75.1 | 75.4667 | 75.575 | 75.425 | 75.82
10.5 | 72.4333 | 73.1 | 74.225 | 74.8333 | 75.1 | 75.3 | 75.175 | 75.55
10.6 | 72.15 | 72.85 | 73.875 | 74.6 | 74.75 | 75 | 74.825 | 75.28
10.7 | 71.75 | 72.4667 | 73.55 | 74.2333 | 74.4 | 74.75 | 74.66 | 75.0333
10.8 | 71.325 | 72.225 | 73.22 | 73.82 | 74.15 | 74.4 | 74.35 | 74.675
10.9 | 71 | 71.75 | 72.85 | 73.5667 | 73.775 | 74.16 | 74.1 | 74.4667
11 | 70.7 | 71.4 | 72.5 | 73.275 | 73.48 | 73.925 | 73.8 | 74.1333
11.1 | 70.4 | 71.2 | 72.2 | 72.85 | 73.225 | 73.675 | 73.55 | 73.9
11.2 | 70 | 70.9 | 71.92 | 72.575 | 72.875 | 73.34 | 73.3 | 73.6333
11.3 | 69.675 | 70.55 | 71.65 | 72.225 | 72.55 | 73 | 72.95 | 73.3286
11.4 | 69.325 | 70.2 | 71.3 | 71.95 | 72.3 | 72.75 | 72.725 | 73.0667
11.5 | 69.15 | 69.95 | 71.05 | 71.7333 | 71.95 | 72.5 | 72.475 | 72.8167
11.6 | 68.65 | 69.525 | 70.6 | 71.375 | 71.65 | 72.2333 | 72.225 | 72.5
11.7 | 68.38 | 69.15 | 70.32 | 71.0333 | 71.275 | 71.975 | 71.9 | 72.22
11.8 | 68.1 | 68.7333 | 70.025 | 70.6667 | 70.98 | 71.625 | 71.7 | 71.94
11.9 | 67.7 | 68.48 | 69.675 | 70.3667 | 70.7 | 71.3667 | 71.35 | 71.7
12 | 67.4 | 68.25 | 69.35 | 70.075 | 70.5 | 71.125 | 71.1 | 71.4667
12.1 | 67 | 67.9 | 69.0667 | 69.85 | 70.3 | 70.8 | 70.8333 | 71.225
12.2 | 66.75 | 67.6 | 68.7 | 69.45 | 69.875 | 70.5 | 70.575 | 70.9667
12.3 | 66.54 | 67.3 | 68.4667 | 69.15 | 69.675 | 70.2 | 70.25 | 70.7167
12.4 | 66.175 | 67.05 | 68.2 | 68.8667 | 69.36 | 69.9667 | 69.975 | 70.38
12.5 | 65.825 | 66.7 | 67.8333 | 68.6 | 69 | 69.7 | 69.8 | 70.1
12.6 | 65.55 | 66.375 | 67.45 | 68.2 | 68.65 | 69.425 | 69.5 | 69.9333
12.7 | 65.15 | 66.125 | 67.2 | 68 | 68.34 | 69.175 | 69.25 | 69.66
12.8 | 64.9 | 65.6667 | 66.95 | 67.7 | 68.2 | 68.925 | 69 | 69.525
12.9 | 64.55 | 65.4 | 66.7 | 67.4 | 67.85 | 68.58 | 68.75 | 69.14
13 | 64.35 | 65.1 | 66.45 | 67.1667 | 67.5 | 68.3 | 68.5 | 68.92
13.1 | 63.9 | 64.8 | 66.15 | 66.775 | 67.275 | 68.1 | 68.175 | 68.7
13.2 | 63.7 | 64.5 | 65.625 | 66.425 | 66.85 | 67.75 | 67.925 | 68.4714
13.3 | 63.45 | 64.225 | 65.35 | 66.2 | 66.7 | 67.5 | 67.575 | 68.2333
13.4 | 63.15 | 63.95 | 65.05 | 65.9 | 66.375 | 67.325 | 67.425 | 67.92
13.5 | 62.8667 | 63.65 | 64.8 | 65.6 | 66.15 | 66.975 | 67.2 | 67.6667
13.6 | 62.575 | 63.25 | 64.525 | 65.35 | 65.8 | 66.675 | 66.9 | 67.4667
13.7 | 62.24 | 63 | 64.175 | 65.15 | 65.5667 | 66.425 | 66.6333 | 67.2
13.8 | 61.98 | 62.66 | 63.925 | 64.75 | 65.3333 | 66.2 | 66.4667 | 66.9
13.9 | 61.65 | 62.425 | 63.675 | 64.4 | 65.05 | 65.95 | 66.2 | 66.7167
14 | 61.2 | 61.975 | 63.425 | 64.225 | 64.65 | 65.7 | 65.9 | 66.45
14.1 | 60.9667 | 61.8 | 63.15 | 63.8 | 64.475 | 65.42 | 65.7 | 66.2
14.2 | 60.825 | 61.5 | 62.7667 | 63.6333 | 64.175 | 65.275 | 65.5 | 65.9667
14.3 | 60.55 | 61.25 | 62.4 | 63.4 | 63.9 | 64.925 | 65.275 | 65.7333
14.4 | 60.125 | 61 | 62.25 | 63.1 | 63.625 | 64.575 | 65 | 65.5
14.5 | 59.84 | 60.76 | 61.9333 | 62.8 | 63.38 | 64.4 | 64.7 | 65.18
14.6 | 59.575 | 60.4 | 61.7 | 62.45 | 63.05 | 64.15 | 64.45 | 64.975
14.7 | 59.3333 | 60.2 | 61.3 | 62.22 | 62.8 | 63.8 | 64.2 | 64.75
14.8 | 59.1 | 59.875 | 60.95 | 61.85 | 62.55 | 63.575 | 63.96 | 64.48
14.9 | 58.75 | 59.625 | 60.7 | 61.625 | 62.2 | 63.38 | 63.7 | 64.2833
15 | 58.5 | 59.2667 | 60.425 | 61.3 | 61.975 | 63.1 | 63.475 | 64.1
15.1 | 58.275 | 59 | 60.2667 | 61.15 | 61.725 | 62.85 | 63.3 | 63.76
15.2 | 57.98 | 58.65 | 60 | 60.8 | 61.45 | 62.6 | 62.98 | 63.55
15.3 | 57.6 | 58.4 | 59.75 | 60.5 | 61.14 | 62.425 | 62.74 | 63.2833
15.4 | 57.35 | 58.15 | 59.5 | 60.3 | 61 | 62.175 | 62.55 | 63
15.5 | 57.1 | 57.9 | 59.2 | 60.05 | 60.58 | 61.925 | 62.3 | 62.8167
15.6 | 56.92 | 57.575 | 58.85 | 59.8 | 60.3 | 61.6 | 61.96 | 62.55
15.7 | 56.675 | 57.325 | 58.65 | 59.45 | 60.05 | 61.45 | 61.725 | 62.36
15.8 | 56.325 | 57.075 | 58.375 | 59.275 | 59.82 | 61.2 | 61.575 | 62.0667
15.9 | 56.1 | 56.8 | 58.1667 | 59.025 | 59.575 | 60.9 | 61.34 | 61.9
16 | 55.8 | 56.65 | 57.85 | 58.6667 | 59.325 | 60.7 | 61.2 | 61.6667
16.1 | 55.5 | 56.3 | 57.52 | 58.475 | 59 | 60.42 | 60.9333 | 61.3833
16.2 | 55.2 | 56.06 | 57.275 | 58.24 | 58.7667 | 60.225 | 60.6 | 61.2167
16.3 | 54.92 | 55.825 | 57.125 | 57.9333 | 58.5 | 59.975 | 60.4 | 60.94
16.4 | 54.725 | 55.575 | 56.8 | 57.65 | 58.3 | 59.725 | 60.16 | 60.75
16.5 | 54.575 | 55.25 | 56.54 | 57.5 | 58 | 59.5 | 59.925 | 60.54
16.6 | 54.325 | 54.975 | 56.3 | 57.275 | 57.825 | 59.25 | 59.675 | 60.2571
16.7 | 53.975 | 54.74 | 56 | 57.025 | 57.575 | 59.1 | 59.5 | 60.1
16.8 | 53.725 | 54.5 | 55.7 | 56.775 | 57.325 | 58.775 | 59.2333 | 59.8667
16.9 | 53.5 | 54.32 | 55.525 | 56.45 | 57 | 58.525 | 59 | 59.5857
17 | 53.3 | 53.95 | 55.2 | 56.25 | 56.825 | 58.275 | 58.85 | 59.3667
17.1 | 53 | 53.7 | 55.05 | 55.8667 | 56.5667 | 58.1 | 58.6 | 59.16
17.2 | 52.7 | 53.4667 | 54.725 | 55.6667 | 56.4 | 57.85 | 58.325 | 58.8833
17.3 | 52.4667 | 53.2 | 54.48 | 55.5 | 56.15 | 57.6 | 58.1 | 58.7167
17.4 | 52.25 | 52.95 | 54.25 | 55.25 | 55.875 | 57.36 | 57.9 | 58.55
17.5 | 52.05 | 52.7 | 54 | 54.925 | 55.6 | 57.125 | 57.7 | 58.25
17.6 | 51.8 | 52.55 | 53.75 | 54.675 | 55.45 | 56.95 | 57.5667 | 58
17.7 | 51.54 | 52.3 | 53.575 | 54.5 | 55.225 | 56.7 | 57.275 | 57.82
17.8 | 51.4 | 52.15 | 53.325 | 54.26 | 54.95 | 56.5 | 57.05 | 57.7
17.9 | 51.05 | 51.825 | 53.2 | 54.05 | 54.7 | 56.3 | 56.9 | 57.4667
18 | 50.8 | 51.575 | 52.95 | 53.8 | 54.45 | 56.05 | 56.65 | 57.26
18.1 | 50.65 | 51.325 | 52.6 | 53.55 | 54.3 | 55.9 | 56.4 | 56.9833
18.2 | 50.4 | 51.175 | 52.4 | 53.42 | 53.95 | 55.55 | 56.2 | 56.8
18.3 | 50.1667 | 50.9 | 52.1 | 53.075 | 53.8 | 55.325 | 56.025 | 56.6167
18.4 | 49.9 | 50.7 | 51.95 | 52.775 | 53.45 | 55.175 | 55.78 | 56.35
18.5 | 49.7667 | 50.55 | 51.65 | 52.625 | 53.2 | 54.925 | 55.55 | 56.18
18.6 | 49.45 | 50.2 | 51.375 | 52.4 | 53.1 | 54.68 | 55.3667 | 55.96
18.7 | 49.2 | 49.96 | 51.175 | 52.15 | 52.9 | 54.5 | 55.2 | 55.7667
18.8 | 49 | 49.7 | 50.925 | 51.9 | 52.6 | 54.4 | 54.95 | 55.5
18.9 | 48.7 | 49.475 | 50.775 | 51.7 | 52.4 | 54.05 | 54.7 | 55.38
19 | 48.525 | 49.2 | 50.5 | 51.525 | 52.2 | 53.9 | 54.5 | 55.175
19.1 | 48.275 | 49.025 | 50.25 | 51.275 | 51.9333 | 53.65 | 54.3 | 54.95
19.2 | 48.05 | 48.825 | 50 | 51.025 | 51.7 | 53.4 | 54.1 | 54.675
19.3 | 47.9 | 48.5667 | 49.76 | 50.7 | 51.6 | 53.175 | 53.86 | 54.55
19.4 | 47.7 | 48.275 | 49.525 | 50.55 | 51.275 | 53 | 53.6 | 54.24
19.5 | 47.425 | 48.0667 | 49.375 | 50.3 | 51.0333 | 52.775 | 53.4667 | 54.05
19.6 | 47.225 | 47.9 | 49.125 | 50.075 | 50.9 | 52.6 | 53.2 | 53.84
19.7 | 47.025 | 47.675 | 48.85 | 49.9 | 50.54 | 52.4 | 53 | 53.6667
19.8 | 46.8 | 47.525 | 48.6 | 49.6333 | 50.4 | 52.15 | 52.85 | 53.5
19.9 | 46.65 | 47.2 | 48.4 | 49.4 | 50.15 | 52 | 52.675 | 53.2667
20 | 46.34 | 47.04 | 48.2 | 49.15 | 49.9 | 51.7333 | 52.4 | 53.0167
20.1 | 46.2 | 46.8 | 48 | 48.96 | 49.7 | 51.55 | 52.2 | 52.85
20.2 | 45.85 | 46.55 | 47.825 | 48.825 | 49.525 | 51.36 | 52.04 | 52.7
20.3 | 45.7 | 46.3 | 47.575 | 48.575 | 49.3 | 51.1 | 51.875 | 52.5
20.4 | 45.475 | 46.2 | 47.38 | 48.325 | 49.125 | 50.9 | 51.64 | 52.3
20.5 | 45.225 | 45.85 | 47.2 | 48.2 | 48.875 | 50.675 | 51.4 | 52.0667
20.6 | 44.975 | 45.675 | 47.1 | 47.95 | 48.55 | 50.5 | 51.2 | 51.925
20.7 | 44.825 | 45.425 | 46.8 | 47.7 | 48.4 | 50.2 | 51 | 51.7
20.8 | 44.58 | 45.275 | 46.6 | 47.56 | 48.25 | 50.0333 | 50.7667 | 51.5
20.9 | 44.45 | 45.05 | 46.375 | 47.4 | 47.975 | 49.9 | 50.58 | 51.3167
21 | 44.2 | 44.9 | 46.125 | 47.175 | 47.8 | 49.72 | 50.425 | 51.1
21.1 | 44.0333 | 44.7 | 45.95 | 46.925 | 47.625 | 49.575 | 50.275 | 50.9571
21.2 | 43.8 | 44.55 | 45.7333 | 46.7 | 47.475 | 49.325 | 50.125 | 50.74
21.3 | 43.625 | 44.3 | 45.55 | 46.5 | 47.24 | 49.075 | 49.8667 | 50.58
21.4 | 43.4667 | 44.1 | 45.375 | 46.25 | 47 | 48.94 | 49.7 | 50.3833
21.5 | 43.3 | 43.925 | 45.125 | 46.1 | 46.85 | 48.7 | 49.475 | 50.1333
21.6 | 43.05 | 43.675 | 44.975 | 45.9 | 46.6 | 48.55 | 49.225 | 49.95
21.7 | 42.9 | 43.5 | 44.825 | 45.6 | 46.4 | 48.4 | 49.1 | 49.78
21.8 | 42.7 | 43.3 | 44.6 | 45.475 | 46.225 | 48.1 | 48.9 | 49.5833
21.9 | 42.6 | 43.125 | 44.35 | 45.325 | 46 | 47.925 | 48.7 | 49.4167
22 | 42.3 | 42.85 | 44.1 | 45.1 | 45.825 | 47.8 | 48.56 | 49.2
22.1 | 42.125 | 42.7 | 43.875 | 44.9 | 45.675 | 47.575 | 48.325 | 49.1
22.2 | 41.9 | 42.56 | 43.725 | 44.65 | 45.45 | 47.375 | 48.175 | 48.9167
22.3 | 41.74 | 42.4 | 43.4667 | 44.5 | 45.2 | 47.175 | 48 | 48.62
22.4 | 41.5 | 42.22 | 43.3 | 44.275 | 45.05 | 47 | 47.8 | 48.5
22.5 | 41.3 | 41.975 | 43.08 | 44.025 | 44.8 | 46.8 | 47.58 | 48.3
22.6 | 41.1 | 41.8 | 42.9 | 43.975 | 44.675 | 46.6 | 47.4 | 48.18
22.7 | 40.85 | 41.6 | 42.75 | 43.7 | 44.525 | 46.4 | 47.2 | 48
22.8 | 40.72 | 41.325 | 42.6 | 43.5 | 44.375 | 46.225 | 47 | 47.78
22.9 | 40.575 | 41.1 | 42.35 | 43.35 | 44.125 | 46.1 | 46.9 | 47.5667
23 | 40.34 | 40.9667 | 42.2 | 43.2 | 43.875 | 45.925 | 46.775 | 47.4
23.1 | 40.2 | 40.8 | 41.975 | 42.975 | 43.725 | 45.675 | 46.425 | 47.25
23.2 | 40 | 40.625 | 41.7333 | 42.825 | 43.55 | 45.5 | 46.3 | 47.0833
23.3 | 39.9 | 40.475 | 41.58 | 42.6 | 43.3 | 45.3 | 46.2 | 46.9
23.4 | 39.72 | 40.3 | 41.45 | 42.45 | 43.175 | 45.1 | 46 | 46.78
23.5 | 39.575 | 40.1 | 41.2 | 42.175 | 42.925 | 45 | 45.8 | 46.5
23.6 | 39.325 | 39.9 | 41.0333 | 42.025 | 42.78 | 44.825 | 45.6333 | 46.3333
23.7 | 39.1 | 39.7 | 40.86 | 41.86 | 42.6 | 44.7 | 45.4 | 46.1667
23.8 | 39 | 39.5 | 40.725 | 41.65 | 42.4 | 44.45 | 45.3 | 46
23.9 | 38.8 | 39.35 | 40.425 | 41.5 | 42.2 | 44.2 | 45.1 | 45.85
24 | 38.675 | 39.125 | 40.3 | 41.325 | 42.1 | 44.1 | 44.92 | 45.6333
24.1 | 38.425 | 38.975 | 40.1 | 41.1 | 41.9 | 43.925 | 44.775 | 45.5
24.2 | 38.2 | 38.9 | 39.9 | 41 | 41.7 | 43.7 | 44.6 | 45.26
24.3 | 38.1 | 38.66 | 39.8 | 40.775 | 41.525 | 43.525 | 44.4 | 45.16
24.4 | 37.95 | 38.4 | 39.625 | 40.6 | 41.4 | 43.325 | 44.25 | 44.9667
24.5 | 37.7333 | 38.3 | 39.4 | 40.4 | 41.15 | 43.15 | 44.1 | 44.8
24.6 | 37.6 | 38.1 | 39.3 | 40.2 | 41 | 43 | 43.8333 | 44.5333
24.7 | 37.48 | 38 | 39.1 | 40.1 | 40.8 | 42.85 | 43.7 | 44.4
24.8 | 37.3 | 37.825 | 39 | 39.925 | 40.625 | 42.65 | 43.6 | 44.3
24.9 | 37.1 | 37.6 | 38.7 | 39.775 | 40.475 | 42.5 | 43.375 | 44.1
25 | 36.9 | 37.5 | 38.6 | 39.525 | 40.3 | 42.325 | 43.2 | 43.8857
25.1 | 36.8 | 37.375 | 38.38 | 39.3333 | 40.175 | 42.175 | 43 | 43.8
25.2 | 36.6 | 37.125 | 38.25 | 39.2 | 40 | 41.975 | 42.92 | 43.6667
25.3 | 36.4 | 37 | 38 | 39.05 | 39.775 | 41.8 | 42.675 | 43.4833
25.4 | 36.3 | 36.8 | 37.8667 | 38.8 | 39.54 | 41.6 | 42.5 | 43.2333
25.5 | 36.1 | 36.6 | 37.7 | 38.75 | 39.4 | 41.425 | 42.375 | 43.15
25.6 | 35.9 | 36.5 | 37.525 | 38.52 | 39.24 | 41.3 | 42.2 | 42.96
25.7 | 35.75 | 36.3 | 37.475 | 38.375 | 39.1 | 41.2 | 42 | 42.8
25.8 | 35.6 | 36.1 | 37.3 | 38.15 | 38.95 | 41 | 41.85 | 42.675
25.9 | 35.46 | 36 | 37.1 | 38 | 38.72 | 40.88 | 41.7 | 42.4667
26 | 35.3 | 35.75 | 36.9 | 37.85 | 38.6 | 40.65 | 41.56 | 42.375
26.1 | 35.1 | 35.6 | 36.7 | 37.7 | 38.425 | 40.5 | 41.425 | 42.1714
26.2 | 34.925 | 35.45 | 36.575 | 37.5 | 38.3 | 40.35 | 41.2 | 42.0167
26.3 | 34.8 | 35.3 | 36.4 | 37.4 | 38.1 | 40.2 | 41.1 | 41.8167
26.4 | 34.6 | 35.2 | 36.275 | 37.26 | 37.9 | 40.05 | 40.925 | 41.62
26.5 | 34.45 | 35.025 | 36.1 | 37.025 | 37.8 | 39.825 | 40.8 | 41.5
26.6 | 34.325 | 34.875 | 36 | 36.875 | 37.62 | 39.675 | 40.65 | 41.36
26.7 | 34.1 | 34.7 | 35.85 | 36.725 | 37.475 | 39.525 | 40.5 | 41.1857
26.8 | 34 | 34.5 | 35.625 | 36.6 | 37.225 | 39.4 | 40.3 | 41.1
26.9 | 33.875 | 34.4 | 35.4 | 36.4 | 37.1333 | 39.3 | 40.1667 | 40.92
27 | 33.7 | 34.3 | 35.22 | 36.2 | 36.975 | 39.1 | 40 | 40.7
27.1 | 33.6 | 34.1 | 35.125 | 36.1 | 36.8 | 38.9 | 39.825 | 40.5429
27.2 | 33.45 | 33.9 | 35 | 35.92 | 36.7 | 38.8 | 39.7 | 40.4833
27.3 | 33.3 | 33.8 | 34.85 | 35.775 | 36.5 | 38.6 | 39.5 | 40.28
27.4 | 33.1 | 33.625 | 34.7 | 35.625 | 36.35 | 38.425 | 39.4 | 40.16
27.5 | 33 | 33.5 | 34.6 | 35.5 | 36.2667 | 38.275 | 39.2 | 40
27.6 | 32.8333 | 33.3 | 34.3667 | 35.4 | 36.1 | 38.125 | 39.04 | 39.84
27.7 | 32.65 | 33.15 | 34.2 | 35.15 | 35.95 | 38 | 38.9 | 39.7333
27.8 | 32.6 | 33 | 34.075 | 35 | 35.8 | 37.85 | 38.8 | 39.5143
27.9 | 32.4 | 32.9 | 33.94 | 34.9 | 35.6 | 37.6 | 38.66 | 39.4
28 | 32.2667 | 32.725 | 33.8 | 34.725 | 35.4 | 37.5 | 38.475 | 39.2
28.1 | 32.1 | 32.6 | 33.65 | 34.6 | 35.25 | 37.4 | 38.3 | 39.1
28.2 | 32 | 32.5 | 33.5 | 34.5 | 35.1 | 37.2 | 38.2 | 38.9429
28.3 | 31.825 | 32.3 | 33.35 | 34.3 | 34.975 | 37.125 | 38 | 38.8
28.4 | 31.7 | 32.2 | 33.2 | 34.1 | 34.825 | 36.975 | 37.925 | 38.675
28.5 | 31.6 | 32 | 33.05 | 33.95 | 34.7333 | 36.74 | 37.7667 | 38.525
28.6 | 31.45 | 31.9 | 32.92 | 33.82 | 34.6 | 36.7 | 37.6 | 38.3429
28.7 | 31.3 | 31.75 | 32.8 | 33.7 | 34.4667 | 36.45 | 37.5 | 38.28
28.8 | 31.15 | 31.6 | 32.64 | 33.6 | 34.3 | 36.325 | 37.325 | 38.1
28.9 | 31.025 | 31.4 | 32.5 | 33.4667 | 34.125 | 36.2 | 37.2 | 37.94
29 | 30.9 | 31.3 | 32.35 | 33.2333 | 34 | 36.1 | 37 | 37.8
29.1 | 30.725 | 31.175 | 32.2 | 33.1 | 33.825 | 35.9667 | 36.86 | 37.6
29.2 | 30.6 | 31.025 | 32.075 | 33 | 33.7 | 35.775 | 36.7667 | 37.46
29.3 | 30.5 | 30.88 | 31.925 | 32.85 | 33.55 | 35.6667 | 36.575 | 37.3
29.4 | 30.35 | 30.8 | 31.775 | 32.7 | 33.4 | 35.5 | 36.425 | 37.2
29.5 | 30.2 | 30.65 | 31.625 | 32.575 | 33.26 | 35.4 | 36.3 | 37
29.6 | 29.96 | 30.56 | 31.6 | 32.425 | 33.125 | 35.2 | 36.2 | 36.9
29.7 | 29.825 | 30.4 | 31.4 | 32.275 | 32.975 | 35.1 | 36.04 | 36.8
29.8 | 29.775 | 30.2 | 31.25 | 32.1 | 32.825 | 35 | 35.9333 | 36.68
29.9 | 29.625 | 30.1 | 31.125 | 32 | 32.7 | 34.8 | 35.7 | 36.6
30 | 29.5 | 29.975 | 31 | 31.8333 | 32.6 | 34.675 | 35.6 | 36.38

## Open Field Profiles at 3cm
Lateral Profiles
Open field
SSD=100cm
Depth=3cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.5 | [7 empty] | 3.9
-25.4 | [7 empty] | 4
-25.3 | [7 empty] | 4.1
-25.2 | [7 empty] | 4.2
-25.1 | [7 empty] | 4.3
-25 | [7 empty] | 4.4
-24.9 | [7 empty] | 4.55
-24.8 | [7 empty] | 4.6
-24.7 | [7 empty] | 4.75
-24.6 | [7 empty] | 4.875
-24.5 | [7 empty] | 5
-24.4 | [7 empty] | 5.1167
-24.3 | [7 empty] | 5.2833
-24.2 | [7 empty] | 5.4
-24.1 | [7 empty] | 5.5143
-24 | [7 empty] | 5.7
-23.9 | [7 empty] | 5.8667
-23.8 | [7 empty] | 6.0333
-23.7 | [7 empty] | 6.2
-23.6 | [7 empty] | 6.4
-23.5 | [7 empty] | 6.5571
-23.4 | [7 empty] | 6.725
-23.3 | [7 empty] | 6.9167
-23.2 | [7 empty] | 7.0833
-23.1 | [7 empty] | 7.3
-23 | [7 empty] | 7.575
-22.9 | [7 empty] | 7.7
-22.8 | [7 empty] | 7.9429
-22.7 | [7 empty] | 8.18
-22.6 | [7 empty] | 8.45
-22.5 | [7 empty] | 8.6667
-22.4 | [7 empty] | 9
-22.3 | [7 empty] | 9.26
-22.2 | [7 empty] | 9.52
-22.1 | [7 empty] | 9.9
-22 | [7 empty] | 10.325
-21.9 | [7 empty] | 10.66
-21.8 | [7 empty] | 11.16
-21.7 | [7 empty] | 11.78
-21.6 | [7 empty] | 12.5
-21.5 | [7 empty] | 13.36
-21.4 | [7 empty] | 14.5
-21.3 | [7 empty] | 15.8667
-21.2 | [7 empty] | 17.96
-21.1 | [7 empty] | 20.8
-21 | [7 empty] | 25
-20.9 | [7 empty] | 31.6
-20.8 | [7 empty] | 40.3667
-20.7 | [7 empty] | 50.2
-20.6 | [7 empty] | 60.9667
-20.5 | [7 empty] | 70.8
-20.4 | [6 empty] | 4.2 | 80.16
-20.3 | [6 empty] | 4.3143 | 87.6833
-20.2 | [6 empty] | 4.4 | 93.0167
-20.1 | [6 empty] | 4.5 | 96.0429
-20 | [6 empty] | 4.6 | 98.15
-19.9 | [6 empty] | 4.7 | 99.5667
-19.8 | [6 empty] | 4.8 | 100.6
-19.7 | [6 empty] | 4.9 | 101.4
-19.6 | [6 empty] | 5.0333 | 102.06
-19.5 | [6 empty] | 5.1333 | 102.54
-19.4 | [6 empty] | 5.3 | 102.96
-19.3 | [6 empty] | 5.4 | 103.425
-19.2 | [6 empty] | 5.5 | 103.76
-19.1 | [6 empty] | 5.65 | 103.98
-19 | [6 empty] | 5.7286 | 104.25
-18.9 | [6 empty] | 5.8833 | 104.35
-18.8 | [6 empty] | 6.05 | 104.58
-18.7 | [6 empty] | 6.2 | 104.7
-18.6 | [6 empty] | 6.3 | 104.8
-18.5 | [6 empty] | 6.5 | 104.8
-18.4 | [6 empty] | 6.62 | 104.9
-18.3 | [6 empty] | 6.75 | 104.9
-18.2 | [6 empty] | 6.975 | 104.8
-18.1 | [6 empty] | 7.1 | 104.8
-18 | [6 empty] | 7.3 | 104.8
-17.9 | [6 empty] | 7.4833 | 104.7
-17.8 | [6 empty] | 7.68 | 104.7
-17.7 | [6 empty] | 7.8 | 104.7
-17.6 | [6 empty] | 8.0167 | 104.7
-17.5 | [6 empty] | 8.2667 | 104.7
-17.4 | [6 empty] | 8.4667 | 104.7
-17.3 | [6 empty] | 8.7333 | 104.8
-17.2 | [6 empty] | 8.94 | 104.8
-17.1 | [6 empty] | 9.2 | 104.8667
-17 | [6 empty] | 9.5 | 104.9
-16.9 | [6 empty] | 9.84 | 104.9
-16.8 | [6 empty] | 10.3 | 105
-16.7 | [6 empty] | 10.7 | 105.1
-16.6 | [6 empty] | 11.24 | 105.1
-16.5 | [6 empty] | 11.8333 | 105.2
-16.4 | [6 empty] | 12.5286 | 105.2
-16.3 | [6 empty] | 13.55 | 105.3
-16.2 | [6 empty] | 14.84 | 105.4
-16.1 | [6 empty] | 16.35 | 105.4
-16 | [6 empty] | 18.6 | 105.4667
-15.9 | [6 empty] | 21.8 | 105.5
-15.8 | [6 empty] | 26.76 | 105.5
-15.7 | [6 empty] | 34.2167 | 105.5
-15.6 | [6 empty] | 43.05 | 105.5
-15.5 | [6 empty] | 53.3429 | 105.5
-15.4 | [6 empty] | 63.74 | 105.5
-15.3 | [5 empty] | 3.2 | 73.74 | 105.5
-15.2 | [5 empty] | 3.3 | 82.24 | 105.5
-15.1 | [5 empty] | 3.375 | 88.8 | 105.5
-15 | [5 empty] | 3.4333 | 93.0333 | 105.5
-14.9 | [5 empty] | 3.5 | 95.575 | 105.5
-14.8 | [5 empty] | 3.6 | 97.38 | 105.4
-14.7 | [5 empty] | 3.7 | 98.7 | 105.4
-14.6 | [5 empty] | 3.75 | 99.62 | 105.4
-14.5 | [5 empty] | 3.825 | 100.425 | 105.4
-14.4 | [5 empty] | 3.9333 | 101.0667 | 105.3
-14.3 | [5 empty] | 4 | 101.4 | 105.3
-14.2 | [5 empty] | 4.1333 | 101.825 | 105.3
-14.1 | [5 empty] | 4.2 | 102.05 | 105.2833
-14 | [5 empty] | 4.35 | 102.2667 | 105.2
-13.9 | [5 empty] | 4.5 | 102.5333 | 105.2
-13.8 | [5 empty] | 4.6 | 102.7 | 105.2
-13.7 | [5 empty] | 4.7 | 102.8 | 105.2
-13.6 | [5 empty] | 4.8 | 103 | 105.2
-13.5 | [5 empty] | 4.9667 | 103.0857 | 105.2
-13.4 | [5 empty] | 5.1 | 103.2 | 105.2
-13.3 | [5 empty] | 5.2 | 103.3 | 105.24
-13.2 | [5 empty] | 5.4 | 103.4 | 105.3
-13.1 | [5 empty] | 5.5 | 103.46 | 105.3
-13 | [5 empty] | 5.675 | 103.6 | 105.3
-12.9 | [5 empty] | 5.84 | 103.7 | 105.3
-12.8 | [5 empty] | 6 | 103.8 | 105.3
-12.7 | [5 empty] | 6.2 | 103.8 | 105.3
-12.6 | [5 empty] | 6.3667 | 103.9 | 105.3
-12.5 | [5 empty] | 6.575 | 103.9 | 105.3
-12.4 | [5 empty] | 6.725 | 104 | 105.3
-12.3 | [5 empty] | 6.9667 | 104 | 105.3
-12.2 | [5 empty] | 7.2 | 104.1 | 105.3
-12.1 | [5 empty] | 7.42 | 104.1 | 105.3
-12 | [5 empty] | 7.7 | 104.1 | 105.3
-11.9 | [5 empty] | 7.9 | 104.2 | 105.3
-11.8 | [5 empty] | 8.3 | 104.2 | 105.3
-11.7 | [5 empty] | 8.7 | 104.2 | 105.25
-11.6 | [5 empty] | 9.0667 | 104.2 | 105.2
-11.5 | [5 empty] | 9.55 | 104.3 | 105.2
-11.4 | [5 empty] | 10.14 | 104.3 | 105.2
-11.3 | [5 empty] | 10.8 | 104.3 | 105.2
-11.2 | [5 empty] | 11.675 | 104.3 | 105.2
-11.1 | [5 empty] | 12.7 | 104.3 | 105.1333
-11 | [5 empty] | 14.1 | 104.3 | 105.0667
-10.9 | [5 empty] | 16 | 104.3 | 105
-10.8 | [5 empty] | 18.7 | 104.3 | 105
-10.7 | [5 empty] | 22.46 | 104.3 | 104.9
-10.6 | [5 empty] | 28.3 | 104.3 | 104.875
-10.5 | [5 empty] | 36.35 | 104.3 | 104.8
-10.4 | [5 empty] | 46.3 | 104.2 | 104.65
-10.3 | [5 empty] | 56.5 | 104.12 | 104.5833
-10.2 | [5 empty] | 67.2 | 104.0333 | 104.5
-10.1 |  |  |  |  | 1.8 | 76.1333 | 104 | 104.4
-10 |  |  |  |  | 1.9 | 83.54 | 103.9 | 104.4
-9.9 |  |  |  |  | 1.9 | 89.175 | 103.9 | 104.4
-9.8 |  |  |  |  | 2 | 92.36 | 103.9 | 104.3
-9.7 |  |  |  |  | 2 | 94.7 | 103.9 | 104.3
-9.6 |  |  |  |  | 2.1 | 96.35 | 103.9 | 104.3
-9.5 |  |  |  |  | 2.1 | 97.6 | 103.9 | 104.3
-9.4 |  |  |  |  | 2.2 | 98.6 | 103.9 | 104.3
-9.3 |  |  |  |  | 2.2 | 99.3 | 103.9 | 104.3
-9.2 |  |  |  |  | 2.3 | 100 | 103.9 | 104.4
-9.1 |  |  |  |  | 2.3 | 100.5 | 104 | 104.4
-9 |  |  |  | 1.4 | 2.4 | 100.775 | 104 | 104.4
-8.9 |  |  |  | 1.5 | 2.5 | 101.1667 | 104.1 | 104.4
-8.8 |  |  |  | 1.5 | 2.5667 | 101.4 | 104.1 | 104.5
-8.7 |  |  |  | 1.6 | 2.6 | 101.7333 | 104.2 | 104.5
-8.6 |  |  |  | 1.6 | 2.7 | 101.9667 | 104.2 | 104.5
-8.5 |  |  |  | 1.6 | 2.8 | 102.2 | 104.2 | 104.4
-8.4 |  |  |  | 1.7 | 2.9 | 102.325 | 104.2 | 104.4
-8.3 |  |  |  | 1.7 | 2.925 | 102.5 | 104.2 | 104.4
-8.2 |  |  |  | 1.8 | 3.05 | 102.6 | 104.2 | 104.4
-8.1 |  |  |  | 1.8 | 3.2 | 102.66 | 104.2167 | 104.4
-8 |  |  | 1.1 | 1.9 | 3.3 | 102.8 | 104.3 | 104.4
-7.9 |  |  | 1.1 | 1.9 | 3.4 | 102.875 | 104.3 | 104.4
-7.8 |  |  | 1.1 | 2 | 3.5 | 102.925 | 104.3 | 104.4
-7.7 |  |  | 1.2 | 2.0333 | 3.6 | 103 | 104.3 | 104.4
-7.6 |  |  | 1.2 | 2.1 | 3.7 | 103.1 | 104.3 | 104.3167
-7.5 |  |  | 1.2 | 2.2 | 3.9 | 103.2 | 104.3 | 104.3
-7.4 |  |  | 1.2 | 2.3 | 4 | 103.2 | 104.3 | 104.3
-7.3 |  |  | 1.3 | 2.3 | 4.175 | 103.3 | 104.3 | 104.3
-7.2 |  |  | 1.3 | 2.4 | 4.3 | 103.3 | 104.3 | 104.3
-7.1 |  |  | 1.4 | 2.5 | 4.5333 | 103.4 | 104.3 | 104.3
-7 |  | 0.7 | 1.4 | 2.6 | 4.7 | 103.4 | 104.3 | 104.3
-6.9 |  | 0.7 | 1.4 | 2.64 | 4.925 | 103.5 | 104.3 | 104.3
-6.8 |  | 0.8 | 1.5 | 2.8 | 5.175 | 103.5 | 104.3 | 104.3
-6.7 |  | 0.8 | 1.5 | 2.8667 | 5.4 | 103.5 | 104.3 | 104.3
-6.6 |  | 0.8 | 1.6 | 3 | 5.75 | 103.6 | 104.3 | 104.3
-6.5 | 0.6 | 0.8 | 1.6 | 3.1 | 6.075 | 103.6 | 104.3 | 104.3
-6.4 | 0.6 | 0.8 | 1.7 | 3.2 | 6.5333 | 103.625 | 104.3 | 104.2833
-6.3 | 0.6 | 0.9 | 1.7 | 3.3667 | 7 | 103.7 | 104.3 | 104.2
-6.2 | 0.6 | 0.9 | 1.8 | 3.5 | 7.575 | 103.725 | 104.3 | 104.2
-6.1 | 0.6 | 0.9 | 1.9 | 3.65 | 8.3 | 103.8 | 104.3 | 104.2
-6 | 0.6 | 0.9 | 1.9 | 3.85 | 9.2 | 103.8 | 104.3 | 104.2
-5.9 | 0.6 | 1 | 2 | 4.04 | 10.4 | 103.8 | 104.3 | 104.2
-5.8 | 0.7 | 1 | 2.1 | 4.225 | 12 | 103.9 | 104.3 | 104.2
-5.7 | 0.7 | 1 | 2.2 | 4.5 | 14 | 103.9 | 104.3 | 104.2
-5.6 | 0.7 | 1 | 2.3 | 4.7667 | 17.1 | 103.9 | 104.3 | 104.1
-5.5 | 0.7 | 1.1 | 2.3 | 5.1 | 21.4667 | 103.9 | 104.3 | 104.1
-5.4 | 0.7 | 1.1 | 2.4 | 5.5 | 28.7 | 103.9 | 104.3 | 104.1
-5.3 | 0.7 | 1.1 | 2.55 | 5.9333 | 37.75 | 103.9 | 104.2 | 104.0667
-5.2 | 0.8 | 1.2 | 2.7 | 6.4 | 47.8667 | 103.9 | 104.2 | 104
-5.1 | 0.8 | 1.2 | 2.8 | 7.1333 | 59.025 | 103.9 | 104.2 | 103.92
-5 | 0.8 | 1.2 | 2.9 | 8 | 69.3 | 103.9 | 104.2 | 103.9
-4.9 | 0.8 | 1.3 | 3.1 | 9.05 | 78.9 | 103.9 | 104.1 | 103.9
-4.8 | 0.9 | 1.4 | 3.3 | 10.45 | 86 | 103.9 | 104.1 | 103.9
-4.7 | 0.9 | 1.4 | 3.5 | 12.35 | 90.9667 | 103.8 | 104.05 | 103.8
-4.6 | 0.9 | 1.5 | 3.75 | 14.8667 | 93.8333 | 103.8 | 104 | 103.72
-4.5 | 0.9 | 1.5 | 4 | 18.775 | 95.7 | 103.7 | 103.9 | 103.7
-4.4 | 1 | 1.6 | 4.3 | 24.9 | 97.2 | 103.6 | 103.9 | 103.6
-4.3 | 1 | 1.6667 | 4.7 | 32.75 | 98.4 | 103.6 | 103.8 | 103.6
-4.2 | 1 | 1.7 | 5.2 | 42.4333 | 99.2 | 103.5 | 103.78 | 103.5
-4.1 | 1.1 | 1.8 | 5.7 | 53.4667 | 99.85 | 103.5 | 103.65 | 103.4333
-4 | 1.1 | 1.9 | 6.5 | 64.7 | 100.275 | 103.4667 | 103.6 | 103.34
-3.9 | 1.2 | 2 | 7.4 | 74.0333 | 100.7 | 103.4 | 103.5333 | 103.3
-3.8 | 1.2 | 2.2 | 8.7 | 82.5333 | 101 | 103.3667 | 103.5 | 103.2143
-3.7 | 1.3 | 2.3 | 10.2 | 88.65 | 101.325 | 103.3 | 103.4 | 103.2
-3.6 | 1.3 | 2.5 | 12.4 | 92.275 | 101.5333 | 103.3 | 103.3167 | 103.1
-3.5 | 1.5 | 2.7 | 15.7 | 94.8667 | 101.7 | 103.2 | 103.3 | 103.1
-3.4 | 1.6 | 3 | 20.5 | 96.55 | 101.8 | 103.2 | 103.2667 | 103
-3.3 | 1.7 | 3.3 | 28.5333 | 97.825 | 101.9 | 103.1 | 103.2 | 102.95
-3.2 | 1.8 | 3.7 | 38.4 | 98.725 | 102 | 103.08 | 103.1 | 102.8
-3.1 | 1.9 | 4.2333 | 48.7 | 99.4333 | 102.0333 | 103 | 103.1 | 102.8
-3 | 2.1 | 4.9 | 59.7 | 99.95 | 102.1 | 102.9 | 102.9667 | 102.7
-2.9 | 2.35 | 5.7 | 70.5333 | 100.3667 | 102.1 | 102.8333 | 102.9 | 102.6333
-2.8 | 2.65 | 6.8 | 79.8 | 100.6 | 102.1 | 102.775 | 102.8 | 102.4667
-2.7 | 3 | 8.3 | 86.6 | 100.8 | 102.0333 | 102.625 | 102.7 | 102.3833
-2.6 | 3.4 | 10.25 | 91 | 100.9 | 102 | 102.5 | 102.5833 | 102.2167
-2.5 | 4 | 13.2 | 94 | 101 | 101.9 | 102.4 | 102.3857 | 102.1
-2.4 | 4.8 | 17.6 | 95.6 | 101.1 | 101.9 | 102.3 | 102.2333 | 102
-2.3 | 5.8 | 24.4 | 97.1 | 101.1 | 101.8 | 102.15 | 102.1 | 101.8714
-2.2 | 7.15 | 33.6 | 98.1 | 101.1 | 101.7 | 102 | 102 | 101.7
-2.1 | 9.15 | 44.05 | 98.7 | 101.1 | 101.6 | 101.9 | 101.85 | 101.55
-2 | 11.8 | 56.3 | 99.2 | 101 | 101.6 | 101.76 | 101.7 | 101.45
-1.9 | 15.7 | 66.75 | 99.65 | 100.9 | 101.4 | 101.6 | 101.5833 | 101.2833
-1.8 | 22.1 | 76.25 | 99.8 | 100.8667 | 101.3 | 101.425 | 101.5 | 101.1167
-1.7 | 31.5 | 84.6 | 100 | 100.8 | 101.3 | 101.2333 | 101.3 | 101.05
-1.6 | 42.2 | 89.3 | 100.2 | 100.7333 | 101.1667 | 101.2 | 101.12 | 101
-1.5 | 53.4 | 92.7 | 100.3 | 100.7 | 101 | 101.1 | 101.08 | 100.88
-1.4 | 65.3 | 94.9 | 100.4 | 100.7 | 100.9 | 101 | 101 | 100.8
-1.3 | 75.4 | 96.4 | 100.4 | 100.7 | 100.8 | 100.9333 | 100.9 | 100.6333
-1.2 | 83.3 | 97.5 | 100.4 | 100.6 | 100.8 | 100.825 | 100.8 | 100.5714
-1.1 | 89.3 | 98.2 | 100.4 | 100.5 | 100.7 | 100.66 | 100.7 | 100.5
-1 | 93.1 | 98.7 | 100.4 | 100.5 | 100.65 | 100.6 | 100.6 | 100.44
-0.9 | 95.1 | 99.1 | 100.4 | 100.4 | 100.575 | 100.5 | 100.48 | 100.34
-0.8 | 96.8 | 99.5 | 100.3 | 100.4 | 100.425 | 100.4 | 100.4 | 100.3
-0.7 | 97.9 | 99.6 | 100.3 | 100.3 | 100.35 | 100.4 | 100.3 | 100.2571
-0.6 | 98.6 | 99.8 | 100.2 | 100.3 | 100.3 | 100.3 | 100.2833 | 100.2
-0.5 | 99.2 | 99.9 | 100.2 | 100.2 | 100.2 | 100.2 | 100.2 | 100.1
-0.4 | 99.65 | 99.9 | 100.15 | 100.1 | 100.1333 | 100.1 | 100.1 | 100.1
-0.3 | 99.9 | 100 | 100.1 | 100.05 | 100.1 | 100.1 | 100.06 | 100.05
-0.2 | 99.95 | 100 | 100 | 100 | 100.05 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100.04 | 100 | 100 | 100 | 100
0.2 | 99.8333 | 100 | 100 | 100.1 | 100 | 100 | 100 | 100
0.3 | 99.65 | 99.9 | 100.1 | 100.1 | 100.05 | 100 | 100 | 100
0.4 | 99.4 | 99.9 | 100.1 | 100.2 | 100.1 | 100 | 100.02 | 100.1
0.5 | 99.1 | 99.8 | 100.2 | 100.2 | 100.125 | 100.1 | 100.1 | 100.1
0.6 | 98.55 | 99.6667 | 100.2 | 100.2 | 100.2 | 100.2 | 100.2 | 100.2
0.7 | 97.8 | 99.5 | 100.2 | 100.3 | 100.3 | 100.2 | 100.22 | 100.2857
0.8 | 96.5667 | 99.3 | 100.2 | 100.3 | 100.4 | 100.3 | 100.3 | 100.3
0.9 | 95 | 99 | 100.3 | 100.4667 | 100.5 | 100.4 | 100.36 | 100.3
1 | 92.5 | 98.6667 | 100.3 | 100.5 | 100.55 | 100.5 | 100.48 | 100.4
1.1 | 89 | 98.1 | 100.3 | 100.6 | 100.6 | 100.575 | 100.6 | 100.5
1.2 | 83.5667 | 97.3 | 100.3 | 100.62 | 100.7 | 100.6 | 100.7 | 100.6
1.3 | 75.1333 | 96.45 | 100.3 | 100.7 | 100.7 | 100.7 | 100.72 | 100.6667
1.4 | 64.9 | 94.9 | 100.25 | 100.7 | 100.7333 | 100.8 | 100.9 | 100.7333
1.5 | 52.9 | 92.8 | 100.2 | 100.75 | 100.8 | 100.9 | 100.9667 | 100.9
1.6 | 42 | 89.5 | 100.1 | 100.8 | 100.8 | 101 | 101 | 101
1.7 | 31.4 | 84 | 99.9 | 100.8 | 100.92 | 101.1 | 101.1167 | 101.1
1.8 | 22.2 | 76.6 | 99.7 | 100.8 | 101 | 101.175 | 101.2 | 101.15
1.9 | 16 | 67.1 | 99.5 | 100.8 | 101.15 | 101.3 | 101.3 | 101.2167
2 | 11.9 | 55.7 | 99.2 | 100.9 | 101.2 | 101.4 | 101.4429 | 101.3833
2.1 | 9.15 | 44.35 | 98.7 | 100.9 | 101.3 | 101.5 | 101.52 | 101.58
2.2 | 7.1667 | 33.1 | 98 | 100.9 | 101.4 | 101.7 | 101.7 | 101.7429
2.3 | 5.7 | 24.35 | 97 | 101 | 101.4 | 101.8 | 101.8 | 101.8
2.4 | 4.75 | 17.4 | 95.7 | 100.9 | 101.5 | 101.9 | 101.9571 | 102
2.5 | 4 | 13.3 | 93.8 | 100.9 | 101.5 | 102 | 102.1 | 102.1
2.6 | 3.4 | 10.3 | 91 | 100.825 | 101.65 | 102.1 | 102.26 | 102.24
2.7 | 3 | 8.2 | 86.75 | 100.6333 | 101.7 | 102.2333 | 102.4 | 102.32
2.8 | 2.65 | 6.8 | 79.8 | 100.4 | 101.7 | 102.4 | 102.5 | 102.5
2.9 | 2.3333 | 5.6 | 70.8 | 100.2 | 101.8 | 102.46 | 102.5667 | 102.5667
3 | 2.1 | 4.8 | 59.8 | 99.8 | 101.8 | 102.5333 | 102.675 | 102.7
3.1 | 1.9 | 4.1667 | 48.9 | 99.3 | 101.7 | 102.6 | 102.7 | 102.8
3.2 | 1.8 | 3.7 | 38.4 | 98.675 | 101.7 | 102.7 | 102.8 | 102.8333
3.3 | 1.6 | 3.3 | 27.9 | 97.7 | 101.55 | 102.8 | 102.9 | 102.9
3.4 | 1.6 | 3 | 20.7 | 96.36 | 101.4667 | 102.8 | 102.9 | 103
3.5 | 1.4667 | 2.75 | 15.7 | 94.8333 | 101.36 | 102.9 | 103 | 103.1
3.6 | 1.4 | 2.55 | 12.5 | 92.3333 | 101.2 | 102.9 | 103.1 | 103.1167
3.7 | 1.3 | 2.3 | 10.2 | 88.55 | 101 | 102.9 | 103.1 | 103.2
3.8 | 1.2 | 2.1667 | 8.5 | 82.7667 | 100.7667 | 103 | 103.2 | 103.25
3.9 | 1.2 | 2 | 7.4 | 73.9667 | 100.45 | 103 | 103.26 | 103.3
4 | 1.1 | 1.9 | 6.4667 | 64.8 | 100.0667 | 103.1 | 103.3 | 103.3
4.1 | 1.1 | 1.8 | 5.7 | 53.6 | 99.5 | 103.1 | 103.35 | 103.4
4.2 | 1 | 1.7 | 5.1 | 42.8 | 98.85 | 103.2 | 103.4 | 103.5
4.3 | 1 | 1.65 | 4.6 | 32.6 | 98.15 | 103.2 | 103.5 | 103.54
4.4 | 1 | 1.6 | 4.3 | 24.3 | 96.95 | 103.3 | 103.5 | 103.6
4.5 | 0.9 | 1.5 | 3.95 | 18.5667 | 95.5667 | 103.3 | 103.54 | 103.7
4.6 | 0.9 | 1.4 | 3.7 | 14.6667 | 93.4 | 103.3 | 103.6 | 103.75
4.7 | 0.9 | 1.4 | 3.5 | 12.2 | 90.54 | 103.375 | 103.6833 | 103.8
4.8 | 0.86667 | 1.3 | 3.2 | 10.375 | 85.8 | 103.4 | 103.7 | 103.9
4.9 | 0.8 | 1.3 | 3.1 | 8.9 | 78.7 | 103.4 | 103.7 | 103.9
5 | 0.8 | 1.2 | 2.9 | 7.875 | 69.9 | 103.5 | 103.8 | 103.9
5.1 | 0.8 | 1.2 | 2.8 | 7.0333 | 59.4 | 103.5 | 103.8 | 103.95
5.2 | 0.8 | 1.2 | 2.7 | 6.4333 | 48.3667 | 103.6 | 103.8667 | 104
5.3 | 0.7 | 1.1 | 2.55 | 5.85 | 38 | 103.6 | 103.9 | 104
5.4 | 0.7 | 1.1 | 2.4 | 5.4 | 28.5667 | 103.5 | 103.9 | 104.1
5.5 | 0.7 | 1.1 | 2.3 | 5 | 21.4 | 103.5 | 103.9 | 104.1
5.6 | 0.7 | 1 | 2.2 | 4.7333 | 17.0667 | 103.5 | 103.9833 | 104.1333
5.7 | 0.7 | 1 | 2.2 | 4.5 | 14.075 | 103.5 | 104 | 104.2
5.8 | 0.7 | 1 | 2.1 | 4.1667 | 12 | 103.5 | 104 | 104.2
5.9 | 0.6 | 0.9 | 2 | 3.975 | 10.3 | 103.5 | 104 | 104.2
6 | 0.6 | 0.9 | 1.9 | 3.825 | 9.2 | 103.4 | 104 | 104.2
6.1 | 0.6 | 0.9 | 1.9 | 3.625 | 8.275 | 103.4 | 104 | 104.2
6.2 | 0.6 | 0.9 | 1.8 | 3.45 | 7.5 | 103.35 | 104 | 104.2
6.3 | 0.6 | 0.85 | 1.7 | 3.3 | 6.9333 | 103.3 | 104 | 104.2
6.4 | 0.6 | 0.8 | 1.7 | 3.2 | 6.4333 | 103.3 | 104 | 104.3
6.5 | 0.6 | 0.8 | 1.6 | 3.0667 | 6 | 103.3 | 104 | 104.3
6.6 |  | 0.8 | 1.6 | 2.9333 | 5.6 | 103.2 | 104 | 104.3
6.7 |  | 0.8 | 1.5 | 2.85 | 5.4 | 103.2 | 104 | 104.3
6.8 |  | 0.8 | 1.5 | 2.7333 | 5.1 | 103.15 | 104 | 104.3
6.9 |  | 0.8 | 1.4 | 2.7 | 4.8 | 103.1 | 104 | 104.3
7 |  | 0.7 | 1.4 | 2.5667 | 4.58 | 103.0667 | 104 | 104.3
7.1 |  |  | 1.4 | 2.5 | 4.4 | 103 | 103.9167 | 104.3
7.2 |  |  | 1.3 | 2.4 | 4.2 | 103 | 103.9 | 104.3
7.3 |  |  | 1.3 | 2.3 | 4.0667 | 102.925 | 103.9 | 104.3
7.4 |  |  | 1.2 | 2.2 | 3.9 | 102.9 | 103.9 | 104.3
7.5 |  |  | 1.2 | 2.2 | 3.78 | 102.9 | 103.9 | 104.3
7.6 |  |  | 1.2 | 2.1 | 3.65 | 102.8 | 103.9 | 104.4
7.7 |  |  | 1.1 | 2 | 3.52 | 102.7 | 103.9 | 104.4
7.8 |  |  | 1.1 | 1.975 | 3.4 | 102.6 | 103.9 | 104.4
7.9 |  |  | 1.1 | 1.9 | 3.3 | 102.575 | 103.9 | 104.4
8 |  |  | 1.1 | 1.85 | 3.2 | 102.5 | 104 | 104.4
8.1 |  |  |  | 1.8 | 3.1 | 102.4 | 104 | 104.4
8.2 |  |  |  | 1.75 | 3 | 102.3 | 104 | 104.4
8.3 |  |  |  | 1.7 | 2.88 | 102.2 | 104 | 104.4
8.4 |  |  |  | 1.7 | 2.8 | 102 | 104 | 104.4
8.5 |  |  |  | 1.6 | 2.725 | 101.95 | 104 | 104.4333
8.6 |  |  |  | 1.6 | 2.66 | 101.7 | 104 | 104.5
8.7 |  |  |  | 1.6 | 2.525 | 101.5 | 103.9 | 104.5
8.8 |  |  |  | 1.5 | 2.5 | 101.3333 | 103.9 | 104.48
8.9 |  |  |  | 1.5 | 2.4 | 100.9 | 103.8857 | 104.4
9 |  |  |  | 1.4 | 2.35 | 100.625 | 103.8 | 104.4
9.1 |  |  |  |  | 2.3 | 100.2 | 103.8 | 104.4
9.2 |  |  |  |  | 2.2 | 99.75 | 103.7667 | 104.4
9.3 |  |  |  |  | 2.2 | 99.14 | 103.7 | 104.35
9.4 |  |  |  |  | 2.1 | 98.3 | 103.65 | 104.3
9.5 |  |  |  |  | 2.025 | 97.2667 | 103.6 | 104.3
9.6 |  |  |  |  | 2 | 96.16 | 103.6 | 104.3
9.7 |  |  |  |  | 2 | 94.4333 | 103.55 | 104.3
9.8 |  |  |  |  | 1.9 | 92.3 | 103.5 | 104.3333
9.9 |  |  |  |  | 1.9 | 88.9 | 103.6 | 104.4
10 |  |  |  |  | 1.8 | 83.9 | 103.6 | 104.4
10.1 |  |  |  |  | 1.8 | 76.1667 | 103.6 | 104.4
10.2 | [5 empty] | 66.45 | 103.7 | 104.5
10.3 | [5 empty] | 56.9 | 103.7 | 104.5167
10.4 | [5 empty] | 46.45 | 103.8 | 104.6833
10.5 | [5 empty] | 36.3 | 103.9 | 104.8
10.6 | [5 empty] | 28 | 103.9 | 104.8
10.7 | [5 empty] | 22.425 | 103.9 | 104.9
10.8 | [5 empty] | 18.6 | 104 | 105
10.9 | [5 empty] | 15.9 | 104 | 105
11 | [5 empty] | 13.975 | 104.1 | 105.1
11.1 | [5 empty] | 12.7 | 104.1 | 105.1667
11.2 | [5 empty] | 11.6 | 104.1 | 105.2
11.3 | [5 empty] | 10.75 | 104.0286 | 105.2
11.4 | [5 empty] | 10.1 | 104 | 105.2
11.5 | [5 empty] | 9.48 | 104 | 105.2
11.6 | [5 empty] | 9 | 104 | 105.2
11.7 | [5 empty] | 8.55 | 103.9 | 105.2
11.8 | [5 empty] | 8.225 | 103.9 | 105.3
11.9 | [5 empty] | 7.875 | 103.9 | 105.3
12 | [5 empty] | 7.575 | 103.9 | 105.3
12.1 | [5 empty] | 7.325 | 103.9 | 105.3
12.2 | [5 empty] | 7.0667 | 103.8 | 105.3
12.3 | [5 empty] | 6.9 | 103.8 | 105.3
12.4 | [5 empty] | 6.7 | 103.7667 | 105.3
12.5 | [5 empty] | 6.4 | 103.7 | 105.3
12.6 | [5 empty] | 6.2667 | 103.7 | 105.3
12.7 | [5 empty] | 6.0333 | 103.6 | 105.3
12.8 | [5 empty] | 5.86 | 103.55 | 105.3
12.9 | [5 empty] | 5.7 | 103.48 | 105.3
13 | [5 empty] | 5.575 | 103.375 | 105.3
13.1 | [5 empty] | 5.4 | 103.3 | 105.3
13.2 | [5 empty] | 5.275 | 103.2 | 105.3
13.3 | [5 empty] | 5.125 | 103.16 | 105.2167
13.4 | [5 empty] | 5 | 103.04 | 105.2
13.5 | [5 empty] | 4.88 | 102.92 | 105.2
13.6 | [5 empty] | 4.7 | 102.8 | 105.2
13.7 | [5 empty] | 4.6 | 102.7 | 105.2
13.8 | [5 empty] | 4.5 | 102.54 | 105.2
13.9 | [5 empty] | 4.4 | 102.3333 | 105.2
14 | [5 empty] | 4.3 | 102.06 | 105.2
14.1 | [5 empty] | 4.15 | 101.9 | 105.3
14.2 | [5 empty] | 4.1 | 101.65 | 105.3
14.3 | [5 empty] | 4 | 101.2333 | 105.3
14.4 | [5 empty] | 3.9 | 100.85 | 105.3
14.5 | [5 empty] | 3.8 | 100.2667 | 105.4
14.6 | [5 empty] | 3.7 | 99.48 | 105.4
14.7 | [5 empty] | 3.6 | 98.5 | 105.4
14.8 | [5 empty] | 3.5 | 97.2333 | 105.4
14.9 | [5 empty] | 3.4 | 95.3 | 105.5
15 | [5 empty] | 3.4 | 92.7667 | 105.5
15.1 | [5 empty] | 3.3 | 88.6 | 105.5
15.2 | [5 empty] | 3.2 | 82.06 | 105.5
15.3 | [5 empty] | 3.1 | 73.36 | 105.5
15.4 | [6 empty] | 63.275 | 105.5
15.5 | [6 empty] | 53.5 | 105.5
15.6 | [6 empty] | 43.3 | 105.5
15.7 | [6 empty] | 34.1833 | 105.5
15.8 | [6 empty] | 26.6833 | 105.5
15.9 | [6 empty] | 21.8 | 105.5
16 | [6 empty] | 18.45 | 105.4333
16.1 | [6 empty] | 16.4 | 105.4
16.2 | [6 empty] | 14.6667 | 105.4
16.3 | [6 empty] | 13.5 | 105.3
16.4 | [6 empty] | 12.5667 | 105.2
16.5 | [6 empty] | 11.7 | 105.2
16.6 | [6 empty] | 11.1667 | 105.16
16.7 | [6 empty] | 10.64 | 105.1
16.8 | [6 empty] | 10.16 | 105
16.9 | [6 empty] | 9.7667 | 104.975
17 | [6 empty] | 9.4333 | 104.9
17.1 | [6 empty] | 9.2 | 104.8333
17.2 | [6 empty] | 8.85 | 104.8
17.3 | [6 empty] | 8.58 | 104.8
17.4 | [6 empty] | 8.4 | 104.7
17.5 | [6 empty] | 8.2 | 104.7
17.6 | [6 empty] | 7.9 | 104.7
17.7 | [6 empty] | 7.7 | 104.7
17.8 | [6 empty] | 7.58 | 104.7
17.9 | [6 empty] | 7.3714 | 104.7286
18 | [6 empty] | 7.2167 | 104.8
18.1 | [6 empty] | 7 | 104.8
18.2 | [6 empty] | 6.8167 | 104.8
18.3 | [6 empty] | 6.7 | 104.9
18.4 | [6 empty] | 6.5 | 104.88
18.5 | [6 empty] | 6.4167 | 104.8
18.6 | [6 empty] | 6.2 | 104.8
18.7 | [6 empty] | 6.1 | 104.7
18.8 | [6 empty] | 5.9667 | 104.6
18.9 | [6 empty] | 5.8 | 104.425
19 | [6 empty] | 5.7 | 104.22
19.1 | [6 empty] | 5.56 | 103.98
19.2 | [6 empty] | 5.4 | 103.725
19.3 | [6 empty] | 5.3 | 103.38
19.4 | [6 empty] | 5.2167 | 103
19.5 | [6 empty] | 5.05 | 102.56
19.6 | [6 empty] | 4.94 | 102.1
19.7 | [6 empty] | 4.85 | 101.4333
19.8 | [6 empty] | 4.725 | 100.64
19.9 | [6 empty] | 4.62 | 99.48
20 | [6 empty] | 4.55 | 98.1333
20.1 | [6 empty] | 4.5 | 96.1333
20.2 | [6 empty] | 4.325 | 92.95
20.3 | [6 empty] | 4.2 | 87.7833
20.4 | [6 empty] | 4.2 | 80.1167
20.5 | [7 empty] | 71.5
20.6 | [7 empty] | 61.2
20.7 | [7 empty] | 50.18
20.8 | [7 empty] | 40.2
20.9 | [7 empty] | 31.6
21 | [7 empty] | 24.9
21.1 | [7 empty] | 20.68
21.2 | [7 empty] | 17.86
21.3 | [7 empty] | 15.9333
21.4 | [7 empty] | 14.46
21.5 | [7 empty] | 13.4
21.6 | [7 empty] | 12.52
21.7 | [7 empty] | 11.8
21.8 | [7 empty] | 11.15
21.9 | [7 empty] | 10.6667
22 | [7 empty] | 10.3
22.1 | [7 empty] | 9.9
22.2 | [7 empty] | 9.5667
22.3 | [7 empty] | 9.25
22.4 | [7 empty] | 8.98
22.5 | [7 empty] | 8.7
22.6 | [7 empty] | 8.42
22.7 | [7 empty] | 8.2
22.8 | [7 empty] | 7.9333
22.9 | [7 empty] | 7.76
23 | [7 empty] | 7.525
23.1 | [7 empty] | 7.3167
23.2 | [7 empty] | 7.1
23.3 | [7 empty] | 6.8833
23.4 | [7 empty] | 6.7
23.5 | [7 empty] | 6.5167
23.6 | [7 empty] | 6.35
23.7 | [7 empty] | 6.1667
23.8 | [7 empty] | 6
23.9 | [7 empty] | 5.8333
24 | [7 empty] | 5.66
24.1 | [7 empty] | 5.5833
24.2 | [7 empty] | 5.4167
24.3 | [7 empty] | 5.25
24.4 | [7 empty] | 5.1
24.5 | [7 empty] | 5.0167
24.6 | [7 empty] | 4.825
24.7 | [7 empty] | 4.7167
24.8 | [7 empty] | 4.64
24.9 | [7 empty] | 4.52
25 | [7 empty] | 4.4
25.1 | [7 empty] | 4.3
25.2 | [7 empty] | 4.2
25.3 | [7 empty] | 4.1
25.4 | [7 empty] | 4
25.5 | [7 empty] | 3.9

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.9 | [7 empty] | 4
-25.8 | [7 empty] | 4.1
-25.7 | [7 empty] | 4.1714
-25.6 | [7 empty] | 4.3143
-25.5 | [7 empty] | 4.4
-25.4 | [7 empty] | 4.5
-25.3 | [7 empty] | 4.6
-25.2 | [7 empty] | 4.7
-25.1 | [7 empty] | 4.84
-25 | [7 empty] | 4.9
-24.9 | [7 empty] | 5.0571
-24.8 | [7 empty] | 5.2
-24.7 | [7 empty] | 5.3375
-24.6 | [7 empty] | 5.4625
-24.5 | [7 empty] | 5.6
-24.4 | [7 empty] | 5.7375
-24.3 | [7 empty] | 5.8625
-24.2 | [7 empty] | 5.9875
-24.1 | [7 empty] | 6.2125
-24 | [7 empty] | 6.3667
-23.9 | [7 empty] | 6.55
-23.8 | [7 empty] | 6.675
-23.7 | [7 empty] | 6.9
-23.6 | [7 empty] | 7.05
-23.5 | [7 empty] | 7.2571
-23.4 | [7 empty] | 7.5
-23.3 | [7 empty] | 7.6667
-23.2 | [7 empty] | 7.925
-23.1 | [7 empty] | 8.15
-23 | [7 empty] | 8.425
-22.9 | [7 empty] | 8.675
-22.8 | [7 empty] | 8.9111
-22.7 | [7 empty] | 9.225
-22.6 | [7 empty] | 9.5571
-22.5 | [7 empty] | 9.95
-22.4 | [7 empty] | 10.3286
-22.3 | [7 empty] | 10.7571
-22.2 | [7 empty] | 11.3714
-22.1 | [7 empty] | 12
-22 | [7 empty] | 12.725
-21.9 | [7 empty] | 13.8
-21.8 | [7 empty] | 14.9875
-21.7 | [7 empty] | 16.5375
-21.6 | [7 empty] | 18.5429
-21.5 | [7 empty] | 21.4667
-21.4 | [7 empty] | 25.7
-21.3 | [7 empty] | 32.48
-21.2 | [7 empty] | 40.5333
-21.1 | [7 empty] | 50.4286
-21 | [7 empty] | 60.5
-20.9 | [7 empty] | 70.0667
-20.8 | [7 empty] | 78.675
-20.7 | [6 empty] | 4.0222 | 85.9
-20.6 | [6 empty] | 4.1429 | 90.9571
-20.5 | [6 empty] | 4.2 | 94.1286
-20.4 | [6 empty] | 4.3 | 96.3
-20.3 | [6 empty] | 4.4 | 97.7714
-20.2 | [6 empty] | 4.48 | 98.9875
-20.1 | [6 empty] | 4.5667 | 99.9
-20 | [6 empty] | 4.7 | 100.48
-19.9 | [6 empty] | 4.8 | 101.1429
-19.8 | [6 empty] | 4.8571 | 101.65
-19.7 | [6 empty] | 5 | 102.2
-19.6 | [6 empty] | 5.1167 | 102.5714
-19.5 | [6 empty] | 5.2 | 102.88
-19.4 | [6 empty] | 5.35 | 103.2286
-19.3 | [6 empty] | 5.4286 | 103.3833
-19.2 | [6 empty] | 5.5833 | 103.65
-19.1 | [6 empty] | 5.7429 | 103.8
-19 | [6 empty] | 5.8 | 103.9167
-18.9 | [6 empty] | 5.9667 | 104.0714
-18.8 | [6 empty] | 6.1333 | 104.1
-18.7 | [6 empty] | 6.3 | 104.2
-18.6 | [6 empty] | 6.4 | 104.2
-18.5 | [6 empty] | 6.5571 | 104.2
-18.4 | [6 empty] | 6.7833 | 104.2
-18.3 | [6 empty] | 6.9 | 104.2
-18.2 | [6 empty] | 7.04 | 104.2
-18.1 | [6 empty] | 7.3167 | 104.2
-18 | [6 empty] | 7.5 | 104.2
-17.9 | [6 empty] | 7.6833 | 104.2
-17.8 | [6 empty] | 7.8667 | 104.25
-17.7 | [6 empty] | 8.1667 | 104.3
-17.6 | [6 empty] | 8.35 | 104.3
-17.5 | [6 empty] | 8.6667 | 104.4
-17.4 | [6 empty] | 8.9 | 104.4286
-17.3 | [6 empty] | 9.3 | 104.5
-17.2 | [6 empty] | 9.6 | 104.5167
-17.1 | [6 empty] | 10.0714 | 104.6
-17 | [6 empty] | 10.55 | 104.7
-16.9 | [6 empty] | 11.1333 | 104.8
-16.8 | [6 empty] | 11.8 | 104.9
-16.7 | [6 empty] | 12.6833 | 104.9429
-16.6 | [6 empty] | 13.6833 | 105
-16.5 | [6 empty] | 14.9286 | 105
-16.4 | [6 empty] | 16.6 | 105.0429
-16.3 | [6 empty] | 18.9 | 105.1
-16.2 | [6 empty] | 22.1667 | 105.1
-16.1 | [6 empty] | 27.2667 | 105.2
-16 | [6 empty] | 34.6 | 105.2571
-15.9 | [6 empty] | 43.3333 | 105.3
-15.8 | [6 empty] | 52.66 | 105.3
-15.7 | [6 empty] | 63.3 | 105.3
-15.6 | [6 empty] | 72.22 | 105.3
-15.5 | [5 empty] | 2.9333 | 80.8 | 105.3
-15.4 | [5 empty] | 3 | 87.05 | 105.2625
-15.3 | [5 empty] | 3.1 | 91.2 | 105.2
-15.2 | [5 empty] | 3.1 | 94.1 | 105.2
-15.1 | [5 empty] | 3.2 | 95.9571 | 105.2
-15 | [5 empty] | 3.3 | 97.3429 | 105.1
-14.9 | [5 empty] | 3.325 | 98.5 | 105.1
-14.8 | [5 empty] | 3.46 | 99.3 | 105.1
-14.7 | [5 empty] | 3.5 | 99.9667 | 105.1
-14.6 | [5 empty] | 3.6 | 100.6 | 105.1
-14.5 | [5 empty] | 3.7 | 101.0571 | 105.1
-14.4 | [5 empty] | 3.8 | 101.46 | 105.1
-14.3 | [5 empty] | 3.9 | 101.775 | 105.1
-14.2 | [5 empty] | 4 | 102 | 105.1
-14.1 | [5 empty] | 4.1 | 102.2167 | 105.1
-14 | [5 empty] | 4.18 | 102.4667 | 105.1
-13.9 | [5 empty] | 4.28 | 102.65 | 105.1
-13.8 | [5 empty] | 4.4 | 102.8 | 105.1
-13.7 | [5 empty] | 4.5 | 102.96 | 105.1
-13.6 | [5 empty] | 4.6333 | 103.1 | 105.1
-13.5 | [5 empty] | 4.78 | 103.2 | 105.1
-13.4 | [5 empty] | 4.9 | 103.3 | 105.1333
-13.3 | [5 empty] | 5.1 | 103.4 | 105.2
-13.2 | [5 empty] | 5.175 | 103.5286 | 105.2
-13.1 | [5 empty] | 5.325 | 103.6 | 105.2
-13 | [5 empty] | 5.5 | 103.75 | 105.3
-12.9 | [5 empty] | 5.725 | 103.8167 | 105.3
-12.8 | [5 empty] | 5.875 | 103.9 | 105.3
-12.7 | [5 empty] | 6.1 | 103.9333 | 105.3
-12.6 | [5 empty] | 6.26 | 104 | 105.3
-12.5 | [5 empty] | 6.52 | 104.0571 | 105.3
-12.4 | [5 empty] | 6.75 | 104.1 | 105.3
-12.3 | [5 empty] | 7 | 104.1 | 105.275
-12.2 | [5 empty] | 7.35 | 104.1571 | 105.2
-12.1 | [5 empty] | 7.65 | 104.2 | 105.275
-12 | [5 empty] | 7.98 | 104.2 | 105.2
-11.9 | [5 empty] | 8.4 | 104.2 | 105.2
-11.8 | [5 empty] | 8.85 | 104.3 | 105.2
-11.7 | [5 empty] | 9.375 | 104.3 | 105.2
-11.6 | [5 empty] | 10.025 | 104.3 | 105.2
-11.5 | [5 empty] | 10.78 | 104.3 | 105.1
-11.4 | [5 empty] | 11.7667 | 104.3 | 105.1
-11.3 | [5 empty] | 12.8 | 104.4 | 105.1
-11.2 | [5 empty] | 14.35 | 104.4 | 105.0833
-11.1 | [5 empty] | 16.32 | 104.4 | 105
-11 | [5 empty] | 19.0333 | 104.4 | 105
-10.9 | [5 empty] | 22.9 | 104.4 | 104.9
-10.8 | [5 empty] | 28.85 | 104.4 | 104.8375
-10.7 | [5 empty] | 36.7 | 104.3 | 104.8
-10.6 | [5 empty] | 45.575 | 104.26 | 104.7
-10.5 | [5 empty] | 56.5 | 104.2 | 104.6429
-10.4 | [5 empty] | 65.9 | 104.1 | 104.6
-10.3 |  |  |  |  | 1.6 | 74.85 | 104.1 | 104.5
-10.2 |  |  |  |  | 1.6 | 82.4 | 104.0167 | 104.5
-10.1 |  |  |  |  | 1.7 | 87.825 | 104 | 104.4
-10 |  |  |  |  | 1.7 | 91.12 | 103.9667 | 104.4
-9.9 |  |  |  |  | 1.7 | 93.58 | 103.9 | 104.4
-9.8 |  |  |  |  | 1.8 | 95.3 | 103.9 | 104.4
-9.7 |  |  |  |  | 1.85 | 96.7 | 103.9 | 104.4
-9.6 |  |  |  |  | 1.9 | 97.7 | 103.95 | 104.4
-9.5 |  |  |  |  | 1.9 | 98.4 | 104 | 104.4
-9.4 |  |  |  |  | 2 | 99.3 | 104 | 104.4
-9.3 |  |  |  |  | 2.02 | 99.875 | 104 | 104.4
-9.2 |  |  |  | 1.3 | 2.1 | 100.35 | 104.1 | 104.4
-9.1 |  |  |  | 1.3 | 2.2 | 100.825 | 104.1 | 104.4
-9 |  |  |  | 1.3 | 2.2 | 101.25 | 104.1667 | 104.4
-8.9 |  |  |  | 1.4 | 2.3 | 101.525 | 104.2 | 104.4
-8.8 |  |  |  | 1.4 | 2.38 | 101.75 | 104.2 | 104.4
-8.7 |  |  |  | 1.4 | 2.4 | 102 | 104.2 | 104.4
-8.6 |  |  |  | 1.5 | 2.5 | 102.2 | 104.3 | 104.4
-8.5 |  |  |  | 1.5 | 2.6 | 102.4 | 104.3 | 104.4
-8.4 |  |  |  | 1.5 | 2.7 | 102.5 | 104.28 | 104.4
-8.3 |  |  |  | 1.6 | 2.8 | 102.68 | 104.3 | 104.4
-8.2 |  |  |  | 1.7 | 2.9 | 102.7 | 104.3 | 104.4
-8.1 |  |  | 1 | 1.7 | 2.98 | 102.8 | 104.3 | 104.4
-8 |  |  | 1 | 1.7 | 3.0667 | 102.9 | 104.3 | 104.4
-7.9 |  |  | 1 | 1.8 | 3.2 | 103 | 104.3 | 104.4
-7.8 |  |  | 1.1 | 1.9 | 3.3 | 103 | 104.3 | 104.3429
-7.7 |  |  | 1.1 | 1.9 | 3.4 | 103.1 | 104.3 | 104.3
-7.6 |  |  | 1.1 | 2 | 3.525 | 103.1 | 104.2429 | 104.3
-7.5 |  |  | 1.1 | 2.05 | 3.7 | 103.2 | 104.3 | 104.3
-7.4 |  |  | 1.2 | 2.1 | 3.9 | 103.2 | 104.2 | 104.3
-7.3 |  |  | 1.2 | 2.2 | 4.025 | 103.3 | 104.3 | 104.3
-7.2 |  |  | 1.2 | 2.3 | 4.24 | 103.3 | 104.3 | 104.3
-7.1 |  |  | 1.3 | 2.4 | 4.4857 | 103.4 | 104.3 | 104.3
-7 |  | 0.7 | 1.3 | 2.5 | 4.72 | 103.46 | 104.2167 | 104.3
-6.9 |  | 0.7 | 1.35 | 2.5 | 5.0333 | 103.5 | 104.3 | 104.3
-6.8 |  | 0.7 | 1.4 | 2.65 | 5.3 | 103.5 | 104.3 | 104.3
-6.7 |  | 0.7 | 1.45 | 2.8 | 5.72 | 103.5 | 104.3 | 104.3
-6.6 |  | 0.8 | 1.5 | 2.9 | 6.1 | 103.525 | 104.3 | 104.3
-6.5 | 0.6 | 0.8 | 1.6 | 3 | 6.6 | 103.6 | 104.3 | 104.3
-6.4 | 0.6 | 0.8 | 1.6 | 3.2 | 7.15 | 103.7 | 104.3 | 104.3
-6.3 | 0.6 | 0.8 | 1.7 | 3.3 | 7.8 | 103.7 | 104.3 | 104.3
-6.2 | 0.6 | 0.8 | 1.7 | 3.5 | 8.65 | 103.7 | 104.3 | 104.2
-6.1 | 0.6 | 0.9 | 1.8 | 3.7 | 9.6 | 103.75 | 104.3 | 104.2
-6 | 0.6 | 0.9 | 1.9 | 3.9 | 10.85 | 103.8 | 104.2833 | 104.2
-5.9 | 0.6 | 0.9 | 1.9333 | 4.2 | 12.6 | 103.8 | 104.2 | 104.2
-5.8 | 0.6 | 0.9 | 2 | 4.5 | 14.725 | 103.8 | 104.2 | 104.2
-5.7 | 0.7 | 1 | 2.1 | 4.7 | 17.875 | 103.8 | 104.2 | 104.1
-5.6 | 0.7 | 1 | 2.2 | 5.1 | 22.625 | 103.8 | 104.2 | 104.1
-5.5 | 0.7 | 1 | 2.3 | 5.55 | 29.575 | 103.8 | 104.2 | 104.1
-5.4 | 0.7 | 1.1 | 2.4 | 6 | 37.7 | 103.8 | 104.2 | 104.1
-5.3 | 0.7 | 1.1 | 2.6 | 6.6333 | 47.75 | 103.8 | 104.2 | 104.0125
-5.2 | 0.8 | 1.1 | 2.7 | 7.4 | 58.3 | 103.8 | 104.1 | 103.9875
-5.1 | 0.8 | 1.2 | 2.9 | 8.2 | 68.7 | 103.8 | 104.1 | 103.9
-5 | 0.8 | 1.2 | 3.05 | 9.3 | 77.5 | 103.8 | 104.1 | 103.9
-4.9 | 0.8 | 1.3 | 3.3 | 10.75 | 84.6 | 103.7 | 104 | 103.8
-4.8 | 0.9 | 1.3 | 3.45 | 12.6 | 89.1667 | 103.7 | 103.9 | 103.7833
-4.7 | 0.9 | 1.4 | 3.8 | 15.15 | 92.36 | 103.7 | 103.9 | 103.7
-4.6 | 0.9 | 1.5 | 4.05 | 18.8 | 94.58 | 103.6 | 103.9 | 103.6857
-4.5 | 1 | 1.55 | 4.4 | 24.3 | 96.3 | 103.6 | 103.8143 | 103.6
-4.4 | 1 | 1.6 | 4.9 | 31.8 | 97.4 | 103.5 | 103.8 | 103.5571
-4.3 | 1 | 1.7 | 5.3 | 40.7 | 98.25 | 103.5 | 103.7 | 103.5
-4.2 | 1.1 | 1.8 | 5.9333 | 50.85 | 99 | 103.5 | 103.6833 | 103.5
-4.1 | 1.1 | 1.9 | 6.7333 | 61.3 | 99.65 | 103.45 | 103.6 | 103.4
-4 | 1.2 | 2.0667 | 7.7 | 71.3 | 100.2 | 103.4 | 103.5 | 103.4
-3.9 | 1.2 | 2.2 | 8.9 | 79.9 | 100.58 | 103.3 | 103.5 | 103.3
-3.8 | 1.3 | 2.4 | 10.5 | 86.15 | 100.875 | 103.2 | 103.5 | 103.2
-3.7 | 1.4 | 2.6 | 12.4 | 90.15 | 101.1 | 103.175 | 103.4 | 103.1167
-3.6 | 1.5 | 2.85 | 15.4 | 92.9 | 101.275 | 103.1 | 103.35 | 103.1
-3.5 | 1.6 | 3.2 | 19.6 | 94.9 | 101.525 | 103.1 | 103.3 | 103.0286
-3.4 | 1.7 | 3.5333 | 25.8 | 96.3 | 101.7 | 103.1 | 103.1857 | 102.9
-3.3 | 1.9 | 3.9 | 34.6 | 97.5 | 101.8 | 103 | 103.1 | 102.9
-3.2 | 2.1 | 4.4 | 44.3 | 98.2667 | 101.9 | 103 | 103 | 102.8111
-3.1 | 2.25 | 5.1 | 54.8 | 98.95 | 102 | 102.82 | 102.98 | 102.8
-3 | 2.5 | 5.95 | 65.2 | 99.5 | 102 | 102.8 | 102.9 | 102.75
-2.9 | 2.8 | 7 | 75.3 | 99.9667 | 102 | 102.725 | 102.8167 | 102.5714
-2.8 | 3.15 | 8.3667 | 82.7 | 100.3 | 102 | 102.6 | 102.75 | 102.4167
-2.7 | 3.6 | 10 | 88.1 | 100.4 | 102 | 102.5 | 102.6833 | 102.34
-2.6 | 4.2 | 12.35 | 91.5 | 100.6 | 102 | 102.4333 | 102.5 | 102.1857
-2.5 | 4.9 | 16.1 | 93.8 | 100.7 | 101.9 | 102.35 | 102.4 | 102.0333
-2.4 | 5.8 | 21.2 | 95.4 | 100.7 | 101.8 | 102.2 | 102.2333 | 101.8833
-2.3 | 7.1 | 28.55 | 96.6 | 100.8 | 101.8 | 102.075 | 102.1 | 101.8
-2.2 | 8.65 | 38.15 | 97.5 | 100.8 | 101.7 | 101.9 | 102.0143 | 101.675
-2.1 | 10.95 | 48.8 | 98.15 | 100.8 | 101.6 | 101.8 | 101.75 | 101.55
-2 | 14 | 60.45 | 98.65 | 100.8 | 101.5 | 101.675 | 101.6833 | 101.425
-1.9 | 18.6 | 69.9 | 99.1 | 100.8 | 101.4 | 101.4667 | 101.5 | 101.3
-1.8 | 26.1 | 78.9 | 99.4 | 100.7 | 101.275 | 101.4 | 101.42 | 101.1333
-1.7 | 35.0667 | 85.7 | 99.6 | 100.7 | 101.2 | 101.3 | 101.3333 | 101
-1.6 | 45.9 | 90 | 99.8 | 100.7 | 101.1 | 101.2 | 101.24 | 100.9625
-1.5 | 57.4 | 92.7 | 99.9 | 100.7 | 101 | 101.05 | 101.1 | 100.8375
-1.4 | 67.9 | 94.6 | 100 | 100.65 | 100.9 | 101 | 101 | 100.8
-1.3 | 77.3 | 96.15 | 100.1 | 100.6 | 100.9 | 100.9 | 100.94 | 100.6333
-1.2 | 84.65 | 97.1667 | 100.2 | 100.5 | 100.8 | 100.8 | 100.85 | 100.6
-1.1 | 89.5 | 97.85 | 100.2 | 100.4 | 100.675 | 100.7 | 100.6667 | 100.5375
-1 | 92.8 | 98.6 | 100.2 | 100.4 | 100.6 | 100.6 | 100.6 | 100.5
-0.9 | 94.9 | 99 | 100.2 | 100.4 | 100.5 | 100.5 | 100.5 | 100.4
-0.8 | 96.6 | 99.2333 | 100.1 | 100.3 | 100.475 | 100.4 | 100.4 | 100.3625
-0.7 | 97.6 | 99.6 | 100.1 | 100.2 | 100.4 | 100.35 | 100.4 | 100.2375
-0.6 | 98.4 | 99.8 | 100.1 | 100.2 | 100.3 | 100.3 | 100.3 | 100.1222
-0.5 | 99.05 | 99.9 | 100.1 | 100.1 | 100.3 | 100.2 | 100.2 | 100.1
-0.4 | 99.5 | 99.9 | 100.1 | 100.1 | 100.2 | 100.2 | 100.1167 | 100
-0.3 | 99.8 | 99.95 | 100 | 100.1 | 100.1 | 100.1 | 100.1 | 100
-0.2 | 100 | 100 | 100 | 100 | 100.1 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100.025 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.7 | 100 | 100 | 100.1 | 100.1 | 100 | 100 | 100
0.4 | 99.35 | 99.9 | 100 | 100.1 | 100.1 | 100 | 100 | 100.025
0.5 | 98.95 | 99.85 | 100.1 | 100.1 | 100.2 | 100.1 | 100.15 | 100.1
0.6 | 98.35 | 99.7 | 100.1 | 100.2 | 100.2 | 100.125 | 100.2167 | 100.1667
0.7 | 97.6 | 99.4 | 100.2 | 100.2 | 100.3 | 100.2 | 100.3 | 100.2
0.8 | 96.8 | 99.3 | 100.2 | 100.3 | 100.375 | 100.3 | 100.4 | 100.3125
0.9 | 95 | 99 | 100.2 | 100.35 | 100.5 | 100.35 | 100.4 | 100.4
1 | 92.8 | 98.6 | 100.2 | 100.4 | 100.6 | 100.4 | 100.5 | 100.4625
1.1 | 89.5 | 97.9333 | 100.2 | 100.4 | 100.6 | 100.54 | 100.5571 | 100.5
1.2 | 84.7 | 97.2667 | 100.2 | 100.5 | 100.7 | 100.6 | 100.64 | 100.6
1.3 | 77.2 | 96.1 | 100.1 | 100.5 | 100.7 | 100.725 | 100.7167 | 100.6833
1.4 | 68.15 | 94.65 | 100 | 100.6 | 100.8 | 100.875 | 100.8 | 100.75
1.5 | 57.5 | 92.5333 | 99.9 | 100.6 | 100.8 | 100.9 | 100.95 | 100.8875
1.6 | 46.1 | 89.9 | 99.8 | 100.6 | 100.875 | 100.95 | 101.1 | 100.9125
1.7 | 35.1 | 85.6 | 99.6 | 100.7 | 101 | 101.1 | 101.1 | 101.0375
1.8 | 26.1 | 78.8 | 99.4 | 100.7 | 101.1 | 101.2 | 101.2 | 101.1833
1.9 | 18.6 | 69.7 | 99 | 100.7 | 101.1 | 101.225 | 101.3 | 101.25
2 | 14 | 59.3 | 98.65 | 100.7 | 101.2 | 101.375 | 101.3857 | 101.45
2.1 | 10.8667 | 49.25 | 98.15 | 100.7 | 101.3 | 101.4 | 101.65 | 101.6
2.2 | 8.6333 | 38.35 | 97.5 | 100.8 | 101.4 | 101.575 | 101.7143 | 101.7
2.3 | 7.05 | 28.75 | 96.6 | 100.7 | 101.4 | 101.8 | 101.88 | 101.7714
2.4 | 5.9 | 21.1 | 95.3 | 100.7 | 101.475 | 101.95 | 101.9667 | 101.9333
2.5 | 4.9 | 15.6 | 93.9 | 100.6 | 101.5 | 102 | 102.1667 | 102.0857
2.6 | 4.1 | 12.35 | 91.3333 | 100.5 | 101.62 | 102.1 | 102.2 | 102.1375
2.7 | 3.6 | 10.05 | 88 | 100.3 | 101.7 | 102.2 | 102.3286 | 102.3667
2.8 | 3.15 | 8.2667 | 82.6 | 100.1 | 101.7 | 102.3 | 102.4833 | 102.4
2.9 | 2.8 | 6.9 | 74.7 | 99.85 | 101.7 | 102.42 | 102.6 | 102.5667
3 | 2.55 | 5.9 | 65.3 | 99.35 | 101.7 | 102.55 | 102.7 | 102.7
3.1 | 2.25 | 5 | 55.2 | 98.95 | 101.7 | 102.6 | 102.8 | 102.8
3.2 | 2.1 | 4.4 | 44.6 | 98.3 | 101.6 | 102.7 | 102.8667 | 102.8625
3.3 | 1.85 | 3.9 | 34.7 | 97.5 | 101.6 | 102.8 | 102.9 | 102.9
3.4 | 1.7 | 3.5 | 25.9 | 96.3 | 101.5 | 102.8 | 103 | 102.92
3.5 | 1.6 | 3.1333 | 19.5 | 94.8 | 101.35 | 102.9 | 103.0833 | 103
3.6 | 1.5 | 2.9 | 15.3 | 92.8 | 101 | 102.9 | 103.1 | 103.1
3.7 | 1.4 | 2.6 | 12.4 | 90.2 | 100.9 | 103 | 103.2 | 103.1667
3.8 | 1.3 | 2.4 | 10.5 | 86.05 | 100.6333 | 103 | 103.2 | 103.2
3.9 | 1.2 | 2.2 | 8.9 | 80.15 | 100.3 | 103.1 | 103.2667 | 103.3
4 | 1.2 | 2.0667 | 7.7 | 71.3 | 99.85 | 103.1 | 103.3167 | 103.3714
4.1 | 1.1 | 2 | 6.7 | 61.4 | 99.45 | 103.2 | 103.4 | 103.4
4.2 | 1.05 | 1.8 | 5.9333 | 51.1 | 98.85 | 103.2 | 103.5 | 103.5
4.3 | 1 | 1.7 | 5.3 | 40.7 | 98.1 | 103.2 | 103.5143 | 103.5
4.4 | 1 | 1.6 | 4.8 | 31.8 | 97.2 | 103.3 | 103.6 | 103.6
4.5 | 0.95 | 1.5 | 4.4 | 24.1 | 96 | 103.3 | 103.6333 | 103.6
4.6 | 0.9 | 1.5 | 4 | 18.7 | 94.5 | 103.4 | 103.7 | 103.7
4.7 | 0.9 | 1.4 | 3.7 | 15.15 | 92.2 | 103.4 | 103.7 | 103.7
4.8 | 0.8 | 1.3 | 3.45 | 12.5 | 89.2 | 103.4 | 103.8 | 103.74
4.9 | 0.8 | 1.3 | 3.3 | 10.65 | 84.425 | 103.5 | 103.8 | 103.8
5 | 0.8 | 1.2 | 3.05 | 9.3 | 77.42 | 103.5 | 103.8833 | 103.8857
5.1 | 0.8 | 1.2 | 2.9 | 8.2 | 68.775 | 103.5 | 103.9 | 103.9125
5.2 | 0.8 | 1.15 | 2.7 | 7.4 | 58.475 | 103.5 | 103.9 | 104
5.3 | 0.7 | 1.1 | 2.6 | 6.65 | 48.175 | 103.5 | 104 | 104
5.4 | 0.7 | 1.0667 | 2.4 | 5.9667 | 37.9333 | 103.5 | 104 | 104.0875
5.5 | 0.7 | 1 | 2.3 | 5.5333 | 29.6 | 103.5 | 104 | 104.1
5.6 | 0.7 | 1 | 2.2 | 5.15 | 22.6 | 103.5 | 104 | 104.1
5.7 | 0.7 | 1 | 2.1 | 4.7 | 17.9 | 103.6 | 104 | 104.1
5.8 | 0.6 | 0.9 | 2 | 4.5 | 14.9 | 103.6 | 104.1 | 104.175
5.9 | 0.6 | 0.9 | 1.9333 | 4.2 | 12.6 | 103.6 | 104.1 | 104.2
6 | 0.6 | 0.9 | 1.9 | 3.95 | 10.9333 | 103.6 | 104.1 | 104.2
6.1 | 0.6 | 0.9 | 1.8 | 3.7 | 9.65 | 103.5 | 104.1 | 104.2
6.2 | 0.6 | 0.8 | 1.7 | 3.5 | 8.6 | 103.475 | 104.1 | 104.2
6.3 | 0.6 | 0.8 | 1.7 | 3.4 | 7.825 | 103.4 | 104.1 | 104.275
6.4 | 0.6 | 0.8 | 1.6 | 3.2 | 7.1 | 103.4 | 104.1 | 104.3
6.5 | 0.6 | 0.8 | 1.5 | 3 | 6.55 | 103.4 | 104.1 | 104.3
6.6 |  | 0.8 | 1.5 | 2.9 | 6.075 | 103.4 | 104.1 | 104.3
6.7 |  | 0.7 | 1.5 | 2.8 | 5.65 | 103.4 | 104.1 | 104.3
6.8 |  | 0.7 | 1.4 | 2.7 | 5.3 | 103.375 | 104.1 | 104.3
6.9 |  | 0.7 | 1.35 | 2.6 | 5 | 103.3 | 104.1 | 104.3
7 |  | 0.7 | 1.3 | 2.5 | 4.7 | 103.275 | 104.1 | 104.3
7.1 |  |  | 1.3 | 2.4 | 4.475 | 103.2 | 104.1 | 104.3
7.2 |  |  | 1.2 | 2.3 | 4.2333 | 103.1333 | 104.1 | 104.3
7.3 |  |  | 1.2 | 2.2 | 4.1 | 103.1 | 104.1 | 104.3
7.4 |  |  | 1.2 | 2.1 | 3.9 | 103 | 104.0333 | 104.3
7.5 |  |  | 1.1 | 2.05 | 3.7 | 103 | 104 | 104.3
7.6 |  |  | 1.1 | 2 | 3.6 | 102.95 | 104 | 104.3
7.7 |  |  | 1.1 | 1.9 | 3.4 | 102.875 | 104.0667 | 104.3
7.8 |  |  | 1.1 | 1.9 | 3.3 | 102.8 | 104.1 | 104.3
7.9 |  |  | 1 | 1.8 | 3.175 | 102.7 | 104.1 | 104.4
8 |  |  | 1 | 1.8 | 3.025 | 102.65 | 104.1 | 104.4
8.1 |  |  | 1 | 1.7 | 2.95 | 102.6 | 104.1 | 104.4
8.2 |  |  |  | 1.65 | 2.8 | 102.5 | 104.1 | 104.4
8.3 |  |  |  | 1.6 | 2.75 | 102.4 | 104.1 | 104.4
8.4 |  |  |  | 1.6 | 2.7 | 102.3 | 104.1 | 104.4
8.5 |  |  |  | 1.5 | 2.6 | 102.1 | 104.1 | 104.4
8.6 |  |  |  | 1.5 | 2.5 | 102 | 104.1 | 104.4
8.7 |  |  |  | 1.45 | 2.44 | 101.8 | 104.1 | 104.4
8.8 |  |  |  | 1.4 | 2.3 | 101.6333 | 104.0143 | 104.4
8.9 |  |  |  | 1.4 | 2.3 | 101.375 | 104 | 104.4
9 |  |  |  | 1.3 | 2.2 | 101.05 | 104 | 104.4
9.1 |  |  |  | 1.3 | 2.1 | 100.7 | 104 | 104.4
9.2 |  |  |  | 1.3 | 2.1 | 100.2 | 103.9 | 104.4
9.3 |  |  |  |  | 2 | 99.7 | 103.9 | 104.4
9.4 |  |  |  |  | 1.96 | 99.1 | 103.8 | 104.4
9.5 |  |  |  |  | 1.9 | 98.3 | 103.8 | 104.4
9.6 |  |  |  |  | 1.9 | 97.5 | 103.8 | 104.4
9.7 |  |  |  |  | 1.8 | 96.5 | 103.7 | 104.4
9.8 |  |  |  |  | 1.8 | 95.1667 | 103.7 | 104.4
9.9 |  |  |  |  | 1.7 | 93.325 | 103.7 | 104.4
10 |  |  |  |  | 1.7 | 91 | 103.7 | 104.4
10.1 |  |  |  |  | 1.7 | 87.575 | 103.7 | 104.4
10.2 | [5 empty] | 82.125 | 103.8 | 104.5
10.3 | [5 empty] | 74.96 | 103.8 | 104.525
10.4 | [5 empty] | 66.025 | 103.84 | 104.6
10.5 | [5 empty] | 56.4 | 103.9167 | 104.6
10.6 | [5 empty] | 46 | 104 | 104.725
10.7 | [5 empty] | 37.05 | 104 | 104.8
10.8 | [5 empty] | 28.6 | 104.1 | 104.8875
10.9 | [5 empty] | 22.8 | 104.1 | 104.9
11 | [5 empty] | 19.08 | 104.1 | 105
11.1 | [5 empty] | 16.34 | 104.1 | 105
11.2 | [5 empty] | 14.3 | 104.1 | 105.1
11.3 | [5 empty] | 12.8 | 104.2 | 105.1
11.4 | [5 empty] | 11.775 | 104.2 | 105.1
11.5 | [5 empty] | 10.78 | 104.2 | 105.1
11.6 | [5 empty] | 9.975 | 104.2 | 105.2
11.7 | [5 empty] | 9.35 | 104.1333 | 105.2
11.8 | [5 empty] | 8.85 | 104.1 | 105.2
11.9 | [5 empty] | 8.35 | 104.1 | 105.2
12 | [5 empty] | 7.925 | 104.1 | 105.2
12.1 | [5 empty] | 7.65 | 104.0167 | 105.275
12.2 | [5 empty] | 7.25 | 103.9167 | 105.2
12.3 | [5 empty] | 7 | 103.9 | 105.3
12.4 | [5 empty] | 6.7 | 103.9 | 105.3
12.5 | [5 empty] | 6.5 | 103.88 | 105.3
12.6 | [5 empty] | 6.25 | 103.8 | 105.3
12.7 | [5 empty] | 6.02 | 103.8 | 105.3
12.8 | [5 empty] | 5.875 | 103.7 | 105.3
12.9 | [5 empty] | 5.7 | 103.6429 | 105.3
13 | [5 empty] | 5.475 | 103.5 | 105.3
13.1 | [5 empty] | 5.325 | 103.4 | 105.2143
13.2 | [5 empty] | 5.1667 | 103.4 | 105.2
13.3 | [5 empty] | 5.04 | 103.3 | 105.2
13.4 | [5 empty] | 4.9 | 103.2 | 105.1833
13.5 | [5 empty] | 4.8 | 103.05 | 105.1
13.6 | [5 empty] | 4.64 | 102.8857 | 105.1
13.7 | [5 empty] | 4.5 | 102.7333 | 105.1
13.8 | [5 empty] | 4.4 | 102.66 | 105.1
13.9 | [5 empty] | 4.275 | 102.5667 | 105.1
14 | [5 empty] | 4.125 | 102.3 | 105.1
14.1 | [5 empty] | 4.08 | 102.0333 | 105.1
14.2 | [5 empty] | 3.925 | 101.8333 | 105.1
14.3 | [5 empty] | 3.9 | 101.5 | 105.1
14.4 | [5 empty] | 3.75 | 101.2 | 105.1
14.5 | [5 empty] | 3.7 | 100.8667 | 105.1
14.6 | [5 empty] | 3.6 | 100.45 | 105.1
14.7 | [5 empty] | 3.5 | 99.9 | 105.1
14.8 | [5 empty] | 3.4 | 99.1167 | 105.1
14.9 | [5 empty] | 3.3 | 98.2 | 105.1
15 | [5 empty] | 3.3 | 97.2 | 105.16
15.1 | [5 empty] | 3.2 | 95.7286 | 105.2
15.2 | [5 empty] | 3.1 | 93.9 | 105.2
15.3 | [5 empty] | 3.06 | 91.22 | 105.2
15.4 | [5 empty] | 3 | 87 | 105.3
15.5 | [6 empty] | 80.58 | 105.3
15.6 | [6 empty] | 72.6 | 105.3
15.7 | [6 empty] | 62.5 | 105.3
15.8 | [6 empty] | 52.7 | 105.3
15.9 | [6 empty] | 42.9 | 105.3
16 | [6 empty] | 34.55 | 105.2
16.1 | [6 empty] | 27.55 | 105.2
16.2 | [6 empty] | 22.25 | 105.15
16.3 | [6 empty] | 18.98 | 105.1
16.4 | [6 empty] | 16.65 | 105
16.5 | [6 empty] | 14.8667 | 105
16.6 | [6 empty] | 13.7 | 105
16.7 | [6 empty] | 12.6667 | 104.9
16.8 | [6 empty] | 11.7667 | 104.85
16.9 | [6 empty] | 11.1 | 104.8
17 | [6 empty] | 10.5 | 104.7
17.1 | [6 empty] | 10 | 104.6
17.2 | [6 empty] | 9.6 | 104.5667
17.3 | [6 empty] | 9.2 | 104.5
17.4 | [6 empty] | 8.8667 | 104.4
17.5 | [6 empty] | 8.5667 | 104.4
17.6 | [6 empty] | 8.2333 | 104.3
17.7 | [6 empty] | 8.05 | 104.3
17.8 | [6 empty] | 7.7833 | 104.2
17.9 | [6 empty] | 7.6167 | 104.2
18 | [6 empty] | 7.3571 | 104.2
18.1 | [6 empty] | 7.2 | 104.2
18.2 | [6 empty] | 7.0333 | 104.2
18.3 | [6 empty] | 6.8333 | 104.2
18.4 | [6 empty] | 6.6667 | 104.2
18.5 | [6 empty] | 6.5 | 104.2
18.6 | [6 empty] | 6.3333 | 104.2
18.7 | [6 empty] | 6.1667 | 104.16
18.8 | [6 empty] | 6.0143 | 104.1
18.9 | [6 empty] | 5.95 | 104.1
19 | [6 empty] | 5.7833 | 103.98
19.1 | [6 empty] | 5.6167 | 103.7833
19.2 | [6 empty] | 5.5 | 103.6
19.3 | [6 empty] | 5.38 | 103.4333
19.4 | [6 empty] | 5.3 | 103.1667
19.5 | [6 empty] | 5.12 | 102.8667
19.6 | [6 empty] | 5.0333 | 102.5
19.7 | [6 empty] | 4.94 | 102.15
19.8 | [6 empty] | 4.8 | 101.6
19.9 | [6 empty] | 4.7 | 101.1333
20 | [6 empty] | 4.6 | 100.54
20.1 | [6 empty] | 4.5 | 99.9
20.2 | [6 empty] | 4.4167 | 98.8286
20.3 | [6 empty] | 4.35 | 97.8
20.4 | [6 empty] | 4.2 | 96.25
20.5 | [6 empty] | 4.1167 | 94.1143
20.6 | [6 empty] | 4.0833 | 90.9714
20.7 | [7 empty] | 85.75
20.8 | [7 empty] | 78.7625
20.9 | [7 empty] | 70.2
21 | [7 empty] | 60.55
21.1 | [7 empty] | 50.4333
21.2 | [7 empty] | 40.3833
21.3 | [7 empty] | 31.9
21.4 | [7 empty] | 25.9
21.5 | [7 empty] | 21.4714
21.6 | [7 empty] | 18.5714
21.7 | [7 empty] | 16.5625
21.8 | [7 empty] | 15.0375
21.9 | [7 empty] | 13.7143
22 | [7 empty] | 12.8
22.1 | [7 empty] | 12.0333
22.2 | [7 empty] | 11.35
22.3 | [7 empty] | 10.82
22.4 | [7 empty] | 10.3
22.5 | [7 empty] | 9.9
22.6 | [7 empty] | 9.5667
22.7 | [7 empty] | 9.1857
22.8 | [7 empty] | 8.9222
22.9 | [7 empty] | 8.675
23 | [7 empty] | 8.4143
23.1 | [7 empty] | 8.1
23.2 | [7 empty] | 7.925
23.3 | [7 empty] | 7.7167
23.4 | [7 empty] | 7.45
23.5 | [7 empty] | 7.3125
23.6 | [7 empty] | 7.075
23.7 | [7 empty] | 6.9
23.8 | [7 empty] | 6.65
23.9 | [7 empty] | 6.5
24 | [7 empty] | 6.375
24.1 | [7 empty] | 6.1625
24.2 | [7 empty] | 6.0375
24.3 | [7 empty] | 5.825
24.4 | [7 empty] | 5.6875
24.5 | [7 empty] | 5.5571
24.6 | [7 empty] | 5.4143
24.7 | [7 empty] | 5.2875
24.8 | [7 empty] | 5.1571
24.9 | [7 empty] | 5
25 | [7 empty] | 4.9571
25.1 | [7 empty] | 4.8
25.2 | [7 empty] | 4.7429
25.3 | [7 empty] | 4.5857
25.4 | [7 empty] | 4.5
25.5 | [7 empty] | 4.4143
25.6 | [7 empty] | 4.3
25.7 | [7 empty] | 4.2167
25.8 | [7 empty] | 4.1
25.9 | [7 empty] | 4.0556

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-26.9 | [7 empty] | 5.5
-26.8 | [7 empty] | 5.6
-26.7 | [7 empty] | 5.6833
-26.6 | [7 empty] | 5.8286
-26.5 | [7 empty] | 5.9714
-26.4 | [7 empty] | 6.1
-26.3 | [7 empty] | 6.2
-26.2 | [7 empty] | 6.325
-26.1 | [7 empty] | 6.45
-26 | [7 empty] | 6.575
-25.9 | [7 empty] | 6.8
-25.8 | [7 empty] | 6.9
-25.7 | [7 empty] | 7.075
-25.6 | [7 empty] | 7.2
-25.5 | [7 empty] | 7.4125
-25.4 | [7 empty] | 7.5375
-25.3 | [7 empty] | 7.6714
-25.2 | [7 empty] | 7.9125
-25.1 | [7 empty] | 8.075
-25 | [7 empty] | 8.2625
-24.9 | [7 empty] | 8.475
-24.8 | [7 empty] | 8.6571
-24.7 | [7 empty] | 8.8714
-24.6 | [7 empty] | 9.0286
-24.5 | [7 empty] | 9.3
-24.4 | [7 empty] | 9.55
-24.3 | [7 empty] | 9.8
-24.2 | [7 empty] | 10.0444
-24.1 | [7 empty] | 10.275
-24 | [7 empty] | 10.525
-23.9 | [7 empty] | 10.8625
-23.8 | [7 empty] | 11.1375
-23.7 | [7 empty] | 11.5
-23.6 | [7 empty] | 11.8625
-23.5 | [7 empty] | 12.2857
-23.4 | [7 empty] | 12.7
-23.3 | [7 empty] | 13.2286
-23.2 | [7 empty] | 13.775
-23.1 | [7 empty] | 14.5
-23 | [7 empty] | 15.22
-22.9 | [7 empty] | 16.225
-22.8 | [7 empty] | 17.3667
-22.7 | [7 empty] | 18.9
-22.6 | [7 empty] | 20.8667
-22.5 | [7 empty] | 23.7
-22.4 | [7 empty] | 27.6429
-22.3 | [7 empty] | 33.475
-22.2 | [7 empty] | 40.52
-22.1 | [7 empty] | 48.9
-22 | [7 empty] | 57.8429
-21.9 | [7 empty] | 67.125
-21.8 | [7 empty] | 74.65
-21.7 | [7 empty] | 81.3714
-21.6 | [7 empty] | 85.8667
-21.5 | [6 empty] | 5.15 | 89.1
-21.4 | [6 empty] | 5.3 | 91.0857
-21.3 | [6 empty] | 5.3833 | 92.65
-21.2 | [6 empty] | 5.4429 | 93.9
-21.1 | [6 empty] | 5.6 | 94.8857
-21 | [6 empty] | 5.7 | 95.64
-20.9 | [6 empty] | 5.86 | 96.4143
-20.8 | [6 empty] | 5.95 | 96.96
-20.7 | [6 empty] | 6.0333 | 97.45
-20.6 | [6 empty] | 6.2 | 97.9125
-20.5 | [6 empty] | 6.35 | 98.35
-20.4 | [6 empty] | 6.4286 | 98.8
-20.3 | [6 empty] | 6.5833 | 99.0571
-20.2 | [6 empty] | 6.68 | 99.325
-20.1 | [6 empty] | 6.86 | 99.6
-20 | [6 empty] | 7.0429 | 99.85
-19.9 | [6 empty] | 7.2 | 100
-19.8 | [6 empty] | 7.2667 | 100.1429
-19.7 | [6 empty] | 7.4333 | 100.2857
-19.6 | [6 empty] | 7.6 | 100.4286
-19.5 | [6 empty] | 7.7667 | 100.5833
-19.4 | [6 empty] | 7.9667 | 100.6286
-19.3 | [6 empty] | 8.2 | 100.7714
-19.2 | [6 empty] | 8.3667 | 100.8286
-19.1 | [6 empty] | 8.5667 | 100.9
-19 | [6 empty] | 8.75 | 100.9
-18.9 | [6 empty] | 9.0167 | 101
-18.8 | [6 empty] | 9.1833 | 101.0667
-18.7 | [6 empty] | 9.4143 | 101.1143
-18.6 | [6 empty] | 9.64 | 101.2667
-18.5 | [6 empty] | 9.9333 | 101.325
-18.4 | [6 empty] | 10.2667 | 101.4
-18.3 | [6 empty] | 10.55 | 101.5286
-18.2 | [6 empty] | 10.86 | 101.6714
-18.1 | [6 empty] | 11.2 | 101.75
-18 | [6 empty] | 11.64 | 101.82
-17.9 | [6 empty] | 12.0429 | 101.9
-17.8 | [6 empty] | 12.5 | 102
-17.7 | [6 empty] | 13.0333 | 102.15
-17.6 | [6 empty] | 13.8 | 102.24
-17.5 | [6 empty] | 14.5333 | 102.325
-17.4 | [6 empty] | 15.4333 | 102.45
-17.3 | [6 empty] | 16.6 | 102.5
-17.2 | [6 empty] | 18.0714 | 102.6
-17.1 | [6 empty] | 20 | 102.7
-17 | [6 empty] | 22.7 | 102.75
-16.9 | [6 empty] | 26.25 | 102.8
-16.8 | [6 empty] | 31.8 | 102.9
-16.7 | [6 empty] | 39.2167 | 102.9429
-16.6 | [6 empty] | 47.65 | 103
-16.5 | [6 empty] | 56.7667 | 103.1
-16.4 | [6 empty] | 65.9333 | 103.1
-16.3 | [6 empty] | 73.8571 | 103.2
-16.2 | [6 empty] | 80.4333 | 103.2
-16.1 | [6 empty] | 85.6429 | 103.2
-16 | [5 empty] | 3.9 | 88.8333 | 103.2
-15.9 | [5 empty] | 4 | 91.0667 | 103.2
-15.8 | [5 empty] | 4 | 92.8 | 103.2
-15.7 | [5 empty] | 4.18 | 94.1333 | 103.275
-15.6 | [5 empty] | 4.225 | 95.0429 | 103.3
-15.5 | [5 empty] | 4.35 | 95.8167 | 103.3
-15.4 | [5 empty] | 4.4 | 96.5857 | 103.3714
-15.3 | [5 empty] | 4.575 | 97.1667 | 103.4
-15.2 | [5 empty] | 4.625 | 97.6667 | 103.4
-15.1 | [5 empty] | 4.7 | 98.1 | 103.4
-15 | [5 empty] | 4.8 | 98.5 | 103.4
-14.9 | [5 empty] | 4.9 | 98.7667 | 103.4
-14.8 | [5 empty] | 5.1 | 99.0857 | 103.4
-14.7 | [5 empty] | 5.2 | 99.3 | 103.5
-14.6 | [5 empty] | 5.35 | 99.5667 | 103.5
-14.5 | [5 empty] | 5.5 | 99.7333 | 103.5375
-14.4 | [5 empty] | 5.6 | 99.9 | 103.6
-14.3 | [5 empty] | 5.7 | 100.1667 | 103.6
-14.2 | [5 empty] | 5.9 | 100.35 | 103.7
-14.1 | [5 empty] | 6.0667 | 100.5167 | 103.7
-14 | [5 empty] | 6.14 | 100.8 | 103.8
-13.9 | [5 empty] | 6.35 | 100.85 | 103.8
-13.8 | [5 empty] | 6.5 | 101.1 | 103.8857
-13.7 | [5 empty] | 6.7 | 101.2667 | 103.9
-13.6 | [5 empty] | 6.9 | 101.35 | 103.9
-13.5 | [5 empty] | 7.1 | 101.5 | 103.9
-13.4 | [5 empty] | 7.3 | 101.6 | 104
-13.3 | [5 empty] | 7.475 | 101.68 | 104
-13.2 | [5 empty] | 7.74 | 101.75 | 104
-13.1 | [5 empty] | 7.975 | 101.8286 | 104
-13 | [5 empty] | 8.225 | 101.9833 | 104
-12.9 | [5 empty] | 8.475 | 102.05 | 104
-12.8 | [5 empty] | 8.8 | 102.1167 | 104.1
-12.7 | [5 empty] | 9.16 | 102.2 | 104.1
-12.6 | [5 empty] | 9.5333 | 102.2333 | 104.1
-12.5 | [5 empty] | 9.9 | 102.3 | 104.1
-12.4 | [5 empty] | 10.35 | 102.4 | 104.1
-12.3 | [5 empty] | 10.8667 | 102.4167 | 104.1857
-12.2 | [5 empty] | 11.4 | 102.5 | 104.2
-12.1 | [5 empty] | 12.125 | 102.6 | 104.2
-12 | [5 empty] | 12.96 | 102.7 | 104.2
-11.9 | [5 empty] | 13.9 | 102.7 | 104.2
-11.8 | [5 empty] | 15.075 | 102.8 | 104.2
-11.7 | [5 empty] | 16.625 | 102.8 | 104.2
-11.6 | [5 empty] | 18.525 | 102.9 | 104.1
-11.5 | [5 empty] | 21.08 | 102.9 | 104.1
-11.4 | [5 empty] | 24.72 | 102.9 | 104.1
-11.3 | [5 empty] | 30.05 | 102.9 | 104.0143
-11.2 | [5 empty] | 37.4 | 102.9 | 103.9333
-11.1 | [5 empty] | 46.22 | 102.9 | 103.9
-11 | [5 empty] | 55.3 | 102.8667 | 103.85
-10.9 | [5 empty] | 64.85 | 102.8 | 103.8
-10.8 | [5 empty] | 73.4 | 102.8 | 103.7
-10.7 | [5 empty] | 80.14 | 102.8 | 103.7
-10.6 | [5 empty] | 84.875 | 102.8 | 103.65
-10.5 |  |  |  |  | 2.2 | 88.325 | 102.8 | 103.6
-10.4 |  |  |  |  | 2.2333 | 90.7 | 102.8 | 103.6
-10.3 |  |  |  |  | 2.3 | 92.525 | 102.8 | 103.6
-10.2 |  |  |  |  | 2.4 | 93.825 | 102.8 | 103.6
-10.1 |  |  |  |  | 2.4 | 95 | 102.85 | 103.6
-10 |  |  |  |  | 2.5 | 95.96 | 102.9167 | 103.6333
-9.9 |  |  |  |  | 2.575 | 96.725 | 103 | 103.7
-9.8 |  |  |  |  | 2.6 | 97.3 | 103 | 103.7
-9.7 |  |  |  |  | 2.7 | 97.9 | 103 | 103.7
-9.6 |  |  |  |  | 2.7 | 98.34 | 103.1 | 103.725
-9.5 |  |  |  |  | 2.825 | 98.85 | 103.1 | 103.8
-9.4 |  |  |  |  | 2.9 | 99.2 | 103.15 | 103.8
-9.3 |  |  |  | 1.8 | 3 | 99.55 | 103.2167 | 103.8
-9.2 |  |  |  | 1.85 | 3.1 | 99.8 | 103.3 | 103.86
-9.1 |  |  |  | 1.9 | 3.2 | 100.14 | 103.3 | 103.9
-9 |  |  |  | 1.9 | 3.2667 | 100.3 | 103.3 | 103.9
-8.9 |  |  |  | 2 | 3.4 | 100.525 | 103.4 | 103.9
-8.8 |  |  |  | 2 | 3.5 | 100.7 | 103.4 | 103.8667
-8.7 |  |  |  | 2.1 | 3.5 | 100.9 | 103.4 | 103.8
-8.6 |  |  |  | 2.1 | 3.7 | 101.12 | 103.4 | 103.9
-8.5 |  |  |  | 2.2 | 3.8 | 101.25 | 103.4 | 103.9
-8.4 |  |  |  | 2.3 | 3.96 | 101.4 | 103.4 | 103.9
-8.3 |  |  |  | 2.3 | 4.1 | 101.4 | 103.5 | 103.9
-8.2 |  |  | 1.4 | 2.4 | 4.26 | 101.5 | 103.5 | 103.9
-8.1 |  |  | 1.4 | 2.5 | 4.4 | 101.6 | 103.5 | 103.8
-8 |  |  | 1.4333 | 2.6 | 4.54 | 101.7 | 103.5 | 103.8
-7.9 |  |  | 1.5 | 2.6 | 4.675 | 101.8 | 103.5 | 103.8
-7.8 |  |  | 1.5 | 2.7 | 4.9 | 101.9 | 103.5 | 103.8
-7.7 |  |  | 1.6 | 2.9 | 5.05 | 102.05 | 103.5 | 103.8
-7.6 |  |  | 1.6 | 2.9333 | 5.3 | 102.125 | 103.5 | 103.8
-7.5 |  |  | 1.7 | 3.05 | 5.475 | 102.2 | 103.5 | 103.8
-7.4 |  |  | 1.7 | 3.2 | 5.725 | 102.275 | 103.6 | 103.8
-7.3 |  |  | 1.7 | 3.2 | 6.06 | 102.32 | 103.6 | 103.8
-7.2 |  |  | 1.8 | 3.35 | 6.4 | 102.4667 | 103.6 | 103.8
-7.1 |  | 0.9 | 1.8 | 3.5 | 6.725 | 102.5 | 103.6 | 103.8
-7 |  | 0.9 | 1.9 | 3.7 | 7.06 | 102.55 | 103.62 | 103.8
-6.9 |  | 1 | 2 | 3.8 | 7.52 | 102.6 | 103.7 | 103.8
-6.8 |  | 1 | 2.05 | 3.9 | 7.92 | 102.7 | 103.7 | 103.8
-6.7 |  | 1 | 2.1 | 4.1 | 8.55 | 102.7 | 103.7 | 103.8
-6.6 | 0.7 | 1 | 2.2 | 4.25 | 9.15 | 102.8 | 103.7 | 103.8
-6.5 | 0.7 | 1.1 | 2.3 | 4.45 | 9.9 | 102.8 | 103.7 | 103.8
-6.4 | 0.8 | 1.1 | 2.35 | 4.6667 | 10.9667 | 102.9 | 103.7 | 103.8
-6.3 | 0.8 | 1.1 | 2.4 | 4.95 | 12.1 | 102.9 | 103.7 | 103.8
-6.2 | 0.8 | 1.2 | 2.5 | 5.15 | 13.65 | 103 | 103.7 | 103.8
-6.1 | 0.8 | 1.2 | 2.6 | 5.5 | 15.6 | 103 | 103.7 | 103.8
-6 | 0.8 | 1.2 | 2.7 | 5.8 | 18.1 | 103 | 103.7 | 103.8
-5.9 | 0.8 | 1.3 | 2.85 | 6.2 | 21.85 | 103.025 | 103.7 | 103.7125
-5.8 | 0.9 | 1.3 | 2.9333 | 6.5667 | 26.875 | 103.1 | 103.7 | 103.7
-5.7 | 0.9 | 1.4 | 3.1 | 7.0333 | 34.325 | 103.1 | 103.7 | 103.7
-5.6 | 0.9 | 1.4 | 3.2 | 7.6 | 43.35 | 103.1 | 103.7 | 103.7
-5.5 | 0.9 | 1.5 | 3.4 | 8.2 | 53.3 | 103.2 | 103.7 | 103.675
-5.4 | 1 | 1.5 | 3.6 | 9 | 63.35 | 103.2 | 103.7 | 103.6
-5.3 | 1 | 1.6 | 3.75 | 9.9 | 72.35 | 103.2 | 103.6 | 103.6
-5.2 | 1 | 1.6 | 4 | 11 | 79.9 | 103.2 | 103.6 | 103.6
-5.1 | 1.1 | 1.7 | 4.2 | 12.55 | 85.4 | 103.2 | 103.6 | 103.575
-5 | 1.1 | 1.75 | 4.4667 | 14.55 | 89.05 | 103.1 | 103.5 | 103.5
-4.9 | 1.1 | 1.9 | 4.7 | 17 | 91.8 | 103.1 | 103.5 | 103.4667
-4.8 | 1.2 | 1.9 | 5.05 | 20.2 | 93.7 | 103.1 | 103.4 | 103.4
-4.7 | 1.2 | 2 | 5.5 | 25.4 | 95.1 | 103.1 | 103.4 | 103.3667
-4.6 | 1.3 | 2.1 | 5.9 | 32.8 | 96.22 | 103.1 | 103.3667 | 103.3
-4.5 | 1.3 | 2.2 | 6.4 | 41.2 | 97.175 | 103 | 103.3 | 103.2
-4.4 | 1.4 | 2.3 | 7 | 51.5 | 97.84 | 103 | 103.24 | 103.15
-4.3 | 1.4 | 2.5 | 7.7 | 61.3 | 98.4 | 103 | 103.2 | 103.1
-4.2 | 1.5 | 2.65 | 8.7 | 70.9 | 98.96 | 103 | 103.1167 | 103.1
-4.1 | 1.6 | 2.8 | 9.8 | 78.4 | 99.375 | 102.9 | 103.1 | 103
-4 | 1.7 | 3 | 11.25 | 84.25 | 99.65 | 102.9 | 103 | 103
-3.9 | 1.75 | 3.2 | 13 | 88.55 | 100.06 | 102.9 | 103 | 102.9
-3.8 | 1.85 | 3.4333 | 15.4 | 91.5 | 100.3 | 102.8 | 103 | 102.9
-3.7 | 1.95 | 3.7 | 18.8 | 93.5 | 100.55 | 102.8 | 103 | 102.9
-3.6 | 2.1 | 4 | 23.9 | 95.05 | 100.7 | 102.8 | 103 | 102.8
-3.5 | 2.2 | 4.4 | 31.1 | 96.15 | 100.825 | 102.725 | 102.9833 | 102.8
-3.4 | 2.4 | 4.95 | 40.3 | 97.1 | 100.98 | 102.7 | 102.9 | 102.7
-3.3 | 2.6 | 5.5 | 50.4 | 97.8 | 101.2 | 102.65 | 102.8 | 102.6625
-3.2 | 2.8 | 6.2 | 60.4 | 98.4 | 101.32 | 102.6 | 102.7 | 102.5375
-3.1 | 3.1 | 7 | 69.65 | 99 | 101.45 | 102.5 | 102.6333 | 102.5
-3 | 3.4 | 8.2 | 78.4 | 99.4 | 101.5 | 102.45 | 102.5667 | 102.38
-2.9 | 3.8 | 9.6 | 84.6 | 99.7 | 101.5 | 102.4 | 102.5 | 102.3
-2.8 | 4.25 | 11.3 | 88.6 | 99.9 | 101.5 | 102.3 | 102.4 | 102.1333
-2.7 | 4.9 | 13.7 | 91.7 | 100 | 101.5 | 102.225 | 102.3 | 102
-2.6 | 5.6 | 17.3 | 93.6 | 100.2 | 101.4 | 102.175 | 102.3 | 101.9143
-2.5 | 6.5 | 22.4 | 95.1 | 100.3 | 101.4 | 102.025 | 102.15 | 101.8
-2.4 | 7.65 | 29.8 | 96.2 | 100.3 | 101.3 | 101.9 | 102.0833 | 101.7
-2.3 | 9.3333 | 39 | 97.1 | 100.4 | 101.2 | 101.8 | 101.9286 | 101.6143
-2.2 | 11.5 | 49.2 | 97.7333 | 100.4 | 101.2 | 101.65 | 101.7667 | 101.475
-2.1 | 14.2 | 59.4 | 98.35 | 100.45 | 101.1 | 101.5 | 101.6 | 101.4
-2 | 18.7 | 69.5 | 98.7667 | 100.5 | 101.1 | 101.4 | 101.45 | 101.325
-1.9 | 24.55 | 77.8 | 99.1 | 100.5 | 101 | 101.4 | 101.3 | 101.2
-1.8 | 33.1 | 84.6 | 99.3 | 100.5 | 101 | 101.3 | 101.2167 | 101.1625
-1.7 | 43 | 89 | 99.5 | 100.5 | 100.9 | 101.15 | 101.15 | 101.0375
-1.6 | 53.9 | 91.7 | 99.6 | 100.5 | 100.9 | 101 | 101.0857 | 101
-1.5 | 64.7 | 93.75 | 99.8 | 100.4 | 100.8 | 100.9333 | 101 | 100.8625
-1.4 | 74.5 | 95.4 | 100 | 100.4 | 100.76 | 100.8 | 100.9 | 100.7375
-1.3 | 82.6 | 96.6 | 100 | 100.4 | 100.7 | 100.8 | 100.8 | 100.7
-1.2 | 87.8333 | 97.45 | 100.1 | 100.4 | 100.6 | 100.7 | 100.7333 | 100.58
-1.1 | 91.7 | 98.1 | 100.1 | 100.3 | 100.54 | 100.65 | 100.6833 | 100.5
-1 | 93.9 | 98.7 | 100.1 | 100.3 | 100.5 | 100.6 | 100.6 | 100.4
-0.9 | 95.75 | 99 | 100.2 | 100.3 | 100.4 | 100.5 | 100.5 | 100.3429
-0.8 | 97.15 | 99.3 | 100.2 | 100.2 | 100.4 | 100.475 | 100.4 | 100.3
-0.7 | 97.9667 | 99.5 | 100.1 | 100.2 | 100.3 | 100.325 | 100.3333 | 100.2333
-0.6 | 98.7 | 99.6 | 100.1 | 100.2 | 100.3 | 100.3 | 100.24 | 100.1667
-0.5 | 99.3333 | 99.7 | 100.1 | 100.2 | 100.2 | 100.2 | 100.2 | 100.1
-0.4 | 99.6 | 99.8 | 100.1 | 100.2 | 100.1 | 100.125 | 100.1167 | 100.1
-0.3 | 99.8 | 99.8 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.0333
-0.2 | 99.9667 | 99.9 | 100 | 100.0333 | 100 | 100.025 | 100.05 | 100
-0.1 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.95 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.8667 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.7 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100
0.4 | 99.5 | 99.8 | 100.0667 | 100.1 | 100 | 100.1 | 100.0167 | 100.1
0.5 | 99.2 | 99.7 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1
0.6 | 98.6 | 99.6 | 100 | 100.2 | 100.1 | 100.175 | 100.15 | 100.2
0.7 | 98 | 99.5 | 100 | 100.2 | 100.2 | 100.225 | 100.22 | 100.2167
0.8 | 96.85 | 99.2667 | 100.1 | 100.3 | 100.2 | 100.3 | 100.3 | 100.3
0.9 | 95.7 | 98.9 | 100.1 | 100.3 | 100.3 | 100.4 | 100.3667 | 100.4
1 | 93.7 | 98.6 | 100.1 | 100.3 | 100.4 | 100.5 | 100.5 | 100.4
1.1 | 91.7 | 98.2 | 100.1 | 100.3 | 100.4 | 100.55 | 100.6 | 100.4857
1.2 | 87.6333 | 97.4 | 100 | 100.3 | 100.5 | 100.6 | 100.6 | 100.55
1.3 | 82.4 | 96.6 | 100 | 100.3 | 100.5 | 100.7 | 100.7 | 100.6875
1.4 | 74.3 | 95.4 | 99.9 | 100.4 | 100.6 | 100.775 | 100.7833 | 100.8
1.5 | 64.6 | 93.7 | 99.8 | 100.5 | 100.7 | 100.8 | 100.85 | 100.8375
1.6 | 53.9 | 91.8 | 99.6 | 100.5 | 100.7 | 100.9 | 100.9167 | 100.9714
1.7 | 43.1 | 89.2 | 99.5 | 100.5 | 100.7 | 101 | 101 | 101.1
1.8 | 33 | 84.7 | 99.3 | 100.5 | 100.8 | 101.075 | 101.1 | 101.1429
1.9 | 24.4667 | 78.1 | 99.15 | 100.5 | 100.8 | 101.1 | 101.2 | 101.2
2 | 18.6 | 69.6 | 98.75 | 100.5 | 100.9 | 101.2 | 101.3 | 101.3
2.1 | 14.15 | 59.6 | 98.3333 | 100.4 | 100.9 | 101.3 | 101.38 | 101.425
2.2 | 11.5 | 49.2 | 97.75 | 100.4 | 101 | 101.44 | 101.4571 | 101.55
2.3 | 9.2 | 38.9333 | 97.1 | 100.4 | 101 | 101.5 | 101.6 | 101.6
2.4 | 7.65 | 29.9 | 96.2 | 100.3 | 101 | 101.6 | 101.7 | 101.7
2.5 | 6.6 | 22.3 | 95 | 100.3 | 101.1 | 101.8 | 101.7833 | 101.85
2.6 | 5.6 | 17.1 | 93.4 | 100.2 | 101.1 | 101.9 | 101.95 | 101.9
2.7 | 4.9 | 13.65 | 91.6 | 100.1 | 101.2 | 102 | 102 | 102.05
2.8 | 4.3 | 11.2 | 88.6 | 99.9 | 101.2 | 102.05 | 102.1 | 102.2
2.9 | 3.8 | 9.5 | 84.65 | 99.55 | 101.2 | 102.1 | 102.2333 | 102.2667
3 | 3.4 | 8.2 | 77.8 | 99.3 | 101.15 | 102.2 | 102.3 | 102.4
3.1 | 3.1 | 7.0333 | 69.85 | 98.9 | 101.1 | 102.3 | 102.35 | 102.4875
3.2 | 2.8 | 6.2 | 59.95 | 98.3 | 101.025 | 102.375 | 102.5 | 102.6
3.3 | 2.6 | 5.5 | 50.4 | 97.8 | 100.9 | 102.5 | 102.5833 | 102.6429
3.4 | 2.4 | 4.95 | 40.4 | 97 | 100.84 | 102.5 | 102.6429 | 102.7
3.5 | 2.2 | 4.4 | 31.3 | 96.2 | 100.7 | 102.6 | 102.7 | 102.8
3.6 | 2.1 | 4 | 23.7 | 95.05 | 100.55 | 102.6 | 102.8 | 102.8
3.7 | 1.9 | 3.7 | 18.8 | 93.5 | 100.3 | 102.6 | 102.8 | 102.8833
3.8 | 1.85 | 3.4333 | 15.3 | 91.45 | 100.175 | 102.7 | 102.8 | 102.9
3.9 | 1.75 | 3.2 | 13.1 | 88.7 | 99.85 | 102.7 | 102.8857 | 102.9
4 | 1.6 | 3 | 11.2333 | 84.6 | 99.575 | 102.8 | 102.9 | 103
4.1 | 1.6 | 2.8 | 9.7 | 78.7 | 99.18 | 102.8 | 102.9 | 103
4.2 | 1.5 | 2.6 | 8.7 | 71.1 | 98.8 | 102.9 | 103 | 103.1
4.3 | 1.4 | 2.5 | 7.7 | 61.6 | 98.3 | 102.9 | 103.0333 | 103.1
4.4 | 1.4 | 2.3 | 7 | 51.5 | 97.6 | 102.9 | 103.1 | 103.1167
4.5 | 1.3 | 2.2 | 6.4 | 41.6 | 96.9667 | 102.9 | 103.1 | 103.2
4.6 | 1.3 | 2.1 | 5.9 | 32.8 | 96.1 | 102.9 | 103.1 | 103.2875
4.7 | 1.2 | 2 | 5.4333 | 25.8 | 94.9 | 102.9 | 103.1833 | 103.4
4.8 | 1.2 | 1.9 | 5.05 | 20.3 | 93.525 | 102.9 | 103.2 | 103.4
4.9 | 1.1 | 1.8 | 4.75 | 16.9667 | 91.55 | 102.9 | 103.2143 | 103.5
5 | 1.1 | 1.8 | 4.4 | 14.5 | 89.1 | 102.9 | 103.3 | 103.5
5.1 | 1.1 | 1.7 | 4.2 | 12.55 | 84.975 | 102.9 | 103.3 | 103.55
5.2 | 1 | 1.6 | 3.9 | 11 | 79.65 | 102.9 | 103.3 | 103.6
5.3 | 1 | 1.6 | 3.7 | 9.9 | 72.1 | 102.9 | 103.3 | 103.6
5.4 | 1 | 1.5 | 3.6 | 9 | 63.05 | 102.9 | 103.4 | 103.625
5.5 | 0.9 | 1.5 | 3.4 | 8.2 | 53.2 | 103 | 103.4 | 103.7
5.6 | 0.9 | 1.4 | 3.2 | 7.6 | 43.55 | 102.9 | 103.4 | 103.7
5.7 | 0.9 | 1.4 | 3.1 | 7 | 34.82 | 102.9 | 103.4 | 103.7
5.8 | 0.9 | 1.3 | 2.95 | 6.5667 | 27.14 | 102.9 | 103.4 | 103.7
5.9 | 0.9 | 1.3 | 2.8 | 6.2 | 21.85 | 102.9 | 103.4 | 103.7
6 | 0.8 | 1.2 | 2.7 | 5.8 | 18.14 | 102.875 | 103.5 | 103.8
6.1 | 0.8 | 1.2 | 2.6 | 5.4 | 15.625 | 102.8 | 103.5 | 103.8
6.2 | 0.8 | 1.2 | 2.5 | 5.15 | 13.625 | 102.8 | 103.5 | 103.8
6.3 | 0.8 | 1.1 | 2.5 | 4.95 | 12.175 | 102.75 | 103.5 | 103.8
6.4 | 0.7 | 1.1 | 2.3333 | 4.65 | 11 | 102.7 | 103.5 | 103.8
6.5 | 0.7 | 1.1 | 2.3 | 4.5 | 10.05 | 102.6 | 103.5 | 103.8
6.6 | 0.7 | 1 | 2.2 | 4.3 | 9.2 | 102.6 | 103.5 | 103.8
6.7 |  | 1 | 2.1 | 4.1 | 8.54 | 102.6 | 103.5 | 103.8
6.8 |  | 1 | 2 | 3.9 | 8 | 102.525 | 103.5 | 103.8
6.9 |  | 1 | 2 | 3.8 | 7.5 | 102.4 | 103.4 | 103.8
7 |  | 0.9 | 1.9 | 3.7 | 7.05 | 102.4 | 103.4 | 103.8
7.1 |  | 0.9 | 1.8 | 3.5 | 6.65 | 102.3 | 103.4 | 103.8
7.2 |  |  | 1.8 | 3.4 | 6.2667 | 102.2 | 103.4 | 103.8
7.3 |  |  | 1.7 | 3.3 | 6.02 | 102.2 | 103.4 | 103.8
7.4 |  |  | 1.7 | 3.1 | 5.775 | 102.1 | 103.4 | 103.8
7.5 |  |  | 1.6 | 3 | 5.48 | 102.1 | 103.3 | 103.8
7.6 |  |  | 1.6 | 2.9 | 5.26 | 102 | 103.3 | 103.8
7.7 |  |  | 1.55 | 2.8 | 5.05 | 101.9 | 103.3 | 103.8
7.8 |  |  | 1.5 | 2.7 | 4.9 | 101.8 | 103.3 | 103.8
7.9 |  |  | 1.5 | 2.7 | 4.7 | 101.6 | 103.2167 | 103.8
8 |  |  | 1.4 | 2.6 | 4.5 | 101.6 | 103.22 | 103.8
8.1 |  |  | 1.4 | 2.5 | 4.375 | 101.48 | 103.2 | 103.8
8.2 |  |  | 1.4 | 2.4 | 4.18 | 101.35 | 103.2 | 103.875
8.3 |  |  |  | 2.3667 | 4.06 | 101.3 | 103.2 | 103.9
8.4 |  |  |  | 2.3 | 3.925 | 101.2 | 103.2 | 103.9
8.5 |  |  |  | 2.2 | 3.8 | 101 | 103.2 | 103.9
8.6 |  |  |  | 2.2 | 3.62 | 100.825 | 103.2 | 103.9
8.7 |  |  |  | 2.1 | 3.5 | 100.7 | 103.15 | 103.8286
8.8 |  |  |  | 2 | 3.5 | 100.54 | 103.1 | 103.8429
8.9 |  |  |  | 2 | 3.375 | 100.3667 | 103.1 | 103.9
9 |  |  |  | 1.9 | 3.225 | 100.1667 | 103.1 | 103.9
9.1 |  |  |  | 1.9 | 3.1 | 100 | 103.1 | 103.9
9.2 |  |  |  | 1.8 | 3.1 | 99.7 | 103.1 | 103.8833
9.3 |  |  |  | 1.8 | 3 | 99.4 | 103.0714 | 103.8
9.4 |  |  |  |  | 2.925 | 99.15 | 103 | 103.8
9.5 |  |  |  |  | 2.8 | 98.7 | 102.95 | 103.8
9.6 |  |  |  |  | 2.7 | 98.25 | 102.9 | 103.75
9.7 |  |  |  |  | 2.675 | 97.74 | 102.8167 | 103.7
9.8 |  |  |  |  | 2.6 | 97.25 | 102.8 | 103.7
9.9 |  |  |  |  | 2.5 | 96.6 | 102.7 | 103.7
10 |  |  |  |  | 2.5 | 95.725 | 102.7 | 103.6556
10.1 |  |  |  |  | 2.4 | 94.8 | 102.6 | 103.6
10.2 |  |  |  |  | 2.375 | 93.8 | 102.6 | 103.6
10.3 |  |  |  |  | 2.3 | 92.38 | 102.6 | 103.6
10.4 |  |  |  |  | 2.2 | 90.6 | 102.6 | 103.6
10.5 | [5 empty] | 88.25 | 102.55 | 103.6
10.6 | [5 empty] | 84.8 | 102.56 | 103.6333
10.7 | [5 empty] | 79.55 | 102.5 | 103.7
10.8 | [5 empty] | 72.6 | 102.5 | 103.775
10.9 | [5 empty] | 64.44 | 102.5 | 103.8
11 | [5 empty] | 55.65 | 102.5833 | 103.9
11.1 | [5 empty] | 45.9667 | 102.6 | 103.9
11.2 | [5 empty] | 37.35 | 102.6 | 104
11.3 | [5 empty] | 30.4 | 102.6 | 104
11.4 | [5 empty] | 24.6 | 102.6 | 104.1
11.5 | [5 empty] | 21.05 | 102.6 | 104.1
11.6 | [5 empty] | 18.5 | 102.6 | 104.1
11.7 | [5 empty] | 16.475 | 102.6 | 104.2
11.8 | [5 empty] | 15.0333 | 102.55 | 104.2
11.9 | [5 empty] | 13.9 | 102.5 | 104.2
12 | [5 empty] | 12.9 | 102.5 | 104.2
12.1 | [5 empty] | 12.1 | 102.4 | 104.2
12.2 | [5 empty] | 11.4 | 102.3 | 104.2
12.3 | [5 empty] | 10.775 | 102.3 | 104.2
12.4 | [5 empty] | 10.3333 | 102.2333 | 104.1
12.5 | [5 empty] | 9.85 | 102.1667 | 104.1
12.6 | [5 empty] | 9.475 | 102 | 104.1
12.7 | [5 empty] | 9.125 | 101.9 | 104.1
12.8 | [5 empty] | 8.78 | 101.9 | 104.0375
12.9 | [5 empty] | 8.54 | 101.8 | 104
13 | [5 empty] | 8.26 | 101.8 | 104
13.1 | [5 empty] | 8 | 101.64 | 104
13.2 | [5 empty] | 7.75 | 101.5667 | 104
13.3 | [5 empty] | 7.4667 | 101.5 | 104
13.4 | [5 empty] | 7.3 | 101.3333 | 103.95
13.5 | [5 empty] | 7.1 | 101.2667 | 103.9
13.6 | [5 empty] | 6.9 | 101.1143 | 103.9
13.7 | [5 empty] | 6.675 | 101 | 103.9
13.8 | [5 empty] | 6.52 | 100.85 | 103.9
13.9 | [5 empty] | 6.375 | 100.6833 | 103.8
14 | [5 empty] | 6.125 | 100.5167 | 103.8
14.1 | [5 empty] | 5.975 | 100.35 | 103.7
14.2 | [5 empty] | 5.825 | 100.1833 | 103.7
14.3 | [5 empty] | 5.7 | 100.0167 | 103.6125
14.4 | [5 empty] | 5.6 | 99.76 | 103.6
14.5 | [5 empty] | 5.475 | 99.5667 | 103.5625
14.6 | [5 empty] | 5.26 | 99.3286 | 103.5
14.7 | [5 empty] | 5.15 | 99.1667 | 103.5
14.8 | [5 empty] | 5.025 | 98.9 | 103.4
14.9 | [5 empty] | 4.9 | 98.5 | 103.4
15 | [5 empty] | 4.8 | 98.1667 | 103.4
15.1 | [5 empty] | 4.7 | 97.8333 | 103.4
15.2 | [5 empty] | 4.6 | 97.4 | 103.4
15.3 | [5 empty] | 4.5 | 96.95 | 103.4
15.4 | [5 empty] | 4.4 | 96.34 | 103.3
15.5 | [5 empty] | 4.3 | 95.6333 | 103.3
15.6 | [5 empty] | 4.2 | 94.8833 | 103.3
15.7 | [5 empty] | 4.1 | 93.8 | 103.3
15.8 | [5 empty] | 4 | 92.45 | 103.2
15.9 | [5 empty] | 3.925 | 90.7833 | 103.2
16 | [6 empty] | 88.55 | 103.2
16.1 | [6 empty] | 85.1833 | 103.2
16.2 | [6 empty] | 80.8 | 103.2
16.3 | [6 empty] | 73.56 | 103.2
16.4 | [6 empty] | 65.9 | 103.1
16.5 | [6 empty] | 56.7 | 103.1
16.6 | [6 empty] | 47.4 | 103.0111
16.7 | [6 empty] | 39.18 | 102.9714
16.8 | [6 empty] | 32.1 | 102.9
16.9 | [6 empty] | 26.4667 | 102.8
17 | [6 empty] | 22.6333 | 102.775
17.1 | [6 empty] | 19.9 | 102.7
17.2 | [6 empty] | 18.04 | 102.625
17.3 | [6 empty] | 16.58 | 102.5
17.4 | [6 empty] | 15.4 | 102.475
17.5 | [6 empty] | 14.4667 | 102.3429
17.6 | [6 empty] | 13.6667 | 102.26
17.7 | [6 empty] | 13 | 102.1
17.8 | [6 empty] | 12.4 | 102
17.9 | [6 empty] | 11.95 | 101.92
18 | [6 empty] | 11.5 | 101.85
18.1 | [6 empty] | 11.14 | 101.78
18.2 | [6 empty] | 10.74 | 101.6
18.3 | [6 empty] | 10.4 | 101.5429
18.4 | [6 empty] | 10.1 | 101.3857
18.5 | [6 empty] | 9.8571 | 101.3
18.6 | [6 empty] | 9.6333 | 101.2
18.7 | [6 empty] | 9.3 | 101.1429
18.8 | [6 empty] | 9.12 | 101.0857
18.9 | [6 empty] | 8.8667 | 101
19 | [6 empty] | 8.7 | 100.9
19.1 | [6 empty] | 8.4333 | 100.9
19.2 | [6 empty] | 8.2333 | 100.8
19.3 | [6 empty] | 8.05 | 100.7
19.4 | [6 empty] | 7.8714 | 100.6571
19.5 | [6 empty] | 7.7167 | 100.6
19.6 | [6 empty] | 7.5167 | 100.45
19.7 | [6 empty] | 7.35 | 100.3143
19.8 | [6 empty] | 7.2667 | 100.1667
19.9 | [6 empty] | 7.0167 | 100.0143
20 | [6 empty] | 6.95 | 99.85
20.1 | [6 empty] | 6.8 | 99.62
20.2 | [6 empty] | 6.6286 | 99.3625
20.3 | [6 empty] | 6.4667 | 99.05
20.4 | [6 empty] | 6.4 | 98.7333
20.5 | [6 empty] | 6.2333 | 98.3
20.6 | [6 empty] | 6.1 | 97.8875
20.7 | [6 empty] | 5.9714 | 97.4143
20.8 | [6 empty] | 5.8167 | 96.98
20.9 | [6 empty] | 5.7333 | 96.4286
21 | [6 empty] | 5.6 | 95.6
21.1 | [6 empty] | 5.5333 | 94.8625
21.2 | [6 empty] | 5.4 | 93.8833
21.3 | [6 empty] | 5.3571 | 92.6333
21.4 | [6 empty] | 5.2333 | 91.1
21.5 | [7 empty] | 88.9333
21.6 | [7 empty] | 85.7714
21.7 | [7 empty] | 81.1
21.8 | [7 empty] | 74.725
21.9 | [7 empty] | 66.7429
22 | [7 empty] | 57.8
22.1 | [7 empty] | 48.66
22.2 | [7 empty] | 40.65
22.3 | [7 empty] | 33.4875
22.4 | [7 empty] | 27.6
22.5 | [7 empty] | 23.5167
22.6 | [7 empty] | 20.75
22.7 | [7 empty] | 18.82
22.8 | [7 empty] | 17.4
22.9 | [7 empty] | 16.25
23 | [7 empty] | 15.2
23.1 | [7 empty] | 14.4429
23.2 | [7 empty] | 13.8
23.3 | [7 empty] | 13.2143
23.4 | [7 empty] | 12.7143
23.5 | [7 empty] | 12.3
23.6 | [7 empty] | 11.8429
23.7 | [7 empty] | 11.4429
23.8 | [7 empty] | 11.175
23.9 | [7 empty] | 10.8375
24 | [7 empty] | 10.5625
24.1 | [7 empty] | 10.225
24.2 | [7 empty] | 9.9889
24.3 | [7 empty] | 9.75
24.4 | [7 empty] | 9.5
24.5 | [7 empty] | 9.275
24.6 | [7 empty] | 9.0857
24.7 | [7 empty] | 8.9
24.8 | [7 empty] | 8.65
24.9 | [7 empty] | 8.4
25 | [7 empty] | 8.275
25.1 | [7 empty] | 8.0625
25.2 | [7 empty] | 7.875
25.3 | [7 empty] | 7.6875
25.4 | [7 empty] | 7.5625
25.5 | [7 empty] | 7.375
25.6 | [7 empty] | 7.2222
25.7 | [7 empty] | 7.1
25.8 | [7 empty] | 6.8714
25.9 | [7 empty] | 6.7286
26 | [7 empty] | 6.6
26.1 | [7 empty] | 6.475
26.2 | [7 empty] | 6.3556
26.3 | [7 empty] | 6.2286
26.4 | [7 empty] | 6.0667
26.5 | [7 empty] | 5.9875
26.6 | [7 empty] | 5.85
26.7 | [7 empty] | 5.7
26.8 | [7 empty] | 5.5625
26.9 | [7 empty] | 5.5

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 8.2111
-28.8 | [7 empty] | 8.3333
-28.7 | [7 empty] | 8.5
-28.6 | [7 empty] | 8.6429
-28.5 | [7 empty] | 8.775
-28.4 | [7 empty] | 8.9
-28.3 | [7 empty] | 9.125
-28.2 | [7 empty] | 9.25
-28.1 | [7 empty] | 9.4
-28 | [7 empty] | 9.5429
-27.9 | [7 empty] | 9.7714
-27.8 | [7 empty] | 9.9125
-27.7 | [7 empty] | 10.05
-27.6 | [7 empty] | 10.2286
-27.5 | [7 empty] | 10.4429
-27.4 | [7 empty] | 10.6286
-27.3 | [7 empty] | 10.8429
-27.2 | [7 empty] | 11.04
-27.1 | [7 empty] | 11.22
-27 | [7 empty] | 11.4
-26.9 | [7 empty] | 11.6667
-26.8 | [7 empty] | 11.9333
-26.7 | [7 empty] | 12.12
-26.6 | [7 empty] | 12.3
-26.5 | [7 empty] | 12.6
-26.4 | [7 empty] | 12.825
-26.3 | [7 empty] | 13.075
-26.2 | [7 empty] | 13.325
-26.1 | [7 empty] | 13.6125
-26 | [7 empty] | 13.925
-25.9 | [7 empty] | 14.2333
-25.8 | [7 empty] | 14.6
-25.7 | [7 empty] | 14.8667
-25.6 | [7 empty] | 15.25
-25.5 | [7 empty] | 15.7143
-25.4 | [7 empty] | 16.15
-25.3 | [7 empty] | 16.6714
-25.2 | [7 empty] | 17.2333
-25.1 | [7 empty] | 17.925
-25 | [7 empty] | 18.6875
-24.9 | [7 empty] | 19.6375
-24.8 | [7 empty] | 20.825
-24.7 | [7 empty] | 22.25
-24.6 | [7 empty] | 24.1222
-24.5 | [7 empty] | 26.55
-24.4 | [7 empty] | 30.2
-24.3 | [7 empty] | 35.05
-24.2 | [7 empty] | 41
-24.1 | [7 empty] | 48.25
-24 | [7 empty] | 55.3111
-23.9 | [7 empty] | 62.3875
-23.8 | [7 empty] | 68.8875
-23.7 | [7 empty] | 74.325
-23.6 | [7 empty] | 78.3125
-23.5 | [7 empty] | 81.2143
-23.4 | [7 empty] | 83.2714
-23.3 | [7 empty] | 84.7833
-23.2 | [7 empty] | 85.925
-23.1 | [7 empty] | 87
-23 | [6 empty] | 7.4286 | 87.7667
-22.9 | [6 empty] | 7.5833 | 88.5167
-22.8 | [6 empty] | 7.7 | 89.1286
-22.7 | [6 empty] | 7.8286 | 89.675
-22.6 | [6 empty] | 8 | 90.1556
-22.5 | [6 empty] | 8.06 | 90.6143
-22.4 | [6 empty] | 8.22 | 91.0667
-22.3 | [6 empty] | 8.3857 | 91.4714
-22.2 | [6 empty] | 8.55 | 91.8625
-22.1 | [6 empty] | 8.7167 | 92.225
-22 | [6 empty] | 8.8833 | 92.5125
-21.9 | [6 empty] | 9.05 | 92.8111
-21.8 | [6 empty] | 9.2143 | 93.05
-21.7 | [6 empty] | 9.3667 | 93.3
-21.6 | [6 empty] | 9.5429 | 93.55
-21.5 | [6 empty] | 9.7143 | 93.75
-21.4 | [6 empty] | 9.8667 | 93.95
-21.3 | [6 empty] | 10.0333 | 94.1
-21.2 | [6 empty] | 10.3 | 94.2222
-21.1 | [6 empty] | 10.4667 | 94.3375
-21 | [6 empty] | 10.6571 | 94.4625
-20.9 | [6 empty] | 10.9 | 94.5875
-20.8 | [6 empty] | 11.0571 | 94.8143
-20.7 | [6 empty] | 11.3286 | 94.9571
-20.6 | [6 empty] | 11.4833 | 95.0778
-20.5 | [6 empty] | 11.75 | 95.2
-20.4 | [6 empty] | 12 | 95.3333
-20.3 | [6 empty] | 12.2333 | 95.475
-20.2 | [6 empty] | 12.4714 | 95.6
-20.1 | [6 empty] | 12.7667 | 95.8
-20 | [6 empty] | 13.1 | 95.975
-19.9 | [6 empty] | 13.3333 | 96.1286
-19.8 | [6 empty] | 13.6667 | 96.2833
-19.7 | [6 empty] | 14 | 96.44
-19.6 | [6 empty] | 14.4333 | 96.5875
-19.5 | [6 empty] | 14.8 | 96.7125
-19.4 | [6 empty] | 15.34 | 96.8857
-19.3 | [6 empty] | 15.7571 | 97.1714
-19.2 | [6 empty] | 16.35 | 97.325
-19.1 | [6 empty] | 16.9833 | 97.4
-19 | [6 empty] | 17.8167 | 97.4857
-18.9 | [6 empty] | 18.7 | 97.65
-18.8 | [6 empty] | 19.8 | 97.85
-18.7 | [6 empty] | 21.2429 | 97.9667
-18.6 | [6 empty] | 22.9571 | 98.0875
-18.5 | [6 empty] | 25.3 | 98.2125
-18.4 | [6 empty] | 28.8 | 98.3375
-18.3 | [6 empty] | 33.4667 | 98.4625
-18.2 | [6 empty] | 39.2 | 98.6
-18.1 | [6 empty] | 46.7 | 98.6
-18 | [6 empty] | 54.3 | 98.7667
-17.9 | [6 empty] | 62.1333 | 98.9
-17.8 | [6 empty] | 69.1 | 98.9
-17.7 | [6 empty] | 75.2 | 98.9857
-17.6 | [6 empty] | 79.6 | 99.05
-17.5 | [6 empty] | 82.6143 | 99.125
-17.4 | [6 empty] | 84.7 | 99.25
-17.3 | [6 empty] | 86.4 | 99.3
-17.2 | [6 empty] | 87.8 | 99.4
-17.1 | [6 empty] | 88.88 | 99.4375
-17 | [5 empty] | 6 | 89.6667 | 99.5
-16.9 | [5 empty] | 6.075 | 90.4 | 99.5875
-16.8 | [5 empty] | 6.15 | 91.02 | 99.625
-16.7 | [5 empty] | 6.28 | 91.5667 | 99.7
-16.6 | [5 empty] | 6.375 | 92.15 | 99.775
-16.5 | [5 empty] | 6.6 | 92.6 | 99.8167
-16.4 | [5 empty] | 6.7 | 92.92 | 99.9
-16.3 | [5 empty] | 6.8 | 93.3429 | 100
-16.2 | [5 empty] | 6.94 | 93.76 | 100
-16.1 | [5 empty] | 7.075 | 94.0333 | 100.0875
-16 | [5 empty] | 7.22 | 94.3333 | 100.15
-15.9 | [5 empty] | 7.4 | 94.5333 | 100.2125
-15.8 | [5 empty] | 7.5 | 94.9 | 100.3
-15.7 | [5 empty] | 7.7 | 95.16 | 100.3857
-15.6 | [5 empty] | 7.88 | 95.35 | 100.5
-15.5 | [5 empty] | 8.04 | 95.55 | 100.525
-15.4 | [5 empty] | 8.175 | 95.8 | 100.6571
-15.3 | [5 empty] | 8.375 | 95.9333 | 100.7167
-15.2 | [5 empty] | 8.5 | 96.2667 | 100.8
-15.1 | [5 empty] | 8.8 | 96.4667 | 100.9
-15 | [5 empty] | 8.9 | 96.7 | 101
-14.9 | [5 empty] | 9.1667 | 96.88 | 101.075
-14.8 | [5 empty] | 9.3 | 97.0333 | 101.125
-14.7 | [5 empty] | 9.55 | 97.2 | 101.2
-14.6 | [5 empty] | 9.8 | 97.4 | 101.275
-14.5 | [5 empty] | 10.04 | 97.58 | 101.3
-14.4 | [5 empty] | 10.225 | 97.7667 | 101.3857
-14.3 | [5 empty] | 10.475 | 97.8286 | 101.5
-14.2 | [5 empty] | 10.725 | 98 | 101.5
-14.1 | [5 empty] | 11.05 | 98.15 | 101.575
-14 | [5 empty] | 11.3 | 98.3167 | 101.6
-13.9 | [5 empty] | 11.65 | 98.4 | 101.6
-13.8 | [5 empty] | 12.02 | 98.5 | 101.6167
-13.7 | [5 empty] | 12.3667 | 98.6714 | 101.7
-13.6 | [5 empty] | 12.75 | 98.8 | 101.7429
-13.5 | [5 empty] | 13.2 | 98.9 | 101.8
-13.4 | [5 empty] | 13.6333 | 99.0429 | 101.8
-13.3 | [5 empty] | 14.14 | 99.2 | 101.9
-13.2 | [5 empty] | 14.7667 | 99.2857 | 101.9
-13.1 | [5 empty] | 15.425 | 99.4 | 101.9625
-13 | [5 empty] | 16.3 | 99.5167 | 102
-12.9 | [5 empty] | 17.25 | 99.6714 | 102.1
-12.8 | [5 empty] | 18.4 | 99.7333 | 102.1
-12.7 | [5 empty] | 19.8 | 99.9 | 102.2
-12.6 | [5 empty] | 21.825 | 99.9 | 102.2
-12.5 | [5 empty] | 24.1333 | 99.9167 | 102.2
-12.4 | [5 empty] | 27.55 | 100 | 102.1167
-12.3 | [5 empty] | 32.5 | 100.0833 | 102.2
-12.2 | [5 empty] | 38.9667 | 100.1 | 102.1
-12.1 | [5 empty] | 46.1 | 100.2 | 102.1
-12 | [5 empty] | 54.6 | 100.2 | 102.1
-11.9 | [5 empty] | 62.75 | 100.2333 | 102.1
-11.8 | [5 empty] | 70.125 | 100.3 | 102.1
-11.7 | [5 empty] | 76.525 | 100.3 | 102.1
-11.6 | [5 empty] | 81 | 100.3 | 102.1
-11.5 | [5 empty] | 84.42 | 100.4 | 102.1
-11.4 | [5 empty] | 86.5 | 100.4 | 102.1
-11.3 | [5 empty] | 88.4333 | 100.5 | 102.1
-11.2 | [5 empty] | 89.9 | 100.52 | 102.1
-11.1 | [5 empty] | 90.9 | 100.6 | 102.15
-11 |  |  |  |  | 3.55 | 91.8 | 100.6833 | 102.2
-10.9 |  |  |  |  | 3.6 | 92.6 | 100.7 | 102.2
-10.8 |  |  |  |  | 3.725 | 93.3 | 100.8 | 102.2
-10.7 |  |  |  |  | 3.86 | 93.9 | 100.8 | 102.3
-10.6 |  |  |  |  | 3.9 | 94.375 | 100.9 | 102.4
-10.5 |  |  |  |  | 4 | 94.9333 | 101 | 102.4
-10.4 |  |  |  |  | 4.1 | 95.375 | 101.1 | 102.4833
-10.3 |  |  |  |  | 4.1333 | 95.7 | 101.2 | 102.5
-10.2 |  |  |  |  | 4.3 | 96.1 | 101.2167 | 102.5
-10.1 |  |  |  |  | 4.4 | 96.4667 | 101.3 | 102.5875
-10 |  |  |  |  | 4.5 | 96.8 | 101.3429 | 102.6
-9.9 |  |  |  |  | 4.6 | 97.1 | 101.4 | 102.6
-9.8 |  |  |  |  | 4.7 | 97.325 | 101.5 | 102.6
-9.7 |  |  |  | 2.9 | 4.9 | 97.48 | 101.6 | 102.6
-9.6 |  |  |  | 3 | 5 | 97.725 | 101.6 | 102.6
-9.5 |  |  |  | 3.1 | 5.125 | 97.975 | 101.6 | 102.6
-9.4 |  |  |  | 3.2 | 5.26 | 98.1 | 101.6167 | 102.6
-9.3 |  |  |  | 3.2333 | 5.4 | 98.3 | 101.7 | 102.625
-9.2 |  |  |  | 3.3 | 5.5 | 98.5 | 101.7429 | 102.7
-9.1 |  |  |  | 3.4 | 5.675 | 98.64 | 101.8 | 102.7
-9 |  |  |  | 3.5 | 5.9 | 98.85 | 101.8 | 102.7
-8.9 |  |  |  | 3.6 | 6 | 99 | 101.8667 | 102.7
-8.8 |  |  |  | 3.7 | 6.2 | 99.075 | 101.9 | 102.7
-8.7 |  |  |  | 3.775 | 6.375 | 99.2 | 101.9833 | 102.7
-8.6 |  |  |  | 3.9 | 6.625 | 99.4 | 102 | 102.7
-8.5 |  |  | 2.2 | 4.05 | 6.775 | 99.6 | 102 | 102.7857
-8.4 |  |  | 2.3 | 4.1 | 7 | 99.68 | 102 | 102.8
-8.3 |  |  | 2.4 | 4.225 | 7.15 | 99.875 | 102.1 | 102.8
-8.2 |  |  | 2.4 | 4.4 | 7.42 | 100 | 102.1 | 102.8
-8.1 |  |  | 2.5 | 4.5 | 7.66 | 100.075 | 102.2 | 102.8
-8 |  |  | 2.5 | 4.6 | 7.9667 | 100.2 | 102.2 | 102.8
-7.9 |  |  | 2.6 | 4.7 | 8.15 | 100.3 | 102.3 | 102.8
-7.8 |  |  | 2.7 | 4.94 | 8.55 | 100.4 | 102.3 | 102.9
-7.7 |  |  | 2.8 | 5.075 | 8.8333 | 100.5 | 102.4 | 102.9
-7.6 |  |  | 2.8 | 5.2 | 9.24 | 100.6 | 102.4 | 102.9
-7.5 |  |  | 2.9 | 5.4 | 9.6667 | 100.7 | 102.4 | 102.9
-7.4 |  |  | 3 | 5.55 | 10.1 | 100.8 | 102.44 | 102.9
-7.3 |  | 1.5 | 3.1 | 5.68 | 10.65 | 100.9 | 102.5 | 102.9
-7.2 |  | 1.5 | 3.1 | 5.925 | 11.3 | 100.9 | 102.5 | 102.9
-7.1 |  | 1.55 | 3.2 | 6.15 | 11.95 | 101 | 102.5 | 103
-7 |  | 1.6 | 3.3 | 6.3 | 12.8 | 101.2 | 102.5167 | 103
-6.9 |  | 1.6 | 3.5 | 6.54 | 13.8 | 101.3 | 102.6 | 103
-6.8 |  | 1.7 | 3.6 | 6.825 | 15.1 | 101.3 | 102.6 | 103
-6.7 | 1.1 | 1.7 | 3.7 | 7.075 | 16.46 | 101.4 | 102.6 | 103
-6.6 | 1.1 | 1.8 | 3.8 | 7.325 | 18.4 | 101.525 | 102.6 | 103
-6.5 | 1.1 | 1.8 | 3.9 | 7.6667 | 20.9 | 101.6 | 102.7 | 103
-6.4 | 1.2 | 1.9 | 4 | 8.04 | 24.55 | 101.7 | 102.7 | 103
-6.3 | 1.2 | 1.9 | 4.2 | 8.5 | 29.475 | 101.7 | 102.7 | 103
-6.2 | 1.2 | 2 | 4.3 | 8.92 | 36.6667 | 101.8 | 102.7 | 103
-6.1 | 1.3 | 2 | 4.5 | 9.4333 | 44.9333 | 101.8 | 102.7 | 103
-6 | 1.3 | 2.1 | 4.65 | 9.95 | 53.74 | 101.8 | 102.7 | 102.9333
-5.9 | 1.4 | 2.2 | 4.8 | 10.675 | 62.5 | 101.8 | 102.7 | 102.9
-5.8 | 1.4 | 2.2 | 5 | 11.5 | 70.9 | 101.9 | 102.7 | 102.9
-5.7 | 1.4 | 2.3 | 5.2 | 12.4 | 77.9 | 101.9333 | 102.7 | 102.9
-5.6 | 1.4 | 2.4 | 5.4333 | 13.65 | 83.3 | 102 | 102.7 | 102.9
-5.5 | 1.5 | 2.4 | 5.7 | 15.225 | 87.1 | 102 | 102.7 | 102.9
-5.4 | 1.5 | 2.5 | 5.9 | 17.1 | 89.6 | 102.1 | 102.7 | 102.8625
-5.3 | 1.6 | 2.6 | 6.3 | 19.5 | 91.45 | 102.1 | 102.7 | 102.8
-5.2 | 1.6 | 2.7 | 6.5333 | 23 | 92.85 | 102.1 | 102.7 | 102.8
-5.1 | 1.7 | 2.8 | 7 | 27.875 | 94.2 | 102.1 | 102.6 | 102.8
-5 | 1.7 | 2.9 | 7.4 | 34.6 | 95.05 | 102.1 | 102.6 | 102.725
-4.9 | 1.8 | 3 | 7.8 | 43 | 95.9 | 102.1 | 102.6 | 102.7
-4.8 | 1.85 | 3.2 | 8.4333 | 52.075 | 96.6 | 102.1 | 102.6 | 102.7
-4.7 | 1.9 | 3.3 | 9.1333 | 61.1 | 97.16 | 102.1 | 102.6 | 102.625
-4.6 | 2 | 3.4 | 9.9 | 69.7667 | 97.55 | 102.1 | 102.6 | 102.6
-4.5 | 2.1 | 3.6 | 10.9 | 77.2 | 98.05 | 102.1 | 102.6 | 102.6
-4.4 | 2.2 | 3.8 | 11.9 | 83 | 98.35 | 102.1 | 102.5 | 102.5
-4.3 | 2.25 | 3.9 | 13.4333 | 87.1 | 98.76 | 102.1 | 102.5 | 102.5
-4.2 | 2.3 | 4.1 | 15.3 | 89.8 | 98.95 | 102.1 | 102.5 | 102.5
-4.1 | 2.4 | 4.4 | 17.6 | 91.8 | 99.3 | 102.1 | 102.4 | 102.4
-4 | 2.5667 | 4.7 | 21.1 | 93.5 | 99.5 | 102.1 | 102.4 | 102.4
-3.9 | 2.7 | 4.9667 | 26 | 94.7333 | 99.7 | 102.1 | 102.4 | 102.3333
-3.8 | 2.8 | 5.3 | 32.9333 | 95.65 | 99.8 | 102.1 | 102.4 | 102.3
-3.7 | 3 | 5.9 | 41.3 | 96.4 | 100.05 | 102.1 | 102.4 | 102.3
-3.6 | 3.1667 | 6.35 | 50.9 | 97.15 | 100.2 | 102.1 | 102.4 | 102.2
-3.5 | 3.3667 | 6.9 | 59.8333 | 97.65 | 100.3 | 102.1 | 102.35 | 102.2
-3.4 | 3.6 | 7.7 | 69.1 | 98.1 | 100.45 | 102 | 102.3 | 102.1
-3.3 | 3.95 | 8.6 | 77 | 98.5 | 100.5 | 102 | 102.2333 | 102
-3.2 | 4.25 | 9.7 | 83.3 | 98.9 | 100.6 | 102 | 102.2 | 102
-3.1 | 4.7 | 11.2 | 87.4 | 99.18 | 100.675 | 101.9 | 102.1143 | 101.9125
-3 | 5.1 | 13 | 90.3 | 99.4 | 100.7 | 101.8 | 102 | 101.8714
-2.9 | 5.7 | 15.5 | 92.4 | 99.525 | 100.7 | 101.8 | 101.9 | 101.7167
-2.8 | 6.4333 | 19 | 94.1 | 99.7 | 100.7 | 101.7 | 101.8167 | 101.6
-2.7 | 7.2667 | 24.1 | 95.5 | 99.8 | 100.78 | 101.6 | 101.7667 | 101.525
-2.6 | 8.3 | 30.55 | 96.2 | 99.85 | 100.75 | 101.5 | 101.7 | 101.5
-2.5 | 9.8 | 39.15 | 97.1 | 99.9 | 100.7 | 101.5 | 101.6 | 101.4
-2.4 | 11.6 | 49.2 | 97.7 | 100 | 100.7 | 101.4 | 101.5833 | 101.3857
-2.3 | 14.1 | 59.8 | 98.2 | 100.1 | 100.7 | 101.3 | 101.4 | 101.3
-2.2 | 17.4333 | 68.9667 | 98.55 | 100.2 | 100.7 | 101.225 | 101.3167 | 101.2
-2.1 | 22.7 | 77.25 | 98.7 | 100.2 | 100.7 | 101.2 | 101.25 | 101.1
-2 | 29.7 | 83.5 | 99 | 100.2 | 100.6 | 101.1 | 101.2 | 101.1
-1.9 | 38.5 | 87.9 | 99.2 | 100.2 | 100.6 | 101 | 101.0286 | 101
-1.8 | 49 | 90.9 | 99.35 | 100.2 | 100.6 | 100.9333 | 101 | 100.9
-1.7 | 59.4 | 93.15 | 99.5 | 100.2 | 100.6 | 100.9 | 100.9143 | 100.9
-1.6 | 68.8667 | 94.7 | 99.6 | 100.2 | 100.6 | 100.9 | 100.9 | 100.8
-1.5 | 77.5 | 96.1 | 99.8 | 100.2 | 100.575 | 100.8 | 100.8833 | 100.7286
-1.4 | 84.1333 | 97.1 | 99.9 | 100.2 | 100.5 | 100.7 | 100.7286 | 100.6
-1.3 | 88.7 | 97.7667 | 99.9 | 100.3 | 100.5 | 100.6 | 100.7 | 100.6
-1.2 | 91.8 | 98.2 | 100 | 100.3 | 100.4 | 100.6 | 100.6 | 100.5
-1.1 | 94 | 98.7 | 100 | 100.3 | 100.4 | 100.5 | 100.6 | 100.4
-1 | 95.7 | 99.1 | 100 | 100.2 | 100.3 | 100.46 | 100.4333 | 100.35
-0.9 | 96.9 | 99.35 | 100 | 100.2 | 100.25 | 100.4 | 100.3833 | 100.3
-0.8 | 97.8 | 99.6 | 100 | 100.1 | 100.2 | 100.3 | 100.3 | 100.2
-0.7 | 98.55 | 99.7 | 100 | 100.1 | 100.1 | 100.3 | 100.25 | 100.1714
-0.6 | 99.05 | 99.8 | 100 | 100 | 100.1 | 100.2 | 100.2 | 100.1
-0.5 | 99.5 | 99.8 | 100 | 100 | 100.08 | 100.2 | 100.2 | 100.0125
-0.4 | 99.7 | 99.9 | 100 | 100 | 100 | 100.1 | 100.1667 | 100
-0.3 | 99.85 | 100 | 100 | 100 | 100 | 100.1 | 100.1 | 100
-0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0333 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.8333 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.4 | 99.6 | 100 | 100 | 100 | 100 | 100.1 | 100.1 | 100
0.5 | 99.3 | 99.9 | 100 | 100 | 100 | 100.1 | 100.1 | 100.0125
0.6 | 98.75 | 99.8 | 100 | 100 | 100 | 100.1 | 100.1 | 100.1
0.7 | 98.35 | 99.6 | 100 | 100.1 | 100.05 | 100.18 | 100.2 | 100.1714
0.8 | 97.8 | 99.5 | 100 | 100.1 | 100.1 | 100.2 | 100.2833 | 100.2
0.9 | 96.8 | 99.3 | 100 | 100.1 | 100.1 | 100.3 | 100.35 | 100.3
1 | 95.6667 | 99 | 100 | 100.2 | 100.2 | 100.4 | 100.4 | 100.35
1.1 | 94 | 98.7 | 100 | 100.2 | 100.2 | 100.4 | 100.4667 | 100.4
1.2 | 91.8 | 98.2 | 100 | 100.2 | 100.3 | 100.5 | 100.5 | 100.5
1.3 | 88.8 | 97.8 | 99.9 | 100.2 | 100.4 | 100.5 | 100.6 | 100.6
1.4 | 84 | 97 | 99.9 | 100.3 | 100.4 | 100.6 | 100.6667 | 100.6
1.5 | 77.4 | 96 | 99.8 | 100.3 | 100.4 | 100.6 | 100.7 | 100.7286
1.6 | 68.65 | 94.75 | 99.7 | 100.2 | 100.4 | 100.7 | 100.7833 | 100.8
1.7 | 59.3 | 93.3 | 99.6 | 100.2 | 100.4 | 100.7 | 100.85 | 100.9
1.8 | 49 | 91.0667 | 99.5 | 100.2 | 100.4 | 100.7 | 100.9167 | 100.9
1.9 | 38.3 | 87.8 | 99.3 | 100.2 | 100.5 | 100.8 | 101 | 101
2 | 29.5 | 83.7 | 99.1 | 100.2 | 100.5 | 100.8667 | 101.0333 | 101.1
2.1 | 22.6 | 77.3667 | 98.7 | 100.2 | 100.5 | 100.95 | 101.1 | 101.1
2.2 | 17.5 | 68.7 | 98.35 | 100.1 | 100.5 | 101.05 | 101.2 | 101.2
2.3 | 14.1 | 59.5 | 98.1 | 100.075 | 100.5 | 101.1 | 101.25 | 101.3
2.4 | 11.55 | 49.25 | 97.6 | 100 | 100.5 | 101.18 | 101.3167 | 101.3857
2.5 | 9.6333 | 39 | 97 | 99.9 | 100.5 | 101.225 | 101.4167 | 101.4
2.6 | 8.2 | 30.35 | 96.1 | 99.8 | 100.5 | 101.4 | 101.5 | 101.5
2.7 | 7.2 | 23.4 | 95.15 | 99.8 | 100.5 | 101.4 | 101.6 | 101.525
2.8 | 6.45 | 18.9 | 94.1 | 99.6 | 100.5 | 101.5 | 101.7 | 101.6
2.9 | 5.7 | 15.4 | 92.4 | 99.475 | 100.5 | 101.6 | 101.7667 | 101.7167
3 | 5.1 | 13 | 90.3 | 99.325 | 100.45 | 101.65 | 101.8333 | 101.8714
3.1 | 4.7 | 11.1 | 87.3 | 99.075 | 100.4 | 101.7 | 101.9 | 101.9125
3.2 | 4.25 | 9.7 | 83.4 | 98.9 | 100.35 | 101.8 | 102 | 102
3.3 | 3.9 | 8.5333 | 77.1 | 98.54 | 100.3 | 101.875 | 102.0167 | 102
3.4 | 3.6 | 7.6 | 69.15 | 98.2 | 100.2 | 101.9 | 102.1 | 102.1
3.5 | 3.4 | 6.9 | 60.25 | 97.78 | 100.125 | 101.95 | 102.1429 | 102.2
3.6 | 3.2 | 6.35 | 50.5 | 97.1 | 100.08 | 102 | 102.2 | 102.2
3.7 | 3 | 5.7667 | 40.9667 | 96.5 | 99.9 | 102 | 102.2 | 102.3
3.8 | 2.8 | 5.3 | 32.8 | 95.72 | 99.8 | 102 | 102.3 | 102.3
3.9 | 2.7 | 5 | 25.9333 | 94.675 | 99.6 | 102 | 102.3 | 102.3333
4 | 2.6 | 4.7 | 21.2 | 93.45 | 99.4 | 102 | 102.3 | 102.4
4.1 | 2.4 | 4.4 | 17.6 | 91.8 | 99.1 | 102 | 102.35 | 102.4
4.2 | 2.4 | 4.2 | 15.3 | 89.7667 | 98.9 | 102 | 102.4 | 102.5
4.3 | 2.2333 | 4 | 13.4667 | 87 | 98.4667 | 102 | 102.4 | 102.5
4.4 | 2.2 | 3.8 | 12 | 83 | 98.1667 | 102 | 102.4 | 102.5
4.5 | 2.1 | 3.6 | 10.9 | 77.5 | 97.8 | 102 | 102.4 | 102.6
4.6 | 2 | 3.4 | 9.9 | 69.6 | 97.4667 | 102 | 102.4 | 102.6
4.7 | 2 | 3.3 | 9.1 | 61.35 | 97.1 | 101.9 | 102.4667 | 102.625
4.8 | 1.9 | 3.2 | 8.4667 | 51.7 | 96.4333 | 101.9 | 102.5 | 102.7
4.9 | 1.8 | 3 | 7.9 | 43.1 | 95.775 | 101.9 | 102.5 | 102.7
5 | 1.8 | 2.9 | 7.4 | 34.8 | 94.9 | 101.9 | 102.5 | 102.725
5.1 | 1.7 | 2.8 | 7.05 | 28.1 | 93.9 | 101.9 | 102.5 | 102.8
5.2 | 1.6 | 2.7 | 6.6 | 22.95 | 92.76 | 101.9 | 102.5 | 102.8
5.3 | 1.6 | 2.6 | 6.3 | 19.4 | 91.3 | 101.8 | 102.6 | 102.8
5.4 | 1.6 | 2.5 | 6 | 17 | 89.4 | 101.8 | 102.6 | 102.8625
5.5 | 1.5 | 2.5 | 5.7 | 15.1333 | 86.62 | 101.8 | 102.6 | 102.9
5.6 | 1.5 | 2.4 | 5.5 | 13.66 | 82.82 | 101.8 | 102.6 | 102.9
5.7 | 1.5 | 2.3 | 5.2333 | 12.45 | 77.45 | 101.8 | 102.6 | 102.9
5.8 | 1.4 | 2.2 | 5 | 11.5 | 70.4 | 101.7 | 102.6 | 102.9
5.9 | 1.3 | 2.1 | 4.8 | 10.68 | 62 | 101.7 | 102.6 | 102.9
6 | 1.3 | 2.1 | 4.6 | 10.0333 | 53 | 101.7 | 102.6 | 102.9333
6.1 | 1.3 | 2 | 4.45 | 9.4333 | 44.175 | 101.6 | 102.6 | 103
6.2 | 1.3 | 2 | 4.3 | 8.92 | 36.64 | 101.6 | 102.6 | 103
6.3 | 1.2 | 1.9 | 4.2 | 8.45 | 29.425 | 101.5 | 102.6 | 103
6.4 | 1.2 | 1.9 | 4 | 8.1 | 24.5 | 101.5 | 102.6 | 103
6.5 | 1.2 | 1.8 | 3.9 | 7.7 | 20.95 | 101.4 | 102.6 | 103
6.6 | 1.2 | 1.8 | 3.8 | 7.4 | 18.45 | 101.4 | 102.5 | 103
6.7 | 1.1 | 1.7 | 3.7 | 7.06 | 16.5 | 101.3 | 102.5 | 103
6.8 |  | 1.7 | 3.6 | 6.825 | 15 | 101.2 | 102.5 | 103
6.9 |  | 1.7 | 3.5 | 6.575 | 13.85 | 101.1 | 102.5 | 103
7 |  | 1.6 | 3.4 | 6.3 | 12.8 | 101.02 | 102.4167 | 103
7.1 |  | 1.6 | 3.3 | 6.1 | 11.95 | 100.9 | 102.4 | 103
7.2 |  | 1.5 | 3.2 | 5.925 | 11.3 | 100.8 | 102.4 | 102.9
7.3 |  | 1.5 | 3 | 5.68 | 10.6667 | 100.725 | 102.36 | 102.9
7.4 |  |  | 3 | 5.5 | 10.2 | 100.575 | 102.3 | 102.9
7.5 |  |  | 2.9 | 5.3 | 9.6667 | 100.5 | 102.3 | 102.9
7.6 |  |  | 2.8 | 5.2 | 9.25 | 100.4 | 102.2 | 102.9
7.7 |  |  | 2.8 | 5.075 | 8.9 | 100.3 | 102.2 | 102.9
7.8 |  |  | 2.7 | 4.9 | 8.55 | 100.22 | 102.1 | 102.9
7.9 |  |  | 2.6 | 4.75 | 8.3 | 100.1 | 102.1 | 102.8
8 |  |  | 2.6 | 4.6 | 7.92 | 100.025 | 102.1 | 102.8
8.1 |  |  | 2.5 | 4.525 | 7.65 | 99.975 | 102.02 | 102.8
8.2 |  |  | 2.4 | 4.3667 | 7.4 | 99.725 | 102 | 102.8
8.3 |  |  | 2.4 | 4.225 | 7.2 | 99.6 | 102 | 102.8
8.4 |  |  | 2.3 | 4.1 | 6.9 | 99.5 | 102 | 102.8
8.5 |  |  | 2.3 | 4 | 6.7 | 99.4 | 101.9 | 102.7857
8.6 |  |  |  | 3.9 | 6.52 | 99.22 | 101.9 | 102.7
8.7 |  |  |  | 3.8333 | 6.4 | 99.1 | 101.9 | 102.7
8.8 |  |  |  | 3.7 | 6.15 | 98.925 | 101.8 | 102.7
8.9 |  |  |  | 3.6 | 6.025 | 98.775 | 101.8 | 102.7
9 |  |  |  | 3.5 | 5.86 | 98.6 | 101.8 | 102.7
9.1 |  |  |  | 3.425 | 5.65 | 98.425 | 101.7 | 102.7
9.2 |  |  |  | 3.4 | 5.575 | 98.3 | 101.7 | 102.7
9.3 |  |  |  | 3.3 | 5.4 | 98.15 | 101.6 | 102.625
9.4 |  |  |  | 3.2 | 5.275 | 97.95 | 101.6 | 102.6
9.5 |  |  |  | 3.125 | 5.1 | 97.7 | 101.5 | 102.6
9.6 |  |  |  | 3.075 | 4.95 | 97.475 | 101.4667 | 102.6
9.7 |  |  |  | 3 | 4.8 | 97.32 | 101.4 | 102.6
9.8 |  |  |  |  | 4.76 | 97.075 | 101.4 | 102.6
9.9 |  |  |  |  | 4.625 | 96.875 | 101.3 | 102.6
10 |  |  |  |  | 4.5 | 96.525 | 101.3 | 102.6
10.1 |  |  |  |  | 4.4 | 96.375 | 101.2 | 102.5875
10.2 |  |  |  |  | 4.3 | 95.95 | 101.1 | 102.5
10.3 |  |  |  |  | 4.2 | 95.575 | 101 | 102.5
10.4 |  |  |  |  | 4.1 | 95.2 | 100.9667 | 102.4833
10.5 |  |  |  |  | 4 | 94.7667 | 100.9 | 102.4
10.6 |  |  |  |  | 3.9 | 94.34 | 100.8429 | 102.4
10.7 |  |  |  |  | 3.8 | 93.8 | 100.7 | 102.3
10.8 |  |  |  |  | 3.7 | 93.2 | 100.6167 | 102.2
10.9 | [5 empty] | 92.45 | 100.5571 | 102.2
11 | [5 empty] | 91.5 | 100.5 | 102.2
11.1 | [5 empty] | 90.625 | 100.3333 | 102.15
11.2 | [5 empty] | 89.5 | 100.3 | 102.1
11.3 | [5 empty] | 88.2 | 100.3 | 102.1
11.4 | [5 empty] | 86.425 | 100.2 | 102.1
11.5 | [5 empty] | 84.125 | 100.15 | 102.1
11.6 | [5 empty] | 80.875 | 100.1 | 102.1
11.7 | [5 empty] | 76.125 | 100 | 102.1
11.8 | [5 empty] | 70.02 | 100 | 102.1
11.9 | [5 empty] | 62.5333 | 100 | 102.1
12 | [5 empty] | 55 | 99.95 | 102.1
12.1 | [5 empty] | 46.65 | 99.9 | 102.1
12.2 | [5 empty] | 39 | 99.9 | 102.1
12.3 | [5 empty] | 32.45 | 99.9 | 102.2
12.4 | [5 empty] | 27.8 | 99.8 | 102.1167
12.5 | [5 empty] | 24.05 | 99.7167 | 102.2
12.6 | [5 empty] | 21.6 | 99.7 | 102.2
12.7 | [5 empty] | 19.88 | 99.6 | 102.2
12.8 | [5 empty] | 18.4 | 99.5 | 102.1
12.9 | [5 empty] | 17.2 | 99.5 | 102.1
13 | [5 empty] | 16.225 | 99.3333 | 102
13.1 | [5 empty] | 15.375 | 99.2 | 101.9625
13.2 | [5 empty] | 14.7 | 99.1 | 101.9
13.3 | [5 empty] | 14.075 | 99 | 101.9
13.4 | [5 empty] | 13.6 | 98.86 | 101.8
13.5 | [5 empty] | 13.1 | 98.8 | 101.8
13.6 | [5 empty] | 12.66 | 98.6167 | 101.7429
13.7 | [5 empty] | 12.3 | 98.5 | 101.7
13.8 | [5 empty] | 11.9 | 98.4 | 101.6167
13.9 | [5 empty] | 11.55 | 98.2429 | 101.6
14 | [5 empty] | 11.32 | 98.05 | 101.6
14.1 | [5 empty] | 10.975 | 97.9 | 101.575
14.2 | [5 empty] | 10.725 | 97.7833 | 101.5
14.3 | [5 empty] | 10.475 | 97.7 | 101.5
14.4 | [5 empty] | 10.2 | 97.6 | 101.3857
14.5 | [5 empty] | 9.975 | 97.35 | 101.3
14.6 | [5 empty] | 9.725 | 97.2 | 101.275
14.7 | [5 empty] | 9.5 | 97.0333 | 101.2
14.8 | [5 empty] | 9.24 | 96.8667 | 101.125
14.9 | [5 empty] | 9.1 | 96.7 | 101.075
15 | [5 empty] | 8.9 | 96.5333 | 101
15.1 | [5 empty] | 8.7 | 96.3667 | 100.9
15.2 | [5 empty] | 8.575 | 96.1 | 100.8
15.3 | [5 empty] | 8.3 | 95.8857 | 100.7167
15.4 | [5 empty] | 8.15 | 95.5714 | 100.6571
15.5 | [5 empty] | 8 | 95.4167 | 100.525
15.6 | [5 empty] | 7.8 | 95.25 | 100.5
15.7 | [5 empty] | 7.62 | 94.9667 | 100.3857
15.8 | [5 empty] | 7.48 | 94.7167 | 100.3
15.9 | [5 empty] | 7.35 | 94.44 | 100.2125
16 | [5 empty] | 7.2 | 94.18 | 100.15
16.1 | [5 empty] | 7.05 | 93.8 | 100.0875
16.2 | [5 empty] | 6.9 | 93.5333 | 100
16.3 | [5 empty] | 6.8 | 93.2333 | 100
16.4 | [5 empty] | 6.65 | 92.8 | 99.9
16.5 | [5 empty] | 6.52 | 92.4667 | 99.8167
16.6 | [5 empty] | 6.4 | 92 | 99.775
16.7 | [5 empty] | 6.2167 | 91.5 | 99.7
16.8 | [5 empty] | 6.1429 | 90.8667 | 99.625
16.9 | [5 empty] | 6.04 | 90.28 | 99.5875
17 | [6 empty] | 89.6 | 99.5
17.1 | [6 empty] | 88.7714 | 99.4375
17.2 | [6 empty] | 87.6 | 99.4
17.3 | [6 empty] | 86.425 | 99.3
17.4 | [6 empty] | 84.6833 | 99.25
17.5 | [6 empty] | 82.35 | 99.125
17.6 | [6 empty] | 79.42 | 99.05
17.7 | [6 empty] | 74.9857 | 98.9857
17.8 | [6 empty] | 69.02 | 98.9
17.9 | [6 empty] | 62.08 | 98.9
18 | [6 empty] | 54.5 | 98.7667
18.1 | [6 empty] | 46.9333 | 98.6
18.2 | [6 empty] | 39.6333 | 98.6
18.3 | [6 empty] | 33.2 | 98.4625
18.4 | [6 empty] | 28.8714 | 98.3375
18.5 | [6 empty] | 25.3 | 98.2125
18.6 | [6 empty] | 23.1 | 98.0875
18.7 | [6 empty] | 21.2 | 97.9667
18.8 | [6 empty] | 19.7833 | 97.85
18.9 | [6 empty] | 18.6167 | 97.65
19 | [6 empty] | 17.7 | 97.4857
19.1 | [6 empty] | 17 | 97.4
19.2 | [6 empty] | 16.2714 | 97.325
19.3 | [6 empty] | 15.7 | 97.1714
19.4 | [6 empty] | 15.2 | 96.8857
19.5 | [6 empty] | 14.8 | 96.7125
19.6 | [6 empty] | 14.3 | 96.5875
19.7 | [6 empty] | 13.9667 | 96.44
19.8 | [6 empty] | 13.6429 | 96.2833
19.9 | [6 empty] | 13.3333 | 96.1286
20 | [6 empty] | 13 | 95.975
20.1 | [6 empty] | 12.6667 | 95.8
20.2 | [6 empty] | 12.4167 | 95.6
20.3 | [6 empty] | 12.1667 | 95.475
20.4 | [6 empty] | 11.9 | 95.3333
20.5 | [6 empty] | 11.72 | 95.2
20.6 | [6 empty] | 11.4667 | 95.0778
20.7 | [6 empty] | 11.2667 | 94.9571
20.8 | [6 empty] | 11 | 94.8143
20.9 | [6 empty] | 10.7667 | 94.5875
21 | [6 empty] | 10.6 | 94.4625
21.1 | [6 empty] | 10.36 | 94.3375
21.2 | [6 empty] | 10.18 | 94.2222
21.3 | [6 empty] | 9.9833 | 94.1
21.4 | [6 empty] | 9.8167 | 93.95
21.5 | [6 empty] | 9.64 | 93.75
21.6 | [6 empty] | 9.45 | 93.55
21.7 | [6 empty] | 9.2833 | 93.3
21.8 | [6 empty] | 9.1 | 93.05
21.9 | [6 empty] | 8.9333 | 92.8111
22 | [6 empty] | 8.7667 | 92.5125
22.1 | [6 empty] | 8.66 | 92.225
22.2 | [6 empty] | 8.5 | 91.8625
22.3 | [6 empty] | 8.3 | 91.4714
22.4 | [6 empty] | 8.1333 | 91.0667
22.5 | [6 empty] | 8.025 | 90.6143
22.6 | [6 empty] | 7.84 | 90.1556
22.7 | [6 empty] | 7.7833 | 89.675
22.8 | [6 empty] | 7.6167 | 89.1286
22.9 | [6 empty] | 7.45 | 88.5167
23 | [7 empty] | 87.7667
23.1 | [7 empty] | 87
23.2 | [7 empty] | 85.925
23.3 | [7 empty] | 84.7833
23.4 | [7 empty] | 83.2714
23.5 | [7 empty] | 81.2143
23.6 | [7 empty] | 78.3125
23.7 | [7 empty] | 74.325
23.8 | [7 empty] | 68.8875
23.9 | [7 empty] | 62.3875
24 | [7 empty] | 55.3111
24.1 | [7 empty] | 48.25
24.2 | [7 empty] | 41
24.3 | [7 empty] | 35.05
24.4 | [7 empty] | 30.2
24.5 | [7 empty] | 26.55
24.6 | [7 empty] | 24.1222
24.7 | [7 empty] | 22.25
24.8 | [7 empty] | 20.825
24.9 | [7 empty] | 19.6375
25 | [7 empty] | 18.6875
25.1 | [7 empty] | 17.925
25.2 | [7 empty] | 17.2333
25.3 | [7 empty] | 16.6714
25.4 | [7 empty] | 16.15
25.5 | [7 empty] | 15.7143
25.6 | [7 empty] | 15.25
25.7 | [7 empty] | 14.8667
25.8 | [7 empty] | 14.6
25.9 | [7 empty] | 14.2333
26 | [7 empty] | 13.925
26.1 | [7 empty] | 13.6125
26.2 | [7 empty] | 13.325
26.3 | [7 empty] | 13.075
26.4 | [7 empty] | 12.825
26.5 | [7 empty] | 12.6
26.6 | [7 empty] | 12.3
26.7 | [7 empty] | 12.12
26.8 | [7 empty] | 11.9333
26.9 | [7 empty] | 11.6667
27 | [7 empty] | 11.4
27.1 | [7 empty] | 11.22
27.2 | [7 empty] | 11.04
27.3 | [7 empty] | 10.8429
27.4 | [7 empty] | 10.6286
27.5 | [7 empty] | 10.4429
27.6 | [7 empty] | 10.2286
27.7 | [7 empty] | 10.05
27.8 | [7 empty] | 9.9125
27.9 | [7 empty] | 9.7714
28 | [7 empty] | 9.5429
28.1 | [7 empty] | 9.4
28.2 | [7 empty] | 9.25
28.3 | [7 empty] | 9.125
28.4 | [7 empty] | 8.9
28.5 | [7 empty] | 8.775
28.6 | [7 empty] | 8.6429
28.7 | [7 empty] | 8.5
28.8 | [7 empty] | 8.3333
28.9 | [7 empty] | 8.2111

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-30.9 | [7 empty] | 10.4
-30.8 | [7 empty] | 10.48
-30.7 | [7 empty] | 10.6333
-30.6 | [7 empty] | 10.8167
-30.5 | [7 empty] | 10.9833
-30.4 | [7 empty] | 11.1333
-30.3 | [7 empty] | 11.32
-30.2 | [7 empty] | 11.4857
-30.1 | [7 empty] | 11.65
-30 | [7 empty] | 11.8
-29.9 | [7 empty] | 11.98
-29.8 | [7 empty] | 12.15
-29.7 | [7 empty] | 12.4
-29.6 | [7 empty] | 12.5571
-29.5 | [7 empty] | 12.7167
-29.4 | [7 empty] | 12.9667
-29.3 | [7 empty] | 13.16
-29.2 | [7 empty] | 13.3333
-29.1 | [7 empty] | 13.475
-29 | [7 empty] | 13.7333
-28.9 | [7 empty] | 13.95
-28.8 | [7 empty] | 14.1667
-28.7 | [7 empty] | 14.375
-28.6 | [7 empty] | 14.6
-28.5 | [7 empty] | 14.8333
-28.4 | [7 empty] | 15.1
-28.3 | [7 empty] | 15.3667
-28.2 | [7 empty] | 15.55
-28.1 | [7 empty] | 15.8667
-28 | [7 empty] | 16.2333
-27.9 | [7 empty] | 16.5
-27.8 | [7 empty] | 16.9
-27.7 | [7 empty] | 17.16
-27.6 | [7 empty] | 17.5
-27.5 | [7 empty] | 17.9
-27.4 | [7 empty] | 18.4
-27.3 | [7 empty] | 18.8333
-27.2 | [7 empty] | 19.4
-27.1 | [7 empty] | 20.0857
-27 | [7 empty] | 20.9
-26.9 | [7 empty] | 21.7625
-26.8 | [7 empty] | 22.825
-26.7 | [7 empty] | 24.25
-26.6 | [7 empty] | 25.95
-26.5 | [7 empty] | 28.38
-26.4 | [7 empty] | 31.78
-26.3 | [7 empty] | 36
-26.2 | [7 empty] | 40.9667
-26.1 | [7 empty] | 46.9857
-26 | [7 empty] | 53.5625
-25.9 | [7 empty] | 59.625
-25.8 | [7 empty] | 65.35
-25.7 | [7 empty] | 70.0625
-25.6 | [7 empty] | 73.4143
-25.5 | [7 empty] | 75.925
-25.4 | [7 empty] | 77.9
-25.3 | [7 empty] | 79.3667
-25.2 | [7 empty] | 80.6
-25.1 | [7 empty] | 81.48
-25 | [7 empty] | 82.3
-24.9 | [7 empty] | 83.05
-24.8 | [7 empty] | 83.7286
-24.7 | [7 empty] | 84.25
-24.6 | [7 empty] | 84.6375
-24.5 | [7 empty] | 85.05
-24.4 | [6 empty] | 9.7833 | 85.55
-24.3 | [6 empty] | 9.86 | 86.05
-24.2 | [6 empty] | 10 | 86.3333
-24.1 | [6 empty] | 10.2 | 86.675
-24 | [6 empty] | 10.4 | 87.05
-23.9 | [6 empty] | 10.5 | 87.35
-23.8 | [6 empty] | 10.6 | 87.7
-23.7 | [6 empty] | 10.9 | 87.975
-23.6 | [6 empty] | 10.9571 | 88.2889
-23.5 | [6 empty] | 11.12 | 88.4625
-23.4 | [6 empty] | 11.3167 | 88.7625
-23.3 | [6 empty] | 11.5 | 88.925
-23.2 | [6 empty] | 11.7 | 89.1375
-23.1 | [6 empty] | 11.9143 | 89.325
-23 | [6 empty] | 12.08 | 89.575
-22.9 | [6 empty] | 12.2667 | 89.7125
-22.8 | [6 empty] | 12.5 | 89.8375
-22.7 | [6 empty] | 12.6667 | 90.025
-22.6 | [6 empty] | 12.8333 | 90.2
-22.5 | [6 empty] | 13.1 | 90.35
-22.4 | [6 empty] | 13.3667 | 90.55
-22.3 | [6 empty] | 13.54 | 90.75
-22.2 | [6 empty] | 13.7667 | 90.9
-22.1 | [6 empty] | 14 | 91.1444
-22 | [6 empty] | 14.2 | 91.2625
-21.9 | [6 empty] | 14.4833 | 91.3875
-21.8 | [6 empty] | 14.82 | 91.6125
-21.7 | [6 empty] | 15.04 | 91.775
-21.6 | [6 empty] | 15.32 | 91.9625
-21.5 | [6 empty] | 15.575 | 92.175
-21.4 | [6 empty] | 15.88 | 92.4125
-21.3 | [6 empty] | 16.24 | 92.5375
-21.2 | [6 empty] | 16.64 | 92.7111
-21.1 | [6 empty] | 16.98 | 92.95
-21 | [6 empty] | 17.48 | 93.1
-20.9 | [6 empty] | 17.9 | 93.1857
-20.8 | [6 empty] | 18.48 | 93.375
-20.7 | [6 empty] | 19.14 | 93.5778
-20.6 | [6 empty] | 19.78 | 93.7286
-20.5 | [6 empty] | 20.6 | 93.9571
-20.4 | [6 empty] | 21.6 | 94.0875
-20.3 | [6 empty] | 22.8333 | 94.25
-20.2 | [6 empty] | 24.325 | 94.32
-20.1 | [6 empty] | 26.25 | 94.5
-20 | [6 empty] | 28.75 | 94.625
-19.9 | [6 empty] | 32.3 | 94.7571
-19.8 | [6 empty] | 36.875 | 94.9143
-19.7 | [6 empty] | 42.925 | 95.0571
-19.6 | [6 empty] | 49.3 | 95.225
-19.5 | [6 empty] | 56.48 | 95.3
-19.4 | [6 empty] | 62.45 | 95.45
-19.3 | [6 empty] | 68.425 | 95.6167
-19.2 | [6 empty] | 73.35 | 95.7
-19.1 | [6 empty] | 77 | 95.8
-19 | [6 empty] | 79.46 | 95.9
-18.9 | [6 empty] | 81.32 | 96.075
-18.8 | [6 empty] | 82.8 | 96.2
-18.7 | [6 empty] | 84.1 | 96.2778
-18.6 | [6 empty] | 85.0167 | 96.4
-18.5 | [6 empty] | 85.7333 | 96.45
-18.4 | [6 empty] | 86.5 | 96.6
-18.3 | [6 empty] | 87.1 | 96.6667
-18.2 | [6 empty] | 87.65 | 96.7286
-18.1 | [6 empty] | 88.22 | 96.8714
-18 | [5 empty] | 7.8167 | 88.7 | 97
-17.9 | [5 empty] | 7.9714 | 89.08 | 97.025
-17.8 | [5 empty] | 8.1 | 89.3667 | 97.1444
-17.7 | [5 empty] | 8.2 | 89.7667 | 97.2714
-17.6 | [5 empty] | 8.3 | 90.16 | 97.3
-17.5 | [5 empty] | 8.5 | 90.5 | 97.4125
-17.4 | [5 empty] | 8.6 | 90.725 | 97.5375
-17.3 | [5 empty] | 8.8 | 91.06 | 97.6
-17.2 | [5 empty] | 8.92 | 91.28 | 97.6875
-17.1 | [5 empty] | 9.1 | 91.6 | 97.8125
-17 | [5 empty] | 9.3 | 91.86 | 97.9429
-16.9 | [5 empty] | 9.45 | 92 | 98
-16.8 | [5 empty] | 9.6 | 92.34 | 98.1
-16.7 | [5 empty] | 9.82 | 92.6 | 98.2222
-16.6 | [5 empty] | 9.925 | 92.78 | 98.3375
-16.5 | [5 empty] | 10.2 | 93.125 | 98.4714
-16.4 | [5 empty] | 10.325 | 93.24 | 98.6125
-16.3 | [5 empty] | 10.5 | 93.4 | 98.7
-16.2 | [5 empty] | 10.775 | 93.6 | 98.7625
-16.1 | [5 empty] | 10.925 | 93.86 | 98.8
-16 | [5 empty] | 11.16 | 94 | 98.9571
-15.9 | [5 empty] | 11.42 | 94.225 | 99.1
-15.8 | [5 empty] | 11.6 | 94.475 | 99.1714
-15.7 | [5 empty] | 11.8 | 94.7 | 99.26
-15.6 | [5 empty] | 12.05 | 94.8 | 99.3
-15.5 | [5 empty] | 12.3 | 94.9333 | 99.35
-15.4 | [5 empty] | 12.55 | 95.2 | 99.5
-15.3 | [5 empty] | 12.8 | 95.3667 | 99.6
-15.2 | [5 empty] | 13.125 | 95.54 | 99.6333
-15.1 | [5 empty] | 13.375 | 95.8167 | 99.7
-15 | [5 empty] | 13.65 | 95.9 | 99.7571
-14.9 | [5 empty] | 13.975 | 96.05 | 99.8
-14.8 | [5 empty] | 14.35 | 96.2 | 99.9
-14.7 | [5 empty] | 14.675 | 96.3667 | 99.9875
-14.6 | [5 empty] | 15.1 | 96.5333 | 100.1
-14.5 | [5 empty] | 15.52 | 96.7 | 100.1375
-14.4 | [5 empty] | 16.04 | 96.8333 | 100.2
-14.3 | [5 empty] | 16.55 | 97 | 100.225
-14.2 | [5 empty] | 17.1 | 97.1 | 100.3
-14.1 | [5 empty] | 17.85 | 97.1833 | 100.3
-14 | [5 empty] | 18.7 | 97.375 | 100.4
-13.9 | [5 empty] | 19.7 | 97.48 | 100.4429
-13.8 | [5 empty] | 20.75 | 97.5571 | 100.5
-13.7 | [5 empty] | 22.175 | 97.6 | 100.525
-13.6 | [5 empty] | 24.075 | 97.6833 | 100.6
-13.5 | [5 empty] | 26.425 | 97.7 | 100.7
-13.4 | [5 empty] | 29.6333 | 97.8167 | 100.7
-13.3 | [5 empty] | 33.9 | 97.9 | 100.7
-13.2 | [5 empty] | 39.84 | 98.08 | 100.7
-13.1 | [5 empty] | 46.8 | 98.1333 | 100.7
-13 | [5 empty] | 54.4333 | 98.2 | 100.7
-12.9 | [5 empty] | 61.65 | 98.3 | 100.7
-12.8 | [5 empty] | 68.8 | 98.3 | 100.7
-12.7 | [5 empty] | 74.25 | 98.4 | 100.7
-12.6 | [5 empty] | 78.6 | 98.4167 | 100.7
-12.5 | [5 empty] | 81.35 | 98.5833 | 100.75
-12.4 | [5 empty] | 83.7 | 98.6 | 100.8
-12.3 | [5 empty] | 85.475 | 98.7 | 100.8
-12.2 | [5 empty] | 87.05 | 98.7 | 100.8
-12.1 | [5 empty] | 88.0667 | 98.75 | 100.8
-12 | [5 empty] | 88.95 | 98.85 | 100.8571
-11.9 | [5 empty] | 89.86 | 98.95 | 100.9
-11.8 | [5 empty] | 90.6 | 99.1 | 100.9429
-11.7 | [5 empty] | 91.1 | 99.2 | 101
-11.6 | [5 empty] | 91.7 | 99.3 | 101.0125
-11.5 |  |  |  |  | 4.9286 | 92.2 | 99.4 | 101.1
-11.4 |  |  |  |  | 5.1 | 92.6 | 99.5 | 101.15
-11.3 |  |  |  |  | 5.2 | 93 | 99.5167 | 101.2
-11.2 |  |  |  |  | 5.3 | 93.425 | 99.6 | 101.2
-11.1 |  |  |  |  | 5.4 | 93.7 | 99.7 | 101.3
-11 |  |  |  |  | 5.5 | 94.14 | 99.8714 | 101.3
-10.9 |  |  |  |  | 5.675 | 94.375 | 100 | 101.3143
-10.8 |  |  |  |  | 5.725 | 94.78 | 100 | 101.4
-10.7 |  |  |  |  | 5.875 | 94.975 | 100.1 | 101.4
-10.6 |  |  |  |  | 6 | 95.225 | 100.1 | 101.5
-10.5 |  |  |  |  | 6.1 | 95.4667 | 100.2 | 101.5375
-10.4 |  |  |  |  | 6.2 | 95.7 | 100.225 | 101.6
-10.3 |  |  |  | 3.9 | 6.4 | 95.975 | 100.3167 | 101.6
-10.2 |  |  |  | 4 | 6.5 | 96.2 | 100.4 | 101.7
-10.1 |  |  |  | 4.1 | 6.6 | 96.3 | 100.5 | 101.7
-10 |  |  |  | 4.2 | 6.8 | 96.4333 | 100.5 | 101.7
-9.9 |  |  |  | 4.3 | 6.9 | 96.7 | 100.6 | 101.7
-9.8 |  |  |  | 4.4 | 7.14 | 96.8 | 100.6 | 101.7
-9.7 |  |  |  | 4.5 | 7.275 | 97.075 | 100.65 | 101.8
-9.6 |  |  |  | 4.6 | 7.4 | 97.225 | 100.72 | 101.8
-9.5 |  |  |  | 4.7 | 7.6 | 97.36 | 100.8 | 101.8
-9.4 |  |  |  | 4.75 | 7.8 | 97.5667 | 100.9 | 101.8
-9.3 |  |  |  | 4.88 | 7.9 | 97.7 | 100.9 | 101.8143
-9.2 |  |  |  | 5 | 8.125 | 97.8 | 100.9 | 101.9
-9.1 |  |  |  | 5.1 | 8.46 | 98.04 | 101 | 101.9
-9 |  |  |  | 5.25 | 8.7 | 98.2 | 101 | 101.9
-8.9 |  |  |  | 5.4 | 8.825 | 98.325 | 101.0333 | 101.9375
-8.8 |  |  | 3.1 | 5.45 | 9.15 | 98.475 | 101.1 | 102
-8.7 |  |  | 3.2 | 5.525 | 9.32 | 98.6 | 101.175 | 102
-8.6 |  |  | 3.3 | 5.7 | 9.55 | 98.7 | 101.2 | 102
-8.5 |  |  | 3.3 | 5.825 | 9.9 | 98.9 | 101.3 | 102.0333
-8.4 |  |  | 3.4 | 6.05 | 10.2 | 99 | 101.4 | 102.1
-8.3 |  |  | 3.5 | 6.2 | 10.52 | 99.1 | 101.425 | 102.1167
-8.2 |  |  | 3.6 | 6.34 | 10.875 | 99.2 | 101.5 | 102.2
-8.1 |  |  | 3.6 | 6.48 | 11.24 | 99.3333 | 101.5 | 102.2
-8 |  |  | 3.7 | 6.7 | 11.7 | 99.5 | 101.5 | 102.2
-7.9 |  |  | 3.8 | 6.825 | 12.175 | 99.6 | 101.54 | 102.2667
-7.8 |  |  | 3.9 | 7 | 12.74 | 99.7 | 101.6 | 102.3
-7.7 |  |  | 4 | 7.3 | 13.3 | 99.88 | 101.6 | 102.3
-7.6 |  |  | 4.1 | 7.52 | 14 | 99.925 | 101.7 | 102.3
-7.5 |  | 2.1 | 4.25 | 7.72 | 14.9 | 100.075 | 101.78 | 102.3
-7.4 |  | 2.1 | 4.35 | 8 | 15.9 | 100.1 | 101.8 | 102.3
-7.3 |  | 2.1 | 4.5 | 8.14 | 17.05 | 100.2 | 101.8333 | 102.4
-7.2 |  | 2.2 | 4.6 | 8.38 | 18.5 | 100.3 | 101.9 | 102.4
-7.1 |  | 2.2 | 4.6667 | 8.64 | 20.6 | 100.4 | 101.9 | 102.4
-7 |  | 2.3 | 4.8 | 9.02 | 23.1 | 100.5 | 101.9 | 102.4
-6.9 |  | 2.4 | 4.9 | 9.4 | 26.3 | 100.525 | 101.9833 | 102.4
-6.8 | 1.5 | 2.4 | 5 | 9.75 | 31 | 100.68 | 102 | 102.4
-6.7 | 1.5667 | 2.5 | 5.1667 | 10.1 | 37.375 | 100.7 | 102 | 102.4
-6.6 | 1.6 | 2.6 | 5.4 | 10.7 | 45 | 100.8 | 102.08 | 102.4
-6.5 | 1.6 | 2.6 | 5.5 | 11.14 | 53.8 | 100.8 | 102.1 | 102.5
-6.4 | 1.7 | 2.6 | 5.7 | 11.82 | 62.275 | 100.9 | 102.1 | 102.5
-6.3 | 1.7 | 2.7 | 5.95 | 12.44 | 70.1 | 100.9 | 102.1 | 102.5
-6.2 | 1.8 | 2.8 | 6.1 | 13.3 | 76.62 | 101 | 102.1 | 102.5
-6.1 | 1.8 | 2.9 | 6.3 | 14.3 | 81.7 | 101.04 | 102.1 | 102.4125
-6 | 1.8 | 2.9 | 6.5 | 15.525 | 85.44 | 101.18 | 102.1 | 102.4
-5.9 | 1.9 | 3 | 6.7 | 16.9 | 88.2 | 101.225 | 102.1 | 102.4
-5.8 | 1.9 | 3.1 | 7 | 18.94 | 90.15 | 101.3 | 102.2 | 102.3286
-5.7 | 2 | 3.2 | 7.2 | 21.3 | 91.8 | 101.3 | 102.2 | 102.3
-5.6 | 2 | 3.3 | 7.6 | 24.8 | 93 | 101.3 | 102.1167 | 102.3
-5.5 | 2.1 | 3.4 | 7.9 | 29.4 | 93.92 | 101.3 | 102.15 | 102.3
-5.4 | 2.1 | 3.5 | 8.35 | 35.85 | 94.7 | 101.3 | 102.1 | 102.3
-5.3 | 2.2 | 3.6 | 8.7 | 44 | 95.35 | 101.4 | 102.1 | 102.3
-5.2 | 2.2667 | 3.7 | 9.3 | 52.32 | 95.9 | 101.4 | 102.1 | 102.3
-5.1 | 2.3 | 3.8 | 9.85 | 61.7 | 96.55 | 101.4 | 102.1 | 102.3
-5 | 2.4 | 4.05 | 10.55 | 69.45 | 96.9 | 101.4 | 102.1 | 102.3
-4.9 | 2.4 | 4.1 | 11.35 | 76.58 | 97.3 | 101.45 | 102.1 | 102.3
-4.8 | 2.5 | 4.35 | 12.3 | 82.25 | 97.75 | 101.5 | 102.1 | 102.3
-4.7 | 2.6 | 4.5 | 13.6 | 86.25 | 98.1 | 101.5 | 102.1 | 102.2
-4.6 | 2.7 | 4.6667 | 15.1 | 88.925 | 98.325 | 101.5 | 102.05 | 102.2
-4.5 | 2.7 | 4.8 | 16.95 | 90.925 | 98.58 | 101.5 | 102 | 102.2
-4.4 | 2.8 | 5.1 | 19.3 | 92.5333 | 98.825 | 101.6 | 102 | 102.2
-4.3 | 2.9 | 5.3 | 22.7 | 93.85 | 99.1 | 101.6 | 102 | 102.2
-4.2 | 3.1 | 5.6 | 27.3 | 95 | 99.3 | 101.6 | 102 | 102.1714
-4.1 | 3.3 | 5.9 | 33.8333 | 95.72 | 99.45 | 101.7 | 101.9857 | 102.1
-4 | 3.45 | 6.3667 | 41.9 | 96.5 | 99.7 | 101.7 | 101.9 | 102.075
-3.9 | 3.6 | 6.85 | 50.8 | 97.075 | 99.9 | 101.7 | 101.9 | 102
-3.8 | 3.8 | 7.3 | 60.3 | 97.575 | 100 | 101.7 | 101.9 | 102
-3.7 | 3.9 | 7.9667 | 69.1 | 98.05 | 100.075 | 101.7 | 101.9 | 102
-3.6 | 4.2 | 8.8 | 76.2 | 98.42 | 100.125 | 101.65 | 101.8 | 102
-3.5 | 4.4 | 9.7 | 82.1 | 98.8 | 100.2 | 101.6 | 101.8 | 101.9
-3.4 | 4.8 | 10.8 | 86.45 | 98.95 | 100.4 | 101.6 | 101.8 | 101.8
-3.3 | 5.2 | 12.3 | 89.55 | 99.18 | 100.5 | 101.6 | 101.7 | 101.7286
-3.2 | 5.6 | 14.2 | 91.7 | 99.3 | 100.5 | 101.5 | 101.7 | 101.6
-3.1 | 6.05 | 16.5 | 93.4 | 99.475 | 100.6 | 101.5 | 101.5714 | 101.6
-3 | 6.8 | 20 | 94.6 | 99.625 | 100.5 | 101.4 | 101.5 | 101.5375
-2.9 | 7.6 | 24.7 | 95.7 | 99.75 | 100.6 | 101.4 | 101.4 | 101.5
-2.8 | 8.7 | 31.25 | 96.4333 | 99.9 | 100.6 | 101.4 | 101.3833 | 101.4
-2.7 | 9.95 | 39.65 | 97.1 | 99.9 | 100.6 | 101.3 | 101.3 | 101.35
-2.6 | 11.6 | 49.2 | 97.55 | 99.9 | 100.6 | 101.225 | 101.2 | 101.2143
-2.5 | 13.7333 | 59 | 98.05 | 99.98 | 100.6 | 101.175 | 101.1 | 101.15
-2.4 | 16.5 | 67.8 | 98.4 | 100 | 100.5 | 101.1 | 101.0333 | 101.1
-2.3 | 20.5 | 76 | 98.6 | 100.125 | 100.5 | 101 | 101 | 101
-2.2 | 26.7 | 82.6 | 98.9 | 100.2 | 100.5 | 101 | 100.9 | 100.9333
-2.1 | 34.3 | 87.1 | 99.1 | 100.2 | 100.5 | 100.9 | 100.9 | 100.9
-2 | 43.4667 | 90.1 | 99.3 | 100.2 | 100.5 | 100.8 | 100.85 | 100.8
-1.9 | 53.2 | 92.35 | 99.4 | 100.2 | 100.5 | 100.8 | 100.7667 | 100.8
-1.8 | 63.9 | 94.2 | 99.6 | 100.2 | 100.5 | 100.7 | 100.7 | 100.7857
-1.7 | 73 | 95.55 | 99.7 | 100.3 | 100.5 | 100.7 | 100.7 | 100.7
-1.6 | 79.9 | 96.4 | 99.7 | 100.3 | 100.4 | 100.625 | 100.625 | 100.64
-1.5 | 85.4 | 97.3 | 99.8 | 100.3 | 100.4 | 100.575 | 100.6 | 100.6
-1.4 | 89.5 | 97.9 | 99.9 | 100.3 | 100.4 | 100.5 | 100.5 | 100.6
-1.3 | 92.1 | 98.4 | 100 | 100.3 | 100.4 | 100.5 | 100.46 | 100.5
-1.2 | 94.2 | 98.8 | 100 | 100.3 | 100.4 | 100.5 | 100.4 | 100.4
-1.1 | 95.7 | 99.1 | 100 | 100.3 | 100.4 | 100.4 | 100.36 | 100.3
-1 | 96.7 | 99.3 | 100 | 100.25 | 100.35 | 100.35 | 100.2 | 100.25
-0.9 | 97.6 | 99.5 | 100 | 100.2 | 100.3 | 100.3 | 100.2 | 100.2
-0.8 | 98.3 | 99.7 | 99.9 | 100.15 | 100.2 | 100.275 | 100.2 | 100.1667
-0.7 | 98.8 | 99.7 | 100 | 100.1 | 100.2 | 100.2 | 100.1 | 100.1
-0.6 | 99.2 | 99.8 | 100 | 100.1 | 100.16 | 100.1 | 100.1 | 100.0875
-0.5 | 99.5 | 99.9 | 100 | 100.1 | 100.1 | 100.1 | 100 | 100
-0.4 | 99.7 | 99.9 | 100 | 100.1 | 100.1 | 100.1 | 100 | 100
-0.3 | 99.8 | 99.9 | 100 | 100.1 | 100.1 | 100.05 | 100 | 100
-0.2 | 99.9 | 99.95 | 100 | 100 | 100.1 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.8 | 100 | 99.9 | 100.1 | 100 | 100 | 100 | 100
0.4 | 99.7 | 100 | 99.9 | 100 | 100 | 100.1 | 100 | 100
0.5 | 99.5 | 100 | 99.9 | 100 | 100.1 | 100.05 | 100 | 100
0.6 | 99.1667 | 100 | 100 | 100 | 100.1 | 100.1 | 100.1 | 100.0875
0.7 | 98.7 | 99.8667 | 100 | 100 | 100 | 100.1 | 100.1 | 100.1
0.8 | 98.2 | 99.7 | 100 | 100.1 | 100.1 | 100.175 | 100.2 | 100.1667
0.9 | 97.4 | 99.6 | 100 | 100.1 | 100.1 | 100.2 | 100.2 | 100.2
1 | 96.6 | 99.4 | 100 | 100.2 | 100.1 | 100.2 | 100.2286 | 100.25
1.1 | 95.6 | 99.15 | 100 | 100.2 | 100.1 | 100.2 | 100.3 | 100.3
1.2 | 94.2 | 98.8 | 100 | 100.2 | 100.2 | 100.3 | 100.4 | 100.4
1.3 | 92.2333 | 98.4333 | 100 | 100.26 | 100.25 | 100.3 | 100.4833 | 100.5
1.4 | 89.6 | 97.9 | 100 | 100.2 | 100.3 | 100.3 | 100.5 | 100.6
1.5 | 85.4 | 97.4 | 99.9 | 100.2 | 100.3 | 100.4 | 100.6 | 100.6
1.6 | 79.7 | 96.5 | 99.8 | 100.2 | 100.3 | 100.4 | 100.6 | 100.64
1.7 | 72.9 | 95.5667 | 99.7 | 100.2 | 100.4 | 100.5 | 100.7 | 100.7
1.8 | 64.1 | 94.3 | 99.6 | 100.2 | 100.4 | 100.6 | 100.7 | 100.7857
1.9 | 53.3 | 92.55 | 99.4 | 100.1 | 100.4 | 100.7 | 100.8 | 100.8
2 | 43.75 | 90.25 | 99.3 | 100.1 | 100.4 | 100.75 | 100.88 | 100.8
2.1 | 34.25 | 87.2 | 99.1 | 100 | 100.4 | 100.8 | 100.9 | 100.9
2.2 | 26.6 | 82.7 | 98.9 | 100 | 100.4 | 100.8 | 100.9 | 100.9333
2.3 | 20.5667 | 76.1 | 98.6 | 100 | 100.375 | 100.8 | 101 | 101
2.4 | 16.5 | 68.3 | 98.3 | 99.9 | 100.3 | 100.825 | 101.0667 | 101.1
2.5 | 13.5667 | 59.2 | 98 | 99.825 | 100.3 | 100.975 | 101.1333 | 101.15
2.6 | 11.5 | 49.2 | 97.55 | 99.7 | 100.3 | 101 | 101.2 | 101.2143
2.7 | 9.95 | 39.6 | 97.1 | 99.7 | 100.3 | 101 | 101.3 | 101.35
2.8 | 8.5333 | 31.2 | 96.4333 | 99.6 | 100.3 | 101.1 | 101.3167 | 101.4
2.9 | 7.6 | 24.75 | 95.7 | 99.56 | 100.3 | 101.2 | 101.4 | 101.5
3 | 6.7 | 20.2 | 94.7 | 99.5 | 100.3 | 101.2 | 101.44 | 101.5375
3.1 | 6.15 | 16.75 | 93.3 | 99.375 | 100.2 | 101.3 | 101.5 | 101.6
3.2 | 5.6 | 14.1333 | 91.7 | 99.2 | 100.175 | 101.375 | 101.7 | 101.6
3.3 | 5.2 | 12.3667 | 89.5 | 99.1 | 100.1 | 101.4 | 101.7 | 101.7286
3.4 | 4.8 | 10.9 | 86.4 | 98.8 | 100.1 | 101.4 | 101.8 | 101.8
3.5 | 4.5 | 9.7 | 82.1 | 98.5 | 100 | 101.4 | 101.8 | 101.9
3.6 | 4.2 | 8.9 | 76.25 | 98.32 | 99.9 | 101.45 | 101.825 | 102
3.7 | 3.8667 | 8.0333 | 68.9667 | 97.9 | 99.86 | 101.5 | 101.9 | 102
3.8 | 3.7 | 7.4 | 60.3 | 97.45 | 99.7 | 101.5 | 101.9 | 102
3.9 | 3.5 | 6.75 | 50.5333 | 96.975 | 99.575 | 101.5 | 101.9 | 102
4 | 3.35 | 6.35 | 41.9 | 96.38 | 99.5 | 101.5 | 101.9 | 102.075
4.1 | 3.2 | 5.9 | 33.9 | 95.62 | 99.4 | 101.5 | 102 | 102.1
4.2 | 3.1 | 5.6 | 27.4 | 94.775 | 99.2 | 101.6 | 102 | 102.1714
4.3 | 3 | 5.2667 | 22.7 | 93.7333 | 99.1 | 101.5 | 102 | 102.2
4.4 | 2.9 | 5 | 19.1333 | 92.54 | 98.85 | 101.5 | 102 | 102.2
4.5 | 2.8 | 4.9 | 16.9 | 90.9 | 98.5 | 101.5 | 102 | 102.2
4.6 | 2.7 | 4.6 | 15.1 | 88.9 | 98.26 | 101.5 | 102.0833 | 102.2
4.7 | 2.6 | 4.4 | 13.5667 | 86 | 97.85 | 101.475 | 102.1 | 102.2
4.8 | 2.5 | 4.3 | 12.3 | 81.95 | 97.6 | 101.4 | 102.1 | 102.3
4.9 | 2.4 | 4.15 | 11.35 | 76.5 | 97.3 | 101.4 | 102.1 | 102.3
5 | 2.3 | 4 | 10.45 | 69.25 | 96.7333 | 101.34 | 102.1 | 102.3
5.1 | 2.3 | 3.9 | 9.85 | 61.625 | 96.2 | 101.3 | 102.1 | 102.3
5.2 | 2.2 | 3.7 | 9.2 | 52.525 | 95.85 | 101.3 | 102.1 | 102.3
5.3 | 2.2 | 3.6 | 8.7 | 43.95 | 95.3333 | 101.2 | 102.1 | 102.3
5.4 | 2.1 | 3.5 | 8.25 | 36.125 | 94.6 | 101.2 | 102.1 | 102.3
5.5 | 2 | 3.4 | 7.9 | 29.525 | 93.775 | 101.2 | 102.1833 | 102.3
5.6 | 2 | 3.2 | 7.5667 | 25 | 92.7 | 101.2 | 102.15 | 102.3
5.7 | 1.9 | 3.15 | 7.3 | 21.3667 | 91.45 | 101.2 | 102.2 | 102.3
5.8 | 1.9 | 3.1 | 7 | 19 | 90.1 | 101.1 | 102.2 | 102.3286
5.9 | 1.8 | 3.05 | 6.8 | 17.1 | 87.85 | 101.1 | 102.12 | 102.4
6 | 1.8 | 3 | 6.5 | 15.5 | 85.4 | 101 | 102.1 | 102.4
6.1 | 1.8 | 2.8 | 6.3 | 14.35 | 81.76 | 100.9 | 102.1 | 102.4125
6.2 | 1.7 | 2.7 | 6.1 | 13.3 | 76.9 | 100.9 | 102.1 | 102.5
6.3 | 1.7 | 2.7 | 6 | 12.5 | 69.7 | 100.8 | 102.1 | 102.5
6.4 | 1.6 | 2.7 | 5.75 | 11.7 | 61.8 | 100.8 | 102.1 | 102.5
6.5 | 1.6 | 2.6 | 5.55 | 11.2 | 53.3 | 100.7 | 102.1 | 102.5
6.6 | 1.6 | 2.5 | 5.4 | 10.65 | 45.525 | 100.65 | 102.05 | 102.4
6.7 | 1.5 | 2.5 | 5.3 | 10.15 | 37.2 | 100.6 | 102 | 102.4
6.8 | 1.5 | 2.4 | 5.1 | 9.8 | 31.575 | 100.5 | 102 | 102.4
6.9 |  | 2.4 | 5 | 9.4 | 26.5 | 100.42 | 101.95 | 102.4
7 |  | 2.3 | 4.85 | 9 | 23 | 100.3 | 101.9 | 102.4
7.1 |  | 2.2 | 4.7 | 8.725 | 20.4 | 100.3 | 101.9 | 102.4
7.2 |  | 2.1 | 4.6 | 8.475 | 18.475 | 100.2 | 101.9 | 102.4
7.3 |  | 2.1 | 4.5 | 8.2 | 17.1 | 100.125 | 101.8 | 102.4
7.4 |  | 2.1 | 4.4 | 8 | 15.85 | 100.05 | 101.8 | 102.3
7.5 |  | 2 | 4.3 | 7.75 | 14.8 | 99.9 | 101.74 | 102.3
7.6 |  |  | 4.2 | 7.525 | 14.1 | 99.76 | 101.7 | 102.3
7.7 |  |  | 4.1 | 7.2 | 13.25 | 99.625 | 101.6 | 102.3
7.8 |  |  | 3.9667 | 7.1 | 12.74 | 99.6 | 101.6 | 102.3
7.9 |  |  | 3.9 | 6.92 | 12.1 | 99.5 | 101.6 | 102.2667
8 |  |  | 3.8 | 6.8 | 11.65 | 99.4 | 101.5 | 102.2
8.1 |  |  | 3.7 | 6.6 | 11.2333 | 99.2667 | 101.5 | 102.2
8.2 |  |  | 3.6 | 6.4 | 10.76 | 99.1 | 101.5 | 102.2
8.3 |  |  | 3.5 | 6.225 | 10.4 | 99 | 101.4 | 102.1167
8.4 |  |  | 3.5 | 6.1 | 10.15 | 98.8667 | 101.3833 | 102.1
8.5 |  |  | 3.4 | 5.9 | 9.84 | 98.7 | 101.3 | 102.0333
8.6 |  |  | 3.3 | 5.8 | 9.5 | 98.6 | 101.2 | 102
8.7 |  |  | 3.2 | 5.6 | 9.25 | 98.4667 | 101.2 | 102
8.8 |  |  | 3.2 | 5.525 | 9.04 | 98.3 | 101.1 | 102
8.9 |  |  |  | 5.4 | 8.775 | 98.1 | 101 | 101.9375
9 |  |  |  | 5.2 | 8.6 | 98 | 101 | 101.9
9.1 |  |  |  | 5.1 | 8.3 | 97.9 | 100.9667 | 101.9
9.2 |  |  |  | 5 | 8.15 | 97.72 | 100.9 | 101.9
9.3 |  |  |  | 4.9 | 7.9 | 97.56 | 100.9 | 101.8143
9.4 |  |  |  | 4.8 | 7.75 | 97.425 | 100.86 | 101.8
9.5 |  |  |  | 4.7 | 7.5 | 97.275 | 100.8 | 101.8
9.6 |  |  |  | 4.525 | 7.32 | 97.1 | 100.7 | 101.8
9.7 |  |  |  | 4.475 | 7.2 | 96.9 | 100.6167 | 101.8
9.8 |  |  |  | 4.4 | 7 | 96.7 | 100.6 | 101.7
9.9 |  |  |  | 4.3 | 6.84 | 96.6 | 100.6 | 101.7
10 |  |  |  | 4.2 | 6.8 | 96.4 | 100.5 | 101.7
10.1 |  |  |  | 4.1 | 6.5333 | 96.2 | 100.5 | 101.7
10.2 |  |  |  |  | 6.475 | 96 | 100.4 | 101.7
10.3 |  |  |  |  | 6.3 | 95.825 | 100.3 | 101.6
10.4 |  |  |  |  | 6.15 | 95.525 | 100.2833 | 101.6
10.5 |  |  |  |  | 6 | 95.325 | 100.2 | 101.5375
10.6 |  |  |  |  | 5.9 | 95.1 | 100.1 | 101.5
10.7 |  |  |  |  | 5.725 | 94.8 | 100.1 | 101.4
10.8 |  |  |  |  | 5.6 | 94.65 | 100 | 101.4
10.9 |  |  |  |  | 5.54 | 94.4 | 100 | 101.3143
11 |  |  |  |  | 5.5 | 93.95 | 99.8333 | 101.3
11.1 |  |  |  |  | 5.375 | 93.64 | 99.7 | 101.3
11.2 |  |  |  |  | 5.2 | 93.35 | 99.6 | 101.2
11.3 |  |  |  |  | 5.1 | 93 | 99.5 | 101.2
11.4 |  |  |  |  | 5 | 92.55 | 99.5 | 101.15
11.5 | [5 empty] | 92.125 | 99.4571 | 101.1
11.6 | [5 empty] | 91.65 | 99.3667 | 101.0125
11.7 | [5 empty] | 91.075 | 99.1667 | 101
11.8 | [5 empty] | 90.4 | 99.1 | 100.9429
11.9 | [5 empty] | 89.76 | 99 | 100.9
12 | [5 empty] | 89 | 98.82 | 100.8571
12.1 | [5 empty] | 88 | 98.72 | 100.8
12.2 | [5 empty] | 86.8 | 98.7 | 100.8
12.3 | [5 empty] | 85.3 | 98.66 | 100.8
12.4 | [5 empty] | 83.68 | 98.6 | 100.8
12.5 | [5 empty] | 81.2 | 98.55 | 100.75
12.6 | [5 empty] | 78.325 | 98.4833 | 100.7
12.7 | [5 empty] | 73.925 | 98.4 | 100.7
12.8 | [5 empty] | 68.7 | 98.3 | 100.7
12.9 | [5 empty] | 61.475 | 98.3 | 100.7
13 | [5 empty] | 54 | 98.2 | 100.7
13.1 | [5 empty] | 46.8 | 98.1 | 100.7
13.2 | [5 empty] | 40.2 | 98.0333 | 100.7
13.3 | [5 empty] | 34.16 | 97.95 | 100.7
13.4 | [5 empty] | 29.7 | 97.7833 | 100.7
13.5 | [5 empty] | 26.35 | 97.7 | 100.7
13.6 | [5 empty] | 23.9 | 97.65 | 100.6
13.7 | [5 empty] | 22.15 | 97.6 | 100.525
13.8 | [5 empty] | 20.8 | 97.5167 | 100.5
13.9 | [5 empty] | 19.68 | 97.45 | 100.4429
14 | [5 empty] | 18.6667 | 97.34 | 100.4
14.1 | [5 empty] | 17.925 | 97.25 | 100.3
14.2 | [5 empty] | 17.175 | 97.0833 | 100.3
14.3 | [5 empty] | 16.55 | 97 | 100.225
14.4 | [5 empty] | 15.95 | 96.8 | 100.2
14.5 | [5 empty] | 15.4667 | 96.68 | 100.1375
14.6 | [5 empty] | 15.0667 | 96.5 | 100.1
14.7 | [5 empty] | 14.58 | 96.35 | 99.9875
14.8 | [5 empty] | 14.3 | 96.1833 | 99.9
14.9 | [5 empty] | 13.9333 | 96.0167 | 99.8
15 | [5 empty] | 13.6 | 95.95 | 99.7571
15.1 | [5 empty] | 13.3333 | 95.7833 | 99.7
15.2 | [5 empty] | 13 | 95.6 | 99.6333
15.3 | [5 empty] | 12.75 | 95.3333 | 99.6
15.4 | [5 empty] | 12.5 | 95.2 | 99.5
15.5 | [5 empty] | 12.2667 | 94.9 | 99.35
15.6 | [5 empty] | 11.975 | 94.8333 | 99.3
15.7 | [5 empty] | 11.775 | 94.6667 | 99.26
15.8 | [5 empty] | 11.525 | 94.45 | 99.1714
15.9 | [5 empty] | 11.3 | 94.2 | 99.1
16 | [5 empty] | 11.1 | 94 | 98.9571
16.1 | [5 empty] | 10.9 | 93.85 | 98.8
16.2 | [5 empty] | 10.7 | 93.6 | 98.7625
16.3 | [5 empty] | 10.5333 | 93.45 | 98.7
16.4 | [5 empty] | 10.3 | 93.3 | 98.6125
16.5 | [5 empty] | 10.1 | 93.08 | 98.4714
16.6 | [5 empty] | 9.95 | 92.78 | 98.3375
16.7 | [5 empty] | 9.8 | 92.55 | 98.2222
16.8 | [5 empty] | 9.6 | 92.3 | 98.1
16.9 | [5 empty] | 9.4 | 92.075 | 98
17 | [5 empty] | 9.2333 | 91.85 | 97.9429
17.1 | [5 empty] | 9.0167 | 91.56 | 97.8125
17.2 | [5 empty] | 8.9 | 91.3 | 97.6875
17.3 | [5 empty] | 8.8 | 91.04 | 97.6
17.4 | [5 empty] | 8.6 | 90.7 | 97.5375
17.5 | [5 empty] | 8.4667 | 90.46 | 97.4125
17.6 | [5 empty] | 8.3 | 90.1 | 97.3
17.7 | [5 empty] | 8.1 | 89.82 | 97.2714
17.8 | [5 empty] | 8 | 89.42 | 97.1444
17.9 | [5 empty] | 7.82 | 89.1 | 97.025
18 | [6 empty] | 88.65 | 97
18.1 | [6 empty] | 88.24 | 96.8714
18.2 | [6 empty] | 87.7 | 96.7286
18.3 | [6 empty] | 87.08 | 96.6667
18.4 | [6 empty] | 86.5 | 96.6
18.5 | [6 empty] | 85.75 | 96.45
18.6 | [6 empty] | 85 | 96.4
18.7 | [6 empty] | 84.0667 | 96.2778
18.8 | [6 empty] | 82.8 | 96.2
18.9 | [6 empty] | 81.26 | 96.075
19 | [6 empty] | 79.45 | 95.9
19.1 | [6 empty] | 76.85 | 95.8
19.2 | [6 empty] | 73.44 | 95.7
19.3 | [6 empty] | 68.6 | 95.6167
19.4 | [6 empty] | 62.65 | 95.45
19.5 | [6 empty] | 56.3 | 95.3
19.6 | [6 empty] | 49.44 | 95.225
19.7 | [6 empty] | 42.96 | 95.0571
19.8 | [6 empty] | 37 | 94.9143
19.9 | [6 empty] | 32.325 | 94.7571
20 | [6 empty] | 28.8 | 94.625
20.1 | [6 empty] | 26.3 | 94.5
20.2 | [6 empty] | 24.175 | 94.32
20.3 | [6 empty] | 22.8667 | 94.25
20.4 | [6 empty] | 21.6 | 94.0875
20.5 | [6 empty] | 20.6 | 93.9571
20.6 | [6 empty] | 19.78 | 93.7286
20.7 | [6 empty] | 19.08 | 93.5778
20.8 | [6 empty] | 18.46 | 93.375
20.9 | [6 empty] | 17.9 | 93.1857
21 | [6 empty] | 17.5333 | 93.1
21.1 | [6 empty] | 17 | 92.95
21.2 | [6 empty] | 16.66 | 92.7111
21.3 | [6 empty] | 16.26 | 92.5375
21.4 | [6 empty] | 15.96 | 92.4125
21.5 | [6 empty] | 15.6 | 92.175
21.6 | [6 empty] | 15.32 | 91.9625
21.7 | [6 empty] | 15.1 | 91.775
21.8 | [6 empty] | 14.78 | 91.6125
21.9 | [6 empty] | 14.52 | 91.3875
22 | [6 empty] | 14.25 | 91.2625
22.1 | [6 empty] | 14 | 91.1444
22.2 | [6 empty] | 13.8 | 90.9
22.3 | [6 empty] | 13.5 | 90.75
22.4 | [6 empty] | 13.3333 | 90.55
22.5 | [6 empty] | 13.0667 | 90.35
22.6 | [6 empty] | 12.8 | 90.2
22.7 | [6 empty] | 12.6333 | 90.025
22.8 | [6 empty] | 12.48 | 89.8375
22.9 | [6 empty] | 12.3 | 89.7125
23 | [6 empty] | 12.04 | 89.575
23.1 | [6 empty] | 11.8833 | 89.325
23.2 | [6 empty] | 11.7167 | 89.1375
23.3 | [6 empty] | 11.48 | 88.925
23.4 | [6 empty] | 11.28 | 88.7625
23.5 | [6 empty] | 11.1 | 88.4625
23.6 | [6 empty] | 11.0286 | 88.2889
23.7 | [6 empty] | 10.8667 | 87.975
23.8 | [6 empty] | 10.66 | 87.7
23.9 | [6 empty] | 10.4667 | 87.35
24 | [6 empty] | 10.38 | 87.05
24.1 | [6 empty] | 10.1667 | 86.675
24.2 | [6 empty] | 10 | 86.3333
24.3 | [6 empty] | 9.9 | 86.05
24.4 | [6 empty] | 9.75 | 85.55
24.5 | [7 empty] | 85.05
24.6 | [7 empty] | 84.6375
24.7 | [7 empty] | 84.25
24.8 | [7 empty] | 83.7286
24.9 | [7 empty] | 83.05
25 | [7 empty] | 82.3
25.1 | [7 empty] | 81.48
25.2 | [7 empty] | 80.6
25.3 | [7 empty] | 79.3667
25.4 | [7 empty] | 77.9
25.5 | [7 empty] | 75.925
25.6 | [7 empty] | 73.4143
25.7 | [7 empty] | 70.0625
25.8 | [7 empty] | 65.35
25.9 | [7 empty] | 59.625
26 | [7 empty] | 53.5625
26.1 | [7 empty] | 46.9857
26.2 | [7 empty] | 40.9667
26.3 | [7 empty] | 36
26.4 | [7 empty] | 31.78
26.5 | [7 empty] | 28.38
26.6 | [7 empty] | 25.95
26.7 | [7 empty] | 24.25
26.8 | [7 empty] | 22.825
26.9 | [7 empty] | 21.7625
27 | [7 empty] | 20.9
27.1 | [7 empty] | 20.0857
27.2 | [7 empty] | 19.4
27.3 | [7 empty] | 18.8333
27.4 | [7 empty] | 18.4
27.5 | [7 empty] | 17.9
27.6 | [7 empty] | 17.5
27.7 | [7 empty] | 17.16
27.8 | [7 empty] | 16.9
27.9 | [7 empty] | 16.5
28 | [7 empty] | 16.2333
28.1 | [7 empty] | 15.8667
28.2 | [7 empty] | 15.55
28.3 | [7 empty] | 15.3667
28.4 | [7 empty] | 15.1
28.5 | [7 empty] | 14.8333
28.6 | [7 empty] | 14.6
28.7 | [7 empty] | 14.375
28.8 | [7 empty] | 14.1667
28.9 | [7 empty] | 13.95
29 | [7 empty] | 13.7333
29.1 | [7 empty] | 13.475
29.2 | [7 empty] | 13.3333
29.3 | [7 empty] | 13.16
29.4 | [7 empty] | 12.9667
29.5 | [7 empty] | 12.7167
29.6 | [7 empty] | 12.5571
29.7 | [7 empty] | 12.4
29.8 | [7 empty] | 12.15
29.9 | [7 empty] | 11.98
30 | [7 empty] | 11.8
30.1 | [7 empty] | 11.65
30.2 | [7 empty] | 11.4857
30.3 | [7 empty] | 11.32
30.4 | [7 empty] | 11.1333
30.5 | [7 empty] | 10.9833
30.6 | [7 empty] | 10.8167
30.7 | [7 empty] | 10.6333
30.8 | [7 empty] | 10.48
30.9 | [7 empty] | 10.4

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:3 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-43.6 |  |  |  |  | 4.3
-43.5 |  |  |  |  | 4.3
-43.4 |  |  |  |  | 4.388
-43.3 |  |  |  |  | 4.4
-43.2 |  |  |  |  | 4.4383
-43.1 |  |  |  |  | 4.5
-43 |  |  |  |  | 4.5
-42.9 |  |  |  |  | 4.6
-42.8 |  |  |  |  | 4.6
-42.7 |  |  |  |  | 4.6808
-42.6 |  |  |  |  | 4.7
-42.5 |  |  |  |  | 4.7114
-42.4 |  |  |  |  | 4.8
-42.3 |  |  |  |  | 4.849
-42.2 |  |  |  |  | 4.9
-42.1 |  |  |  |  | 4.9
-42 |  |  |  |  | 4.9593
-41.9 |  |  |  |  | 5
-41.8 |  |  |  |  | 5.0537
-41.7 |  |  |  |  | 5.142
-41.6 |  |  |  |  | 5.2
-41.5 |  |  |  |  | 5.2215
-41.4 |  |  |  |  | 5.3197
-41.3 |  |  |  |  | 5.4
-41.2 |  |  |  |  | 5.4
-41.1 |  |  |  |  | 5.4724
-41 |  |  |  |  | 5.554
-40.9 |  |  |  | 3.3 | 5.6419
-40.8 |  |  |  | 3.3 | 5.7334
-40.7 |  |  |  | 3.3512 | 5.8
-40.6 |  |  |  | 3.4 | 5.8019
-40.5 |  |  |  | 3.4 | 5.8902
-40.4 |  |  |  | 3.4898 | 5.9786
-40.3 |  |  |  | 3.5 | 6
-40.2 |  |  |  | 3.5347 | 6.0493
-40.1 |  |  |  | 3.6 | 6.1313
-40 |  |  |  | 3.6 | 6.2175
-39.9 |  |  |  | 3.6607 | 6.2961
-39.8 |  |  |  | 3.7 | 6.3746
-39.7 |  |  |  | 3.7 | 6.4684
-39.6 |  |  |  | 3.7476 | 6.554
-39.5 |  |  |  | 3.8 | 6.6366
-39.4 |  |  |  | 3.8799 | 6.7222
-39.3 |  |  |  | 3.9 | 6.8009
-39.2 |  |  |  | 3.9357 | 6.8893
-39.1 |  |  |  | 4 | 6.969
-39 |  |  |  | 4.0656 | 7.0476
-38.9 |  |  |  | 4.1 | 7.1294
-38.8 |  |  |  | 4.1404 | 7.2159
-38.7 |  |  |  | 4.2 | 7.2944
-38.6 |  |  |  | 4.3 | 7.3821
-38.5 |  |  |  | 4.3091 | 7.4627
-38.4 |  |  |  | 4.4 | 7.553
-38.3 |  |  |  | 4.413 | 7.6473
-38.2 |  |  |  | 4.5 | 7.7356
-38.1 |  |  | 2.3 | 4.5132 | 7.8275
-38 |  |  | 2.3 | 4.6142 | 7.9443
-37.9 |  |  | 2.4 | 4.7 | 8.1008
-37.8 |  |  | 2.4 | 4.7163 | 8.1892
-37.7 |  |  | 2.4 | 4.8 | 8.269
-37.6 |  |  | 2.4213 | 4.8469 | 8.3951
-37.5 |  |  | 2.5 | 4.9479 | 8.5294
-37.4 |  |  | 2.5 | 5 | 8.6178
-37.3 |  |  | 2.5356 | 5.0499 | 8.7141
-37.2 |  |  | 2.6 | 5.1509 | 8.9094
-37.1 |  |  | 2.6 | 5.2519 | 9.0204
-37 |  |  | 2.6101 | 5.3 | 9.1157
-36.9 |  |  | 2.7 | 5.3796 | 9.2728
-36.8 |  |  | 2.7 | 5.4835 | 9.3731
-36.7 |  | 1.8486 | 2.7131 | 5.5 | 9.5093
-36.6 |  | 1.9 | 2.8 | 5.6141 | 9.6374
-36.5 |  | 1.9 | 2.8 | 5.7152 | 9.7458
-36.4 |  | 1.9 | 2.8266 | 5.8162 | 9.9019
-36.3 |  | 1.9 | 2.9 | 5.9 | 10.0051
-36.2 |  | 1.9 | 2.9284 | 5.9182 | 10.1818
-36.1 | 1.7868 | 1.9906 | 3 | 6.0192 | 10.3812
-36 | 1.8 | 2 | 3 | 6.1283 | 10.5832
-35.9 | 1.8 | 2 | 3.0212 | 6.2 | 10.6811
-35.8 | 1.8 | 2 | 3.1 | 6.2652 | 10.8235
-35.7 | 1.8 | 2.0946 | 3.1232 | 6.3662 | 11.0209
-35.6 | 1.8713 | 2.1 | 3.2 | 6.4672 | 11.2242
-35.5 | 1.9 | 2.1 | 3.2128 | 6.5682 | 11.3442
-35.4 | 1.9 | 2.1 | 3.3 | 6.6807 | 11.521
-35.3 | 1.9 | 2.1988 | 3.3364 | 6.7986 | 11.6979
-35.2 | 1.983 | 2.2 | 3.4 | 6.9 | 11.9618
-35.1 | 2 | 2.2 | 3.4509 | 6.958 | 12.1346
-35 | 2 | 2.2875 | 3.5 | 7.059 | 12.3548
-34.9 | 2 | 2.3 | 3.565 | 7.1525 | 12.5799
-34.8 | 2.1 | 2.3 | 3.6 | 7.2933 | 12.7851
-34.7 | 2.1 | 2.4 | 3.689 | 7.4477 | 13.0446
-34.6 | 2.1051 | 2.4051 | 3.8 | 7.5487 | 13.2804
-34.5 | 2.1935 | 2.4935 | 3.8 | 7.6497 | 13.543
-34.4 | 2.2 | 2.5 | 3.8259 | 7.8014 | 13.8776
-34.3 | 2.2089 | 2.5 | 3.9 | 7.9453 | 14.1982
-34.2 | 2.3 | 2.5099 | 3.9671 | 8.0449 | 14.5935
-34.1 | 2.3 | 2.6 | 4.0794 | 8.1753 | 15.0093
-34 | 2.3119 | 2.6119 | 4.1 | 8.2834 | 15.4148
-33.9 | 2.4 | 2.7 | 4.1987 | 8.4689 | 15.8906
-33.8 | 2.4 | 2.714 | 4.2997 | 8.5747 | 16.5979
-33.7 | 2.4293 | 2.8 | 4.3132 | 8.7442 | 17.305
-33.6 | 2.5 | 2.8017 | 4.4017 | 8.9024 | 18.1889
-33.5 | 2.5313 | 2.9 | 4.5 | 9.0626 | 19.2589
-33.4 | 2.6 | 2.9044 | 4.5037 | 9.2323 | 20.5526
-33.3 | 2.6667 | 3 | 4.6055 | 9.3667 | 22.0544
-33.2 | 2.7 | 3.0234 | 4.7201 | 9.5481 | 23.8453
-33.1 | 2.7783 | 3.1 | 4.8211 | 9.7279 | 26.0071
-33 | 2.8 | 3.1364 | 4.9221 | 9.9299 | 28.4038
-32.9 | 2.8803 | 3.2 | 5.0231 | 10.066 | 31.3881
-32.8 | 2.9 | 3.2336 | 5.1241 | 10.2564 | 34.4875
-32.7 | 2.968 | 3.3252 | 5.222 | 10.492 | 38.2036
-32.6 | 3 | 3.4 | 5.3 | 10.6334 | 41.9768
-32.5 | 3.07 | 3.4272 | 5.338 | 10.8581 | 45.9655
-32.4 | 3.171 | 3.5 | 5.4568 | 11.0989 | 49.5389
-32.3 | 3.2 | 3.5435 | 5.5506 | 11.2761 | 52.7191
-32.2 | 3.264 | 3.6445 | 5.6445 | 11.5057 | 55.783
-32.1 | 3.3598 | 3.7 | 5.7456 | 11.8475 | 58.1565
-32 | 3.4 | 3.7466 | 5.8466 | 12.1086 | 60.2523
-31.9 | 3.4722 | 3.8416 | 5.9476 | 12.3731 | 61.6617
-31.8 | 3.5772 | 3.9343 | 6.0851 | 12.7049 | 62.9518
-31.7 | 3.6 | 4.0354 | 6.2 | 13.0788 | 63.871
-31.6 | 3.6792 | 4.1 | 6.2709 | 13.4925 | 64.88
-31.5 | 3.7 | 4.1374 | 6.3802 | 13.941 | 65.6615
-31.4 | 3.7955 | 4.2384 | 6.5894 | 14.4905 | 66.4421
-31.3 | 3.8965 | 4.3394 | 6.6292 | 15.0933 | 67.1756
-31.2 | 3.9976 | 4.4353 | 6.7471 | 15.8659 | 67.7371
-31.1 | 4.0862 | 4.5271 | 6.8974 | 16.7374 | 68.451
-31 | 4.1853 | 4.6 | 7.0848 | 17.828 | 69.1472
-30.9 | 4.3 | 4.6292 | 7.2506 | 19.1102 | 69.8039
-30.8 | 4.3222 | 4.7422 | 7.3587 | 20.849 | 70.4426
-30.7 | 4.4636 | 4.8598 | 7.5195 | 22.7951 | 71.1183
-30.6 | 4.5 | 4.9608 | 7.6608 | 25.0181 | 71.8837
-30.5 | 4.6054 | 5.0618 | 7.8236 | 27.6536 | 72.688
-30.4 | 4.7199 | 5.1733 | 8.0257 | 31.2245 | 73.4396
-30.3 | 4.8209 | 5.2781 | 8.1744 | 34.9313 | 74.2246
-30.2 | 4.9256 | 5.3077 | 8.3582 | 39.1338 | 75.0231
-30.1 | 5.0435 | 5.4087 | 8.5602 | 43.373 | 75.8184
-30 | 5.1 | 5.5194 | 8.742 | 48.0364 | 76.6292
-29.9 | 5.1959 | 5.7107 | 8.9358 | 52.5125 | 77.4541
-29.8 | 5.3137 | 5.8137 | 9.1378 | 56.3382 | 78.1823
-29.7 | 5.4316 | 5.9379 | 9.4097 | 59.8121 | 79.0891
-29.6 | 5.5494 | 6.0496 | 9.6654 | 62.7078 | 79.7962
-29.5 | 5.6673 | 6.1433 | 9.8725 | 64.8761 | 80.5031
-29.4 | 5.7851 | 6.2887 | 10.1618 | 66.7182 | 81.1316
-29.3 | 5.8772 | 6.4635 | 10.4648 | 67.9295 | 81.6883
-29.2 | 5.9749 | 6.5749 | 10.8572 | 69.0668 | 82.2186
-29.1 | 6.1519 | 6.7519 | 11.2212 | 70.0313 | 82.7212
-29 | 6.2899 | 6.877 | 11.6995 | 71.0189 | 83.187
-28.9 | 6.3923 | 7.056 | 12.2255 | 71.8463 | 83.6306
-28.8 | 6.5867 | 7.179 | 12.8278 | 72.704 | 84.102
-28.7 | 6.6943 | 7.3602 | 13.5472 | 73.4501 | 84.5113
-28.6 | 6.8192 | 7.5622 | 14.5149 | 74.1094 | 84.8378
-28.5 | 7.0124 | 7.7642 | 15.7534 | 74.9694 | 85.1715
-28.4 | 7.1519 | 7.9454 | 16.9933 | 75.7181 | 85.5605
-28.3 | 7.3314 | 8.163 | 18.7698 | 76.5668 | 85.8082
-28.2 | 7.4844 | 8.3703 | 21.116 | 77.4345 | 86.0839
-28.1 | 7.6605 | 8.6585 | 23.5987 | 78.3895 | 86.3508
-28 | 7.87 | 8.9615 | 27.14 | 79.262 | 86.5473
-27.9 | 8.0544 | 9.2645 | 31.0094 | 80.2676 | 86.8293
-27.8 | 8.2499 | 9.6568 | 35.4487 | 81.2321 | 87.009
-27.7 | 8.4519 | 10.0948 | 40.6932 | 82.1691 | 87.2743
-27.6 | 8.6539 | 10.5848 | 46.5769 | 83.1218 | 87.4419
-27.5 | 8.9339 | 11.3365 | 51.7051 | 84.0952 | 87.5636
-27.4 | 9.2763 | 12.0526 | 57.3152 | 84.9941 | 87.7005
-27.3 | 9.63 | 13.0424 | 62.5061 | 85.853 | 87.8577
-27.2 | 10.024 | 14.3383 | 66.3347 | 86.5948 | 88.0192
-27.1 | 10.5921 | 15.9102 | 69.8673 | 87.3352 | 88.2212
-27 | 11.1982 | 18.0518 | 72.2444 | 87.9344 | 88.4116
-26.9 | 11.9723 | 20.8082 | 74.0866 | 88.5106 | 88.5126
-26.8 | 12.9505 | 23.8793 | 75.5205 | 89.0764 | 88.6106
-26.7 | 14.3048 | 27.559 | 76.69 | 89.5729 | 88.6892
-26.6 | 15.7195 | 32.4247 | 77.7141 | 89.977 | 88.8354
-26.5 | 17.9081 | 37.9645 | 78.7218 | 90.381 | 89.0042
-26.4 | 20.3382 | 43.9243 | 79.628 | 90.8313 | 89.1405
-26.3 | 23.5733 | 50.5653 | 80.4682 | 91.2418 | 89.2385
-26.2 | 27.9378 | 57.0066 | 81.4071 | 91.5352 | 89.3
-26.1 | 33.0618 | 62.8102 | 82.3339 | 91.8226 | 89.3863
-26 | 38.8685 | 68.0556 | 83.4433 | 92.1434 | 89.5631
-25.9 | 45.5645 | 72.4164 | 84.4792 | 92.3796 | 89.7265
-25.8 | 52.0164 | 75.352 | 85.5774 | 92.6533 | 89.8381
-25.7 | 58.5247 | 77.5496 | 86.7102 | 92.8578 | 89.9684
-25.6 | 64.945 | 79.161 | 87.7885 | 93.0161 | 90.1225
-25.5 | 70.3149 | 80.5224 | 88.9976 | 93.1959 | 90.2195
-25.4 | 74.191 | 81.7067 | 90.2422 | 93.3564 | 90.3766
-25.3 | 77.4877 | 82.8032 | 91.3369 | 93.5149 | 90.486
-25.2 | 79.5912 | 83.8414 | 92.4638 | 93.6511 | 90.587
-25.1 | 81.2178 | 84.8064 | 93.5244 | 93.7452 | 90.677
-25 | 82.4311 | 85.8078 | 94.4727 | 93.8462 | 90.8309
-24.9 | 83.428 | 86.8717 | 95.4075 | 93.9472 | 91.023
-24.8 | 84.5129 | 87.9061 | 96.1468 | 94.0482 | 91.2251
-24.7 | 85.4594 | 88.9668 | 96.8871 | 94.1431 | 91.3635
-24.6 | 86.3822 | 90.2917 | 97.4227 | 94.2419 | 91.4645
-24.5 | 87.3381 | 91.4356 | 97.9933 | 94.3 | 91.6836
-24.4 | 88.3657 | 92.9273 | 98.4663 | 94.4132 | 91.9082
-24.3 | 89.457 | 94.2353 | 98.8703 | 94.5 | 91.9966
-24.2 | 90.7304 | 95.5466 | 99.2401 | 94.5 | 92.1133
-24.1 | 91.9921 | 96.8121 | 99.6214 | 94.6375 | 92.2312
-24 | 93.2887 | 97.9415 | 100.0368 | 94.7 | 92.4177
-23.9 | 94.6231 | 99.1019 | 100.3005 | 94.8002 | 92.5627
-23.8 | 96.0177 | 100.0244 | 100.5363 | 94.9272 | 92.7362
-23.7 | 97.3127 | 100.8545 | 100.7616 | 95.0832 | 92.8645
-23.6 | 98.5044 | 101.7228 | 100.9836 | 95.189 | 93.0057
-23.5 | 99.7716 | 102.3581 | 101.2217 | 95.29 | 93.1734
-23.4 | 100.7921 | 102.8837 | 101.3474 | 95.3457 | 93.3171
-23.3 | 101.7822 | 103.4667 | 101.4 | 95.5 | 93.4
-23.2 | 102.6303 | 103.8756 | 101.4 | 95.5585 | 93.4
-23.1 | 103.2788 | 104.2348 | 101.4 | 95.6917 | 93.4823
-23 | 103.8711 | 104.6061 | 101.4 | 95.7951 | 93.5808
-22.9 | 104.4416 | 104.8575 | 101.4 | 95.8961 | 93.6954
-22.8 | 104.8457 | 105.1173 | 101.4 | 96.03 | 93.7133
-22.7 | 105.2373 | 105.339 | 101.4981 | 96.1956 | 93.8311
-22.6 | 105.5537 | 105.4823 | 101.5 | 96.3563 | 93.942
-22.5 | 105.9337 | 105.6001 | 101.5 | 96.5145 | 94
-22.4 | 106.1832 | 105.718 | 101.5 | 96.651 | 94.044
-22.3 | 106.383 | 105.8 | 101.5307 | 96.7901 | 94.1394
-22.2 | 106.5045 | 105.8 | 101.6 | 96.9645 | 94.2317
-22.1 | 106.6 | 105.8 | 101.6 | 97.1042 | 94.3327
-22 | 106.6873 | 105.8 | 101.6 | 97.2046 | 94.4295
-21.9 | 106.7 | 105.7513 | 101.6063 | 97.2929 | 94.5319
-21.8 | 106.7 | 105.6357 | 101.7 | 97.4085 | 94.689
-21.7 | 106.7 | 105.5346 | 101.7 | 97.5 | 94.7822
-21.6 | 106.6051 | 105.5 | 101.7093 | 97.5521 | 94.8628
-21.5 | 106.504 | 105.4325 | 101.8103 | 97.6744 | 94.9826
-21.4 | 106.5 | 105.4 | 101.9 | 97.7 | 95.1224
-21.3 | 106.4019 | 105.4 | 101.9 | 97.8 | 95.2192
-21.2 | 106.3009 | 105.4 | 101.9419 | 97.8 | 95.3763
-21.1 | 106.2 | 105.36 | 102.0429 | 97.8501 | 95.4751
-21 | 106.1989 | 105.3 | 102.1 | 97.9 | 95.6128
-20.9 | 106.1 | 105.3 | 102.1525 | 97.9 | 95.7449
-20.8 | 106.0963 | 105.3 | 102.2703 | 97.946 | 95.846
-20.7 | 106 | 105.3 | 102.3756 | 98.047 | 95.9
-20.6 | 105.9806 | 105.3 | 102.4894 | 98.1 | 95.9766
-20.5 | 105.9 | 105.3 | 102.6204 | 98.149 | 96.1552
-20.4 | 105.9 | 105.3 | 102.7214 | 98.2 | 96.2786
-20.3 | 105.9 | 105.3 | 102.8224 | 98.2894 | 96.4393
-20.2 | 105.9 | 105.3 | 102.9 | 98.3 | 96.5774
-20.1 | 105.9 | 105.3119 | 102.9388 | 98.4245 | 96.6
-20 | 105.9 | 105.4 | 103 | 98.5298 | 96.8
-19.9 | 105.9 | 105.4571 | 103.0476 | 98.6571 | 96.8122
-19.8 | 105.9 | 105.5 | 103.1 | 98.7 | 96.9132
-19.7 | 105.9 | 105.5 | 103.1 | 98.8 | 97.0166
-19.6 | 105.901 | 105.6345 | 103.1 | 98.8614 | 97.1345
-19.5 | 106.002 | 105.7524 | 103.0571 | 98.9857 | 97.2
-19.4 | 106.1 | 105.8702 | 103 | 99.0888 | 97.2744
-19.3 | 106.1035 | 105.9 | 103 | 99.2 | 97.3755
-19.2 | 106.192 | 105.9765 | 102.9092 | 99.2393 | 97.4
-19.1 | 106.2918 | 106 | 102.9 | 99.3685 | 97.4678
-19 | 106.4 | 106 | 102.9 | 99.4916 | 97.59
-18.9 | 106.4081 | 105.9062 | 102.9 | 99.6 | 97.6821
-18.8 | 106.5 | 105.9 | 102.9 | 99.6 | 97.7
-18.7 | 106.4882 | 105.9 | 102.9 | 99.6959 | 97.7673
-18.6 | 106.4 | 105.8031 | 102.9255 | 99.7969 | 97.8
-18.5 | 106.4 | 105.7 | 103.0265 | 99.9 | 97.8607
-18.4 | 106.3846 | 105.6884 | 103.1 | 99.9584 | 97.949
-18.3 | 106.3 | 105.6001 | 103.1 | 100 | 98.0374
-18.2 | 106.2387 | 105.5786 | 103.1383 | 100.1011 | 98.1229
-18.1 | 106.1978 | 105.5 | 103.2 | 100.2 | 98.2
-18 | 106.1 | 105.4298 | 103.2368 | 100.2441 | 98.2034
-17.9 | 106.0144 | 105.4 | 103.3 | 100.3 | 98.3183
-17.8 | 105.895 | 105.4 | 103.4 | 100.3225 | 98.4
-17.7 | 105.8 | 105.4 | 103.4061 | 100.4 | 98.4
-17.6 | 105.8 | 105.4582 | 103.5 | 100.4 | 98.4213
-17.5 | 105.7805 | 105.5 | 103.5366 | 100.4446 | 98.5261
-17.4 | 105.7 | 105.5 | 103.6 | 100.5 | 98.6439
-17.3 | 105.7101 | 105.5529 | 103.6951 | 100.5 | 98.7529
-17.2 | 105.8 | 105.6 | 103.7156 | 100.5539 | 98.8472
-17.1 | 105.8 | 105.6 | 103.8 | 100.6 | 98.9317
-17 | 105.8274 | 105.649 | 103.8 | 100.6 | 99
-16.9 | 105.9 | 105.7 | 103.8 | 100.6999 | 99.0165
-16.8 | 105.9 | 105.7 | 103.8612 | 100.7 | 99.1343
-16.7 | 105.9 | 105.7 | 103.9 | 100.7826 | 99.2
-16.6 | 105.9867 | 105.77 | 103.9 | 100.8 | 99.2401
-16.5 | 106.0055 | 105.8 | 103.9 | 100.8468 | 99.3284
-16.4 | 106.1 | 105.8269 | 103.9 | 100.9669 | 99.4168
-16.3 | 106.1 | 105.9 | 103.9 | 101 | 99.5052
-16.2 | 106.1 | 105.9 | 103.9 | 101 | 99.5936
-16.1 | 106.1 | 105.9311 | 103.9 | 101.0711 | 99.6729
-16 | 106.1 | 105.9395 | 103.9 | 101.1 | 99.7579
-15.9 | 106.1 | 105.9 | 103.9 | 101.1 | 99.8
-15.8 | 106.1 | 105.9 | 103.9 | 101.1824 | 99.8396
-15.7 | 106.1 | 105.9 | 103.9 | 101.2 | 99.9
-15.6 | 106.1 | 105.9 | 103.9 | 101.2844 | 99.9324
-15.5 | 106.1 | 105.9 | 103.9 | 101.3748 | 100.0123
-15.4 | 106.1 | 105.8 | 103.9 | 101.4722 | 100.1008
-15.3 | 106.0982 | 105.7982 | 103.9 | 101.5 | 100.2
-15.2 | 106 | 105.7 | 103.9 | 101.5742 | 100.2133
-15.1 | 106 | 105.7 | 103.9 | 101.6 | 100.2919
-15 | 106 | 105.7 | 103.9 | 101.6763 | 100.3
-14.9 | 105.9941 | 105.6518 | 103.9 | 101.7 | 100.3773
-14.8 | 105.9 | 105.6 | 103.9 | 101.781 | 100.4914
-14.7 | 105.8921 | 105.6 | 103.9 | 101.8 | 100.5092
-14.6 | 105.8 | 105.5675 | 103.9 | 101.8 | 100.6
-14.5 | 105.8 | 105.5 | 103.9 | 101.8957 | 100.6339
-14.4 | 105.7891 | 105.5 | 104 | 101.9 | 100.7
-14.3 | 105.7 | 105.5 | 104 | 101.9 | 100.7605
-14.2 | 105.6761 | 105.5 | 104.0318 | 101.9987 | 100.8
-14.1 | 105.6 | 105.5 | 104.1 | 102.0873 | 100.8
-14 | 105.6 | 105.5 | 104.1 | 102.1 | 100.8721
-13.9 | 105.6 | 105.5 | 104.1 | 102.1225 | 100.9569
-13.8 | 105.6 | 105.5 | 104.1839 | 102.2 | 101
-13.7 | 105.6 | 105.5 | 104.2 | 102.2 | 101.0251
-13.6 | 105.6 | 105.5 | 104.2 | 102.2267 | 101.1
-13.5 | 105.6 | 105.5 | 104.2 | 102.3 | 101.1
-13.4 | 105.6 | 105.5 | 104.2 | 102.3 | 101.1
-13.3 | 105.6 | 105.5091 | 104.2 | 102.3935 | 101.1
-13.2 | 105.6 | 105.6 | 104.2 | 102.4 | 101.1
-13.1 | 105.6 | 105.6 | 104.2 | 102.4956 | 101.1
-13 | 105.6 | 105.5606 | 104.2 | 102.5 | 101.1
-12.9 | 105.6 | 105.5 | 104.2 | 102.5 | 101.0881
-12.8 | 105.6 | 105.5 | 104.2 | 102.5 | 101.0101
-12.7 | 105.6 | 105.5 | 104.2622 | 102.5 | 101.0886
-12.6 | 105.6 | 105.5 | 104.3 | 102.5 | 101.1
-12.5 | 105.5698 | 105.5 | 104.3 | 102.5 | 101.1
-12.4 | 105.5 | 105.5 | 104.3 | 102.4116 | 101.1
-12.3 | 105.5 | 105.4963 | 104.3 | 102.4 | 101.1376
-12.2 | 105.5 | 105.4 | 104.3 | 102.4 | 101.2
-12.1 | 105.47 | 105.4 | 104.3 | 102.4 | 101.2
-12 | 105.4 | 105.4 | 104.3 | 102.4 | 101.2638
-11.9 | 105.4 | 105.4 | 104.3 | 102.4 | 101.3
-11.8 | 105.4 | 105.4 | 104.3 | 102.4 | 101.3
-11.7 | 105.4 | 105.4 | 104.3 | 102.3719 | 101.3955
-11.6 | 105.4 | 105.4 | 104.3 | 102.3 | 101.4844
-11.5 | 105.4 | 105.4 | 104.2 | 102.3 | 101.5
-11.4 | 105.4 | 105.3 | 104.185 | 102.3 | 101.5487
-11.3 | 105.3406 | 105.3 | 104.1 | 102.3 | 101.6
-11.2 | 105.3 | 105.2708 | 104.0327 | 102.3292 | 101.6255
-11.1 | 105.2984 | 105.2 | 103.9148 | 102.4 | 101.7
-11 | 105.2 | 105.1727 | 103.8969 | 102.4 | 101.7
-10.9 | 105.1964 | 105.0791 | 103.8 | 102.4465 | 101.7
-10.8 | 105.0954 | 104.9668 | 103.7535 | 102.5 | 101.7041
-10.7 | 104.9943 | 104.9 | 103.7 | 102.5424 | 101.7925
-10.6 | 104.9 | 104.8505 | 103.7 | 102.6 | 101.8
-10.5 | 104.8781 | 104.7293 | 103.7 | 102.6 | 101.8
-10.4 | 104.777 | 104.6848 | 103.7 | 102.6435 | 101.8451
-10.3 | 104.676 | 104.6 | 103.7 | 102.7 | 101.9
-10.2 | 104.5708 | 104.5 | 103.7 | 102.7 | 101.9
-10.1 | 104.4597 | 104.4836 | 103.7 | 102.7975 | 101.9
-10 | 104.4 | 104.4 | 103.7 | 102.8 | 101.9099
-9.9 | 104.4 | 104.4 | 103.7 | 102.8 | 101.9885
-9.8 | 104.4 | 104.4 | 103.7 | 102.8 | 102
-9.7 | 104.4 | 104.4 | 103.7301 | 102.8 | 102
-9.6 | 104.4 | 104.4 | 103.8 | 102.8 | 102
-9.5 | 104.4 | 104.4 | 103.8 | 102.8 | 102.0178
-9.4 | 104.4 | 104.4189 | 103.8 | 102.8553 | 102.1
-9.3 | 104.4 | 104.5 | 103.8 | 102.9 | 102.1
-9.2 | 104.4 | 104.5 | 103.8 | 102.9 | 102.1
-9.1 | 104.4922 | 104.5 | 103.8 | 102.9 | 102.1
-9 | 104.5 | 104.5658 | 103.8 | 102.9 | 102.1
-8.9 | 104.5 | 104.6 | 103.8 | 102.9 | 102.1
-8.8 | 104.5 | 104.6 | 103.8 | 102.9 | 102.1
-8.7 | 104.5 | 104.6 | 103.8 | 102.9 | 102.147
-8.6 | 104.5 | 104.5352 | 103.8 | 102.9 | 102.2
-8.5 | 104.5 | 104.5 | 103.8 | 102.9 | 102.262
-8.4 | 104.4995 | 104.5 | 103.8 | 102.9 | 102.3
-8.3 | 104.4 | 104.5 | 103.8 | 102.9 | 102.3
-8.2 | 104.4 | 104.5 | 103.8 | 102.9 | 102.3
-8.1 | 104.4 | 104.4965 | 103.8 | 102.9035 | 102.3
-8 | 104.4 | 104.4 | 103.8 | 103 | 102.3
-7.9 | 104.3102 | 104.4 | 103.8 | 103 | 102.3173
-7.8 | 104.3 | 104.3943 | 103.8 | 103 | 102.4
-7.7 | 104.3 | 104.3059 | 103.8 | 103 | 102.4
-7.6 | 104.3 | 104.3 | 103.8 | 103 | 102.4622
-7.5 | 104.3 | 104.3 | 103.8 | 103 | 102.5
-7.4 | 104.3 | 104.3 | 103.8 | 103 | 102.5
-7.3 | 104.3 | 104.3 | 103.8 | 103 | 102.5
-7.2 | 104.3 | 104.3 | 103.8 | 103 | 102.5
-7.1 | 104.3 | 104.3 | 103.8 | 103.0326 | 102.5
-7 | 104.3 | 104.3 | 103.8 | 103.1 | 102.5
-6.9 | 104.3 | 104.3 | 103.8 | 103.1 | 102.5
-6.8 | 104.3 | 104.3 | 103.8 | 103.1 | 102.5
-6.7 | 104.2537 | 104.3 | 103.8 | 103.1 | 102.5
-6.6 | 104.2 | 104.3 | 103.8 | 103.1 | 102.5
-6.5 | 104.2 | 104.2937 | 103.8 | 103.0946 | 102.5
-6.4 | 104.2 | 104.2 | 103.8 | 103 | 102.5
-6.3 | 104.2 | 104.2 | 103.8 | 103 | 102.5
-6.2 | 104.2 | 104.2 | 103.7629 | 103 | 102.5
-6.1 | 104.2 | 104.2 | 103.7 | 103 | 102.5
-6 | 104.2 | 104.2 | 103.7 | 103 | 102.5
-5.9 | 104.1885 | 104.2 | 103.7 | 103 | 102.5
-5.8 | 104.1 | 104.1187 | 103.6354 | 103 | 102.5
-5.7 | 104.1 | 104.1 | 103.6 | 103 | 102.4756
-5.6 | 104.1 | 104.1 | 103.6 | 103 | 102.4
-5.5 | 104.1 | 104 | 103.6 | 102.9558 | 102.4
-5.4 | 104.0168 | 104 | 103.6 | 102.9 | 102.4
-5.3 | 103.9961 | 103.9553 | 103.5961 | 102.8538 | 102.4
-5.2 | 103.9 | 103.9 | 103.5 | 102.8 | 102.4
-5.1 | 103.9 | 103.9 | 103.4604 | 102.8 | 102.4
-5 | 103.865 | 103.8936 | 103.4 | 102.8 | 102.4
-4.9 | 103.8 | 103.7914 | 103.3914 | 102.8 | 102.372
-4.8 | 103.7773 | 103.7 | 103.2773 | 102.7 | 102.3
-4.7 | 103.7 | 103.7 | 103.2 | 102.6468 | 102.3
-4.6 | 103.6753 | 103.6181 | 103.1753 | 102.6 | 102.261
-4.5 | 103.564 | 103.5 | 103.1 | 102.5 | 102.2
-4.4 | 103.5 | 103.4521 | 103.059 | 102.5 | 102.1447
-4.3 | 103.5 | 103.4 | 103 | 102.5 | 102.1
-4.2 | 103.3998 | 103.3998 | 102.9426 | 102.5 | 102.1
-4.1 | 103.3 | 103.3 | 102.9 | 102.5 | 102.1
-4 | 103.2835 | 103.2974 | 102.9 | 102.4474 | 102.1
-3.9 | 103.2 | 103.2 | 102.8396 | 102.4 | 102.1
-3.8 | 103.1617 | 103.1814 | 102.7283 | 102.3243 | 102.0463
-3.7 | 103.0518 | 103.0772 | 102.6926 | 102.2233 | 102
-3.6 | 103 | 102.9651 | 102.6 | 102.2 | 102
-3.5 | 102.9415 | 102.9 | 102.5355 | 102.2 | 101.9581
-3.4 | 102.9 | 102.8631 | 102.5 | 102.1283 | 101.9
-3.3 | 102.8058 | 102.8 | 102.4224 | 102.1 | 101.8478
-3.2 | 102.7712 | 102.8 | 102.4 | 102 | 101.7534
-3.1 | 102.7 | 102.7041 | 102.4 | 101.9367 | 101.66
-3 | 102.6189 | 102.6689 | 102.3019 | 101.9 | 101.559
-2.9 | 102.5152 | 102.551 | 102.1017 | 101.8 | 101.4673
-2.8 | 102.4249 | 102.3998 | 102.0832 | 101.7198 | 101.3874
-2.7 | 102.3274 | 102.2846 | 101.9703 | 101.5865 | 101.2991
-2.6 | 102.1528 | 102.177 | 101.8693 | 101.4978 | 101.2205
-2.5 | 101.9508 | 102.0355 | 101.7259 | 101.3968 | 101.142
-2.4 | 101.8244 | 101.9244 | 101.5529 | 101.2916 | 101.0588
-2.3 | 101.7234 | 101.8234 | 101.4327 | 101.1 | 101
-2.2 | 101.632 | 101.7094 | 101.2927 | 101.0927 | 100.9652
-2.1 | 101.5249 | 101.4855 | 101.2 | 101 | 100.8582
-2 | 101.3406 | 101.2904 | 101.1684 | 101 | 100.8
-1.9 | 101.2193 | 101.2 | 101.0392 | 100.9479 | 100.727
-1.8 | 101.2 | 101.1897 | 100.9326 | 100.9 | 100.7
-1.7 | 101.0842 | 101.0887 | 100.8315 | 100.8315 | 100.7
-1.6 | 100.9523 | 100.9877 | 100.8 | 100.7027 | 100.6
-1.5 | 100.8344 | 100.8867 | 100.6867 | 100.6813 | 100.6
-1.4 | 100.7285 | 100.7874 | 100.6 | 100.6 | 100.5333
-1.3 | 100.6275 | 100.6989 | 100.5846 | 100.5321 | 100.449
-1.2 | 100.6 | 100.5975 | 100.4857 | 100.5 | 100.4
-1.1 | 100.5 | 100.5 | 100.4 | 100.4 | 100.3397
-1 | 100.453 | 100.4452 | 100.3959 | 100.3542 | 100.2285
-0.9 | 100.352 | 100.3377 | 100.2949 | 100.3 | 100.2
-0.8 | 100.3 | 100.2446 | 100.1946 | 100.2 | 100.2
-0.7 | 100.2083 | 100.2 | 100.1062 | 100.15 | 100.1437
-0.6 | 100.1061 | 100.1404 | 100.1 | 100.1 | 100.0489
-0.5 | 100.1 | 100.0336 | 100.0051 | 100.0479 | 100
-0.4 | 100.1 | 100 | 100 | 100 | 100
-0.3 | 100.0643 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100.0443 | 100 | 100 | 100 | 100
0.4 | 100.1 | 100.0161 | 100 | 100 | 100
0.5 | 100.1 | 100.1 | 100.0479 | 100.0051 | 100
0.6 | 100.1 | 100.1204 | 100.1 | 100.1061 | 100.0347
0.7 | 100.1916 | 100.2 | 100.1357 | 100.2 | 100.1312
0.8 | 100.3 | 100.2224 | 100.2 | 100.2 | 100.2
0.9 | 100.3234 | 100.3234 | 100.2377 | 100.2728 | 100.2
1 | 100.4285 | 100.4285 | 100.3339 | 100.4 | 100.2118
1.1 | 100.5556 | 100.5 | 100.4255 | 100.4 | 100.3255
1.2 | 100.6 | 100.5693 | 100.5 | 100.4693 | 100.4
1.3 | 100.7 | 100.6703 | 100.5275 | 100.5985 | 100.4325
1.4 | 100.7142 | 100.7714 | 100.6285 | 100.6166 | 100.5124
1.5 | 100.8178 | 100.8724 | 100.7295 | 100.7 | 100.6
1.6 | 100.9356 | 100.9734 | 100.8 | 100.7827 | 100.6
1.7 | 101.0642 | 101.0651 | 100.8701 | 100.8 | 100.6868
1.8 | 101.188 | 101.1611 | 100.9754 | 100.9 | 100.7
1.9 | 101.2088 | 101.2621 | 101 | 100.905 | 100.707
2 | 101.3855 | 101.3737 | 101.106 | 101 | 100.8
2.1 | 101.4927 | 101.4785 | 101.2082 | 101 | 100.8356
2.2 | 101.6081 | 101.6855 | 101.3313 | 101.0594 | 100.9427
2.3 | 101.7181 | 101.7948 | 101.4 | 101.1579 | 101
2.4 | 101.9101 | 101.8958 | 101.5916 | 101.2529 | 101.0824
2.5 | 102.0194 | 102.0155 | 101.6968 | 101.4079 | 101.2197
2.6 | 102.1962 | 102.157 | 101.8121 | 101.555 | 101.2983
2.7 | 102.2988 | 102.2703 | 101.9262 | 101.656 | 101.3865
2.8 | 102.3999 | 102.3832 | 102.1 | 101.7 | 101.4
2.9 | 102.5009 | 102.5294 | 102.1437 | 101.8 | 101.4508
3 | 102.6022 | 102.6426 | 102.2447 | 101.8826 | 101.5447
3.1 | 102.7 | 102.7 | 102.3458 | 102 | 101.6458
3.2 | 102.7546 | 102.8 | 102.4 | 102.0255 | 101.7409
3.3 | 102.8 | 102.8 | 102.4724 | 102.1 | 101.8335
3.4 | 102.8917 | 102.8488 | 102.5 | 102.1 | 101.9
3.5 | 102.9248 | 102.9 | 102.5784 | 102.2 | 101.9355
3.6 | 103 | 102.9508 | 102.6 | 102.2 | 102
3.7 | 103.0376 | 103.0605 | 102.7 | 102.2772 | 102
3.8 | 103.145 | 103.1588 | 102.7 | 102.3814 | 102.0243
3.9 | 103.2 | 103.2629 | 102.7682 | 102.4 | 102.1
4 | 103.2692 | 103.3 | 102.8807 | 102.4978 | 102.1
4.1 | 103.3702 | 103.3 | 102.9 | 102.5 | 102.1
4.2 | 103.4 | 103.3831 | 102.9 | 102.5 | 102.1
4.3 | 103.4865 | 103.4 | 102.9865 | 102.5 | 102.1
4.4 | 103.5 | 103.4355 | 103.0018 | 102.5 | 102.1188
4.5 | 103.5366 | 103.5 | 103.1033 | 102.5533 | 102.2
4.6 | 103.6467 | 103.5878 | 103.2 | 102.6 | 102.2467
4.7 | 103.7 | 103.7 | 103.2191 | 102.7 | 102.3
4.8 | 103.7487 | 103.7 | 103.3 | 102.7 | 102.3
4.9 | 103.8 | 103.764 | 103.3296 | 102.7896 | 102.356
5 | 103.8508 | 103.865 | 103.4 | 102.8 | 102.4
5.1 | 103.9 | 103.9 | 103.4104 | 102.8 | 102.4
5.2 | 103.9 | 103.9 | 103.5 | 102.8 | 102.4
5.3 | 103.9794 | 104 | 103.5395 | 102.9 | 102.4
5.4 | 104 | 104 | 103.6 | 102.9 | 102.4
5.5 | 104.1 | 104 | 103.6 | 102.913 | 102.4
5.6 | 104.1 | 104.0997 | 103.6 | 103 | 102.4
5.7 | 104.1 | 104.1 | 103.6 | 103 | 102.4561
5.8 | 104.1 | 104.1 | 103.6 | 103 | 102.5
5.9 | 104.1699 | 104.2 | 103.6866 | 103 | 102.5
6 | 104.2 | 104.2 | 103.7 | 103 | 102.5
6.1 | 104.2 | 104.2 | 103.7 | 103 | 102.5
6.2 | 104.2 | 104.2 | 103.7051 | 103 | 102.5
6.3 | 104.2 | 104.2 | 103.7935 | 103 | 102.5
6.4 | 104.2 | 104.2 | 103.8 | 103.0507 | 102.5
6.5 | 104.2 | 104.277 | 103.8 | 103.1 | 102.5
6.6 | 104.2 | 104.3 | 103.8 | 103.1 | 102.5
6.7 | 104.2294 | 104.3 | 103.8 | 103.1 | 102.5
6.8 | 104.3 | 104.3 | 103.8 | 103.1 | 102.5
6.9 | 104.3 | 104.3 | 103.8 | 103.1 | 102.5
7 | 104.3 | 104.3 | 103.8 | 103.1 | 102.5
7.1 | 104.3 | 104.3 | 103.8 | 103.0791 | 102.5
7.2 | 104.3 | 104.3 | 103.8 | 103 | 102.5
7.3 | 104.3 | 104.3 | 103.8 | 103 | 102.5
7.4 | 104.3 | 104.3 | 103.8 | 103 | 102.5
7.5 | 104.3 | 104.3 | 103.8 | 103 | 102.5
7.6 | 104.3 | 104.3 | 103.8 | 103 | 102.4825
7.7 | 104.3 | 104.3 | 103.8 | 103 | 102.4
7.8 | 104.3 | 104.3792 | 103.8 | 103 | 102.4
7.9 | 104.3 | 104.4 | 103.8 | 103 | 102.3376
8 | 104.4 | 104.4 | 103.8 | 103 | 102.3
8.1 | 104.4 | 104.4679 | 103.8 | 102.9463 | 102.3
8.2 | 104.4 | 104.5 | 103.8 | 102.9 | 102.3
8.3 | 104.4 | 104.5 | 103.8 | 102.9 | 102.3
8.4 | 104.4828 | 104.5 | 103.8 | 102.9 | 102.3
8.5 | 104.5 | 104.5 | 103.8 | 102.9 | 102.2745
8.6 | 104.5 | 104.6 | 103.8 | 102.9 | 102.2
8.7 | 104.5 | 104.6 | 103.8 | 102.9 | 102.1636
8.8 | 104.5 | 104.6 | 103.8 | 102.9 | 102.1
8.9 | 104.5 | 104.5954 | 103.8 | 102.9 | 102.1
9 | 104.5 | 104.5 | 103.8 | 102.9 | 102.1
9.1 | 104.5 | 104.5 | 103.8 | 102.8255 | 102.1
9.2 | 104.4 | 104.5 | 103.8 | 102.8923 | 102.1
9.3 | 104.4 | 104.5 | 103.8 | 102.9 | 102.1
9.4 | 104.4 | 104.4331 | 103.8 | 102.8887 | 102.1
9.5 | 104.4 | 104.4 | 103.8 | 102.8 | 102.0406
9.6 | 104.4 | 104.4 | 103.8 | 102.8 | 102
9.7 | 104.4 | 104.4 | 103.7 | 102.8 | 102
9.8 | 104.4 | 104.4 | 103.7 | 102.8 | 102
9.9 | 104.4 | 104.4 | 103.7 | 102.8 | 102
10 | 104.4 | 104.4 | 103.7 | 102.8 | 101.9237
10.1 | 104.4454 | 104.4635 | 103.7 | 102.8 | 101.9
10.2 | 104.5541 | 104.5 | 103.7 | 102.7251 | 101.9
10.3 | 104.6617 | 104.6 | 103.7 | 102.7 | 101.8409
10.4 | 104.7549 | 104.6597 | 103.7 | 102.6801 | 101.8
10.5 | 104.8577 | 104.7077 | 103.7 | 102.6 | 101.8
10.6 | 104.9647 | 104.8256 | 103.7 | 102.5781 | 101.8
10.7 | 105 | 104.9372 | 103.7 | 102.5 | 101.8
10.8 | 105.0668 | 105 | 103.7935 | 102.4903 | 101.7165
10.9 | 105.1678 | 105.0535 | 103.8 | 102.4 | 101.7
11 | 105.2 | 105.1545 | 103.8469 | 102.4 | 101.7
11.1 | 105.2841 | 105.2556 | 103.9648 | 102.4 | 101.7
11.2 | 105.3 | 105.3 | 104.0708 | 102.372 | 101.6449
11.3 | 105.4 | 105.3 | 104.1 | 102.3 | 101.6
11.4 | 105.4 | 105.3 | 104.1263 | 102.3 | 101.5655
11.5 | 105.4 | 105.3862 | 104.2196 | 102.3 | 101.5
11.6 | 105.4 | 105.4 | 104.3 | 102.3 | 101.4973
11.7 | 105.4 | 105.4 | 104.3 | 102.3219 | 101.4187
11.8 | 105.4 | 105.4 | 104.3 | 102.4 | 101.3102
11.9 | 105.4 | 105.4 | 104.3 | 102.4 | 101.2923
12 | 105.4 | 105.4 | 104.3 | 102.4 | 101.2
12.1 | 105.4514 | 105.4 | 104.3 | 102.4 | 101.2
12.2 | 105.5 | 105.4 | 104.3 | 102.4 | 101.2
12.3 | 105.5 | 105.482 | 104.3 | 102.4 | 101.1543
12.4 | 105.5 | 105.5 | 104.3 | 102.4727 | 101.1
12.5 | 105.5481 | 105.5 | 104.3 | 102.5 | 101.1
12.6 | 105.6 | 105.5 | 104.3 | 102.5 | 101.0894
12.7 | 105.6 | 105.5 | 104.3 | 102.5 | 101.0108
12.8 | 105.6 | 105.5 | 104.2151 | 102.5 | 101
12.9 | 105.6 | 105.5 | 104.2 | 102.5 | 101.0646
13 | 105.6 | 105.5463 | 104.2 | 102.5 | 101.1
13.1 | 105.6 | 105.6 | 104.2 | 102.4384 | 101.1
13.2 | 105.6 | 105.6 | 104.2 | 102.4 | 101.1
13.3 | 105.6 | 105.531 | 104.2 | 102.4 | 101.1
13.4 | 105.6 | 105.5 | 104.2 | 102.3246 | 101.1
13.5 | 105.6 | 105.5 | 104.2 | 102.2901 | 101.1
13.6 | 105.6 | 105.5 | 104.2 | 102.2 | 101.1
13.7 | 105.6 | 105.5 | 104.2 | 102.2 | 101.0408
13.8 | 105.6 | 105.5 | 104.2 | 102.2 | 101
13.9 | 105.6 | 105.5 | 104.1024 | 102.1 | 100.9765
14 | 105.6 | 105.5 | 104.1 | 102.1 | 100.8882
14.1 | 105.6 | 105.5 | 104.0996 | 102.0283 | 100.8
14.2 | 105.6585 | 105.5 | 104 | 102 | 100.8
14.3 | 105.7694 | 105.5 | 104 | 101.9355 | 100.773
14.4 | 105.8 | 105.5 | 103.9681 | 101.9 | 100.7
14.5 | 105.8 | 105.5 | 103.9 | 101.8385 | 100.6616
14.6 | 105.8 | 105.5475 | 103.9 | 101.8 | 100.6
14.7 | 105.8635 | 105.6 | 103.9 | 101.8 | 100.5259
14.8 | 105.9 | 105.6 | 103.9 | 101.7212 | 100.5
14.9 | 105.9764 | 105.6318 | 103.9 | 101.7 | 100.3916
15 | 106 | 105.7 | 103.9 | 101.7 | 100.3
15.1 | 106 | 105.7 | 103.9 | 101.6158 | 100.3
15.2 | 106 | 105.7 | 103.9 | 101.6 | 100.2244
15.3 | 106.0839 | 105.7839 | 103.9 | 101.5161 | 100.2
15.4 | 106.1 | 105.8 | 103.9 | 101.4151 | 100.1151
15.5 | 106.1 | 105.8836 | 103.9 | 101.4 | 100.0332
15.6 | 106.1 | 105.9 | 103.9 | 101.3239 | 99.9489
15.7 | 106.1 | 105.9 | 103.9 | 101.2264 | 99.8548
15.8 | 106.1 | 105.9 | 103.9 | 101.2 | 99.8
15.9 | 106.1 | 105.9 | 103.9 | 101.1242 | 99.7587
16 | 106.1 | 105.9229 | 103.9 | 101.1 | 99.7
16.1 | 106.1 | 105.9512 | 103.9 | 101.1 | 99.6945
16.2 | 106.1 | 105.9 | 103.9 | 101 | 99.6061
16.3 | 106.1 | 105.9 | 103.9 | 101 | 99.4354
16.4 | 106.0558 | 105.847 | 103.9 | 100.9087 | 99.4
16.5 | 106 | 105.8 | 103.9 | 100.8879 | 99.3475
16.6 | 106 | 105.7867 | 103.9 | 100.8 | 99.26
16.7 | 105.9 | 105.7 | 103.9 | 100.8 | 99.159
16.8 | 105.9 | 105.7 | 103.9 | 100.7 | 99.1
16.9 | 105.9 | 105.7 | 103.8 | 100.7 | 99.0331
17 | 105.8417 | 105.6703 | 103.8 | 100.6 | 99
17.1 | 105.8 | 105.6 | 103.8 | 100.5978 | 98.9481
17.2 | 105.8 | 105.5683 | 103.7 | 100.5 | 98.8284
17.3 | 105.7243 | 105.5 | 103.6529 | 100.5 | 98.7
17.4 | 105.7 | 105.5 | 103.6 | 100.4805 | 98.6605
17.5 | 105.7634 | 105.5 | 103.5 | 100.4 | 98.5427
17.6 | 105.8 | 105.4749 | 103.4749 | 100.4 | 98.4356
17.7 | 105.8 | 105.4 | 103.4 | 100.4 | 98.4
17.8 | 105.8807 | 105.4 | 103.4 | 100.3558 | 98.4
17.9 | 105.9954 | 105.4 | 103.3455 | 100.3 | 98.3325
18 | 106.1 | 105.4 | 103.3 | 100.2 | 98.2201
18.1 | 106.181 | 105.5311 | 103.2019 | 100.2 | 98.1267
18.2 | 106.2187 | 105.6 | 103.1009 | 100.1438 | 98.1
18.3 | 106.3 | 105.6715 | 103.1 | 100.0333 | 98.05
18.4 | 106.368 | 105.7 | 103.0846 | 99.9 | 97.9615
18.5 | 106.4 | 105.7735 | 103 | 99.8642 | 97.8731
18.6 | 106.4 | 105.8 | 102.9826 | 99.8 | 97.8
18.7 | 106.4658 | 105.8882 | 102.9 | 99.7387 | 97.7839
18.8 | 106.5 | 105.9 | 102.9 | 99.6377 | 97.6955
18.9 | 106.4321 | 105.9 | 102.9 | 99.5367 | 97.6071
19 | 106.4 | 105.9917 | 102.9 | 99.4357 | 97.6
19.1 | 106.3204 | 106 | 102.9 | 99.4 | 97.4803
19.2 | 106.2194 | 106 | 102.9378 | 99.2726 | 97.4
19.3 | 106.1184 | 105.9 | 103 | 99.1547 | 97.3911
19.4 | 106.1 | 105.8869 | 103 | 99.0369 | 97.3027
19.5 | 106.0268 | 105.769 | 103.0972 | 99 | 97.1829
19.6 | 105.9296 | 105.6512 | 103.1 | 98.9014 | 97.1
19.7 | 105.9 | 105.5199 | 103.1 | 98.7801 | 97.0333
19.8 | 105.9 | 105.5 | 103.1 | 98.7 | 96.9276
19.9 | 105.9 | 105.4809 | 103.098 | 98.5972 | 96.8266
20 | 105.9 | 105.3557 | 103 | 98.4684 | 96.7256
20.1 | 105.9 | 105.3 | 102.9953 | 98.4 | 96.6119
20.2 | 105.9 | 105.3 | 102.8806 | 98.3441 | 96.5941
20.3 | 105.9 | 105.3 | 102.7822 | 98.3 | 96.4643
20.4 | 105.9 | 105.3 | 102.6929 | 98.1786 | 96.2929
20.5 | 105.9 | 105.3 | 102.5887 | 98.1 | 96.1859
20.6 | 105.9663 | 105.3 | 102.4472 | 98.0909 | 96.009
20.7 | 106 | 105.3 | 102.3327 | 98 | 95.8857
20.8 | 106.0797 | 105.3 | 102.3 | 97.9869 | 95.7603
20.9 | 106.1 | 105.3 | 102.1327 | 97.9 | 95.6644
21 | 106.1846 | 105.3 | 102.1 | 97.9 | 95.576
21.1 | 106.275 | 105.3399 | 102.0001 | 97.8 | 95.4778
21.2 | 106.3 | 105.4 | 101.9991 | 97.8 | 95.3207
21.3 | 106.3856 | 105.4 | 101.9 | 97.7644 | 95.2233
21.4 | 106.4887 | 105.4 | 101.8685 | 97.7 | 95.1421
21.5 | 106.5897 | 105.416 | 101.8 | 97.7 | 95.018
21.6 | 106.6 | 105.5 | 101.7665 | 97.5949 | 94.8739
21.7 | 106.6917 | 105.5061 | 101.7 | 97.4915 | 94.7953
21.8 | 106.7 | 105.6082 | 101.6643 | 97.4 | 94.7167
21.9 | 106.7 | 105.7261 | 101.6 | 97.3406 | 94.561
22 | 106.7 | 105.8 | 101.6 | 97.1675 | 94.4485
22.1 | 106.6 | 105.8 | 101.6 | 97.0327 | 94.3549
22.2 | 106.5245 | 105.8 | 101.5889 | 96.9204 | 94.2461
22.3 | 106.4025 | 105.7451 | 101.5 | 96.7879 | 94.1519
22.4 | 106.1617 | 105.7 | 101.5 | 96.6737 | 94.0635
22.5 | 105.9005 | 105.6202 | 101.5 | 96.4717 | 94
22.6 | 105.583 | 105.4989 | 101.5 | 96.2867 | 93.9125
22.7 | 105.28 | 105.3717 | 101.4644 | 96.1973 | 93.8
22.8 | 104.8283 | 105.1912 | 101.4 | 96.0633 | 93.73
22.9 | 104.3986 | 104.8908 | 101.4 | 95.9454 | 93.7
23 | 103.8548 | 104.6463 | 101.4 | 95.8379 | 93.5951
23.1 | 103.2583 | 104.2722 | 101.4 | 95.7369 | 93.4954
23.2 | 102.5512 | 103.8306 | 101.4 | 95.6102 | 93.4168
23.3 | 101.8106 | 103.4521 | 101.4 | 95.4741 | 93.4
23.4 | 100.7843 | 102.8551 | 101.3061 | 95.3562 | 93.2749
23.5 | 99.7475 | 102.3416 | 101.27 | 95.2383 | 93.1538
23.6 | 98.4789 | 101.7227 | 100.978 | 95.1318 | 93.0384
23.7 | 97.3317 | 100.9608 | 100.7719 | 95.0193 | 92.888
23.8 | 95.9755 | 100.0127 | 100.5362 | 94.9 | 92.7696
23.9 | 94.6028 | 99.1187 | 100.3147 | 94.843 | 92.5752
24 | 93.2544 | 97.8915 | 100.0778 | 94.7188 | 92.4579
24.1 | 91.9089 | 96.7788 | 99.658 | 94.6 | 92.2479
24.2 | 90.7137 | 95.4604 | 99.2126 | 94.516 | 92.0601
24.3 | 89.5165 | 94.197 | 98.8742 | 94.4788 | 92
24.4 | 88.3081 | 92.8316 | 98.4949 | 94.4 | 91.8415
24.5 | 87.3115 | 91.5783 | 97.9364 | 94.3117 | 91.6863
24.6 | 86.4021 | 90.2947 | 97.4086 | 94.1788 | 91.4788
24.7 | 85.4173 | 89.0354 | 96.8371 | 94.0778 | 91.3778
24.8 | 84.5438 | 87.9503 | 96.1375 | 94 | 91.2535
24.9 | 83.5032 | 86.8152 | 95.3767 | 93.9901 | 91.0576
25 | 82.3847 | 85.7656 | 94.4873 | 93.8891 | 90.8782
25.1 | 81.2271 | 84.7355 | 93.5091 | 93.788 | 90.6895
25.2 | 79.5569 | 83.7848 | 92.4634 | 93.6772 | 90.6011
25.3 | 77.3095 | 82.7532 | 91.4032 | 93.5005 | 90.5003
25.4 | 74.1752 | 81.7427 | 90.2893 | 93.3983 | 90.4118
25.5 | 70.3454 | 80.6196 | 89.0006 | 93.1813 | 90.247
25.6 | 64.9706 | 79.2152 | 87.7694 | 93.0562 | 90.0702
25.7 | 59.1247 | 77.5133 | 86.6379 | 92.8578 | 89.9467
25.8 | 52.5267 | 75.2604 | 85.4771 | 92.6867 | 89.8046
25.9 | 45.6646 | 72.3104 | 84.4792 | 92.3797 | 89.6513
26 | 38.5081 | 68.0559 | 83.4232 | 92.1131 | 89.5131
26.1 | 32.7182 | 62.559 | 82.3945 | 91.8472 | 89.3636
26.2 | 27.9777 | 56.9044 | 81.4171 | 91.5153 | 89.2626
26.3 | 23.9199 | 50.9309 | 80.5069 | 91.2604 | 89.2
26.4 | 20.5657 | 44.0871 | 79.6651 | 90.8564 | 89.1059
26.5 | 17.8622 | 37.8607 | 78.6949 | 90.4523 | 88.937
26.6 | 15.8283 | 32.5958 | 77.7264 | 90.0483 | 88.7887
26.7 | 14.2065 | 27.9164 | 76.7236 | 89.635 | 88.7004
26.8 | 12.8929 | 23.8783 | 75.5345 | 89.0956 | 88.6217
26.9 | 11.8895 | 20.6147 | 74.0299 | 88.553 | 88.5269
27 | 11.2711 | 18.057 | 72.2167 | 87.9152 | 88.4259
27.1 | 10.5742 | 15.9642 | 69.8965 | 87.3741 | 88.2497
27.2 | 10.048 | 14.3455 | 66.486 | 86.6405 | 88.0668
27.3 | 9.677 | 13.1598 | 62.28 | 85.7918 | 87.9022
27.4 | 9.2797 | 12.1182 | 57.5208 | 84.9618 | 87.7382
27.5 | 8.9767 | 11.3459 | 52.1636 | 84.1184 | 87.578
27.6 | 8.6824 | 10.6382 | 46.3774 | 83.1118 | 87.4642
27.7 | 8.4743 | 10.1222 | 40.4786 | 82.2162 | 87.3069
27.8 | 8.209 | 9.7137 | 35.3684 | 81.2139 | 87.0747
27.9 | 8.0048 | 9.3097 | 30.9273 | 80.2564 | 86.793
28 | 7.8028 | 9.0042 | 26.9767 | 79.4116 | 86.5776
28.1 | 7.6612 | 8.7387 | 23.7284 | 78.3893 | 86.3866
28.2 | 7.462 | 8.4412 | 21.0994 | 77.5179 | 86.1267
28.3 | 7.3648 | 8.2255 | 18.8717 | 76.589 | 85.8458
28.4 | 7.1939 | 7.9947 | 17.1727 | 75.7583 | 85.5408
28.5 | 7.0291 | 7.7927 | 15.696 | 74.9747 | 85.1873
28.6 | 6.8479 | 7.5907 | 14.5347 | 74.1675 | 84.8892
28.7 | 6.71 | 7.3887 | 13.5802 | 73.4603 | 84.5955
28.8 | 6.6076 | 7.1883 | 12.7739 | 72.7534 | 84.0689
28.9 | 6.4366 | 7.0115 | 12.1679 | 71.9537 | 83.5976
29 | 6.2398 | 6.9056 | 11.6079 | 70.9927 | 83.1883
29.1 | 6.1053 | 6.8045 | 11.2159 | 70.0836 | 82.7239
29.2 | 6.0035 | 6.6071 | 10.8165 | 69.1194 | 82.2167
29.3 | 5.9025 | 6.4863 | 10.4361 | 67.9442 | 81.718
29.4 | 5.8015 | 6.3459 | 10.133 | 66.6212 | 81.0307
29.5 | 5.6862 | 6.1719 | 9.83 | 64.9092 | 80.4803
29.6 | 5.5827 | 6.0709 | 9.5655 | 62.8191 | 79.7965
29.7 | 5.4579 | 5.9579 | 9.3541 | 59.9766 | 79.0166
29.8 | 5.3304 | 5.8403 | 9.1521 | 56.3288 | 78.1958
29.9 | 5.2125 | 5.6786 | 8.9688 | 52.3808 | 77.4711
30 | 5.0947 | 5.5383 | 8.7766 | 48.2636 | 76.6458
30.1 | 5 | 5.4373 | 8.5745 | 43.388 | 75.831
30.2 | 4.9423 | 5.3256 | 8.3725 | 39.2865 | 75.0537
30.3 | 4.8353 | 5.2078 | 8.1705 | 34.8258 | 74.3467
30.4 | 4.7342 | 5.1 | 8.0232 | 31.2294 | 73.4291
30.5 | 4.6221 | 5.0903 | 7.8378 | 27.9083 | 72.6046
30.6 | 4.5042 | 4.9893 | 7.7179 | 25.1648 | 71.8255
30.7 | 4.4 | 4.8863 | 7.5546 | 22.7612 | 71.118
30.8 | 4.3422 | 4.7685 | 7.3604 | 20.7861 | 70.4258
30.9 | 4.3 | 4.6577 | 7.2173 | 19.149 | 69.804
31 | 4.1996 | 4.6 | 7.1138 | 17.7777 | 69.0968
31.1 | 4.1112 | 4.5113 | 6.9257 | 16.7283 | 68.4134
31.2 | 4.0118 | 4.4 | 6.8118 | 15.8064 | 67.7854
31.3 | 3.9108 | 4.3536 | 6.6951 | 15.1535 | 67.1753
31.4 | 3.8098 | 4.2526 | 6.5537 | 14.5442 | 66.4619
31.5 | 3.7088 | 4.1516 | 6.4374 | 13.8741 | 65.6616
31.6 | 3.7 | 4.0568 | 6.3257 | 13.4417 | 64.8968
31.7 | 3.6068 | 4 | 6.2079 | 13.0883 | 64.0182
31.8 | 3.5058 | 3.9628 | 6.1058 | 12.77 | 62.9335
31.9 | 3.4904 | 3.8618 | 6.0047 | 12.3798 | 61.6469
32 | 3.3894 | 3.7608 | 5.8075 | 12.0753 | 60.1697
32.1 | 3.3 | 3.7 | 5.8 | 11.8276 | 58.2857
32.2 | 3.2874 | 3.6686 | 5.714 | 11.5891 | 55.911
32.3 | 3.2 | 3.5578 | 5.6149 | 11.3682 | 52.9758
32.4 | 3.1854 | 3.5 | 5.5139 | 11.0845 | 49.5864
32.5 | 3.0863 | 3.4488 | 5.398 | 10.896 | 45.8325
32.6 | 3 | 3.3547 | 5.2729 | 10.6277 | 41.9941
32.7 | 2.9966 | 3.3 | 5.1823 | 10.4352 | 38.0809
32.8 | 2.9 | 3.2527 | 5.0813 | 10.2231 | 34.6222
32.9 | 2.8945 | 3.2 | 4.9803 | 10.1088 | 31.4208
33 | 2.7935 | 3.1507 | 4.8793 | 9.9156 | 28.4191
33.1 | 2.7 | 3.0497 | 4.7783 | 9.7135 | 25.9939
33.2 | 2.7 | 3 | 4.7 | 9.6058 | 24.0251
33.3 | 2.6 | 2.9333 | 4.6722 | 9.3762 | 22.184
33.4 | 2.6 | 2.9 | 4.5658 | 9.2503 | 20.5525
33.5 | 2.5456 | 2.8274 | 4.4741 | 9.0397 | 19.2156
33.6 | 2.5 | 2.8 | 4.3731 | 8.9 | 18.1582
33.7 | 2.4435 | 2.8 | 4.3 | 8.7351 | 17.3049
33.8 | 2.4 | 2.7283 | 4.2569 | 8.614 | 16.5979
33.9 | 2.4 | 2.7 | 4.2 | 8.4259 | 15.959
34 | 2.3263 | 2.6355 | 4.1548 | 8.3119 | 15.4559
34.1 | 2.3 | 2.6 | 4.0461 | 8.1218 | 14.998
34.2 | 2.3 | 2.5385 | 4 | 7.9813 | 14.5559
34.3 | 2.2271 | 2.5 | 3.9104 | 7.8828 | 14.1409
34.4 | 2.2 | 2.5 | 3.8911 | 7.7936 | 13.8457
34.5 | 2.2 | 2.4212 | 3.8 | 7.6853 | 13.5828
34.6 | 2.1201 | 2.4 | 3.7568 | 7.4916 | 13.3469
34.7 | 2.1 | 2.4 | 3.7 | 7.3906 | 13.0877
34.8 | 2.1 | 2.3045 | 3.6324 | 7.2896 | 12.8423
34.9 | 2.0033 | 2.3 | 3.5314 | 7.1899 | 12.58
35 | 2 | 2.3 | 3.5 | 7.1016 | 12.3362
35.1 | 2 | 2.2132 | 3.4176 | 7.0008 | 12.1645
35.2 | 2 | 2.2 | 3.4 | 6.8796 | 11.8997
35.3 | 1.9 | 2.2 | 3.3988 | 6.7383 | 11.7426
35.4 | 1.9 | 2.112 | 3.3 | 6.7 | 11.5526
35.5 | 1.9 | 2.1 | 3.2824 | 6.5968 | 11.3693
35.6 | 1.8958 | 2.1 | 3.2 | 6.4963 | 11.1925
35.7 | 1.8 | 2.1 | 3.1804 | 6.4079 | 11.0251
35.8 | 1.8 | 2.0222 | 3.1 | 6.308 | 10.8639
35.9 | 1.8 | 2 | 3.0784 | 6.207 | 10.6871
36 | 1.8 | 2 | 3 | 6.1 | 10.5104
36.1 | 1.8 | 2 | 2.9793 | 6.0557 | 10.4049
36.2 |  | 1.9039 | 2.9 | 5.9534 | 10.2079
36.3 |  | 1.9 | 2.8886 | 5.86 | 10.0301
36.4 |  | 1.9 | 2.8 | 5.759 | 9.9162
36.5 |  | 1.9 | 2.8 | 5.7 | 9.768
36.6 |  | 1.9 | 2.7832 | 5.657 | 9.6109
36.7 |  |  | 2.7 | 5.5383 | 9.474
36.8 |  |  | 2.7 | 5.4264 | 9.3744
36.9 |  |  | 2.7 | 5.4 | 9.2173
37 |  |  | 2.6617 | 5.2958 | 9.1213
37.1 |  |  | 2.6 | 5.1947 | 9.0329
37.2 |  |  | 2.6 | 5.1 | 8.8521
37.3 |  |  | 2.5 | 5.0927 | 8.7311
37.4 |  |  | 2.5 | 4.9917 | 8.6345
37.5 |  |  | 2.5 | 4.9 | 8.4966
37.6 |  |  | 2.4897 | 4.888 | 8.366
37.7 |  |  | 2.4 | 4.8 | 8.28
37.8 |  |  | 2.4 | 4.7591 | 8.2015
37.9 |  |  | 2.4 | 4.6581 | 8.0459
38 |  |  | 2.3499 | 4.6 | 7.9374
38.1 |  |  | 2.3 | 4.5561 | 7.8417
38.2 |  |  |  | 4.5 | 7.7539
38.3 |  |  |  | 4.4463 | 7.6683
38.4 |  |  |  | 4.4 | 7.5673
38.5 |  |  |  | 4.3455 | 7.4738
38.6 |  |  |  | 4.3 | 7.3952
38.7 |  |  |  | 4.2099 | 7.3167
38.8 |  |  |  | 4.1737 | 7.2303
38.9 |  |  |  | 4.1 | 7.1484
39 |  |  |  | 4.1 | 7.0698
39.1 |  |  |  | 4 | 6.9902
39.2 |  |  |  | 3.9734 | 6.9018
39.3 |  |  |  | 3.9 | 6.823
39.4 |  |  |  | 3.9 | 6.7375
39.5 |  |  |  | 3.8 | 6.6548
39.6 |  |  |  | 3.7809 | 6.5732
39.7 |  |  |  | 3.7 | 6.4848
39.8 |  |  |  | 3.7 | 6.3965
39.9 |  |  |  | 3.6949 | 6.308
40 |  |  |  | 3.6 | 6.2286
40.1 |  |  |  | 3.5786 | 6.1438
40.2 |  |  |  | 3.5 | 6.0603
40.3 |  |  |  | 3.5 | 6
40.4 |  |  |  | 3.4257 | 6
40.5 |  |  |  | 3.4 | 5.9031
40.6 |  |  |  | 3.4 | 5.8143
40.7 |  |  |  | 3.3868 | 5.7342
40.8 |  |  |  | 3.3 | 5.7
40.9 |  |  |  | 3.3 | 5.6617
41 |  |  |  |  | 5.5733
41.1 |  |  |  |  | 5.4866
41.2 |  |  |  |  | 5.408
41.3 |  |  |  |  | 5.4
41.4 |  |  |  |  | 5.3397
41.5 |  |  |  |  | 5.2418
41.6 |  |  |  |  | 5.2
41.7 |  |  |  |  | 5.1596
41.8 |  |  |  |  | 5.0757
41.9 |  |  |  |  | 5
42 |  |  |  |  | 4.9736
42.1 |  |  |  |  | 4.9
42.2 |  |  |  |  | 4.8626
42.3 |  |  |  |  | 4.8
42.4 |  |  |  |  | 4.8
42.5 |  |  |  |  | 4.7257
42.6 |  |  |  |  | 4.7
42.7 |  |  |  |  | 4.6
42.8 |  |  |  |  | 4.6
42.9 |  |  |  |  | 4.6
43 |  |  |  |  | 4.5073
43.1 |  |  |  |  | 4.5
43.2 |  |  |  |  | 4.455
43.3 |  |  |  |  | 4.4
43.4 |  |  |  |  | 4.4
43.5 |  |  |  |  | 4.3014
43.6 |  |  |  |  | 4.3

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.877 | 0.894 | 0.902 | 0.91 | 0.916 | 0.919 | 0.921 | 0.923 | 0.923
| 4 | 0.896 | 0.915 | 0.925 | 0.936 | 0.942 | 0.947 | 0.949 | 0.952 | 0.952
| 5 | 0.906 | 0.927 | 0.939 | 0.951 | 0.959 | 0.965 | 0.968 | 0.971 | 0.971
| 7 | 0.917 | 0.941 | 0.955 | 0.97 | 0.98 | 0.988 | 0.992 | 0.996 | 0.997
| 10 | 0.927 | 0.953 | 0.969 | 0.987 | 1 | 1.01 | 1.016 | 1.021 | 1.022
| 15 | 0.934 | 0.963 | 0.98 | 1.001 | 1.017 | 1.03 | 1.038 | 1.045 | 1.047
| 20 | 0.938 | 0.968 | 0.986 | 1.008 | 1.026 | 1.04 | 1.05 | 1.059 | 1.061
| 30 | 0.942 | 0.973 | 0.992 | 1.016 | 1.035 | 1.052 | 1.064 | 1.075 | 1.078
| 40 | 0.945 | 0.976 | 0.995 | 1.02 | 1.04 | 1.058 | 1.07 | 1.082 | 1.085

## Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 22.7 | 30.1 | 42.7 | 54.1 | 58.7
0.1 | 25.6 | 32.82 | 45.66 | 57.02 | 61.5667
0.2 | 29.9 | 37.1333 | 49.825 | 61.2333 | 65.28
0.3 | 36.7 | 44.1 | 56.14 | 66.6 | 70.5
0.4 | 44.46 | 51.5333 | 62.92 | 72.5 | 76.3
0.5 | 52.2 | 58.1 | 68.7 | 77.72 | 81.14
0.6 | 58.48 | 64.16 | 74.1 | 82.1 | 85.1
0.7 | 64.65 | 69.5 | 78.36 | 85.62 | 88.3
0.8 | 69.1 | 73.8 | 81.95 | 88.6 | 90.8
0.9 | 73.32 | 77.425 | 84.74 | 91 | 92.8
1 | 76.95 | 80.65 | 87.5 | 92.9 | 94.5
1.1 | 80.18 | 83.5 | 89.8 | 94.46 | 95.9
1.2 | 82.8 | 86.05 | 91.65 | 95.86 | 97
1.3 | 85.6 | 88.26 | 93.35 | 96.9 | 97.95
1.4 | 87.86 | 90.04 | 94.55 | 97.72 | 98.475
1.5 | 89.68 | 91.9 | 95.68 | 98.3 | 99.15
1.6 | 91.28 | 93.25 | 96.65 | 98.9 | 99.5
1.7 | 92.85 | 94.56 | 97.52 | 99.35 | 99.725
1.8 | 93.975 | 95.5 | 98.225 | 99.5667 | 99.9333
1.9 | 95.18 | 96.38 | 98.62 | 99.875 | 100
2 | 96.2 | 97.06 | 99.12 | 99.9833 | 100
2.1 | 97.175 | 97.82 | 99.46 | 100 | 100
2.2 | 97.65 | 98.32 | 99.7 | 100 | 99.9
2.3 | 98.2 | 98.85 | 99.8 | 100 | 99.8
2.4 | 98.65 | 99.2 | 99.92 | 99.9 | 99.6667
2.5 | 99.1 | 99.45 | 100 | 99.8 | 99.5
2.6 | 99.475 | 99.65 | 100 | 99.66 | 99.325
2.7 | 99.7 | 99.8 | 100 | 99.5 | 99
2.8 | 99.8 | 100 | 99.9 | 99.26 | 98.75
2.9 | 99.9 | 100 | 99.8 | 98.96 | 98.54
3 | 100 | 100 | 99.7 | 98.74 | 98.2333
3.1 | 100 | 100 | 99.5 | 98.5 | 97.9333
3.2 | 100 | 99.8 | 99.3 | 98.26 | 97.6333
3.3 | 99.9 | 99.7 | 99.16 | 97.8 | 97.3
3.4 | 99.7 | 99.6 | 98.85 | 97.62 | 97.02
3.5 | 99.6 | 99.42 | 98.6 | 97.2 | 96.65
3.6 | 99.4167 | 99.325 | 98.25 | 96.925 | 96.24
3.7 | 99.22 | 99.1 | 97.975 | 96.65 | 96.0167
3.8 | 99.02 | 98.86 | 97.68 | 96.3 | 95.7
3.9 | 98.84 | 98.55 | 97.4 | 95.9667 | 95.45
4 | 98.5 | 98.24 | 97.14 | 95.66 | 95.1
4.1 | 98.15 | 97.95 | 96.8 | 95.38 | 94.7
4.2 | 97.88 | 97.84 | 96.44 | 95 | 94.4
4.3 | 97.54 | 97.45 | 96.1667 | 94.64 | 94
4.4 | 97.18 | 97 | 95.825 | 94.325 | 93.6857
4.5 | 96.68 | 96.75 | 95.4333 | 94.02 | 93.35
4.6 | 96.3429 | 96.5 | 95.2 | 93.6 | 93.0333
4.7 | 96.1 | 96.1667 | 94.9 | 93.35 | 92.7
4.8 | 95.56 | 95.7 | 94.52 | 92.98 | 92.38
4.9 | 95.3 | 95.4333 | 94.0857 | 92.74 | 92.05
5 | 94.92 | 95.0667 | 93.8667 | 92.38 | 91.78
5.1 | 94.5333 | 94.65 | 93.52 | 91.9333 | 91.3667
5.2 | 94.15 | 94.3 | 93.1 | 91.675 | 91
5.3 | 93.725 | 93.95 | 92.9 | 91.32 | 90.6667
5.4 | 93.25 | 93.625 | 92.4 | 90.88 | 90.42
5.5 | 92.85 | 93.1667 | 92.0333 | 90.525 | 90.1143
5.6 | 92.5 | 92.8333 | 91.7333 | 90.3333 | 89.8
5.7 | 92.175 | 92.55 | 91.5 | 90 | 89.52
5.8 | 91.82 | 92.05 | 91.15 | 89.75 | 89.08
5.9 | 91.425 | 91.84 | 90.82 | 89.525 | 88.8
6 | 90.9 | 91.5 | 90.45 | 89.0667 | 88.5333
6.1 | 90.54 | 91 | 90.1 | 88.775 | 88.1667
6.2 | 90.05 | 90.54 | 89.7667 | 88.48 | 87.9
6.3 | 89.58 | 90.2 | 89.4 | 88.1333 | 87.6
6.4 | 89.2667 | 89.775 | 89.1 | 87.85 | 87.175
6.5 | 88.98 | 89.525 | 88.7 | 87.5 | 86.925
6.6 | 88.54 | 89.2 | 88.375 | 87.1 | 86.675
6.7 | 87.96 | 88.86 | 88.125 | 86.84 | 86.32
6.8 | 87.55 | 88.6 | 87.775 | 86.54 | 86.0333
6.9 | 87.125 | 88.25 | 87.4667 | 86.25 | 85.74
7 | 86.86 | 87.8667 | 87.1 | 85.88 | 85.3286
7.1 | 86.52 | 87.56 | 86.8333 | 85.58 | 85.15
7.2 | 86.08 | 87.2 | 86.55 | 85.25 | 84.7667
7.3 | 85.55 | 86.925 | 86.25 | 85 | 84.5286
7.4 | 85.1667 | 86.5 | 85.88 | 84.7 | 84.22
7.5 | 84.92 | 86.12 | 85.5667 | 84.4 | 83.8667
7.6 | 84.5 | 85.7667 | 85.25 | 84.02 | 83.52
7.7 | 84.08 | 85.4429 | 84.9 | 83.725 | 83.275
7.8 | 83.75 | 85.15 | 84.46 | 83.3667 | 82.94
7.9 | 83.4 | 84.76 | 84.22 | 83.0333 | 82.62
8 | 83.1 | 84.45 | 83.96 | 82.74 | 82.32
8.1 | 82.6 | 83.98 | 83.65 | 82.5 | 82
8.2 | 82.2 | 83.65 | 83.4 | 82.3 | 81.74
8.3 | 81.8 | 83.3 | 83 | 81.94 | 81.42
8.4 | 81.425 | 82.975 | 82.7 | 81.58 | 81.08
8.5 | 81.1667 | 82.66 | 82.3 | 81.275 | 80.8667
8.6 | 80.8333 | 82.3 | 82.0333 | 80.95 | 80.5333
8.7 | 80.36 | 81.825 | 81.6 | 80.64 | 80.3
8.8 | 79.9667 | 81.54 | 81.3 | 80.4 | 79.98
8.9 | 79.6 | 81.2 | 81.04 | 80.1167 | 79.62
9 | 79.2 | 80.92 | 80.7 | 79.8 | 79.3
9.1 | 78.9 | 80.56 | 80.5 | 79.475 | 79
9.2 | 78.62 | 80.3 | 80.1333 | 79.26 | 78.7
9.3 | 78.35 | 80 | 79.86 | 78.9 | 78.48
9.4 | 78 | 79.55 | 79.4833 | 78.6 | 78.15
9.5 | 77.62 | 79.26 | 79.2571 | 78.3333 | 77.8
9.6 | 77.28 | 78.9 | 78.92 | 77.96 | 77.5333
9.7 | 76.7667 | 78.475 | 78.54 | 77.68 | 77.3
9.8 | 76.3286 | 78.25 | 78.22 | 77.4 | 77.0333
9.9 | 75.95 | 78.04 | 77.98 | 77.12 | 76.74
10 | 75.58 | 77.68 | 77.65 | 76.8429 | 76.475
10.1 | 75.2667 | 77.38 | 77.425 | 76.55 | 76.15
10.2 | 74.9333 | 76.98 | 77.05 | 76.2 | 75.9167
10.3 | 74.6 | 76.66 | 76.725 | 75.9 | 75.6
10.4 | 74.2333 | 76.3 | 76.5167 | 75.6 | 75.275
10.5 | 73.96 | 76 | 76.1714 | 75.38 | 75.025
10.6 | 73.58 | 75.74 | 75.92 | 75.05 | 74.82
10.7 | 73.18 | 75.38 | 75.6 | 74.8 | 74.56
10.8 | 72.94 | 75.12 | 75.2667 | 74.5 | 74.175
10.9 | 72.66 | 74.7 | 74.9 | 74.275 | 73.925
11 | 72.25 | 74.5 | 74.6 | 74 | 73.6
11.1 | 71.9 | 74.175 | 74.28 | 73.66 | 73.275
11.2 | 71.64 | 73.8 | 74.075 | 73.38 | 73.1
11.3 | 71.2333 | 73.45 | 73.725 | 73.15 | 72.775
11.4 | 70.9 | 73.2167 | 73.4667 | 72.9 | 72.5333
11.5 | 70.6 | 72.84 | 73.22 | 72.58 | 72.16
11.6 | 70.32 | 72.6 | 72.9 | 72.4 | 71.9
11.7 | 70.08 | 72.3 | 72.6 | 72 | 71.78
11.8 | 69.65 | 71.8333 | 72.38 | 71.875 | 71.45
11.9 | 69.3 | 71.58 | 72.12 | 71.56 | 71.1667
12 | 68.9857 | 71.3333 | 71.86 | 71.36 | 70.9
12.1 | 68.74 | 71 | 71.5667 | 71.08 | 70.65
12.2 | 68.4333 | 70.8 | 71.3167 | 70.725 | 70.4143
12.3 | 68.05 | 70.375 | 71.05 | 70.4 | 70.1
12.4 | 67.6 | 70.06 | 70.7667 | 70.225 | 69.8667
12.5 | 67.26 | 69.7714 | 70.48 | 69.9667 | 69.7143
12.6 | 67 | 69.52 | 70.16 | 69.6333 | 69.4
12.7 | 66.725 | 69.22 | 69.9 | 69.4 | 69.12
12.8 | 66.36 | 68.94 | 69.7 | 69.0857 | 68.76
12.9 | 66 | 68.6 | 69.425 | 68.92 | 68.48
13 | 65.8 | 68.2667 | 69.175 | 68.5667 | 68.275
13.1 | 65.475 | 67.95 | 68.825 | 68.3 | 68.025
13.2 | 65.225 | 67.7 | 68.55 | 68.0333 | 67.78
13.3 | 65.0429 | 67.5333 | 68.2 | 67.8667 | 67.54
13.4 | 64.68 | 67.2 | 67.8857 | 67.625 | 67.24
13.5 | 64.4667 | 66.9 | 67.6667 | 67.3 | 67.04
13.6 | 64.2 | 66.675 | 67.4333 | 67.025 | 66.74
13.7 | 63.74 | 66.32 | 67.175 | 66.8 | 66.58
13.8 | 63.5 | 66.025 | 66.925 | 66.5667 | 66.275
13.9 | 63.3 | 65.82 | 66.6667 | 66.32 | 66.06
14 | 62.9 | 65.55 | 66.4 | 66.12 | 65.7714
14.1 | 62.7 | 65.1 | 66.1333 | 65.8 | 65.5333
14.2 | 62.4667 | 64.9 | 65.86 | 65.6 | 65.25
14.3 | 62.1 | 64.72 | 65.575 | 65.3 | 65
14.4 | 61.8333 | 64.4 | 65.25 | 65.0333 | 64.82
14.5 | 61.54 | 64.16 | 65.1 | 64.75 | 64.5
14.6 | 61.22 | 63.85 | 64.88 | 64.4833 | 64.35
14.7 | 60.96 | 63.625 | 64.66 | 64.3 | 64.0667
14.8 | 60.62 | 63.3333 | 64.3667 | 64.12 | 63.75
14.9 | 60.275 | 63.025 | 64 | 63.74 | 63.5
15 | 60.0333 | 62.7 | 63.7667 | 63.475 | 63.25
15.1 | 59.76 | 62.42 | 63.52 | 63.2833 | 63.04
15.2 | 59.5143 | 62.125 | 63.3 | 63 | 62.85
15.3 | 59.2167 | 61.9 | 63.08 | 62.7333 | 62.5
15.4 | 58.98 | 61.72 | 62.86 | 62.56 | 62.35
15.5 | 58.65 | 61.4 | 62.6 | 62.36 | 62.075
15.6 | 58.46 | 61.1 | 62.375 | 61.98 | 61.84
15.7 | 58.16 | 60.9 | 62.06 | 61.8 | 61.525
15.8 | 57.8667 | 60.74 | 61.8667 | 61.62 | 61.36
15.9 | 57.66 | 60.35 | 61.6 | 61.45 | 61.175
16 | 57.2667 | 60.16 | 61.3 | 61.15 | 61
16.1 | 57.06 | 59.9 | 61.1333 | 60.9 | 60.78
16.2 | 56.8833 | 59.64 | 60.9 | 60.66 | 60.4667
16.3 | 56.6333 | 59.4333 | 60.6667 | 60.45 | 60.2286
16.4 | 56.325 | 59.3 | 60.4167 | 60.24 | 60
16.5 | 56.1 | 59 | 60.14 | 60 | 59.8
16.6 | 55.85 | 58.8 | 59.94 | 59.76 | 59.56
16.7 | 55.5 | 58.475 | 59.66 | 59.5 | 59.35
16.8 | 55.3 | 58.18 | 59.44 | 59.3 | 59.16
16.9 | 55.1 | 57.9 | 59.2 | 59.04 | 58.96
17 | 54.82 | 57.76 | 59 | 58.875 | 58.6667
17.1 | 54.55 | 57.4 | 58.7333 | 58.6667 | 58.4333
17.2 | 54.24 | 57.1167 | 58.5 | 58.4 | 58.22
17.3 | 54.025 | 56.9167 | 58.3 | 58.18 | 58
17.4 | 53.8 | 56.6333 | 58.08 | 58.06 | 57.8667
17.5 | 53.5 | 56.425 | 57.76 | 57.75 | 57.6
17.6 | 53.2667 | 56.22 | 57.6 | 57.5 | 57.3333
17.7 | 53.08 | 56.02 | 57.38 | 57.2667 | 57.1
17.8 | 52.7714 | 55.74 | 57.1 | 57.06 | 56.9
17.9 | 52.6167 | 55.4333 | 56.9 | 56.925 | 56.8
18 | 52.375 | 55.225 | 56.62 | 56.64 | 56.5833
18.1 | 52.16 | 54.98 | 56.5 | 56.42 | 56.3333
18.2 | 51.975 | 54.7667 | 56.25 | 56.2 | 56.2
18.3 | 51.7 | 54.56 | 56.1 | 55.9 | 55.9
18.4 | 51.45 | 54.25 | 55.75 | 55.72 | 55.74
18.5 | 51.1667 | 54.04 | 55.525 | 55.56 | 55.45
18.6 | 50.9 | 53.85 | 55.3333 | 55.3833 | 55.3
18.7 | 50.66 | 53.575 | 55.12 | 55.1667 | 55.0333
18.8 | 50.4833 | 53.4 | 54.94 | 54.9 | 54.8167
18.9 | 50.2 | 53.2333 | 54.76 | 54.7167 | 54.55
19 | 50.0167 | 52.9833 | 54.525 | 54.5333 | 54.4
19.1 | 49.7167 | 52.7333 | 54.3333 | 54.32 | 54.25
19.2 | 49.5 | 52.4857 | 54.12 | 54.1 | 54.05
19.3 | 49.3 | 52.24 | 53.9167 | 53.9 | 53.8333
19.4 | 49.1 | 52.075 | 53.65 | 53.675 | 53.5714
19.5 | 48.875 | 51.825 | 53.5 | 53.5 | 53.4167
19.6 | 48.65 | 51.5 | 53.22 | 53.24 | 53.2
19.7 | 48.4 | 51.375 | 52.9857 | 53.02 | 53
19.8 | 48.2167 | 51.14 | 52.85 | 52.8333 | 52.825
19.9 | 48 | 50.92 | 52.56 | 52.66 | 52.5
20 | 47.74 | 50.6667 | 52.4 | 52.45 | 52.3
20.1 | 47.6 | 50.52 | 52.14 | 52.225 | 52.125
20.2 | 47.4 | 50.325 | 51.9 | 52.05 | 52
20.3 | 47.1 | 50.1 | 51.76 | 51.85 | 51.775
20.4 | 46.9 | 49.875 | 51.4833 | 51.7 | 51.6
20.5 | 46.6833 | 49.625 | 51.275 | 51.5 | 51.4167
20.6 | 46.5 | 49.4 | 51.04 | 51.2333 | 51.1
20.7 | 46.225 | 49.18 | 50.9 | 51.02 | 50.98
20.8 | 45.975 | 48.96 | 50.68 | 50.825 | 50.85
20.9 | 45.8 | 48.82 | 50.46 | 50.6 | 50.6
21 | 45.6 | 48.6 | 50.32 | 50.4333 | 50.45
21.1 | 45.4833 | 48.4 | 50.1333 | 50.1667 | 50.2
21.2 | 45.3 | 48.2 | 49.925 | 50.0833 | 50.025
21.3 | 45.1 | 47.9429 | 49.7 | 49.9 | 49.84
21.4 | 44.9 | 47.8 | 49.6 | 49.7333 | 49.6
21.5 | 44.6667 | 47.6 | 49.4 | 49.5333 | 49.4333
21.6 | 44.4 | 47.4 | 49.2 | 49.26 | 49.25
21.7 | 44.25 | 47.175 | 48.9333 | 49.04 | 49.1
21.8 | 44 | 46.95 | 48.72 | 48.8714 | 48.86
21.9 | 43.82 | 46.78 | 48.5 | 48.7 | 48.66
22 | 43.64 | 46.6 | 48.3 | 48.6 | 48.5
22.1 | 43.425 | 46.4 | 48.1 | 48.4 | 48.3
22.2 | 43.175 | 46.2 | 47.9167 | 48.2 | 48.0833
22.3 | 43.05 | 46 | 47.72 | 48 | 47.975
22.4 | 42.85 | 45.8167 | 47.52 | 47.78 | 47.7833
22.5 | 42.64 | 45.6 | 47.4 | 47.6 | 47.6167
22.6 | 42.5333 | 45.4 | 47.15 | 47.45 | 47.475
22.7 | 42.25 | 45.2 | 47 | 47.3 | 47.2667
22.8 | 42.0833 | 45 | 46.8 | 47.05 | 47.05
22.9 | 41.975 | 44.82 | 46.7 | 46.9 | 46.85
23 | 41.8 | 44.625 | 46.5167 | 46.74 | 46.62
23.1 | 41.55 | 44.4333 | 46.24 | 46.56 | 46.48
23.2 | 41.3333 | 44.24 | 46.025 | 46.34 | 46.3
23.3 | 41.18 | 44.04 | 45.84 | 46.28 | 46.12
23.4 | 40.9286 | 43.9 | 45.7 | 46.0333 | 45.9167
23.5 | 40.8 | 43.75 | 45.5 | 45.8333 | 45.7
23.6 | 40.6 | 43.525 | 45.3333 | 45.62 | 45.6
23.7 | 40.425 | 43.35 | 45.2 | 45.5 | 45.42
23.8 | 40.2 | 43.15 | 45.02 | 45.2 | 45.3333
23.9 | 40.0667 | 43.05 | 44.82 | 45.1 | 45.14
24 | 39.9 | 42.8 | 44.6 | 45 | 44.9
24.1 | 39.8 | 42.675 | 44.4167 | 44.75 | 44.7833
24.2 | 39.6 | 42.44 | 44.3 | 44.6 | 44.5571
24.3 | 39.2667 | 42.225 | 44.0833 | 44.45 | 44.46
24.4 | 39.08 | 42.05 | 43.9167 | 44.3 | 44.26
24.5 | 38.96 | 41.9 | 43.7667 | 44.1333 | 44.06
24.6 | 38.675 | 41.7 | 43.6 | 43.9 | 43.96
24.7 | 38.6 | 41.55 | 43.4 | 43.7 | 43.75
24.8 | 38.38 | 41.35 | 43.275 | 43.58 | 43.575
24.9 | 38.3 | 41.1333 | 43.1 | 43.4143 | 43.34
25 | 38.05 | 41.025 | 42.9333 | 43.2 | 43.225
25.1 | 37.9 | 40.85 | 42.8 | 43.075 | 43.1
25.2 | 37.7833 | 40.625 | 42.55 | 42.825 | 42.9
25.3 | 37.6 | 40.475 | 42.4 | 42.775 | 42.75
25.4 | 37.46 | 40.2833 | 42.18 | 42.525 | 42.6
25.5 | 37.225 | 40.1 | 42.0833 | 42.4 | 42.4333
25.6 | 37.075 | 39.95 | 41.875 | 42.225 | 42.3
25.7 | 36.9 | 39.82 | 41.76 | 42 | 42.1
25.8 | 36.6833 | 39.6 | 41.55 | 41.9 | 41.95
25.9 | 36.6 | 39.5 | 41.35 | 41.8 | 41.8
26 | 36.5 | 39.325 | 41.24 | 41.6 | 41.575
26.1 | 36.38 | 39.15 | 41.125 | 41.44 | 41.44
26.2 | 36.2 | 38.975 | 40.9333 | 41.3333 | 41.28
26.3 | 35.975 | 38.85 | 40.74 | 41.16 | 41.1
26.4 | 35.8 | 38.6 | 40.6 | 41 | 41
26.5 | 35.6667 | 38.46 | 40.44 | 40.85 | 40.775
26.6 | 35.5 | 38.3 | 40.3 | 40.66 | 40.7
26.7 | 35.35 | 38.16 | 40.1 | 40.5 | 40.5
26.8 | 35.175 | 38.025 | 39.95 | 40.4 | 40.38
26.9 | 35 | 37.86 | 39.775 | 40.2 | 40.1833
27 | 34.86 | 37.74 | 39.54 | 40.06 | 40.0167
27.1 | 34.74 | 37.56 | 39.45 | 39.8667 | 39.9
27.2 | 34.55 | 37.36 | 39.3 | 39.7 | 39.76
27.3 | 34.4667 | 37.2 | 39.2167 | 39.6 | 39.575
27.4 | 34.2667 | 37.1 | 39 | 39.42 | 39.45
27.5 | 34.1 | 37 | 38.8 | 39.3 | 39.2667
27.6 | 33.95 | 36.82 | 38.7 | 39.1 | 39.2
27.7 | 33.8 | 36.62 | 38.575 | 39 | 39.0143
27.8 | 33.7 | 36.38 | 38.4 | 38.82 | 38.9167
27.9 | 33.6 | 36.3 | 38.225 | 38.7 | 38.725
28 | 33.42 | 36.15 | 38.1 | 38.5 | 38.54
28.1 | 33.2 | 35.9667 | 37.9333 | 38.4 | 38.44
28.2 | 33.02 | 35.9 | 37.72 | 38.3 | 38.26
28.3 | 32.8833 | 35.6667 | 37.64 | 38.1 | 38.075
28.4 | 32.8 | 35.5 | 37.45 | 37.9333 | 37.96
28.5 | 32.675 | 35.4333 | 37.3667 | 37.825 | 37.825
28.6 | 32.52 | 35.3 | 37.2 | 37.7 | 37.7
28.7 | 32.35 | 35.1333 | 37.02 | 37.5 | 37.5
28.8 | 32.2 | 35 | 36.82 | 37.3333 | 37.4
28.9 | 32.0333 | 34.8333 | 36.74 | 37.2 | 37.24
29 | 31.9667 | 34.6667 | 36.6 | 37.1 | 37.1
29.1 | 31.875 | 34.56 | 36.5 | 37 | 37
29.2 | 31.65 | 34.3833 | 36.34 | 36.825 | 36.8
29.3 | 31.44 | 34.2 | 36.12 | 36.675 | 36.7
29.4 | 31.32 | 34.1 | 36.04 | 36.4833 | 36.55
29.5 | 31.2 | 34 | 35.9 | 36.4 | 36.45
29.6 | 31.1 | 33.8 | 35.7667 | 36.24 | 36.3
29.7 | 31 | 33.6 | 35.6333 | 36.0333 | 36.2
29.8 | 30.825 | 33.48 | 35.44 | 35.96 | 36.05
29.9 | 30.7 | 33.375 | 35.35 | 35.85 | 35.9
30 | 30.5333 | 33.24 | 35.18 | 35.7 | 35.7

## SSD80 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 25.6 | 35.7 | 54.7 | 69.4 | 74.3
0.1 | 29.3125 | 39.9875 | 58.6 | 73.15 | 77.925
0.2 | 35.8 | 45.65 | 64.1667 | 77.8 | 82.08
0.3 | 44.3 | 54.1 | 71.04 | 83.35 | 87.0111
0.4 | 52.1 | 61.05 | 76.525 | 87.5 | 90.4375
0.5 | 59 | 66.825 | 81.25 | 90.8667 | 93.3125
0.6 | 64.82 | 72.2 | 85.15 | 93.34 | 95.4125
0.7 | 69.95 | 76.46 | 88.35 | 95.4 | 97.0125
0.8 | 74.5 | 80.25 | 90.85 | 96.725 | 98.1
0.9 | 78.3 | 83.45 | 92.8 | 97.9 | 99.0167
1 | 81.48 | 86.175 | 94.5 | 98.675 | 99.55
1.1 | 84.35 | 88.6 | 95.8 | 99.35 | 99.8
1.2 | 86.8 | 90.6 | 96.825 | 99.65 | 100
1.3 | 89 | 92.3 | 97.825 | 99.875 | 100
1.4 | 90.9 | 93.7 | 98.5125 | 100 | 100
1.5 | 92.4714 | 95 | 99.05 | 100 | 99.8667
1.6 | 93.9889 | 96 | 99.4375 | 100 | 99.725
1.7 | 95.1 | 96.95 | 99.7667 | 99.9 | 99.5
1.8 | 96.1875 | 97.7 | 99.9571 | 99.7 | 99.15
1.9 | 97.1571 | 98.32 | 100 | 99.4 | 98.85
2 | 97.8 | 98.8 | 100 | 99.25 | 98.5
2.1 | 98.5 | 99.3 | 99.975 | 98.9 | 98.225
2.2 | 99.04 | 99.525 | 99.9 | 98.6 | 97.725
2.3 | 99.3333 | 99.775 | 99.75 | 98.2 | 97.4
2.4 | 99.65 | 99.9 | 99.6143 | 97.85 | 97.05
2.5 | 99.8667 | 99.9778 | 99.375 | 97.5 | 96.7
2.6 | 99.9667 | 100 | 99.125 | 97.1 | 96.2
2.7 | 100 | 100 | 98.8714 | 96.7 | 95.8
2.8 | 100 | 99.95 | 98.5857 | 96.32 | 95.375
2.9 | 99.9 | 99.7857 | 98.26 | 95.9571 | 94.9
3 | 99.9 | 99.6429 | 97.925 | 95.54 | 94.55
3.1 | 99.8 | 99.46 | 97.575 | 95.05 | 94.1
3.2 | 99.575 | 99.2 | 97.325 | 94.675 | 93.8
3.3 | 99.3667 | 98.975 | 96.9 | 94.3 | 93.35
3.4 | 99.2 | 98.725 | 96.5375 | 93.85 | 92.88
3.5 | 98.9 | 98.5 | 96.2 | 93.4625 | 92.525
3.6 | 98.66 | 98.15 | 95.7571 | 93.125 | 92.1
3.7 | 98.3 | 97.8 | 95.4 | 92.7 | 91.725
3.8 | 97.9 | 97.55 | 95 | 92.3333 | 91.35
3.9 | 97.575 | 97.1 | 94.6 | 91.95 | 90.975
4 | 97.24 | 96.75 | 94.25 | 91.54 | 90.6
4.1 | 96.85 | 96.4 | 93.8 | 91.1556 | 90.225
4.2 | 96.475 | 96.1 | 93.475 | 90.7333 | 89.8222
4.3 | 96.1375 | 95.64 | 93.05 | 90.4 | 89.4125
4.4 | 95.6625 | 95.175 | 92.75 | 90 | 89
4.5 | 95.25 | 94.9 | 92.25 | 89.6 | 88.6571
4.6 | 94.8 | 94.45 | 91.875 | 89.2333 | 88.3
4.7 | 94.35 | 94.0571 | 91.575 | 88.88 | 87.9
4.8 | 93.85 | 93.7 | 91.15 | 88.5 | 87.6
4.9 | 93.425 | 93.3143 | 90.7667 | 88.12 | 87.2
5 | 92.95 | 92.9 | 90.35 | 87.7 | 86.8
5.1 | 92.44 | 92.45 | 89.85 | 87.275 | 86.5
5.2 | 92.1 | 92.0667 | 89.6 | 86.95 | 86.14
5.3 | 91.7 | 91.6143 | 89.2 | 86.575 | 85.775
5.4 | 91.2 | 91.1875 | 88.9 | 86.2 | 85.44
5.5 | 90.85 | 90.8 | 88.45 | 85.8 | 85.1
5.6 | 90.35 | 90.4 | 88.0667 | 85.5 | 84.6667
5.7 | 89.85 | 89.9 | 87.65 | 85.1 | 84.4
5.8 | 89.425 | 89.55 | 87.275 | 84.825 | 83.9333
5.9 | 88.9143 | 89.1 | 86.8 | 84.45 | 83.6
6 | 88.6 | 88.7667 | 86.525 | 84.125 | 83.25
6.1 | 88.14 | 88.35 | 86.1667 | 83.775 | 83
6.2 | 87.7714 | 87.95 | 85.75 | 83.35 | 82.65
6.3 | 87.3 | 87.525 | 85.425 | 83.075 | 82.325
6.4 | 86.9 | 87.15 | 84.95 | 82.75 | 81.9
6.5 | 86.45 | 86.725 | 84.6333 | 82.4 | 81.5
6.6 | 85.9 | 86.4 | 84.3 | 82.05 | 81.225
6.7 | 85.4333 | 85.95 | 83.95 | 81.65 | 80.88
6.8 | 85.025 | 85.525 | 83.66 | 81.3 | 80.58
6.9 | 84.675 | 85.15 | 83.25 | 81 | 80.2
7 | 84.25 | 84.7333 | 82.9 | 80.6 | 79.8
7.1 | 83.775 | 84.3375 | 82.55 | 80.35 | 79.55
7.2 | 83.4 | 83.92 | 82.15 | 79.9 | 79.225
7.3 | 82.9 | 83.625 | 81.775 | 79.65 | 78.86
7.4 | 82.4 | 83.16 | 81.425 | 79.32 | 78.55
7.5 | 82.12 | 82.85 | 81.1 | 78.875 | 78.2
7.6 | 81.75 | 82.4 | 80.7333 | 78.58 | 77.9
7.7 | 81.35 | 82.05 | 80.3667 | 78.3 | 77.6
7.8 | 80.95 | 81.65 | 80 | 78.05 | 77.32
7.9 | 80.625 | 81.35 | 79.725 | 77.7 | 76.88
8 | 80.0625 | 81.05 | 79.4 | 77.3 | 76.6
8.1 | 79.725 | 80.5667 | 79.04 | 77 | 76.325
8.2 | 79.3 | 80.2 | 78.6 | 76.7 | 75.925
8.3 | 79 | 79.9 | 78.35 | 76.3625 | 75.65
8.4 | 78.5143 | 79.4 | 78 | 76.025 | 75.4
8.5 | 78.1 | 79 | 77.6 | 75.675 | 74.95
8.6 | 77.7 | 78.65 | 77.25 | 75.35 | 74.6333
8.7 | 77.3 | 78.325 | 76.925 | 75.1 | 74.3667
8.8 | 76.9333 | 77.875 | 76.6625 | 74.725 | 74.1
8.9 | 76.45 | 77.6 | 76.25 | 74.3857 | 73.7333
9 | 76.075 | 77.2375 | 75.98 | 74.08 | 73.46
9.1 | 75.7667 | 76.9714 | 75.65 | 73.825 | 73.22
9.2 | 75.35 | 76.6 | 75.3 | 73.525 | 72.875
9.3 | 75.1 | 76.1 | 75 | 73.15 | 72.525
9.4 | 74.6 | 75.8 | 74.65 | 72.9 | 72.25
9.5 | 74.2 | 75.4 | 74.3 | 72.5375 | 71.94
9.6 | 73.925 | 75.1 | 73.95 | 72.1857 | 71.7
9.7 | 73.475 | 74.6667 | 73.62 | 71.9 | 71.25
9.8 | 73.15 | 74.325 | 73.3333 | 71.5857 | 71
9.9 | 72.675 | 73.975 | 73 | 71.3 | 70.7
10 | 72.425 | 73.625 | 72.675 | 70.95 | 70.3
10.1 | 71.95 | 73.175 | 72.325 | 70.6 | 70.05
10.2 | 71.55 | 72.95 | 72.05 | 70.3 | 69.7
10.3 | 71.3 | 72.6 | 71.75 | 70.02 | 69.45
10.4 | 70.95 | 72.25 | 71.4 | 69.775 | 69.22
10.5 | 70.7 | 71.92 | 71.05 | 69.425 | 68.875
10.6 | 70.3 | 71.525 | 70.75 | 69.1 | 68.55
10.7 | 69.9 | 71.25 | 70.4667 | 68.85 | 68.225
10.8 | 69.52 | 70.825 | 70.2 | 68.6 | 67.9857
10.9 | 69.2 | 70.5 | 69.8 | 68.25 | 67.68
11 | 68.75 | 70.25 | 69.4 | 68 | 67.4
11.1 | 68.3375 | 70 | 69.25 | 67.625 | 67.14
11.2 | 68.0667 | 69.6 | 68.925 | 67.35 | 66.8
11.3 | 67.65 | 69.3 | 68.575 | 67.15 | 66.55
11.4 | 67.5 | 69 | 68.325 | 66.85 | 66.3
11.5 | 67.14 | 68.6 | 68.06 | 66.575 | 66.06
11.6 | 66.7 | 68.2 | 67.65 | 66.325 | 65.7714
11.7 | 66.4 | 68.02 | 67.4 | 65.975 | 65.45
11.8 | 66.12 | 67.625 | 67.1 | 65.725 | 65.225
11.9 | 65.85 | 67.275 | 66.85 | 65.4125 | 64.95
12 | 65.425 | 66.925 | 66.5 | 65.125 | 64.7
12.1 | 65.15 | 66.675 | 66.3 | 64.9 | 64.4
12.2 | 64.75 | 66.325 | 65.9 | 64.625 | 64.1
12.3 | 64.48 | 66.1 | 65.66 | 64.2667 | 63.8
12.4 | 64.1333 | 65.7 | 65.3 | 63.9 | 63.56
12.5 | 63.84 | 65.4 | 65.075 | 63.76 | 63.325
12.6 | 63.5 | 65.1 | 64.825 | 63.525 | 63.05
12.7 | 63.15 | 64.8 | 64.475 | 63.275 | 62.725
12.8 | 62.9 | 64.55 | 64.225 | 62.96 | 62.475
12.9 | 62.525 | 64.225 | 63.88 | 62.725 | 62.225
13 | 62.175 | 63.875 | 63.6 | 62.375 | 61.95
13.1 | 61.925 | 63.525 | 63.4 | 62.1143 | 61.65
13.2 | 61.55 | 63.2429 | 63.05 | 61.95 | 61.4
13.3 | 61.15 | 62.9667 | 62.7 | 61.6 | 61.15
13.4 | 60.88 | 62.75 | 62.45 | 61.35 | 60.9
13.5 | 60.65 | 62.3333 | 62.2 | 61.1 | 60.65
13.6 | 60.4 | 62.1 | 61.9333 | 60.86 | 60.42
13.7 | 60.025 | 61.8 | 61.65 | 60.525 | 60.175
13.8 | 59.775 | 61.5 | 61.4 | 60.3 | 59.925
13.9 | 59.4375 | 61.2333 | 61.15 | 60.1 | 59.675
14 | 59.125 | 60.8875 | 60.84 | 59.875 | 59.325
14.1 | 58.7833 | 60.575 | 60.575 | 59.525 | 59.2
14.2 | 58.6286 | 60.3667 | 60.325 | 59.275 | 58.8
14.3 | 58.24 | 60.075 | 60.075 | 59.05 | 58.6
14.4 | 58 | 59.7 | 59.825 | 58.8 | 58.3667
14.5 | 57.65 | 59.45 | 59.5 | 58.5 | 58.1
14.6 | 57.4 | 59.2 | 59.225 | 58.25 | 57.9
14.7 | 57.15 | 58.9 | 58.98 | 58 | 57.6
14.8 | 56.8 | 58.6 | 58.74 | 57.75 | 57.3
14.9 | 56.52 | 58.3667 | 58.4 | 57.475 | 57.1
15 | 56.225 | 58.125 | 58.25 | 57.225 | 56.825
15.1 | 55.875 | 57.85 | 57.9 | 56.975 | 56.7
15.2 | 55.625 | 57.54 | 57.65 | 56.725 | 56.4
15.3 | 55.375 | 57.3 | 57.4 | 56.5 | 56.15
15.4 | 55.14 | 57 | 57.175 | 56.25 | 55.8
15.5 | 54.9 | 56.7 | 56.925 | 56 | 55.7
15.6 | 54.55 | 56.45 | 56.675 | 55.825 | 55.4
15.7 | 54.3 | 56.2 | 56.4 | 55.5 | 55.175
15.8 | 54.05 | 55.875 | 56.15 | 55.275 | 55
15.9 | 53.8 | 55.625 | 55.9 | 55.0857 | 54.7667
16 | 53.4 | 55.4333 | 55.675 | 54.8625 | 54.525
16.1 | 53.225 | 55.15 | 55.425 | 54.5625 | 54.3
16.2 | 52.975 | 54.9 | 55.1857 | 54.375 | 54.0333
16.3 | 52.725 | 54.7 | 55 | 54.125 | 53.8
16.4 | 52.475 | 54.425 | 54.6667 | 54 | 53.6
16.5 | 52.25 | 54.1333 | 54.4 | 53.75 | 53.4
16.6 | 51.9375 | 53.875 | 54.25 | 53.5 | 53.15
16.7 | 51.7 | 53.625 | 53.96 | 53.175 | 52.9
16.8 | 51.425 | 53.38 | 53.725 | 53 | 52.675
16.9 | 51.1 | 53.1 | 53.475 | 52.775 | 52.425
17 | 50.85 | 52.825 | 53.25 | 52.525 | 52.25
17.1 | 50.675 | 52.575 | 53 | 52.3 | 52.05
17.2 | 50.425 | 52.325 | 52.8 | 52.05 | 51.8
17.3 | 50.15 | 52.0889 | 52.5 | 51.9 | 51.56
17.4 | 49.92 | 51.925 | 52.375 | 51.6 | 51.325
17.5 | 49.7 | 51.6 | 52.125 | 51.45 | 51.16
17.6 | 49.45 | 51.35 | 51.875 | 51.275 | 51
17.7 | 49.1 | 51.15 | 51.65 | 51.025 | 50.8
17.8 | 48.9 | 50.95 | 51.4 | 50.775 | 50.5
17.9 | 48.66 | 50.66 | 51.225 | 50.525 | 50.3
18 | 48.44 | 50.4 | 50.9571 | 50.3667 | 50.05
18.1 | 48.275 | 50.175 | 50.7667 | 50.2 | 49.9
18.2 | 48.025 | 49.925 | 50.55 | 49.95 | 49.675
18.3 | 47.775 | 49.7 | 50.2 | 49.7 | 49.425
18.4 | 47.525 | 49.525 | 50.05 | 49.475 | 49.2167
18.5 | 47.32 | 49.3 | 49.85 | 49.3 | 49.075
18.6 | 47.1 | 49.075 | 49.6286 | 49.125 | 48.8625
18.7 | 46.8333 | 48.85 | 49.475 | 48.8333 | 48.6
18.8 | 46.56 | 48.6 | 49.2 | 48.6833 | 48.45
18.9 | 46.4 | 48.35 | 49 | 48.475 | 48.2333
19 | 46.1 | 48.2 | 48.8 | 48.3 | 48
19.1 | 45.85 | 47.92 | 48.55 | 48.075 | 47.8
19.2 | 45.7 | 47.625 | 48.4 | 47.85 | 47.6
19.3 | 45.475 | 47.4 | 48.15 | 47.7 | 47.35
19.4 | 45.225 | 47.225 | 47.9 | 47.44 | 47.22
19.5 | 44.975 | 46.975 | 47.675 | 47.2 | 46.975
19.6 | 44.825 | 46.75 | 47.5 | 47 | 46.825
19.7 | 44.6 | 46.65 | 47.275 | 46.8 | 46.675
19.8 | 44.325 | 46.3333 | 47.1 | 46.625 | 46.425
19.9 | 44.1333 | 46.125 | 46.875 | 46.425 | 46.175
20 | 44 | 45.98 | 46.64 | 46.2 | 45.94
20.1 | 43.7 | 45.725 | 46.5 | 45.95 | 45.8
20.2 | 43.5 | 45.5 | 46.25 | 45.8 | 45.55
20.3 | 43.3 | 45.25 | 46 | 45.6 | 45.4
20.4 | 43.1 | 45.1 | 45.85 | 45.4 | 45.16
20.5 | 42.85 | 44.85 | 45.6 | 45.2 | 45.0143
20.6 | 42.7 | 44.7 | 45.375 | 45.025 | 44.8333
20.7 | 42.35 | 44.45 | 45.2 | 44.875 | 44.625
20.8 | 42.2 | 44.2 | 45.0143 | 44.6 | 44.46
20.9 | 42 | 44.05 | 44.7833 | 44.45 | 44.3
21 | 41.82 | 43.8 | 44.5833 | 44.3 | 44.1
21.1 | 41.675 | 43.6 | 44.4375 | 44.075 | 43.925
21.2 | 41.425 | 43.4 | 44.225 | 43.925 | 43.78
21.3 | 41.275 | 43.225 | 44.0857 | 43.75 | 43.55
21.4 | 41.025 | 42.975 | 43.875 | 43.525 | 43.4
21.5 | 40.85 | 42.78 | 43.625 | 43.4 | 43.2
21.6 | 40.6 | 42.55 | 43.475 | 43.1333 | 43
21.7 | 40.4 | 42.4 | 43.225 | 43 | 42.78
21.8 | 40.25 | 42.25 | 43.1 | 42.8 | 42.5833
21.9 | 40.02 | 42.02 | 42.9 | 42.6 | 42.4375
22 | 39.825 | 41.8667 | 42.8 | 42.38 | 42.25
22.1 | 39.6 | 41.7 | 42.5 | 42.25 | 42.1
22.2 | 39.4 | 41.475 | 42.35 | 42.1 | 41.925
22.3 | 39.3 | 41.2667 | 42.2 | 41.95 | 41.8
22.4 | 39.1 | 41.1 | 41.9667 | 41.7 | 41.575
22.5 | 38.9 | 40.9 | 41.75 | 41.6 | 41.425
22.6 | 38.8 | 40.7667 | 41.6 | 41.36 | 41.2
22.7 | 38.7 | 40.525 | 41.42 | 41.2 | 41.04
22.8 | 38.35 | 40.3 | 41.275 | 40.975 | 40.9
22.9 | 38.15 | 40.125 | 41.1 | 40.825 | 40.64
23 | 37.95 | 39.975 | 40.875 | 40.675 | 40.5
23.1 | 37.875 | 39.8 | 40.625 | 40.525 | 40.35
23.2 | 37.625 | 39.58 | 40.4 | 40.3 | 40.2
23.3 | 37.475 | 39.35 | 40.3 | 40.15 | 40
23.4 | 37.3 | 39.2 | 40.2 | 40 | 39.9
23.5 | 37.1 | 39.05 | 39.9667 | 39.85 | 39.6625
23.6 | 37 | 38.84 | 39.825 | 39.6 | 39.5375
23.7 | 36.7 | 38.7 | 39.675 | 39.46 | 39.325
23.8 | 36.55 | 38.5 | 39.425 | 39.3 | 39.1889
23.9 | 36.42 | 38.4 | 39.2875 | 39.175 | 39.05
24 | 36.2 | 38.1 | 39.125 | 38.95 | 38.85
24.1 | 36.0667 | 37.9 | 38.9375 | 38.7833 | 38.725
24.2 | 35.82 | 37.8 | 38.75 | 38.6833 | 38.5
24.3 | 35.675 | 37.6 | 38.6 | 38.5167 | 38.35
24.4 | 35.525 | 37.425 | 38.44 | 38.275 | 38.2
24.5 | 35.3 | 37.275 | 38.28 | 38.14 | 38.075
24.6 | 35.2 | 37.125 | 38.15 | 37.975 | 37.9125
24.7 | 35 | 36.9 | 37.98 | 37.84 | 37.6875
24.8 | 34.9 | 36.8 | 37.7667 | 37.6667 | 37.5625
24.9 | 34.65 | 36.6 | 37.6667 | 37.5 | 37.4444
25 | 34.5 | 36.4 | 37.45 | 37.3 | 37.25
25.1 | 34.3333 | 36.3 | 37.32 | 37.225 | 37.0667
25.2 | 34.15 | 36.1 | 37.1 | 37.05 | 36.925
25.3 | 34.02 | 35.9 | 37 | 36.9 | 36.75
25.4 | 33.8667 | 35.75 | 36.85 | 36.775 | 36.65
25.5 | 33.7 | 35.6 | 36.6 | 36.525 | 36.5
25.6 | 33.5714 | 35.4 | 36.45 | 36.4 | 36.3714
25.7 | 33.3571 | 35.24 | 36.3 | 36.25 | 36.12
25.8 | 33.2 | 35.1 | 36.1 | 36.025 | 36.05
25.9 | 33 | 35 | 36.02 | 35.9 | 35.9
26 | 32.9 | 34.8 | 35.875 | 35.825 | 35.7
26.1 | 32.78 | 34.675 | 35.7 | 35.675 | 35.5333
26.2 | 32.55 | 34.5 | 35.5 | 35.54 | 35.475
26.3 | 32.4 | 34.325 | 35.375 | 35.4 | 35.3125
26.4 | 32.275 | 34.1 | 35.225 | 35.15 | 35.1
26.5 | 32.125 | 34 | 35.075 | 35 | 34.95
26.6 | 31.975 | 33.9 | 34.94 | 34.9 | 34.8
26.7 | 31.825 | 33.6667 | 34.775 | 34.8 | 34.7
26.8 | 31.6667 | 33.52 | 34.625 | 34.6 | 34.6
26.9 | 31.5 | 33.4 | 34.5 | 34.475 | 34.4
27 | 31.3 | 33.225 | 34.35 | 34.2625 | 34.3
27.1 | 31.2 | 33.1 | 34.2 | 34.1375 | 34.1
27.2 | 31.125 | 32.9 | 34.05 | 34.0125 | 33.925
27.3 | 30.975 | 32.75 | 33.9 | 33.9 | 33.88
27.4 | 30.8 | 32.6333 | 33.75 | 33.675 | 33.65
27.5 | 30.6 | 32.5 | 33.6 | 33.55 | 33.6
27.6 | 30.5 | 32.34 | 33.4 | 33.5 | 33.45
27.7 | 30.375 | 32.18 | 33.3 | 33.35 | 33.3
27.8 | 30.2 | 32.0333 | 33.175 | 33.2 | 33.175
27.9 | 30.1 | 31.9 | 33 | 33.075 | 33.025
28 | 29.9 | 31.725 | 32.8 | 32.9 | 32.875
28.1 | 29.8 | 31.6 | 32.7 | 32.8 | 32.725
28.2 | 29.675 | 31.5 | 32.6 | 32.7 | 32.6
28.3 | 29.5333 | 31.3 | 32.45 | 32.5 | 32.425
28.4 | 29.375 | 31.15 | 32.3 | 32.4 | 32.3
28.5 | 29.2 | 31.1 | 32.1667 | 32.22 | 32.225
28.6 | 29.1 | 30.9 | 32.025 | 32.1 | 32
28.7 | 29 | 30.725 | 31.9 | 32.025 | 31.9
28.8 | 28.8 | 30.6 | 31.75 | 31.875 | 31.8
28.9 | 28.65 | 30.5 | 31.62 | 31.7 | 31.6
29 | 28.6 | 30.375 | 31.45 | 31.6 | 31.5
29.1 | 28.4 | 30.2 | 31.3 | 31.4 | 31.4
29.2 | 28.3 | 30.075 | 31.2 | 31.3 | 31.2375
29.3 | 28.2125 | 29.925 | 31.1 | 31.1 | 31.1222
29.4 | 28.075 | 29.8 | 30.9 | 31 | 31
29.5 | 27.9 | 29.7 | 30.8 | 30.875 | 30.85
29.6 | 27.8375 | 29.575 | 30.6333 | 30.75 | 30.8
29.7 | 27.7125 | 29.4 | 30.5667 | 30.625 | 30.6
29.8 | 27.575 | 29.275 | 30.4 | 30.5 | 30.5
29.9 | 27.3625 | 29.15 | 30.3 | 30.3778 | 30.375
30 | 27.2 | 29.0125 | 30.175 | 30.2625 | 30.2

## SSD120 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 24.1 | 30.4 | 41 | 51 | 54.8
0.1 | 27.75 | 34.05 | 44.8333 | 54.525 | 58.7
0.2 | 34 | 40.4 | 50.4667 | 60 | 64
0.3 | 42.75 | 48.5 | 57.9 | 66.65 | 70.525
0.4 | 50.2 | 55.6 | 64.6 | 72.65 | 75.7
0.5 | 56.8 | 61.65 | 69.8 | 77 | 79.85
0.6 | 62.4 | 66.4444 | 74.5 | 81.3 | 83.6
0.7 | 67.3 | 71.2625 | 78.05 | 84.3333 | 86.65
0.8 | 71.6 | 74.9625 | 81.6 | 87.1 | 89.1
0.9 | 75.06 | 78.4625 | 84.275 | 89.35 | 91.25
1 | 78.45 | 81.675 | 86.85 | 91.5 | 93
1.1 | 81.425 | 84.2375 | 88.9 | 93.1 | 94.45
1.2 | 83.825 | 86.4667 | 90.9 | 94.6 | 95.8
1.3 | 86.275 | 88.4571 | 92.55 | 95.65 | 96.75
1.4 | 88.4 | 90.325 | 93.8 | 96.7 | 97.6
1.5 | 90.1 | 91.8429 | 95 | 97.45 | 98.25
1.6 | 91.8 | 93.26 | 96 | 98.16 | 98.78
1.7 | 92.9333 | 94.4 | 96.8 | 98.7 | 99.15
1.8 | 94.3 | 95.35 | 97.55 | 99.1 | 99.525
1.9 | 95.35 | 96.2 | 98.15 | 99.45 | 99.775
2 | 96.3 | 97.05 | 98.7 | 99.7333 | 99.9
2.1 | 97.05 | 97.5 | 99 | 99.8 | 99.9667
2.2 | 97.6 | 98.225 | 99.4 | 99.9 | 100
2.3 | 98.25 | 98.7 | 99.65 | 100 | 100
2.4 | 98.8 | 99.0667 | 99.8 | 100 | 99.975
2.5 | 99.125 | 99.4625 | 99.9 | 100 | 99.9
2.6 | 99.38 | 99.6 | 100 | 99.92 | 99.8
2.7 | 99.625 | 99.7714 | 100 | 99.8 | 99.6
2.8 | 99.875 | 99.9 | 100 | 99.7 | 99.5
2.9 | 100 | 99.9429 | 99.9 | 99.52 | 99.25
3 | 100 | 100 | 99.8 | 99.325 | 99.1
3.1 | 100 | 100 | 99.7 | 99.1 | 98.85
3.2 | 100 | 100 | 99.625 | 98.95 | 98.6
3.3 | 99.94 | 99.9 | 99.475 | 98.74 | 98.25
3.4 | 99.775 | 99.8 | 99.325 | 98.4 | 98.1
3.5 | 99.65 | 99.625 | 99.2 | 98.25 | 97.9
3.6 | 99.58 | 99.475 | 99.025 | 97.92 | 97.46
3.7 | 99.4 | 99.325 | 98.6 | 97.7 | 97.225
3.8 | 99.2 | 99.08 | 98.44 | 97.45 | 96.95
3.9 | 98.85 | 98.95 | 98.1 | 97.1 | 96.625
4 | 98.7 | 98.75 | 97.85 | 96.8 | 96.375
4.1 | 98.45 | 98.5 | 97.55 | 96.5 | 96.025
4.2 | 98.2 | 98.1 | 97.2 | 96.3 | 95.775
4.3 | 97.85 | 97.825 | 96.95 | 95.975 | 95.425
4.4 | 97.5 | 97.6 | 96.7 | 95.6 | 95.2
4.5 | 97.25 | 97.2429 | 96.35 | 95.45 | 94.85
4.6 | 96.82 | 96.8111 | 96.1 | 95.12 | 94.62
4.7 | 96.45 | 96.6 | 95.86 | 94.775 | 94.35
4.8 | 96.125 | 96.42 | 95.575 | 94.525 | 94.05
4.9 | 95.75 | 96.05 | 95.25 | 94.175 | 93.8
5 | 95.425 | 95.725 | 94.95 | 93.925 | 93.55
5.1 | 95.175 | 95.375 | 94.6 | 93.65 | 93.2
5.2 | 94.88 | 95.025 | 94.26 | 93.42 | 93
5.3 | 94.4 | 94.78 | 93.8333 | 93.05 | 92.55
5.4 | 94.0333 | 94.45 | 93.5 | 92.725 | 92.22
5.5 | 93.7 | 94 | 93.4 | 92.4667 | 91.95
5.6 | 93.4 | 93.75 | 93 | 92.05 | 91.625
5.7 | 93 | 93.4 | 92.66 | 91.6875 | 91.325
5.8 | 92.62 | 92.96 | 92.4 | 91.5143 | 91.1
5.9 | 92.25 | 92.6 | 92.075 | 91.25 | 90.8
6 | 91.7667 | 92.375 | 91.825 | 90.9 | 90.55
6.1 | 91.4 | 92.025 | 91.475 | 90.6333 | 90.3
6.2 | 91.2 | 91.75 | 91.225 | 90.35 | 89.95
6.3 | 90.8333 | 91.35 | 90.7 | 90.075 | 89.62
6.4 | 90.3 | 91.075 | 90.55 | 89.8 | 89.375
6.5 | 89.95 | 90.85 | 90.2 | 89.575 | 89.1
6.6 | 89.6 | 90.44 | 89.85 | 89.075 | 88.75
6.7 | 89.2 | 90.05 | 89.6 | 88.9 | 88.45
6.8 | 88.8 | 89.8 | 89.25 | 88.6 | 88.1
6.9 | 88.375 | 89.525 | 89 | 88.2 | 87.8
7 | 88.05 | 89.1 | 88.675 | 87.9 | 87.55
7.1 | 87.7833 | 88.8 | 88.425 | 87.72 | 87.32
7.2 | 87.425 | 88.5333 | 88.025 | 87.35 | 87
7.3 | 87.025 | 88.1 | 87.8 | 87.075 | 86.6333
7.4 | 86.775 | 87.9 | 87.45 | 86.75 | 86.4
7.5 | 86.45 | 87.525 | 87.3 | 86.525 | 86.1333
7.6 | 86 | 87.26 | 86.94 | 86.3 | 85.8
7.7 | 85.6 | 86.925 | 86.575 | 86 | 85.5
7.8 | 85.1 | 86.475 | 86.35 | 85.74 | 85.1667
7.9 | 84.72 | 86.125 | 85.975 | 85.4 | 85
8 | 84.46 | 85.85 | 85.68 | 85.15 | 84.8
8.1 | 84.125 | 85.525 | 85.4 | 84.84 | 84.45
8.2 | 83.875 | 85.16 | 85.15 | 84.6 | 84.175
8.3 | 83.525 | 84.85 | 84.9 | 84.275 | 83.95
8.4 | 83.075 | 84.48 | 84.55 | 83.9875 | 83.7
8.5 | 82.675 | 84.2 | 84.325 | 83.675 | 83.25
8.6 | 82.375 | 83.9 | 83.98 | 83.275 | 83
8.7 | 82.05 | 83.65 | 83.64 | 83.04 | 82.75
8.8 | 81.675 | 83.3 | 83.4 | 82.8 | 82.45
8.9 | 81.38 | 83 | 83.14 | 82.66 | 82.3
9 | 81.1 | 82.775 | 82.8 | 82.35 | 82
9.1 | 80.6 | 82.525 | 82.55 | 81.975 | 81.65
9.2 | 80.3 | 82.1 | 82.3 | 81.725 | 81.3
9.3 | 79.8 | 81.75 | 82 | 81.475 | 81.05
9.4 | 79.4 | 81.4333 | 81.7 | 81.225 | 80.9
9.5 | 79.2333 | 81.1 | 81.35 | 80.88 | 80.525
9.6 | 79 | 80.9 | 81.1 | 80.525 | 80.35
9.7 | 78.5333 | 80.4667 | 80.85 | 80.3 | 80
9.8 | 78.175 | 80.28 | 80.525 | 80.05 | 79.75
9.9 | 77.925 | 79.95 | 80.25 | 79.8 | 79.5
10 | 77.575 | 79.6 | 79.925 | 79.6 | 79.2
10.1 | 77.25 | 79.45 | 79.575 | 79.2 | 78.875
10.2 | 76.9 | 79.1 | 79.325 | 78.96 | 78.625
10.3 | 76.6 | 78.85 | 79.15 | 78.7143 | 78.375
10.4 | 76.3 | 78.5 | 78.825 | 78.3833 | 78.125
10.5 | 76.05 | 78.1667 | 78.575 | 78.175 | 77.85
10.6 | 75.6714 | 77.85 | 78.325 | 77.8429 | 77.64
10.7 | 75.12 | 77.5 | 78.05 | 77.675 | 77.4
10.8 | 74.8333 | 77.24 | 77.725 | 77.4625 | 77.1
10.9 | 74.5125 | 76.975 | 77.42 | 77.0875 | 76.85
11 | 74.225 | 76.68 | 77.1333 | 76.7125 | 76.5
11.1 | 74.075 | 76.35 | 76.9 | 76.5125 | 76.26
11.2 | 73.7714 | 76.1 | 76.5 | 76.35 | 76
11.3 | 73.38 | 75.8 | 76.3 | 76.15 | 75.825
11.4 | 73.15 | 75.4 | 76.05 | 75.7333 | 75.575
11.5 | 72.9 | 75.15 | 75.8 | 75.525 | 75.325
11.6 | 72.4 | 74.92 | 75.5 | 75.25 | 75.075
11.7 | 72.1 | 74.575 | 75.3 | 74.9 | 74.825
11.8 | 71.775 | 74.3 | 74.95 | 74.675 | 74.48
11.9 | 71.525 | 74.1 | 74.72 | 74.425 | 74.25
12 | 71.175 | 73.925 | 74.475 | 74.2 | 74
12.1 | 70.925 | 73.55 | 74.225 | 74.05 | 73.85
12.2 | 70.65 | 73.225 | 73.975 | 73.7667 | 73.4
12.3 | 70.325 | 72.875 | 73.7 | 73.5 | 73.2333
12.4 | 70.05 | 72.55 | 73.475 | 73.2333 | 73
12.5 | 69.725 | 72.275 | 73.14 | 72.9 | 72.725
12.6 | 69.4 | 72.025 | 72.9 | 72.7 | 72.46
12.7 | 69.0333 | 71.88 | 72.65 | 72.475 | 72.1
12.8 | 68.75 | 71.55 | 72.275 | 72.225 | 72
12.9 | 68.5 | 71.2 | 72.125 | 72 | 71.7
13 | 68.06 | 71 | 71.95 | 71.75 | 71.45
13.1 | 67.725 | 70.75 | 71.7 | 71.5 | 71.2
13.2 | 67.5 | 70.3 | 71.425 | 71.25 | 70.95
13.3 | 67.325 | 70.05 | 71.1125 | 71 | 70.7222
13.4 | 67.1 | 69.8 | 70.9222 | 70.75 | 70.56
13.5 | 66.7 | 69.55 | 70.6 | 70.45 | 70.225
13.6 | 66.3 | 69.3 | 70.3571 | 70.2 | 69.975
13.7 | 66.15 | 69.1286 | 70.1 | 69.9 | 69.725
13.8 | 65.9 | 68.8 | 69.85 | 69.775 | 69.4875
13.9 | 65.66 | 68.45 | 69.6 | 69.525 | 69.325
14 | 65.42 | 68.2111 | 69.3667 | 69.35 | 69.075
14.1 | 65.1333 | 67.975 | 69.15 | 69.025 | 68.9222
14.2 | 64.725 | 67.725 | 68.85 | 68.7 | 68.6
14.3 | 64.475 | 67.38 | 68.52 | 68.55 | 68.325
14.4 | 64.225 | 67.1286 | 68.3 | 68.28 | 68.2
14.5 | 64 | 66.8857 | 68.025 | 68 | 67.85
14.6 | 63.7333 | 66.7125 | 67.8 | 67.85 | 67.6429
14.7 | 63.2 | 66.42 | 67.55 | 67.58 | 67.425
14.8 | 63 | 66.125 | 67.325 | 67.35 | 67.15
14.9 | 62.7 | 65.9 | 67.075 | 67.1 | 66.925
15 | 62.4 | 65.575 | 66.825 | 67 | 66.675
15.1 | 62.25 | 65.225 | 66.62 | 66.7 | 66.425
15.2 | 62 | 65 | 66.325 | 66.4 | 66.2
15.3 | 61.5667 | 64.725 | 66.075 | 66.16 | 66
15.4 | 61.4 | 64.475 | 65.9 | 65.925 | 65.7
15.5 | 61.2 | 64.25 | 65.7 | 65.775 | 65.5
15.6 | 60.9 | 64 | 65.4667 | 65.525 | 65.2
15.7 | 60.6 | 63.7 | 65.2 | 65.25 | 64.975
15.8 | 60.275 | 63.5 | 64.9 | 64.975 | 64.82
15.9 | 60.025 | 63.3667 | 64.7 | 64.725 | 64.675
16 | 59.8 | 63.12 | 64.5 | 64.575 | 64.425
16.1 | 59.58 | 62.8 | 64.2 | 64.34 | 64.075
16.2 | 59.35 | 62.6 | 64.05 | 64.1 | 63.9
16.3 | 59.0167 | 62.2667 | 63.65 | 63.9 | 63.7
16.4 | 58.875 | 62.1 | 63.3833 | 63.6 | 63.55
16.5 | 58.625 | 61.85 | 63.2375 | 63.45 | 63.3
16.6 | 58.45 | 61.6 | 63.04 | 63.2 | 63.05
16.7 | 58.05 | 61.35 | 62.7833 | 62.95 | 62.8
16.8 | 57.8 | 61.1 | 62.6 | 62.75 | 62.66
16.9 | 57.55 | 60.9 | 62.4 | 62.6 | 62.4
17 | 57.3 | 60.6 | 62.2 | 62.375 | 62.275
17.1 | 57.05 | 60.36 | 61.975 | 62.125 | 62.025
17.2 | 56.82 | 60.125 | 61.74 | 61.875 | 61.775
17.3 | 56.675 | 59.975 | 61.5 | 61.625 | 61.525
17.4 | 56.425 | 59.725 | 61.3 | 61.5 | 61.38
17.5 | 56.075 | 59.4 | 61.1 | 61.225 | 61.2
17.6 | 55.84 | 59.225 | 60.775 | 61.0667 | 60.9
17.7 | 55.7 | 58.975 | 60.575 | 60.9 | 60.7
17.8 | 55.4 | 58.8 | 60.325 | 60.6 | 60.5
17.9 | 55.1 | 58.5 | 60.1 | 60.35 | 60.275
18 | 54.9333 | 58.2429 | 59.92 | 60.2 | 60.1
18.1 | 54.675 | 58.0375 | 59.7 | 59.95 | 59.9
18.2 | 54.5 | 57.85 | 59.575 | 59.8 | 59.65
18.3 | 54.2667 | 57.6 | 59.325 | 59.55 | 59.4
18.4 | 54.0375 | 57.375 | 59.1375 | 59.325 | 59.15
18.5 | 53.8444 | 57.2 | 58.925 | 59.1 | 59
18.6 | 53.6 | 56.9667 | 58.675 | 58.925 | 58.875
18.7 | 53.325 | 56.65 | 58.425 | 58.575 | 58.525
18.8 | 53.05 | 56.5 | 58.175 | 58.425 | 58.375
18.9 | 52.85 | 56.325 | 58.0222 | 58.275 | 58.1625
19 | 52.6 | 56.075 | 57.8 | 58 | 57.975
19.1 | 52.275 | 55.8 | 57.5429 | 57.8 | 57.725
19.2 | 52.0857 | 55.5333 | 57.3444 | 57.6 | 57.54
19.3 | 51.85 | 55.4 | 57.1286 | 57.5 | 57.275
19.4 | 51.675 | 55.1 | 56.9 | 57.1333 | 57.0667
19.5 | 51.5 | 54.975 | 56.6286 | 56.88 | 56.925
19.6 | 51.225 | 54.7 | 56.5 | 56.8 | 56.7571
19.7 | 51.075 | 54.475 | 56.275 | 56.6 | 56.55
19.8 | 50.9 | 54.225 | 56.0625 | 56.4 | 56.3333
19.9 | 50.7 | 54.08 | 55.875 | 56.22 | 56.125
20 | 50.35 | 53.85 | 55.75 | 56.1 | 55.9667
20.1 | 50.2 | 53.6 | 55.5 | 55.875 | 55.675
20.2 | 49.95 | 53.4 | 55.25 | 55.525 | 55.425
20.3 | 49.8 | 53.2 | 55 | 55.375 | 55.275
20.4 | 49.575 | 52.95 | 54.8 | 55.125 | 55.0625
20.5 | 49.425 | 52.8 | 54.7 | 54.975 | 54.9375
20.6 | 49.075 | 52.66 | 54.45 | 54.8 | 54.7444
20.7 | 48.95 | 52.5 | 54.2 | 54.7 | 54.55
20.8 | 48.7 | 52.275 | 54.1 | 54.4286 | 54.35
20.9 | 48.55 | 52.025 | 53.825 | 54.2875 | 54.2
21 | 48.3667 | 51.875 | 53.6 | 54.125 | 54.05
21.1 | 48.1 | 51.525 | 53.5 | 53.8889 | 53.8
21.2 | 47.9 | 51.375 | 53.3333 | 53.65 | 53.575
21.3 | 47.7667 | 51.225 | 53.2 | 53.5 | 53.4
21.4 | 47.5 | 51.05 | 52.95 | 53.35 | 53.175
21.5 | 47.25 | 50.825 | 52.7111 | 53.1 | 53.1
21.6 | 47 | 50.575 | 52.5375 | 52.95 | 52.88
21.7 | 46.75 | 50.4 | 52.3 | 52.75 | 52.7
21.8 | 46.6 | 50.2 | 52.0875 | 52.5667 | 52.5
21.9 | 46.4286 | 49.95 | 51.9 | 52.325 | 52.35
22 | 46.3 | 49.775 | 51.7 | 52.1 | 52.1
22.1 | 46.05 | 49.54 | 51.6 | 52 | 51.96
22.2 | 45.85 | 49.4 | 51.35 | 51.9 | 51.7
22.3 | 45.675 | 49.15 | 51.1 | 51.65 | 51.56
22.4 | 45.425 | 48.9 | 50.975 | 51.4 | 51.325
22.5 | 45.275 | 48.75 | 50.9 | 51.3 | 51.175
22.6 | 45.025 | 48.6 | 50.65 | 51 | 51
22.7 | 44.8 | 48.46 | 50.4 | 50.85 | 50.8
22.8 | 44.725 | 48.225 | 50.1667 | 50.62 | 50.625
22.9 | 44.6 | 48.05 | 50 | 50.575 | 50.4
23 | 44.3667 | 47.9 | 49.95 | 50.3 | 50.25
23.1 | 44.2 | 47.65 | 49.72 | 50.0833 | 50.1
23.2 | 44.025 | 47.4 | 49.55 | 50 | 49.95
23.3 | 43.775 | 47.2333 | 49.32 | 49.8 | 49.8
23.4 | 43.625 | 47 | 49.175 | 49.625 | 49.6
23.5 | 43.3333 | 46.9 | 48.925 | 49.4 | 49.425
23.6 | 43.18 | 46.775 | 48.7333 | 49.25 | 49.2
23.7 | 43.05 | 46.5 | 48.6 | 49.1 | 49.1
23.8 | 42.8 | 46.2833 | 48.4 | 48.95 | 48.975
23.9 | 42.65 | 46.1 | 48.1 | 48.7 | 48.66
24 | 42.5 | 46 | 48 | 48.55 | 48.5
24.1 | 42.3 | 45.7 | 47.9 | 48.375 | 48.35
24.2 | 42.1 | 45.575 | 47.7 | 48.2 | 48.2714
24.3 | 41.9 | 45.425 | 47.575 | 48 | 48
24.4 | 41.8 | 45.1875 | 47.4 | 47.8 | 47.875
24.5 | 41.6 | 45.04 | 47.18 | 47.68 | 47.7333
24.6 | 41.475 | 44.8875 | 47.1 | 47.5 | 47.5
24.7 | 41.2 | 44.74 | 46.9 | 47.3 | 47.36
24.8 | 41.0875 | 44.5429 | 46.675 | 47.1 | 47.2
24.9 | 40.925 | 44.3444 | 46.4625 | 47 | 46.9667
25 | 40.7 | 44.2 | 46.2875 | 46.825 | 46.775
25.1 | 40.6125 | 44.0286 | 46.125 | 46.675 | 46.675
25.2 | 40.4 | 43.9 | 45.9125 | 46.525 | 46.5
25.3 | 40.175 | 43.65 | 45.8 | 46.375 | 46.3
25.4 | 40.0429 | 43.5 | 45.6714 | 46.125 | 46.15
25.5 | 39.86 | 43.25 | 45.5 | 45.98 | 46
25.6 | 39.7 | 43.125 | 45.24 | 45.85 | 45.85
25.7 | 39.5 | 42.975 | 45.1 | 45.7 | 45.72
25.8 | 39.4 | 42.825 | 44.9 | 45.525 | 45.55
25.9 | 39.275 | 42.7 | 44.7 | 45.4 | 45.325
26 | 39.1333 | 42.525 | 44.5667 | 45.1875 | 45.175
26.1 | 38.88 | 42.375 | 44.4 | 45.0143 | 45.1
26.2 | 38.7 | 42.14 | 44.3 | 44.8333 | 44.875
26.3 | 38.6 | 42.1 | 44.1 | 44.7 | 44.725
26.4 | 38.4 | 41.85 | 43.9 | 44.5 | 44.5429
26.5 | 38.2 | 41.7 | 43.7 | 44.325 | 44.34
26.6 | 38.075 | 41.55 | 43.55 | 44.2 | 44.2
26.7 | 37.925 | 41.4 | 43.4 | 44.1 | 44.0333
26.8 | 37.875 | 41.18 | 43.275 | 43.975 | 43.9
26.9 | 37.7 | 40.95 | 43.125 | 43.8 | 43.7667
27 | 37.425 | 40.9 | 42.96 | 43.6 | 43.6
27.1 | 37.3 | 40.625 | 42.825 | 43.5 | 43.4667
27.2 | 37.1 | 40.525 | 42.7 | 43.275 | 43.3
27.3 | 37 | 40.325 | 42.5667 | 43.1 | 43.125
27.4 | 36.8429 | 40.175 | 42.375 | 43 | 42.9
27.5 | 36.6375 | 40 | 42.225 | 42.75 | 42.8
27.6 | 36.5 | 39.86 | 42.08 | 42.6 | 42.6
27.7 | 36.4 | 39.6 | 41.95 | 42.5 | 42.5
27.8 | 36.225 | 39.5 | 41.8 | 42.4 | 42.4
27.9 | 36.1 | 39.325 | 41.55 | 42.25 | 42.2
28 | 35.9 | 39.2 | 41.4 | 42 | 42.1
28.1 | 35.75 | 39.18 | 41.25 | 41.8333 | 41.875
28.2 | 35.62 | 38.95 | 41.1 | 41.7 | 41.75
28.3 | 35.4 | 38.8 | 40.9 | 41.6 | 41.6
28.4 | 35.325 | 38.65 | 40.82 | 41.45 | 41.5
28.5 | 35.2 | 38.5 | 40.66 | 41.3 | 41.4
28.6 | 35.025 | 38.375 | 40.6 | 41.1 | 41.2
28.7 | 34.875 | 38.2 | 40.375 | 41 | 41
28.8 | 34.7 | 38.075 | 40.225 | 40.85 | 40.875
28.9 | 34.5 | 37.925 | 40.0875 | 40.65 | 40.725
29 | 34.3 | 37.75 | 39.9571 | 40.525 | 40.6
29.1 | 34.24 | 37.68 | 39.84 | 40.4 | 40.425
29.2 | 34.1 | 37.45 | 39.7 | 40.3 | 40.3
29.3 | 33.95 | 37.3 | 39.45 | 40.175 | 40.1333
29.4 | 33.8 | 37.15 | 39.3 | 40 | 40
29.5 | 33.7 | 37 | 39.15 | 39.8 | 39.8
29.6 | 33.6 | 36.8 | 39 | 39.7 | 39.7
29.7 | 33.425 | 36.6875 | 38.8 | 39.5 | 39.6
29.8 | 33.325 | 36.5625 | 38.7667 | 39.4 | 39.4
29.9 | 33.1167 | 36.375 | 38.6429 | 39.275 | 39.3
30 | 33 | 36.2125 | 38.475 | 39.125 | 39.2

## Wedge 15 Lateral Profiles
Lateral Profiles
Wedge 15
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:3 cm FieldSize:4x4 cm2 | Depth:3 cm FieldSize:10x10 cm2 | Depth:3 cm FieldSize:20x20 cm2 | Depth:3 cm FieldSize:30x30 cm2 | Depth:3 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.9 | [8 empty] | 4.4333 | 5.1 |  |  |  | 5.8 | 6.5667 |  |  |  | 9.775 | 10.6 |  |  |  | 14.75 | 15.66
-22.8 | [8 empty] | 4.5 | 5.2 |  |  |  | 5.9 | 6.6333 |  |  |  | 9.925 | 10.8 |  |  |  | 14.94 | 15.875
-22.7 | [8 empty] | 4.5429 | 5.24 |  |  |  | 5.98 | 6.8 |  |  |  | 10.1 | 10.95 |  |  |  | 15.12 | 16.0333
-22.6 |  |  |  | 5.075 | 5.7992 |  |  |  | 4.6 | 5.3 |  |  |  | 6.1 | 6.9 |  |  |  | 10.28 | 11.1 |  |  |  | 15.34 | 16.35
-22.5 |  |  |  | 5.1 | 5.8825 |  |  |  | 4.7 | 5.4 |  |  |  | 6.3 | 7.025 |  |  |  | 10.425 | 11.25 |  |  |  | 15.62 | 16.5667
-22.4 |  |  |  | 5.2 | 5.9134 |  |  |  | 4.775 | 5.4 |  |  |  | 6.3667 | 7.1 |  |  |  | 10.6 | 11.4 |  |  |  | 15.8 | 16.82
-22.3 |  |  |  | 5.2 | 5.9991 |  |  |  | 4.82 | 5.5 |  |  |  | 6.45 | 7.2167 |  |  |  | 10.775 | 11.5833 |  |  |  | 16.1 | 17
-22.2 |  |  |  | 5.3 | 6.0848 |  |  |  | 4.9 | 5.6 |  |  |  | 6.5 | 7.3 |  |  |  | 10.95 | 11.75 |  |  |  | 16.32 | 17.3
-22.1 |  |  |  | 5.4 | 6.1491 |  |  |  | 5 | 5.65 |  |  |  | 6.6833 | 7.46 |  |  |  | 11.1167 | 11.975 |  |  |  | 16.6 | 17.5333
-22 |  |  |  | 5.44 | 6.2191 |  |  |  | 5.0143 | 5.74 |  |  |  | 6.75 | 7.575 |  |  |  | 11.3 | 12.18 |  |  | 5.3 | 16.8 | 17.7833
-21.9 |  |  |  | 5.575 | 6.2991 |  |  |  | 5.1 | 5.8167 |  |  |  | 6.95 | 7.7 |  |  |  | 11.4333 | 12.3333 |  |  | 5.3 | 17.1667 | 18.1
-21.8 |  |  |  | 5.6 | 6.3991 |  |  |  | 5.2 | 5.925 |  |  |  | 7.06 | 7.8 |  |  |  | 11.7 | 12.5 |  |  | 5.4 | 17.38 | 18.35
-21.7 |  |  |  | 5.7 | 6.4991 |  |  |  | 5.3 | 6 |  |  |  | 7.15 | 7.9 |  |  |  | 11.9 | 12.7 |  |  | 5.5 | 17.75 | 18.68
-21.6 |  |  |  | 5.8 | 6.5991 |  |  |  | 5.4 | 6.1 |  |  |  | 7.3 | 8.1 |  |  |  | 12.0667 | 12.9 |  |  | 5.5833 | 18 | 19.0167
-21.5 |  |  |  | 5.9 | 6.6324 |  |  |  | 5.5125 | 6.2 |  |  |  | 7.4 | 8.125 |  |  |  | 12.2667 | 13.12 |  |  | 5.66 | 18.4 | 19.3
-21.4 |  |  |  | 5.9833 | 6.799 |  |  |  | 5.6429 | 6.3 |  |  |  | 7.5 | 8.35 |  |  |  | 12.5 | 13.3333 |  |  | 5.76 | 18.7 | 19.66
-21.3 |  |  |  | 6.1 | 6.8823 |  |  |  | 5.7 | 6.4 |  |  |  | 7.7 | 8.44 |  |  |  | 12.725 | 13.5 |  |  | 5.8 | 19.02 | 20
-21.2 |  |  |  | 6.2 | 6.959 |  |  |  | 5.8 | 6.5 |  |  |  | 7.8 | 8.6 |  |  |  | 12.86 | 13.76 |  |  | 5.8286 | 19.4 | 20.4333
-21.1 |  |  |  | 6.3 | 7.099 |  |  |  | 5.9125 | 6.6667 |  |  |  | 7.9 | 8.7 |  |  |  | 13.1286 | 14 |  |  | 5.9833 | 19.85 | 20.7667
-21 |  |  |  | 6.4 | 7.2133 |  |  |  | 6.05 | 6.7333 |  |  |  | 8.08 | 8.825 |  |  |  | 13.3667 | 14.2 |  |  | 6 | 20.25 | 21.24
-20.9 |  |  |  | 6.5 | 7.299 |  |  |  | 6.2 | 6.8667 |  |  |  | 8.18 | 9 |  |  |  | 13.575 | 14.38 |  |  | 6.1 | 20.6667 | 21.7
-20.8 |  |  |  | 6.6 | 7.4239 |  |  |  | 6.25 | 6.96 |  |  |  | 8.3667 | 9.14 |  |  |  | 13.84 | 14.7 |  |  | 6.1 | 21.3286 | 22.24
-20.7 |  |  |  | 6.78 | 7.5989 |  |  |  | 6.4 | 7.1 |  |  |  | 8.5 | 9.3 |  |  |  | 14.04 | 15 |  |  | 6.2 | 21.9667 | 22.9
-20.6 |  |  |  | 6.8667 | 7.6989 |  |  |  | 6.5 | 7.2 |  |  |  | 8.68 | 9.44 |  |  |  | 14.3167 | 15.2 |  |  | 6.3 | 22.6333 | 23.55
-20.5 |  |  |  | 7 | 7.7989 |  |  |  | 6.6 | 7.35 |  |  |  | 8.775 | 9.6143 |  |  |  | 14.6 | 15.5 |  |  | 6.4 | 23.5 | 24.35
-20.4 |  |  |  | 7.12 | 7.9589 |  |  |  | 6.7 | 7.48 |  |  |  | 8.9 | 9.7 |  |  |  | 14.825 | 15.78 |  |  | 6.5 | 24.4167 | 25.375
-20.3 |  |  |  | 7.2 | 7.9989 |  |  |  | 6.8167 | 7.58 |  |  |  | 9.0833 | 9.88 |  |  | 4.54 | 15.15 | 16.0333 |  |  | 6.6 | 25.5 | 26.45
-20.2 |  |  |  | 7.32 | 8.1655 |  |  |  | 6.9833 | 7.68 |  |  |  | 9.22 | 10.0667 |  |  | 4.6 | 15.5 | 16.3 |  |  | 6.64 | 27.05 | 27.8833
-20.1 |  |  |  | 7.5 | 8.2417 |  |  |  | 7.06 | 7.85 |  |  |  | 9.44 | 10.2 |  |  | 4.7 | 15.7333 | 16.6333 |  |  | 6.7333 | 28.8 | 29.5857
-20 |  |  |  | 7.6 | 8.4155 |  |  |  | 7.2 | 7.94 |  |  |  | 9.6 | 10.38 |  |  | 4.7 | 16.08 | 16.94 |  |  | 6.8 | 30.85 | 31.82
-19.9 |  |  |  | 7.7 | 8.5988 |  |  |  | 7.3 | 8.05 |  |  |  | 9.8 | 10.6 |  |  | 4.8 | 16.4 | 17.24 |  |  | 6.9 | 33.8667 | 34.78
-19.8 |  |  |  | 7.8 | 8.6654 |  |  |  | 7.4 | 8.18 |  |  |  | 9.9667 | 10.8 |  |  | 4.8 | 16.825 | 17.55 |  |  | 7 | 37.94 | 38.86
-19.7 |  |  |  | 7.98 | 8.8321 |  |  |  | 7.5833 | 8.3 |  |  |  | 10.1333 | 10.96 |  |  | 4.9 | 17.15 | 18.02 |  |  | 7.1 | 43.4 | 44.3
-19.6 |  |  |  | 8.1 | 8.9154 |  |  |  | 7.7 | 8.48 |  |  |  | 10.4 | 11.1 |  |  | 5 | 17.55 | 18.4333 |  |  | 7.2 | 50.8 | 51.78
-19.5 |  |  |  | 8.3 | 9.082 |  |  |  | 7.82 | 8.58 |  |  |  | 10.6 | 11.3 |  |  | 5 | 18.05 | 18.9 |  |  | 7.3 | 58.9 | 59.15
-19.4 |  |  |  | 8.4 | 9.2387 |  |  |  | 8 | 8.72 |  |  |  | 10.7 | 11.58 |  |  | 5.1 | 18.55 | 19.36 |  |  | 7.4 | 67.3667 | 67.7
-19.3 |  |  |  | 8.5 | 9.3987 |  |  |  | 8.1667 | 8.9 |  |  |  | 10.9667 | 11.78 |  |  | 5.2 | 19.12 | 19.9 |  |  | 7.4667 | 74.8 | 75.32
-19.2 |  |  |  | 8.625 | 9.5486 |  |  |  | 8.3 | 9 |  |  |  | 11.18 | 12.02 |  |  | 5.2 | 19.7667 | 20.6 |  |  | 7.55 | 82.2 | 82.35
-19.1 |  |  |  | 8.84 | 9.7153 |  |  |  | 8.4286 | 9.18 |  |  |  | 11.38 | 12.24 |  |  | 5.25 | 20.44 | 21.3 |  |  | 7.7 | 87.8 | 87.68
-19 |  |  |  | 8.96 | 9.8819 |  |  |  | 8.5833 | 9.3833 |  |  |  | 11.6333 | 12.44 |  |  | 5.4 | 21.4 | 22.2 |  |  | 7.8 | 91.7 | 91.5833
-18.9 |  |  |  | 9.14 | 9.9986 |  |  |  | 8.75 | 9.5 |  |  |  | 11.85 | 12.72 |  |  | 5.4 | 22.3 | 23.1 |  |  | 7.8833 | 94.6167 | 94.3
-18.8 |  |  |  | 9.2333 | 10.1386 |  |  |  | 8.92 | 9.65 |  |  |  | 12.16 | 13 |  |  | 5.5 | 23.5 | 24.26 |  |  | 7.95 | 96.775 | 96.5167
-18.7 |  |  |  | 9.44 | 10.2652 |  |  |  | 9.0857 | 9.875 |  |  |  | 12.4 | 13.22 |  |  | 5.6 | 25.04 | 25.65 |  |  | 8.1167 | 98.1 | 97.8667
-18.6 |  |  |  | 9.65 | 10.4985 |  |  |  | 9.25 | 10.08 |  |  | 3.7 | 12.7 | 13.5 |  |  | 5.7 | 26.9 | 27.6833 |  |  | 8.2833 | 99.46 | 99
-18.5 |  |  |  | 9.78 | 10.6818 |  |  |  | 9.5 | 10.2 |  |  | 3.7 | 12.98 | 13.825 |  |  | 5.7667 | 29.4 | 30.1 |  |  | 8.36 | 100.3333 | 99.88
-18.4 |  |  |  | 9.98 | 10.8651 |  |  |  | 9.6833 | 10.5 |  |  | 3.8 | 13.32 | 14.1571 |  |  | 5.8 | 32.7333 | 33.32 |  |  | 8.5 | 101.3 | 100.7
-18.3 |  |  |  | 10.12 | 11.0484 |  |  |  | 9.85 | 10.62 |  |  | 3.8 | 13.65 | 14.5 |  |  | 5.9 | 37.7 | 38.15 |  |  | 8.66 | 102.02 | 101.5
-18.2 |  |  |  | 10.35 | 11.2184 |  |  |  | 10.1 | 10.8333 |  |  | 3.88 | 13.975 | 14.9 |  |  | 6 | 44.66 | 44.64 |  |  | 8.8 | 102.54 | 101.9333
-18.1 |  |  |  | 10.58 | 11.3984 |  |  |  | 10.2667 | 11.0333 |  |  | 3.9 | 14.45 | 15.25 |  |  | 6.1 | 52.7 | 52.7 |  |  | 8.9 | 102.9 | 102.36
-18 |  |  |  | 10.725 | 11.565 |  |  |  | 10.4857 | 11.3 |  |  | 3.9167 | 14.85 | 15.6667 |  |  | 6.2 | 62.0333 | 61.96 |  |  | 9 | 103.26 | 102.7714
-17.9 |  |  |  | 10.96 | 11.7983 |  |  |  | 10.8167 | 11.52 |  |  | 4 | 15.35 | 16.18 |  |  | 6.3 | 71.3 | 71.24 |  |  | 9.2 | 103.5333 | 103.14
-17.8 |  |  |  | 11.175 | 11.9983 |  |  | 3.3 | 10.9833 | 11.75 |  |  | 4.1 | 15.94 | 16.8 |  |  | 6.4 | 79.94 | 79.8333 |  |  | 9.25 | 103.775 | 103.44
-17.7 |  |  |  | 11.3 | 12.1783 |  |  | 3.3 | 11.25 | 11.98 |  |  | 4.1 | 16.48 | 17.3857 |  |  | 6.475 | 87.9 | 87.4 |  |  | 9.4 | 104 | 103.6167
-17.6 |  |  |  | 11.5833 | 12.4649 |  |  | 3.3 | 11.6 | 12.32 |  |  | 4.2 | 17.25 | 18.125 |  |  | 6.575 | 94 | 93.34 |  |  | 9.62 | 104.3 | 103.8
-17.5 |  |  |  | 11.8 | 12.6582 |  |  | 3.4 | 11.8667 | 12.65 |  |  | 4.2 | 18.1 | 19 |  |  | 6.65 | 97.8 | 97.2 |  |  | 9.725 | 104.5 | 104
-17.4 |  |  | 3.8 | 12.1 | 12.9315 |  |  | 3.4 | 12.2333 | 13 |  |  | 4.3 | 19.175 | 20 |  |  | 6.7333 | 100.5667 | 99.95 |  |  | 9.9167 | 104.7 | 104.2
-17.3 |  |  | 3.9 | 12.28 | 13.1981 |  |  | 3.5 | 12.6 | 13.375 |  |  | 4.4 | 20.4 | 21.2667 |  |  | 6.9 | 102.4 | 101.8 |  |  | 10.1 | 104.82 | 104.4
-17.2 |  |  | 3.9 | 12.6 | 13.4781 |  |  | 3.5 | 13.0375 | 13.8 |  |  | 4.4 | 22.1 | 22.86 |  |  | 7 | 103.9167 | 103.28 |  |  | 10.24 | 104.92 | 104.5
-17.1 |  |  | 4 | 12.9 | 13.798 |  |  | 3.5 | 13.45 | 14.24 |  |  | 4.5 | 24.2 | 24.95 |  |  | 7.1 | 105 | 104.325 |  |  | 10.5 | 105.0167 | 104.6
-17 |  |  | 4 | 13.225 | 14.138 |  |  | 3.6 | 14.0667 | 14.7667 |  |  | 4.6 | 27 | 27.75 |  |  | 7.2 | 105.8333 | 105.38 |  |  | 10.65 | 105.2 | 104.7
-16.9 |  |  | 4.1 | 13.6 | 14.5379 |  |  | 3.6 | 14.65 | 15.425 |  |  | 4.64 | 30.6 | 31.5 |  |  | 7.3667 | 106.66 | 106.15 |  |  | 10.8333 | 105.3 | 104.8
-16.8 |  |  | 4.1 | 14.02 | 14.9312 |  |  | 3.7 | 15.5 | 16.175 |  |  | 4.7 | 36.96 | 37.86 |  |  | 7.5 | 107.2 | 106.7 |  |  | 11 | 105.3 | 104.94
-16.7 |  |  | 4.2 | 14.54 | 15.3835 |  |  | 3.7 | 16.4 | 17.2 |  |  | 4.75 | 45.28 | 45.9 |  |  | 7.6 | 107.6 | 107.1 |  |  | 11.1667 | 105.3833 | 105.0333
-16.6 |  |  | 4.2 | 15.175 | 15.9977 |  |  | 3.8 | 17.6 | 18.34 |  |  | 4.8 | 55.14 | 55.5 |  |  | 7.7 | 108 | 107.52 |  |  | 11.3667 | 105.5 | 105.1
-16.5 |  |  | 4.3 | 15.825 | 16.6576 |  |  | 3.8333 | 19 | 19.7333 |  |  | 4.9 | 65.98 | 66.94 |  |  | 7.85 | 108.3 | 107.7 |  |  | 11.5 | 105.6 | 105.1857
-16.4 |  |  | 4.32 | 16.6 | 17.5641 |  |  | 3.9 | 21.0667 | 21.56 |  |  | 4.9833 | 77.04 | 77.55 |  |  | 8.04 | 108.58 | 107.95 |  |  | 11.7667 | 105.66 | 105.2
-16.3 |  |  | 4.4 | 17.7167 | 18.583 |  |  | 3.9 | 23.6286 | 24 |  |  | 5 | 87.25 | 86.98 |  |  | 8.175 | 108.8 | 108.2 |  |  | 11.98 | 105.74 | 105.3
-16.2 |  |  | 4.5 | 19.1 | 19.9472 |  |  | 4 | 27.2857 | 27.6 |  |  | 5.1 | 95.8 | 95.46 |  |  | 8.225 | 108.9333 | 108.4 |  |  | 12.1 | 105.8 | 105.4
-16.1 |  |  | 4.55 | 20.9 | 21.8218 |  |  | 4.06 | 32.6857 | 32.95 |  |  | 5.2 | 101.8 | 101.5 |  |  | 8.4429 | 109.1833 | 108.5 |  |  | 12.375 | 105.9 | 105.5
-16 |  |  | 4.6 | 23.3 | 24.1632 |  |  | 4.1 | 40.875 | 41.075 |  |  | 5.3 | 105.84 | 105.8 |  |  | 8.6 | 109.3 | 108.6 |  |  | 12.52 | 105.9 | 105.5
-15.9 |  |  | 4.64 | 27.1 | 27.6961 |  |  | 4.14 | 51 | 51.5 |  |  | 5.4 | 108.4 | 108 |  |  | 8.8 | 109.35 | 108.7 |  |  | 12.8167 | 106 | 105.5
-15.8 |  |  | 4.7 | 32.44 | 33.5552 |  |  | 4.2167 | 62.6 | 62.575 |  |  | 5.5 | 110.24 | 109.8 |  |  | 8.9 | 109.4 | 108.825 |  |  | 13 | 106 | 105.6
-15.7 |  |  | 4.8 | 40.7 | 41.9607 |  |  | 4.3 | 74.2333 | 74.075 |  |  | 5.5 | 111.6 | 111.1 |  |  | 9.025 | 109.5 | 108.9 |  |  | 13.2333 | 106 | 105.6
-15.6 |  |  | 4.8833 | 51.46 | 52.5525 |  |  | 4.4 | 86.4 | 85.74 |  |  | 5.68 | 112.525 | 112.16 |  |  | 9.2167 | 109.6 | 108.9 |  |  | 13.56 | 106.0667 | 105.62
-15.5 |  |  | 4.9 | 64.15 | 63.5909 |  |  | 4.425 | 96.4333 | 95.6 |  |  | 5.8 | 113.34 | 112.82 |  |  | 9.4667 | 109.6 | 109 |  | 3 | 13.85 | 106.1 | 105.7
-15.4 |  |  | 4.9667 | 76.28 | 76.2891 |  |  | 4.5 | 103.9571 | 103.75 |  |  | 5.88 | 113.9333 | 113.55 |  |  | 9.65 | 109.6 | 109 |  | 3 | 14.0667 | 106.1 | 105.8
-15.3 |  |  | 5 | 88.2 | 87.8541 |  |  | 4.6 | 109.04 | 108.8 |  |  | 6 | 114.28 | 113.95 |  |  | 9.8 | 109.7 | 109 |  | 3.1 | 14.2857 | 106.1833 | 105.9
-15.2 |  |  | 5.2 | 98.18 | 98.4692 |  |  | 4.7 | 112.2 | 111.64 |  |  | 6.1 | 114.66 | 114.2 |  |  | 9.98 | 109.7 | 109.1 |  | 3.1 | 14.6 | 106.25 | 105.9
-15.1 |  |  | 5.275 | 106.15 | 105.8136 |  |  | 4.8 | 114.35 | 113.8 |  |  | 6.3 | 114.78 | 114.4833 |  |  | 10.275 | 109.7 | 109.1 |  | 3.2 | 14.95 | 106.3 | 105.9
-15 |  |  | 5.3 | 111.1 | 110.6126 |  |  | 4.8833 | 115.84 | 115.25 |  |  | 6.35 | 114.86 | 114.66 |  |  | 10.4667 | 109.7 | 109.1 |  | 3.2 | 15.175 | 106.3833 | 106
-14.9 |  |  | 5.4 | 114.2 | 113.6007 |  |  | 4.9833 | 116.9 | 116.4 |  |  | 6.5 | 114.96 | 114.7 |  |  | 10.575 | 109.7 | 109.1 |  | 3.2 | 15.55 | 106.5 | 106.1
-14.8 |  |  | 5.5 | 116 | 115.3835 |  |  | 5 | 117.5286 | 117 |  |  | 6.6 | 115.0667 | 114.8 |  |  | 10.8333 | 109.7 | 109.1 |  | 3.3 | 15.95 | 106.52 | 106.1
-14.7 |  |  | 5.6 | 117.3 | 116.583 |  |  | 5.1 | 118.1 | 117.55 |  |  | 6.7 | 115.1 | 114.8 |  |  | 11.04 | 109.7 | 109.1 |  | 3.3 | 16.275 | 106.6 | 106.2
-14.6 |  |  | 5.7 | 118.06 | 117.583 |  |  | 5.275 | 118.48 | 117.875 |  |  | 6.88 | 115.2 | 114.8 |  |  | 11.25 | 109.6 | 109.1 |  | 3.36 | 16.6667 | 106.7 | 106.3
-14.5 |  |  | 5.74 | 118.64 | 118.0497 |  |  | 5.3 | 118.6286 | 118.1 |  |  | 7.1 | 115.2 | 114.8 |  |  | 11.48 | 109.6 | 109.1 |  | 3.4 | 17.1 | 106.8 | 106.4
-14.4 |  |  | 5.9 | 119 | 118.4163 |  |  | 5.4 | 118.7714 | 118.2 |  |  | 7.14 | 115.2 | 114.7 |  |  | 11.7667 | 109.6 | 109.04 |  | 3.4 | 17.625 | 106.8 | 106.46
-14.3 |  |  | 6 | 119.2 | 118.4997 |  |  | 5.5 | 118.8286 | 118.28 |  |  | 7.3 | 115.1 | 114.7 |  | 2.5 | 12.02 | 109.6 | 109 |  | 3.5 | 18.125 | 106.9 | 106.6
-14.2 |  |  | 6.1 | 119.2 | 118.583 |  |  | 5.6 | 118.9 | 118.3 |  |  | 7.5 | 115.05 | 114.6 |  | 2.5 | 12.34 | 109.6 | 109 |  | 3.5 | 18.8 | 106.98 | 106.6
-14.1 |  |  | 6.2 | 119.2 | 118.623 |  |  | 5.7 | 118.8833 | 118.3 |  |  | 7.6 | 114.9667 | 114.6 |  | 2.6 | 12.625 | 109.6 | 109 |  | 3.5 | 19.5667 | 107.0333 | 106.7
-14 |  |  | 6.3 | 119.2 | 118.583 |  |  | 5.875 | 118.8 | 118.2 |  |  | 7.8 | 114.9 | 114.525 |  | 2.6 | 12.925 | 109.6 | 109 |  | 3.6 | 20.2429 | 107.1 | 106.7
-13.9 |  |  | 6.42 | 119.12 | 118.483 |  |  | 5.98 | 118.6714 | 118.14 |  |  | 7.975 | 114.8 | 114.4167 |  | 2.6 | 13.2 | 109.6 | 109 |  | 3.6 | 21.3 | 107.16 | 106.8
-13.8 |  |  | 6.6 | 118.96 | 118.383 |  |  | 6.1 | 118.5 | 118 |  |  | 8.175 | 114.7 | 114.3 |  | 2.6333 | 13.6333 | 109.6 | 109 |  | 3.7 | 22.5 | 107.2 | 106.8
-13.7 |  |  | 6.7 | 118.8 | 118.2163 |  |  | 6.2 | 118.4429 | 117.9 |  |  | 8.35 | 114.54 | 114.2 |  | 2.7 | 14.02 | 109.6 | 109.05 |  | 3.7 | 23.85 | 107.3 | 106.9
-13.6 |  |  | 6.8 | 118.6 | 118.083 |  |  | 6.38 | 118.2 | 117.7833 |  |  | 8.56 | 114.4 | 114.1 |  | 2.7 | 14.3667 | 109.6 | 109.1 |  | 3.8 | 25.45 | 107.3 | 106.9167
-13.5 |  |  | 7 | 118.44 | 117.8687 |  |  | 6.5 | 118.1 | 117.6 |  |  | 8.78 | 114.3 | 114 |  | 2.7 | 14.82 | 109.7 | 109.15 |  | 3.9 | 27.7667 | 107.4 | 107.025
-13.4 |  |  | 7.075 | 118.2667 | 117.703 |  |  | 6.68 | 117.85 | 117.4667 |  |  | 9 | 114.3 | 113.9 |  | 2.8 | 15.3 | 109.7 | 109.2 |  | 3.9 | 30.6833 | 107.4 | 107.1
-13.3 |  |  | 7.26 | 118 | 117.423 |  |  | 6.8 | 117.6833 | 117.3 |  |  | 9.18 | 114.1 | 113.8 |  | 2.8 | 15.875 | 109.7 | 109.2 |  | 4 | 34.7 | 107.4 | 107.1
-13.2 |  |  | 7.4 | 117.85 | 117.243 |  |  | 6.98 | 117.5 | 117.08 |  |  | 9.42 | 114 | 113.64 |  | 2.8 | 16.54 | 109.8 | 109.3 |  | 4.1 | 40 | 107.4 | 107.1
-13.1 |  |  | 7.52 | 117.525 | 116.983 |  |  | 7.18 | 117.3333 | 116.775 |  | 1.9 | 9.7333 | 113.85 | 113.5 |  | 2.8 | 17.2429 | 109.8 | 109.35 |  | 4.1 | 47.24 | 107.4 | 107.1
-13 |  |  | 7.66 | 117.28 | 116.7997 |  |  | 7.325 | 117.1714 | 116.55 |  | 1.9 | 10 | 113.7 | 113.4 |  | 2.9 | 18 | 109.8 | 109.4 |  | 4.2 | 55.1 | 107.4 | 107
-12.9 |  |  | 7.9 | 117.1333 | 116.6163 |  |  | 7.5 | 117 | 116.425 |  | 1.9 | 10.32 | 113.5 | 113.3 |  | 3 | 19.06 | 109.8 | 109.4 |  | 4.3 | 63.6 | 107.4 | 107
-12.8 |  |  | 8.0833 | 116.96 | 116.3544 |  |  | 7.72 | 116.7429 | 116.2 |  | 1.94 | 10.6 | 113.5 | 113.2 |  | 3 | 20.24 | 109.9 | 109.4 |  | 4.3 | 71.3 | 107.4 | 107
-12.7 |  |  | 8.2333 | 116.775 | 116.1163 |  |  | 7.94 | 116.6 | 116 |  | 2 | 10.95 | 113.4 | 113.1 |  | 3 | 21.6 | 109.9 | 109.4 |  | 4.4 | 77.8 | 107.4 | 107
-12.6 |  |  | 8.44 | 116.5 | 115.983 |  |  | 8.1333 | 116.3667 | 115.82 |  | 2 | 11.35 | 113.3 | 113.1 |  | 3.1 | 23.5 | 109.9 | 109.4 |  | 4.5 | 83.6 | 107.4 | 107
-12.5 |  |  | 8.6 | 116.275 | 115.6687 |  | 1.7 | 8.3 | 116.08 | 115.62 |  | 2 | 11.75 | 113.3 | 113 |  | 3.1 | 25.8167 | 109.9 | 109.4 |  | 4.5 | 87.5 | 107.4 | 107
-12.4 |  |  | 8.86 | 116.1 | 115.4503 |  | 1.7 | 8.62 | 115.9143 | 115.44 |  | 2 | 12.22 | 113.22 | 113 |  | 3.2 | 29.0667 | 109.9 | 109.4 |  | 4.6 | 90.25 | 107.4 | 107
-12.3 |  | 2 | 9.025 | 115.875 | 115.2507 |  | 1.7 | 8.94 | 115.75 | 115.3 |  | 2.1 | 12.74 | 113.2 | 113 |  | 3.2 | 33.9667 | 109.9 | 109.4 |  | 4.7 | 92.4 | 107.375 | 107
-12.2 |  | 2 | 9.26 | 115.6 | 115.084 |  | 1.7 | 9.22 | 115.5857 | 115.2 |  | 2.1 | 13.42 | 113.1833 | 112.9 |  | 3.3 | 39.98 | 109.8833 | 109.4 |  | 4.8 | 94.1167 | 107.3 | 107.1
-12.1 |  | 2 | 9.56 | 115.5 | 114.9173 |  | 1.7 | 9.54 | 115.42 | 115 |  | 2.1 | 14.08 | 113.1 | 112.84 |  | 3.3 | 47.825 | 109.8 | 109.35 |  | 4.9 | 95.3667 | 107.3 | 107.1
-12 |  | 2 | 9.8 | 115.3 | 114.7507 |  | 1.8 | 9.9 | 115.35 | 114.9 |  | 2.2 | 14.9 | 113.1 | 112.8 |  | 3.4 | 56.3667 | 109.8 | 109.3 |  | 4.9 | 96.4 | 107.4 | 107.1
-11.9 |  | 2 | 10.15 | 115 | 114.684 |  | 1.8 | 10.36 | 115.24 | 114.825 |  | 2.2 | 15.9333 | 113.1 | 112.8 |  | 3.4 | 65.5 | 109.7 | 109.3 |  | 5.025 | 97.2 | 107.4 | 107.1
-11.8 |  | 2.075 | 10.46 | 114.9 | 114.5269 |  | 1.8 | 10.8 | 115.18 | 114.7 |  | 2.2 | 17.1 | 113.05 | 112.8 |  | 3.5 | 74 | 109.6333 | 109.25 |  | 5.15 | 98.1 | 107.5 | 107.1
-11.7 |  | 2.1 | 10.85 | 114.8 | 114.3507 |  | 1.8 | 11.45 | 115.1 | 114.6 |  | 2.3 | 18.7 | 113 | 112.8 |  | 3.6 | 81.6 | 109.6 | 109.2 |  | 5.2 | 98.8 | 107.5 | 107.1
-11.6 |  | 2.1 | 11.325 | 114.7 | 114.284 |  | 1.8 | 12.125 | 115 | 114.5 |  | 2.3 | 20.7 | 112.94 | 112.7 |  | 3.7 | 87.24 | 109.6 | 109.175 | 1.3 | 5.3 | 99.4 | 107.5 | 107.14
-11.5 |  | 2.1 | 11.7667 | 114.64 | 114.184 |  | 1.85 | 12.9 | 114.9 | 114.5 |  | 2.3 | 23.34 | 112.9 | 112.65 |  | 3.78 | 91.02 | 109.6 | 109.1 | 1.225 | 5.4 | 99.8667 | 107.5 | 107.2
-11.4 |  | 2.1333 | 12.4 | 114.6 | 114.084 |  | 1.9 | 13.85 | 114.8 | 114.4 |  | 2.4 | 26.9667 | 112.8 | 112.58 |  | 3.8 | 93.86 | 109.5 | 109.025 | 1.275 | 5.475 | 100.275 | 107.5 | 107.25
-11.3 |  | 2.2 | 13.0833 | 114.5 | 113.984 |  | 1.9 | 15.12 | 114.7333 | 114.3 |  | 2.4 | 32.1 | 112.7333 | 112.4 |  | 3.9 | 95.9333 | 109.4 | 109 | 1.3 | 5.5 | 100.65 | 107.5 | 107.3
-11.2 |  | 2.2 | 14 | 114.4 | 113.904 |  | 1.9 | 16.66 | 114.6 | 114.2 |  | 2.5 | 39.76 | 112.65 | 112.3 |  | 3.975 | 97.42 | 109.4 | 109 | 1.2 | 5.6 | 100.96 | 107.6 | 107.3
-11.1 |  | 2.2 | 15.1 | 114.3 | 113.834 |  | 2 | 18.7 | 114.5 | 114.1333 |  | 2.5 | 48.62 | 112.6 | 112.2 |  | 4.02 | 98.6 | 109.4 | 108.9 | 1.3 | 5.7 | 101.36 | 107.6 | 107.3
-11 |  | 2.3 | 16.54 | 114.2 | 113.684 |  | 2 | 21.64 | 114.32 | 114.05 |  | 2.6 | 58.45 | 112.4 | 112.1667 |  | 4.1 | 99.7 | 109.4 | 108.98 | 1.3 | 5.8667 | 101.7 | 107.6 | 107.3
-10.9 |  | 2.3 | 18.575 | 114.2 | 113.584 |  | 2 | 25.4 | 114.25 | 113.9 |  | 2.6 | 69.35 | 112.3 | 112.1 |  | 4.2 | 100.6 | 109.4 | 109 | 1.3 | 5.975 | 101.95 | 107.7 | 107.4
-10.8 |  | 2.3 | 21.04 | 114 | 113.5269 |  | 2 | 31.54 | 114.2 | 113.8 |  | 2.6 | 78.5 | 112.2 | 111.9 |  | 4.25 | 101.38 | 109.4 | 109 | 1.3 | 6.1333 | 102.2 | 107.7 | 107.4
-10.7 |  | 2.3 | 25 | 113.9 | 113.384 |  | 2.1 | 39.8 | 114.1 | 113.675 |  | 2.7 | 86.525 | 112.1 | 111.8 | 1 | 4.3667 | 101.85 | 109.4 | 109 | 1.3 | 6.2 | 102.3429 | 107.7 | 107.4
-10.6 |  | 2.4 | 30.8 | 113.8 | 113.3173 |  | 2.1 | 49.26 | 114 | 113.575 |  | 2.7 | 92.48 | 112 | 111.7 | 1 | 4.4 | 102.4667 | 109.4 | 108.9 | 1.3 | 6.4 | 102.6 | 107.6 | 107.4167
-10.5 |  | 2.4 | 39.78 | 113.66 | 113.1507 |  | 2.1 | 61 | 113.8167 | 113.4 |  | 2.8 | 96.3 | 111.9 | 111.6 | 1 | 4.525 | 103 | 109.4 | 108.9 | 1.375 | 6.4 | 102.7714 | 107.6 | 107.5
-10.4 |  | 2.5 | 50.55 | 113.55 | 112.984 |  | 2.1 | 71.75 | 113.64 | 113.2 |  | 2.8 | 98.8 | 111.8 | 111.525 | 1 | 4.6 | 103.44 | 109.45 | 108.9 | 1.4 | 6.5667 | 102.9333 | 107.6 | 107.4
-10.3 |  | 2.5 | 61.38 | 113.325 | 112.864 |  | 2.2 | 81.6 | 113.4714 | 113.06 |  | 2.9 | 100.74 | 111.7 | 111.5 | 1 | 4.7 | 103.7 | 109.4 | 108.9 | 1.4 | 6.7 | 103.2 | 107.6 | 107.4
-10.2 |  | 2.5333 | 72.8 | 113.12 | 112.684 |  | 2.2 | 89.58 | 113.3167 | 112.95 |  | 2.9 | 102.3 | 111.7 | 111.4 | 1 | 4.8333 | 104.02 | 109.4833 | 108.9 | 1.4 | 6.9 | 103.3 | 107.6 | 107.4
-10.1 |  | 2.6 | 82.9 | 113 | 112.5411 |  | 2.3 | 95.75 | 113.1714 | 112.84 |  | 3 | 103.325 | 111.6 | 111.32 | 1 | 4.9667 | 104.24 | 109.4286 | 108.9 | 1.4 | 7 | 103.4333 | 107.6 | 107.4
-10 |  | 2.6 | 91 | 112.86 | 112.484 |  | 2.3 | 99.22 | 113.1 | 112.7 |  | 3.1 | 104.4 | 111.5333 | 111.2833 | 1.1 | 5.075 | 104.55 | 109.4 | 108.9 | 1.5 | 7.225 | 103.6 | 107.6 | 107.4
-9.9 |  | 2.7 | 97.05 | 112.725 | 112.3697 |  | 2.4 | 102 | 113 | 112.56 |  | 3.1 | 105.24 | 111.42 | 111.2 | 1.1 | 5.2 | 104.78 | 109.4 | 108.9 | 1.45 | 7.35 | 103.76 | 107.5333 | 107.4
-9.8 |  | 2.7 | 100.86 | 112.6 | 112.2173 |  | 2.4 | 103.9 | 112.875 | 112.5 | 0.8 | 3.2 | 105.775 | 111.4 | 111.2 | 1.1 | 5.3 | 105 | 109.3 | 108.9 | 1.5 | 7.6 | 103.8167 | 107.6 | 107.4
-9.7 |  | 2.8 | 103.5 | 112.48 | 112.124 |  | 2.5 | 105.3 | 112.72 | 112.4 | 0.7 | 3.25 | 106.35 | 111.3 | 111.2 | 1.1 | 5.4 | 105.18 | 109.3 | 108.9 | 1.5 | 7.75 | 104 | 107.6 | 107.4
-9.6 |  | 2.8 | 105.0667 | 112.4 | 112.034 |  | 2.5 | 106.45 | 112.6 | 112.34 | 0.7 | 3.32 | 106.8333 | 111.3 | 111.1 | 1.1 | 5.55 | 105.38 | 109.3 | 108.9 | 1.5 | 7.8 | 104.1 | 107.575 | 107.4
-9.5 |  | 2.9 | 106.45 | 112.3 | 111.9715 |  | 2.6 | 107.225 | 112.6 | 112.25 | 0.8 | 3.4 | 107.25 | 111.3 | 111.1 | 1.1 | 5.7 | 105.4667 | 109.2 | 108.8 | 1.5 | 7.975 | 104.25 | 107.6 | 107.3
-9.4 |  | 2.9 | 107.44 | 112.2 | 111.884 | 0.7 | 2.6 | 108 | 112.5 | 112.2 | 0.8 | 3.5 | 107.625 | 111.225 | 111.1 | 1.1 | 5.8 | 105.5857 | 109.2 | 108.8 | 1.6 | 8.2333 | 104.3 | 107.6 | 107.3
-9.3 |  | 3 | 108.0333 | 112.1 | 111.7983 | 0.7 | 2.7 | 108.5 | 112.5 | 112.2 | 0.8 | 3.6 | 107.8167 | 111.2 | 111 | 1.1 | 5.9 | 105.8 | 109.1 | 108.7 | 1.6 | 8.44 | 104.4167 | 107.55 | 107.3
-9.2 | 0.7 | 3.08 | 108.65 | 112.1 | 111.784 | 0.7 | 2.7667 | 109.1 | 112.5 | 112.18 | 0.8 | 3.6667 | 108 | 111.2 | 110.9 | 1.1 | 6.025 | 105.96 | 109.1 | 108.7 | 1.6 | 8.625 | 104.5 | 107.5 | 107.2667
-9.1 | 0.7 | 3.1 | 109.075 | 112.1 | 111.684 | 0.7 | 2.8 | 109.3667 | 112.4 | 112.1 | 0.8 | 3.75 | 108.3 | 111.14 | 110.9 | 1.2 | 6.3 | 106 | 109.075 | 108.7 | 1.7 | 8.9 | 104.5833 | 107.5 | 107.2
-9 | 0.76667 | 3.2 | 109.4 | 112.02 | 111.684 | 0.7 | 2.9 | 109.65 | 112.3 | 112 | 0.8 | 3.8 | 108.5 | 111.1 | 110.9 | 1.2 | 6.4 | 106.1 | 109 | 108.65 | 1.7 | 9.0667 | 104.7 | 107.5 | 107.1
-8.9 | 0.8 | 3.2 | 109.8333 | 112 | 111.684 | 0.7 | 2.95 | 109.9 | 112.2375 | 111.95 | 0.8 | 3.9333 | 108.6 | 111.1 | 110.8 | 1.2 | 6.6 | 106.2 | 108.9 | 108.6 | 1.7 | 9.275 | 104.7714 | 107.4 | 107.2
-8.8 | 0.8 | 3.3 | 110 | 111.9 | 111.584 | 0.7 | 3.025 | 110.05 | 112.1857 | 111.9 | 0.9 | 4.05 | 108.7 | 111 | 110.8 | 1.2 | 6.7 | 106.3 | 108.9 | 108.6 | 1.7667 | 9.525 | 104.8 | 107.4 | 107.2
-8.7 | 0.8 | 3.4 | 110.16 | 111.9 | 111.584 | 0.7 | 3.1 | 110.3 | 112.0333 | 111.8 | 0.9 | 4.2 | 108.8 | 110.9 | 110.7 | 1.3 | 6.9 | 106.3 | 108.9 | 108.58 | 1.7 | 9.75 | 104.84 | 107.4 | 107.2
-8.6 | 0.8 | 3.5 | 110.28 | 111.86 | 111.484 | 0.7 | 3.2 | 110.4 | 112 | 111.7 | 0.9 | 4.3 | 108.825 | 110.9 | 110.6 | 1.3 | 7.175 | 106.3 | 108.8 | 108.5 | 1.8 | 10.075 | 104.9333 | 107.4 | 107.2
-8.5 | 0.8 | 3.6 | 110.38 | 111.8 | 111.4007 | 0.7 | 3.3 | 110.48 | 111.9 | 111.7 | 0.9 | 4.4 | 108.96 | 110.8 | 110.54 | 1.3 | 7.3 | 106.4 | 108.8 | 108.5 | 1.8 | 10.4 | 105 | 107.4 | 107.2
-8.4 | 0.8 | 3.6 | 110.48 | 111.7 | 111.334 | 0.7 | 3.325 | 110.5 | 111.9 | 111.6 | 0.9 | 4.6 | 109 | 110.7 | 110.425 | 1.3 | 7.5 | 106.42 | 108.8 | 108.5 | 1.85 | 10.575 | 105.1 | 107.32 | 107.1
-8.3 | 0.8 | 3.7 | 110.5 | 111.6 | 111.284 | 0.7 | 3.475 | 110.6 | 111.8 | 111.5 | 0.9 | 4.7 | 109 | 110.6833 | 110.38 | 1.4 | 7.7 | 106.5 | 108.8 | 108.4 | 1.9 | 10.925 | 105.175 | 107.3 | 107.1
-8.2 | 0.8 | 3.8 | 110.5 | 111.5 | 111.184 | 0.8 | 3.5667 | 110.6 | 111.6571 | 111.4 | 0.9 | 4.9 | 109.02 | 110.6 | 110.3 | 1.4 | 7.9333 | 106.6 | 108.7 | 108.3 | 1.9 | 11.2667 | 105.2 | 107.3 | 107.1
-8.1 | 0.8 | 3.9 | 110.5 | 111.45 | 111.084 | 0.8 | 3.7 | 110.6 | 111.6 | 111.3 | 0.9 | 4.95 | 109.1 | 110.5 | 110.2 | 1.4 | 8.2 | 106.6 | 108.7 | 108.3 | 1.9 | 11.64 | 105.3 | 107.3 | 107
-8 | 0.86667 | 4 | 110.5 | 111.4 | 110.984 | 0.8 | 3.75 | 110.6 | 111.5 | 111.2 | 0.9 | 5.1 | 109.1 | 110.4 | 110.175 | 1.4 | 8.5 | 106.6333 | 108.6 | 108.2 | 2 | 12.05 | 105.3 | 107.3 | 107.05
-7.9 | 0.9 | 4.1 | 110.475 | 111.3 | 110.984 | 0.8 | 3.9 | 110.6 | 111.4 | 111.1 | 1 | 5.3333 | 109 | 110.3 | 110.1 | 1.4 | 8.8 | 106.7 | 108.5 | 108.2 | 2 | 12.6 | 105.3 | 107.3 | 107
-7.8 | 0.9 | 4.2 | 110.4 | 111.2 | 110.824 | 0.8 | 4 | 110.6 | 111.3 | 111 | 1 | 5.5 | 109.1 | 110.3 | 110 | 1.425 | 9.1 | 106.7667 | 108.5 | 108.2 | 2.025 | 13 | 105.4 | 107.3 | 107
-7.7 | 0.9 | 4.375 | 110.4 | 111.1 | 110.724 | 0.8 | 4.1 | 110.5 | 111.1625 | 110.95 | 1 | 5.675 | 109.1 | 110.2 | 109.94 | 1.5 | 9.4333 | 106.8 | 108.4 | 108.1333 | 2.1 | 13.7 | 105.4 | 107.3 | 106.9167
-7.6 | 0.9 | 4.5 | 110.4 | 111 | 110.684 | 0.8 | 4.25 | 110.5 | 111.0286 | 110.86 | 1 | 5.9 | 109.02 | 110.1 | 109.9 | 1.5 | 9.775 | 106.8 | 108.4 | 108.1 | 2.1 | 14.45 | 105.44 | 107.18 | 106.9
-7.5 | 0.9 | 4.6 | 110.3 | 110.9 | 110.584 | 0.8 | 4.4 | 110.4 | 110.9714 | 110.7667 | 1 | 6.1333 | 109 | 110 | 109.8 | 1.6 | 10.225 | 106.8 | 108.3 | 108.1 | 2.125 | 15.1 | 105.5 | 107.1 | 106.86
-7.4 | 0.9 | 4.775 | 110.3 | 110.7 | 110.484 | 0.8 | 4.6 | 110.4 | 110.9 | 110.6 | 1 | 6.3333 | 109 | 109.9 | 109.7 | 1.6 | 10.7667 | 106.8 | 108.3 | 108 | 2.2 | 16.175 | 105.6 | 107.1 | 106.8
-7.3 | 0.9 | 4.95 | 110.3 | 110.6 | 110.384 | 0.8 | 4.7 | 110.32 | 110.8 | 110.6 | 1 | 6.65 | 109 | 109.8 | 109.6 | 1.6 | 11.2667 | 106.8 | 108.2 | 108 | 2.3 | 17.225 | 105.6 | 107.0333 | 106.8
-7.2 | 0.9 | 5.1 | 110.2 | 110.6 | 110.3507 | 0.85 | 5 | 110.3 | 110.7333 | 110.46 | 1.075 | 6.95 | 109 | 109.8 | 109.5167 | 1.7 | 11.8 | 106.8 | 108.2 | 107.9167 | 2.3 | 18.625 | 105.6 | 107 | 106.8
-7.1 | 0.9 | 5.3 | 110.2 | 110.6 | 110.2697 | 0.9 | 5.2 | 110.2 | 110.6667 | 110.4 | 1.1 | 7.25 | 109 | 109.68 | 109.4167 | 1.7 | 12.5667 | 106.8 | 108.1 | 107.9 | 2.3333 | 20.5 | 105.6 | 106.9667 | 106.8
-7 | 0.9 | 5.4 | 110.1 | 110.5 | 110.184 | 0.9 | 5.5 | 110.2 | 110.6 | 110.3 | 1.1 | 7.675 | 109 | 109.55 | 109.3667 | 1.7 | 13.3667 | 106.8 | 108.1 | 107.8 | 2.4 | 22.5 | 105.6 | 106.9 | 106.8
-6.9 | 0.9 | 5.7 | 110.0333 | 110.4 | 110.084 | 0.9 | 5.775 | 110.16 | 110.5 | 110.2 | 1.1 | 8.1 | 109 | 109.5 | 109.3 | 1.8 | 14.3 | 106.9 | 108.0333 | 107.8 | 2.4 | 25.5 | 105.6 | 106.8571 | 106.8
-6.8 | 1 | 5.9667 | 110 | 110.32 | 109.984 | 0.9 | 6.125 | 110.1 | 110.38 | 110.125 | 1.2 | 8.6 | 108.9 | 109.4 | 109.2 | 1.8 | 15.6 | 106.9 | 108 | 107.7 | 2.5 | 29.6667 | 105.6 | 106.8 | 106.775
-6.7 | 1 | 6.15 | 109.9 | 110.2 | 109.9173 | 0.9 | 6.4 | 110 | 110.3 | 110.1 | 1.2 | 9.15 | 108.9 | 109.325 | 109.2 | 1.8 | 16.9667 | 106.9 | 108 | 107.7 | 2.5 | 35.05 | 105.7 | 106.75 | 106.7
-6.6 | 1 | 6.5 | 109.9 | 110.1 | 109.844 | 0.9 | 6.8 | 110 | 110.2 | 110.0167 | 1.2 | 9.8 | 108.9 | 109.3 | 109.1 | 1.9 | 18.8 | 106.9 | 107.95 | 107.6 | 2.6 | 42.7333 | 105.7 | 106.7 | 106.6
-6.5 | 1 | 6.8 | 109.8 | 110.02 | 109.7507 | 0.9 | 7.3 | 109.9 | 110.1 | 109.9 | 1.2 | 10.6 | 108.8167 | 109.2 | 109.1 | 1.9 | 21.1667 | 106.9 | 107.9 | 107.6 | 2.7 | 50.8 | 105.7 | 106.7 | 106.575
-6.4 | 1 | 7.225 | 109.7 | 110 | 109.684 | 0.96667 | 7.8667 | 109.88 | 110.1 | 109.8167 | 1.22 | 11.5333 | 108.8 | 109.125 | 109 | 2 | 24.5 | 106.9 | 107.8 | 107.54 | 2.7 | 60.05 | 105.7 | 106.6 | 106.5
-6.3 | 1.025 | 7.7667 | 109.7 | 109.9 | 109.584 | 1 | 8.5667 | 109.75 | 109.9857 | 109.7833 | 1.3 | 12.7667 | 108.7 | 109 | 108.9 | 2.0333 | 29.1 | 106.8667 | 107.7 | 107.42 | 2.8333 | 68.7333 | 105.7 | 106.6 | 106.4
-6.2 | 1.1 | 8.35 | 109.7 | 109.8 | 109.5173 | 1 | 9.4333 | 109.7 | 109.8429 | 109.68 | 1.3 | 14.1667 | 108.7 | 109 | 108.9 | 2.1 | 35.9 | 106.8 | 107.66 | 107.3333 | 2.9 | 76.6333 | 105.6571 | 106.6 | 106.4
-6.1 | 1.1 | 9.1 | 109.6167 | 109.7833 | 109.484 | 1 | 10.5 | 109.6 | 109.7833 | 109.575 | 1.3 | 16.15 | 108.65 | 108.9 | 108.8 | 2.2 | 44.05 | 106.8 | 107.6 | 107.3 | 2.95 | 82.875 | 105.6 | 106.5333 | 106.4
-6 | 1.1 | 10.1 | 109.5 | 109.6 | 109.334 | 1 | 11.725 | 109.5 | 109.6444 | 109.5 | 1.3333 | 18.65 | 108.6 | 108.8 | 108.7 | 2.26 | 52.95 | 106.8 | 107.5 | 107.2 | 3 | 87.4 | 105.6 | 106.4667 | 106.3333
-5.9 | 1.1 | 11.2 | 109.5 | 109.5 | 109.2673 | 1.1 | 13.45 | 109.5 | 109.5143 | 109.4 | 1.4 | 22 | 108.5333 | 108.8 | 108.56 | 2.3 | 62.7333 | 106.75 | 107.4 | 107.2 | 3.125 | 90.825 | 105.6 | 106.4 | 106.2667
-5.8 | 1.2 | 12.7333 | 109.42 | 109.4 | 109.1007 | 1.1 | 15.7333 | 109.36 | 109.4667 | 109.3 | 1.475 | 27.025 | 108.5 | 108.6333 | 108.5 | 2.35 | 71.6667 | 106.7 | 107.4 | 107.1 | 3.2 | 93.1333 | 105.6 | 106.28 | 106.2
-5.7 | 1.2 | 15 | 109.325 | 109.3 | 109.084 | 1.1 | 18.875 | 109.3 | 109.3833 | 109.175 | 1.5 | 34.4 | 108.44 | 108.6 | 108.4 | 2.4 | 79.3 | 106.7 | 107.3167 | 107 | 3.25 | 94.7667 | 105.525 | 106.175 | 106.1333
-5.6 | 1.2 | 17.8667 | 109.225 | 109.2167 | 108.984 | 1.15 | 23.2 | 109.2 | 109.3 | 109.0667 | 1.56 | 43 | 108.4 | 108.42 | 108.3 | 2.5 | 85.7333 | 106.7 | 107.2 | 106.9 | 3.3 | 96.1667 | 105.5 | 106.0833 | 106
-5.5 | 1.3 | 22.2333 | 109.12 | 109.1 | 108.904 | 1.2 | 30.0667 | 109.16 | 109.2625 | 108.975 | 1.6 | 53 | 108.3 | 108.3167 | 108.3 | 2.6 | 90 | 106.6 | 107.1 | 106.8667 | 3.5 | 97.2 | 105.5 | 106 | 105.975
-5.4 | 1.3 | 29.4 | 109.1 | 109 | 108.784 | 1.2 | 39.7 | 109.06 | 109.1167 | 108.9 | 1.7 | 63.9 | 108.2 | 108.2 | 108.16 | 2.6 | 92.8667 | 106.5667 | 107 | 106.8 | 3.6 | 98.2 | 105.4667 | 105.9 | 105.9
-5.3 | 1.3 | 38.8333 | 109 | 108.88 | 108.634 | 1.3 | 49.7667 | 108.94 | 108.9571 | 108.8 | 1.7 | 73.675 | 108.1 | 108.1 | 108.1 | 2.725 | 95.05 | 106.5 | 106.9 | 106.7 | 3.7 | 98.9 | 105.4 | 105.82 | 105.8
-5.2 | 1.4 | 49.1667 | 108.8 | 108.78 | 108.5642 | 1.3 | 60.6333 | 108.85 | 108.8143 | 108.7 | 1.8 | 82.1 | 108 | 108 | 107.9667 | 2.8667 | 96.7 | 106.4 | 106.8 | 106.6 | 3.8 | 99.4 | 105.35 | 105.7 | 105.725
-5.1 | 1.4 | 60.7 | 108.6833 | 108.68 | 108.485 | 1.3 | 71.3333 | 108.7833 | 108.68 | 108.575 | 1.8333 | 88.4333 | 107.9 | 107.86 | 107.8 | 2.9333 | 97.9 | 106.3 | 106.6833 | 106.5 | 3.925 | 100.0667 | 105.2833 | 105.7 | 105.7
-5 | 1.4 | 71.85 | 108.575 | 108.575 | 108.3017 | 1.4 | 80.1 | 108.6 | 108.5 | 108.4 | 1.9 | 92.5667 | 107.8 | 107.75 | 107.6333 | 3.0667 | 98.8667 | 106.3 | 106.5833 | 106.4 | 4.1 | 100.4 | 105.2 | 105.6 | 105.6
-4.9 | 1.5 | 81.725 | 108.5 | 108.4 | 108.185 | 1.425 | 88.05 | 108.44 | 108.3833 | 108.275 | 2 | 95.2 | 107.675 | 107.64 | 107.5 | 3.2 | 99.55 | 106.2 | 106.5 | 106.2833 | 4.2 | 100.75 | 105.15 | 105.5 | 105.5
-4.8 | 1.5 | 89.5 | 108.4 | 108.2833 | 108.085 | 1.5 | 93.1 | 108.35 | 108.2 | 108.1 | 2.1 | 97.35 | 107.55 | 107.5 | 107.5 | 3.3 | 100.4 | 106.1 | 106.4 | 106.2 | 4.4 | 101.075 | 105.1 | 105.5 | 105.4
-4.7 | 1.6 | 94.8 | 108.3 | 108.2 | 107.9707 | 1.5 | 96.3 | 108.2 | 108.1333 | 108 | 2.1 | 98.9 | 107.45 | 107.38 | 107.32 | 3.425 | 100.85 | 106 | 106.34 | 106.1 | 4.5 | 101.325 | 105.1 | 105.4571 | 105.3167
-4.6 | 1.6 | 97.9 | 108.1 | 108.0667 | 107.785 | 1.6 | 98.6 | 108.025 | 107.9833 | 107.8 | 2.2667 | 100.075 | 107.3 | 107.2 | 107.22 | 3.575 | 101.3333 | 105.96 | 106.3 | 106 | 4.7 | 101.575 | 105 | 105.4 | 105.3
-4.5 | 1.7 | 100 | 108.025 | 107.9 | 107.5993 | 1.7 | 100.2 | 107.92 | 107.8286 | 107.675 | 2.3 | 100.975 | 107.2 | 107.125 | 107.025 | 3.7 | 101.6 | 105.775 | 106.1333 | 105.9 | 4.8 | 101.725 | 104.9857 | 105.2333 | 105.2
-4.4 | 1.7 | 101.6 | 107.84 | 107.8 | 107.485 | 1.7 | 101.4 | 107.8 | 107.7714 | 107.54 | 2.475 | 101.7 | 107.1 | 106.9667 | 106.9 | 3.9 | 102.02 | 105.7 | 106 | 105.8 | 5.1 | 101.96 | 104.9 | 105.14 | 105.1
-4.3 | 1.8 | 102.65 | 107.75 | 107.7 | 107.3993 | 1.8333 | 102.3 | 107.7 | 107.6 | 107.4 | 2.5 | 102.2667 | 107 | 106.8 | 106.78 | 4.05 | 102.3 | 105.6 | 105.9 | 105.7 | 5.2 | 102.1 | 104.8667 | 105.1 | 105.1
-4.2 | 1.9 | 103.425 | 107.6 | 107.5 | 107.26 | 2 | 103 | 107.5 | 107.5 | 107.3 | 2.7 | 102.6333 | 106.85 | 106.68 | 106.65 | 4.25 | 102.5 | 105.54 | 105.82 | 105.6 | 5.45 | 102.3 | 104.8 | 105.0143 | 105
-4.1 | 2 | 104.1 | 107.4 | 107.3333 | 107.0421 | 2.1 | 103.5 | 107.375 | 107.2 | 107.1833 | 2.875 | 103.0333 | 106.7 | 106.6 | 106.52 | 4.45 | 102.625 | 105.46 | 105.7 | 105.5 | 5.9 | 102.45 | 104.8 | 104.95 | 104.88
-4 | 2.1 | 104.6 | 107.28 | 107.2 | 106.8993 | 2.225 | 103.8 | 107.2429 | 107.05 | 106.96 | 3.025 | 103.4333 | 106.6 | 106.44 | 106.38 | 4.7 | 102.8 | 105.4 | 105.6 | 105.4 | 6.3 | 102.6 | 104.7 | 104.9 | 104.8
-3.9 | 2.2 | 104.825 | 107.1 | 107.06 | 106.7707 | 2.3667 | 104.3 | 107 | 107 | 106.8 | 3.25 | 103.6 | 106.5 | 106.3 | 106.2667 | 5.15 | 102.9 | 105.2833 | 105.475 | 105.3 | 6.8 | 102.6 | 104.6 | 104.825 | 104.7
-3.8 | 2.3 | 105.075 | 107 | 106.86 | 106.635 | 2.6 | 104.5 | 106.9 | 106.85 | 106.64 | 3.5 | 103.8 | 106.4 | 106.2 | 106.15 | 5.3667 | 103.1 | 105.2 | 105.34 | 105.225 | 7.2 | 102.7 | 104.525 | 104.72 | 104.64
-3.7 | 2.42 | 105.1667 | 106.825 | 106.74 | 106.4993 | 2.775 | 104.7 | 106.8 | 106.7 | 106.5 | 3.7333 | 104 | 106.3 | 106.1 | 106.0333 | 5.925 | 103.2 | 105.2 | 105.3 | 105.1 | 7.65 | 102.8 | 104.5 | 104.6 | 104.55
-3.6 | 2.6 | 105.3 | 106.7 | 106.6 | 106.3421 | 3.05 | 104.8 | 106.7 | 106.5571 | 106.36 | 4.1 | 104.0333 | 106.1 | 106 | 105.94 | 6.3667 | 103.2667 | 105.1 | 105.18 | 105 | 8.3 | 102.875 | 104.4 | 104.5 | 104.4833
-3.5 | 2.8333 | 105.4 | 106.6 | 106.38 | 106.1993 | 3.3 | 104.8 | 106.575 | 106.3833 | 106.25 | 4.5333 | 104.1 | 106 | 105.82 | 105.7 | 6.94 | 103.3 | 105 | 105.06 | 104.8333 | 9.325 | 102.9 | 104.3 | 104.375 | 104.4
-3.4 | 3.05 | 105.4 | 106.46 | 106.275 | 106.085 | 3.6 | 104.9 | 106.4 | 106.2286 | 106.05 | 4.8667 | 104.2 | 105.8833 | 105.64 | 105.6333 | 7.7 | 103.3 | 104.9 | 104.925 | 104.8 | 10.375 | 102.9 | 104.22 | 104.25 | 104.24
-3.3 | 3.3333 | 105.4 | 106.26 | 106.15 | 105.9183 | 4.1333 | 104.9 | 106.2 | 106.06 | 106 | 5.4 | 104.3 | 105.7 | 105.5 | 105.48 | 8.55 | 103.4 | 104.75 | 104.84 | 104.7 | 11.75 | 102.9 | 104.0167 | 104.1 | 104.15
-3.2 | 3.8 | 105.35 | 106.1 | 105.96 | 105.785 | 4.6 | 104.9 | 106.075 | 105.9833 | 105.9 | 6.1 | 104.3 | 105.5833 | 105.4 | 105.3833 | 9.5 | 103.32 | 104.6 | 104.7 | 104.55 | 13.46 | 102.9 | 103.9 | 104 | 104
-3.1 | 4.3 | 105.3 | 105.8833 | 105.7833 | 105.61 | 5.2667 | 105 | 105.88 | 105.8286 | 105.65 | 6.9 | 104.2667 | 105.48 | 105.26 | 105.2 | 10.9667 | 103.3 | 104.475 | 104.6 | 104.3833 | 15.5 | 102.9 | 103.8 | 103.9 | 103.82
-3 | 4.84 | 105.2 | 105.78 | 105.7 | 105.465 | 6 | 104.9 | 105.7 | 105.6143 | 105.5 | 8 | 104.2 | 105.2333 | 105.05 | 105.075 | 12.65 | 103.275 | 104.36 | 104.475 | 104.2286 | 18.45 | 102.9 | 103.6833 | 103.7 | 103.62
-2.9 | 5.7 | 105.2 | 105.6 | 105.46 | 105.285 | 7.025 | 104.9 | 105.6 | 105.4143 | 105.3 | 9.275 | 104.125 | 105.1 | 104.92 | 104.86 | 14.9 | 103.2 | 104.3 | 104.25 | 104.05 | 22.4 | 102.9 | 103.6 | 103.54 | 103.54
-2.8 | 6.7667 | 105.1 | 105.4 | 105.34 | 105.185 | 8.3667 | 104.8 | 105.5 | 105.2667 | 105.1 | 11.15 | 104.1 | 105 | 104.6667 | 104.75 | 17.95 | 103.2 | 104.1 | 104.1333 | 103.9 | 27.9667 | 102.8 | 103.45 | 103.44 | 103.5
-2.7 | 8.1 | 105 | 105.2333 | 105.12 | 104.9517 | 10.2 | 104.7 | 105.24 | 105.1 | 104.9 | 13.4 | 104 | 104.85 | 104.4 | 104.56 | 22 | 103.2 | 103.8667 | 103.95 | 103.7667 | 35.8 | 102.6667 | 103.25 | 103.325 | 103.4
-2.6 | 10.1 | 104.9 | 105.02 | 104.975 | 104.7564 | 12.4667 | 104.525 | 105.025 | 104.9375 | 104.8 | 16.5667 | 104 | 104.65 | 104.3 | 104.34 | 28.4 | 103.1 | 103.7 | 103.8 | 103.6143 | 44.35 | 102.5 | 103.16 | 103.18 | 103.2333
-2.5 | 12.8333 | 104.75 | 104.85 | 104.74 | 104.5183 | 15.7333 | 104.3333 | 104.8 | 104.68 | 104.6 | 21.1 | 103.8 | 104.34 | 104.125 | 104.12 | 36.5 | 103.0667 | 103.52 | 103.7 | 103.45 | 54.7 | 102.4 | 103 | 102.95 | 103.02
-2.4 | 17.1 | 104.6 | 104.65 | 104.5 | 104.335 | 20.5333 | 104.175 | 104.5 | 104.5 | 104.35 | 27.9333 | 103.7 | 104.12 | 104.0167 | 104 | 46.4 | 102.9667 | 103.4 | 103.5167 | 103.2833 | 63.6 | 102.4 | 102.84 | 102.9 | 102.8333
-2.3 | 23.55 | 104.3333 | 104.325 | 104.275 | 104.085 | 27.725 | 104.025 | 104.375 | 104.1875 | 104.175 | 36.7667 | 103.6 | 104 | 103.725 | 103.8 | 56.4 | 102.85 | 103.2333 | 103.25 | 103.2 | 72.9 | 102.3 | 102.78 | 102.8 | 102.6333
-2.2 | 31.9 | 104.1 | 104.125 | 104.1 | 103.8183 | 37.5 | 103.85 | 104.16 | 103.9286 | 103.98 | 46.6 | 103.3667 | 103.82 | 103.55 | 103.525 | 66.6 | 102.6333 | 103.075 | 103.0714 | 103.0167 | 80.5 | 102.2 | 102.6 | 102.64 | 102.5
-2.1 | 43.2 | 103.9 | 103.9 | 103.8 | 103.635 | 48.3 | 103.7 | 103.85 | 103.75 | 103.7 | 57.8 | 103.2 | 103.5333 | 103.3 | 103.3 | 75.5 | 102.45 | 102.9 | 102.9 | 102.82 | 86.025 | 102.125 | 102.5 | 102.52 | 102.42
-2 | 54.9 | 103.7 | 103.7 | 103.62 | 103.445 | 59.4 | 103.5 | 103.625 | 103.5714 | 103.55 | 67.7 | 103.1 | 103.32 | 103.12 | 103.2 | 83 | 102.4 | 102.7 | 102.72 | 102.7 | 89.9 | 102.075 | 102.375 | 102.3 | 102.34
-1.9 | 65.75 | 103.5 | 103.4 | 103.35 | 103.2183 | 69.9 | 103.2 | 103.4167 | 103.3 | 103.3 | 77.4333 | 102.925 | 103.06 | 102.84 | 103 | 87.9667 | 102.3 | 102.56 | 102.6 | 102.55 | 92.55 | 101.925 | 102.2833 | 102.2 | 102.225
-1.8 | 76 | 103.3 | 103.2167 | 103.04 | 102.9683 | 78.6667 | 103.0333 | 103.225 | 103.2 | 103 | 84.4 | 102.8 | 102.9 | 102.7 | 102.72 | 91.425 | 102.25 | 102.3667 | 102.42 | 102.3 | 94.72 | 101.8 | 102.175 | 102.1333 | 102.16
-1.7 | 84.6 | 103.1 | 103.0167 | 102.9 | 102.785 | 85.9 | 102.8333 | 102.95 | 102.9286 | 102.725 | 89.3 | 102.6 | 102.75 | 102.58 | 102.6 | 93.7 | 102.175 | 102.2833 | 102.3 | 102.175 | 95.95 | 101.725 | 102 | 101.9667 | 102.04
-1.6 | 89.9 | 102.8 | 102.8 | 102.72 | 102.6421 | 90.65 | 102.6 | 102.825 | 102.675 | 102.6 | 92.7 | 102.4667 | 102.5833 | 102.4 | 102.3333 | 95.3333 | 102.075 | 102.1833 | 102.175 | 102.05 | 97.2 | 101.7 | 101.9 | 101.9 | 101.8667
-1.5 | 93.54 | 102.6 | 102.66 | 102.48 | 102.435 | 93.4 | 102.4 | 102.58 | 102.5429 | 102.425 | 94.65 | 102.3 | 102.4 | 102.2 | 102.15 | 96.8 | 101.9 | 102.1 | 102 | 101.9 | 98.0333 | 101.6 | 101.76 | 101.8 | 101.7
-1.4 | 95.6667 | 102.475 | 102.38 | 102.3 | 102.265 | 95.65 | 102.2 | 102.4 | 102.4 | 102.3 | 96.4 | 102.1 | 102.2167 | 102.06 | 102 | 97.7 | 101.8 | 101.88 | 101.875 | 101.8 | 98.5667 | 101.55 | 101.6833 | 101.6 | 101.5333
-1.3 | 97.3667 | 102.2 | 102.2 | 102.1333 | 102.0993 | 96.9 | 102 | 102.2 | 102.1429 | 102.1 | 97.4 | 102 | 102.1 | 101.84 | 101.8833 | 98.3 | 101.68 | 101.78 | 101.6833 | 101.6 | 99 | 101.5 | 101.5167 | 101.46 | 101.4
-1.2 | 98.3 | 102.1 | 102.04 | 101.9667 | 101.935 | 98 | 101.9 | 102.02 | 101.9833 | 101.9 | 98.175 | 101.825 | 101.78 | 101.7 | 101.7 | 98.94 | 101.5 | 101.7 | 101.5167 | 101.4 | 99.35 | 101.4 | 101.4 | 101.3667 | 101.3143
-1.1 | 99.1333 | 101.9667 | 101.85 | 101.8 | 101.6683 | 98.8 | 101.7 | 101.78 | 101.8167 | 101.7 | 99.15 | 101.7 | 101.65 | 101.48 | 101.6 | 99.44 | 101.4 | 101.5 | 101.45 | 101.3 | 99.65 | 101.375 | 101.2667 | 101.3 | 101.175
-1 | 99.5667 | 101.725 | 101.7 | 101.5333 | 101.585 | 99.3333 | 101.6 | 101.66 | 101.62 | 101.55 | 99.5667 | 101.5333 | 101.575 | 101.35 | 101.4 | 99.75 | 101.3 | 101.4 | 101.3 | 101.15 | 99.9 | 101.225 | 101.1 | 101 | 101.025
-0.9 | 99.9333 | 101.5 | 101.54 | 101.3833 | 101.4027 | 99.725 | 101.4 | 101.42 | 101.44 | 101.34 | 99.8 | 101.4 | 101.375 | 101.2 | 101.32 | 100 | 101.2 | 101.18 | 101.175 | 101.1 | 99.95 | 101.1 | 101 | 100.975 | 101
-0.8 | 100.1 | 101.3 | 101.2667 | 101.2 | 101.186 | 99.9 | 101.3 | 101.22 | 101.25 | 101.2 | 100.1 | 101.2 | 101.225 | 101.04 | 101.12 | 100.1 | 101 | 101 | 101 | 100.98 | 100.08 | 101 | 100.9 | 100.8 | 100.9
-0.7 | 100.2 | 101.075 | 101.125 | 101.1 | 101.0193 | 100 | 101.05 | 101.0333 | 101 | 101.08 | 100.2 | 101.1 | 101.08 | 100.92 | 100.92 | 100.125 | 100.9 | 100.9 | 100.9 | 100.7833 | 100.1 | 100.775 | 100.7667 | 100.6 | 100.8
-0.6 | 100.3 | 100.875 | 100.9 | 100.96 | 100.826 | 100.1667 | 100.8333 | 100.92 | 100.9 | 100.8667 | 100.3 | 100.875 | 100.88 | 100.8 | 100.78 | 100.2 | 100.7 | 100.7 | 100.7 | 100.6 | 100.2 | 100.7 | 100.62 | 100.525 | 100.55
-0.5 | 100.3 | 100.65 | 100.8167 | 100.7833 | 100.7527 | 100.3 | 100.65 | 100.7 | 100.7 | 100.74 | 100.3 | 100.725 | 100.7 | 100.5 | 100.575 | 100.2 | 100.6 | 100.6 | 100.6333 | 100.5 | 100.2 | 100.5 | 100.4667 | 100.4 | 100.4
-0.4 | 100.2 | 100.5667 | 100.6 | 100.6 | 100.5527 | 100.2 | 100.5 | 100.5 | 100.6 | 100.525 | 100.3 | 100.5 | 100.5 | 100.42 | 100.45 | 100.2 | 100.5 | 100.5167 | 100.5 | 100.34 | 100.2 | 100.5 | 100.3143 | 100.34 | 100.3
-0.3 | 100.2 | 100.4 | 100.42 | 100.46 | 100.3304 | 100.2 | 100.4 | 100.4167 | 100.4857 | 100.34 | 100.3 | 100.4 | 100.4 | 100.3 | 100.36 | 100.2 | 100.375 | 100.4 | 100.3 | 100.2 | 100.2 | 100.4 | 100.24 | 100.2667 | 100.2
-0.2 | 100.2 | 100.2 | 100.2 | 100.225 | 100.211 | 100.1 | 100.25 | 100.2333 | 100.3333 | 100.24 | 100.2333 | 100.25 | 100.2833 | 100.2 | 100.2 | 100.2 | 100.25 | 100.2667 | 100.1333 | 100.18 | 100.2 | 100.3 | 100.2 | 100.1857 | 100.1167
-0.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.086 | 100.1 | 100.1333 | 100.1 | 100.1571 | 100.14 | 100.1 | 100.1333 | 100.1167 | 100.1 | 100.08 | 100.1 | 100.1 | 100.175 | 100.06 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.8 | 99.9 | 99.88 | 99.9 | 99.8857 | 99.9 | 99.9 | 99.9 | 99.9333 | 99.9 | 99.9 | 99.8333 | 99.9 | 99.88 | 99.9 | 99.85 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.94 | 99.9 | 99.9
0.2 | 99.7 | 99.8 | 99.8 | 99.8 | 99.8107 | 99.8 | 99.9 | 99.8 | 99.8714 | 99.8 | 99.7333 | 99.7 | 99.8 | 99.7 | 99.8167 | 99.8 | 99.8 | 99.8 | 99.9 | 99.7143 | 99.8 | 99.8 | 99.9 | 99.8 | 99.9
0.3 | 99.5 | 99.7 | 99.72 | 99.8 | 99.7857 | 99.6667 | 99.75 | 99.72 | 99.8 | 99.7 | 99.65 | 99.6 | 99.7167 | 99.6 | 99.7 | 99.7 | 99.8 | 99.7333 | 99.8 | 99.7 | 99.7 | 99.8 | 99.8 | 99.7 | 99.7333
0.4 | 99.3333 | 99.6 | 99.7 | 99.7 | 99.6858 | 99.5 | 99.7 | 99.62 | 99.6875 | 99.7 | 99.475 | 99.6 | 99.7 | 99.5 | 99.6 | 99.5 | 99.7 | 99.6 | 99.7 | 99.6 | 99.6 | 99.8 | 99.7 | 99.6 | 99.6714
0.5 | 99.1333 | 99.6 | 99.6 | 99.675 | 99.6025 | 99.375 | 99.6 | 99.6 | 99.6 | 99.6 | 99.3 | 99.6 | 99.6 | 99.5 | 99.6 | 99.3 | 99.6667 | 99.6 | 99.6833 | 99.6 | 99.4 | 99.7 | 99.6 | 99.6 | 99.6
0.6 | 98.925 | 99.5 | 99.6 | 99.6 | 99.5858 | 99.05 | 99.52 | 99.5 | 99.6 | 99.6 | 99.2 | 99.5 | 99.6 | 99.4 | 99.525 | 99.1 | 99.5333 | 99.5429 | 99.6 | 99.55 | 99.3 | 99.675 | 99.6 | 99.55 | 99.5
0.7 | 98.6 | 99.5 | 99.56 | 99.6 | 99.5858 | 98.7 | 99.5 | 99.5 | 99.6 | 99.5 | 98.9 | 99.5 | 99.6 | 99.4 | 99.5 | 98.9 | 99.5 | 99.5 | 99.6 | 99.5 | 99.2 | 99.575 | 99.5 | 99.5 | 99.5
0.8 | 98.3 | 99.5 | 99.5 | 99.5 | 99.5858 | 98.45 | 99.45 | 99.5 | 99.6 | 99.5 | 98.65 | 99.5 | 99.5 | 99.3 | 99.5 | 98.75 | 99.4 | 99.5 | 99.6 | 99.5 | 99.1 | 99.5 | 99.5 | 99.4 | 99.5
0.9 | 98.1 | 99.4 | 99.45 | 99.5 | 99.5858 | 98.2 | 99.4 | 99.5 | 99.6 | 99.5 | 98.3 | 99.475 | 99.5 | 99.3 | 99.5 | 98.6 | 99.375 | 99.5 | 99.6 | 99.5 | 98.8 | 99.4 | 99.5 | 99.4 | 99.4
1 | 97.8 | 99.4 | 99.4 | 99.5 | 99.5858 | 97.45 | 99.4 | 99.4 | 99.5571 | 99.5 | 97.8 | 99.4 | 99.4333 | 99.3 | 99.5 | 98.2 | 99.2 | 99.4 | 99.6 | 99.5 | 98.625 | 99.4 | 99.5 | 99.4 | 99.4
1.1 | 97 | 99.3 | 99.4 | 99.5 | 99.4858 | 96.7667 | 99.4 | 99.5 | 99.5 | 99.4 | 97.2667 | 99.4 | 99.4 | 99.3 | 99.5 | 97.8 | 99.2 | 99.4 | 99.55 | 99.4 | 98.3 | 99.4 | 99.4 | 99.32 | 99.4
1.2 | 95.95 | 99.275 | 99.35 | 99.5 | 99.4858 | 96 | 99.4 | 99.4 | 99.5 | 99.4 | 96.5 | 99.3 | 99.3 | 99.3 | 99.5 | 97.3667 | 99.2 | 99.4 | 99.5 | 99.4 | 97.9 | 99.3667 | 99.3 | 99.3 | 99.4
1.3 | 95 | 99.2 | 99.3 | 99.5 | 99.4858 | 94.725 | 99.3 | 99.4 | 99.5 | 99.4 | 95.7 | 99.3 | 99.3 | 99.3 | 99.5 | 96.7667 | 99.1 | 99.3167 | 99.4 | 99.4 | 97.6333 | 99.3 | 99.3 | 99.3 | 99.3333
1.4 | 93.4 | 99.2 | 99.3 | 99.5 | 99.3858 | 93.15 | 99.25 | 99.4 | 99.475 | 99.4 | 94.7 | 99.2 | 99.4 | 99.3 | 99.4167 | 96 | 99.1 | 99.3 | 99.4 | 99.3 | 97.1 | 99.2333 | 99.3 | 99.3 | 99.3143
1.5 | 91.425 | 99.1 | 99.3 | 99.48 | 99.3858 | 91.25 | 99.2 | 99.4 | 99.5 | 99.4 | 92.9 | 99.1 | 99.4 | 99.3 | 99.4 | 94.95 | 99.1 | 99.3 | 99.4 | 99.3 | 96.6 | 99.2 | 99.3 | 99.3 | 99.3
1.6 | 88.3 | 99 | 99.3 | 99.4 | 99.3858 | 88.4 | 99.1 | 99.4 | 99.5 | 99.4 | 91.15 | 99.0667 | 99.4 | 99.3 | 99.4 | 93.7333 | 99.0667 | 99.3 | 99.4 | 99.24 | 95.5333 | 99.1 | 99.3 | 99.3 | 99.2167
1.7 | 84 | 99.1 | 99.3 | 99.4 | 99.3858 | 84.4 | 99.1 | 99.4 | 99.4 | 99.4 | 88.5667 | 99 | 99.3 | 99.275 | 99.3 | 92.2 | 99 | 99.3 | 99.4 | 99.2 | 94.8 | 99 | 99.3 | 99.3 | 99.2
1.8 | 76.625 | 99.1 | 99.3 | 99.4 | 99.3858 | 78.3 | 99.1 | 99.3 | 99.4 | 99.4 | 84.3 | 98.9 | 99.3 | 99.2 | 99.3 | 90.3 | 98.9 | 99.2 | 99.3 | 99.2 | 93.4 | 99 | 99.26 | 99.2333 | 99.2
1.9 | 67.0667 | 99 | 99.3 | 99.4 | 99.3858 | 70.5 | 99 | 99.3 | 99.4333 | 99.4 | 78.6 | 98.9 | 99.3 | 99.2 | 99.3 | 87.55 | 98.9 | 99.2 | 99.28 | 99.2 | 92.2333 | 99 | 99.2 | 99.2 | 99.2
2 | 56.725 | 99 | 99.2 | 99.42 | 99.3858 | 60.1 | 99 | 99.3 | 99.4 | 99.4 | 71 | 98.8 | 99.3 | 99.2 | 99.3 | 83.875 | 98.8 | 99.2 | 99.2 | 99.2 | 90.04 | 98.9 | 99.15 | 99.125 | 99.2
2.1 | 45.4 | 99 | 99.3 | 99.5 | 99.4858 | 49.7 | 99 | 99.3 | 99.4 | 99.4 | 61.6667 | 98.8 | 99.2 | 99.2 | 99.3 | 78.3333 | 98.75 | 99.16 | 99.2 | 99.2 | 87.55 | 98.8 | 99.1 | 99.1 | 99.2
2.2 | 35.0333 | 99 | 99.3 | 99.5 | 99.4858 | 39.0333 | 99 | 99.3 | 99.4333 | 99.4 | 51.5667 | 98.8 | 99.2 | 99.2 | 99.3 | 70.8 | 98.7 | 99.1 | 99.1 | 99.2 | 83.8 | 98.8 | 99.0167 | 99.1 | 99.2
2.3 | 25.6 | 98.9 | 99.3 | 99.5 | 99.4858 | 29.775 | 98.9 | 99.3 | 99.4875 | 99.4 | 41 | 98.7667 | 99.2 | 99.2 | 99.3 | 61.7 | 98.6 | 99.1 | 99.2 | 99.2 | 78.6 | 98.725 | 99 | 99.1 | 99.2
2.4 | 18.5333 | 98.9 | 99.3 | 99.5 | 99.4858 | 21.9 | 98.9 | 99.4 | 99.45 | 99.5 | 31.8333 | 98.7 | 99.2 | 99.2 | 99.3 | 51.9 | 98.55 | 99 | 99.2 | 99.1 | 71.6 | 98.6667 | 99 | 99.1 | 99.1667
2.5 | 14.25 | 98.8 | 99.3 | 99.5 | 99.4858 | 16.525 | 98.9 | 99.4 | 99.5 | 99.5 | 24.1 | 98.675 | 99.2 | 99.2 | 99.3333 | 42.85 | 98.4 | 99 | 99.2 | 99.1 | 63.4 | 98.5 | 99 | 99.1 | 99.1
2.6 | 10.8 | 98.8 | 99.3 | 99.5 | 99.4858 | 12.8333 | 98.8 | 99.4 | 99.5 | 99.5 | 18.5333 | 98.6 | 99.2 | 99.25 | 99.3 | 33.4 | 98.4 | 99 | 99.2 | 99.06 | 54.4 | 98.4 | 99 | 99.1 | 99.1
2.7 | 8.6 | 98.7 | 99.2 | 99.5 | 99.4858 | 10.4 | 98.7667 | 99.4 | 99.5 | 99.5 | 14.9 | 98.5 | 99.2 | 99.3 | 99.3 | 26.2 | 98.3 | 99 | 99.3 | 99.04 | 44.4 | 98.4 | 99 | 99.1 | 99
2.8 | 7.2 | 98.6 | 99.2 | 99.5 | 99.4858 | 8.5 | 98.7 | 99.3 | 99.5 | 99.5 | 12.1 | 98.4 | 99.2 | 99.2 | 99.3 | 20.5 | 98.2 | 99.0333 | 99.3 | 99 | 35.65 | 98.3 | 99 | 99.1 | 99
2.9 | 5.925 | 98.5 | 99.2 | 99.5 | 99.4858 | 7 | 98.6 | 99.3 | 99.5 | 99.5 | 10.175 | 98.3 | 99.2 | 99.2 | 99.3 | 16.5333 | 98.1 | 99.1 | 99.2 | 99.0667 | 28.6 | 98.2 | 98.9 | 99.1 | 99
3 | 5.05 | 98.4333 | 99.2 | 99.5 | 99.4858 | 6 | 98.5 | 99.2667 | 99.5 | 99.4 | 8.6 | 98.15 | 99.2 | 99.2 | 99.3 | 13.975 | 98 | 99.1 | 99.2 | 99 | 22.6 | 98.1 | 98.9 | 99 | 99
3.1 | 4.35 | 98.3 | 99.1 | 99.4 | 99.4658 | 5.25 | 98.3 | 99.2 | 99.5 | 99.4 | 7.4 | 98 | 99.1 | 99.2 | 99.2 | 11.8 | 97.85 | 99.1 | 99.1833 | 99 | 18.45 | 98 | 98.9 | 99 | 99
3.2 | 3.85 | 98.225 | 99.1 | 99.4 | 99.3858 | 4.55 | 98.2 | 99.2 | 99.5 | 99.4 | 6.5 | 97.9 | 99.1 | 99.1 | 99.2 | 10.325 | 97.7 | 99.05 | 99.1 | 99 | 15.5333 | 97.9 | 98.9 | 98.9571 | 99
3.3 | 3.425 | 98.1 | 99 | 99.35 | 99.3858 | 4.05 | 98.0667 | 99.2 | 99.4 | 99.4 | 5.775 | 97.725 | 99.1 | 99.175 | 99.2 | 9.05 | 97.55 | 98.9 | 99 | 99 | 13.5 | 97.8 | 98.8 | 98.9 | 99
3.4 | 3.1 | 97.875 | 98.98 | 99.3 | 99.2858 | 3.65 | 97.6 | 99.16 | 99.4 | 99.3 | 5.15 | 97.6 | 99 | 99.1 | 99.2 | 8.075 | 97.5 | 98.9 | 99 | 99 | 11.925 | 97.675 | 98.725 | 98.9 | 99
3.5 | 2.8333 | 97.6 | 98.9 | 99.3 | 99.2691 | 3.2333 | 97.45 | 99.1 | 99.3 | 99.3 | 4.7 | 97.3333 | 99 | 99.1 | 99.1 | 7.175 | 97.36 | 98.8 | 98.9 | 98.92 | 10.5667 | 97.5 | 98.7 | 98.9 | 98.9
3.6 | 2.6 | 97.3333 | 98.875 | 99.2 | 99.1858 | 2.9333 | 97.1333 | 99.1 | 99.3 | 99.2 | 4.2667 | 97.1333 | 98.9 | 99.0667 | 99.1 | 6.52 | 97.1667 | 98.8 | 98.9 | 98.9 | 9.24 | 97.3 | 98.7 | 98.8 | 98.9
3.7 | 2.45 | 97.025 | 98.8 | 99.1 | 99.0858 | 2.74 | 96.85 | 99 | 99.2 | 99.2 | 3.9 | 96.85 | 98.8333 | 98.98 | 99 | 6 | 97 | 98.7 | 98.8 | 98.8 | 8.5333 | 97.1 | 98.6 | 98.8 | 98.9
3.8 | 2.3 | 96.8 | 98.7 | 99 | 99.0858 | 2.575 | 96.65 | 98.9 | 99.2 | 99.1 | 3.6 | 96.5 | 98.8 | 98.9 | 99 | 5.6333 | 96.7 | 98.6 | 98.8 | 98.8 | 7.85 | 96.9 | 98.6 | 98.7 | 98.8
3.9 | 2.15 | 96.5 | 98.6 | 98.92 | 99.0359 | 2.3 | 96.1 | 98.88 | 99.0875 | 99.1 | 3.3 | 96.2 | 98.7 | 98.8167 | 98.875 | 5.1333 | 96.5 | 98.5167 | 98.8 | 98.7 | 7.25 | 96.7667 | 98.6 | 98.7 | 98.8
4 | 2 | 95.9 | 98.5 | 98.9 | 98.9859 | 2.2 | 95.45 | 98.78 | 99 | 99 | 3.1 | 95.8667 | 98.7 | 98.8 | 98.8 | 4.8 | 96.15 | 98.45 | 98.8 | 98.64 | 6.7 | 96.5 | 98.5 | 98.68 | 98.7
4.1 | 1.9 | 95.2 | 98.4 | 98.8 | 98.8859 | 2.0333 | 95.0333 | 98.7 | 99 | 98.95 | 2.9 | 95.42 | 98.5 | 98.7 | 98.7 | 4.55 | 95.8 | 98.4 | 98.7 | 98.6 | 6.325 | 96.3 | 98.46 | 98.6 | 98.7
4.2 | 1.8 | 94.7 | 98.3 | 98.8 | 98.7859 | 1.9 | 94.4 | 98.6 | 98.9 | 98.9 | 2.7 | 94.95 | 98.5 | 98.65 | 98.7 | 4.3 | 95.5 | 98.3 | 98.7 | 98.5 | 5.925 | 96.1 | 98.4 | 98.5 | 98.6
4.3 | 1.7 | 93.8 | 98.3 | 98.7 | 98.6859 | 1.8 | 93.5667 | 98.5 | 98.8 | 98.8 | 2.6 | 94.32 | 98.4 | 98.6 | 98.7 | 4.1 | 95.1 | 98.2 | 98.58 | 98.48 | 5.575 | 95.9 | 98.3 | 98.4 | 98.5
4.4 | 1.6333 | 92.5 | 98.2 | 98.64 | 98.6859 | 1.7 | 92.7 | 98.4 | 98.8 | 98.7 | 2.5 | 93.6 | 98.3 | 98.5 | 98.6 | 3.9 | 94.9 | 98.02 | 98.5 | 98.4 | 5.3 | 95.5 | 98.1 | 98.4 | 98.425
4.5 | 1.6 | 91.1 | 98.1 | 98.6 | 98.5859 | 1.6333 | 91.4333 | 98.38 | 98.7 | 98.68 | 2.3 | 92.8 | 98.3 | 98.4 | 98.5 | 3.7667 | 94.45 | 98 | 98.4 | 98.38 | 5.075 | 95.2333 | 98.1 | 98.3167 | 98.4
4.6 | 1.5 | 89.1 | 98.1 | 98.5 | 98.5526 | 1.6 | 89.7 | 98.28 | 98.6667 | 98.6 | 2.2 | 91.8333 | 98.175 | 98.35 | 98.48 | 3.5 | 93.9 | 97.9 | 98.32 | 98.25 | 4.8 | 94.9 | 98 | 98.3 | 98.3
4.7 | 1.5 | 86.2 | 97.9167 | 98.4167 | 98.4859 | 1.5 | 87.5 | 98.2 | 98.6 | 98.55 | 2.1 | 90.9 | 98 | 98.3 | 98.3667 | 3.4 | 93.1667 | 97.8 | 98.3 | 98.2 | 4.6 | 94.6333 | 97.9 | 98.2 | 98.3
4.8 | 1.5 | 81.6667 | 97.9 | 98.3 | 98.3859 | 1.45 | 84.5 | 98.1 | 98.5 | 98.48 | 2 | 89.4 | 97.9 | 98.3 | 98.3 | 3.275 | 92.7 | 97.7 | 98.3 | 98.1167 | 4.3333 | 94.275 | 97.8 | 98.2 | 98.2
4.9 | 1.4 | 75.3 | 97.75 | 98.3 | 98.286 | 1.4 | 80.2 | 98 | 98.4 | 98.4 | 1.9 | 87.5 | 97.8 | 98.2 | 98.2 | 3.1 | 91.9 | 97.7 | 98.2 | 98.1 | 4.2 | 93.95 | 97.7 | 98.1 | 98.1
5 | 1.4 | 66.6 | 97.66 | 98.2 | 98.2693 | 1.3333 | 73.575 | 97.9 | 98.4 | 98.3 | 1.9 | 85.1333 | 97.74 | 98.1 | 98.1167 | 3 | 91.2 | 97.5 | 98.1 | 98 | 4.0333 | 93.55 | 97.575 | 98 | 98.0333
5.1 | 1.3 | 57.2 | 97.55 | 98.12 | 98.186 | 1.3 | 65.45 | 97.86 | 98.3 | 98.2 | 1.8 | 81.5333 | 97.6667 | 98 | 98.1 | 2.9 | 90.225 | 97.425 | 98.0667 | 97.925 | 4 | 93 | 97.5 | 97.9333 | 98
5.2 | 1.3 | 46.8667 | 97.5 | 98 | 98.086 | 1.2333 | 56.4333 | 97.8 | 98.28 | 98.2 | 1.7 | 76.5333 | 97.5 | 97.9333 | 98 | 2.8 | 89.2 | 97.3 | 97.98 | 97.8333 | 3.9 | 92.4667 | 97.38 | 97.8 | 97.9
5.3 | 1.25 | 37.1 | 97.375 | 97.9333 | 97.986 | 1.2 | 46.3 | 97.7 | 98.2 | 98.1 | 1.7 | 69.5667 | 97.5 | 97.9 | 97.9 | 2.7 | 87.55 | 97.2 | 97.8143 | 97.7667 | 3.7 | 91.8333 | 97.3 | 97.7 | 97.8333
5.4 | 1.2 | 28.15 | 97.225 | 97.88 | 97.986 | 1.2 | 36.4 | 97.575 | 98 | 98 | 1.6 | 61.3 | 97.4 | 97.8 | 97.8 | 2.6 | 85.8 | 97.2 | 97.7 | 97.6 | 3.55 | 91.1 | 97.1667 | 97.6 | 97.8
5.5 | 1.2 | 21.95 | 97.1 | 97.8 | 97.8146 | 1.1 | 28.7667 | 97.45 | 97.9333 | 97.925 | 1.6 | 52.6333 | 97.3 | 97.7667 | 97.78 | 2.5 | 83.3 | 97.1 | 97.68 | 97.5333 | 3.4 | 90.4 | 97 | 97.55 | 97.68
5.6 | 1.1667 | 17.2 | 97 | 97.7 | 97.646 | 1.1 | 22.5 | 97.4 | 97.85 | 97.8 | 1.5 | 43.3667 | 97.2 | 97.64 | 97.7 | 2.5 | 80.1333 | 97 | 97.6 | 97.5 | 3.3 | 89.4 | 96.9 | 97.46 | 97.6
5.7 | 1.1 | 14.3 | 96.9 | 97.55 | 97.5528 | 1.1 | 17.7 | 97.2 | 97.7714 | 97.75 | 1.5 | 34.5 | 97.1 | 97.525 | 97.6333 | 2.4 | 75.5333 | 96.82 | 97.5 | 97.4 | 3.3 | 88.2333 | 96.8 | 97.4 | 97.5
5.8 | 1.1 | 12.125 | 96.8 | 97.4 | 97.4861 | 1.1 | 14.8 | 97.1 | 97.6167 | 97.6333 | 1.4 | 27.5667 | 96.98 | 97.5 | 97.5 | 2.3 | 69.3 | 96.75 | 97.36 | 97.3 | 3.1667 | 86.8667 | 96.7 | 97.3 | 97.4
5.9 | 1.1 | 10.55 | 96.7 | 97.3 | 97.3718 | 1 | 12.7 | 97 | 97.5333 | 97.5 | 1.4 | 22.05 | 96.9 | 97.375 | 97.42 | 2.22 | 61.85 | 96.6 | 97.26 | 97.275 | 3.1 | 85.2333 | 96.6 | 97.225 | 97.4
6 | 1.1 | 9.4 | 96.525 | 97.2 | 97.2194 | 1 | 11.1 | 96.8 | 97.46 | 97.4 | 1.3 | 18.5667 | 96.68 | 97.225 | 97.3 | 2.2 | 53.3 | 96.5 | 97.2 | 97.2 | 3 | 82.9 | 96.55 | 97.1 | 97.2167
6.1 | 1 | 8.5 | 96.3 | 97.1 | 97.1861 | 1 | 9.8 | 96.6833 | 97.3 | 97.22 | 1.3 | 16 | 96.48 | 97.0833 | 97.26 | 2.1 | 45.2 | 96.4 | 97.1 | 97.1167 | 2.9 | 80.1 | 96.375 | 97.1 | 97.2
6.2 | 1 | 7.85 | 96.15 | 97 | 97.029 | 1 | 8.8 | 96.5 | 97.225 | 97.1 | 1.275 | 13.9 | 96.3 | 97 | 97.18 | 2.1 | 36.8333 | 96.1857 | 97 | 97 | 2.9 | 75.825 | 96.2667 | 97.1 | 97.1167
6.3 | 1 | 7.2333 | 96.1 | 96.8833 | 96.8862 | 0.9 | 8.0667 | 96.4 | 97.1 | 97.06 | 1.2 | 12.4667 | 96.2 | 96.85 | 96.9857 | 2 | 30.575 | 96.02 | 96.9 | 96.9 | 2.8 | 70.5 | 96.175 | 96.9 | 96.95
6.4 | 1 | 6.8 | 95.95 | 96.7 | 96.7662 | 0.9 | 7.35 | 96.3 | 97.025 | 96.9833 | 1.2 | 11.2 | 96.1 | 96.8 | 96.82 | 1.9333 | 25 | 95.88 | 96.84 | 96.85 | 2.8 | 63.5 | 96 | 96.8 | 96.8833
6.5 | 0.9 | 6.3333 | 95.84 | 96.6 | 96.6862 | 0.9 | 6.8333 | 96.1 | 96.8857 | 96.8 | 1.2 | 10.2 | 95.98 | 96.7 | 96.74 | 1.9 | 21.5667 | 95.7667 | 96.725 | 96.7 | 2.7 | 55.4333 | 95.82 | 96.8 | 96.7
6.6 | 0.9 | 6.05 | 95.6333 | 96.5 | 96.5262 | 0.9 | 6.3667 | 96 | 96.8 | 96.7286 | 1.1 | 9.45 | 95.775 | 96.6 | 96.64 | 1.8 | 18.85 | 95.6 | 96.6333 | 96.6167 | 2.6 | 47.4333 | 95.6667 | 96.6 | 96.62
6.7 | 0.9 | 5.75 | 95.44 | 96.35 | 96.3862 | 0.9 | 5.95 | 95.9 | 96.6833 | 96.55 | 1.1 | 8.8 | 95.625 | 96.5 | 96.55 | 1.8 | 16.8333 | 95.45 | 96.4714 | 96.45 | 2.6 | 39.7 | 95.55 | 96.6 | 96.6
6.8 | 0.9 | 5.475 | 95.32 | 96.24 | 96.2862 | 0.8 | 5.6667 | 95.68 | 96.6 | 96.46 | 1.1 | 8.2 | 95.4 | 96.4 | 96.4571 | 1.7 | 15.175 | 95.3 | 96.4 | 96.34 | 2.5 | 33.2667 | 95.4 | 96.4333 | 96.5
6.9 | 0.9 | 5.2667 | 95.1167 | 96.1 | 96.1363 | 0.8 | 5.3 | 95.55 | 96.5 | 96.34 | 1.1 | 7.7333 | 95.3 | 96.3 | 96.3 | 1.7 | 13.975 | 95.1 | 96.3 | 96.2 | 2.5 | 27.9 | 95.2 | 96.36 | 96.46
7 | 0.9 | 5.1 | 95.04 | 96.02 | 95.9863 | 0.8 | 5.05 | 95.4 | 96.2667 | 96.24 | 1.0667 | 7.34 | 95.2 | 96.1 | 96.2 | 1.65 | 12.925 | 95 | 96.2 | 96.1 | 2.4 | 24.1 | 95.1333 | 96.2 | 96.3
7.1 | 0.9 | 4.825 | 94.9 | 95.975 | 95.8863 | 0.8 | 4.8667 | 95.2 | 96.15 | 96.18 | 1 | 6.95 | 95.075 | 96 | 96.05 | 1.6 | 12.1333 | 94.78 | 96.05 | 96 | 2.3 | 21.6 | 94.96 | 96.1 | 96.2
7.2 | 0.8 | 4.7 | 94.74 | 95.8 | 95.8006 | 0.8 | 4.6 | 95.1 | 96.08 | 96.06 | 1 | 6.6 | 94.88 | 95.875 | 96 | 1.5 | 11.3667 | 94.66 | 96 | 95.9 | 2.3 | 19.2 | 94.8 | 96 | 96.15
7.3 | 0.8 | 4.6 | 94.6 | 95.7 | 95.6577 | 0.8 | 4.5 | 94.9 | 96 | 95.95 | 1 | 6.3 | 94.625 | 95.7833 | 95.8333 | 1.5 | 10.8333 | 94.5 | 95.875 | 95.8 | 2.2 | 17.6 | 94.6 | 95.9 | 96.0571
7.4 | 0.8 | 4.4 | 94.4 | 95.5667 | 95.5031 | 0.8 | 4.3 | 94.8 | 95.9 | 95.74 | 1 | 6.1 | 94.475 | 95.6 | 95.7 | 1.5 | 10.3333 | 94.32 | 95.74 | 95.7 | 2.2 | 16.3 | 94.425 | 95.8167 | 95.8857
7.5 | 0.8 | 4.3 | 94.175 | 95.4 | 95.4531 | 0.8 | 4.1 | 94.7 | 95.75 | 95.625 | 1 | 5.8 | 94.24 | 95.4833 | 95.6 | 1.5 | 9.8667 | 94.1833 | 95.625 | 95.5 | 2.1 | 15.15 | 94.325 | 95.7 | 95.8
7.6 | 0.8 | 4.1667 | 93.94 | 95.32 | 95.3114 | 0.76667 | 4 | 94.5 | 95.5714 | 95.5167 | 0.9 | 5.6 | 94.1333 | 95.4 | 95.5 | 1.5 | 9.4 | 93.9 | 95.5 | 95.4857 | 2.1 | 14.3333 | 94.18 | 95.5333 | 95.6
7.7 | 0.8 | 4 | 93.8 | 95.2 | 95.1864 | 0.7 | 3.9 | 94.2 | 95.5 | 95.4 | 0.9 | 5.38 | 93.9 | 95.3 | 95.4 | 1.4 | 9 | 93.7667 | 95.3167 | 95.3333 | 2 | 13.675 | 93.88 | 95.4 | 95.5
7.8 | 0.8 | 3.9 | 93.575 | 95.1 | 95.0864 | 0.7 | 3.7 | 94.05 | 95.4 | 95.3 | 0.9 | 5.15 | 93.75 | 95.2 | 95.3 | 1.4 | 8.7 | 93.525 | 95.2333 | 95.2 | 2 | 12.9 | 93.68 | 95.3 | 95.4
7.9 | 0.8 | 3.8 | 93.45 | 94.975 | 94.9864 | 0.7 | 3.6 | 93.9 | 95.22 | 95.15 | 0.9 | 5 | 93.5667 | 95.1 | 95.2667 | 1.4 | 8.4333 | 93.46 | 95.1667 | 95.1 | 2 | 12.35 | 93.5 | 95.225 | 95.2
8 | 0.8 | 3.725 | 93.36 | 94.84 | 94.8364 | 0.7 | 3.5 | 93.7833 | 95.1333 | 95.1 | 0.9 | 4.8 | 93.36 | 95 | 95.1 | 1.3 | 8.1 | 93.2 | 95 | 94.96 | 2 | 11.9 | 93.3833 | 95.12 | 95.1429
8.1 | 0.7 | 3.6 | 93.125 | 94.72 | 94.7532 | 0.7 | 3.4 | 93.5571 | 95 | 95 | 0.9 | 4.7 | 93.175 | 94.8 | 95 | 1.3 | 7.8 | 93.0333 | 94.8833 | 94.8 | 2 | 11.4 | 93.2167 | 95 | 95.025
8.2 | 0.725 | 3.5333 | 92.975 | 94.5667 | 94.5698 | 0.7 | 3.3 | 93.325 | 94.9 | 94.84 | 0.9 | 4.525 | 92.98 | 94.7 | 94.8 | 1.3 | 7.6 | 92.8 | 94.7 | 94.6 | 1.9 | 10.975 | 93.075 | 94.86 | 94.95
8.3 | 0.7 | 3.5 | 92.725 | 94.4 | 94.3865 | 0.7 | 3.2 | 93.12 | 94.8 | 94.7 | 0.9 | 4.4 | 92.8667 | 94.6 | 94.6667 | 1.3 | 7.4 | 92.58 | 94.6 | 94.5 | 1.9 | 10.6333 | 92.85 | 94.6667 | 94.8
8.4 | 0.7 | 3.375 | 92.5 | 94.3 | 94.2865 | 0.7 | 3.1 | 92.9429 | 94.7 | 94.6 | 0.8 | 4.3 | 92.6167 | 94.5 | 94.5333 | 1.2 | 7.125 | 92.45 | 94.525 | 94.4 | 1.9 | 10.4 | 92.62 | 94.6 | 94.7
8.5 | 0.7 | 3.3 | 92.26 | 94.2 | 94.1266 | 0.7 | 3.05 | 92.68 | 94.55 | 94.5 | 0.8 | 4.15 | 92.45 | 94.32 | 94.4 | 1.2 | 6.9 | 92.3 | 94.4 | 94.3 | 1.875 | 10.05 | 92.4 | 94.475 | 94.5
8.6 | 0.7 | 3.2 | 91.9 | 94.05 | 94.0366 | 0.65 | 2.96 | 92.4 | 94.4 | 94.4 | 0.8 | 4 | 92.225 | 94.24 | 94.35 | 1.2 | 6.8 | 92.06 | 94.26 | 94.125 | 1.8 | 9.75 | 92.2 | 94.4 | 94.48
8.7 | 0.7 | 3.15 | 91.65 | 93.92 | 93.8866 | 0.6 | 2.9 | 92.125 | 94.3286 | 94.275 | 0.8 | 3.9333 | 91.9 | 94.125 | 94.3 | 1.2 | 6.5 | 91.8 | 94.1667 | 94 | 1.8 | 9.4 | 92.0333 | 94.25 | 94.3
8.8 | 0.7 | 3.1 | 91.35 | 93.8 | 93.7866 | 0.6 | 2.8 | 91.82 | 94.1833 | 94.1 | 0.8 | 3.8 | 91.6333 | 94 | 94.1 | 1.2 | 6.4 | 91.6167 | 94.08 | 93.9 | 1.7 | 9.175 | 91.9 | 94.1 | 94.18
8.9 | 0.7 | 3 | 90.9 | 93.6 | 93.7033 | 0.6 | 2.7 | 91.525 | 94 | 94 | 0.8 | 3.7 | 91.46 | 93.88 | 93.9714 | 1.2 | 6.275 | 91.24 | 93.9 | 93.82 | 1.7 | 9 | 91.7333 | 93.9833 | 94
9 | 0.7 | 3 | 90.475 | 93.475 | 93.4367 | 0.6 | 2.7 | 91.05 | 93.9333 | 93.9 | 0.8 | 3.6 | 91.2 | 93.775 | 93.9 | 1.1 | 6.1 | 91.0833 | 93.8 | 93.75 | 1.7 | 8.8 | 91.45 | 93.8 | 93.94
9.1 | 0.7 | 2.9 | 90.2 | 93.4 | 93.3581 | 0.6 | 2.6 | 90.6 | 93.7714 | 93.8 | 0.8 | 3.5 | 90.9 | 93.64 | 93.725 | 1.1 | 6 | 90.875 | 93.66 | 93.5857 | 1.7 | 8.5667 | 91.2667 | 93.6857 | 93.8
9.2 | 0.7 | 2.8 | 89.5 | 93.275 | 93.2867 | 0.6 | 2.6 | 90.2 | 93.6167 | 93.7 | 0.8 | 3.475 | 90.5 | 93.5 | 93.6 | 1.1 | 5.8667 | 90.725 | 93.525 | 93.4 | 1.6 | 8.4 | 90.9 | 93.6 | 93.68
9.3 |  | 2.8 | 89 | 93.2 | 93.1438 | 0.6 | 2.5 | 89.7 | 93.5 | 93.575 | 0.75 | 3.38 | 90.15 | 93.35 | 93.475 | 1.1 | 5.6333 | 90.48 | 93.42 | 93.3 | 1.6 | 8.2 | 90.8333 | 93.5 | 93.5
9.4 |  | 2.7 | 88.225 | 93.1 | 93.0867 | 0.6 | 2.425 | 89.1 | 93.4167 | 93.48 | 0.7 | 3.3 | 89.825 | 93.26 | 93.3 | 1.1 | 5.475 | 90.16 | 93.3 | 93.15 | 1.6 | 7.9 | 90.5667 | 93.3833 | 93.42
9.5 |  | 2.7 | 87.475 | 93 | 92.9867 |  | 2.4 | 88.48 | 93.35 | 93.3667 | 0.7 | 3.2 | 89.36 | 93.1 | 93.25 | 1.1 | 5.325 | 89.85 | 93.1 | 93.04 | 1.5 | 7.8333 | 90.375 | 93.3 | 93.3429
9.6 |  | 2.6 | 86.275 | 92.9 | 92.9296 |  | 2.3333 | 87.6 | 93.3 | 93.2 | 0.7 | 3.1333 | 88.875 | 93 | 93.1 | 1.1 | 5.2 | 89.5333 | 93 | 92.9 | 1.5 | 7.6 | 90.125 | 93.2 | 93.1833
9.7 |  | 2.6 | 84.84 | 92.9 | 92.8867 |  | 2.3 | 86.7 | 93.2167 | 93.1 | 0.7 | 3 | 88.375 | 92.9 | 93 | 1 | 5.1 | 89.35 | 92.86 | 92.7 | 1.5 | 7.5 | 89.9 | 93.0667 | 93.0167
9.8 |  | 2.5 | 82.725 | 92.8 | 92.7867 |  | 2.3 | 85.4 | 93.2 | 93.075 | 0.7 | 3 | 87.85 | 92.76 | 92.8 | 1 | 5 | 89.06 | 92.7833 | 92.6 | 1.5 | 7.3 | 89.72 | 93 | 92.95
9.9 |  | 2.5 | 79.7 | 92.8 | 92.7867 |  | 2.2 | 83.75 | 93.1 | 93 |  | 2.9 | 87.25 | 92.66 | 92.7167 | 1 | 4.9 | 88.7667 | 92.7 | 92.6 | 1.45 | 7.1 | 89.54 | 92.8333 | 92.8
10 |  | 2.5 | 75.3 | 92.8 | 92.6868 |  | 2.2 | 81.675 | 93.1 | 93 |  | 2.9 | 86.55 | 92.6 | 92.6 | 1 | 4.8333 | 88.4 | 92.54 | 92.42 | 1.4 | 7 | 89.22 | 92.66 | 92.7
10.1 |  | 2.4 | 68.625 | 92.8 | 92.6868 |  | 2.1667 | 78.475 | 93.1 | 93 |  | 2.8 | 85.7 | 92.5 | 92.6 | 1 | 4.7 | 88.1333 | 92.4 | 92.2 | 1.4 | 6.9 | 89 | 92.4833 | 92.6
10.2 |  | 2.4 | 60.575 | 92.8 | 92.6868 |  | 2.1 | 74.05 | 93 | 92.9 |  | 2.725 | 84.72 | 92.4 | 92.46 | 0.93333 | 4.6 | 87.7 | 92.25 | 92.1 | 1.3 | 6.7 | 88.7333 | 92.4 | 92.38
10.3 |  | 2.4 | 51.8333 | 92.8 | 92.6868 |  | 2.1 | 67.875 | 93 | 92.9 |  | 2.7 | 83.4 | 92.3333 | 92.4 | 0.9 | 4.5 | 87.4333 | 92.2 | 92.0167 | 1.3 | 6.6 | 88.5 | 92.3 | 92.275
10.4 |  | 2.3 | 42.7 | 92.8 | 92.6868 |  | 2 | 60.025 | 93 | 92.9 |  | 2.7 | 81.72 | 92.3 | 92.4 | 0.9 | 4.4 | 86.9 | 92.05 | 91.8833 | 1.3 | 6.5 | 88.26 | 92.16 | 92.2
10.5 |  | 2.3 | 34.475 | 92.8 | 92.6868 |  | 2 | 51.2333 | 93 | 92.9 |  | 2.6 | 79.9 | 92.2 | 92.3 | 0.9 | 4.3 | 86.46 | 91.8667 | 91.7167 | 1.3 | 6.375 | 87.95 | 92.05 | 92.06
10.6 |  | 2.25 | 27.6 | 92.8 | 92.6868 |  | 2 | 42.4 | 93 | 92.9 |  | 2.6 | 76.775 | 92.2 | 92.3 | 0.95 | 4.2 | 85.95 | 91.775 | 91.6 | 1.3 | 6.2 | 87.62 | 91.875 | 91.94
10.7 |  | 2.2 | 22.1333 | 92.8 | 92.6868 |  | 1.9 | 34.3167 | 93 | 92.9 |  | 2.5 | 72.5 | 92.2 | 92.2 | 1 | 4.1333 | 85.4 | 91.64 | 91.48 | 1.3 | 6.1 | 87.35 | 91.75 | 91.82
10.8 |  | 2.2 | 18.76 | 92.7333 | 92.6868 |  | 1.9 | 27.6 | 93 | 92.9 |  | 2.5 | 66.85 | 92.1 | 92.12 |  | 4.08 | 84.88 | 91.4667 | 91.4 | 1.2 | 6 | 87.05 | 91.55 | 91.6
10.9 |  | 2.1 | 16.325 | 92.7 | 92.6868 |  | 1.9 | 22.5 | 93 | 92.9 |  | 2.4 | 59.66 | 92.1 | 92.1 |  | 3.975 | 84.3333 | 91.3143 | 91.2333 | 1.2 | 5.825 | 86.74 | 91.5 | 91.46
11 |  | 2.1 | 14.625 | 92.7 | 92.6368 |  | 1.8333 | 18.9 | 93 | 92.9 |  | 2.4 | 51.575 | 92.1 | 92.1 |  | 3.9 | 83.5 | 91.24 | 91.12 | 1.1333 | 5.7 | 86.4 | 91.32 | 91.35
11.1 |  | 2.1 | 13.4 | 92.7 | 92.5868 |  | 1.8 | 16.54 | 93 | 92.9 |  | 2.3 | 43.5 | 92.1 | 92.1 |  | 3.825 | 82.8 | 91.125 | 91.1 | 1.2 | 5.6 | 86.12 | 91.2 | 91.2
11.2 |  | 2.1 | 12.325 | 92.6 | 92.5868 |  | 1.8 | 14.75 | 93 | 92.9 |  | 2.3 | 36.08 | 92 | 92.1 |  | 3.725 | 81.66 | 91.05 | 91 | 1.2 | 5.5 | 85.74 | 91.1 | 91.1
11.3 |  | 2.1 | 11.575 | 92.6 | 92.4868 |  | 1.8 | 13.3 | 92.9 | 92.9 |  | 2.2667 | 29.4 | 92 | 92.1 |  | 3.7 | 80.54 | 90.9833 | 90.9 | 1.2 | 5.4 | 85.3 | 90.925 | 90.9
11.4 |  | 2 | 10.9333 | 92.5 | 92.3868 |  | 1.7 | 12.3 | 92.9 | 92.8 |  | 2.2 | 24.5 | 92 | 92 |  | 3.6 | 78.95 | 90.9 | 90.8 | 1.2 | 5.3 | 85.02 | 90.8 | 90.8
11.5 |  | 2 | 10.4 | 92.4 | 92.3035 |  | 1.7 | 11.36 | 92.8 | 92.775 |  | 2.2 | 20.9 | 91.975 | 92 |  | 3.5 | 76.92 | 90.8 | 90.7 | 1.2 | 5.2 | 84.54 | 90.62 | 90.575
11.6 |  | 2 | 9.975 | 92.3 | 92.2868 |  | 1.7 | 10.7 | 92.7333 | 92.7 |  | 2.175 | 18.5 | 91.9 | 91.9 |  | 3.5 | 74.32 | 90.7333 | 90.6 | 1.2 | 5.2 | 84.05 | 90.44 | 90.5
11.7 |  | 1.9 | 9.6 | 92.3 | 92.2582 |  | 1.7 | 10.125 | 92.7 | 92.6667 |  | 2.1 | 16.65 | 91.9 | 91.9 |  | 3.4 | 70.48 | 90.7 | 90.6 |  | 5.1 | 83.55 | 90.25 | 90.3667
11.8 |  | 1.9 | 9.35 | 92.2 | 92.1868 |  | 1.7 | 9.55 | 92.64 | 92.6 |  | 2.1 | 15.22 | 91.8 | 91.8 |  | 3.3 | 65.55 | 90.6 | 90.5 |  | 5 | 83.04 | 90.1 | 90.3
11.9 |  | 1.9 | 9.02 | 92.125 | 92.0868 |  | 1.7 | 9.16 | 92.6 | 92.5 |  | 2 | 14.1 | 91.7 | 91.8 |  | 3.275 | 59.1 | 90.5 | 90.5 |  | 4.9 | 82.28 | 90.06 | 90.1667
12 |  | 1.9 | 8.78 | 92.1 | 92.0369 |  | 1.6 | 8.7667 | 92.5143 | 92.5 |  | 2 | 13.1667 | 91.66 | 91.7 |  | 3.2 | 52 | 90.5 | 90.4 |  | 4.8 | 81.62 | 89.9857 | 90
12.1 |  | 1.875 | 8.58 | 92.1 | 91.9869 |  | 1.6 | 8.46 | 92.5 | 92.4 |  | 2 | 12.46 | 91.6 | 91.6 |  | 3.2 | 44.96 | 90.42 | 90.4 |  | 4.7 | 80.7 | 89.8333 | 89.9167
12.2 |  | 1.8 | 8.3 | 92 | 91.9012 |  | 1.6 | 8.18 | 92.4 | 92.3 |  | 2 | 11.82 | 91.5 | 91.5 |  | 3.1 | 38.15 | 90.32 | 90.3 |  | 4.6 | 79.86 | 89.7667 | 89.8
12.3 |  | 1.8 | 8.1 | 91.96 | 91.8202 |  | 1.6 | 7.85 | 92.3 | 92.2 |  | 2 | 11.3 | 91.46 | 91.4 |  | 3.1 | 32.34 | 90.3 | 90.2 |  | 4.6 | 78.65 | 89.7 | 89.7
12.4 |  |  | 7.925 | 91.85 | 91.7469 |  | 1.6 | 7.6 | 92.2 | 92.1833 |  | 1.9 | 10.78 | 91.3 | 91.3 |  | 3 | 27.7 | 90.3 | 90.2 |  | 4.4667 | 77.12 | 89.6 | 89.7
12.5 |  |  | 7.7 | 91.8 | 91.5869 |  | 1.6 | 7.45 | 92.1429 | 92.1 |  | 1.9 | 10.34 | 91.28 | 91.22 |  | 3 | 23.9 | 90.2 | 90.1 |  | 4.4 | 75.4 | 89.5 | 89.6
12.6 |  |  | 7.6 | 91.7 | 91.5536 |  |  | 7.2 | 92.08 | 91.95 |  | 1.9 | 9.98 | 91.1 | 91.1333 |  | 2.9 | 21.7 | 90.2 | 90 |  | 4.3 | 72.84 | 89.4 | 89.5
12.7 |  |  | 7.4167 | 91.6 | 91.4869 |  |  | 7 | 92 | 91.9 |  | 1.825 | 9.72 | 91.05 | 91.0667 |  | 2.9 | 19.84 | 90.16 | 90 |  | 4.2 | 69.3667 | 89.3 | 89.4
12.8 |  |  | 7.2833 | 91.5 | 91.307 |  |  | 6.84 | 91.9 | 91.8 |  | 1.8 | 9.36 | 90.925 | 90.9 |  | 2.8667 | 18.4 | 90 | 89.86 |  | 4.2 | 65 | 89.24 | 89.3
12.9 |  |  | 7.075 | 91.4 | 91.287 |  |  | 6.7 | 91.85 | 91.7 |  | 1.8 | 9.05 | 90.8 | 90.8 |  | 2.8 | 17.2 | 89.8667 | 89.7 |  | 4.1 | 59.2667 | 89.125 | 89.2
13 |  |  | 6.95 | 91.3 | 91.187 |  |  | 6.525 | 91.6667 | 91.625 |  | 1.8 | 8.8 | 90.7 | 90.7 |  | 2.8 | 16.26 | 89.7 | 89.6167 |  | 4 | 53.08 | 89.05 | 89.1
13.1 |  |  | 6.8 | 91.2 | 91.0203 |  |  | 6.4 | 91.6 | 91.5 |  | 1.8 | 8.6 | 90.7 | 90.6 |  | 2.7 | 15.46 | 89.6 | 89.5 |  | 4 | 46.65 | 89 | 89.1
13.2 |  |  | 6.7 | 91.075 | 90.9703 |  |  | 6.25 | 91.5 | 91.4 |  |  | 8.36 | 90.6 | 90.5 |  | 2.7 | 14.7667 | 89.48 | 89.44 |  | 3.9 | 39.875 | 88.8833 | 88.975
13.3 |  |  | 6.575 | 90.925 | 90.8037 |  |  | 6.125 | 91.3571 | 91.275 |  |  | 8.14 | 90.42 | 90.4 |  | 2.6 | 14.24 | 89.375 | 89.34 |  | 3.9 | 34.7833 | 88.8 | 88.85
13.4 |  |  | 6.44 | 90.9 | 90.707 |  |  | 6 | 91.2 | 91.1 |  |  | 8 | 90.375 | 90.2 |  | 2.6 | 13.74 | 89.225 | 89.1667 |  | 3.8 | 29.9 | 88.7833 | 88.7
13.5 |  |  | 6.3 | 90.7 | 90.5071 |  |  | 5.8 | 91.12 | 90.98 |  |  | 7.8 | 90.175 | 90.1 |  | 2.6 | 13.25 | 89.1 | 89 |  | 3.8 | 26.7 | 88.6 | 88.6
13.6 |  |  | 6.2 | 90.575 | 90.4071 |  |  | 5.76 | 90.9429 | 90.875 |  |  | 7.625 | 90 | 90 |  | 2.5 | 12.92 | 88.925 | 88.9 |  | 3.7 | 24.1714 | 88.58 | 88.54
13.7 |  |  | 6.1 | 90.35 | 90.2204 |  |  | 5.625 | 90.78 | 90.625 |  |  | 7.475 | 89.9 | 89.8 |  | 2.5 | 12.5 | 88.825 | 88.7 |  | 3.7 | 22.25 | 88.5 | 88.5
13.8 |  |  | 5.9833 | 90.225 | 90.0371 |  |  | 5.5 | 90.6833 | 90.525 |  |  | 7.325 | 89.775 | 89.64 |  | 2.5 | 12.2 | 88.7 | 88.5 |  | 3.6 | 20.65 | 88.4 | 88.3833
13.9 |  |  | 5.8833 | 90.08 | 89.8539 |  |  | 5.4 | 90.5286 | 90.3833 |  |  | 7.2 | 89.64 | 89.52 |  | 2.4 | 11.84 | 88.6 | 88.45 |  | 3.6 | 19.6 | 88.2667 | 88.2167
14 |  |  | 5.8 | 89.85 | 89.6672 |  |  | 5.3 | 90.3333 | 90.1333 |  |  | 6.98 | 89.5 | 89.375 |  | 2.4 | 11.6 | 88.5 | 88.32 |  | 3.5 | 18.6 | 88.1 | 88.1167
14.1 |  |  | 5.66 | 89.6 | 89.3872 |  |  | 5.2 | 90.1 | 89.98 |  |  | 6.8333 | 89.26 | 89.24 |  | 2.4 | 11.28 | 88.325 | 88.175 |  | 3.45 | 17.95 | 88.1 | 88
14.2 |  |  | 5.6 | 89.26 | 89.1273 |  |  | 5.1167 | 89.8 | 89.7333 |  |  | 6.7 | 89.025 | 89.05 |  | 2.3 | 11 | 88.15 | 88 |  | 3.4 | 17.2 | 87.92 | 87.9
14.3 |  |  | 5.5 | 88.95 | 88.7873 |  |  | 5 | 89.6 | 89.42 |  |  | 6.58 | 88.82 | 88.82 |  | 2.3 | 10.775 | 87.96 | 87.88 |  | 3.4 | 16.525 | 87.825 | 87.74
14.4 |  |  | 5.45 | 88.52 | 88.4274 |  |  | 4.925 | 89.2333 | 89.075 |  |  | 6.4 | 88.6 | 88.6 |  |  | 10.525 | 87.7833 | 87.65 |  | 3.4 | 16 | 87.7 | 87.6
14.5 |  |  | 5.3667 | 87.9667 | 87.8874 |  |  | 4.8 | 88.8714 | 88.7333 |  |  | 6.3 | 88.425 | 88.36 |  |  | 10.38 | 87.7 | 87.55 |  | 3.3 | 15.58 | 87.48 | 87.48
14.6 |  |  | 5.2 | 87.3 | 87.1875 |  |  | 4.7 | 88.4429 | 88.34 |  |  | 6.2 | 88.3 | 88.16 |  |  | 10.16 | 87.48 | 87.35 |  | 3.2 | 15.2 | 87.375 | 87.375
14.7 |  |  | 5.15 | 86.45 | 86.1877 |  |  | 4.7 | 87.9 | 87.9 |  |  | 6.1 | 88.075 | 87.96 |  |  | 9.95 | 87.4 | 87.22 |  | 3.1 | 14.72 | 87.18 | 87.14
14.8 |  |  | 5.05 | 85.24 | 85.0279 |  |  | 4.6 | 87.3 | 87.125 |  |  | 6 | 87.825 | 87.7 |  |  | 9.725 | 87.2 | 87 |  | 3.1 | 14.36 | 87 | 87.02
14.9 |  |  | 4.9833 | 83.5 | 83.4881 |  |  | 4.5 | 86.4286 | 86.3 |  |  | 5.8833 | 87.5 | 87.38 |  |  | 9.475 | 87 | 86.86 |  | 3 | 14.02 | 86.8667 | 86.82
15 |  |  | 4.9 | 81.18 | 80.8885 |  |  | 4.4 | 85.5 | 85.3 |  |  | 5.7167 | 87.075 | 87.02 |  |  | 9.34 | 86.7667 | 86.6 |  | 3 | 13.68 | 86.7 | 86.6167
15.1 |  |  | 4.8 | 77.6 | 77.2764 |  |  | 4.3667 | 84.1 | 84.0667 |  |  | 5.65 | 86.78 | 86.7333 |  |  | 9.2 | 86.6 | 86.45 |  | 3 | 13.5 | 86.5 | 86.5
15.2 |  |  | 4.775 | 72.18 | 72.2897 |  |  | 4.3 | 82.5714 | 82.42 |  |  | 5.5857 | 86.425 | 86.425 |  |  | 9 | 86.3 | 86.22 |  | 3 | 13.1 | 86.4 | 86.2667
15.3 |  |  | 4.7 | 65.1 | 64.924 |  |  | 4.2 | 80.1714 | 80 |  |  | 5.5 | 85.925 | 85.9 |  |  | 8.8 | 86.18 | 86.05 |  | 3 | 13 | 86.2333 | 86.06
15.4 |  |  | 4.6 | 56.975 | 56.9347 |  |  | 4.14 | 76.5833 | 76.62 |  |  | 5.3667 | 85.4 | 85.36 |  |  | 8.7 | 85.9833 | 85.85 |  | 2.9 | 12.675 | 86 | 85.925
15.5 |  |  | 4.5 | 48.2 | 48.1931 |  |  | 4.1 | 71.1667 | 71.38 |  |  | 5.3 | 84.78 | 84.7 |  |  | 8.525 | 85.8167 | 85.7 |  | 2.9 | 12.46 | 85.8 | 85.8
15.6 |  |  | 4.5 | 39.95 | 39.9943 |  |  | 4.04 | 64.22 | 65 |  |  | 5.2 | 84.06 | 83.975 |  |  | 8.4 | 85.58 | 85.36 |  |  | 12.275 | 85.68 | 85.625
15.7 |  |  | 4.45 | 32.36 | 32.8953 |  |  | 4 | 56.84 | 57.08 |  |  | 5.1 | 83.04 | 83.1 |  |  | 8.2167 | 85.36 | 85.125 |  |  | 12.04 | 85.475 | 85.5
15.8 |  |  | 4.4 | 26.34 | 27.1561 |  |  | 3.9 | 48.6 | 48.46 |  |  | 5 | 81.96 | 81.9 |  |  | 8.04 | 85.06 | 84.88 |  |  | 11.775 | 85.26 | 85.38
15.9 |  |  | 4.3 | 22.1333 | 22.8634 |  |  | 3.9 | 39.7 | 40.42 |  |  | 4.9 | 80.52 | 80.45 |  |  | 7.9 | 84.88 | 84.66 |  |  | 11.6333 | 85.15 | 85.25
16 |  |  | 4.3 | 19.22 | 20.0305 |  |  | 3.8 | 32.6 | 33.6 |  |  | 4.9 | 78.42 | 78.4 |  |  | 7.8 | 84.55 | 84.425 |  |  | 11.4 | 84.88 | 84.98
16.1 |  |  | 4.2 | 17.3333 | 18.1974 |  |  | 3.8 | 26.3 | 27.15 |  |  | 4.8 | 75.68 | 75.625 |  |  | 7.675 | 84.3333 | 84.1 |  |  | 11.2 | 84.65 | 84.7667
16.2 |  |  | 4.2 | 15.98 | 16.6976 |  |  | 3.7 | 22.3333 | 23.15 |  |  | 4.7 | 71.525 | 71.6 |  |  | 7.575 | 83.925 | 83.9 |  |  | 10.975 | 84.52 | 84.5667
16.3 |  |  | 4.1 | 14.9 | 15.7978 |  |  | 3.7 | 19.4 | 20.1 |  |  | 4.7 | 66.15 | 66.04 |  |  | 7.4 | 83.7 | 83.5667 |  |  | 10.8333 | 84.4 | 84.3167
16.4 |  |  | 4.1 | 14.1 | 14.9408 |  |  | 3.6333 | 17.3333 | 18.1 |  |  | 4.5833 | 59.6 | 59.1 |  |  | 7.2 | 83.44 | 83.1667 |  |  | 10.64 | 84.0857 | 84.1167
16.5 |  |  | 4.0833 | 13.4333 | 14.2646 |  |  | 3.6 | 15.8333 | 16.5 |  |  | 4.5 | 51.6 | 51.8333 |  |  | 7.1 | 83 | 82.85 |  |  | 10.4833 | 83.9333 | 83.9
16.6 |  |  | 4 | 12.86 | 13.748 |  |  | 3.6 | 14.7 | 15.4833 |  |  | 4.42 | 43.94 | 44.4 |  |  | 7 | 82.6 | 82.48 |  |  | 10.3167 | 83.725 | 83.7
16.7 |  |  | 4 | 12.4333 | 13.3314 |  |  | 3.5 | 13.7667 | 14.58 |  |  | 4.4 | 36.95 | 37.0333 |  |  | 6.9 | 82.2 | 82.125 |  |  | 10.1833 | 83.5333 | 83.3333
16.8 |  |  | 3.9 | 12.04 | 12.9648 |  |  | 3.5 | 12.9857 | 13.78 |  |  | 4.3 | 30.8 | 31.06 |  |  | 6.8 | 81.7 | 81.5333 |  |  | 10.0167 | 83.26 | 83.15
16.9 |  |  | 3.9 | 11.725 | 12.6315 |  |  | 3.5 | 12.4 | 13.25 |  |  | 4.3 | 25.9 | 26.32 |  |  | 6.7 | 80.9 | 80.9333 |  |  | 9.9 | 83.0333 | 82.96
17 |  |  | 3.9 | 11.45 | 12.3482 |  |  | 3.4 | 11.8714 | 12.75 |  |  | 4.2 | 22.5 | 23.15 |  |  | 6.6 | 80.45 | 80.3 |  |  | 9.7167 | 82.8 | 82.75
17.1 |  |  | 3.8 | 11.2 | 12.1316 |  |  | 3.4 | 11.4667 | 12.35 |  |  | 4.1 | 20 | 20.875 |  |  | 6.5167 | 79.6667 | 79.44 |  |  | 9.6 | 82.45 | 82.52
17.2 |  |  | 3.8 | 11 | 11.8583 |  |  | 3.4 | 11.0857 | 11.92 |  |  | 4.1 | 18.26 | 19.05 |  |  | 6.4 | 78.775 | 78.68 |  |  | 9.4 | 82.24 | 82.25
17.3 |  |  | 3.7 | 10.78 | 11.6783 |  |  | 3.3 | 10.76 | 11.625 |  |  | 4.1 | 16.9 | 17.75 |  |  | 6.3 | 77.58 | 77.55 |  |  | 9.2 | 81.9714 | 82.02
17.4 |  |  | 3.7 | 10.58 | 11.4984 |  |  | 3.3 | 10.4857 | 11.3 |  |  | 4 | 15.9167 | 16.72 |  |  | 6.24 | 76.3 | 76.15 |  |  | 9.08 | 81.7333 | 81.72
17.5 |  |  |  | 10.4 | 11.3184 |  |  | 3.2833 | 10.2667 | 11.075 |  |  | 4 | 15.0833 | 15.9 |  |  | 6.1 | 74.3 | 74.4 |  |  | 8.95 | 81.5 | 81.4
17.6 |  |  |  | 10.2 | 11.2184 |  |  | 3.2 | 10.0667 | 10.875 |  |  | 3.9 | 14.46 | 15.18 |  |  | 6 | 71.9 | 71.775 |  |  | 8.8333 | 81.1333 | 81.06
17.7 |  |  |  | 10.1 | 11.0484 |  |  | 3.2 | 9.8167 | 10.65 |  |  | 3.9 | 13.82 | 14.68 |  |  | 5.95 | 68.32 | 68.4 |  |  | 8.66 | 80.82 | 80.8
17.8 |  |  |  | 10 | 10.8841 |  |  | 3.2 | 9.64 | 10.425 |  |  | 3.8 | 13.25 | 14.1 |  |  | 5.8333 | 63.55 | 63.58 |  |  | 8.5 | 80.6 | 80.475
17.9 |  |  |  | 9.85 | 10.7185 |  |  |  | 9.4714 | 10.3 |  |  | 3.8 | 12.8667 | 13.6667 |  |  | 5.7667 | 57.76 | 57.925 |  |  | 8.38 | 80.2 | 80.175
18 |  |  |  | 9.7 | 10.5985 |  |  |  | 9.3 | 10.16 |  |  | 3.7 | 12.5 | 13.3 |  |  | 5.6143 | 50.8 | 51.4 |  |  | 8.275 | 79.78 | 79.84
18.1 |  |  |  | 9.55 | 10.4818 |  |  |  | 9.2 | 9.94 |  |  | 3.7 | 12.075 | 12.9667 |  |  | 5.6 | 44.3 | 44.85 |  |  | 8.16 | 79.3 | 79.35
18.2 |  |  |  | 9.42 | 10.3152 |  |  |  | 9 | 9.84 |  |  | 3.6667 | 11.78 | 12.64 |  |  | 5.5 | 38.32 | 39 |  |  | 8 | 78.94 | 78.7714
18.3 |  |  |  | 9.3 | 10.1985 |  |  |  | 8.875 | 9.7 |  |  | 3.6 | 11.5 | 12.425 |  |  | 5.4 | 32.75 | 33.6 |  |  | 7.8833 | 78.3333 | 78.3
18.4 |  |  |  | 9.18 | 10.0653 |  |  |  | 8.78 | 9.5 |  |  | 3.6 | 11.275 | 12.12 |  |  | 5.4 | 28.38 | 29.2 |  |  | 7.8 | 77.78 | 77.8
18.5 |  |  |  | 9.0667 | 9.9786 |  |  |  | 8.6 | 9.35 |  |  | 3.5 | 11.02 | 11.9 |  |  | 5.3 | 25.15 | 26.05 |  |  | 7.7 | 77.1 | 77.18
18.6 |  |  |  | 8.95 | 9.7986 |  |  |  | 8.5 | 9.22 |  |  | 3.5 | 10.82 | 11.7 |  |  | 5.3 | 22.7 | 23.7 |  |  | 7.62 | 76.4 | 76.3667
18.7 |  |  |  | 8.8 | 9.7653 |  |  |  | 8.4 | 9.175 |  |  |  | 10.675 | 11.5167 |  |  | 5.2 | 21.15 | 21.975 |  |  | 7.5571 | 75.4 | 75.3667
18.8 |  |  |  | 8.7 | 9.5986 |  |  |  | 8.25 | 9.025 |  |  |  | 10.425 | 11.32 |  |  | 5.1 | 19.84 | 20.625 |  |  | 7.4 | 74.34 | 74.44
18.9 |  |  |  | 8.6 | 9.4986 |  |  |  | 8.1 | 8.9167 |  |  |  | 10.28 | 11.1 |  |  | 5.06 | 18.7 | 19.7 |  |  | 7.3 | 73.05 | 72.96
19 |  |  |  | 8.5 | 9.3653 |  |  |  | 8.0167 | 8.8 |  |  |  | 10.1 | 10.96 |  |  | 5 | 17.925 | 18.85 |  |  | 7.16 | 71.375 | 71.3
19.1 |  |  |  | 8.4 | 9.213 |  |  |  | 7.9 | 8.68 |  |  |  | 9.94 | 10.725 |  |  | 4.9333 | 17.225 | 18 |  |  | 7.04 | 68.85 | 69
19.2 |  |  |  | 8.25 | 9.0987 |  |  |  | 7.8 | 8.55 |  |  |  | 9.7333 | 10.6333 |  |  | 4.8667 | 16.6 | 17.5 |  |  | 6.9833 | 65.54 | 65.6
19.3 |  |  |  | 8.2 | 9.0487 |  |  |  | 7.7 | 8.5 |  |  |  | 9.6 | 10.5 |  |  | 4.8 | 16.05 | 16.94 |  |  | 6.9 | 61.475 | 61.6
19.4 |  |  |  | 8.025 | 8.8821 |  |  |  | 7.6 | 8.38 |  |  |  | 9.5 | 10.3333 |  |  | 4.7 | 15.55 | 16.45 |  |  | 6.88 | 56.24 | 56.625
19.5 |  |  |  | 7.975 | 8.7987 |  |  |  | 7.4167 | 8.2833 |  |  |  | 9.3 | 10.2 |  |  | 4.65 | 15.1 | 16.025 |  |  | 6.75 | 50.8 | 50.975
19.6 |  |  |  | 7.825 | 8.6988 |  |  |  | 7.3 | 8.1 |  |  |  | 9.2 | 10.075 |  |  | 4.6 | 14.85 | 15.64 |  |  | 6.7 | 44.92 | 44.9
19.7 |  |  |  | 7.7 | 8.5988 |  |  |  | 7.25 | 8.0333 |  |  |  | 9.05 | 9.875 |  |  | 4.55 | 14.45 | 15.28 |  |  | 6.575 | 39.1 | 39.7
19.8 |  |  |  | 7.6 | 8.4655 |  |  |  | 7.18 | 7.9 |  |  |  | 8.94 | 9.75 |  |  | 4.5 | 14.14 | 15.1 |  |  | 6.5 | 34.3 | 34.68
19.9 |  |  |  | 7.5 | 8.3988 |  |  |  | 7.08 | 7.8 |  |  |  | 8.8 | 9.6 |  |  | 4.5 | 13.85 | 14.75 |  |  | 6.4 | 30.2 | 30.86
20 |  |  |  | 7.42 | 8.2417 |  |  |  | 6.9833 | 7.75 |  |  |  | 8.7 | 9.48 |  |  | 4.4 | 13.6 | 14.48 |  |  | 6.3 | 27 | 27.9
20.1 |  |  |  | 7.3 | 8.1845 |  |  |  | 6.9 | 7.6 |  |  |  | 8.575 | 9.38 |  |  | 4.4 | 13.36 | 14.22 |  |  | 6.2 | 24.88 | 25.7333
20.2 |  |  |  | 7.2 | 8.0417 |  |  |  | 6.8 | 7.5 |  |  |  | 8.45 | 9.2 |  |  | 4.3 | 13.1333 | 14.0167 |  |  | 6.1333 | 23.28 | 24.02
20.3 |  |  |  | 7.1 | 7.9655 |  |  |  | 6.7 | 7.45 |  |  |  | 8.32 | 9.1 |  |  | 4.3 | 12.9 | 13.8 |  |  | 6.1 | 21.85 | 22.8
20.4 |  |  |  | 7.0167 | 7.8132 |  |  |  | 6.5833 | 7.3 |  |  |  | 8.2 | 9 |  |  |  | 12.66 | 13.55 |  |  | 6 | 20.9 | 21.775
20.5 |  |  |  | 6.9 | 7.6989 |  |  |  | 6.5 | 7.2 |  |  |  | 8.1 | 8.9 |  |  |  | 12.5 | 13.35 |  |  | 5.9333 | 19.96 | 20.8667
20.6 |  |  |  | 6.8 | 7.6114 |  |  |  | 6.4 | 7.1 |  |  |  | 8 | 8.8 |  |  |  | 12.3 | 13.125 |  |  | 5.9 | 19.18 | 20.225
20.7 |  |  |  | 6.7 | 7.4989 |  |  |  | 6.3 | 7.1 |  |  |  | 7.875 | 8.65 |  |  |  | 12.1 | 12.975 |  |  | 5.8 | 18.575 | 19.55
20.8 |  |  |  | 6.5333 | 7.3789 |  |  |  | 6.2143 | 7 |  |  |  | 7.725 | 8.5 |  |  |  | 11.9 | 12.75 |  |  | 5.7167 | 18.05 | 19
20.9 |  |  |  | 6.46 | 7.299 |  |  |  | 6.1 | 6.85 |  |  |  | 7.6 | 8.4667 |  |  |  | 11.7 | 12.6 |  |  | 5.7 | 17.55 | 18.55
21 |  |  |  | 6.35 | 7.139 |  |  |  | 6.0143 | 6.8 |  |  |  | 7.5 | 8.28 |  |  |  | 11.5 | 12.45 |  |  | 5.6 | 17.14 | 18.15
21.1 |  |  |  | 6.225 | 7.0704 |  |  |  | 5.85 | 6.7 |  |  |  | 7.4 | 8.2 |  |  |  | 11.36 | 12.22 |  |  | 5.6 | 16.7 | 17.74
21.2 |  |  |  | 6.1167 | 6.9365 |  |  |  | 5.8 | 6.58 |  |  |  | 7.3 | 8.0833 |  |  |  | 11.225 | 12.1 |  |  | 5.5 | 16.4333 | 17.4
21.3 |  |  |  | 6.0167 | 6.8823 |  |  |  | 5.7 | 6.4 |  |  |  | 7.2 | 7.98 |  |  |  | 11.025 | 11.9 |  |  | 5.4286 | 16.1 | 17.075
21.4 |  |  |  | 5.9167 | 6.799 |  |  |  | 5.6 | 6.3833 |  |  |  | 7.1 | 7.9 |  |  |  | 10.875 | 11.7 |  |  | 5.3 | 15.86 | 16.8333
21.5 |  |  |  | 5.85 | 6.699 |  |  |  | 5.5 | 6.3 |  |  |  | 7 | 7.8 |  |  |  | 10.6833 | 11.58 |  |  | 5.2 | 15.55 | 16.475
21.6 |  |  |  | 5.7667 | 6.5991 |  |  |  | 5.5 | 6.2 |  |  |  | 6.9 | 7.675 |  |  |  | 10.58 | 11.4 |  |  | 5.2 | 15.32 | 16.26
21.7 |  |  |  | 5.7 | 6.4991 |  |  |  | 5.4 | 6.1 |  |  |  | 6.8 | 7.6 |  |  |  | 10.4 | 11.26 |  |  | 5.2 | 15.025 | 16
21.8 |  |  |  | 5.6 | 6.3991 |  |  |  | 5.2667 | 6 |  |  |  | 6.7 | 7.475 |  |  |  | 10.26 | 11.15 |  |  | 5.1 | 14.82 | 15.76
21.9 |  |  |  | 5.52 | 6.2991 |  |  |  | 5.1833 | 5.9 |  |  |  | 6.6 | 7.3667 |  |  |  | 10.15 | 11 |  |  | 5.1 | 14.625 | 15.525
22 |  |  |  | 5.5 | 6.2991 |  |  |  | 5.1 | 5.9 |  |  |  | 6.5 | 7.2143 |  |  |  | 10 | 10.8167 |  |  | 5 | 14.375 | 15.325
22.1 |  |  |  | 5.4 | 6.1991 |  |  |  | 5.1 | 5.7167 |  |  |  | 6.38 | 7.12 |  |  |  | 9.8 | 10.72 |  |  |  | 14.1 | 15.1
22.2 |  |  |  | 5.3 | 6.0991 |  |  |  | 4.95 | 5.65 |  |  |  | 6.28 | 7 |  |  |  | 9.725 | 10.6 |  |  |  | 13.9 | 14.85
22.3 |  |  |  | 5.2333 | 6.0991 |  |  |  | 4.9 | 5.6 |  |  |  | 6.2 | 7 |  |  |  | 9.5 | 10.4667 |  |  |  | 13.75 | 14.725
22.4 |  |  |  | 5.2 | 5.9563 |  |  |  | 4.8 | 5.5 |  |  |  | 6.1 | 6.88 |  |  |  | 9.4 | 10.3 |  |  |  | 13.55 | 14.48
22.5 |  |  |  | 5.1 | 5.8992 |  |  |  | 4.8 | 5.5 |  |  |  | 6 | 6.7 |  |  |  | 9.3 | 10.2 |  |  |  | 13.3333 | 14.36
22.6 |  |  |  | 5.1 |  |  |  |  | 4.7 | 5.4 |  |  |  | 5.9 | 6.7 |  |  |  | 9.1667 | 10.1 |  |  |  | 13.2 | 14.2
22.7 | [8 empty] | 4.6333 | 5.4 |  |  |  | 5.8 | 6.55 |  |  |  | 9 | 9.9 |  |  |  | 13 | 13.9667
22.8 | [8 empty] | 4.6 | 5.3 |  |  |  | 5.7333 | 6.5 |  |  |  | 8.9 | 9.8 |  |  |  | 12.8667 | 13.7667
22.9 | [8 empty] | 4.5 | 5.2667 |  |  |  | 5.66 | 6.4 |  |  |  | 8.8 | 9.6 |  |  |  | 12.65 | 13.68

## Wedge 15 Longitudinal Profiles
Longitudinal Profiles
Wedge 15
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:3 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 7.8556
-34.9 |  |  |  |  | 7.9
-34.8 |  |  |  |  | 7.97
-34.7 |  |  |  |  | 8
-34.6 |  |  |  |  | 8.1
-34.5 |  |  |  |  | 8.2
-34.4 |  |  |  |  | 8.2571
-34.3 |  |  |  |  | 8.3
-34.2 |  |  |  |  | 8.4
-34.1 |  |  |  |  | 8.5
-34 |  |  |  |  | 8.5714
-33.9 |  |  |  |  | 8.7
-33.8 |  |  |  |  | 8.8
-33.7 |  |  |  |  | 8.82
-33.6 |  |  |  |  | 8.92
-33.5 |  |  |  |  | 9.02
-33.4 |  |  |  |  | 9.12
-33.3 |  |  |  |  | 9.2182
-33.2 |  |  |  |  | 9.3
-33.1 |  |  |  |  | 9.31
-33 |  |  |  |  | 9.41
-32.9 |  |  |  |  | 9.5091
-32.8 |  |  |  |  | 9.6
-32.7 |  |  |  |  | 9.7
-32.6 |  |  |  |  | 9.8
-32.5 |  |  |  |  | 9.9
-32.4 |  |  |  |  | 10
-32.3 |  |  |  | 7.0875 | 10.1125
-32.2 |  |  |  | 7.1125 | 10.26
-32.1 |  |  |  | 7.2 | 10.43
-32 |  |  |  | 7.2625 | 10.53
-31.9 |  |  |  | 7.3875 | 10.63
-31.8 |  |  |  | 7.4 | 10.73
-31.7 |  |  |  | 7.5 | 10.83
-31.6 |  |  |  | 7.5 | 10.93
-31.5 |  |  |  | 7.6429 | 11.0429
-31.4 |  |  |  | 7.7 | 11.1857
-31.3 |  |  |  | 7.8 | 11.3
-31.2 |  |  |  | 7.85 | 11.3714
-31.1 |  |  |  | 7.9 | 11.5222
-31 |  |  |  | 8 | 11.6333
-30.9 |  |  |  | 8.1125 | 11.8143
-30.8 |  |  |  | 8.2 | 11.97
-30.7 |  |  |  | 8.2625 | 12.14
-30.6 |  |  |  | 8.3875 | 12.27
-30.5 |  |  |  | 8.46 | 12.37
-30.4 |  |  |  | 8.6 | 12.47
-30.3 |  |  |  | 8.7 | 12.64
-30.2 |  |  |  | 8.7667 | 12.8273
-30.1 |  |  |  | 8.8429 | 12.96
-30 |  |  |  | 8.9 | 13.06
-29.9 |  |  |  | 9.0222 | 13.22
-29.8 |  |  |  | 9.1375 | 13.42
-29.7 |  |  |  | 9.2 | 13.56
-29.6 |  |  | 6.3125 | 9.2875 | 13.72
-29.5 |  |  | 6.4 | 9.4375 | 13.9091
-29.4 |  |  | 6.4 | 9.5625 | 14.1
-29.3 |  |  | 6.5 | 9.7 | 14.25
-29.2 |  |  | 6.5 | 9.7375 | 14.4
-29.1 |  |  | 6.6 | 9.8625 | 14.6
-29 |  |  | 6.7 | 9.9778 | 14.8
-28.9 |  |  | 6.7 | 10.1111 | 14.9909
-28.8 |  |  | 6.775 | 10.2333 | 15.18
-28.7 |  |  | 6.8429 | 10.4 | 15.4
-28.6 |  |  | 6.9 | 10.4714 | 15.62
-28.5 |  |  | 7 | 10.625 | 15.82
-28.4 |  |  | 7.0714 | 10.7571 | 16.02
-28.3 |  | 5.9 | 7.1375 | 10.9143 | 16.28
-28.2 |  | 6 | 7.2 | 11 | 16.52
-28.1 |  | 6 | 7.2667 | 11.1667 | 16.7091
-28 |  | 6.0875 | 7.3143 | 11.3125 | 16.95
-27.9 |  | 6.1 | 7.4 | 11.4429 | 17.25
-27.8 |  | 6.1667 | 7.5 | 11.5857 | 17.55
-27.7 | 6.3 | 6.2 | 7.58 | 11.7333 | 17.825
-27.6 | 6.4 | 6.28 | 7.7 | 11.875 | 18.11
-27.5 | 6.4143 | 6.3 | 7.7 | 12.1 | 18.48
-27.4 | 6.5 | 6.3143 | 7.8 | 12.2 | 18.88
-27.3 | 6.6 | 6.4 | 7.9286 | 12.4 | 19.3182
-27.2 | 6.6 | 6.5 | 8 | 12.55 | 19.8333
-27.1 | 6.65 | 6.5 | 8.1 | 12.7 | 20.3889
-27 | 6.7 | 6.6 | 8.1875 | 12.925 | 21.06
-26.9 | 6.8 | 6.6714 | 8.3 | 13.0667 | 21.84
-26.8 | 6.8167 | 6.7286 | 8.3833 | 13.2778 | 22.8
-26.7 | 6.9 | 6.8 | 8.5 | 13.46 | 23.8
-26.6 | 6.9667 | 6.9 | 8.5571 | 13.65 | 25.2667
-26.5 | 7.0286 | 6.9143 | 8.7 | 13.84 | 27.05
-26.4 | 7.1 | 7 | 8.7667 | 14.0333 | 29.175
-26.3 | 7.2 | 7.1 | 8.9286 | 14.2833 | 32.1778
-26.2 | 7.2875 | 7.175 | 9 | 14.525 | 35.9714
-26.1 | 7.3143 | 7.2167 | 9.1 | 14.7143 | 40.1857
-26 | 7.4 | 7.3 | 9.2429 | 15.0286 | 45.0429
-25.9 | 7.5 | 7.4 | 9.4 | 15.2571 | 50.6556
-25.8 | 7.5875 | 7.5 | 9.4833 | 15.5 | 56.5375
-25.7 | 7.6571 | 7.5667 | 9.6 | 15.86 | 61.7
-25.6 | 7.7333 | 7.6286 | 9.7 | 16.1375 | 66.45
-25.5 | 7.8 | 7.7833 | 9.8667 | 16.5125 | 70.2625
-25.4 | 7.9 | 7.85 | 10.0286 | 16.8875 | 73.4143
-25.3 | 8 | 7.9 | 10.1 | 17.3625 | 75.5
-25.2 | 8.0857 | 8 | 10.3 | 17.95 | 77.1727
-25.1 | 8.2 | 8.0875 | 10.425 | 18.5125 | 78.6222
-25 | 8.3 | 8.2 | 10.55 | 19.1375 | 79.725
-24.9 | 8.3286 | 8.3 | 10.75 | 20.0429 | 80.6667
-24.8 | 8.4714 | 8.4 | 10.9167 | 21 | 81.3556
-24.7 | 8.5286 | 8.5 | 11.0714 | 22.2286 | 82.125
-24.6 | 8.6714 | 8.6167 | 11.3 | 23.8 | 82.75
-24.5 | 8.7167 | 8.7 | 11.45 | 25.7429 | 83.18
-24.4 | 8.8833 | 8.8143 | 11.6 | 28.4 | 83.72
-24.3 | 8.95 | 8.9 | 11.8143 | 31.65 | 84.3
-24.2 | 9.1 | 9.0167 | 11.9667 | 35.9125 | 84.7
-24.1 | 9.2 | 9.1833 | 12.1571 | 41.225 | 85.05
-24 | 9.2571 | 9.3333 | 12.4667 | 47.3857 | 85.38
-23.9 | 9.4 | 9.4857 | 12.675 | 53.8 | 85.65
-23.8 | 9.5167 | 9.6 | 12.9667 | 59.9 | 85.9667
-23.7 | 9.6167 | 9.6857 | 13.2 | 65.9 | 86.325
-23.6 | 9.7714 | 9.85 | 13.5 | 70.7 | 86.575
-23.5 | 9.9 | 10.0167 | 13.875 | 74.625 | 86.8667
-23.4 | 10.0333 | 10.1833 | 14.2143 | 77.9 | 87.12
-23.3 | 10.2 | 10.3429 | 14.7 | 80.3 | 87.35
-23.2 | 10.2833 | 10.5167 | 15.2 | 82.05 | 87.6667
-23.1 | 10.4429 | 10.6833 | 15.8333 | 83.6 | 87.85
-23 | 10.6167 | 10.84 | 16.5 | 84.5 | 88.1
-22.9 | 10.7714 | 11.1167 | 17.4 | 85.4 | 88.2667
-22.8 | 10.9333 | 11.3429 | 18.5625 | 86.15 | 88.4556
-22.7 | 11.1 | 11.54 | 19.775 | 86.825 | 88.58
-22.6 | 11.26 | 11.8286 | 21.5429 | 87.3571 | 88.76
-22.5 | 11.52 | 12.1333 | 23.7571 | 87.9286 | 88.96
-22.4 | 11.65 | 12.4286 | 26.8333 | 88.4375 | 89.1
-22.3 | 11.9143 | 12.8571 | 31.075 | 88.85 | 89.2818
-22.2 | 12.14 | 13.3 | 36.0429 | 89.3667 | 89.48
-22.1 | 12.3857 | 13.9 | 42.7 | 89.7 | 89.59
-22 | 12.7286 | 14.6 | 49.9286 | 89.9857 | 89.78
-21.9 | 13.1 | 15.5 | 57.4 | 90.25 | 89.98
-21.8 | 13.55 | 16.5 | 64.9286 | 90.5667 | 90.18
-21.7 | 14.0667 | 17.8714 | 72 | 90.8286 | 90.38
-21.6 | 14.7571 | 19.6667 | 77.85 | 91.16 | 90.6111
-21.5 | 15.6714 | 22 | 82.2 | 91.35 | 90.75
-21.4 | 16.7 | 25.3571 | 85.1286 | 91.5286 | 90.9889
-21.3 | 17.98 | 30 | 87.3 | 91.7667 | 91.225
-21.2 | 19.8 | 36.1 | 89 | 91.9333 | 91.3778
-21.1 | 22.1833 | 43.1333 | 90.28 | 92.0857 | 91.5
-21 | 25.7 | 50.8714 | 91.2 | 92.2286 | 91.7
-20.9 | 30.4 | 59.6 | 92.0333 | 92.3625 | 91.8429
-20.8 | 37 | 68.1 | 92.8333 | 92.5 | 91.9889
-20.7 | 44.4333 | 75.8143 | 93.5 | 92.6833 | 92.18
-20.6 | 52.5833 | 81.9 | 93.9714 | 92.85 | 92.4
-20.5 | 61.5833 | 86.1667 | 94.425 | 93 | 92.5333
-20.4 | 70.08 | 89.3 | 94.875 | 93.0667 | 92.675
-20.3 | 77.7333 | 91.38 | 95.2 | 93.25 | 92.9
-20.2 | 83.8286 | 93.0667 | 95.4714 | 93.4 | 93.0125
-20.1 | 88.2857 | 94.2 | 95.8143 | 93.6667 | 93.1375
-20 | 91 | 95.05 | 96.1167 | 93.8286 | 93.3
-19.9 | 92.9667 | 95.9571 | 96.3429 | 93.9714 | 93.5111
-19.8 | 94.4667 | 96.5667 | 96.625 | 94.1286 | 93.66
-19.7 | 95.5714 | 97.1375 | 96.7571 | 94.2833 | 93.7857
-19.6 | 96.3333 | 97.55 | 96.92 | 94.4571 | 93.92
-19.5 | 97.0333 | 97.9286 | 97.075 | 94.6 | 94.0364
-19.4 | 97.6 | 98.3 | 97.12 | 94.7429 | 94.2111
-19.3 | 97.9667 | 98.6571 | 97.2 | 94.8857 | 94.3222
-19.2 | 98.42 | 98.8833 | 97.3 | 95.1571 | 94.43
-19.1 | 98.75 | 99.2143 | 97.4 | 95.3 | 94.53
-19 | 99.0167 | 99.4333 | 97.4714 | 95.425 | 94.6375
-18.9 | 99.2667 | 99.6 | 97.6 | 95.5571 | 94.7714
-18.8 | 99.45 | 99.66 | 97.7 | 95.7 | 94.8143
-18.7 | 99.6 | 99.7 | 97.8 | 95.8333 | 94.9444
-18.6 | 99.7667 | 99.8143 | 97.8429 | 95.9667 | 95.0625
-18.5 | 99.825 | 99.9 | 98.0167 | 96.175 | 95.1
-18.4 | 99.9 | 99.96 | 98.2 | 96.3125 | 95.2429
-18.3 | 99.9333 | 100 | 98.3 | 96.4375 | 95.3857
-18.2 | 100 | 100 | 98.425 | 96.5625 | 95.5222
-18.1 | 100 | 100.02 | 98.5167 | 96.7 | 95.6429
-18 | 100 | 100.1 | 98.6 | 96.7429 | 95.7
-17.9 | 100.04 | 100.1 | 98.7571 | 96.8 | 95.78
-17.8 | 100.1 | 100.2 | 98.9 | 96.95 | 95.8727
-17.7 | 100.1667 | 100.2714 | 99.05 | 97.1 | 95.9778
-17.6 | 100.2 | 100.3 | 99.1857 | 97.1857 | 96.1
-17.5 | 100.2 | 100.32 | 99.2429 | 97.3 | 96.2333
-17.4 | 100.3 | 100.5 | 99.4 | 97.3375 | 96.375
-17.3 | 100.4 | 100.55 | 99.5 | 97.4714 | 96.4
-17.2 | 100.46 | 100.6 | 99.6143 | 97.55 | 96.525
-17.1 | 100.52 | 100.7833 | 99.7 | 97.6875 | 96.68
-17 | 100.6 | 100.9 | 99.8 | 97.8 | 96.8
-16.9 | 100.7 | 101.0286 | 99.9667 | 97.8625 | 96.875
-16.8 | 100.775 | 101.1 | 100 | 97.9 | 97
-16.7 | 100.9 | 101.2 | 100.1 | 98 | 97.12
-16.6 | 101 | 101.2857 | 100.15 | 98 | 97.2222
-16.5 | 101 | 101.4286 | 100.2 | 98.0833 | 97.33
-16.4 | 101.1333 | 101.5 | 100.2 | 98.15 | 97.4333
-16.3 | 101.2 | 101.6 | 100.3 | 98.3 | 97.55
-16.2 | 101.3 | 101.6 | 100.3833 | 98.4286 | 97.6667
-16.1 | 101.4 | 101.6857 | 100.5 | 98.5 | 97.84
-16 | 101.4 | 101.74 | 100.5 | 98.6 | 97.9
-15.9 | 101.5 | 101.8 | 100.52 | 98.7 | 98
-15.8 | 101.6 | 101.8 | 100.6 | 98.8 | 98.0375
-15.7 | 101.7 | 101.9 | 100.6 | 98.8571 | 98.1556
-15.6 | 101.7833 | 101.9 | 100.65 | 98.9 | 98.26
-15.5 | 101.8 | 101.94 | 100.7 | 99 | 98.36
-15.4 | 101.8 | 102 | 100.8 | 99.075 | 98.4
-15.3 | 101.8 | 102 | 100.8 | 99.125 | 98.46
-15.2 | 101.9 | 102 | 100.8 | 99.2667 | 98.5667
-15.1 | 101.9 | 102.1 | 100.9 | 99.35 | 98.6778
-15 | 101.9 | 102.1 | 100.95 | 99.5 | 98.8
-14.9 | 101.9 | 102.1 | 101 | 99.6 | 98.825
-14.8 | 101.9875 | 102.1 | 101.1 | 99.6 | 98.95
-14.7 | 102 | 102.1 | 101.1333 | 99.7 | 99
-14.6 | 102 | 102.1 | 101.2 | 99.7 | 99.1
-14.5 | 102 | 102.1 | 101.2 | 99.8 | 99.1333
-14.4 | 102 | 102.1375 | 101.2875 | 99.875 | 99.2
-14.3 | 102 | 102.2 | 101.4 | 99.9167 | 99.2875
-14.2 | 102 | 102.3 | 101.4 | 100 | 99.3
-14.1 | 102.0333 | 102.3 | 101.5 | 100.1 | 99.4
-14 | 102.1 | 102.3 | 101.6 | 100.1875 | 99.42
-13.9 | 102.1 | 102.4 | 101.6 | 100.2125 | 99.52
-13.8 | 102.1 | 102.4 | 101.6857 | 100.3375 | 99.625
-13.7 | 102.2 | 102.4 | 101.7 | 100.4 | 99.7
-13.6 | 102.2 | 102.4333 | 101.8 | 100.4667 | 99.7875
-13.5 | 102.3 | 102.5 | 101.8571 | 100.5 | 99.8
-13.4 | 102.35 | 102.55 | 101.9 | 100.5857 | 99.8
-13.3 | 102.4 | 102.6 | 101.9429 | 100.625 | 99.8
-13.2 | 102.4 | 102.7 | 102 | 100.7 | 99.8
-13.1 | 102.5 | 102.7 | 102 | 100.7 | 99.8545
-13 | 102.5125 | 102.8 | 102 | 100.8 | 99.9
-12.9 | 102.6 | 102.8667 | 102 | 100.8286 | 99.9
-12.8 | 102.6 | 102.9 | 102.1 | 100.9 | 99.9
-12.7 | 102.7 | 102.9 | 102.1 | 100.9 | 99.9
-12.6 | 102.7 | 102.9571 | 102.1 | 100.9857 | 99.9
-12.5 | 102.7833 | 102.9 | 102.2 | 101 | 99.9
-12.4 | 102.8 | 103 | 102.26 | 101 | 99.9636
-12.3 | 102.8 | 103 | 102.3 | 101 | 100
-12.2 | 102.8 | 103 | 102.3 | 101 | 100.075
-12.1 | 102.8 | 103 | 102.4 | 101 | 100.1
-12 | 102.8 | 103 | 102.4 | 101 | 100.2
-11.9 | 102.8571 | 103 | 102.4429 | 101 | 100.2
-11.8 | 102.9 | 103.1 | 102.5 | 101 | 100.3
-11.7 | 102.9 | 103.1 | 102.5 | 101 | 100.33
-11.6 | 102.9 | 103.2 | 102.5 | 101 | 100.4375
-11.5 | 102.98 | 103.2 | 102.5 | 101 | 100.5
-11.4 | 103 | 103.2 | 102.5 | 101.1 | 100.5
-11.3 | 103.0167 | 103.2 | 102.4429 | 101.1 | 100.5889
-11.2 | 103.1 | 103.2 | 102.4 | 101.1 | 100.6
-11.1 | 103.1 | 103.2 | 102.4 | 101.1143 | 100.7
-11 | 103.1 | 103.1667 | 102.3571 | 101.2 | 100.7222
-10.9 | 103 | 103.1 | 102.3 | 101.2 | 100.8
-10.8 | 103 | 103.1 | 102.3 | 101.3 | 100.8
-10.7 | 103 | 103.1 | 102.3 | 101.3 | 100.8875
-10.6 | 102.9833 | 103.0286 | 102.3 | 101.4 | 100.9
-10.5 | 102.9 | 103 | 102.3 | 101.4429 | 100.9
-10.4 | 102.86 | 102.9833 | 102.3 | 101.5 | 100.9111
-10.3 | 102.8 | 102.9 | 102.3 | 101.5429 | 101
-10.2 | 102.8 | 102.8571 | 102.35 | 101.6 | 101
-10.1 | 102.7 | 102.8 | 102.4 | 101.6667 | 101
-10 | 102.7 | 102.8 | 102.4 | 101.7 | 101
-9.9 | 102.7 | 102.85 | 102.4714 | 101.7 | 101.1
-9.8 | 102.7 | 102.9 | 102.5 | 101.8 | 101.1
-9.7 | 102.7 | 102.9 | 102.6 | 101.8 | 101.1
-9.6 | 102.8 | 103 | 102.6 | 101.8 | 101.2
-9.5 | 102.8 | 103 | 102.6333 | 101.8 | 101.2125
-9.4 | 102.9 | 103.1 | 102.7 | 101.8 | 101.3429
-9.3 | 102.9 | 103.1 | 102.7 | 101.8375 | 101.4
-9.2 | 102.9 | 103.1 | 102.7 | 101.9 | 101.4
-9.1 | 102.9857 | 103.1 | 102.7 | 101.925 | 101.4
-9 | 103 | 103.2 | 102.7 | 102 | 101.4
-8.9 | 103.0125 | 103.2 | 102.7167 | 102 | 101.4
-8.8 | 103.1 | 103.2 | 102.8 | 102 | 101.4
-8.7 | 103.1 | 103.2 | 102.8 | 102 | 101.5
-8.6 | 103.1 | 103.2 | 102.8 | 102 | 101.5
-8.5 | 103.1 | 103.2 | 102.86 | 102 | 101.5
-8.4 | 103.1 | 103.22 | 102.8833 | 102.0625 | 101.56
-8.3 | 103.1 | 103.3 | 102.8 | 102.1 | 101.6
-8.2 | 103.2 | 103.25 | 102.8 | 102.1 | 101.6
-8.1 | 103.2 | 103.2286 | 102.9 | 102.1 | 101.6667
-8 | 103.2 | 103.3 | 102.9 | 102.2 | 101.7
-7.9 | 103.2 | 103.3 | 102.9 | 102.2 | 101.8
-7.8 | 103.2 | 103.3 | 102.9 | 102.3 | 101.8
-7.7 | 103.2 | 103.3 | 102.9 | 102.3 | 101.85
-7.6 | 103.2 | 103.3 | 102.9 | 102.3 | 101.9
-7.5 | 103.2 | 103.3 | 102.9143 | 102.3 | 101.9
-7.4 | 103.2 | 103.3 | 103 | 102.3 | 101.9
-7.3 | 103.3 | 103.3 | 103 | 102.3 | 101.9
-7.2 | 103.3 | 103.3 | 103.1 | 102.3 | 101.9
-7.1 | 103.3 | 103.3 | 103.1 | 102.3571 | 101.9
-7 | 103.3 | 103.3 | 103.1 | 102.4 | 101.9571
-6.9 | 103.3 | 103.3 | 103.1 | 102.4 | 102
-6.8 | 103.3 | 103.3 | 103.1 | 102.4 | 102
-6.7 | 103.3 | 103.4 | 103.1 | 102.44 | 102
-6.6 | 103.3 | 103.4 | 103.1 | 102.5 | 102
-6.5 | 103.3 | 103.4 | 103.1 | 102.5 | 102
-6.4 | 103.3 | 103.4 | 103.1 | 102.4429 | 102
-6.3 | 103.3 | 103.4 | 103.1 | 102.5 | 101.9625
-6.2 | 103.3 | 103.4 | 103.1 | 102.5 | 101.9
-6.1 | 103.3 | 103.4 | 103.1 | 102.4667 | 101.9
-6 | 103.3 | 103.4 | 103.1 | 102.4 | 101.9
-5.9 | 103.3 | 103.4 | 103.1 | 102.4571 | 101.9
-5.8 | 103.4 | 103.4 | 103.1 | 102.4 | 101.9
-5.7 | 103.3 | 103.4 | 103.1 | 102.4 | 101.9
-5.6 | 103.3 | 103.4 | 103.1 | 102.4 | 102
-5.5 | 103.3 | 103.4 | 103.1 | 102.4 | 102
-5.4 | 103.3 | 103.4 | 103 | 102.3167 | 101.97
-5.3 | 103.3 | 103.3667 | 102.9333 | 102.3 | 101.9
-5.2 | 103.3 | 103.3 | 102.9 | 102.3 | 101.9
-5.1 | 103.3 | 103.3 | 102.9 | 102.3 | 101.9
-5 | 103.2 | 103.2125 | 102.9 | 102.3 | 101.8667
-4.9 | 103.2 | 103.2 | 102.9 | 102.3 | 101.8
-4.8 | 103.2 | 103.2 | 102.9 | 102.2 | 101.8
-4.7 | 103.1 | 103.1 | 102.8 | 102.2 | 101.8
-4.6 | 103.1 | 103.1 | 102.8 | 102.2 | 101.8
-4.5 | 103.0125 | 103 | 102.76 | 102.1333 | 101.8
-4.4 | 103 | 103 | 102.7 | 102.1 | 101.8
-4.3 | 102.9 | 102.9 | 102.6286 | 102.1 | 101.8
-4.2 | 102.9 | 102.9 | 102.6 | 102.1 | 101.8
-4.1 | 102.8 | 102.8667 | 102.5 | 102 | 101.8
-4 | 102.7833 | 102.8 | 102.5 | 102 | 101.7444
-3.9 | 102.7 | 102.8 | 102.5 | 102 | 101.7
-3.8 | 102.7 | 102.7 | 102.5 | 102 | 101.7
-3.7 | 102.6 | 102.7 | 102.4 | 102 | 101.7
-3.6 | 102.6 | 102.6 | 102.4 | 101.9667 | 101.7
-3.5 | 102.6 | 102.6 | 102.35 | 101.9 | 101.625
-3.4 | 102.5 | 102.54 | 102.3 | 101.8667 | 101.6
-3.3 | 102.4833 | 102.5 | 102.2625 | 101.8 | 101.5
-3.2 | 102.4 | 102.4167 | 102.1167 | 101.7 | 101.4
-3.1 | 102.35 | 102.3667 | 102.1 | 101.65 | 101.4
-3 | 102.3 | 102.3 | 102.025 | 101.6 | 101.325
-2.9 | 102.2 | 102.2 | 102 | 101.5 | 101.2
-2.8 | 102.1 | 102.1167 | 101.9 | 101.46 | 101.1857
-2.7 | 102 | 101.9571 | 101.74 | 101.3833 | 101.05
-2.6 | 101.9167 | 101.9 | 101.6 | 101.3 | 101
-2.5 | 101.8333 | 101.8 | 101.5167 | 101.2 | 100.9333
-2.4 | 101.6833 | 101.68 | 101.4625 | 101.1 | 100.9
-2.3 | 101.5167 | 101.5143 | 101.3833 | 101.1 | 100.8273
-2.2 | 101.4429 | 101.34 | 101.3 | 100.95 | 100.8
-2.1 | 101.3 | 101.2667 | 101.2 | 100.8 | 100.8
-2 | 101.2143 | 101.2 | 101.1143 | 100.8 | 100.7111
-1.9 | 101.06 | 101.1333 | 101 | 100.7429 | 100.61
-1.8 | 100.9667 | 101 | 100.86 | 100.7 | 100.6
-1.7 | 100.8833 | 101 | 100.78 | 100.65 | 100.5667
-1.6 | 100.8 | 100.92 | 100.7 | 100.6 | 100.5
-1.5 | 100.7 | 100.8429 | 100.6286 | 100.6 | 100.5
-1.4 | 100.7 | 100.7833 | 100.6 | 100.5 | 100.5
-1.3 | 100.6 | 100.6286 | 100.5 | 100.5 | 100.4
-1.2 | 100.5 | 100.5 | 100.5 | 100.4 | 100.3
-1.1 | 100.4833 | 100.5 | 100.4 | 100.3 | 100.2556
-1 | 100.4 | 100.4 | 100.3 | 100.3 | 100.2
-0.9 | 100.3 | 100.4 | 100.2571 | 100.2 | 100.2
-0.8 | 100.3 | 100.3 | 100.1857 | 100.1 | 100.1556
-0.7 | 100.2 | 100.2 | 100.1 | 100.1 | 100.1
-0.6 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1
-0.5 | 100.1 | 100.1 | 100.0143 | 100.1 | 100.1
-0.4 | 100.1 | 100.0667 | 100 | 100.1 | 100.025
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100 | 100
0.4 | 100.0667 | 100.1 | 100 | 100.0857 | 100
0.5 | 100.1 | 100.1 | 100 | 100.1 | 100.1
0.6 | 100.1 | 100.1167 | 100.1 | 100.1 | 100.1
0.7 | 100.2 | 100.2 | 100.1 | 100.1 | 100.1
0.8 | 100.28 | 100.3 | 100.1667 | 100.14 | 100.125
0.9 | 100.3 | 100.3333 | 100.2286 | 100.2125 | 100.2
1 | 100.3833 | 100.4 | 100.3 | 100.3 | 100.2
1.1 | 100.46 | 100.5 | 100.4 | 100.3333 | 100.2222
1.2 | 100.5 | 100.5 | 100.48 | 100.4 | 100.3375
1.3 | 100.6 | 100.6 | 100.56 | 100.5 | 100.4
1.4 | 100.6833 | 100.7143 | 100.6 | 100.5143 | 100.475
1.5 | 100.7375 | 100.8667 | 100.7 | 100.6 | 100.5
1.6 | 100.8 | 100.94 | 100.7 | 100.6 | 100.5125
1.7 | 100.9 | 101 | 100.75 | 100.7 | 100.6
1.8 | 101 | 101 | 100.84 | 100.7 | 100.6
1.9 | 101.1 | 101.1571 | 101 | 100.7 | 100.6625
2 | 101.2 | 101.2 | 101.0833 | 100.8 | 100.7
2.1 | 101.35 | 101.2857 | 101.2 | 100.8571 | 100.7875
2.2 | 101.4167 | 101.3667 | 101.2833 | 101 | 100.8
2.3 | 101.5 | 101.5286 | 101.36 | 101.08 | 100.8818
2.4 | 101.66 | 101.7 | 101.4429 | 101.14 | 100.9
2.5 | 101.8 | 101.7833 | 101.5 | 101.2286 | 101
2.6 | 102 | 101.9 | 101.6571 | 101.3 | 101
2.7 | 102 | 101.9857 | 101.7167 | 101.3333 | 101.1
2.8 | 102.1 | 102.14 | 101.9 | 101.4167 | 101.1444
2.9 | 102.1833 | 102.2 | 101.9667 | 101.5 | 101.2625
3 | 102.2667 | 102.3 | 102 | 101.575 | 101.3
3.1 | 102.3167 | 102.3857 | 102.1 | 101.7 | 101.4
3.2 | 102.4 | 102.4 | 102.2 | 101.74 | 101.45
3.3 | 102.5 | 102.5 | 102.2375 | 101.8 | 101.5
3.4 | 102.54 | 102.5 | 102.3 | 101.9 | 101.5667
3.5 | 102.6 | 102.6 | 102.3286 | 101.9 | 101.6
3.6 | 102.6 | 102.6 | 102.4 | 102 | 101.7
3.7 | 102.6333 | 102.7 | 102.4 | 102 | 101.7
3.8 | 102.7 | 102.72 | 102.5 | 102 | 101.7
3.9 | 102.7 | 102.8 | 102.5 | 102 | 101.7
4 | 102.75 | 102.8 | 102.5 | 102 | 101.7
4.1 | 102.8 | 102.9 | 102.5 | 102 | 101.8
4.2 | 102.88 | 102.9 | 102.6 | 102.0714 | 101.8
4.3 | 102.95 | 102.9 | 102.6 | 102.1 | 101.8
4.4 | 103 | 102.95 | 102.7 | 102.1 | 101.8
4.5 | 103 | 103 | 102.8 | 102.1833 | 101.8
4.6 | 103.1 | 103.1 | 102.8 | 102.2 | 101.8
4.7 | 103.1 | 103.1167 | 102.8 | 102.2 | 101.8
4.8 | 103.2 | 103.2 | 102.9 | 102.2143 | 101.8
4.9 | 103.2 | 103.2 | 102.9 | 102.3 | 101.8
5 | 103.2 | 103.2375 | 102.9 | 102.3 | 101.9
5.1 | 103.2667 | 103.3 | 102.9 | 102.3 | 101.9
5.2 | 103.3 | 103.3 | 102.9 | 102.3 | 101.9
5.3 | 103.3 | 103.4 | 103 | 102.3 | 101.93
5.4 | 103.3 | 103.4 | 103 | 102.3 | 102
5.5 | 103.3 | 103.4 | 103.1 | 102.4 | 102
5.6 | 103.3 | 103.4 | 103.1 | 102.4 | 102
5.7 | 103.3167 | 103.4 | 103.1 | 102.4 | 101.9333
5.8 | 103.4 | 103.4 | 103.1 | 102.4 | 101.9
5.9 | 103.3 | 103.4 | 103.1 | 102.4 | 101.9
6 | 103.3 | 103.4 | 103.1 | 102.4 | 101.9
6.1 | 103.3 | 103.4 | 103.1 | 102.5 | 101.9
6.2 | 103.3 | 103.4 | 103.1 | 102.5 | 101.9
6.3 | 103.3 | 103.4 | 103.1 | 102.5 | 101.9222
6.4 | 103.3 | 103.4 | 103.1 | 102.4 | 102
6.5 | 103.3 | 103.4 | 103.1 | 102.45 | 102
6.6 | 103.3 | 103.4 | 103.1 | 102.5 | 102
6.7 | 103.3 | 103.3833 | 103.1 | 102.4 | 102
6.8 | 103.3 | 103.3 | 103.1 | 102.4 | 102
6.9 | 103.3 | 103.3 | 103.1 | 102.4 | 102
7 | 103.3 | 103.3 | 103.1 | 102.4 | 102
7.1 | 103.3 | 103.3 | 103.1 | 102.3143 | 101.9
7.2 | 103.3 | 103.3 | 103.1 | 102.3 | 101.9
7.3 | 103.2375 | 103.3 | 103.025 | 102.3 | 101.9
7.4 | 103.2 | 103.3 | 103 | 102.3 | 101.9
7.5 | 103.2 | 103.3 | 102.9333 | 102.3 | 101.9
7.6 | 103.2 | 103.3 | 102.9 | 102.3 | 101.9
7.7 | 103.2 | 103.3 | 102.9 | 102.3 | 101.8875
7.8 | 103.2 | 103.3 | 102.9 | 102.2625 | 101.8
7.9 | 103.2 | 103.3 | 102.9 | 102.2 | 101.7571
8 | 103.2 | 103.3 | 102.9 | 102.2 | 101.7
8.1 | 103.2 | 103.2 | 102.9 | 102.1167 | 101.6111
8.2 | 103.2 | 103.2 | 102.8 | 102.1 | 101.6
8.3 | 103.1 | 103.3 | 102.8 | 102.1 | 101.6
8.4 | 103.1 | 103.2 | 102.8571 | 102.1 | 101.6
8.5 | 103.1 | 103.2 | 102.8833 | 102 | 101.5
8.6 | 103.1 | 103.3 | 102.8 | 102 | 101.5
8.7 | 103.1 | 103.2 | 102.8 | 102 | 101.5
8.8 | 103.0714 | 103.2 | 102.8 | 102 | 101.45
8.9 | 103 | 103.2 | 102.7333 | 102 | 101.4
9 | 103 | 103.2 | 102.7 | 102 | 101.4
9.1 | 103 | 103.1 | 102.7 | 101.975 | 101.4
9.2 | 102.9 | 103.1 | 102.7 | 101.9 | 101.4
9.3 | 102.9 | 103.1 | 102.7 | 101.8875 | 101.4
9.4 | 102.9 | 103.08 | 102.7 | 101.8 | 101.3
9.5 | 102.8 | 103 | 102.65 | 101.8 | 101.25
9.6 | 102.8 | 103 | 102.6 | 101.8 | 101.2
9.7 | 102.7 | 102.9 | 102.5571 | 101.7333 | 101.14
9.8 | 102.7 | 102.9 | 102.5 | 101.8 | 101.1
9.9 | 102.7 | 102.8167 | 102.5 | 101.7 | 101.1
10 | 102.7 | 102.8 | 102.4 | 101.7 | 101.0333
10.1 | 102.7286 | 102.8 | 102.4 | 101.6167 | 101
10.2 | 102.8 | 102.8833 | 102.3714 | 101.6 | 101
10.3 | 102.8 | 102.9 | 102.3 | 101.5 | 101
10.4 | 102.9 | 103 | 102.3 | 101.5 | 100.95
10.5 | 102.9 | 103 | 102.3 | 101.4 | 100.9
10.6 | 102.9667 | 103.05 | 102.3 | 101.4 | 100.9
10.7 | 103 | 103.1 | 102.3 | 101.3 | 100.9
10.8 | 103 | 103.1 | 102.3 | 101.2857 | 100.8
10.9 | 103 | 103.1 | 102.3 | 101.2 | 100.8
11 | 103.0667 | 103.1833 | 102.4 | 101.2 | 100.7625
11.1 | 103.1 | 103.2 | 102.4 | 101.1 | 100.7
11.2 | 103.1 | 103.2 | 102.4 | 101.1 | 100.6444
11.3 | 103 | 103.2 | 102.4143 | 101.1 | 100.6
11.4 | 103 | 103.2 | 102.5 | 101.1 | 100.5125
11.5 | 103 | 103.2 | 102.5 | 101.0167 | 100.5
11.6 | 102.9 | 103.2 | 102.5 | 101 | 100.4778
11.7 | 102.9 | 103.1 | 102.5 | 101 | 100.3667
11.8 | 102.9 | 103.1 | 102.5 | 101 | 100.3
11.9 | 102.875 | 103.0714 | 102.4625 | 101 | 100.24
12 | 102.8 | 103 | 102.4 | 101 | 100.2
12.1 | 102.8 | 103 | 102.4 | 101 | 100.1222
12.2 | 102.8 | 103 | 102.3167 | 101 | 100.1
12.3 | 102.8 | 103 | 102.3 | 101 | 100
12.4 | 102.8 | 103 | 102.2833 | 101 | 99.9091
12.5 | 102.8 | 103 | 102.2 | 101 | 99.9
12.6 | 102.7 | 102.9333 | 102.12 | 101 | 99.9
12.7 | 102.6667 | 102.9 | 102.1 | 100.9 | 99.9
12.8 | 102.6 | 102.9 | 102.1 | 100.9 | 99.9
12.9 | 102.6 | 102.84 | 102 | 100.8857 | 99.9
13 | 102.5286 | 102.7833 | 102 | 100.8 | 99.9
13.1 | 102.5 | 102.7 | 102 | 100.725 | 99.8
13.2 | 102.4143 | 102.7 | 102 | 100.7 | 99.8
13.3 | 102.4 | 102.6 | 101.9714 | 100.6 | 99.8
13.4 | 102.3714 | 102.6 | 101.9 | 100.6 | 99.8
13.5 | 102.3 | 102.5 | 101.8833 | 100.5 | 99.8
13.6 | 102.22 | 102.4 | 101.8 | 100.4167 | 99.725
13.7 | 102.2 | 102.4 | 101.7 | 100.4 | 99.6875
13.8 | 102.125 | 102.4 | 101.7 | 100.3 | 99.6
13.9 | 102.1 | 102.4 | 101.6 | 100.2625 | 99.56
14 | 102.1 | 102.3 | 101.6 | 100.1375 | 99.4636
14.1 | 102 | 102.3 | 101.4833 | 100.1 | 99.4
14.2 | 102 | 102.2875 | 101.4 | 100 | 99.3444
14.3 | 102 | 102.2 | 101.3429 | 99.9714 | 99.3
14.4 | 102 | 102.1143 | 101.3 | 99.9 | 99.2
14.5 | 102 | 102.1 | 101.2 | 99.7833 | 99.1778
14.6 | 102 | 102.1 | 101.2 | 99.7 | 99.0571
14.7 | 102 | 102.1 | 101.16 | 99.7 | 99
14.8 | 102 | 102.1 | 101.1 | 99.6 | 98.9
14.9 | 101.9 | 102.1 | 101.0143 | 99.6 | 98.8714
15 | 101.9 | 102.1 | 100.975 | 99.5 | 98.7375
15.1 | 101.9 | 102.1 | 100.9 | 99.3714 | 98.7
15.2 | 101.8571 | 102 | 100.8 | 99.2 | 98.6111
15.3 | 101.8 | 102 | 100.8 | 99.1714 | 98.5
15.4 | 101.8 | 102 | 100.8 | 99.0375 | 98.4
15.5 | 101.8 | 101.9167 | 100.7 | 98.9857 | 98.4
15.6 | 101.7 | 101.9 | 100.675 | 98.9 | 98.3
15.7 | 101.6429 | 101.9 | 100.6 | 98.9 | 98.1889
15.8 | 101.6 | 101.8375 | 100.6 | 98.8 | 98.075
15.9 | 101.5 | 101.8 | 100.5 | 98.6667 | 98
16 | 101.42 | 101.72 | 100.5 | 98.6 | 97.9143
16.1 | 101.4 | 101.7 | 100.4429 | 98.5 | 97.81
16.2 | 101.28 | 101.625 | 100.4 | 98.375 | 97.71
16.3 | 101.2 | 101.5857 | 100.3 | 98.25 | 97.5
16.4 | 101.1 | 101.5 | 100.2 | 98.1143 | 97.47
16.5 | 101 | 101.4 | 100.2 | 98.0333 | 97.37
16.6 | 101 | 101.2667 | 100.1667 | 98 | 97.2667
16.7 | 100.8667 | 101.2 | 100.1 | 98 | 97.16
16.8 | 100.8 | 101.1 | 100.0333 | 97.9 | 97.0429
16.9 | 100.7 | 101.0143 | 99.9 | 97.8125 | 96.9125
17 | 100.6125 | 100.9 | 99.8286 | 97.7833 | 96.7889
17.1 | 100.5 | 100.8 | 99.6714 | 97.7 | 96.6778
17.2 | 100.4833 | 100.6833 | 99.6 | 97.6 | 96.5667
17.3 | 100.3167 | 100.5167 | 99.5 | 97.5 | 96.4429
17.4 | 100.3 | 100.4667 | 99.4286 | 97.3875 | 96.3125
17.5 | 100.2167 | 100.4 | 99.2714 | 97.2625 | 96.2429
17.6 | 100.2 | 100.3 | 99.1 | 97.1375 | 96.0375
17.7 | 100.1 | 100.2429 | 99 | 97.1 | 95.9125
17.8 | 100.1 | 100.18 | 98.9333 | 96.975 | 95.81
17.9 | 100.06 | 100.1 | 98.7714 | 96.8333 | 95.7182
18 | 100 | 100.1 | 98.6167 | 96.7857 | 95.7
18.1 | 100 | 100 | 98.5333 | 96.7 | 95.6
18.2 | 100 | 100 | 98.475 | 96.6143 | 95.5556
18.3 | 99.9571 | 100 | 98.2625 | 96.4875 | 95.4286
18.4 | 99.9 | 99.9333 | 98.1167 | 96.3625 | 95.2857
18.5 | 99.85 | 99.9 | 97.94 | 96.1889 | 95.1429
18.6 | 99.8 | 99.7833 | 97.8714 | 96.025 | 95.1
18.7 | 99.6333 | 99.7 | 97.8 | 95.8 | 94.9778
18.8 | 99.4667 | 99.6333 | 97.7 | 95.7 | 94.8571
18.9 | 99.2 | 99.5667 | 97.625 | 95.6 | 94.8
19 | 99 | 99.4 | 97.5 | 95.45 | 94.67
19.1 | 98.7333 | 99.2333 | 97.4 | 95.25 | 94.57
19.2 | 98.46 | 98.92 | 97.3 | 95.1 | 94.47
19.3 | 98.0429 | 98.7 | 97.2 | 94.9333 | 94.3625
19.4 | 97.6 | 98.32 | 97.15 | 94.7857 | 94.2091
19.5 | 97.0333 | 97.9 | 97.1 | 94.6333 | 94.02
19.6 | 96.3 | 97.4857 | 96.95 | 94.4833 | 93.9
19.7 | 95.55 | 97.1429 | 96.7714 | 94.3286 | 93.8167
19.8 | 94.4 | 96.5714 | 96.6 | 94.0833 | 93.7
19.9 | 92.9333 | 96 | 96.4 | 93.9286 | 93.5
20 | 90.9714 | 95.1 | 96.1 | 93.7714 | 93.2889
20.1 | 88.2857 | 94.2 | 95.8667 | 93.6167 | 93.175
20.2 | 84 | 93 | 95.4 | 93.4571 | 93.0111
20.3 | 77.7857 | 91.35 | 95.25 | 93.3 | 92.8444
20.4 | 70.66 | 89.4 | 94.8714 | 93.1167 | 92.7143
20.5 | 61.7286 | 86.2 | 94.5 | 92.9571 | 92.5667
20.6 | 52.95 | 81.8286 | 93.9857 | 92.8 | 92.3375
20.7 | 44.8 | 75.6143 | 93.45 | 92.7333 | 92.16
20.8 | 36.74 | 67.7429 | 92.8286 | 92.56 | 91.96
20.9 | 30.6667 | 59.3 | 92.0857 | 92.4125 | 91.8125
21 | 25.5429 | 50.8 | 91.3 | 92.2714 | 91.675
21.1 | 22.14 | 43.16 | 90.26 | 92.0429 | 91.4857
21.2 | 19.7 | 36.26 | 88.9333 | 91.9 | 91.3222
21.3 | 17.8667 | 30.1 | 87.1833 | 91.7333 | 91.2
21.4 | 16.7 | 25.3333 | 85.1857 | 91.5571 | 90.9778
21.5 | 15.6 | 22.0143 | 82.1857 | 91.3 | 90.7429
21.6 | 14.8 | 19.6 | 78 | 91.12 | 90.5444
21.7 | 14.0714 | 17.9333 | 72 | 90.86 | 90.3667
21.8 | 13.5857 | 16.5 | 65.1714 | 90.4833 | 90.16
21.9 | 13.0857 | 15.4167 | 57.7333 | 90.2571 | 89.96
22 | 12.7333 | 14.5667 | 49.4 | 89.975 | 89.76
22.1 | 12.4286 | 13.8 | 42.6 | 89.7 | 89.63
22.2 | 12.14 | 13.3 | 36.2 | 89.35 | 89.46
22.3 | 11.9714 | 12.8 | 30.8286 | 88.9 | 89.3364
22.4 | 11.6857 | 12.4286 | 26.6857 | 88.4143 | 89.1
22.5 | 11.4833 | 12.0857 | 23.8286 | 87.9571 | 88.9222
22.6 | 11.3 | 11.7833 | 21.5 | 87.35 | 88.74
22.7 | 11.1167 | 11.5333 | 19.8333 | 86.725 | 88.62
22.8 | 10.95 | 11.2857 | 18.4857 | 86.1 | 88.44
22.9 | 10.8 | 11.1 | 17.4667 | 85.3 | 88.3
23 | 10.6333 | 10.9 | 16.5333 | 84.5 | 88.1
23.1 | 10.4667 | 10.6571 | 15.7714 | 83.4571 | 87.8714
23.2 | 10.3143 | 10.4833 | 15.2 | 82.025 | 87.675
23.3 | 10.14 | 10.3167 | 14.7143 | 80.3 | 87.425
23.4 | 10.0667 | 10.15 | 14.2571 | 77.9571 | 87.0818
23.5 | 9.925 | 9.9833 | 13.85 | 74.6 | 86.9
23.6 | 9.8 | 9.8286 | 13.55 | 70.5143 | 86.5714
23.7 | 9.6333 | 9.7571 | 13.2333 | 65.6571 | 86.3
23.8 | 9.54 | 9.6143 | 13 | 60 | 86
23.9 | 9.4 | 9.4571 | 12.725 | 53.8 | 85.725
24 | 9.2857 | 9.3143 | 12.4 | 47.0857 | 85.37
24.1 | 9.1333 | 9.15 | 12.2143 | 41.2375 | 85.0667
24.2 | 9.1 | 9 | 12 | 35.9125 | 84.7
24.3 | 8.9833 | 8.95 | 11.7857 | 31.6714 | 84.2857
24.4 | 8.9 | 8.7875 | 11.6222 | 28.26 | 83.7625
24.5 | 8.75 | 8.7 | 11.4714 | 25.6 | 83.25
24.6 | 8.6 | 8.6 | 11.2571 | 23.6667 | 82.8
24.7 | 8.5571 | 8.5 | 11.1 | 22.2 | 82.1444
24.8 | 8.5 | 8.4 | 10.9333 | 21 | 81.4
24.9 | 8.3571 | 8.3 | 10.7 | 20 | 80.65
25 | 8.225 | 8.2167 | 10.575 | 19.25 | 79.7
25.1 | 8.1714 | 8.1 | 10.45 | 18.475 | 78.5727
25.2 | 8.1 | 8.0286 | 10.2833 | 17.9375 | 77.3
25.3 | 8 | 7.9 | 10.1167 | 17.3714 | 75.3143
25.4 | 7.9 | 7.8286 | 10.05 | 16.9429 | 73.2
25.5 | 7.82 | 7.7667 | 9.8 | 16.4625 | 70.4
25.6 | 7.7 | 7.6 | 9.72 | 16.125 | 66.3714
25.7 | 7.6833 | 7.6 | 9.6 | 15.8333 | 61.8
25.8 | 7.6 | 7.4714 | 9.5143 | 15.5 | 56.7
25.9 | 7.5 | 7.4 | 9.4 | 15.3 | 50.725
26 | 7.4 | 7.3 | 9.2714 | 15 | 45.1571
26.1 | 7.3429 | 7.2 | 9.1286 | 14.725 | 40.1857
26.2 | 7.3 | 7.15 | 9 | 14.55 | 36.0286
26.3 | 7.1857 | 7.1 | 8.9 | 14.2667 | 32.2875
26.4 | 7.1 | 7 | 8.8 | 14.0667 | 29.1875
26.5 | 7.05 | 6.9 | 8.7 | 13.88 | 27.1
26.6 | 6.9833 | 6.9 | 8.575 | 13.7 | 25.3
26.7 | 6.9 | 6.8 | 8.5 | 13.5167 | 23.8667
26.8 | 6.8429 | 6.7 | 8.4 | 13.2857 | 22.8
26.9 | 6.8 | 6.6429 | 8.3 | 13.1125 | 21.88
27 | 6.72 | 6.6 | 8.2125 | 12.95 | 21.06
27.1 | 6.675 | 6.55 | 8.1 | 12.7429 | 20.34
27.2 | 6.6 | 6.4857 | 8.0167 | 12.5 | 19.8
27.3 | 6.55 | 6.4 | 7.9571 | 12.35 | 19.3
27.4 | 6.5 | 6.4 | 7.8 | 12.2 | 18.8455
27.5 | 6.4375 | 6.3 | 7.7143 | 12.05 | 18.44
27.6 | 6.4 | 6.25 | 7.7 | 11.925 | 18.04
27.7 | 6.3 | 6.2 | 7.6 | 11.7 | 17.8
27.8 |  | 6.2 | 7.5 | 11.5857 | 17.47
27.9 |  | 6.1 | 7.4 | 11.3875 | 17.18
28 |  | 6.0625 | 7.3429 | 11.2625 | 16.9727
28.1 |  | 6 | 7.2833 | 11.1 | 16.7
28.2 |  | 6 | 7.2 | 11 | 16.5
28.3 |  | 5.9 | 7.1571 | 10.9 | 16.2
28.4 |  |  | 7.1 | 10.8 | 16.1
28.5 |  |  | 7 | 10.6714 | 15.8
28.6 |  |  | 6.9 | 10.5167 | 15.6
28.7 |  |  | 6.8667 | 10.36 | 15.3889
28.8 |  |  | 6.7 | 10.2 | 15.2636
28.9 |  |  | 6.7 | 10.1625 | 15.07
29 |  |  | 6.65 | 10.0375 | 14.78
29.1 |  |  | 6.6 | 9.9125 | 14.58
29.2 |  |  | 6.5 | 9.7875 | 14.39
29.3 |  |  | 6.5 | 9.65 | 14.28
29.4 |  |  | 6.4 | 9.5125 | 14.08
29.5 |  |  | 6.4 | 9.4857 | 13.8909
29.6 |  |  | 6.3286 | 9.3429 | 13.8
29.7 |  |  |  | 9.2125 | 13.6
29.8 |  |  |  | 9.1 | 13.4
29.9 |  |  |  | 9.0625 | 13.2
30 |  |  |  | 8.9286 | 13.1
30.1 |  |  |  | 8.8 | 12.9
30.2 |  |  |  | 8.8 | 12.8091
30.3 |  |  |  | 8.675 | 12.62
30.4 |  |  |  | 8.55 | 12.51
30.5 |  |  |  | 8.5 | 12.41
30.6 |  |  |  | 8.3429 | 12.31
30.7 |  |  |  | 8.3 | 12.12
30.8 |  |  |  | 8.1875 | 12.01
30.9 |  |  |  | 8.1 | 11.8
31 |  |  |  | 8.0167 | 11.6778
31.1 |  |  |  | 7.9 | 11.5625
31.2 |  |  |  | 7.9 | 11.4167
31.3 |  |  |  | 7.8 | 11.2714
31.4 |  |  |  | 7.7286 | 11.2
31.5 |  |  |  | 7.6 | 11.0778
31.6 |  |  |  | 7.5286 | 10.94
31.7 |  |  |  | 7.5 | 10.8
31.8 |  |  |  | 7.4 | 10.74
31.9 |  |  |  | 7.3375 | 10.57
32 |  |  |  | 7.3 | 10.47
32.1 |  |  |  | 7.1875 | 10.37
32.2 |  |  |  | 7.1 | 10.2667
32.3 |  |  |  | 7.1 | 10.1556
32.4 |  |  |  |  | 10.0444
32.5 |  |  |  |  | 9.94
32.6 |  |  |  |  | 9.84
32.7 |  |  |  |  | 9.74
32.8 |  |  |  |  | 9.64
32.9 |  |  |  |  | 9.5455
33 |  |  |  |  | 9.45
33.1 |  |  |  |  | 9.35
33.2 |  |  |  |  | 9.25
33.3 |  |  |  |  | 9.1545
33.4 |  |  |  |  | 9.1
33.5 |  |  |  |  | 9.06
33.6 |  |  |  |  | 8.96
33.7 |  |  |  |  | 8.86
33.8 |  |  |  |  | 8.8
33.9 |  |  |  |  | 8.7375
34 |  |  |  |  | 8.6125
34.1 |  |  |  |  | 8.5
34.2 |  |  |  |  | 8.4444
34.3 |  |  |  |  | 8.3333
34.4 |  |  |  |  | 8.2889
34.5 |  |  |  |  | 8.1714
34.6 |  |  |  |  | 8.1
34.7 |  |  |  |  | 8.0125
34.8 |  |  |  |  | 8
34.9 |  |  |  |  | 7.91
35 |  |  |  |  | 7.9

## Wedge 15 Output Factors
Output factors
Wedge 15 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.871 | 0.888 | 0.897 | 0.905 | 0.911 | 0.916 | 0.919 | 0.922
| 4 | 0.89 | 0.91 | 0.92 | 0.931 | 0.939 | 0.945 | 0.949 | 0.953
| 5 | 0.9 | 0.922 | 0.934 | 0.947 | 0.956 | 0.963 | 0.968 | 0.974
| 7 | 0.912 | 0.937 | 0.951 | 0.967 | 0.978 | 0.988 | 0.995 | 1.002
| 10 | 0.922 | 0.949 | 0.965 | 0.985 | 1 | 1.013 | 1.023 | 1.033
| 15 | 0.93 | 0.96 | 0.978 | 1.001 | 1.02 | 1.038 | 1.051 | 1.065
| 20 | 0.935 | 0.966 | 0.985 | 1.01 | 1.031 | 1.052 | 1.068 | 1.086
| 30 | 0.941 | 0.973 | 0.993 | 1.021 | 1.045 | 1.071 | 1.091 | 1.114
| 40 | 0.944 | 0.977 | 0.998 | 1.027 | 1.054 | 1.082 | 1.104 | 1.129

## Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 23.4 | 30.8 | 44 | 55.9 | 61.6
0.1 | 26.86 | 34.32 | 47.66 | 59.8 | 65.78
0.2 | 33.06 | 40.24 | 53.2 | 64.88 | 71.26
0.3 | 41.82 | 48.62 | 60.7 | 71.32 | 77.16
0.4 | 49.3 | 55.48 | 66.94 | 76.86 | 81.78
0.5 | 56.24 | 62.04 | 72.56 | 81.3 | 85.74
0.6 | 62.04 | 67.42 | 77.1 | 85.18 | 88.84
0.7 | 67.18 | 72.06 | 80.8 | 88.18 | 91.36
0.8 | 71.54 | 76.06 | 84.08 | 90.62 | 93.3
0.9 | 75.42 | 79.4 | 86.84 | 92.68 | 94.98
1 | 78.9 | 82.46 | 89.32 | 94.28 | 96.26
1.1 | 81.94 | 85.12 | 91.24 | 95.72 | 97.34
1.2 | 84.44 | 87.36 | 92.92 | 96.86 | 98.08
1.3 | 86.76 | 89.56 | 94.38 | 97.68 | 98.72
1.4 | 88.7 | 91.28 | 95.5 | 98.38 | 99.24
1.5 | 90.64 | 92.62 | 96.5 | 99.02 | 99.56
1.6 | 92.16 | 94.12 | 97.34 | 99.42 | 99.86
1.7 | 93.58 | 95.26 | 97.98 | 99.66 | 99.96
1.8 | 94.56 | 96.06 | 98.48 | 99.9 | 100
1.9 | 95.74 | 96.92 | 98.92 | 100 | 100
2 | 96.72 | 97.62 | 99.38 | 100 | 99.96
2.1 | 97.36 | 98.12 | 99.64 | 100 | 99.86
2.2 | 97.92 | 98.62 | 99.74 | 99.96 | 99.72
2.3 | 98.56 | 99.06 | 99.9 | 99.86 | 99.52
2.4 | 99.08 | 99.44 | 100 | 99.76 | 99.36
2.5 | 99.44 | 99.64 | 100 | 99.56 | 99.12
2.6 | 99.68 | 99.8 | 100 | 99.36 | 98.82
2.7 | 99.84 | 99.94 | 99.9 | 99.06 | 98.52
2.8 | 99.9 | 100 | 99.86 | 98.86 | 98.26
2.9 | 99.9 | 100 | 99.66 | 98.62 | 98.02
3 | 99.94 | 100 | 99.46 | 98.36 | 97.66
3.1 | 100 | 99.9 | 99.3 | 98.06 | 97.42
3.2 | 99.96 | 99.8 | 99.1 | 97.76 | 97.02
3.3 | 99.9 | 99.7 | 98.86 | 97.52 | 96.76
3.4 | 99.76 | 99.6 | 98.52 | 97.16 | 96.42
3.5 | 99.56 | 99.46 | 98.26 | 96.8 | 96.06
3.6 | 99.36 | 99.16 | 98.02 | 96.56 | 95.76
3.7 | 99.12 | 98.92 | 97.76 | 96.18 | 95.42
3.8 | 98.86 | 98.66 | 97.46 | 95.82 | 95.08
3.9 | 98.62 | 98.46 | 97.06 | 95.56 | 94.76
4 | 98.32 | 98.16 | 96.76 | 95.22 | 94.38
4.1 | 98.02 | 97.82 | 96.52 | 94.82 | 94.02
4.2 | 97.62 | 97.56 | 96.12 | 94.56 | 93.66
4.3 | 97.22 | 97.32 | 95.82 | 94.16 | 93.36
4.4 | 96.96 | 97.02 | 95.46 | 93.82 | 93.02
4.5 | 96.66 | 96.52 | 95.12 | 93.52 | 92.66
4.6 | 96.26 | 96.22 | 94.86 | 93.16 | 92.32
4.7 | 95.92 | 95.82 | 94.38 | 92.86 | 92.06
4.8 | 95.52 | 95.52 | 94.06 | 92.52 | 91.68
4.9 | 94.98 | 95.16 | 93.72 | 92.12 | 91.32
5 | 94.56 | 94.78 | 93.32 | 91.78 | 90.96
5.1 | 94.28 | 94.42 | 93.02 | 91.46 | 90.66
5.2 | 93.8 | 94.0333 | 92.7667 | 91.1667 | 90.3333
5.3 | 93.46 | 93.66 | 92.36 | 90.78 | 90.06
5.4 | 93.16 | 93.28 | 92.06 | 90.46 | 89.68
5.5 | 92.66 | 92.98 | 91.76 | 90.16 | 89.34
5.6 | 92.26 | 92.66 | 91.38 | 89.68 | 88.88
5.7 | 91.86 | 92.26 | 91.06 | 89.46 | 88.68
5.8 | 91.46 | 91.96 | 90.68 | 89.08 | 88.38
5.9 | 90.96 | 91.56 | 90.36 | 88.76 | 87.98
6 | 90.56 | 91.28 | 90.06 | 88.46 | 87.76
6.1 | 90.26 | 90.86 | 89.76 | 88.06 | 87.36
6.2 | 89.94 | 90.46 | 89.36 | 87.76 | 86.98
6.3 | 89.38 | 90.16 | 88.96 | 87.48 | 86.78
6.4 | 89.18 | 89.76 | 88.58 | 87.18 | 86.48
6.5 | 88.76 | 89.46 | 88.36 | 86.86 | 86.16
6.6 | 88.36 | 88.98 | 87.98 | 86.58 | 85.78
6.7 | 88.06 | 88.68 | 87.66 | 86.26 | 85.48
6.8 | 87.54 | 88.36 | 87.28 | 85.96 | 85.16
6.9 | 87.06 | 87.96 | 87.06 | 85.64 | 84.78
7 | 86.68 | 87.66 | 86.66 | 85.28 | 84.56
7.1 | 86.36 | 87.36 | 86.36 | 84.98 | 84.26
7.2 | 85.96 | 87.06 | 85.98 | 84.66 | 83.98
7.3 | 85.48 | 86.72 | 85.76 | 84.38 | 83.74
7.4 | 85.14 | 86.14 | 85.46 | 84.16 | 83.38
7.5 | 84.78 | 85.76 | 85.14 | 83.78 | 83.06
7.6 | 84.36 | 85.46 | 84.78 | 83.48 | 82.78
7.7 | 84.06 | 85.08 | 84.48 | 83.18 | 82.48
7.8 | 83.68 | 84.88 | 84.18 | 82.86 | 82.16
7.9 | 83.38 | 84.58 | 83.88 | 82.56 | 81.86
8 | 83.04 | 84.08 | 83.56 | 82.18 | 81.58
8.1 | 82.56 | 83.86 | 83.16 | 81.98 | 81.18
8.2 | 82.16 | 83.48 | 82.88 | 81.66 | 80.88
8.3 | 81.76 | 83.16 | 82.56 | 81.38 | 80.66
8.4 | 81.46 | 82.78 | 82.28 | 81.06 | 80.46
8.5 | 80.98 | 82.48 | 82.06 | 80.68 | 80.06
8.6 | 80.56 | 82.16 | 81.68 | 80.38 | 79.76
8.7 | 80.3 | 81.9 | 81.4 | 80.1 | 79.5
8.8 | 80 | 81.5 | 81 | 79.8 | 79.1
8.9 | 79.7 | 81.2 | 80.7 | 79.5 | 79
9 | 79.2 | 80.8 | 80.4 | 79.2 | 78.5
9.1 | 78.8 | 80.5 | 80.1 | 78.9 | 78.3
9.2 | 78.4 | 80.1 | 79.8 | 78.7 | 78
9.3 | 78.1 | 79.7 | 79.5 | 78.3 | 77.7
9.4 | 77.8 | 79.4 | 79.1 | 78 | 77.4
9.5 | 77.3 | 79.2 | 78.7 | 77.7 | 77.1
9.6 | 77 | 78.9 | 78.5 | 77.5 | 76.7
9.7 | 76.6 | 78.5 | 78.2 | 77.1 | 76.6
9.8 | 76.2 | 78 | 77.9 | 76.8 | 76.2
9.9 | 75.9 | 77.7 | 77.6 | 76.5 | 76
10 | 75.7 | 77.5 | 77.2 | 76.2 | 75.6
10.1 | 75.3 | 77.1 | 77 | 76 | 75.3
10.2 | 74.8 | 76.8 | 76.6 | 75.7 | 75
10.3 | 74.6 | 76.5 | 76.4 | 75.4 | 74.7
10.4 | 74.2 | 76.2 | 76 | 75.1 | 74.5
10.5 | 73.9 | 75.9 | 75.8 | 74.8 | 74.2
10.6 | 73.6 | 75.7 | 75.3 | 74.5 | 74
10.7 | 73.2 | 75.2 | 75.1 | 74.2 | 73.6
10.8 | 72.8 | 74.8 | 74.9 | 74 | 73.4
10.9 | 72.4 | 74.5 | 74.6 | 73.7 | 73.1
11 | 72.1 | 74.3 | 74.3 | 73.4 | 72.8
11.1 | 71.9 | 73.9 | 73.9 | 73.2 | 72.5
11.2 | 71.5 | 73.6 | 73.7 | 72.8 | 72.2
11.3 | 71.2 | 73.3 | 73.4 | 72.5 | 72
11.4 | 71 | 73 | 73.2 | 72.3 | 71.7
11.5 | 70.7 | 72.8 | 72.9 | 72 | 71.4
11.6 | 70.3 | 72.4 | 72.6 | 71.7 | 71.1
11.7 | 69.8 | 72.1 | 72.3 | 71.5 | 70.9
11.8 | 69.4 | 71.8 | 72 | 71.2 | 70.7
11.9 | 69.2 | 71.5 | 71.7 | 70.9 | 70.4
12 | 69 | 71.2 | 71.4 | 70.6 | 70.1
12.1 | 68.6 | 70.9 | 71.2 | 70.4 | 69.9
12.2 | 68.2 | 70.6 | 70.9 | 70.2 | 69.5
12.3 | 68 | 70.3 | 70.7 | 69.9 | 69.3
12.4 | 67.7 | 70 | 70.4 | 69.6 | 69.1
12.5 | 67.3 | 69.8 | 70.1 | 69.4 | 68.7
12.6 | 67 | 69.5 | 69.8 | 69.1 | 68.6
12.7 | 66.7 | 69.1 | 69.5 | 68.8 | 68.3
12.8 | 66.4 | 68.9 | 69.3 | 68.5 | 68.1
12.9 | 66.1 | 68.6 | 69 | 68.3 | 67.9
13 | 65.9 | 68.4 | 68.8 | 68 | 67.6
13.1 | 65.6 | 68 | 68.5 | 67.8 | 67.3
13.2 | 65.2 | 67.7 | 68.2 | 67.5 | 67
13.3 | 65 | 67.5 | 67.9 | 67.3 | 66.8
13.4 | 64.7 | 67.3 | 67.7 | 67 | 66.5
13.5 | 64.4 | 66.9 | 67.5 | 66.8 | 66.3
13.6 | 64.1 | 66.6 | 67.2 | 66.5 | 66
13.7 | 63.8 | 66.3 | 66.8 | 66.2 | 65.7
13.8 | 63.5 | 66 | 66.6 | 66 | 65.6
13.9 | 63.3 | 65.8 | 66.3 | 65.7 | 65.2
14 | 62.9 | 65.5 | 66.1 | 65.6 | 65
14.1 | 62.7 | 65.2 | 65.9 | 65.3 | 64.8
14.2 | 62.3 | 65 | 65.6 | 65.1 | 64.6
14.3 | 62 | 64.7 | 65.3 | 64.8 | 64.3
14.4 | 61.7 | 64.4 | 65 | 64.5 | 64
14.5 | 61.4 | 64.1 | 64.8 | 64.2 | 63.8
14.6 | 61.1 | 63.8 | 64.5 | 64 | 63.5
14.7 | 60.8 | 63.6 | 64.3 | 63.8 | 63.3
14.8 | 60.6 | 63.2 | 64 | 63.5 | 63
14.9 | 60.4 | 63 | 63.7 | 63.2 | 62.7
15 | 60.1 | 62.8 | 63.5 | 63 | 62.5
15.1 | 59.9 | 62.5 | 63.2 | 62.7 | 62.3
15.2 | 59.5 | 62.2 | 63 | 62.5 | 62
15.3 | 59.3 | 61.9 | 62.7 | 62.2 | 61.7
15.4 | 59 | 61.7 | 62.6 | 62 | 61.6
15.5 | 58.7 | 61.4 | 62.2 | 61.8 | 61.3
15.6 | 58.4 | 61.1 | 61.9 | 61.6 | 61.2
15.7 | 58.1 | 60.9 | 61.7 | 61.4 | 60.9
15.8 | 57.9 | 60.7 | 61.6 | 61.1 | 60.6
15.9 | 57.6 | 60.5 | 61.3 | 60.8 | 60.5
16 | 57.4 | 60.1 | 61 | 60.7 | 60.3
16.1 | 57.1 | 59.9 | 60.8 | 60.4 | 59.9
16.2 | 56.8 | 59.7 | 60.5 | 60.2 | 59.8
16.3 | 56.5 | 59.4 | 60.3 | 60 | 59.5
16.4 | 56.3 | 59.1 | 60.1 | 59.7 | 59.3
16.5 | 56 | 58.9 | 59.9 | 59.4 | 59.1
16.6 | 55.8 | 58.6 | 59.6 | 59.2 | 58.9
16.7 | 55.5 | 58.4 | 59.4 | 59 | 58.6
16.8 | 55.2 | 58.2 | 59.1 | 58.8 | 58.4
16.9 | 54.9 | 58 | 58.9 | 58.6 | 58.2
17 | 54.7 | 57.7 | 58.7 | 58.4 | 57.9
17.1 | 54.5 | 57.5 | 58.5 | 58.2 | 57.8
17.2 | 54.3 | 57.1 | 58.2 | 57.9 | 57.4
17.3 | 54.1 | 56.9 | 58 | 57.8 | 57.3
17.4 | 53.8 | 56.6 | 57.8 | 57.5 | 57.1
17.5 | 53.5 | 56.4 | 57.5 | 57.2 | 56.9
17.6 | 53.3 | 56.2 | 57.3 | 57 | 56.7
17.7 | 53.1 | 56 | 57.1 | 56.8 | 56.4
17.8 | 52.9 | 55.7 | 56.8 | 56.6 | 56.2
17.9 | 52.6 | 55.5 | 56.7 | 56.4 | 56
18 | 52.4 | 55.2 | 56.4 | 56.2 | 55.8
18.1 | 52.1 | 55 | 56.2 | 56 | 55.5
18.2 | 51.9 | 54.8 | 55.9 | 55.8 | 55.4
18.3 | 51.7 | 54.6 | 55.7 | 55.5 | 55.1
18.4 | 51.4 | 54.4 | 55.6 | 55.3 | 54.9
18.5 | 51 | 54.1 | 55.3 | 55.1 | 54.8
18.6 | 50.9 | 53.9 | 55.1 | 55 | 54.6
18.7 | 50.7 | 53.7 | 54.9 | 54.8 | 54.4
18.8 | 50.5 | 53.4 | 54.6 | 54.5 | 54.1
18.9 | 50.2 | 53.2 | 54.4 | 54.3 | 53.9
19 | 50 | 53 | 54.3 | 54.1 | 53.7
19.1 | 49.8 | 52.7 | 54 | 53.9 | 53.5
19.2 | 49.6 | 52.5 | 53.8 | 53.7 | 53.3
19.3 | 49.4 | 52.3 | 53.5 | 53.5 | 53.1
19.4 | 49.2 | 52 | 53.3 | 53.3 | 52.8
19.5 | 48.9 | 51.8 | 53.2 | 53.1 | 52.6
19.6 | 48.6 | 51.6 | 52.9 | 52.8 | 52.5
19.7 | 48.4 | 51.4 | 52.7 | 52.6 | 52.3
19.8 | 48.1 | 51.2 | 52.5 | 52.4 | 52.1
19.9 | 47.9 | 51 | 52.3 | 52.2 | 51.9
20 | 47.7 | 50.7 | 52.2 | 52 | 51.6
20.1 | 47.5 | 50.5 | 51.9 | 51.8 | 51.4
20.2 | 47.3 | 50.3 | 51.6 | 51.6 | 51.2
20.3 | 47.1 | 50.1 | 51.4 | 51.4 | 51
20.4 | 46.8 | 49.8 | 51.3 | 51.2 | 50.8
20.5 | 46.7 | 49.6 | 51 | 51 | 50.6
20.6 | 46.6 | 49.4 | 50.9 | 50.8 | 50.4
20.7 | 46.3 | 49.2 | 50.6 | 50.6 | 50.3
20.8 | 46 | 49 | 50.4 | 50.4 | 50.1
20.9 | 45.8 | 48.8 | 50.2 | 50.2 | 49.9
21 | 45.6 | 48.6 | 50 | 50 | 49.7
21.1 | 45.4 | 48.4 | 49.8 | 49.8 | 49.5
21.2 | 45.2 | 48.2 | 49.6 | 49.6 | 49.3
21.3 | 45 | 48 | 49.4 | 49.5 | 49.1
21.4 | 44.8 | 47.8 | 49.3 | 49.3 | 48.9
21.5 | 44.5 | 47.6 | 49.1 | 49.1 | 48.7
21.6 | 44.3 | 47.4 | 48.9 | 48.9 | 48.5
21.7 | 44.2 | 47.2 | 48.6 | 48.7 | 48.4
21.8 | 44 | 47 | 48.4 | 48.6 | 48.2
21.9 | 43.8 | 46.8 | 48.3 | 48.4 | 48.1
22 | 43.7 | 46.6 | 48.1 | 48.2 | 47.8
22.1 | 43.5 | 46.4 | 47.9 | 48 | 47.7
22.2 | 43.3 | 46.3 | 47.7 | 47.8 | 47.4
22.3 | 43 | 46 | 47.5 | 47.6 | 47.3
22.4 | 42.9 | 45.8 | 47.4 | 47.4 | 47.1
22.5 | 42.6 | 45.6 | 47.2 | 47.3 | 47
22.6 | 42.4 | 45.5 | 47 | 47.1 | 46.7
22.7 | 42.2 | 45.3 | 46.8 | 46.9 | 46.6
22.8 | 42 | 45.1 | 46.6 | 46.7 | 46.4
22.9 | 41.9 | 44.8 | 46.4 | 46.5 | 46.3
23 | 41.7 | 44.7 | 46.2 | 46.4 | 46
23.1 | 41.5 | 44.5 | 46 | 46.2 | 45.9
23.2 | 41.3 | 44.3 | 45.9 | 46 | 45.7
23.3 | 41.1 | 44.1 | 45.7 | 45.9 | 45.5
23.4 | 40.9 | 43.9 | 45.6 | 45.6 | 45.3
23.5 | 40.7 | 43.8 | 45.4 | 45.5 | 45.2
23.6 | 40.6 | 43.6 | 45.2 | 45.3 | 45
23.7 | 40.4 | 43.4 | 45 | 45.1 | 44.8
23.8 | 40.2 | 43.3 | 44.8 | 44.9 | 44.6
23.9 | 40 | 43.1 | 44.6 | 44.7 | 44.5
24 | 39.8 | 42.9 | 44.5 | 44.6 | 44.4
24.1 | 39.6 | 42.7 | 44.3 | 44.4 | 44.2
24.2 | 39.5 | 42.5 | 44.1 | 44.2 | 44
24.3 | 39.3 | 42.3 | 43.9 | 44.1 | 43.8
24.4 | 39.2 | 42.2 | 43.7 | 43.9 | 43.7
24.5 | 39 | 42 | 43.6 | 43.7 | 43.5
24.6 | 38.8 | 41.8 | 43.4 | 43.6 | 43.3
24.7 | 38.6 | 41.6 | 43.2 | 43.4 | 43.1
24.8 | 38.5 | 41.5 | 43 | 43.3 | 43
24.9 | 38.3 | 41.3 | 42.9 | 43.1 | 42.8
25 | 38.1 | 41 | 42.7 | 42.9 | 42.6
25.1 | 37.9 | 40.9 | 42.5 | 42.8 | 42.5
25.2 | 37.8 | 40.6 | 42.3 | 42.6 | 42.3
25.3 | 37.6 | 40.5 | 42.2 | 42.4 | 42.1
25.4 | 37.5 | 40.3 | 42 | 42.2 | 41.9
25.5 | 37.3 | 40.2 | 41.8 | 42 | 41.8
25.6 | 37.1 | 40.1 | 41.7 | 41.9 | 41.6
25.7 | 37 | 39.9 | 41.5 | 41.7 | 41.5
25.8 | 36.8 | 39.8 | 41.3 | 41.6 | 41.3
25.9 | 36.6 | 39.6 | 41.2 | 41.4 | 41.2
26 | 36.5 | 39.3 | 41 | 41.3 | 41
26.1 | 36.3 | 39.2 | 40.8 | 41.1 | 40.9
26.2 | 36.1 | 39 | 40.7 | 41 | 40.7
26.3 | 36 | 38.8 | 40.6 | 40.8 | 40.5
26.4 | 35.8 | 38.7 | 40.4 | 40.6 | 40.4
26.5 | 35.6 | 38.6 | 40.2 | 40.5 | 40.2
26.6 | 35.5 | 38.4 | 40.1 | 40.3 | 40.1
26.7 | 35.4 | 38.2 | 39.9 | 40.2 | 40
26.8 | 35.2 | 38.1 | 39.8 | 40.1 | 39.8
26.9 | 35 | 37.9 | 39.6 | 39.9 | 39.6
27 | 34.8 | 37.7 | 39.4 | 39.7 | 39.5
27.1 | 34.7 | 37.6 | 39.3 | 39.6 | 39.3
27.2 | 34.5 | 37.5 | 39.1 | 39.4 | 39.2
27.3 | 34.4 | 37.3 | 39 | 39.3 | 39.1
27.4 | 34.3 | 37.1 | 38.8 | 39.1 | 38.9
27.5 | 34.1 | 36.9 | 38.7 | 39 | 38.8
27.6 | 34 | 36.8 | 38.5 | 38.8 | 38.6
27.7 | 33.9 | 36.6 | 38.3 | 38.7 | 38.4
27.8 | 33.7 | 36.5 | 38.2 | 38.5 | 38.3
27.9 | 33.5 | 36.4 | 38.1 | 38.4 | 38.1
28 | 33.4 | 36.2 | 37.9 | 38.2 | 37.9
28.1 | 33.2 | 36.1 | 37.8 | 38 | 37.8
28.2 | 33.1 | 35.9 | 37.6 | 37.9 | 37.7
28.3 | 33 | 35.8 | 37.5 | 37.8 | 37.5
28.4 | 32.8 | 35.6 | 37.2 | 37.7 | 37.4
28.5 | 32.7 | 35.4 | 37.1 | 37.5 | 37.3
28.6 | 32.5 | 35.3 | 37 | 37.4 | 37.1
28.7 | 32.4 | 35.2 | 36.9 | 37.2 | 37
28.8 | 32.2 | 35 | 36.8 | 37.1 | 36.9
28.9 | 32.1 | 34.9 | 36.6 | 37 | 36.7
29 | 32 | 34.7 | 36.4 | 36.8 | 36.6
29.1 | 31.8 | 34.5 | 36.3 | 36.7 | 36.4
29.2 | 31.7 | 34.4 | 36.2 | 36.5 | 36.3
29.3 | 31.6 | 34.3 | 36 | 36.4 | 36.2
29.4 | 31.4 | 34.2 | 35.9 | 36.3 | 36
29.5 | 31.3 | 34 | 35.8 | 36 | 35.9
29.6 | 31.2 | 33.8 | 35.6 | 35.9 | 35.8
29.7 | 31 | 33.7 | 35.4 | 35.8 | 35.6
29.8 | 30.8 | 33.6 | 35.3 | 35.7 | 35.4
29.9 | 30.7 | 33.4 | 35.2 | 35.6 | 35.3
30 | 30.6 | 33.3 | 35 | 35.4 | 35.2

## SSD80 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 24.9 | 35.5 | 55.9 | 71.1 | 75.9
0.1 | 28.6 | 39.75 | 59.775 | 74.6 | 79.6
0.2 | 34.1 | 45.6667 | 64.8 | 79.1 | 83.7
0.3 | 43.425 | 53.4667 | 71.5 | 84.4 | 88
0.4 | 50.98 | 60.3 | 77.5 | 88.5 | 91.55
0.5 | 58 | 66.5 | 82.1 | 91.8 | 94.325
0.6 | 64.0167 | 71.8 | 85.625 | 94.15 | 96.125
0.7 | 69.2167 | 76.35 | 88.65 | 96 | 97.55
0.8 | 73.55 | 80.1 | 91.15 | 97.35 | 98.6
0.9 | 77.5 | 83.46 | 93.275 | 98.3 | 99.25
1 | 80.7125 | 86.15 | 94.9 | 99.05 | 99.75
1.1 | 83.62 | 88.425 | 96.2333 | 99.45 | 100
1.2 | 86.32 | 90.575 | 97.3 | 99.925 | 100
1.3 | 88.4857 | 92.275 | 98.125 | 100 | 100
1.4 | 90.35 | 93.825 | 98.74 | 100 | 100
1.5 | 92.0667 | 95 | 99.2 | 100 | 99.74
1.6 | 93.6 | 96 | 99.55 | 99.9 | 99.6
1.7 | 94.9 | 97.0667 | 99.82 | 99.675 | 99.2714
1.8 | 96 | 97.8 | 100 | 99.44 | 98.9
1.9 | 96.96 | 98.3333 | 100 | 99.2 | 98.6
2 | 97.6 | 98.9 | 100 | 98.825 | 98.2
2.1 | 98.32 | 99.3 | 99.9667 | 98.6 | 97.8
2.2 | 98.8 | 99.6 | 99.86 | 98.15 | 97.35
2.3 | 99.25 | 99.85 | 99.725 | 97.8 | 96.925
2.4 | 99.575 | 100 | 99.475 | 97.4 | 96.5
2.5 | 99.7 | 100 | 99.225 | 97 | 96.1
2.6 | 99.9 | 100 | 98.92 | 96.6 | 95.65
2.7 | 99.975 | 100 | 98.675 | 96.2 | 95.25
2.8 | 100 | 99.9 | 98.275 | 95.8 | 94.825
2.9 | 100 | 99.8 | 98.05 | 95.3 | 94.35
3 | 99.9 | 99.625 | 97.68 | 94.95 | 93.95
3.1 | 99.775 | 99.4857 | 97.34 | 94.55 | 93.45
3.2 | 99.65 | 99.25 | 97 | 94.12 | 93.05
3.3 | 99.4 | 99.1 | 96.6 | 93.65 | 92.55
3.4 | 99.2 | 98.7 | 96.2 | 93.25 | 92.2
3.5 | 99 | 98.4 | 95.8333 | 92.825 | 91.9
3.6 | 98.7 | 98.1 | 95.4 | 92.45 | 91.4
3.7 | 98.3 | 97.7 | 95.1 | 91.95 | 91.1
3.8 | 98 | 97.425 | 94.72 | 91.7 | 90.6
3.9 | 97.6 | 97 | 94.2 | 91.2333 | 90.2
4 | 97.3 | 96.7571 | 93.7667 | 90.9 | 89.9
4.1 | 96.9 | 96.42 | 93.525 | 90.55 | 89.4
4.2 | 96.5333 | 96.025 | 93.075 | 90 | 89
4.3 | 96.12 | 95.6 | 92.725 | 89.7 | 88.7
4.4 | 95.62 | 95.24 | 92.2 | 89.2 | 88.35
4.5 | 95.25 | 94.85 | 91.9 | 88.9 | 87.875
4.6 | 94.8333 | 94.4667 | 91.45 | 88.525 | 87.5
4.7 | 94.45 | 93.96 | 91.1 | 88.1143 | 87.12
4.8 | 93.925 | 93.6 | 90.65 | 87.825 | 86.75
4.9 | 93.5667 | 93.275 | 90.2125 | 87.375 | 86.375
5 | 93.1 | 92.85 | 89.8375 | 87.05 | 85.95
5.1 | 92.55 | 92.35 | 89.55 | 86.625 | 85.6667
5.2 | 92.2 | 92.05 | 89.175 | 86.1375 | 85.3
5.3 | 91.75 | 91.55 | 88.74 | 85.8571 | 84.8667
5.4 | 91.3 | 91.15 | 88.375 | 85.525 | 84.5
5.5 | 90.85 | 90.7 | 87.95 | 85.15 | 84.2333
5.6 | 90.4 | 90.3 | 87.475 | 84.7 | 83.8333
5.7 | 89.875 | 89.9 | 87.12 | 84.3333 | 83.5333
5.8 | 89.375 | 89.4 | 86.7 | 83.975 | 83.05
5.9 | 88.95 | 89.05 | 86.4 | 83.65 | 82.675
6 | 88.58 | 88.65 | 85.95 | 83.25 | 82.425
6.1 | 88.2 | 88.2333 | 85.65 | 82.95 | 82
6.2 | 87.7 | 87.775 | 85.25 | 82.575 | 81.7
6.3 | 87.2 | 87.425 | 84.86 | 82.25 | 81.3
6.4 | 86.85 | 87.025 | 84.45 | 81.8 | 81.02
6.5 | 86.6 | 86.65 | 84.1375 | 81.45 | 80.65
6.6 | 86.025 | 86.225 | 83.8667 | 81.12 | 80.275
6.7 | 85.65 | 85.9 | 83.525 | 80.82 | 80.025
6.8 | 85.175 | 85.55 | 83.15 | 80.475 | 79.65
6.9 | 84.725 | 85 | 82.7 | 80.15 | 79.34
7 | 84.35 | 84.75 | 82.2 | 79.675 | 78.9857
7.1 | 83.98 | 84.35 | 81.975 | 79.425 | 78.65
7.2 | 83.5 | 84.05 | 81.65 | 79.05 | 78.3
7.3 | 83.05 | 83.6125 | 81.25 | 78.725 | 77.95
7.4 | 82.7 | 83.15 | 80.925 | 78.5 | 77.64
7.5 | 82.3333 | 82.7625 | 80.6 | 78.05 | 77.275
7.6 | 81.925 | 82.4 | 80.175 | 77.7667 | 77.025
7.7 | 81.45 | 82 | 79.8429 | 77.4 | 76.75
7.8 | 81 | 81.6 | 79.6 | 77.12 | 76.325
7.9 | 80.7 | 81.24 | 79.1 | 76.82 | 75.9778
8 | 80.25 | 80.8667 | 78.7333 | 76.4 | 75.725
8.1 | 79.9 | 80.5 | 78.45 | 76.075 | 75.4
8.2 | 79.375 | 80.15 | 78.025 | 75.8 | 75.1
8.3 | 79.125 | 79.725 | 77.775 | 75.4625 | 74.7667
8.4 | 78.75 | 79.4 | 77.45 | 75.1333 | 74.45
8.5 | 78.35 | 79.04 | 77.075 | 74.85 | 74.12
8.6 | 77.85 | 78.6 | 76.75 | 74.55 | 73.6857
8.7 | 77.45 | 78.2 | 76.4 | 74.2 | 73.425
8.8 | 77.1 | 77.95 | 76.05 | 73.8333 | 73.125
8.9 | 76.7 | 77.55 | 75.74 | 73.4875 | 72.8
9 | 76.4 | 77.175 | 75.375 | 73.175 | 72.5333
9.1 | 75.95 | 76.85 | 75.05 | 72.925 | 72.2
9.2 | 75.62 | 76.5 | 74.65 | 72.55 | 71.9
9.3 | 75.25 | 76.05 | 74.325 | 72.25 | 71.6
9.4 | 74.75 | 75.7 | 74.05 | 72 | 71.275
9.5 | 74.375 | 75.25 | 73.6667 | 71.675 | 71.025
9.6 | 74.0875 | 74.9 | 73.3667 | 71.325 | 70.575
9.7 | 73.65 | 74.6 | 73.0333 | 71.05 | 70.35
9.8 | 73.2375 | 74.3 | 72.65 | 70.625 | 69.975
9.9 | 72.8667 | 73.9 | 72.425 | 70.4 | 69.725
10 | 72.5143 | 73.5 | 72.04 | 70.04 | 69.4
10.1 | 72.175 | 73.25 | 71.75 | 69.7 | 69.1333
10.2 | 71.78 | 72.85 | 71.325 | 69.45 | 68.8
10.3 | 71.5 | 72.4667 | 71.075 | 69.2 | 68.5
10.4 | 71.1143 | 72.175 | 70.825 | 68.85 | 68.2
10.5 | 70.7 | 71.85 | 70.4625 | 68.5 | 67.925
10.6 | 70.35 | 71.525 | 70.14 | 68.2667 | 67.6
10.7 | 69.975 | 71.15 | 69.8 | 67.925 | 67.32
10.8 | 69.65 | 70.825 | 69.52 | 67.675 | 67.025
10.9 | 69.2 | 70.46 | 69.15 | 67.35 | 66.7333
11 | 69 | 70.1 | 68.875 | 67.1 | 66.4
11.1 | 68.6 | 69.8 | 68.55 | 66.75 | 66.16
11.2 | 68.2333 | 69.55 | 68.275 | 66.5 | 65.85
11.3 | 67.9333 | 69.2 | 67.94 | 66.2 | 65.675
11.4 | 67.6 | 68.8 | 67.7333 | 65.95 | 65.35
11.5 | 67.35 | 68.5 | 67.35 | 65.6 | 65.05
11.6 | 67 | 68.25 | 67.025 | 65.375 | 64.74
11.7 | 66.6 | 67.9 | 66.825 | 65.05 | 64.48
11.8 | 66.24 | 67.55 | 66.475 | 64.85 | 64.25
11.9 | 65.96 | 67.225 | 66.24 | 64.475 | 64
12 | 65.65 | 66.975 | 65.875 | 64.25 | 63.65
12.1 | 65.25 | 66.6375 | 65.5333 | 64 | 63.4
12.2 | 64.925 | 66.3333 | 65.3 | 63.65 | 63.15
12.3 | 64.675 | 66.05 | 65.075 | 63.34 | 62.82
12.4 | 64.35 | 65.7 | 64.75 | 63.1 | 62.6
12.5 | 63.9 | 65.35 | 64.475 | 62.925 | 62.3333
12.6 | 63.65 | 65.05 | 64.25 | 62.65 | 62.1
12.7 | 63.375 | 64.675 | 63.8875 | 62.325 | 61.775
12.8 | 63.05 | 64.4222 | 63.575 | 62 | 61.5
12.9 | 62.7 | 64.15 | 63.2375 | 61.825 | 61.25
13 | 62.4333 | 63.85 | 63 | 61.48 | 61
13.1 | 62.1 | 63.575 | 62.725 | 61.25 | 60.7333
13.2 | 61.75 | 63.25 | 62.45 | 61 | 60.4667
13.3 | 61.4 | 62.9 | 62.175 | 60.7 | 60.2
13.4 | 61.15 | 62.6286 | 61.9 | 60.4333 | 59.9667
13.5 | 60.8 | 62.3714 | 61.7 | 60.1 | 59.7
13.6 | 60.55 | 62 | 61.35 | 59.95 | 59.5
13.7 | 60.225 | 61.75 | 61.125 | 59.64 | 59.1333
13.8 | 59.975 | 61.5 | 60.775 | 59.375 | 58.9667
13.9 | 59.725 | 61.22 | 60.55 | 59.125 | 58.7
14 | 59.375 | 60.925 | 60.275 | 58.875 | 58.5
14.1 | 59.125 | 60.65 | 60.025 | 58.625 | 58.2667
14.2 | 58.85 | 60.325 | 59.775 | 58.38 | 57.9667
14.3 | 58.45 | 60.05 | 59.44 | 58.15 | 57.6667
14.4 | 58.2 | 59.75 | 59.175 | 57.9 | 57.4571
14.5 | 57.9 | 59.5 | 58.94 | 57.65 | 57.1667
14.6 | 57.6 | 59.3333 | 58.7 | 57.4 | 56.95
14.7 | 57.375 | 58.9 | 58.4333 | 57.1 | 56.7
14.8 | 57.125 | 58.65 | 58.1 | 56.9 | 56.45
14.9 | 56.775 | 58.3 | 57.85 | 56.6 | 56.2
15 | 56.45 | 58.1143 | 57.6 | 56.4 | 55.9
15.1 | 56.175 | 57.8167 | 57.4 | 56.175 | 55.7
15.2 | 55.85 | 57.575 | 57.12 | 55.925 | 55.4429
15.3 | 55.58 | 57.2875 | 56.85 | 55.6 | 55.1833
15.4 | 55.35 | 56.975 | 56.6 | 55.4 | 55.0375
15.5 | 55.1 | 56.6667 | 56.2667 | 55.15 | 54.7375
15.6 | 54.8 | 56.5 | 56.1 | 54.9333 | 54.56
15.7 | 54.55 | 56.15 | 55.9 | 54.6714 | 54.25
15.8 | 54.1 | 55.9 | 55.6 | 54.5 | 54
15.9 | 53.925 | 55.65 | 55.3 | 54.2429 | 53.8
16 | 53.7333 | 55.425 | 55.1429 | 54 | 53.5
16.1 | 53.38 | 55.12 | 54.85 | 53.775 | 53.35
16.2 | 53.2 | 54.8625 | 54.62 | 53.575 | 53.125
16.3 | 52.975 | 54.5857 | 54.375 | 53.25 | 52.875
16.4 | 52.725 | 54.4 | 54.14 | 53.075 | 52.725
16.5 | 52.5 | 54.1 | 53.88 | 52.825 | 52.475
16.6 | 52.1667 | 53.875 | 53.6333 | 52.6 | 52.225
16.7 | 52 | 53.625 | 53.4 | 52.4 | 52
16.8 | 51.75 | 53.4 | 53.15 | 52.1 | 51.8
16.9 | 51.5 | 53.1 | 52.9 | 51.95 | 51.5
17 | 51.1857 | 52.9 | 52.75 | 51.725 | 51.36
17.1 | 50.925 | 52.6 | 52.5 | 51.575 | 51.1
17.2 | 50.66 | 52.35 | 52.25 | 51.225 | 50.975
17.3 | 50.5 | 52.1 | 52.025 | 51.075 | 50.725
17.4 | 50.175 | 51.9 | 51.8 | 50.825 | 50.475
17.5 | 49.925 | 51.65 | 51.55 | 50.6 | 50.225
17.6 | 49.775 | 51.4 | 51.3 | 50.425 | 50.075
17.7 | 49.525 | 51.15 | 51.1 | 50.2 | 49.9
17.8 | 49.26 | 50.9 | 50.9 | 50 | 49.675
17.9 | 48.95 | 50.65 | 50.7 | 49.8 | 49.4
18 | 48.775 | 50.46 | 50.4667 | 49.58 | 49.2
18.1 | 48.425 | 50.2 | 50.2 | 49.35 | 49.0333
18.2 | 48.25 | 50 | 49.96 | 49.1 | 48.8
18.3 | 48.025 | 49.725 | 49.7 | 48.95 | 48.6
18.4 | 47.8 | 49.475 | 49.5833 | 48.7 | 48.4
18.5 | 47.55 | 49.2333 | 49.4167 | 48.475 | 48.2
18.6 | 47.2 | 49.05 | 49.1 | 48.325 | 47.95
18.7 | 47.1 | 48.85 | 48.9125 | 48.075 | 47.775
18.8 | 46.9 | 48.6 | 48.675 | 47.825 | 47.55
18.9 | 46.6 | 48.375 | 48.425 | 47.675 | 47.375
19 | 46.375 | 48.1333 | 48.275 | 47.425 | 47.15
19.1 | 46.125 | 47.9 | 48.025 | 47.275 | 46.95
19.2 | 45.975 | 47.7 | 47.775 | 47.025 | 46.775
19.3 | 45.75 | 47.4667 | 47.525 | 46.775 | 46.55
19.4 | 45.525 | 47.275 | 47.4 | 46.65 | 46.3
19.5 | 45.25 | 47 | 47.15 | 46.38 | 46.14
19.6 | 45.04 | 46.8 | 47 | 46.25 | 45.925
19.7 | 44.9 | 46.65 | 46.775 | 46 | 45.7
19.8 | 44.55 | 46.4 | 46.525 | 45.8 | 45.6
19.9 | 44.3 | 46.15 | 46.3167 | 45.6167 | 45.3667
20 | 44.16 | 45.95 | 46.1 | 45.4 | 45.2
20.1 | 43.975 | 45.6667 | 45.95 | 45.3 | 44.95
20.2 | 43.7 | 45.4 | 45.7 | 45.075 | 44.72
20.3 | 43.575 | 45.25 | 45.55 | 44.9 | 44.575
20.4 | 43.325 | 45.1 | 45.4 | 44.675 | 44.42
20.5 | 43.075 | 44.85 | 45.075 | 44.4333 | 44.175
20.6 | 42.825 | 44.62 | 44.9 | 44.2 | 44.025
20.7 | 42.575 | 44.375 | 44.675 | 44.0333 | 43.7833
20.8 | 42.425 | 44.225 | 44.5125 | 43.925 | 43.675
20.9 | 42.18 | 44 | 44.375 | 43.7333 | 43.44
21 | 42.04 | 43.825 | 44.15 | 43.5 | 43.3
21.1 | 41.9 | 43.58 | 44 | 43.35 | 43.1
21.2 | 41.65 | 43.45 | 43.8 | 43.1 | 42.9
21.3 | 41.475 | 43.2 | 43.58 | 42.95 | 42.725
21.4 | 41.34 | 43.025 | 43.3 | 42.8 | 42.575
21.5 | 41.15 | 42.8333 | 43.15 | 42.6 | 42.325
21.6 | 40.875 | 42.6 | 42.9667 | 42.425 | 42.1875
21.7 | 40.7 | 42.4 | 42.8 | 42.275 | 42.025
21.8 | 40.55 | 42.2 | 42.6 | 42.1 | 41.8375
21.9 | 40.3 | 42.05 | 42.475 | 41.9667 | 41.625
22 | 40.15 | 41.8 | 42.2 | 41.7 | 41.48
22.1 | 39.9 | 41.675 | 42.075 | 41.575 | 41.3
22.2 | 39.65 | 41.425 | 41.925 | 41.3667 | 41.1667
22.3 | 39.5 | 41.275 | 41.775 | 41.22 | 41
22.4 | 39.36 | 41.1 | 41.525 | 41 | 40.8
22.5 | 39.2 | 40.875 | 41.375 | 40.825 | 40.675
22.6 | 39.075 | 40.7 | 41.1 | 40.68 | 40.425
22.7 | 38.725 | 40.5 | 41 | 40.5 | 40.3
22.8 | 38.575 | 40.325 | 40.775 | 40.4 | 40.1
22.9 | 38.425 | 40.175 | 40.5857 | 40.15 | 39.94
23 | 38.2 | 39.95 | 40.42 | 40 | 39.8
23.1 | 38 | 39.7 | 40.275 | 39.85 | 39.65
23.2 | 37.8 | 39.55 | 40.025 | 39.6 | 39.4
23.3 | 37.675 | 39.4 | 39.9 | 39.46 | 39.3
23.4 | 37.5 | 39.2 | 39.75 | 39.225 | 39.1
23.5 | 37.35 | 39 | 39.6 | 39.1 | 38.95
23.6 | 37.15 | 38.8 | 39.35 | 38.925 | 38.8
23.7 | 37 | 38.6 | 39.2 | 38.8 | 38.65
23.8 | 36.8 | 38.5 | 39.05 | 38.64 | 38.5
23.9 | 36.6 | 38.3 | 38.9 | 38.5 | 38.3
24 | 36.475 | 38.1667 | 38.775 | 38.35 | 38.175
24.1 | 36.22 | 38 | 38.525 | 38.2 | 38
24.2 | 36.1 | 37.8 | 38.375 | 38 | 37.825
24.3 | 35.9 | 37.6 | 38.2 | 37.8 | 37.7
24.4 | 35.775 | 37.4333 | 38.025 | 37.6333 | 37.4875
24.5 | 35.6 | 37.275 | 37.875 | 37.5 | 37.3625
24.6 | 35.4 | 37.125 | 37.725 | 37.35 | 37.175
24.7 | 35.225 | 36.975 | 37.5 | 37.2 | 37.0125
24.8 | 35.1 | 36.8 | 37.3667 | 36.9833 | 36.8429
24.9 | 34.85 | 36.625 | 37.2 | 36.875 | 36.7
25 | 34.7 | 36.45 | 37.025 | 36.7333 | 36.52
25.1 | 34.6333 | 36.2 | 36.9 | 36.58 | 36.36
25.2 | 34.4 | 36.1 | 36.7 | 36.425 | 36.225
25.3 | 34.3 | 35.9 | 36.54 | 36.2 | 36.0667
25.4 | 34.1167 | 35.8 | 36.4 | 36.1125 | 35.95
25.5 | 33.875 | 35.6 | 36.3 | 35.8875 | 35.725
25.6 | 33.725 | 35.325 | 36.1 | 35.8 | 35.6
25.7 | 33.575 | 35.275 | 35.9 | 35.675 | 35.475
25.8 | 33.44 | 35.1 | 35.75 | 35.4444 | 35.3
25.9 | 33.3 | 34.975 | 35.625 | 35.3 | 35.2
26 | 33.1333 | 34.725 | 35.5 | 35.14 | 35.05
26.1 | 33 | 34.68 | 35.3 | 35 | 34.9
26.2 | 32.8 | 34.45 | 35.15 | 34.9 | 34.8
26.3 | 32.725 | 34.3 | 35 | 34.7 | 34.6
26.4 | 32.5 | 34.15 | 34.85 | 34.6 | 34.5
26.5 | 32.325 | 34 | 34.7 | 34.475 | 34.325
26.6 | 32.2 | 33.85 | 34.56 | 34.3 | 34.175
26.7 | 32 | 33.7 | 34.325 | 34.2 | 34.05
26.8 | 31.9 | 33.5 | 34.2 | 34.1 | 33.825
26.9 | 31.775 | 33.4 | 34.1 | 33.9 | 33.6889
27 | 31.625 | 33.25 | 34.0111 | 33.725 | 33.575
27.1 | 31.4714 | 33.1 | 33.8 | 33.6 | 33.45
27.2 | 31.25 | 32.975 | 33.6625 | 33.5 | 33.325
27.3 | 31.2 | 32.8 | 33.5167 | 33.3 | 33.2
27.4 | 30.975 | 32.675 | 33.4 | 33.2 | 33.04
27.5 | 30.825 | 32.525 | 33.275 | 33 | 32.9
27.6 | 30.7625 | 32.375 | 33.1 | 32.95 | 32.8
27.7 | 30.6375 | 32.2 | 32.95 | 32.8 | 32.7
27.8 | 30.5125 | 32.075 | 32.8 | 32.6 | 32.5
27.9 | 30.2875 | 31.9 | 32.65 | 32.5 | 32.3
28 | 30.1625 | 31.8 | 32.52 | 32.34 | 32.2
28.1 | 30.0375 | 31.625 | 32.375 | 32.2 | 32.1
28.2 | 29.825 | 31.5 | 32.225 | 32.1 | 31.95
28.3 | 29.7 | 31.4 | 32.075 | 31.95 | 31.8
28.4 | 29.6571 | 31.2 | 31.925 | 31.8333 | 31.7
28.5 | 29.5 | 31.05 | 31.8 | 31.7 | 31.6
28.6 | 29.3 | 30.9 | 31.7 | 31.575 | 31.5
28.7 | 29.2 | 30.76 | 31.5 | 31.4 | 31.3
28.8 | 29.1 | 30.625 | 31.4 | 31.2 | 31.2
28.9 | 28.9 | 30.475 | 31.2571 | 31.1 | 31
29 | 28.8 | 30.325 | 31.1143 | 31.025 | 30.9
29.1 | 28.7 | 30.2 | 30.98 | 30.875 | 30.8
29.2 | 28.6 | 30.1 | 30.825 | 30.725 | 30.6333
29.3 | 28.4 | 30 | 30.775 | 30.6333 | 30.6
29.4 | 28.3 | 29.8 | 30.6 | 30.5 | 30.4
29.5 | 28.1 | 29.7 | 30.5 | 30.4 | 30.3
29.6 | 28 | 29.6 | 30.4 | 30.2 | 30.2
29.7 | 27.8667 | 29.45 | 30.2143 | 30.16 | 30.1
29.8 | 27.7 | 29.3 | 30.1 | 30.025 | 29.9
29.9 | 27.6 | 29.15 | 29.9833 | 29.875 | 29.8
30 | 27.5 | 29 | 29.875 | 29.725 | 29.65

## SSD120 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 22.9 | 29.7 | 40.6 | 51.1 | 55.1
0.1 | 26.7 | 33.05 | 44.475 | 54.6333 | 58.85
0.2 | 32.6333 | 39.02 | 49.84 | 59.8 | 63.825
0.3 | 41.175 | 46.85 | 57.3 | 66.3 | 70.125
0.4 | 48.525 | 54.06 | 63.7 | 72.3 | 75.65
0.5 | 55.175 | 60.3 | 69.15 | 76.86 | 79.9
0.6 | 60.62 | 65.55 | 73.425 | 80.75 | 83.45
0.7 | 65.4 | 69.9333 | 77.66 | 84.25 | 86.5
0.8 | 70 | 74.05 | 80.825 | 86.925 | 89.175
0.9 | 74.05 | 77.85 | 83.875 | 89.5 | 91.3
1 | 77.4 | 80.775 | 86.65 | 91.325 | 93.025
1.1 | 80.5333 | 83.5857 | 88.725 | 93 | 94.55
1.2 | 83.3 | 86.1 | 90.725 | 94.4333 | 95.7667
1.3 | 85.5 | 87.9 | 92.14 | 95.7667 | 96.8
1.4 | 87.7 | 89.85 | 93.5667 | 96.7 | 97.68
1.5 | 89.625 | 91.35 | 94.7857 | 97.4571 | 98.3
1.6 | 91.2 | 92.9167 | 95.7 | 98.04 | 98.85
1.7 | 92.4 | 94 | 96.575 | 98.68 | 99.3
1.8 | 93.85 | 95.175 | 97.4 | 99.1714 | 99.55
1.9 | 95 | 96.1 | 98 | 99.3571 | 99.8
2 | 95.6333 | 97.0125 | 98.6 | 99.7 | 99.9
2.1 | 96.7 | 97.5 | 98.95 | 99.9 | 99.9
2.2 | 97.5 | 98.075 | 99.3 | 99.9875 | 100
2.3 | 97.95 | 98.6 | 99.6333 | 100 | 100
2.4 | 98.525 | 99.1 | 99.7 | 100 | 99.9
2.5 | 98.975 | 99.5 | 99.9 | 100 | 99.825
2.6 | 99.3 | 99.6667 | 99.975 | 99.9 | 99.7
2.7 | 99.6 | 99.8 | 100 | 99.8333 | 99.6
2.8 | 99.75 | 99.9 | 100 | 99.6333 | 99.4
2.9 | 99.9 | 99.9 | 100 | 99.4 | 99.28
3 | 100 | 100 | 99.9 | 99.275 | 98.98
3.1 | 100 | 100 | 99.8 | 99.025 | 98.675
3.2 | 99.9 | 100 | 99.675 | 98.775 | 98.45
3.3 | 99.9 | 99.9167 | 99.425 | 98.55 | 98.2
3.4 | 99.9 | 99.78 | 99.275 | 98.4 | 97.95
3.5 | 99.8 | 99.6333 | 99.125 | 98.15 | 97.725
3.6 | 99.5 | 99.5833 | 98.8 | 98 | 97.5333
3.7 | 99.375 | 99.375 | 98.625 | 97.6 | 97.1667
3.8 | 99.125 | 99.24 | 98.5 | 97.2 | 96.8333
3.9 | 98.9 | 98.95 | 98.2333 | 96.95 | 96.475
4 | 98.72 | 98.65 | 97.775 | 96.72 | 96.225
4.1 | 98.4667 | 98.4667 | 97.5 | 96.42 | 95.9125
4.2 | 98.225 | 98.22 | 97.4 | 96.075 | 95.625
4.3 | 97.975 | 97.85 | 97.05 | 95.7875 | 95.4
4.4 | 97.65 | 97.5857 | 96.7 | 95.4333 | 95.0429
4.5 | 97.25 | 97.36 | 96.45 | 95.3 | 94.7125
4.6 | 97 | 97.12 | 96.2 | 95.0125 | 94.4
4.7 | 96.65 | 96.675 | 95.9333 | 94.625 | 94.0667
4.8 | 96.4 | 96.525 | 95.625 | 94.275 | 93.76
4.9 | 96.05 | 96.175 | 95.375 | 94.025 | 93.525
5 | 95.9 | 95.8333 | 95.1444 | 93.78 | 93.3
5.1 | 95.54 | 95.5 | 94.5 | 93.55 | 92.975
5.2 | 95.05 | 95.25 | 94.15 | 93.0333 | 92.625
5.3 | 94.4667 | 94.9 | 93.9 | 92.8667 | 92.275
5.4 | 94.22 | 94.45 | 93.65 | 92.5 | 92.15
5.5 | 93.75 | 94.2 | 93.4 | 92.2333 | 91.8
5.6 | 93.4333 | 93.9 | 93.05 | 91.9 | 91.45
5.7 | 93.14 | 93.65 | 92.82 | 91.5333 | 91.2
5.8 | 92.8625 | 93.2 | 92.4625 | 91.3 | 90.96
5.9 | 92.5 | 92.86 | 92.1625 | 91.05 | 90.625
6 | 92.15 | 92.55 | 91.9333 | 90.74 | 90.275
6.1 | 91.7 | 92.225 | 91.575 | 90.375 | 90.05
6.2 | 91.3 | 91.875 | 91.2 | 90.125 | 89.75
6.3 | 90.95 | 91.55 | 90.825 | 89.86 | 89.35
6.4 | 90.6 | 91.3 | 90.5 | 89.625 | 89.2
6.5 | 90.05 | 90.85 | 90.1857 | 89.25 | 88.85
6.6 | 89.72 | 90.5 | 89.875 | 88.925 | 88.5
6.7 | 89.45 | 90.15 | 89.7 | 88.75 | 88.25
6.8 | 89.1 | 89.85 | 89.475 | 88.45 | 87.925
6.9 | 88.72 | 89.55 | 89.2 | 88 | 87.675
7 | 88.3 | 89.225 | 88.7 | 87.75 | 87.425
7.1 | 87.92 | 88.875 | 88.4 | 87.5 | 87.08
7.2 | 87.625 | 88.6 | 88.1 | 87.22 | 86.85
7.3 | 87.375 | 88.25 | 87.775 | 86.925 | 86.6
7.4 | 86.875 | 87.9 | 87.625 | 86.575 | 86.2
7.5 | 86.38 | 87.5 | 87.3 | 86.425 | 86
7.6 | 86.05 | 87.2 | 87.125 | 86.25 | 85.7
7.7 | 85.75 | 86.975 | 86.575 | 85.7625 | 85.325
7.8 | 85.45 | 86.55 | 86.4 | 85.4333 | 85.1
7.9 | 85.2 | 86.375 | 86.2 | 85.1857 | 84.7625
8 | 84.65 | 86.125 | 86 | 84.85 | 84.4625
8.1 | 84.4 | 85.7 | 85.6 | 84.6111 | 84.2889
8.2 | 84.05 | 85.4 | 85.2 | 84.375 | 83.975
8.3 | 83.625 | 85 | 85 | 84.125 | 83.7
8.4 | 83.3333 | 84.75 | 84.65 | 83.8625 | 83.425
8.5 | 82.9714 | 84.5 | 84.1833 | 83.5143 | 83.1
8.6 | 82.55 | 84.26 | 83.975 | 83.28 | 82.9143
8.7 | 82.26 | 83.85 | 83.8 | 83.05 | 82.65
8.8 | 81.9333 | 83.475 | 83.5 | 82.65 | 82.325
8.9 | 81.65 | 83.225 | 83.2667 | 82.4 | 81.975
9 | 81.3 | 83 | 83 | 82.15 | 81.75
9.1 | 80.825 | 82.6 | 82.65 | 82 | 81.5
9.2 | 80.45 | 82.175 | 82.4 | 81.75 | 81.15
9.3 | 80.15 | 81.94 | 82 | 81.3 | 80.9333
9.4 | 79.6875 | 81.6 | 81.675 | 80.95 | 80.7
9.5 | 79.275 | 81.35 | 81.38 | 80.725 | 80.4
9.6 | 78.98 | 81.05 | 81.1222 | 80.375 | 80.1
9.7 | 78.55 | 80.8 | 80.875 | 80.225 | 79.74
9.8 | 78.3 | 80.4 | 80.5875 | 79.9667 | 79.46
9.9 | 78.15 | 80 | 80.275 | 79.55 | 79.3
10 | 77.8 | 79.725 | 80.025 | 79.3 | 79.0333
10.1 | 77.45 | 79.5 | 79.6714 | 79.05 | 78.8
10.2 | 77.025 | 79.35 | 79.38 | 78.825 | 78.5
10.3 | 76.75 | 79 | 79.15 | 78.575 | 78.25
10.4 | 76.35 | 78.375 | 78.9 | 78.325 | 78
10.5 | 75.96 | 78.225 | 78.56 | 78 | 77.75
10.6 | 75.625 | 77.975 | 78.3778 | 77.78 | 77.5
10.7 | 75.36 | 77.65 | 78.2 | 77.5 | 77.1
10.8 | 75.05 | 77.375 | 77.74 | 77.16 | 76.95
10.9 | 74.8 | 77.05 | 77.45 | 76.925 | 76.625
11 | 74.5 | 76.775 | 77.2 | 76.675 | 76.375
11.1 | 74.1 | 76.525 | 76.9 | 76.325 | 76.025
11.2 | 73.85 | 76.1667 | 76.6667 | 76.075 | 75.7
11.3 | 73.625 | 75.9 | 76.3667 | 75.825 | 75.6
11.4 | 73.275 | 75.6667 | 76.1 | 75.7 | 75.4
11.5 | 72.95 | 75.4 | 75.85 | 75.3333 | 75.05
11.6 | 72.4875 | 75.1333 | 75.625 | 75.1 | 74.8667
11.7 | 72.2875 | 74.825 | 75.3 | 74.85 | 74.675
11.8 | 71.9125 | 74.55 | 75.075 | 74.7 | 74.325
11.9 | 71.625 | 74.325 | 74.825 | 74.35 | 74.075
12 | 71.4 | 74.075 | 74.575 | 74.1 | 73.75
12.1 | 71.05 | 73.7 | 74.325 | 73.85 | 73.48
12.2 | 70.8 | 73.3 | 73.9 | 73.62 | 73.25
12.3 | 70.5667 | 73.05 | 73.725 | 73.35 | 73
12.4 | 70.2375 | 72.8 | 73.5 | 73.025 | 72.7
12.5 | 69.85 | 72.45 | 73.18 | 72.875 | 72.5
12.6 | 69.4625 | 72.25 | 72.9 | 72.625 | 72.375
12.7 | 69.2125 | 71.9 | 72.7 | 72.35 | 72.125
12.8 | 68.8375 | 71.6143 | 72.5 | 72.1 | 71.775
12.9 | 68.6778 | 71.375 | 72.3 | 71.9 | 71.55
13 | 68.425 | 71.2 | 71.9 | 71.6667 | 71.35
13.1 | 68.15 | 70.9 | 71.6 | 71.38 | 71.125
13.2 | 67.875 | 70.725 | 71.35 | 71.15 | 70.8
13.3 | 67.5 | 70.475 | 71.1 | 70.875 | 70.55
13.4 | 67.25 | 70.05 | 70.85 | 70.7 | 70.3
13.5 | 66.975 | 69.775 | 70.625 | 70.425 | 70.1
13.6 | 66.65 | 69.6 | 70.425 | 70.1429 | 69.82
13.7 | 66.275 | 69.375 | 70.2 | 69.925 | 69.5571
13.8 | 65.925 | 69.025 | 69.875 | 69.65 | 69.375
13.9 | 65.8 | 68.7 | 69.55 | 69.275 | 69.1833
14 | 65.45 | 68.375 | 69.4 | 69.1 | 68.98
14.1 | 65.1 | 68.14 | 69.075 | 68.95 | 68.625
14.2 | 64.96 | 67.9 | 68.9 | 68.62 | 68.3
14.3 | 64.6 | 67.6 | 68.6 | 68.4 | 68.1143
14.4 | 64.25 | 67.4 | 68.3 | 68.15 | 67.8
14.5 | 64.02 | 67.15 | 68.15 | 67.9 | 67.65
14.6 | 63.7875 | 66.92 | 67.9 | 67.65 | 67.45
14.7 | 63.5875 | 66.675 | 67.66 | 67.4 | 67.2
14.8 | 63.15 | 66.425 | 67.5 | 67.16 | 67
14.9 | 62.825 | 66.15 | 67.175 | 66.975 | 66.7667
15 | 62.575 | 65.825 | 66.925 | 66.725 | 66.525
15.1 | 62.3571 | 65.675 | 66.675 | 66.475 | 66.3
15.2 | 62.125 | 65.325 | 66.425 | 66.225 | 66.05
15.3 | 62 | 65 | 66.175 | 66.05 | 65.775
15.4 | 61.725 | 64.65 | 65.925 | 65.74 | 65.6
15.5 | 61.35 | 64.4 | 65.8 | 65.6 | 65.3
15.6 | 61.1 | 64.25 | 65.56 | 65.45 | 65.15
15.7 | 60.9333 | 64 | 65.25 | 65.3 | 64.8
15.8 | 60.6333 | 63.8 | 64.975 | 64.9 | 64.575
15.9 | 60.375 | 63.5 | 64.725 | 64.7 | 64.4
16 | 60.0875 | 63.3 | 64.6 | 64.5 | 64.225
16.1 | 59.9 | 63.08 | 64.35 | 64.25 | 63.875
16.2 | 59.55 | 62.85 | 64.2 | 64.025 | 63.74
16.3 | 59.3 | 62.7 | 63.85 | 63.8 | 63.6
16.4 | 59 | 62.45 | 63.6 | 63.55 | 63.3286
16.5 | 58.85 | 62.1 | 63.375 | 63.3 | 63.1889
16.6 | 58.525 | 61.75 | 63.2 | 63.1 | 63.025
16.7 | 58.3 | 61.6 | 62.95 | 63 | 62.7
16.8 | 58.125 | 61.35 | 62.75 | 62.75 | 62.45
16.9 | 57.775 | 61.1 | 62.4 | 62.4667 | 62.2
17 | 57.6 | 60.825 | 62.25 | 62.225 | 62.05
17.1 | 57.4 | 60.575 | 62.05 | 62.05 | 61.8429
17.2 | 57.15 | 60.325 | 61.9 | 61.825 | 61.625
17.3 | 56.8 | 60.075 | 61.6 | 61.6 | 61.475
17.4 | 56.6 | 59.75 | 61.5 | 61.4 | 61.2667
17.5 | 56.3 | 59.6 | 61.2 | 61.15 | 61
17.6 | 56.05 | 59.4 | 60.925 | 60.925 | 60.7
17.7 | 55.725 | 59.2 | 60.675 | 60.725 | 60.5
17.8 | 55.575 | 58.95 | 60.525 | 60.5 | 60.36
17.9 | 55.325 | 58.7 | 60.3 | 60.375 | 60.125
18 | 55.08 | 58.55 | 60.1 | 60.08 | 59.875
18.1 | 54.85 | 58.3 | 59.875 | 59.875 | 59.725
18.2 | 54.5 | 58.1 | 59.675 | 59.7 | 59.475
18.3 | 54.4 | 57.95 | 59.4 | 59.45 | 59.225
18.4 | 54.3 | 57.6 | 59.2 | 59.2 | 59
18.5 | 54.0556 | 57.45 | 58.975 | 59.075 | 58.8
18.6 | 53.825 | 57.2 | 58.75 | 58.8667 | 58.6667
18.7 | 53.575 | 56.95 | 58.55 | 58.6375 | 58.45
18.8 | 53.325 | 56.625 | 58.325 | 58.425 | 58.2
18.9 | 53.175 | 56.475 | 58.0778 | 58.2625 | 57.9857
19 | 52.8375 | 56.3125 | 57.85 | 57.9625 | 57.85
19.1 | 52.6 | 56.075 | 57.6 | 57.8167 | 57.65
19.2 | 52.425 | 55.9 | 57.4143 | 57.6 | 57.3667
19.3 | 52.22 | 55.68 | 57.2 | 57.35 | 57.1167
19.4 | 52 | 55.5 | 57.025 | 57.1 | 56.9
19.5 | 51.825 | 55.3 | 56.8 | 56.975 | 56.75
19.6 | 51.5667 | 55 | 56.6 | 56.72 | 56.62
19.7 | 51.275 | 54.7 | 56.375 | 56.5 | 56.375
19.8 | 51.05 | 54.56 | 56.125 | 56.325 | 56.2
19.9 | 50.9333 | 54.4 | 55.9 | 56.175 | 56
20 | 50.7667 | 54.175 | 55.7 | 55.925 | 55.9
20.1 | 50.45 | 53.94 | 55.6 | 55.7 | 55.575
20.2 | 50.3 | 53.74 | 55.35 | 55.5 | 55.4
20.3 | 50.05 | 53.3857 | 55.2 | 55.4 | 55.275
20.4 | 49.8 | 53.22 | 54.9667 | 55.15 | 55
20.5 | 49.6667 | 53.1 | 54.8 | 54.9667 | 54.825
20.6 | 49.34 | 52.825 | 54.5 | 54.9 | 54.6
20.7 | 49.0333 | 52.7 | 54.35 | 54.65 | 54.45
20.8 | 48.8 | 52.55 | 54.1 | 54.42 | 54.3
20.9 | 48.725 | 52.3667 | 53.9 | 54.2 | 54.1
21 | 48.5 | 52.1 | 53.775 | 54.025 | 53.9
21.1 | 48.25 | 51.95 | 53.525 | 53.8 | 53.7
21.2 | 48.1 | 51.7 | 53.275 | 53.625 | 53.5
21.3 | 47.95 | 51.4 | 53.1 | 53.425 | 53.375
21.4 | 47.7 | 51.25 | 52.975 | 53.175 | 53.14
21.5 | 47.5 | 51.1 | 52.725 | 53.025 | 53
21.6 | 47.3 | 50.9 | 52.6 | 52.775 | 52.75
21.7 | 47.15 | 50.65 | 52.4 | 52.6333 | 52.5
21.8 | 46.8 | 50.475 | 52.2 | 52.5 | 52.35
21.9 | 46.6 | 50.225 | 52.1 | 52.35 | 52.2
22 | 46.5 | 50.05 | 51.9 | 52.1 | 52.05
22.1 | 46.3 | 49.825 | 51.65 | 51.95 | 51.9
22.2 | 46.12 | 49.675 | 51.4 | 51.8 | 51.675
22.3 | 45.975 | 49.44 | 51.16 | 51.6 | 51.525
22.4 | 45.725 | 49.3 | 51 | 51.35 | 51.275
22.5 | 45.475 | 49.05 | 50.8667 | 51.2 | 51.1
22.6 | 45.325 | 48.8 | 50.7 | 50.975 | 50.9
22.7 | 45.1 | 48.65 | 50.525 | 50.725 | 50.65
22.8 | 44.925 | 48.5 | 50.375 | 50.575 | 50.45
22.9 | 44.775 | 48.3 | 50.1 | 50.4571 | 50.35
23 | 44.5 | 48.1 | 49.95 | 50.375 | 50.2
23.1 | 44.4 | 47.875 | 49.8 | 50.125 | 50.05
23.2 | 44.125 | 47.7 | 49.6333 | 49.975 | 49.8
23.3 | 43.975 | 47.525 | 49.4667 | 49.8 | 49.7
23.4 | 43.8 | 47.3 | 49.275 | 49.6 | 49.475
23.5 | 43.6 | 47.1 | 49.125 | 49.45 | 49.3
23.6 | 43.4 | 46.925 | 48.975 | 49.28 | 49.2
23.7 | 43.3 | 46.76 | 48.75 | 49 | 49.05
23.8 | 43.15 | 46.6 | 48.575 | 48.8333 | 48.8
23.9 | 42.9 | 46.475 | 48.3571 | 48.7 | 48.65
24 | 42.75 | 46.275 | 48.1778 | 48.6 | 48.5
24.1 | 42.52 | 46.0667 | 47.975 | 48.35 | 48.25
24.2 | 42.3 | 45.8714 | 47.85 | 48.12 | 48.02
24.3 | 42.2 | 45.7 | 47.7 | 47.95 | 47.975
24.4 | 42.075 | 45.4833 | 47.5 | 47.82 | 47.725
24.5 | 41.9 | 45.3833 | 47.3 | 47.7 | 47.575
24.6 | 41.68 | 45.2375 | 47.1667 | 47.52 | 47.4
24.7 | 41.35 | 45.05 | 46.975 | 47.4 | 47.3
24.8 | 41.2429 | 44.9 | 46.825 | 47.175 | 47.1
24.9 | 41.04 | 44.7 | 46.7 | 46.95 | 46.925
25 | 40.95 | 44.6 | 46.425 | 46.775 | 46.8
25.1 | 40.7 | 44.3333 | 46.3 | 46.6 | 46.625
25.2 | 40.6 | 44.175 | 46.2 | 46.4833 | 46.4
25.3 | 40.5 | 43.925 | 45.95 | 46.275 | 46.25
25.4 | 40.35 | 43.775 | 45.775 | 46.1125 | 46.1
25.5 | 40.14 | 43.6 | 45.6 | 45.9 | 45.95
25.6 | 39.92 | 43.4 | 45.5 | 45.75 | 45.8
25.7 | 39.7667 | 43.3 | 45.3 | 45.6 | 45.6
25.8 | 39.6 | 43.2 | 45.15 | 45.5 | 45.4
25.9 | 39.4 | 42.95 | 44.9 | 45.4 | 45.275
26 | 39.3 | 42.8 | 44.825 | 45.2 | 45.1
26.1 | 39.175 | 42.575 | 44.7 | 45.1 | 45
26.2 | 39 | 42.4833 | 44.5 | 44.875 | 44.825
26.3 | 38.775 | 42.3375 | 44.35 | 44.7 | 44.675
26.4 | 38.625 | 42.125 | 44.1 | 44.5 | 44.425
26.5 | 38.48 | 41.9875 | 43.975 | 44.38 | 44.375
26.6 | 38.4 | 41.8333 | 43.825 | 44.25 | 44.225
26.7 | 38.2 | 41.65 | 43.575 | 44 | 44.1
26.8 | 37.95 | 41.525 | 43.425 | 43.875 | 43.875
26.9 | 37.8 | 41.3 | 43.3 | 43.75 | 43.75
27 | 37.65 | 41.15 | 43.2 | 43.55 | 43.55
27.1 | 37.5 | 40.9429 | 43.1 | 43.4 | 43.4
27.2 | 37.3 | 40.8714 | 42.8 | 43.3 | 43.2778
27.3 | 37.2 | 40.7 | 42.6 | 43.2 | 43.1625
27.4 | 37.075 | 40.6 | 42.5 | 42.98 | 42.975
27.5 | 36.8625 | 40.4 | 42.32 | 42.85 | 42.8125
27.6 | 36.7375 | 40.25 | 42.275 | 42.7 | 42.5875
27.7 | 36.525 | 40.12 | 42.1 | 42.5 | 42.4571
27.8 | 36.3889 | 40 | 41.975 | 42.4 | 42.35
27.9 | 36.275 | 39.8 | 41.725 | 42.225 | 42.2
28 | 36.1 | 39.575 | 41.575 | 42 | 42.0778
28.1 | 35.925 | 39.425 | 41.4125 | 41.9 | 41.8667
28.2 | 35.8 | 39.2667 | 41.2889 | 41.8 | 41.75
28.3 | 35.675 | 39.1 | 41.15 | 41.6 | 41.7
28.4 | 35.5 | 38.875 | 40.9429 | 41.5 | 41.475
28.5 | 35.375 | 38.7 | 40.85 | 41.325 | 41.225
28.6 | 35.125 | 38.5 | 40.7 | 41.1 | 41.1
28.7 | 35 | 38.325 | 40.55 | 41 | 41
28.8 | 34.9 | 38.2 | 40.4 | 40.875 | 40.8667
28.9 | 34.775 | 38.125 | 40.2333 | 40.7 | 40.7
29 | 34.7 | 38 | 40.1 | 40.6 | 40.6
29.1 | 34.6 | 37.825 | 39.9 | 40.5 | 40.4
29.2 | 34.4 | 37.74 | 39.8 | 40.275 | 40.3
29.3 | 34.2 | 37.6 | 39.65 | 40.125 | 40.2
29.4 | 34.1 | 37.45 | 39.5 | 40 | 40
29.5 | 33.95 | 37.3 | 39.3 | 39.825 | 39.9
29.6 | 33.7 | 37.125 | 39.2 | 39.775 | 39.7
29.7 | 33.65 | 36.98 | 39.1 | 39.6 | 39.6
29.8 | 33.55 | 36.8 | 38.94 | 39.5 | 39.4333
29.9 | 33.4 | 36.6 | 38.8 | 39.3 | 39.3
30 | 33.2333 | 36.5 | 38.6 | 39.175 | 39.2

## Wedge 30 Lateral Profiles
Lateral Profiles
Wedge 30
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:3 cm FieldSize:4x4 cm2 | Depth:3 cm FieldSize:10x10 cm2 | Depth:3 cm FieldSize:20x20 cm2 | Depth:3 cm FieldSize:30x30 cm2 | Depth:3 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.9 | [18 empty] | 12.05 | 13.1522 |  |  |  | 17.72 | 18.88
-22.8 | [18 empty] | 12.2333 | 13.3022 |  |  |  | 17.94 | 19.1
-22.7 | [18 empty] | 12.4714 | 13.488 |  |  |  | 18.15 | 19.2625
-22.6 | [18 empty] | 12.64 | 13.7023 |  |  |  | 18.44 | 19.6333
-22.5 | [18 empty] | 12.8333 | 13.9523 |  |  |  | 18.72 | 19.8833
-22.4 | [18 empty] | 13.0833 | 14.1024 |  |  |  | 19.025 | 20.175
-22.3 | [18 empty] | 13.1667 | 14.3691 |  |  |  | 19.26 | 20.4875
-22.2 | [18 empty] | 13.45 | 14.5024 |  |  |  | 19.6 | 20.8
-22.1 | [18 empty] | 13.625 | 14.7692 |  |  |  | 19.9 | 21.05
-22 | [18 empty] | 13.85 | 14.9358 |  |  |  | 20.28 | 21.4
-21.9 | [18 empty] | 14.1 | 15.2025 |  |  |  | 20.65 | 21.7333
-21.8 | [18 empty] | 14.3333 | 15.4312 |  |  |  | 20.875 | 22.08
-21.7 | [18 empty] | 14.5167 | 15.6455 |  |  |  | 21.26 | 22.4667
-21.6 | [18 empty] | 14.7667 | 15.836 |  |  |  | 21.6 | 22.7429
-21.5 | [18 empty] | 15.025 | 16.0741 |  |  |  | 21.9 | 23.15
-21.4 | [13 empty] | 9.6 | 10.54 |  |  |  | 15.26 | 16.3456 |  |  |  | 22.44 | 23.525
-21.3 | [13 empty] | 9.72 | 10.775 |  |  |  | 15.5286 | 16.6361 |  |  |  | 22.85 | 24.0125
-21.2 | [13 empty] | 9.94 | 10.9 |  |  |  | 15.8 | 16.9028 |  |  |  | 23.32 | 24.45
-21.1 | [13 empty] | 10.075 | 11.1 |  |  |  | 16 | 17.1171 |  |  |  | 23.82 | 24.94
-21 | [13 empty] | 10.2 | 11.22 |  |  |  | 16.2571 | 17.4029 |  |  |  | 24.35 | 25.4429
-20.9 | [13 empty] | 10.4 | 11.45 |  |  |  | 16.5667 | 17.6697 |  |  |  | 24.94 | 26.0333
-20.8 | [13 empty] | 10.5 | 11.6 |  |  |  | 16.8571 | 17.923 |  |  |  | 25.65 | 26.7
-20.7 | [8 empty] | 8.3 | 9.2 |  |  |  | 10.725 | 11.74 |  |  |  | 17.1667 | 18.243 |  |  |  | 26.3167 | 27.45
-20.6 | [8 empty] | 8.4 | 9.4 |  |  |  | 10.86 | 11.9 |  |  |  | 17.4 | 18.5031 |  |  |  | 27.3 | 28.26
-20.5 | [8 empty] | 8.6 | 9.5 |  |  |  | 11.1 | 12.125 |  |  |  | 17.7333 | 18.8531 |  |  |  | 28.2 | 29.2143
-20.4 |  |  |  | 8.9768 | 9.975 |  |  |  | 8.7 | 9.66 |  |  |  | 11.26 | 12.3 |  |  |  | 18.0429 | 19.2199 |  |  |  | 29.4 | 30.45
-20.3 |  |  |  | 9.1018 | 10.15 |  |  |  | 8.875 | 9.8 |  |  |  | 11.5 | 12.5 |  |  |  | 18.4143 | 19.47 |  |  |  | 30.92 | 31.9
-20.2 |  |  |  | 9.2519 | 10.2667 |  |  |  | 9 | 9.95 |  |  |  | 11.65 | 12.65 |  |  |  | 18.8333 | 19.8533 |  |  |  | 32.7 | 33.5
-20.1 |  |  |  | 9.4019 | 10.4 |  |  |  | 9.125 | 10.1 |  |  |  | 11.9 | 12.9 |  |  |  | 19.1667 | 20.1784 |  |  |  | 34.7 | 35.75
-20 |  |  |  | 9.5519 | 10.6 |  |  |  | 9.24 | 10.2 |  |  |  | 12.1 | 13.125 |  |  |  | 19.5857 | 20.6034 |  |  |  | 37.6833 | 38.4333
-19.9 |  |  |  | 9.7019 | 10.7 |  |  |  | 9.35 | 10.425 |  |  |  | 12.3 | 13.26 |  |  |  | 19.9 | 21.0368 |  |  |  | 41.42 | 42.3857
-19.8 |  |  |  | 9.827 | 10.88 |  |  |  | 9.575 | 10.525 |  |  |  | 12.5 | 13.5 |  |  |  | 20.3286 | 21.4235 |  |  |  | 47.06 | 47.5857
-19.7 |  |  |  | 9.952 | 11.08 |  |  |  | 9.72 | 10.7 |  |  |  | 12.7 | 13.75 |  |  |  | 20.85 | 21.8608 |  |  |  | 54.1 | 54.6714
-19.6 |  |  |  | 10.202 | 11.2 |  |  |  | 9.85 | 10.9 |  |  |  | 13 | 14 |  |  |  | 21.2875 | 22.3752 |  |  |  | 62.7 | 62.7
-19.5 |  |  |  | 10.3021 | 11.35 |  |  |  | 10 | 11.02 |  |  |  | 13.225 | 14.2 |  |  |  | 21.8 | 22.9538 |  |  |  | 72.56 | 72.2286
-19.4 |  |  |  | 10.4354 | 11.5 |  |  |  | 10.18 | 11.2 |  |  |  | 13.45 | 14.42 |  |  |  | 22.4875 | 23.4896 |  |  |  | 82.08 | 81.875
-19.3 |  |  |  | 10.6271 | 11.7 |  |  |  | 10.3667 | 11.3 |  |  |  | 13.66 | 14.725 |  |  |  | 23.0857 | 24.1754 |  |  |  | 91.7667 | 90.825
-19.2 |  |  |  | 10.8622 | 11.875 |  |  |  | 10.5 | 11.5333 |  |  |  | 13.96 | 14.92 |  |  |  | 23.85 | 24.8899 |  |  |  | 99.5 | 98.4
-19.1 |  |  |  | 10.9222 | 12.1 |  |  |  | 10.725 | 11.7 |  |  |  | 14.2 | 15.2 |  |  |  | 24.7 | 25.7043 |  |  |  | 105.3 | 104.3
-19 |  |  |  | 11.1622 | 12.2 |  |  |  | 10.975 | 11.9 |  |  |  | 14.4667 | 15.45 |  |  |  | 25.8 | 26.8045 |  |  |  | 109.8 | 108.7429
-18.9 |  |  |  | 11.4223 | 12.35 |  |  |  | 11.1 | 12.1333 |  |  |  | 14.775 | 15.76 |  |  |  | 27.0571 | 28.0618 |  |  |  | 112.82 | 111.7333
-18.8 |  |  |  | 11.5623 | 12.625 |  |  |  | 11.35 | 12.275 |  |  |  | 15.05 | 16.075 |  |  |  | 28.4 | 29.4549 |  |  |  | 114.8167 | 114.1
-18.7 |  |  |  | 11.7023 | 12.8 |  |  |  | 11.6 | 12.55 |  |  |  | 15.425 | 16.4 |  |  |  | 30.3833 | 31.3052 |  |  |  | 116.8 | 115.7
-18.6 |  |  |  | 11.9524 | 13.0333 |  |  |  | 11.775 | 12.775 |  |  |  | 15.775 | 16.7 |  |  |  | 32.55 | 33.5484 |  |  |  | 118 | 116.9
-18.5 |  |  |  | 12.2024 | 13.2 |  |  |  | 11.98 | 13 |  |  |  | 16 | 17.05 |  |  |  | 35.6714 | 36.5489 |  |  |  | 119.14 | 118.1
-18.4 |  |  |  | 12.3525 | 13.425 |  |  |  | 12.2 | 13.18 |  |  |  | 16.5333 | 17.45 |  |  |  | 40.0286 | 40.8568 |  |  |  | 120.05 | 118.95
-18.3 |  |  |  | 12.6025 | 13.675 |  |  |  | 12.44 | 13.4333 |  |  |  | 16.9 | 17.85 |  |  |  | 45.92 | 46.8578 |  |  |  | 120.8 | 119.65
-18.2 |  |  |  | 12.7525 | 13.9 |  |  |  | 12.66 | 13.725 |  |  |  | 17.35 | 18.3 |  |  |  | 53.7 | 54.6948 |  |  |  | 121.4 | 120.2333
-18.1 |  |  |  | 13.0026 | 14.1 |  |  |  | 13 | 13.975 |  |  |  | 17.8 | 18.85 |  |  |  | 64.425 | 65.0108 |  |  |  | 121.8 | 120.7333
-18 |  |  |  | 13.2359 | 14.3 |  |  |  | 13.275 | 14.225 |  |  |  | 18.3 | 19.35 |  |  |  | 74.9667 | 75.584 |  |  |  | 122.3 | 121.0667
-17.9 |  |  |  | 13.5277 | 14.55 |  |  |  | 13.5 | 14.52 |  |  |  | 18.95 | 19.8667 |  |  |  | 86 | 86.3144 |  |  | 10.55 | 122.58 | 121.3714
-17.8 |  |  |  | 13.7628 | 14.875 |  |  |  | 13.78 | 14.8333 |  |  |  | 19.6 | 20.6 |  |  |  | 96.7714 | 96.5037 |  |  | 10.66 | 122.86 | 121.7333
-17.7 |  |  |  | 14.0028 | 15.125 |  |  |  | 14.22 | 15.2 |  |  |  | 20.38 | 21.38 |  |  |  | 105.76 | 105.2463 |  |  | 10.8 | 123.06 | 121.9833
-17.6 |  |  |  | 14.2829 | 15.34 |  |  |  | 14.48 | 15.46 |  |  |  | 21.35 | 22.2 |  |  |  | 112.0857 | 111.7857 |  |  | 10.9333 | 123.3571 | 122.1667
-17.5 |  |  |  | 14.6529 | 15.65 |  |  |  | 14.9 | 15.8333 |  |  |  | 22.4 | 23.25 |  |  |  | 116.55 | 115.869 |  |  | 11.1 | 123.525 | 122.3857
-17.4 |  |  |  | 14.928 | 15.925 |  |  |  | 15.26 | 16.25 |  |  |  | 23.7 | 24.6 |  |  |  | 119.55 | 118.86 |  |  | 11.3 | 123.76 | 122.575
-17.3 |  |  |  | 15.163 | 16.3 |  |  |  | 15.78 | 16.75 |  |  |  | 25.24 | 26 |  |  |  | 121.8 | 120.9629 |  |  | 11.4667 | 123.9 | 122.7429
-17.2 |  |  |  | 15.5231 | 16.625 |  |  |  | 16.2 | 17.25 |  |  |  | 27.1667 | 28.0667 |  |  |  | 123.4333 | 122.7493 |  |  | 11.6667 | 123.92 | 122.8
-17.1 |  |  |  | 15.8782 | 16.94 |  |  |  | 16.78 | 17.8 |  |  |  | 29.6 | 30.45 |  |  |  | 124.6833 | 123.861 |  |  | 11.8 | 124.1 | 122.9571
-17 |  |  |  | 16.3366 | 17.4 |  |  |  | 17.475 | 18.5 |  |  |  | 33.4 | 34 |  |  |  | 125.75 | 124.921 |  |  | 12.0333 | 124.1833 | 123.1
-16.9 |  |  |  | 16.8034 | 17.88 |  |  |  | 18.225 | 19.275 |  |  |  | 38.4 | 38.74 |  |  | 8.66 | 126.525 | 125.7377 |  |  | 12.225 | 124.2 | 123.2
-16.8 |  |  |  | 17.2035 | 18.28 |  |  |  | 19.2 | 20.2 |  |  |  | 45.45 | 46.2 |  |  | 8.78 | 127.2333 | 126.301 |  |  | 12.35 | 124.36 | 123.2
-16.7 |  |  |  | 17.8636 | 18.875 |  |  |  | 20.3 | 21.3 |  |  |  | 56.14 | 56.15 |  |  | 8.975 | 127.7 | 126.8067 |  |  | 12.6167 | 124.42 | 123.3
-16.6 |  |  |  | 18.6037 | 19.625 |  |  |  | 21.7667 | 22.74 |  |  |  | 67.8 | 68.5 |  |  | 9.1 | 128.15 | 127.271 |  |  | 12.8667 | 124.5 | 123.3625
-16.5 |  |  |  | 19.4789 | 20.475 |  |  |  | 23.7 | 24.5 |  |  |  | 80.8 | 80.86 |  |  | 9.2 | 128.6 | 127.571 |  |  | 13.08 | 124.6 | 123.4
-16.4 |  |  |  | 20.5041 | 21.5 |  |  |  | 25.8 | 26.7 |  |  |  | 93.5 | 93.3 |  |  | 9.4 | 128.7571 | 127.7781 |  |  | 13.28 | 124.65 | 123.5
-16.3 |  |  |  | 21.7543 | 22.775 |  |  |  | 29.18 | 29.92 |  |  |  | 105.32 | 105.1 |  |  | 9.54 | 129 | 128.0353 |  |  | 13.52 | 124.7 | 123.6
-16.2 |  |  |  | 23.5047 | 24.5 |  |  |  | 33.55 | 34.175 |  |  |  | 115.2 | 115.1 |  |  | 9.6667 | 129.1667 | 128.171 |  |  | 13.68 | 124.7 | 123.6
-16.1 |  |  |  | 25.8052 | 26.6 |  |  |  | 40.2 | 41.2 |  |  |  | 122.36 | 121.12 |  |  | 9.875 | 129.2333 | 128.321 |  |  | 14 | 124.7833 | 123.625
-16 |  |  |  | 28.9058 | 29.7 |  |  |  | 50.3 | 50.2 |  |  |  | 126.48 | 125.65 |  |  | 10.08 | 129.3 | 128.4924 |  |  | 14.3167 | 124.7833 | 123.7
-15.9 |  |  |  | 33.7067 | 34.25 |  |  |  | 62.5 | 63.05 |  |  | 6.675 | 129.54 | 128.7 |  |  | 10.2167 | 129.4 | 128.521 |  |  | 14.5 | 124.8 | 123.7
-15.8 |  |  |  | 40.4581 | 40.96 |  |  |  | 76.4 | 76.4 |  |  | 6.725 | 131.42 | 130.55 |  |  | 10.3833 | 129.5 | 128.6043 |  |  | 14.75 | 124.8 | 123.7
-15.7 |  |  |  | 50.7852 | 50.9 |  |  |  | 91.05 | 91.4 |  |  | 6.85 | 133.05 | 132.05 |  |  | 10.58 | 129.5 | 128.621 |  |  | 15.05 | 124.8 | 123.7
-15.6 |  |  |  | 64.0378 | 64.1 |  |  |  | 104.125 | 104.42 |  |  | 6.98 | 134.2 | 133.1 |  |  | 10.825 | 129.6 | 128.661 |  |  | 15.32 | 124.8 | 123.7
-15.5 |  |  |  | 78.2957 | 78.2 |  |  |  | 116.35 | 115.45 |  |  | 7.06 | 135.02 | 133.98 |  |  | 11 | 129.6 | 128.721 |  |  | 15.5 | 124.725 | 123.7
-15.4 |  |  |  | 92.5185 | 93.15 |  |  | 5.775 | 124.725 | 123.94 |  |  | 7.1833 | 135.6 | 134.55 |  |  | 11.125 | 129.6 | 128.721 |  |  | 15.9 | 124.8 | 123.7
-15.3 |  |  |  | 106.921 | 106.025 |  |  | 5.875 | 130.5 | 129.58 |  |  | 7.26 | 136.05 | 134.95 |  |  | 11.35 | 129.6 | 128.721 |  |  | 16.15 | 124.7 | 123.64
-15.2 |  |  | 6.4 | 118.7738 | 117.7 |  |  | 6 | 133.875 | 132.9333 |  |  | 7.4 | 136.45 | 135.42 |  |  | 11.56 | 129.6 | 128.721 |  |  | 16.48 | 124.64 | 123.5875
-15.1 |  |  | 6.425 | 127.7752 | 126.05 |  |  | 6.1 | 136.4 | 135.1 |  |  | 7.55 | 136.6 | 135.55 |  |  | 11.82 | 129.6 | 128.721 |  |  | 16.8 | 124.6 | 123.5
-15 |  |  | 6.575 | 132.827 | 131.62 |  |  | 6.15 | 137.9 | 136.9 |  |  | 7.7 | 136.7667 | 135.7 |  |  | 12.05 | 129.6 | 128.721 |  |  | 17.12 | 124.6 | 123.5
-14.9 |  |  | 6.7 | 136.027 | 134.72 |  |  | 6.25 | 139.1 | 138 |  |  | 7.82 | 136.8 | 135.8 |  |  | 12.275 | 129.6 | 128.721 |  |  | 17.54 | 124.52 | 123.5
-14.8 |  |  | 6.775 | 137.9676 | 136.64 |  |  | 6.3667 | 139.675 | 138.74 |  |  | 8 | 136.9 | 135.9 |  |  | 12.52 | 129.5375 | 128.6639 |  |  | 18 | 124.5 | 123.45
-14.7 |  |  | 6.9 | 139.428 | 138 |  |  | 6.4333 | 140.4 | 139.3 |  |  | 8.14 | 136.9 | 135.9 |  |  | 12.74 | 129.4833 | 128.621 |  |  | 18.3333 | 124.4 | 123.4
-14.6 |  |  | 7 | 140.228 | 138.94 |  |  | 6.54 | 140.78 | 139.625 |  |  | 8.3 | 136.9 | 135.9 |  |  | 13.06 | 129.4 | 128.5496 |  |  | 18.8 | 124.3 | 123.4
-14.5 |  |  | 7.12 | 140.928 | 139.6 |  |  | 6.7 | 140.98 | 139.8667 |  |  | 8.46 | 136.9 | 135.825 |  |  | 13.35 | 129.3429 | 128.4877 |  |  | 19.34 | 124.3 | 123.4
-14.4 |  |  | 7.2 | 141.228 | 139.925 |  |  | 6.8 | 141.1 | 140 |  |  | 8.68 | 136.8 | 135.8 |  |  | 13.58 | 129.1875 | 128.3353 |  |  | 19.94 | 124.3 | 123.3
-14.3 |  |  | 7.325 | 141.428 | 140.2 |  |  | 6.9 | 141.125 | 140.1 |  |  | 8.7833 | 136.7 | 135.725 |  |  | 13.9 | 129.1 | 128.221 |  |  | 20.5 | 124.2 | 123.2571
-14.2 |  |  | 7.475 | 141.528 | 140.24 |  |  | 7.06 | 141.2 | 140.1 |  |  | 9 | 136.6 | 135.65 |  |  | 14.1667 | 129 | 128.1496 |  |  | 21.2 | 124.1 | 123.2
-14.1 |  |  | 7.58 | 141.528 | 140.3 |  |  | 7.2 | 141.1 | 140 |  |  | 9.16 | 136.5 | 135.6 |  |  | 14.525 | 128.9 | 128.081 |  |  | 21.9833 | 124.1 | 123.1333
-14 |  |  | 7.7333 | 141.428 | 140.2 |  |  | 7.3 | 141 | 139.95 |  |  | 9.375 | 136.375 | 135.475 |  |  | 14.9167 | 128.8 | 127.9924 |  |  | 22.9 | 123.98 | 122.9833
-13.9 |  |  | 7.84 | 141.328 | 140.025 |  |  | 7.42 | 140.9 | 139.8 |  |  | 9.62 | 136.225 | 135.3 |  |  | 15.25 | 128.75 | 127.921 |  |  | 24 | 123.9 | 122.9
-13.8 |  |  | 8 | 141.088 | 139.9 |  |  | 7.54 | 140.7 | 139.64 |  |  | 9.75 | 136.1 | 135.18 |  |  | 15.55 | 128.6143 | 127.821 |  |  | 25.4 | 123.78 | 122.8
-13.7 |  |  | 8.2 | 140.968 | 139.76 |  |  | 7.75 | 140.5 | 139.5 |  |  | 10 | 135.9 | 134.95 |  |  | 16.04 | 128.5 | 127.721 |  |  | 27.1 | 123.6 | 122.7167
-13.6 |  |  | 8.3 | 140.668 | 139.5 |  |  | 7.9 | 140.3 | 139.24 |  |  | 10.225 | 135.8 | 134.8 |  |  | 16.475 | 128.3667 | 127.6639 |  |  | 28.9 | 123.5429 | 122.6
-13.5 |  |  | 8.4 | 140.528 | 139.22 |  |  | 8.02 | 140.1 | 139 |  |  | 10.4 | 135.7 | 134.6333 |  |  | 17.05 | 128.2 | 127.501 |  |  | 31.7 | 123.4 | 122.5
-13.4 |  |  | 8.7 | 140.1613 | 138.925 |  |  | 8.24 | 139.9333 | 138.8 |  |  | 10.65 | 135.475 | 134.4 |  |  | 17.64 | 128.04 | 127.3353 |  |  | 35.4 | 123.3167 | 122.4333
-13.3 |  |  | 8.825 | 139.953 | 138.7 |  |  | 8.4 | 139.65 | 138.525 |  |  | 10.95 | 135.25 | 134.3 |  |  | 18.18 | 127.9571 | 127.171 |  |  | 40.2333 | 123.18 | 122.35
-13.2 |  |  | 8.95 | 139.708 | 138.44 |  |  | 8.6 | 139.325 | 138.3 |  |  | 11.2333 | 134.96 | 134.04 |  |  | 18.88 | 127.825 | 127.121 |  |  | 46.9 | 123 | 122.0857
-13.1 |  |  | 9.175 | 139.388 | 138.18 |  |  | 8.725 | 139.12 | 138 |  |  | 11.5667 | 134.825 | 133.7 |  |  | 19.775 | 127.6833 | 126.9639 |  |  | 54.8 | 122.8333 | 121.9333
-13 |  |  | 9.3 | 139.1613 | 137.85 |  |  | 8.975 | 138.85 | 137.78 |  |  | 11.8 | 134.6 | 133.56 |  |  | 20.7 | 127.5286 | 126.821 |  |  | 64.6 | 122.66 | 121.7714
-12.9 |  |  | 9.54 | 138.828 | 137.62 |  |  | 9.125 | 138.525 | 137.5667 |  |  | 12.1667 | 134.34 | 133.4 |  |  | 21.8333 | 127.4 | 126.6924 |  |  | 73.4 | 122.4833 | 121.55
-12.8 |  |  | 9.7 | 138.468 | 137.375 |  |  | 9.45 | 138.3 | 137.25 |  |  | 12.5 | 134.1 | 133.2 |  |  | 23.26 | 127.2714 | 126.5377 |  |  | 82.8 | 122.3286 | 121.4833
-12.7 |  |  | 9.95 | 138.1613 | 137.0333 |  |  | 9.62 | 137.95 | 136.9 |  |  | 12.88 | 133.875 | 132.9 |  |  | 24.8857 | 127.1 | 126.421 |  |  | 89.7 | 122.3 | 121.3
-12.6 |  |  | 10.2 | 137.888 | 136.7 |  |  | 9.84 | 137.72 | 136.7 |  |  | 13.3 | 133.64 | 132.78 |  |  | 27.25 | 126.98 | 126.2877 |  |  | 95.8 | 122.08 | 121.05
-12.5 |  |  | 10.36 | 137.5613 | 136.475 |  |  | 10.1667 | 137.3667 | 136.28 |  |  | 13.84 | 133.35 | 132.4833 |  |  | 30.04 | 126.7 | 126.021 |  |  | 99.9 | 121.82 | 120.9143
-12.4 |  |  | 10.625 | 137.152 | 136.22 |  |  | 10.525 | 137.1 | 136.04 |  |  | 14.4 | 133.1 | 132.25 |  |  | 34.08 | 126.55 | 125.821 |  |  | 103 | 121.625 | 120.75
-12.3 |  |  | 10.94 | 136.902 | 135.85 |  |  | 10.8 | 136.675 | 135.825 |  |  | 14.9667 | 132.9 | 131.925 |  |  | 39.7333 | 126.36 | 125.721 |  |  | 104.96 | 121.4 | 120.6
-12.2 |  |  | 11.1 | 136.607 | 135.45 |  |  | 11.125 | 136.35 | 135.4 |  |  | 15.68 | 132.7 | 131.7 |  |  | 47 | 126.1 | 125.4496 |  |  | 106.78 | 121.3 | 120.5
-12.1 |  |  | 11.45 | 136.227 | 135.15 |  |  | 11.5 | 136 | 135.075 |  |  | 16.5 | 132.4 | 131.45 |  |  | 56.84 | 125.9 | 125.196 |  |  | 108.0333 | 121.1333 | 120.4286
-12 |  |  | 11.7 | 135.902 | 134.7 |  |  | 11.95 | 135.725 | 134.7667 |  |  | 17.45 | 132.16 | 131.24 |  |  | 66.6 | 125.6286 | 124.9353 |  |  | 109 | 121 | 120.1833
-11.9 |  |  | 12.08 | 135.577 | 134.38 |  |  | 12.5 | 135.52 | 134.5 |  |  | 18.675 | 131.9333 | 130.94 |  |  | 77.5 | 125.3571 | 124.6496 |  |  | 109.9 | 120.8286 | 120.0375
-11.8 |  |  | 12.45 | 135.127 | 134.15 |  |  | 13.05 | 135.1 | 134.16 |  |  | 20.14 | 131.7 | 130.58 |  |  | 86.425 | 125.1286 | 124.4877 |  |  | 110.6 | 120.7 | 119.9125
-11.7 |  |  | 12.9 | 134.777 | 133.8 |  |  | 13.7 | 134.75 | 133.85 |  |  | 21.8 | 131.275 | 130.45 |  |  | 94.25 | 124.8429 | 124.196 |  |  | 111.075 | 120.58 | 119.7714
-11.6 |  |  | 13.375 | 134.377 | 133.38 |  |  | 14.4 | 134.4 | 133.5 |  |  | 24.075 | 130.925 | 130.12 |  |  | 100.3 | 124.62 | 123.9639 |  |  | 111.62 | 120.5 | 119.7
-11.5 |  |  | 13.95 | 134.177 | 133.1 |  |  | 15.4 | 134.08 | 133.12 |  |  | 27.15 | 130.8 | 129.85 |  |  | 104.5 | 124.4 | 123.821 |  |  | 112.0167 | 120.4 | 119.6
-11.4 |  |  | 14.7 | 133.707 | 132.82 |  |  | 16.6 | 133.8 | 132.825 |  |  | 31.325 | 130.4 | 129.58 |  |  | 107.15 | 124.2143 | 123.641 |  | 5.6 | 112.1833 | 120.2 | 119.475
-11.3 |  |  | 15.575 | 133.452 | 132.52 |  |  | 18 | 133.425 | 132.5 |  |  | 37.85 | 130.08 | 129.2 |  |  | 109.3 | 124 | 123.2877 |  | 5.725 | 112.5667 | 120 | 119.3429
-11.2 |  |  | 16.54 | 133.077 | 132.2 |  |  | 19.7333 | 133.05 | 132.25 |  |  | 46.8 | 129.82 | 128.9333 |  |  | 110.66 | 123.72 | 123.0875 |  | 5.9 | 112.8 | 119.925 | 119.2
-11.1 |  |  | 17.85 | 132.727 | 131.8 |  |  | 22.2333 | 132.75 | 131.92 |  |  | 57.5 | 129.44 | 128.6 |  |  | 112 | 123.5667 | 122.8367 |  | 5.975 | 112.925 | 119.85 | 119
-11 |  |  | 19.6 | 132.326 | 131.425 |  |  | 25.3 | 132.425 | 131.5 |  |  | 68.55 | 129.04 | 128.25 |  |  | 112.94 | 123.3286 | 122.6867 |  | 6.1 | 113.25 | 119.6833 | 118.95
-10.9 |  |  | 21.7 | 132.026 | 131.1 |  |  | 29.9 | 132 | 131.05 |  |  | 80.4 | 128.725 | 127.85 |  | 4.4 | 113.68 | 123.1143 | 122.5343 |  | 6.2 | 113.325 | 119.6 | 118.8
-10.8 |  |  | 24.8 | 131.726 | 130.775 |  |  | 36.9 | 131.5667 | 130.8 |  |  | 90.98 | 128.28 | 127.54 |  | 4.5 | 114.275 | 122.9 | 122.3867 |  | 6.2 | 113.5 | 119.48 | 118.6429
-10.7 |  |  | 29.65 | 131.326 | 130.35 |  |  | 46.6 | 131.16 | 130.42 |  |  | 99.6 | 128 | 127.2 |  | 4.6 | 114.7 | 122.7625 | 122.12 |  | 6.38 | 113.6 | 119.3 | 118.5
-10.6 |  |  | 36.6 | 130.826 | 130 |  |  | 57.9 | 130.76 | 130.08 |  |  | 105.9 | 127.6 | 126.88 |  | 4.7 | 115.2 | 122.6286 | 121.9057 |  | 6.5 | 113.7 | 119.0667 | 118.425
-10.5 |  |  | 46.4 | 130.4927 | 129.64 |  |  | 71 | 130.3667 | 129.64 |  |  | 110.125 | 127.25 | 126.62 |  | 4.75 | 115.625 | 122.4714 | 121.72 |  | 6.7 | 113.7 | 118.9 | 118.2833
-10.4 |  |  | 58.85 | 130.0927 | 129.325 |  |  | 82.9 | 130 | 129.24 |  | 3.2 | 112.9 | 126.9 | 126.26 |  | 4.8667 | 115.96 | 122.1857 | 121.5825 |  | 6.8 | 113.7 | 118.7 | 118.2
-10.3 |  |  | 71.3 | 129.666 | 128.74 |  |  | 94.65 | 129.64 | 128.8 |  | 3.3 | 114.75 | 126.675 | 126 |  | 5 | 116.08 | 122 | 121.3771 |  | 6.9 | 113.7 | 118.65 | 118.0333
-10.2 |  |  | 84.05 | 129.126 | 128.35 |  | 2.8 | 103.4 | 129.2 | 128.4667 |  | 3.3 | 116.3 | 126.425 | 125.7 |  | 5.1 | 116.22 | 121.8625 | 121.2629 |  | 7.075 | 113.7 | 118.5167 | 117.8
-10.1 |  | 3 | 95.7 | 128.776 | 128.025 |  | 2.8 | 109.625 | 128.76 | 128.04 |  | 3.4 | 117.4 | 126.1 | 125.46 |  | 5.125 | 116.4 | 121.6571 | 121.12 |  | 7.325 | 113.7 | 118.35 | 117.6429
-10 |  | 3.1 | 104.9 | 128.3593 | 127.8 |  | 2.8 | 113.8 | 128.45 | 127.62 |  | 3.4 | 118.34 | 125.85 | 125.175 |  | 5.3 | 116.475 | 121.46 | 120.84 |  | 7.475 | 113.7 | 118.1667 | 117.48
-9.9 |  | 3.1 | 111.475 | 128.051 | 127.25 |  | 2.8 | 116.35 | 128.1 | 127.26 |  | 3.5 | 118.96 | 125.64 | 124.8 |  | 5.4 | 116.5833 | 121.2667 | 120.7057 |  | 7.6 | 113.7 | 118 | 117.3
-9.8 |  | 3.2 | 115.3 | 127.606 | 126.88 |  | 2.9 | 118.08 | 127.65 | 126.94 |  | 3.6 | 119.45 | 125.32 | 124.66 |  | 5.6 | 116.6 | 121.1 | 120.52 |  | 7.7 | 113.7 | 117.7 | 117.1625
-9.7 |  | 3.2 | 117.7 | 127.275 | 126.48 |  | 2.9 | 119.6 | 127.3 | 126.575 |  | 3.7 | 119.85 | 125.075 | 124.42 |  | 5.7 | 116.6 | 120.8857 | 120.3533 |  | 7.9 | 113.6 | 117.6 | 116.9333
-9.6 |  | 3.3 | 119.5 | 126.945 | 126.2 |  | 3 | 120.6 | 127.02 | 126.34 |  | 3.7 | 120.125 | 124.9 | 124.1667 |  | 5.8 | 116.6 | 120.6667 | 120.1867 |  | 8.1333 | 113.6 | 117.42 | 116.7714
-9.5 |  | 3.3 | 120.7 | 126.645 | 125.85 |  | 3.1 | 121.45 | 126.75 | 126.0667 |  | 3.8 | 120.375 | 124.6 | 123.95 |  | 6 | 116.6 | 120.4 | 119.97 |  | 8.35 | 113.5 | 117.2167 | 116.6
-9.4 |  | 3.4 | 121.6 | 126.3 | 125.55 |  | 3.1 | 121.98 | 126.4 | 125.7 |  | 3.9 | 120.6 | 124.36 | 123.75 |  | 6.075 | 116.6 | 120.1429 | 119.745 |  | 8.425 | 113.5 | 117.2 | 116.4429
-9.3 |  | 3.4 | 122.15 | 126.05 | 125.32 |  | 3.2 | 122.3667 | 126.12 | 125.475 |  | 3.975 | 120.6 | 124.0333 | 123.4 |  | 6.2333 | 116.6 | 119.925 | 119.42 |  | 8.65 | 113.5 | 117 | 116.2833
-9.2 |  | 3.5 | 122.65 | 125.785 | 125.08 |  | 3.2 | 122.625 | 125.9 | 125.08 |  | 4.1 | 120.7 | 123.8667 | 123.18 |  | 6.375 | 116.5 | 119.6833 | 119.2343 |  | 8.9 | 113.5 | 116.8167 | 116.1286
-9.1 |  | 3.6 | 122.9 | 125.525 | 124.8 |  | 3.3 | 122.82 | 125.58 | 124.85 |  | 4.2 | 120.7 | 123.46 | 122.9 |  | 6.5667 | 116.4 | 119.5167 | 119.0057 |  | 9.05 | 113.4 | 116.58 | 115.9625
-9 |  | 3.6 | 123 | 125.125 | 124.48 |  | 3.4 | 122.9667 | 125.275 | 124.6 |  | 4.3 | 120.7 | 123.175 | 122.7 |  | 6.7 | 116.3 | 119.2714 | 118.7867 |  | 9.3 | 113.36 | 116.4 | 115.8375
-8.9 |  | 3.7 | 123.1 | 124.825 | 124.2 |  | 3.4 | 123 | 125.025 | 124.325 |  | 4.4 | 120.7 | 122.95 | 122.4 |  | 6.9 | 116.28 | 119.0333 | 118.595 |  | 9.575 | 113.3 | 116.3667 | 115.6
-8.8 |  | 3.8 | 123.1 | 124.505 | 123.9 |  | 3.5 | 123 | 124.65 | 124 |  | 4.5 | 120.55 | 122.65 | 122.125 |  | 7.0667 | 116.18 | 118.8 | 118.37 |  | 9.75 | 113.3 | 116.2 | 115.4429
-8.7 |  | 3.9 | 123.1 | 124.285 | 123.65 |  | 3.6 | 123 | 124.34 | 123.85 |  | 4.6 | 120.4 | 122.4 | 121.86 |  | 7.3 | 116.05 | 118.6 | 118.1867 |  | 10.05 | 113.2 | 116 | 115.3125
-8.6 |  | 3.9 | 123 | 124.075 | 123.32 |  | 3.7 | 122.9 | 124.04 | 123.4667 |  | 4.7 | 120.375 | 122.14 | 121.55 |  | 7.5 | 115.9 | 118.46 | 117.9367 |  | 10.4 | 113.1 | 115.7333 | 115.1714
-8.5 |  | 4.1 | 123 | 123.725 | 123 |  | 3.7333 | 122.7 | 123.75 | 123.2 |  | 4.875 | 120.25 | 121.85 | 121.3 |  | 7.7 | 115.8 | 118.2143 | 117.77 |  | 10.6 | 112.98 | 115.525 | 114.9429
-8.4 |  | 4.125 | 122.775 | 123.325 | 122.675 |  | 3.8667 | 122.575 | 123.45 | 122.8 |  | 5 | 120.02 | 121.52 | 121.025 |  | 7.9 | 115.75 | 118.0429 | 117.58 |  | 10.8667 | 112.9 | 115.3667 | 114.7833
-8.3 |  | 4.2667 | 122.625 | 123.025 | 122.4 |  | 4 | 122.35 | 123.15 | 122.525 |  | 5.2 | 119.9 | 121.3 | 120.675 |  | 8.1333 | 115.6 | 117.7571 | 117.3033 |  | 11.225 | 112.8 | 115.2 | 114.6286
-8.2 |  | 4.3 | 122.36 | 122.7 | 122.1 |  | 4.1 | 122.2 | 122.88 | 122.2667 |  | 5.3 | 119.75 | 120.96 | 120.38 |  | 8.4 | 115.46 | 117.55 | 117.1575 |  | 11.6333 | 112.7333 | 115.1 | 114.45
-8.1 |  | 4.4 | 122.2 | 122.4745 | 121.825 |  | 4.2 | 121.94 | 122.6 | 121.925 |  | 5.4333 | 119.62 | 120.6667 | 120.14 |  | 8.65 | 115.3 | 117.3 | 117.02 |  | 12.1 | 112.65 | 114.9 | 114.3125
-8 |  | 4.5 | 121.925 | 122.099 | 121.5 |  | 4.3 | 121.7 | 122.12 | 121.56 |  | 5.6 | 119.32 | 120.5 | 119.9333 |  | 8.95 | 115.2 | 117 | 116.7523 |  | 12.5667 | 112.52 | 114.8 | 114.05
-7.9 |  | 4.625 | 121.6 | 121.849 | 121.225 |  | 4.5 | 121.575 | 121.86 | 121.3333 |  | 5.725 | 119.075 | 120.14 | 119.58 |  | 9.2667 | 115.1 | 116.8167 | 116.494 |  | 13 | 112.4833 | 114.6167 | 113.9857
-7.8 |  | 4.7333 | 121.3833 | 121.524 | 120.9 |  | 4.6 | 121.15 | 121.54 | 120.975 |  | 5.95 | 118.84 | 119.94 | 119.3 |  | 9.6 | 114.92 | 116.6625 | 116.3047 |  | 13.5333 | 112.3167 | 114.34 | 113.8429
-7.7 |  | 4.9 | 121.18 | 121.024 | 120.58 |  | 4.7 | 120.9 | 121.3 | 120.6 |  | 6.2 | 118.7 | 119.6 | 119 |  | 9.875 | 114.82 | 116.4571 | 116.0023 |  | 14.2 | 112.225 | 114.125 | 113.6
-7.6 |  | 5.0333 | 121 | 120.774 | 120.325 |  | 4.825 | 120.75 | 120.85 | 120.4 |  | 6.4 | 118.5 | 119.3 | 118.8 |  | 10.3333 | 114.66 | 116.2667 | 115.819 | 1.5 | 14.9 | 112.1 | 114 | 113.4667
-7.5 |  | 5.2 | 120.65 | 120.424 | 119.975 |  | 5.0667 | 120.5 | 120.475 | 120.05 |  | 6.6 | 118.32 | 119 | 118.525 |  | 10.7333 | 114.5 | 116.0286 | 115.6904 | 1.9 | 15.8 | 112 | 113.8 | 113.3714
-7.4 |  | 5.35 | 120.4 | 120.084 | 119.72 |  | 5.1 | 120.25 | 120.15 | 119.625 |  | 6.9 | 118.075 | 118.7 | 118.2 |  | 11.2 | 114.325 | 115.9 | 115.4523 | 1.8 | 16.875 | 111.9143 | 113.7 | 113.1429
-7.3 |  | 5.5 | 120.075 | 119.874 | 119.4 |  | 5.4 | 119.95 | 119.95 | 119.375 |  | 7.1667 | 117.86 | 118.5 | 117.9 | 1.4 | 11.8 | 114.2 | 115.65 | 115.2023 | 1.9 | 18.1333 | 111.8 | 113.4833 | 113.0125
-7.2 |  | 5.6333 | 119.78 | 119.584 | 119.06 |  | 5.625 | 119.7 | 119.65 | 119.08 | 0.8 | 7.55 | 117.65 | 118.16 | 117.65 | 1.4 | 12.4 | 114.1 | 115.48 | 115.0357 | 1.9 | 19.6333 | 111.7 | 113.275 | 112.84
-7.1 |  | 5.8 | 119.55 | 119.2907 | 118.75 | 0.8 | 5.8 | 119.375 | 119.24 | 118.72 | 1 | 7.9 | 117.46 | 117.86 | 117.4 | 1.5 | 13.2 | 114 | 115.2 | 114.859 | 2 | 21.65 | 111.6 | 113.1333 | 112.65
-7 | 0.9 | 6.0667 | 119.25 | 119.024 | 118.42 | 0.8 | 6.1 | 119.12 | 119 | 118.5 | 1 | 8.2 | 117.18 | 117.625 | 117.175 | 1.5 | 14.1 | 113.8 | 114.92 | 114.6619 | 2.1 | 24.1 | 111.5 | 113 | 112.5
-6.9 | 0.9 | 6.2 | 118.94 | 118.649 | 118.2167 | 0.8 | 6.4 | 118.8 | 118.725 | 118.15 | 1 | 8.6667 | 116.88 | 117.275 | 117.0333 | 1.5 | 15.0333 | 113.6 | 114.7429 | 114.4357 | 2.1 | 27.4 | 111.4333 | 112.8 | 112.4333
-6.8 | 0.9 | 6.575 | 118.7 | 118.374 | 117.95 | 0.9 | 6.6 | 118.52 | 118.4 | 118 | 1 | 9.2333 | 116.65 | 116.95 | 116.7 | 1.6 | 16.3 | 113.525 | 114.525 | 114.259 | 2.1 | 32.4 | 111.4 | 112.8 | 112.2571
-6.7 | 1 | 6.875 | 118.4 | 117.999 | 117.5 | 0.9 | 7.1 | 118.28 | 118.125 | 117.72 | 1.1 | 9.8 | 116.425 | 116.8 | 116.4 | 1.6 | 17.9 | 113.3333 | 114.3667 | 114.059 | 2.2 | 38.85 | 111.35 | 112.64 | 112.1
-6.6 | 1 | 7.1333 | 118.14 | 117.764 | 117.3 | 0.9 | 7.5333 | 117.95 | 117.8333 | 117.275 | 1.1 | 10.6 | 116.22 | 116.56 | 116.18 | 1.7 | 20 | 113.2 | 114.2143 | 113.8619 | 2.2 | 46.38 | 111.3 | 112.5 | 112.0167
-6.5 | 1 | 7.5 | 117.875 | 117.4903 | 117 | 0.9 | 8.05 | 117.6 | 117.5 | 117.05 | 1.2 | 11.4 | 116 | 116.2667 | 115.9 | 1.7333 | 22.5 | 113.1 | 113.9333 | 113.719 | 2.3 | 56 | 111.26 | 112.325 | 111.8857
-6.4 | 1 | 7.9667 | 117.55 | 117.173 | 116.75 | 0.9 | 8.65 | 117.375 | 117.25 | 116.8 | 1.2 | 12.4333 | 115.85 | 116.06 | 115.66 | 1.8 | 26.2 | 113 | 113.7667 | 113.5523 | 2.4 | 65.7 | 111.1 | 112.25 | 111.7125
-6.3 | 1 | 8.5333 | 117.32 | 116.873 | 116.4 | 0.9 | 9.4667 | 117.16 | 116.95 | 116.475 | 1.2 | 13.65 | 115.625 | 115.75 | 115.35 | 1.9 | 31.25 | 112.84 | 113.5 | 113.269 | 2.5 | 74 | 110.925 | 112.06 | 111.6
-6.2 | 1.0333 | 9.0667 | 116.98 | 116.573 | 116.15 | 1 | 10.4 | 116.92 | 116.625 | 116.26 | 1.2 | 15.3 | 115.3 | 115.45 | 115.075 | 1.9 | 38.3667 | 112.7 | 113.3333 | 113.019 | 2.5 | 82.15 | 110.8 | 111.86 | 111.4
-6.1 | 1.1 | 9.9 | 116.65 | 116.443 | 115.8333 | 1 | 11.55 | 116.675 | 116.3 | 115.98 | 1.2 | 17.2667 | 115.0667 | 115.25 | 114.825 | 2 | 47.4667 | 112.5333 | 113.225 | 112.9523 | 2.7 | 88.3 | 110.75 | 111.6667 | 111.25
-6 | 1.1 | 10.9 | 116.45 | 115.998 | 115.6 | 1 | 12.85 | 116.32 | 116.08 | 115.7 | 1.3 | 20 | 114.9 | 115.02 | 114.54 | 2.1 | 57.325 | 112.4 | 113.0625 | 112.8047 | 2.8 | 92.5 | 110.6 | 111.5 | 111.05
-5.9 | 1.1 | 12.1667 | 116.15 | 115.623 | 115.35 | 1 | 14.7 | 116.05 | 115.8 | 115.4 | 1.3 | 23.8 | 114.725 | 114.875 | 114.4 | 2.2 | 67.35 | 112.2 | 112.9286 | 112.6523 | 2.8 | 95.375 | 110.5 | 111.32 | 110.9
-5.8 | 1.1 | 13.9667 | 115.86 | 115.473 | 115 | 1.1 | 17.0667 | 115.8 | 115.5 | 115.1 | 1.4 | 29.1 | 114.54 | 114.5 | 114.2 | 2.2 | 76.6667 | 112.08 | 112.6714 | 112.3857 | 2.85 | 97.975 | 110.35 | 111.2 | 110.775
-5.7 | 1.2 | 16.3 | 115.65 | 115.223 | 114.8 | 1.1 | 20.4667 | 115.58 | 115.28 | 114.96 | 1.4 | 37.5333 | 114.3667 | 114.28 | 114 | 2.3 | 84.3 | 111.95 | 112.5375 | 112.219 | 2.9 | 99.4 | 110.28 | 111.05 | 110.6833
-5.6 | 1.2 | 19.4 | 115.475 | 114.863 | 114.6 | 1.1 | 25.2333 | 115.3333 | 115 | 114.725 | 1.5 | 47 | 114.14 | 114.075 | 113.775 | 2.3 | 90.5 | 111.8 | 112.28 | 112.059 | 3 | 100.6667 | 110.2 | 110.88 | 110.5
-5.5 | 1.2 | 24.2333 | 115.225 | 114.723 | 114.26 | 1.1 | 33 | 115.075 | 114.8333 | 114.375 | 1.5 | 57.0333 | 113.9 | 113.875 | 113.6 | 2.4 | 94.7 | 111.7333 | 112.2 | 111.919 | 3.1 | 101.8 | 110.1 | 110.8 | 110.3375
-5.4 | 1.3 | 31.5 | 114.95 | 114.498 | 114.1 | 1.2 | 42.1 | 114.9 | 114.525 | 114.2 | 1.6 | 68.5667 | 113.8 | 113.7 | 113.4 | 2.4 | 97.55 | 111.46 | 112 | 111.8023 | 3.2 | 102.55 | 109.95 | 110.6 | 110.1833
-5.3 | 1.3 | 42.3333 | 114.72 | 114.2897 | 113.9 | 1.2 | 53.3333 | 114.7 | 114.375 | 113.96 | 1.6 | 78.6 | 113.5 | 113.4 | 113.15 | 2.5 | 99.5667 | 111.275 | 111.7857 | 111.619 | 3.35 | 103.125 | 109.82 | 110.4333 | 110.0444
-5.2 | 1.3 | 52.8 | 114.5 | 114.023 | 113.66 | 1.3 | 65.1 | 114.5 | 114.08 | 113.75 | 1.7 | 86.6 | 113.35 | 113.3 | 112.92 | 2.6 | 101.2333 | 111.08 | 111.6429 | 111.369 | 3.5 | 103.65 | 109.65 | 110.24 | 109.9
-5.1 | 1.4 | 64.5 | 114.2 | 113.798 | 113.46 | 1.3 | 75.8 | 114.275 | 113.82 | 113.5 | 1.7667 | 93.2 | 113.2 | 113.06 | 112.7 | 2.7 | 102.3333 | 111 | 111.46 | 111.144 | 3.6 | 104.1667 | 109.475 | 110.025 | 109.6333
-5 | 1.4 | 76.3 | 114.1 | 113.548 | 113.25 | 1.4 | 85.4333 | 114 | 113.58 | 113.325 | 1.8 | 97.4 | 113.02 | 112.75 | 112.5 | 2.8 | 103.3 | 110.88 | 111.1 | 110.9757 | 3.75 | 104.525 | 109.35 | 109.9 | 109.5
-4.9 | 1.5 | 85.9 | 113.9 | 113.2563 | 113.02 | 1.4 | 92.6 | 113.7 | 113.4 | 113 | 1.9 | 100.1 | 112.7667 | 112.6 | 112.3 | 2.9 | 104.1333 | 110.7 | 110.9429 | 110.7894 | 3.85 | 104.9 | 109.225 | 109.8 | 109.4143
-4.8 | 1.5 | 94 | 113.575 | 113.023 | 112.775 | 1.5 | 97.675 | 113.4 | 113.2 | 112.76 | 2 | 102.1333 | 112.54 | 112.35 | 112.02 | 3 | 104.625 | 110.5833 | 110.7667 | 110.5513 | 4 | 105 | 109.1 | 109.575 | 109.2875
-4.7 | 1.6 | 99.15 | 113.325 | 112.873 | 112.5 | 1.5 | 101.1 | 113.2833 | 112.88 | 112.525 | 2.1 | 103.6 | 112.3 | 112.1 | 111.76 | 3.15 | 105 | 110.375 | 110.6143 | 110.368 | 4.2 | 105.3 | 108.95 | 109.3571 | 109.125
-4.6 | 1.65 | 102.5 | 113.0667 | 112.573 | 112.25 | 1.6 | 103.2667 | 113.05 | 112.65 | 112.3 | 2.2 | 104.6 | 112.075 | 111.94 | 111.5333 | 3.3 | 105.3 | 110.15 | 110.3857 | 110.138 | 4.3 | 105.4 | 108.9 | 109.2 | 108.9286
-4.5 | 1.7 | 104.6667 | 112.8 | 112.297 | 112 | 1.7 | 104.925 | 112.74 | 112.325 | 112.06 | 2.25 | 105.4 | 111.86 | 111.64 | 111.325 | 3.4 | 105.6 | 110.02 | 110.2 | 109.943 | 4.5 | 105.5667 | 108.68 | 109.1 | 108.8
-4.4 | 1.7667 | 106.2333 | 112.575 | 112.122 | 111.8 | 1.8 | 106.1333 | 112.6 | 112.1 | 111.8333 | 2.3 | 106.1 | 111.5 | 111.3333 | 111.18 | 3.7 | 105.8 | 109.8 | 109.9 | 109.7466 | 4.85 | 105.7 | 108.6 | 108.98 | 108.68
-4.3 | 1.8 | 107.2 | 112.3 | 111.872 | 111.5333 | 1.8 | 106.9 | 112.325 | 111.86 | 111.575 | 2.5 | 106.6 | 111.38 | 111.2 | 110.88 | 3.9 | 106.0333 | 109.6167 | 109.6286 | 109.618 | 5.1 | 105.8 | 108.3 | 108.76 | 108.5
-4.2 | 1.9 | 107.8667 | 112.05 | 111.542 | 111.3 | 1.9 | 107.4 | 112 | 111.64 | 111.24 | 2.6 | 106.9333 | 111.16 | 110.875 | 110.66 | 4.1 | 106.225 | 109.5 | 109.5 | 109.4847 | 5.4 | 105.8 | 108.125 | 108.6 | 108.3167
-4.1 | 2 | 108.5333 | 111.9 | 111.272 | 111.04 | 2 | 107.9333 | 111.72 | 111.4 | 111.1 | 2.85 | 107.2667 | 110.9 | 110.725 | 110.46 | 4.2 | 106.4 | 109.32 | 109.3833 | 109.2347 | 5.8 | 105.8 | 108.0167 | 108.5167 | 108.1571
-4 | 2.1 | 108.8 | 111.55 | 111.122 | 110.8 | 2.15 | 108.3333 | 111.56 | 111.08 | 110.85 | 3.1 | 107.4 | 110.7 | 110.45 | 110.28 | 4.5 | 106.4 | 109.2 | 109.2 | 109.038 | 6.1 | 105.9 | 107.9 | 108.32 | 108.0143
-3.9 | 2.2 | 109.1 | 111.4 | 110.822 | 110.55 | 2.3 | 108.5667 | 111.3 | 110.82 | 110.575 | 3.25 | 107.6 | 110.42 | 110.2 | 109.96 | 4.9 | 106.5 | 109 | 109.0375 | 108.8513 | 6.55 | 105.82 | 107.75 | 108.1 | 107.8571
-3.8 | 2.4 | 109.2 | 111.1 | 110.602 | 110.32 | 2.5 | 108.7 | 110.9 | 110.625 | 110.325 | 3.5 | 107.6 | 110.22 | 109.975 | 109.76 | 5.3 | 106.5 | 108.75 | 108.8 | 108.6751 | 7.05 | 105.9 | 107.6 | 107.9667 | 107.7
-3.7 | 2.5 | 109.3 | 110.7 | 110.347 | 110.08 | 2.7 | 108.8 | 110.725 | 110.32 | 110.1 | 3.7 | 107.7 | 110 | 109.74 | 109.54 | 5.7 | 106.6 | 108.6 | 108.6333 | 108.468 | 7.7333 | 105.9 | 107.5167 | 107.8 | 107.4857
-3.6 | 2.7 | 109.3 | 110.45 | 110.002 | 109.85 | 2.9 | 108.9 | 110.475 | 110.08 | 109.85 | 4.1 | 107.7 | 109.86 | 109.6 | 109.325 | 6.2 | 106.6 | 108.4 | 108.4 | 108.2894 | 8.5 | 105.9 | 107.3 | 107.55 | 107.35
-3.5 | 2.9 | 109.3 | 110.2 | 109.672 | 109.6 | 3.2 | 108.9 | 110.24 | 109.84 | 109.64 | 4.45 | 107.7 | 109.58 | 109.36 | 109.0833 | 6.65 | 106.6 | 108.2 | 108.1375 | 108.1347 | 9.4 | 105.8 | 107.1333 | 107.4 | 107.1667
-3.4 | 3.2333 | 109.2 | 110 | 109.522 | 109.35 | 3.55 | 108.9 | 110 | 109.6 | 109.4 | 5 | 107.7 | 109.35 | 109.05 | 108.9 | 7.4 | 106.5 | 108.075 | 107.9571 | 107.938 | 10.7 | 105.7 | 106.9 | 107.2 | 106.8857
-3.3 | 3.5 | 109.1 | 109.775 | 109.282 | 109.2 | 4 | 108.8 | 109.8 | 109.3333 | 109.0833 | 5.6 | 107.7 | 108.975 | 108.8333 | 108.65 | 8.4 | 106.475 | 107.825 | 107.8 | 107.7323 | 11.9 | 105.6 | 106.775 | 107.0333 | 106.6667
-3.2 | 3.9 | 108.9333 | 109.475 | 109.062 | 108.9333 | 4.55 | 108.65 | 109.4333 | 109.14 | 108.9 | 6.2 | 107.6 | 108.74 | 108.58 | 108.4 | 9.5 | 106.3 | 107.5 | 107.55 | 107.518 | 13.7 | 105.5 | 106.55 | 106.85 | 106.45
-3.1 | 4.4 | 108.8 | 109.2333 | 108.822 | 108.675 | 5.1 | 108.5 | 109.2 | 108.76 | 108.62 | 7 | 107.5 | 108.525 | 108.325 | 108.1667 | 10.8 | 106.2 | 107.28 | 107.275 | 107.2847 | 15.9 | 105.4 | 106.3 | 106.5429 | 106.2667
-3 | 5.075 | 108.6333 | 108.84 | 108.5553 | 108.4 | 5.9 | 108.3667 | 108.94 | 108.475 | 108.4 | 8.4 | 107.4 | 108.32 | 108.02 | 107.9667 | 12.7 | 106.0333 | 107.05 | 107.05 | 107.038 | 19 | 105.3 | 106.18 | 106.38 | 105.9833
-2.9 | 5.9 | 108.3333 | 108.62 | 108.222 | 108.1 | 6.95 | 108.2 | 108.6 | 108.2 | 108.175 | 9.6 | 107.2333 | 108.1 | 107.8 | 107.64 | 14.9 | 105.9667 | 106.8 | 106.9 | 106.7894 | 23.3 | 105.125 | 106 | 106.1667 | 105.8167
-2.8 | 7.15 | 108.2 | 108.325 | 107.897 | 107.875 | 8.4 | 108 | 108.35 | 108 | 107.725 | 11.5 | 106.925 | 107.7667 | 107.5667 | 107.275 | 18 | 105.8 | 106.6 | 106.5667 | 106.538 | 29.6 | 105 | 105.8 | 105.8 | 105.64
-2.7 | 8.4667 | 107.9667 | 107.975 | 107.622 | 107.5 | 10.2 | 107.7 | 108.1 | 107.7 | 107.5 | 13.9 | 106.75 | 107.5 | 107.3 | 107.06 | 22.5 | 105.6 | 106.4 | 106.2714 | 106.343 | 37.7 | 104.8 | 105.6 | 105.62 | 105.3714
-2.6 | 10.7 | 107.75 | 107.7333 | 107.421 | 107.3 | 12.4 | 107.55 | 107.75 | 107.38 | 107.1667 | 17.3 | 106.5333 | 107.275 | 106.875 | 106.78 | 29.15 | 105.4 | 106.1 | 106.1167 | 106.1347 | 47.3 | 104.7 | 105.32 | 105.4 | 105.12
-2.5 | 13.5 | 107.5 | 107.475 | 107.171 | 106.9667 | 15.7 | 107.25 | 107.42 | 107.1 | 106.85 | 22.3 | 106.3 | 106.85 | 106.64 | 106.5 | 37.5 | 105.2 | 105.9 | 105.82 | 105.843 | 57.6 | 104.55 | 105.1333 | 105.26 | 104.9571
-2.4 | 17.75 | 107.05 | 107.075 | 106.821 | 106.6 | 20.3 | 107 | 107 | 106.82 | 106.6 | 29.5 | 106 | 106.54 | 106.5 | 106.14 | 47.3 | 105 | 105.5667 | 105.5333 | 105.5751 | 66.85 | 104.4 | 104.9 | 105.04 | 104.7
-2.3 | 24.575 | 106.8 | 106.8 | 106.471 | 106.4 | 27.7 | 106.8 | 106.65 | 106.5 | 106.28 | 38.55 | 105.8 | 106.25 | 106.1 | 106 | 57.4 | 104.7 | 105.32 | 105.3143 | 105.2805 | 76.25 | 104.2 | 104.5667 | 104.85 | 104.5429
-2.2 | 33.9 | 106.5 | 106.4 | 106.121 | 106.1 | 37.15 | 106.4 | 106.36 | 106.2 | 106 | 49.1 | 105.6667 | 106.02 | 105.775 | 105.64 | 67.7 | 104.6 | 105 | 105.1 | 104.9426 | 82.85 | 104.1 | 104.3167 | 104.5857 | 104.4
-2.1 | 44.1 | 106.1333 | 106.08 | 105.771 | 105.7 | 48 | 106.1 | 106.08 | 105.85 | 105.75 | 59.6 | 105.3 | 105.72 | 105.4667 | 105.3 | 77 | 104.3667 | 104.8333 | 104.7833 | 104.717 | 88.2 | 103.875 | 104.0571 | 104.4333 | 104.1714
-2 | 55.775 | 105.8 | 105.8 | 105.546 | 105.425 | 59.6 | 105.7667 | 105.7 | 105.4667 | 105.42 | 70.5 | 105.1333 | 105.35 | 105.125 | 105.075 | 84.25 | 104.1333 | 104.55 | 104.6 | 104.617 | 91.7 | 103.8 | 103.9 | 104.2333 | 104.0286
-1.9 | 67.6 | 105.4667 | 105.4 | 105.221 | 105.15 | 70 | 105.4333 | 105.4 | 105.1 | 105.05 | 79.5 | 104.8 | 105.04 | 104.94 | 104.76 | 89.1 | 104 | 104.35 | 104.3857 | 104.342 | 94.4 | 103.6 | 103.66 | 104.06 | 103.78
-1.8 | 77.2667 | 105.2 | 105.02 | 104.861 | 104.85 | 79.2 | 105.2 | 105.1 | 104.78 | 104.775 | 86.3 | 104.6 | 104.74 | 104.6667 | 104.45 | 92.5 | 103.7 | 104.02 | 104.08 | 104.0027 | 96.15 | 103.4 | 103.54 | 103.7 | 103.5714
-1.7 | 86.2 | 104.85 | 104.75 | 104.621 | 104.4 | 86.8 | 104.85 | 104.74 | 104.46 | 104.48 | 91.3 | 104.3 | 104.4333 | 104.325 | 104.24 | 95 | 103.525 | 103.8 | 103.8 | 103.8003 | 97.3 | 103.16 | 103.32 | 103.525 | 103.3333
-1.6 | 91.75 | 104.6333 | 104.425 | 104.2543 | 104.12 | 91.4 | 104.55 | 104.45 | 104.24 | 104.15 | 93.9 | 104.0667 | 104.3 | 104.075 | 103.95 | 96.5 | 103.3 | 103.6 | 103.6286 | 103.6337 | 98.5 | 103 | 103.12 | 103.3 | 103.15
-1.5 | 94.9 | 104.3667 | 104.16 | 103.961 | 103.9 | 94.6 | 104.1 | 104.225 | 103.94 | 103.8 | 96.2 | 103.8 | 103.925 | 103.82 | 103.6 | 98.1 | 103.1667 | 103.3 | 103.3667 | 103.417 | 99.2 | 102.8 | 102.8667 | 103.075 | 102.86
-1.4 | 97.4 | 104.1333 | 103.875 | 103.621 | 103.66 | 96.6 | 103.9 | 103.8667 | 103.8 | 103.55 | 97.5 | 103.6667 | 103.6167 | 103.58 | 103.4 | 98.8 | 103 | 103.1 | 103.1 | 103.1837 | 99.65 | 102.65 | 102.6167 | 102.85 | 102.6667
-1.3 | 98.9333 | 103.7333 | 103.62 | 103.346 | 103.375 | 98.5 | 103.7 | 103.5 | 103.525 | 103.32 | 98.7 | 103.3 | 103.38 | 103.25 | 103.2333 | 99.5 | 102.9 | 102.88 | 102.8333 | 103.017 | 100.1 | 102.4333 | 102.44 | 102.625 | 102.5
-1.2 | 99.8667 | 103.3 | 103.275 | 103.071 | 102.98 | 99.1 | 103.4 | 103.3 | 103.075 | 103.12 | 99.3 | 103 | 103.18 | 102.925 | 102.9333 | 99.8 | 102.6333 | 102.6667 | 102.64 | 102.817 | 100.4 | 102.25 | 102.3 | 102.4 | 102.28
-1.1 | 100.5 | 103.1 | 103.025 | 102.846 | 102.8 | 99.6 | 103.1 | 103.06 | 102.82 | 102.7 | 99.8 | 102.8333 | 102.9333 | 102.7 | 102.68 | 100.2 | 102.4 | 102.375 | 102.3778 | 102.457 | 100.6 | 102.1 | 102.1 | 102.18 | 102.1
-1 | 100.8667 | 102.8333 | 102.675 | 102.641 | 102.525 | 100.2 | 102.8333 | 102.7 | 102.45 | 102.4333 | 100.3 | 102.5 | 102.64 | 102.4 | 102.38 | 100.5 | 102.06 | 102.16 | 102.18 | 102.217 | 100.7 | 101.95 | 101.9 | 102.0143 | 101.9571
-0.9 | 101.1 | 102.4667 | 102.425 | 102.32 | 102.28 | 100.55 | 102.4 | 102.3667 | 102.25 | 102.1 | 100.6 | 102.2 | 102.3 | 102.2333 | 102.14 | 100.6 | 101.875 | 101.95 | 101.88 | 102.017 | 100.8 | 101.7 | 101.6429 | 101.8 | 101.8
-0.8 | 101.2333 | 102.1 | 102.2 | 102.02 | 102.05 | 100.7 | 102.2333 | 102.2 | 101.98 | 101.86 | 100.8 | 102 | 102 | 101.92 | 102 | 100.7 | 101.6667 | 101.7 | 101.775 | 101.7795 | 100.8 | 101.5333 | 101.4833 | 101.66 | 101.5
-0.7 | 101.3 | 101.8 | 101.88 | 101.72 | 101.8 | 100.8 | 101.85 | 101.875 | 101.76 | 101.625 | 100.8 | 101.7333 | 101.7 | 101.72 | 101.675 | 100.7 | 101.4 | 101.4 | 101.525 | 101.517 | 100.8 | 101.325 | 101.2333 | 101.425 | 101.3167
-0.6 | 101.3 | 101.6 | 101.58 | 101.3867 | 101.45 | 100.8 | 101.5333 | 101.55 | 101.38 | 101.44 | 100.8 | 101.4667 | 101.6 | 101.475 | 101.375 | 100.8 | 101.1667 | 101.2333 | 101.3 | 101.3837 | 100.7 | 101.0667 | 101.02 | 101.1667 | 101.025
-0.5 | 101.2 | 101.3333 | 101.3333 | 101.14 | 101.175 | 100.8 | 101.3 | 101.3 | 101.15 | 101.1 | 100.7 | 101.1667 | 101.275 | 101.26 | 101.125 | 100.7 | 100.9 | 101 | 101.1167 | 101.217 | 100.6 | 100.8 | 100.9 | 101 | 100.8167
-0.4 | 101 | 101 | 100.94 | 101.02 | 101.02 | 100.7 | 101 | 100.975 | 100.9 | 100.9 | 100.6 | 100.9333 | 101.05 | 101 | 100.925 | 100.6 | 100.7 | 100.7667 | 100.85 | 100.9337 | 100.5 | 100.6667 | 100.64 | 100.9 | 100.6857
-0.3 | 100.8333 | 100.775 | 100.72 | 100.695 | 100.675 | 100.5 | 100.7667 | 100.65 | 100.65 | 100.65 | 100.5 | 100.6667 | 100.7 | 100.8 | 100.7 | 100.5 | 100.5 | 100.575 | 100.6 | 100.6837 | 100.4 | 100.5 | 100.5 | 100.65 | 100.44
-0.2 | 100.55 | 100.4333 | 100.45 | 100.44 | 100.52 | 100.3 | 100.45 | 100.44 | 100.4 | 100.42 | 100.35 | 100.35 | 100.52 | 100.5 | 100.5 | 100.3 | 100.3667 | 100.425 | 100.3167 | 100.497 | 100.35 | 100.3333 | 100.275 | 100.4333 | 100.275
-0.1 | 100.1 | 100.1 | 100.1 | 100.22 | 100.25 | 100.2 | 100.1667 | 100.2 | 100.175 | 100.175 | 100.2 | 100.2 | 100.22 | 100.2 | 100.28 | 100.1 | 100.2 | 100.22 | 100.175 | 100.377 | 100.2 | 100.175 | 100.15 | 100.24 | 100.15
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0003 | 100 | 100 | 100 | 100 | 100
0.1 | 99.7333 | 99.8 | 99.82 | 99.8533 | 99.8 | 99.8 | 99.7667 | 99.9 | 99.8 | 99.8 | 99.8 | 99.7333 | 99.9 | 99.775 | 99.8 | 99.7 | 99.82 | 99.84 | 99.8167 | 99.8452 | 99.8 | 99.86 | 99.82 | 99.84 | 99.8
0.2 | 99.5667 | 99.5 | 99.6 | 99.6199 | 99.62 | 99.5 | 99.5667 | 99.7 | 99.64 | 99.62 | 99.6 | 99.6 | 99.7 | 99.7 | 99.72 | 99.5 | 99.6 | 99.64 | 99.7 | 99.8166 | 99.7 | 99.725 | 99.625 | 99.7333 | 99.7143
0.3 | 99.4 | 99.3 | 99.45 | 99.4699 | 99.475 | 99.2 | 99.4 | 99.575 | 99.52 | 99.45 | 99.4 | 99.4 | 99.48 | 99.5 | 99.55 | 99.3 | 99.4 | 99.5 | 99.5375 | 99.6666 | 99.55 | 99.5 | 99.5333 | 99.6333 | 99.5571
0.4 | 99.1 | 99.2 | 99.2 | 99.3199 | 99.3 | 98.9 | 99.3 | 99.32 | 99.375 | 99.275 | 99.1 | 99.2 | 99.35 | 99.4 | 99.3 | 99.1 | 99.2 | 99.4 | 99.38 | 99.5166 | 99.4 | 99.4 | 99.34 | 99.5 | 99.4
0.5 | 98.9 | 99 | 99 | 99.1698 | 99.2 | 98.55 | 99.2 | 99.1333 | 99.2 | 99.125 | 98.7 | 99.0333 | 99.22 | 99.3 | 99.2 | 98.8 | 99.1 | 99.2 | 99.2143 | 99.3916 | 99.1 | 99.2 | 99.2 | 99.4 | 99.3
0.6 | 98.5667 | 98.7667 | 98.9 | 99.1198 | 99.08 | 98.3 | 99 | 98.975 | 99 | 99 | 98.4 | 98.8667 | 99.1 | 99.1 | 99.1 | 98.6 | 99.05 | 99.1 | 99.0571 | 99.2499 | 98.85 | 99.1 | 99 | 99.2833 | 99.2167
0.7 | 98.2 | 98.6 | 98.8 | 98.9948 | 98.975 | 97.75 | 98.7333 | 98.825 | 98.875 | 98.9 | 98.1 | 98.7 | 98.85 | 98.98 | 99 | 98.2667 | 98.9 | 99 | 98.9 | 99.1832 | 98.6 | 99 | 98.9333 | 99.2 | 99.0333
0.8 | 97.7667 | 98.4 | 98.6 | 98.8198 | 98.85 | 97.25 | 98.6 | 98.7 | 98.75 | 98.7333 | 97.6 | 98.525 | 98.72 | 98.9 | 98.9 | 97.8 | 98.75 | 98.8333 | 98.9 | 99.0165 | 98.25 | 98.8667 | 98.8667 | 99 | 98.9571
0.9 | 97.3333 | 98.3 | 98.6 | 98.6697 | 98.7 | 96.8 | 98.5 | 98.6 | 98.7 | 98.66 | 97.2 | 98.4667 | 98.7 | 98.8 | 98.8 | 97.45 | 98.7 | 98.7 | 98.8 | 98.9332 | 98 | 98.76 | 98.7667 | 98.9 | 98.8
1 | 96.9 | 98.2 | 98.5 | 98.5447 | 98.6 | 96.15 | 98.4 | 98.5 | 98.58 | 98.54 | 96.7 | 98.4 | 98.54 | 98.7 | 98.7 | 96.9 | 98.525 | 98.6 | 98.6714 | 98.7831 | 97.7 | 98.625 | 98.6833 | 98.8 | 98.75
1.1 | 96.1 | 98.0333 | 98.325 | 98.4197 | 98.475 | 95.5 | 98.3 | 98.375 | 98.4 | 98.5 | 96 | 98.2 | 98.475 | 98.6 | 98.525 | 96.5 | 98.4 | 98.55 | 98.5167 | 98.6164 | 97.3 | 98.5 | 98.6 | 98.7 | 98.6
1.2 | 95.0667 | 97.975 | 98.175 | 98.3197 | 98.4 | 94.6 | 98.1 | 98.26 | 98.325 | 98.4 | 95.2 | 98.2 | 98.3 | 98.5 | 98.4 | 96.2 | 98.225 | 98.4333 | 98.45 | 98.5164 | 96.85 | 98.4 | 98.54 | 98.6 | 98.6
1.3 | 93.8333 | 97.9 | 98.04 | 98.2596 | 98.3 | 93.55 | 98 | 98.16 | 98.2 | 98.24 | 94.25 | 98 | 98.18 | 98.36 | 98.3 | 95.3 | 98.1 | 98.225 | 98.3 | 98.4164 | 96.4 | 98.2667 | 98.4 | 98.5 | 98.4
1.4 | 92.2 | 97.6 | 97.9 | 98.0396 | 98.25 | 91.9 | 97.8 | 98 | 98.1 | 98.1 | 93.1 | 97.8 | 98 | 98.25 | 98.2 | 94.4667 | 98 | 98.1 | 98.2 | 98.3164 | 95.85 | 98.175 | 98.275 | 98.4 | 98.3667
1.5 | 89.9 | 97.475 | 97.78 | 97.9196 | 98.1 | 90.1 | 97.7 | 97.9 | 98 | 98 | 91.6 | 97.7 | 97.9667 | 98.1 | 98.1 | 93.4 | 97.85 | 98.06 | 98.1143 | 98.2331 | 94.85 | 98 | 98.15 | 98.3 | 98.2143
1.6 | 86.8 | 97.35 | 97.6 | 97.8196 | 98 | 87.25 | 97.5 | 97.78 | 97.8 | 97.9167 | 89.5 | 97.5 | 97.8 | 97.875 | 97.9667 | 92.45 | 97.7 | 97.9 | 98.04 | 98.1164 | 94.1 | 97.95 | 98 | 98.2 | 98.175
1.7 | 82 | 97.1333 | 97.525 | 97.7195 | 97.85 | 83.15 | 97.325 | 97.6 | 97.75 | 97.8 | 86.3 | 97.4 | 97.7 | 97.72 | 97.825 | 90.4 | 97.5 | 97.8 | 97.86 | 97.9763 | 93.2 | 97.86 | 97.9 | 98.1 | 98.05
1.8 | 74.9 | 97 | 97.4 | 97.5995 | 97.7 | 77.5 | 97.2 | 97.5 | 97.6 | 97.65 | 82.4 | 97.15 | 97.56 | 97.7 | 97.7333 | 88.6 | 97.4 | 97.66 | 97.7667 | 97.8996 | 91.7 | 97.6667 | 97.8 | 98 | 97.9
1.9 | 65.925 | 96.8 | 97.2 | 97.4695 | 97.55 | 69.1 | 97.0333 | 97.35 | 97.5 | 97.575 | 77.1 | 97 | 97.4333 | 97.6 | 97.65 | 85.7 | 97.3 | 97.55 | 97.7 | 97.8163 | 90 | 97.5667 | 97.7 | 97.88 | 97.9
2 | 55.9 | 96.7 | 97.125 | 97.4195 | 97.5 | 59.65 | 96.9 | 97.3 | 97.42 | 97.5 | 67.9 | 96.875 | 97.4 | 97.5 | 97.55 | 81.9 | 97.1 | 97.4 | 97.6 | 97.7163 | 87.9 | 97.3 | 97.6 | 97.76 | 97.8
2.1 | 44.4 | 96.5667 | 97.075 | 97.3195 | 97.4 | 49.3 | 96.8 | 97.16 | 97.3 | 97.4 | 58.7 | 96.7 | 97.2 | 97.4 | 97.425 | 76.5 | 96.9333 | 97.3 | 97.475 | 97.633 | 85.4 | 97.1 | 97.45 | 97.7 | 97.7
2.2 | 33.7 | 96.4 | 96.9 | 97.2194 | 97.3 | 39.2 | 96.7 | 97 | 97.3 | 97.3 | 48.9 | 96.6 | 97.1 | 97.3 | 97.3667 | 69.3 | 96.675 | 97.2 | 97.35 | 97.5363 | 81.05 | 97 | 97.35 | 97.6 | 97.6
2.3 | 24.4 | 96.3 | 96.8 | 97.1194 | 97.22 | 29.5 | 96.575 | 96.9 | 97.2 | 97.24 | 38.7667 | 96.4 | 97 | 97.2 | 97.28 | 60 | 96.5 | 97.1167 | 97.2857 | 97.4877 | 75.6 | 96.9 | 97.2 | 97.5 | 97.5
2.4 | 18 | 96.2 | 96.7 | 97.1194 | 97.1 | 21.95 | 96.3667 | 96.8 | 97.1 | 97.2 | 29.7 | 96.3333 | 96.9 | 97.2 | 97.2 | 50.4 | 96.34 | 97 | 97.2 | 97.3162 | 68.1 | 96.75 | 97.1333 | 97.4 | 97.4667
2.5 | 13.4 | 96.1 | 96.6 | 97.0194 | 97 | 16.3 | 96.2 | 96.8 | 97.1 | 97.1 | 22.5 | 96.2 | 96.8 | 97.1 | 97.1 | 41.2 | 96.2 | 96.9 | 97.0833 | 97.2162 | 59.5 | 96.5 | 97.0333 | 97.3 | 97.3
2.6 | 10.4333 | 95.9 | 96.48 | 96.8944 | 96.9 | 13 | 96 | 96.625 | 96.95 | 97 | 17.2 | 96 | 96.7 | 97 | 97 | 32.3 | 96.0333 | 96.8 | 97 | 97.1448 | 50.6 | 96.3 | 97 | 97.2 | 97.1714
2.7 | 8.4 | 95.8 | 96.4 | 96.7443 | 96.8833 | 10.3 | 95.9 | 96.5 | 96.82 | 96.9 | 13.7 | 95.8 | 96.6 | 96.9 | 96.9667 | 24.8 | 95.9 | 96.7 | 96.8667 | 97.1162 | 41 | 96.2 | 96.9 | 97.12 | 97
2.8 | 6.8 | 95.5333 | 96.25 | 96.6193 | 96.7 | 8.5 | 95.7 | 96.4 | 96.7 | 96.8 | 11.35 | 95.6 | 96.5 | 96.76 | 96.86 | 19.6 | 95.775 | 96.6 | 96.8 | 96.9912 | 33.15 | 96.025 | 96.8 | 97.1 | 96.9333
2.9 | 5.6333 | 95.4 | 96.12 | 96.5193 | 96.6 | 7.2 | 95.5 | 96.3 | 96.575 | 96.7 | 9.45 | 95.4333 | 96.4 | 96.7 | 96.75 | 15.8 | 95.525 | 96.525 | 96.7 | 96.859 | 25.8 | 95.8 | 96.7 | 97 | 96.8
3 | 4.7667 | 95.1667 | 96 | 96.4193 | 96.5 | 6 | 95.2667 | 96.2 | 96.46 | 96.5 | 8.1 | 95.2 | 96.3 | 96.6 | 96.6 | 13.4 | 95.275 | 96.4 | 96.5667 | 96.7161 | 20.55 | 95.6 | 96.6 | 96.9 | 96.7375
3.1 | 4.1667 | 94.8667 | 95.85 | 96.3193 | 96.4 | 5.15 | 95 | 96.1 | 96.32 | 96.38 | 6.9 | 94.9 | 96.2 | 96.4 | 96.5 | 11.55 | 95.1 | 96.3 | 96.5 | 96.6494 | 17.2 | 95.4 | 96.42 | 96.7 | 96.6857
3.2 | 3.6667 | 94.6 | 95.68 | 96.1525 | 96.3 | 4.55 | 94.8 | 96 | 96.3 | 96.3 | 6 | 94.7 | 96 | 96.3 | 96.4 | 9.9 | 95.0333 | 96.1667 | 96.3857 | 96.5828 | 14.5 | 95.25 | 96.4 | 96.6667 | 96.55
3.3 | 3.3 | 94.4 | 95.575 | 96.0192 | 96.1 | 4 | 94.4333 | 95.8 | 96.175 | 96.2 | 5.3 | 94.45 | 95.85 | 96.2 | 96.2 | 8.7 | 94.8 | 95.9667 | 96.3 | 96.4161 | 12.3 | 95.025 | 96.2333 | 96.6 | 96.5
3.4 | 3 | 94.2 | 95.35 | 95.8525 | 95.9 | 3.5 | 94.2 | 95.62 | 95.925 | 96 | 4.75 | 93.9667 | 95.7 | 96.1 | 96.1 | 7.7333 | 94.5 | 95.86 | 96.15 | 96.3161 | 10.9 | 94.875 | 96.1 | 96.425 | 96.4
3.5 | 2.7333 | 93.9333 | 95.2 | 95.6191 | 95.775 | 3.2 | 93.85 | 95.475 | 95.75 | 95.86 | 4.3 | 93.7 | 95.5167 | 95.9667 | 95.975 | 7 | 94.1 | 95.8 | 96.0833 | 96.1785 | 9.7 | 94.625 | 96 | 96.3 | 96.3
3.6 | 2.5 | 93.3 | 95 | 95.5191 | 95.65 | 2.9 | 93.4 | 95.3 | 95.7 | 95.725 | 4 | 93.475 | 95.42 | 95.8 | 95.84 | 6.2 | 93.8 | 95.675 | 96 | 96.116 | 8.6 | 94.4 | 95.9 | 96.2 | 96.1
3.7 | 2.25 | 92.9333 | 94.8 | 95.4191 | 95.525 | 2.6 | 93.1 | 95.1 | 95.5 | 95.6 | 3.5 | 93.175 | 95.2 | 95.62 | 95.6833 | 5.55 | 93.575 | 95.45 | 95.74 | 95.916 | 7.8 | 94.3 | 95.8 | 96.1 | 96.0429
3.8 | 2.2 | 92.6 | 94.7 | 95.244 | 95.3 | 2.4 | 92.6 | 94.92 | 95.3 | 95.45 | 3.4 | 92.8 | 95.025 | 95.5 | 95.5167 | 5.2 | 93.35 | 95.3 | 95.55 | 95.7302 | 7 | 94.1 | 95.6 | 96 | 95.9125
3.9 | 2 | 92.2 | 94.5 | 95.019 | 95.2 | 2.2 | 92.2 | 94.8 | 95.2 | 95.225 | 3.1 | 92.3667 | 94.88 | 95.3 | 95.4 | 4.9 | 93 | 95.1833 | 95.4857 | 95.6159 | 6.65 | 93.8 | 95.48 | 95.9 | 95.8714
4 | 1.9 | 91.525 | 94.4 | 94.879 | 94.975 | 2.1 | 91.6 | 94.65 | 95.025 | 95.06 | 2.85 | 91.8 | 94.7 | 95.125 | 95.2 | 4.45 | 92.6 | 95.1 | 95.3429 | 95.5302 | 6.05 | 93.6 | 95.26 | 95.7667 | 95.7
4.1 | 1.8 | 90.85 | 94.175 | 94.6939 | 94.86 | 2 | 90.9667 | 94.44 | 94.8333 | 94.9 | 2.7 | 91.4 | 94.58 | 95 | 95.075 | 4.2 | 92.3 | 94.875 | 95.1714 | 95.3588 | 5.75 | 93.2333 | 95.1667 | 95.5833 | 95.6
4.2 | 1.75 | 90.1 | 94.025 | 94.5189 | 94.6 | 1.8 | 90.2 | 94.2167 | 94.68 | 94.8 | 2.6 | 90.8 | 94.46 | 94.85 | 94.9333 | 3.9 | 91.85 | 94.64 | 95.0375 | 95.2159 | 5.4 | 92.875 | 95.06 | 95.4167 | 95.4167
4.3 | 1.7 | 89.1667 | 93.74 | 94.3989 | 94.475 | 1.7 | 89.0667 | 94.1 | 94.46 | 94.65 | 2.4 | 90.1 | 94.3 | 94.7 | 94.775 | 3.7 | 91.4333 | 94.475 | 94.8833 | 95.1016 | 5.1 | 92.575 | 94.84 | 95.3 | 95.34
4.4 | 1.6 | 87.9 | 93.6 | 94.2188 | 94.3 | 1.6 | 88.1333 | 93.85 | 94.4 | 94.525 | 2.3 | 89.3333 | 94.05 | 94.5 | 94.66 | 3.6 | 90.9667 | 94.28 | 94.7 | 94.9358 | 4.8 | 92.2 | 94.7 | 95.2 | 95.1857
4.5 | 1.5 | 86.25 | 93.375 | 94.0688 | 94.2 | 1.5 | 86.7 | 93.66 | 94.26 | 94.275 | 2.2 | 88.4 | 93.84 | 94.34 | 94.425 | 3.4 | 90.4333 | 94.1 | 94.5 | 94.7558 | 4.5 | 91.8 | 94.5 | 95.1333 | 95.0333
4.6 | 1.5 | 84.1667 | 93.175 | 93.9188 | 94 | 1.5 | 85.1 | 93.5 | 94 | 94.15 | 2 | 87.525 | 93.7 | 94.2 | 94.28 | 3.3 | 90 | 93.9 | 94.3714 | 94.5758 | 4.3 | 91.5333 | 94.34 | 95 | 94.975
4.7 | 1.4 | 81.4 | 92.94 | 93.7187 | 93.825 | 1.4 | 83.125 | 93.28 | 93.9 | 94.02 | 1.9 | 86.3 | 93.45 | 94 | 94.16 | 3.1 | 89.4 | 93.7833 | 94.3 | 94.4032 | 4.1 | 91.2 | 94.125 | 94.775 | 94.8
4.8 | 1.4 | 77.25 | 92.8333 | 93.5687 | 93.7 | 1.3 | 80.0667 | 93.05 | 93.8 | 93.9 | 1.8 | 84.9 | 93.225 | 93.86 | 93.98 | 2.9 | 88.7333 | 93.6167 | 94.14 | 94.2728 | 3.95 | 90.7 | 93.92 | 94.64 | 94.6143
4.9 | 1.3 | 70.7667 | 92.6 | 93.4187 | 93.5 | 1.3 | 75.8667 | 92.9 | 93.66 | 93.66 | 1.8 | 83.05 | 93.075 | 93.725 | 93.8 | 2.8 | 88.0333 | 93.44 | 93.9667 | 94.1157 | 3.8 | 90.275 | 93.72 | 94.525 | 94.475
5 | 1.3 | 63.1 | 92.45 | 93.1936 | 93.35 | 1.2 | 69.6667 | 92.7 | 93.425 | 93.54 | 1.7 | 80.55 | 92.925 | 93.6 | 93.62 | 2.7 | 87.2 | 93.2 | 93.8 | 93.9728 | 3.7 | 89.8 | 93.6 | 94.38 | 94.3333
5.1 | 1.2 | 53.975 | 92.2 | 92.9686 | 93.18 | 1.2 | 62 | 92.5 | 93.3 | 93.25 | 1.6 | 77.3 | 92.7 | 93.44 | 93.4 | 2.6 | 86.1 | 93 | 93.6429 | 93.9157 | 3.6 | 88.9667 | 93.425 | 94.2 | 94.2
5.2 | 1.2 | 44 | 92 | 92.8586 | 92.95 | 1.1 | 52.9 | 92.375 | 93.075 | 93.1 | 1.6 | 72.4333 | 92.5 | 93.2 | 93.25 | 2.5 | 84.9 | 92.9 | 93.5 | 93.7823 | 3.5 | 88.4667 | 93.3 | 94.0167 | 94
5.3 | 1.1 | 35.4333 | 91.82 | 92.6185 | 92.825 | 1.1 | 43.9 | 92.2 | 92.9 | 92.925 | 1.5 | 65.9 | 92.4 | 93.1 | 93.1 | 2.4 | 83.4 | 92.6667 | 93.4571 | 93.587 | 3.3 | 87.9 | 93.08 | 94 | 93.9286
5.4 | 1.1 | 27.1 | 91.5 | 92.4352 | 92.6 | 1.1 | 35.3 | 91.9 | 92.7 | 92.8 | 1.4 | 58 | 92.16 | 92.9 | 92.975 | 2.3 | 81.7 | 92.4833 | 93.2 | 93.4156 | 3.2 | 87.25 | 92.9 | 93.76 | 93.74
5.5 | 1.1 | 20.9667 | 91.38 | 92.2384 | 92.4 | 1.1 | 27.7667 | 91.7 | 92.475 | 92.58 | 1.4 | 49.7667 | 91.94 | 92.75 | 92.825 | 2.2 | 79.275 | 92.3 | 93.1 | 93.203 | 3.1 | 86.3 | 92.8 | 93.575 | 93.5
5.6 | 1.1 | 17 | 91.2 | 92.0934 | 92.2 | 1 | 21.4 | 91.58 | 92.34 | 92.48 | 1.3 | 40.9 | 91.75 | 92.6 | 92.7 | 2.2 | 76.0333 | 92.08 | 92.9286 | 93.0655 | 2.9 | 85.3 | 92.62 | 93.425 | 93.3833
5.7 | 1 | 14.1667 | 90.96 | 91.8184 | 92.0333 | 1 | 17.4333 | 91.275 | 92.2 | 92.25 | 1.3 | 32.9 | 91.6 | 92.4 | 92.5 | 2.1 | 71.325 | 91.9 | 92.7571 | 92.8869 | 2.9 | 83.9 | 92.45 | 93.3 | 93.2286
5.8 | 1 | 12.15 | 90.74 | 91.635 | 91.8 | 0.9 | 14.6 | 91.025 | 91.95 | 92 | 1.2 | 26.1 | 91.2833 | 92.25 | 92.425 | 2 | 65.2667 | 91.675 | 92.6429 | 92.7321 | 2.8 | 82.6 | 92.22 | 93.1833 | 93.0667
5.9 | 1 | 10.6 | 90.45 | 91.4183 | 91.575 | 0.9 | 12.6333 | 90.8 | 91.725 | 91.85 | 1.2 | 21.1333 | 91.1 | 92.02 | 92.1 | 2 | 57.9 | 91.48 | 92.4167 | 92.6154 | 2.7 | 80.8667 | 92.075 | 93.1 | 92.9833
6 | 1 | 9.4667 | 90.2 | 91.2432 | 91.4 | 0.9 | 11.1 | 90.65 | 91.6167 | 91.6 | 1.2 | 17.8 | 90.975 | 91.78 | 91.9 | 1.9 | 50.0333 | 91.38 | 92.2143 | 92.4954 | 2.6 | 78.5 | 91.88 | 92.9 | 92.775
6.1 | 1 | 8.6 | 90.06 | 91.0849 | 91.2 | 0.9 | 9.8 | 90.4 | 91.34 | 91.4333 | 1.2 | 15.4 | 90.68 | 91.5667 | 91.725 | 1.8 | 42.2333 | 91.18 | 92.0143 | 92.244 | 2.5 | 75.675 | 91.6667 | 92.7 | 92.68
6.2 | 0.9 | 7.9 | 89.7 | 90.8182 | 91.05 | 0.8 | 8.9 | 90.2 | 91.075 | 91.2 | 1.1 | 13.5 | 90.4 | 91.375 | 91.48 | 1.8 | 34.65 | 90.95 | 91.84 | 91.982 | 2.5 | 71.2333 | 91.4 | 92.52 | 92.5667
6.3 | 0.9 | 7.35 | 89.4 | 90.6681 | 90.9 | 0.8 | 8.1333 | 90 | 90.925 | 91 | 1.1 | 12.1 | 90.175 | 91.2 | 91.3 | 1.7 | 28.7 | 90.74 | 91.7 | 91.8153 | 2.4 | 65.9 | 91.275 | 92.4 | 92.3167
6.4 | 0.9 | 6.85 | 89.16 | 90.4181 | 90.6 | 0.8 | 7.5 | 89.7 | 90.675 | 90.76 | 1 | 11 | 89.925 | 91 | 91.175 | 1.7 | 23.9333 | 90.5 | 91.5714 | 91.732 | 2.3 | 59 | 91.04 | 92.24 | 92.1667
6.5 | 0.9 | 6.55 | 89 | 90.168 | 90.375 | 0.8 | 6.9667 | 89.45 | 90.525 | 90.6 | 1 | 10.0333 | 89.7 | 90.88 | 90.92 | 1.6 | 20.4 | 90.225 | 91.4167 | 91.5153 | 2.3 | 51.325 | 90.88 | 92.1 | 92
6.6 | 0.8 | 6.2 | 88.725 | 89.938 | 90.15 | 0.8 | 6.6 | 89.32 | 90.3667 | 90.45 | 1 | 9.3 | 89.575 | 90.6 | 90.7 | 1.6 | 18.05 | 89.975 | 91.1833 | 91.3952 | 2.2 | 44.05 | 90.7 | 91.9 | 91.8286
6.7 | 0.8 | 5.9333 | 88.4 | 89.818 | 89.9 | 0.8 | 6.1 | 89.08 | 90.2 | 90.22 | 1 | 8.7 | 89.3 | 90.4333 | 90.56 | 1.5 | 16.15 | 89.76 | 91 | 91.2152 | 2.1 | 36.7667 | 90.5333 | 91.8 | 91.7
6.8 | 0.8 | 5.6 | 88.24 | 89.5679 | 89.675 | 0.7 | 5.8 | 88.7333 | 89.9 | 90 | 1 | 8.1333 | 89.05 | 90.2 | 90.26 | 1.5 | 14.7 | 89.54 | 90.8167 | 91.0552 | 2.1 | 31 | 90.25 | 91.6667 | 91.5833
6.9 | 0.8 | 5.4 | 88.075 | 89.3179 | 89.5 | 0.7 | 5.5 | 88.4667 | 89.7 | 89.775 | 0.9 | 7.7 | 88.9 | 90 | 90.1 | 1.4 | 13.6 | 89.38 | 90.6429 | 90.7651 | 2.05 | 26.3 | 90.0667 | 91.48 | 91.4
7 | 0.8 | 5.25 | 87.76 | 89.0428 | 89.3 | 0.7 | 5.225 | 88.2 | 89.48 | 89.55 | 0.9 | 7.3 | 88.5667 | 89.85 | 89.9 | 1.4 | 12.5 | 89.15 | 90.4667 | 90.5951 | 2 | 22.6 | 89.84 | 91.3 | 91.2
7.1 |  | 5.1 | 87.56 | 88.8928 | 89 | 0.7 | 5 | 88.0333 | 89.275 | 89.36 | 0.9 | 7 | 88.3 | 89.6333 | 89.8 | 1.4 | 11.7667 | 88.925 | 90.2286 | 90.3865 | 1.9 | 20 | 89.6 | 91.1667 | 90.9
7.2 |  | 4.9 | 87.3 | 88.7177 | 88.9 |  | 4.8 | 87.85 | 89.1 | 89.15 |  | 6.5667 | 88.05 | 89.4 | 89.5 | 1.35 | 11 | 88.6 | 90.0571 | 90.1483 | 1.9 | 18.3 | 89.34 | 90.9 | 90.8
7.3 |  | 4.8 | 87 | 88.5177 | 88.625 |  | 4.6 | 87.525 | 88.9 | 88.975 |  | 6.325 | 87.8 | 89.16 | 89.4 | 1.3 | 10.45 | 88.36 | 89.88 | 90.065 | 1.8 | 16.8 | 89.18 | 90.76 | 90.7
7.4 |  | 4.6667 | 86.725 | 88.2677 | 88.425 |  | 4.5 | 87.26 | 88.675 | 88.8 |  | 6.05 | 87.56 | 88.94 | 89.15 |  | 10.0333 | 88.175 | 89.7 | 89.7549 | 1.8 | 15.5 | 88.9 | 90.5667 | 90.5167
7.5 |  | 4.5 | 86.46 | 88.1376 | 88.24 |  | 4.3 | 86.95 | 88.48 | 88.6 |  | 5.9 | 87.25 | 88.8 | 88.9 |  | 9.65 | 87.84 | 89.5 | 89.5863 | 1.8 | 14.4 | 88.7 | 90.35 | 90.3
7.6 |  | 4.4 | 86.175 | 87.8176 | 88 |  | 4.2 | 86.74 | 88.3 | 88.3 |  | 5.6 | 87.08 | 88.56 | 88.675 |  | 9.2 | 87.5 | 89.2571 | 89.4149 |  | 13.7667 | 88.5333 | 90.125 | 90.1
7.7 |  | 4.2333 | 86 | 87.6508 | 87.85 |  | 4.05 | 86.5 | 88 | 88.1 |  | 5.4 | 86.78 | 88.3667 | 88.5 |  | 8.7667 | 87.3 | 89.1143 | 89.2149 |  | 13.0667 | 88.32 | 89.975 | 89.9571
7.8 |  | 4.125 | 85.7 | 87.4508 | 87.54 |  | 3.9 | 86.25 | 87.825 | 87.9 |  | 5.2 | 86.4 | 88.125 | 88.25 |  | 8.5 | 87.12 | 88.8667 | 89.1149 |  | 12.3667 | 88.02 | 89.82 | 89.8
7.9 |  | 4.1 | 85.425 | 87.1174 | 87.32 |  | 3.8 | 85.975 | 87.58 | 87.7 |  | 5 | 86.15 | 87.9 | 88.05 |  | 8.2 | 86.85 | 88.7 | 88.9148 |  | 11.8 | 87.82 | 89.62 | 89.52
8 |  | 4 | 85.15 | 86.9774 | 87.1667 |  | 3.7 | 85.65 | 87.38 | 87.5 |  | 4.9 | 85.92 | 87.8 | 87.925 |  | 8 | 86.6667 | 88.4857 | 88.7291 |  | 11.4 | 87.475 | 89.4 | 89.4667
8.1 |  | 3.9 | 84.85 | 86.8174 | 86.9333 |  | 3.6667 | 85.425 | 87.15 | 87.35 |  | 4.8 | 85.68 | 87.575 | 87.7 |  | 7.65 | 86.36 | 88.1667 | 88.5291 |  | 10.9333 | 87.325 | 89.28 | 89.3143
8.2 |  | 3.8 | 84.5 | 86.584 | 86.68 |  | 3.5 | 85.1833 | 87 | 87.1 |  | 4.625 | 85.375 | 87.32 | 87.5 |  | 7.4333 | 86.1333 | 88.1 | 88.2861 |  | 10.6667 | 87.12 | 89.05 | 89.0667
8.3 |  | 3.7 | 84.3333 | 86.3423 | 86.45 |  | 3.5 | 84.875 | 86.8 | 86.875 |  | 4.45 | 85 | 87.2 | 87.2 |  | 7.3 | 85.8 | 87.92 | 88.1147 |  | 10.2 | 86.8333 | 88.94 | 88.8667
8.4 |  | 3.6 | 84.1 | 86.1339 | 86.225 |  | 3.4 | 84.6 | 86.5667 | 86.7 |  | 4.3 | 84.75 | 87 | 87.05 |  | 7.1 | 85.48 | 87.85 | 87.8896 |  | 9.85 | 86.5 | 88.7333 | 88.7
8.5 |  | 3.5333 | 83.575 | 85.8172 | 85.975 |  | 3.2333 | 84.225 | 86.3 | 86.45 |  | 4.2 | 84.5 | 86.825 | 86.9 |  | 6.875 | 85.26 | 87.5667 | 87.686 |  | 9.58 | 86.275 | 88.56 | 88.4667
8.6 |  | 3.5 | 83.325 | 85.6771 | 85.78 |  | 3.2 | 83.875 | 86.15 | 86.3 |  | 4.1 | 84.25 | 86.6 | 86.7 |  | 6.6 | 85.0667 | 87.3286 | 87.4896 |  | 9.44 | 85.95 | 88.34 | 88.2833
8.7 |  | 3.4 | 82.94 | 85.4421 | 85.56 |  | 3.1 | 83.6 | 85.9 | 86.05 |  | 4 | 83.94 | 86.42 | 86.48 |  | 6.4 | 84.8 | 87.1429 | 87.3146 |  | 9.1 | 85.7333 | 88.2 | 88.1167
8.8 |  | 3.325 | 82.4 | 85.217 | 85.36 |  | 3.025 | 83.2 | 85.7 | 85.825 |  | 3.9 | 83.6 | 86.2 | 86.26 |  | 6.3 | 84.45 | 87.0125 | 87.1812 |  | 8.825 | 85.46 | 87.96 | 87.85
8.9 |  | 3.275 | 82.1 | 84.992 | 85.15 |  | 3 | 82.675 | 85.4667 | 85.6 |  | 3.8 | 83.25 | 86.04 | 86.04 |  | 6.1667 | 84.14 | 86.775 | 86.9812 |  | 8.6 | 85.24 | 87.825 | 87.7
9 |  | 3.2 | 81.7 | 84.7369 | 84.85 |  | 2.9 | 82.2 | 85.325 | 85.45 |  | 3.7667 | 82.975 | 85.8 | 85.825 |  | 6 | 83.9 | 86.5625 | 86.7144 |  | 8.4 | 84.88 | 87.6 | 87.52
9.1 |  | 3.1 | 81.0667 | 84.5369 | 84.58 |  | 2.9 | 81.825 | 85.12 | 85.22 |  | 3.675 | 82.6 | 85.5667 | 85.6333 |  | 5.85 | 83.625 | 86.3 | 86.5144 |  | 8.2 | 84.675 | 87.4 | 87.3429
9.2 |  | 3.1 | 80.475 | 84.2835 | 84.48 |  | 2.8 | 81.325 | 84.84 | 85 |  | 3.6 | 82.2333 | 85.375 | 85.45 |  | 5.7 | 83.375 | 86.1333 | 86.3311 |  | 8 | 84.45 | 87.28 | 87.2125
9.3 |  | 3 | 79.8 | 84.0418 | 84.26 |  | 2.7 | 80.75 | 84.575 | 84.7 |  | 3.5 | 81.85 | 85.2 | 85.22 |  | 5.6 | 83.05 | 85.98 | 86.1811 |  | 7.8667 | 84.225 | 87.05 | 87.0429
9.4 |  | 3 | 79.04 | 83.8168 | 83.95 |  | 2.7 | 80.05 | 84.4 | 84.525 |  | 3.4 | 81.5 | 84.925 | 85.1 |  | 5.4 | 82.74 | 85.74 | 85.981 |  | 7.6333 | 83.9 | 86.85 | 86.8167
9.5 |  | 2.9 | 78.2 | 83.6667 | 83.8 |  | 2.6 | 79.375 | 84.22 | 84.3333 |  | 3.3333 | 80.98 | 84.78 | 84.825 |  | 5.3 | 82.325 | 85.5667 | 85.7393 |  | 7.475 | 83.65 | 86.625 | 86.5857
9.6 |  | 2.8 | 77.2 | 83.5167 | 83.6 |  | 2.6 | 78.6333 | 84.05 | 84.1 |  | 3.3 | 80.35 | 84.58 | 84.6 |  | 5.175 | 82 | 85.4143 | 85.581 |  | 7.4 | 83.3667 | 86.5 | 86.4333
9.7 |  | 2.8 | 75.68 | 83.3917 | 83.5 |  | 2.5 | 77.475 | 83.825 | 84 |  | 3.2 | 79.76 | 84.3 | 84.45 |  | 5.1 | 81.675 | 85.15 | 85.4142 |  | 7.2 | 83.22 | 86.275 | 86.26
9.8 |  | 2.8 | 73.7 | 83.2166 | 83.3 |  | 2.5 | 76.3 | 83.68 | 83.82 |  | 3.1333 | 79.2 | 84.1 | 84.2 |  | 4.9 | 81.46 | 84.9667 | 85.1571 |  | 7 | 82.84 | 86.075 | 86.1
9.9 |  | 2.7 | 71.2 | 83.1166 | 83.25 |  | 2.5 | 74.84 | 83.56 | 83.65 |  | 3.1 | 78.5 | 83.9333 | 83.96 |  | 4.8 | 81.12 | 84.8143 | 84.9892 |  | 6.9 | 82.5 | 85.95 | 85.9333
10 |  | 2.7 | 67.2 | 83.0166 | 83.1 |  | 2.4 | 72.86 | 83.425 | 83.5 |  | 3 | 77.8333 | 83.675 | 83.8 |  | 4.7 | 80.625 | 84.65 | 84.7941 |  | 6.8 | 82.2 | 85.78 | 85.7667
10.1 |  | 2.6 | 61.48 | 82.9166 | 83.06 |  | 2.4 | 70.275 | 83.2667 | 83.4 |  | 3 | 77.1 | 83.525 | 83.625 |  | 4.6 | 80.3 | 84.4 | 84.6141 |  | 6.6 | 81.95 | 85.56 | 85.5
10.2 |  |  | 55 | 82.8166 | 82.94 |  | 2.3 | 66.1 | 83.15 | 83.3 |  | 2.9 | 76 | 83.4 | 83.4 |  | 4.5 | 79.875 | 84.24 | 84.3474 |  | 6.45 | 81.68 | 85.34 | 85.4167
10.3 |  |  | 47.475 | 82.7165 | 82.8833 |  |  | 60.525 | 83.1 | 83.2 |  | 2.8667 | 74.8 | 83.25 | 83.3 |  | 4.4 | 79.425 | 84 | 84.1807 |  | 6.325 | 81.4 | 85.18 | 85.2333
10.4 |  |  | 39.65 | 82.6165 | 82.75 |  |  | 53.88 | 83 | 83.1 |  | 2.8 | 73.4 | 83 | 83.2 |  | 4.35 | 78.96 | 83.72 | 84.0283 |  | 6.2 | 81.0667 | 85 | 85
10.5 |  |  | 32.425 | 82.5165 | 82.7 |  |  | 46.85 | 82.94 | 83.025 |  |  | 71.32 | 82.9 | 83 |  | 4.2 | 78.6 | 83.56 | 83.814 |  | 6.1 | 80.82 | 84.85 | 84.7875
10.6 |  |  | 26.125 | 82.4765 | 82.58 |  |  | 38.9 | 82.9 | 83 |  |  | 68.725 | 82.8 | 82.9 |  | 4.1667 | 78.0333 | 83.36 | 83.4939 |  | 5.975 | 80.25 | 84.62 | 84.65
10.7 |  |  | 21.4 | 82.3365 | 82.5 |  |  | 31.98 | 82.8 | 82.9 |  |  | 65 | 82.7 | 82.8 |  | 4.1 | 77.46 | 83.2 | 83.3139 |  | 5.78 | 79.975 | 84.4 | 84.4
10.8 |  |  | 18.22 | 82.3165 | 82.4 |  |  | 25.9 | 82.7 | 82.8 |  |  | 60.2 | 82.6 | 82.7 |  | 4 | 76.78 | 82.9167 | 83.1472 |  | 5.6333 | 79.74 | 84.2 | 84.24
10.9 |  |  | 16.3 | 82.1164 | 82.3 |  |  | 21.4 | 82.6 | 82.675 |  |  | 53.8 | 82.5 | 82.6 |  | 3.9 | 76.24 | 82.75 | 82.9538 |  | 5.525 | 79.3833 | 84.075 | 83.9667
11 |  |  | 14.58 | 82.0414 | 82.14 |  |  | 18.3 | 82.5 | 82.6 |  |  | 47 | 82.4 | 82.5 |  |  | 75.5 | 82.6 | 82.7638 |  | 5.5 | 79.05 | 83.825 | 83.75
11.1 |  |  | 13.525 | 81.9164 | 82.025 |  |  | 16.15 | 82.4 | 82.5 |  |  | 39.8 | 82.3 | 82.4 |  |  | 74.7 | 82.5143 | 82.6138 |  | 5.4 | 78.58 | 83.6333 | 83.5571
11.2 |  |  | 12.675 | 81.8164 | 81.8 |  |  | 14.625 | 82.3 | 82.425 |  |  | 33.2 | 82.2 | 82.3 |  |  | 73.6 | 82.35 | 82.4471 |  | 5.2 | 78.225 | 83.42 | 83.4
11.3 |  |  | 11.9 | 81.7163 | 81.78 |  |  | 13.3 | 82.2 | 82.25 |  |  | 27.5 | 82.1167 | 82.2 |  |  | 72.5 | 82.2 | 82.3137 |  | 5.1 | 77.85 | 83.22 | 83.25
11.4 |  |  | 11.35 | 81.4913 | 81.65 |  |  | 12.4 | 82.02 | 82.125 |  |  | 22.8 | 82 | 82.1 |  |  | 71.05 | 81.9833 | 82.0937 |  | 4.95 | 77.38 | 83 | 82.98
11.5 |  |  | 10.775 | 81.3413 | 81.5 |  |  | 11.6 | 81.9 | 82 |  |  | 19.98 | 81.9 | 82 |  |  | 69.175 | 81.8286 | 81.9137 |  |  | 76.9 | 82.8 | 82.78
11.6 |  |  | 10.44 | 81.2913 | 81.375 |  |  | 11 | 81.8 | 81.9 |  |  | 17.8 | 81.775 | 81.9 |  |  | 66.8 | 81.6714 | 81.8136 |  |  | 76.4333 | 82.6333 | 82.6
11.7 |  |  | 10.075 | 81.1662 | 81.2 |  |  | 10.4667 | 81.7 | 81.775 |  |  | 16.2 | 81.625 | 81.8 |  |  | 63.5 | 81.6 | 81.7136 |  |  | 75.9 | 82.4 | 82.3667
11.8 |  |  | 9.8 | 81.0162 | 81.05 |  |  | 9.975 | 81.6 | 81.625 |  |  | 14.8 | 81.5 | 81.6 |  |  | 58.7 | 81.5 | 81.6636 |  |  | 75.35 | 82.2 | 82.25
11.9 |  |  | 9.54 | 80.8762 | 81 |  |  | 9.625 | 81.4 | 81.5 |  |  | 13.9 | 81.4 | 81.5 |  |  | 53.3 | 81.35 | 81.5136 |  |  | 74.625 | 82.08 | 81.9857
12 |  |  | 9.275 | 80.7161 | 80.825 |  |  | 9.275 | 81.3 | 81.3 |  |  | 13.1 | 81.3 | 81.4 |  |  | 46.74 | 81.2 | 81.4136 |  |  | 73.9333 | 81.86 | 81.88
12.1 |  |  | 9 | 80.5911 | 80.625 |  |  | 8.94 | 81.0667 | 81.1 |  |  | 12.4 | 81.1 | 81.25 |  |  | 40.54 | 81.1333 | 81.3303 |  |  | 73.1 | 81.7 | 81.6833
12.2 |  |  | 8.85 | 80.4161 | 80.45 |  |  | 8.7 | 80.9 | 81 |  |  | 11.8 | 81.025 | 81.1 |  |  | 34.54 | 81 | 81.1535 |  |  | 72.025 | 81.55 | 81.5167
12.3 |  |  | 8.7 | 80.216 | 80.34 |  |  | 8.45 | 80.8 | 80.9 |  |  | 11.3 | 80.9 | 80.95 |  |  | 29 | 80.8714 | 81.0135 |  |  | 71 | 81.36 | 81.36
12.4 |  |  | 8.46 | 80.091 | 80.2 |  |  | 8.2 | 80.7 | 80.78 |  |  | 10.86 | 80.7 | 80.7 |  |  | 25.3 | 80.8 | 80.9135 |  |  | 69.44 | 81.2 | 81.2
12.5 |  |  | 8.34 | 79.936 | 79.975 |  |  | 8.025 | 80.5 | 80.55 |  |  | 10.575 | 80.54 | 80.6 |  |  | 22.3 | 80.6429 | 80.8302 |  |  | 67.85 | 81.02 | 81.05
12.6 |  |  | 8.2 | 79.816 | 79.85 |  |  | 7.8 | 80.35 | 80.4 |  |  | 10.1 | 80.425 | 80.45 |  |  | 20.24 | 80.5 | 80.7635 |  |  | 65.4 | 80.92 | 80.9
12.7 |  |  | 8 | 79.6759 | 79.65 |  |  | 7.625 | 80.22 | 80.3 |  |  | 9.85 | 80.2 | 80.3 |  |  | 18.72 | 80.4 | 80.4967 |  |  | 62.2 | 80.8 | 80.8667
12.8 |  |  | 7.85 | 79.4959 | 79.5 |  |  | 7.475 | 80.08 | 80.2 |  |  | 9.6333 | 80.05 | 80.075 |  |  | 17.5 | 80.25 | 80.3134 |  |  | 58 | 80.72 | 80.7143
12.9 |  |  | 7.7 | 79.3159 | 79.3333 |  |  | 7.3 | 79.8667 | 80 |  |  | 9.38 | 79.925 | 79.925 |  |  | 16.6 | 80.0833 | 80.2334 |  |  | 52.8 | 80.6 | 80.5333
13 |  |  | 7.6 | 79.2492 | 79.2 |  |  | 7.2 | 79.7 | 79.8 |  |  | 9.15 | 79.76 | 79.75 |  |  | 15.725 | 79.9167 | 80.0705 |  |  | 47.375 | 80.45 | 80.4
13.1 |  |  | 7.5 | 78.8758 | 79 |  |  | 7.025 | 79.6 | 79.6 |  |  | 8.9 | 79.5333 | 79.64 |  |  | 14.92 | 79.7667 | 79.9133 |  |  | 41.55 | 80.3 | 80.3
13.2 |  |  | 7.3 | 78.7157 | 78.7667 |  |  | 6.875 | 79.375 | 79.4 |  |  | 8.7 | 79.4 | 79.5 |  |  | 14.35 | 79.58 | 79.7533 |  |  | 35.92 | 80.2 | 80.2
13.3 |  |  | 7.2 | 78.5157 | 78.6 |  |  | 6.8 | 79.225 | 79.3 |  |  | 8.5 | 79.2 | 79.3 |  |  | 13.85 | 79.4 | 79.5704 |  |  | 31.1 | 80.1 | 80.1
13.4 |  |  | 7.075 | 78.3407 | 78.3333 |  |  | 6.6 | 78.95 | 79.125 |  |  | 8.3 | 79.1 | 79.2 |  |  | 13.425 | 79.3 | 79.4132 |  |  | 27.5667 | 80 | 79.98
13.5 |  |  | 7 | 78.1356 | 78.175 |  |  | 6.525 | 78.8 | 78.9 |  |  | 8.1667 | 78.92 | 78.925 |  |  | 13.05 | 79.0667 | 79.2561 |  |  | 24.64 | 79.9 | 79.8
13.6 |  |  | 6.9 | 77.9656 | 78 |  |  | 6.4 | 78.6 | 78.625 |  |  | 8 | 78.68 | 78.775 |  |  | 12.675 | 78.8167 | 79.0965 |  |  | 22.5 | 79.725 | 79.7
13.7 |  |  | 6.8 | 77.7355 | 77.8 |  |  | 6.3 | 78.4 | 78.4 |  |  | 7.82 | 78.5333 | 78.525 |  |  | 12.34 | 78.7667 | 78.8703 |  |  | 20.82 | 79.6 | 79.5833
13.8 |  |  | 6.68 | 77.6155 | 77.625 |  |  | 6.22 | 78.15 | 78.26 |  |  | 7.68 | 78.3 | 78.3 |  |  | 12.02 | 78.6 | 78.6931 |  |  | 19.7 | 79.5 | 79.4375
13.9 |  |  | 6.58 | 77.3655 | 77.35 |  |  | 6.1 | 77.9 | 78 |  |  | 7.5 | 78.1 | 78.2 |  |  | 11.675 | 78.34 | 78.5131 |  |  | 18.6 | 79.25 | 79.2
14 |  |  | 6.46 | 77.0154 | 77 |  |  | 6 | 77.64 | 77.725 |  |  | 7.4 | 77.925 | 78 |  |  | 11.425 | 78.1833 | 78.3131 |  |  | 17.76 | 79.15 | 79.0333
14.1 |  |  | 6.325 | 76.682 | 76.75 |  |  | 5.9 | 77.52 | 77.55 |  |  | 7.3 | 77.65 | 77.675 |  |  | 11.18 | 78 | 78.0963 |  |  | 17.1 | 78.95 | 78.8714
14.2 |  |  | 6.2 | 76.4153 | 76.425 |  |  | 5.8 | 77.1 | 77.2 |  |  | 7.1 | 77.42 | 77.5 |  |  | 10.96 | 77.8286 | 77.913 |  |  | 16.5 | 78.725 | 78.675
14.3 |  |  | 6.15 | 76.0152 | 75.9667 |  |  | 5.75 | 76.8333 | 76.95 |  |  | 7 | 77.175 | 77.3 |  |  | 10.775 | 77.58 | 77.833 |  |  | 15.82 | 78.575 | 78.5125
14.4 |  |  | 6.1 | 75.5151 | 75.55 |  |  | 5.7 | 76.5 | 76.5 |  |  | 6.88 | 76.95 | 77 |  |  | 10.55 | 77.4143 | 77.5796 |  |  | 15.4 | 78.35 | 78.34
14.5 |  |  | 6 | 75.065 | 75.05 |  |  | 5.54 | 76.1 | 76.15 |  |  | 6.8 | 76.7 | 76.8 |  |  | 10.325 | 77.2143 | 77.3129 |  |  | 15 | 78.22 | 78.16
14.6 |  |  | 5.9 | 74.4149 | 74.35 |  |  | 5.5 | 75.58 | 75.7333 |  |  | 6.64 | 76.4667 | 76.45 |  |  | 10.1 | 77.04 | 77.1296 |  |  | 14.58 | 78.1 | 78
14.7 |  |  | 5.8 | 73.6147 | 73.6 |  |  | 5.4 | 75.1 | 75.1 |  |  | 6.6 | 76.2 | 76.2 |  |  | 9.95 | 76.9 | 76.9328 |  |  | 14.2333 | 77.82 | 77.8667
14.8 |  |  | 5.7 | 72.6145 | 72.575 |  |  | 5.3 | 74.6 | 74.7 |  |  | 6.5 | 75.925 | 75.95 |  |  | 9.75 | 76.7 | 76.7528 |  |  | 13.95 | 77.7 | 77.6667
14.9 |  |  | 5.625 | 71.1142 | 71.28 |  |  | 5.24 | 73.8 | 73.9 |  |  | 6.4 | 75.6 | 75.6 |  |  | 9.62 | 76.4167 | 76.4927 |  |  | 13.6 | 77.4 | 77.4167
15 |  |  | 5.6 | 69.1638 | 69.35 |  |  | 5.2 | 72.86 | 73 |  |  | 6.3 | 75.25 | 75.325 |  |  | 9.475 | 76.15 | 76.3794 |  |  | 13.35 | 77.2 | 77.24
15.1 |  |  | 5.5 | 66.5133 | 66.8 |  |  | 5.1 | 71.8 | 71.85 |  |  | 6.15 | 74.85 | 74.95 |  |  | 9.325 | 75.9857 | 76.1127 |  |  | 13.1 | 77 | 77.04
15.2 |  |  | 5.425 | 62.0624 | 62.36 |  |  | 5 | 70.28 | 70.48 |  |  | 6.1 | 74.475 | 74.45 |  |  | 9.2 | 75.74 | 75.846 |  |  | 12.85 | 76.9 | 76.8833
15.3 |  |  |  | 56.5447 | 57.04 |  |  | 4.925 | 68.45 | 68.475 |  |  | 6 | 74.0333 | 74.1 |  |  | 8.98 | 75.4625 | 75.5526 |  |  | 12.6 | 76.7 | 76.7
15.4 |  |  |  | 50.3101 | 50.3 |  |  | 4.9 | 65.52 | 65.7 |  |  | 5.9 | 73.45 | 73.625 |  |  | 8.8 | 75.2333 | 75.4126 |  |  | 12.45 | 76.5 | 76.5
15.5 |  |  |  | 42.9086 | 43 |  |  |  | 61.6 | 61.725 |  |  | 5.8 | 72.9 | 73 |  |  | 8.7 | 75.0333 | 75.1792 |  |  | 12.225 | 76.3 | 76.24
15.6 |  |  |  | 36.1072 | 36.8 |  |  |  | 55.9333 | 56.475 |  |  | 5.7 | 72.175 | 72.46 |  |  | 8.5 | 74.7833 | 74.8458 |  |  | 11.975 | 76.08 | 76.0429
15.7 |  |  |  | 29.806 | 30.4 |  |  |  | 49.2 | 49.8 |  |  | 5.6667 | 71.45 | 71.5 |  |  | 8.3 | 74.5333 | 74.5838 |  |  | 11.86 | 75.9 | 75.8
15.8 |  |  |  | 24.7049 | 25.84 |  |  |  | 42.3 | 43.05 |  |  | 5.6 | 70.4 | 70.5 |  |  | 8.2 | 74.25 | 74.3695 |  |  | 11.54 | 75.7 | 75.7
15.9 |  |  |  | 21.4043 | 22.4 |  |  |  | 36.0667 | 36.7 |  |  | 5.475 | 68.9667 | 69.25 |  |  | 8.1 | 73.98 | 74.0724 |  |  | 11.34 | 75.48 | 75.4333
16 |  |  |  | 18.9038 | 20.0667 |  |  |  | 29.8 | 30.9 |  |  |  | 67.4 | 67.58 |  |  | 7.925 | 73.7286 | 73.8323 |  |  | 11.2 | 75.375 | 75.2833
16.1 |  |  |  | 17.3535 | 18.4 |  |  |  | 24.6 | 25.8 |  |  |  | 65.1 | 65.25 |  |  | 7.8 | 73.4333 | 73.4789 |  |  | 11.08 | 75.12 | 75.0571
16.2 |  |  |  | 16.1232 | 17.2667 |  |  |  | 21.3333 | 22.2 |  |  |  | 62.06 | 61.95 |  |  | 7.7 | 73.2 | 73.2408 |  |  | 10.9 | 74.8 | 74.88
16.3 |  |  |  | 15.223 | 16.4 |  |  |  | 18.9 | 20 |  |  |  | 57.56 | 57.7 |  |  | 7.6 | 72.8 | 72.9121 |  |  | 10.74 | 74.62 | 74.5667
16.4 |  |  |  | 14.5029 | 15.65 |  |  |  | 17.25 | 18.275 |  |  |  | 51.7 | 51.925 |  |  | 7.5 | 72.4667 | 72.5721 |  |  | 10.6 | 74.4 | 74.3714
16.5 |  |  |  | 13.9028 | 15.1 |  |  |  | 15.92 | 16.9 |  |  |  | 45.48 | 46.1 |  |  | 7.4 | 72.22 | 72.2406 |  |  | 10.3 | 74.15 | 74.1333
16.6 |  |  |  | 13.4627 | 14.6333 |  |  |  | 14.925 | 16.05 |  |  |  | 39.3 | 40.02 |  |  | 7.275 | 71.78 | 71.852 |  |  | 10.22 | 73.9333 | 73.875
16.7 |  |  |  | 13.1226 | 14.3 |  |  |  | 14.1 | 15.225 |  |  |  | 33.575 | 33.9 |  |  | 7.2 | 71.3 | 71.5119 |  |  | 10.1 | 73.7 | 73.6
16.8 |  |  |  | 12.8026 | 13.9667 |  |  |  | 13.5 | 14.65 |  |  |  | 28.25 | 28.925 |  |  | 7.1 | 70.86 | 70.8976 |  |  | 10 | 73.4333 | 73.3833
16.9 |  |  |  | 12.4275 | 13.675 |  |  |  | 12.9667 | 14.1 |  |  |  | 23.9 | 25 |  |  | 7 | 70.25 | 70.4117 |  |  | 9.8 | 73.2 | 73.2286
17 |  |  |  | 12.2024 | 13.4 |  |  |  | 12.6 | 13.6667 |  |  |  | 21.3 | 22.375 |  |  |  | 69.6667 | 69.745 |  |  | 9.65 | 72.8667 | 72.9333
17.1 |  |  |  | 12.0024 | 13.1667 |  |  |  | 12.2 | 13.28 |  |  |  | 19.24 | 20.38 |  |  |  | 68.975 | 69.1449 |  |  | 9.5 | 72.7 | 72.6714
17.2 |  |  |  | 11.8024 | 13 |  |  |  | 11.9333 | 13.025 |  |  |  | 17.85 | 18.925 |  |  |  | 68.1833 | 68.4114 |  |  | 9.3 | 72.36 | 72.3667
17.3 |  |  |  | 11.6823 | 12.8 |  |  |  | 11.65 | 12.72 |  |  |  | 16.8 | 17.8 |  |  |  | 67.2 | 67.4112 |  |  | 9.2 | 72.1 | 72
17.4 |  |  |  | 11.5023 | 12.675 |  |  |  | 11.3333 | 12.425 |  |  |  | 15.8 | 16.875 |  |  |  | 65.9 | 66.111 |  |  | 9.1 | 71.825 | 71.8
17.5 |  |  |  | 11.3023 | 12.5 |  |  |  | 11.125 | 12.2 |  |  |  | 15.15 | 16.24 |  |  |  | 64.3 | 64.4774 |  |  | 8.94 | 71.575 | 71.48
17.6 |  |  |  | 11.2022 | 12.4 |  |  |  | 10.9 | 12.0333 |  |  |  | 14.475 | 15.65 |  |  |  | 62.1333 | 62.2604 |  |  | 8.8 | 71.3 | 71.2
17.7 |  |  |  | 11.0522 | 12.2 |  |  |  | 10.8 | 11.875 |  |  |  | 14.05 | 15.08 |  |  |  | 59.3 | 59.3099 |  |  | 8.68 | 70.975 | 70.9
17.8 |  |  |  | 10.9022 | 12.075 |  |  |  | 10.675 | 11.66 |  |  |  | 13.56 | 14.7 |  |  |  | 55.22 | 55.3092 |  |  | 8.5667 | 70.55 | 70.5333
17.9 |  |  |  | 10.8022 | 11.9 |  |  |  | 10.42 | 11.525 |  |  |  | 13.2 | 14.3667 |  |  |  | 50.4 | 50.8751 |  |  | 8.44 | 70.2 | 70.2286
18 |  |  |  | 10.7021 | 11.8333 |  |  |  | 10.3 | 11.35 |  |  |  | 12.9167 | 14 |  |  |  | 44.7333 | 45.2742 |  |  |  | 69.8 | 69.88
18.1 |  |  |  | 10.6021 | 11.7 |  |  |  | 10.2 | 11.2 |  |  |  | 12.6 | 13.65 |  |  |  | 39.28 | 40.1067 |  |  |  | 69.46 | 69.3714
18.2 |  |  |  | 10.5021 | 11.6 |  |  |  | 10.1 | 11.15 |  |  |  | 12.4 | 13.4 |  |  |  | 34.15 | 34.9058 |  |  |  | 69.125 | 69
18.3 |  |  |  | 10.4021 | 11.5 |  |  |  | 9.9 | 11 |  |  |  | 12.1 | 13.18 |  |  |  | 29.5 | 30.4051 |  |  |  | 68.54 | 68.55
18.4 |  |  |  | 10.3021 | 11.4 |  |  |  | 9.8 | 10.9 |  |  |  | 11.88 | 12.95 |  |  |  | 25.9143 | 26.9378 |  |  |  | 67.98 | 67.9
18.5 |  |  |  | 10.202 | 11.3 |  |  |  | 9.7 | 10.75 |  |  |  | 11.7 | 12.78 |  |  |  | 23.45 | 24.4041 |  |  |  | 67.3333 | 67.4
18.6 |  |  |  | 10.102 | 11.2 |  |  |  | 9.6 | 10.7 |  |  |  | 11.58 | 12.6 |  |  |  | 21.4333 | 22.4913 |  |  |  | 66.725 | 66.7333
18.7 |  |  |  | 10.002 | 11.1 |  |  |  | 9.5 | 10.6 |  |  |  | 11.35 | 12.4 |  |  |  | 20.15 | 21.0035 |  |  |  | 65.84 | 65.9
18.8 |  |  |  | 9.927 | 11.025 |  |  |  | 9.4 | 10.5 |  |  |  | 11.2 | 12.275 |  |  |  | 19 | 20.0034 |  |  |  | 64.8 | 64.8
18.9 |  |  |  | 9.877 | 10.9 |  |  |  | 9.35 | 10.4 |  |  |  | 11 | 12.1 |  |  |  | 18 | 19.2032 |  |  |  | 63.66 | 63.6
19 |  |  |  | 9.7269 | 10.85 |  |  |  | 9.3 | 10.3 |  |  |  | 10.9 | 11.9 |  |  |  | 17.4 | 18.423 |  |  |  | 61.96 | 62.0429
19.1 |  |  |  | 9.6619 | 10.775 |  |  |  | 9.2 | 10.2 |  |  |  | 10.775 | 11.75 |  |  |  | 16.8 | 17.803 |  |  |  | 59.76 | 59.9
19.2 |  |  |  | 9.5269 | 10.6 |  |  |  | 9.1 | 10.1 |  |  |  | 10.6 | 11.6 |  |  |  | 16.2714 | 17.3362 |  |  |  | 57.08 | 57.22
19.3 |  |  |  | 9.5019 | 10.55 |  |  |  | 9 | 10 |  |  |  | 10.5 | 11.55 |  |  |  | 15.8 | 16.9028 |  |  |  | 53.325 | 53.8
19.4 |  |  |  | 9.4019 | 10.5 |  |  |  | 8.9 | 10 |  |  |  | 10.3 | 11.4 |  |  |  | 15.5 | 16.5599 |  |  |  | 49.025 | 49.2
19.5 |  |  |  | 9.3019 | 10.375 |  |  |  | 8.825 | 9.9 |  |  |  | 10.225 | 11.3 |  |  |  | 15.1667 | 16.2527 |  |  |  | 44.0333 | 44.7714
19.6 |  |  |  | 9.2018 | 10.3 |  |  |  | 8.8 | 9.8 |  |  |  | 10.1 | 11.2 |  |  |  | 14.78 | 15.9027 |  |  |  | 39.28 | 39.7714
19.7 |  |  |  | 9.1018 | 10.1667 |  |  |  | 8.7 | 9.7 |  |  |  | 10 | 11.1 |  |  |  | 14.4667 | 15.6193 |  |  |  | 34.6333 | 35.12
19.8 |  |  |  | 9.0518 | 10.1 |  |  |  | 8.6 | 9.6 |  |  |  | 9.9 | 10.975 |  |  |  | 14.3 | 15.3359 |  |  |  | 30.5 | 31.48
19.9 |  |  |  | 8.9018 | 10 |  |  |  | 8.5 | 9.5 |  |  |  | 9.8 | 10.9 |  |  |  | 14 | 15.0525 |  |  |  | 27.475 | 28.2143
20 |  |  |  | 8.8268 | 9.9 |  |  |  | 8.4 | 9.45 |  |  |  | 9.7 | 10.725 |  |  |  | 13.8 | 14.8625 |  |  |  | 24.925 | 25.9
20.1 |  |  |  | 8.7017 | 9.76 |  |  |  | 8.3 | 9.4 |  |  |  | 9.6 | 10.66 |  |  |  | 13.5333 | 14.5857 |  |  |  | 23.175 | 24.1333
20.2 |  |  |  | 8.6267 | 9.625 |  |  |  | 8.3 | 9.3 |  |  |  | 9.5 | 10.525 |  |  |  | 13.35 | 14.4024 |  |  |  | 21.9 | 22.76
20.3 |  |  |  | 8.5817 | 9.6 |  |  |  | 8.2 | 9.2 |  |  |  | 9.4 | 10.48 |  |  |  | 13.1 | 14.231 |  |  |  | 20.7 | 21.74
20.4 |  |  |  | 8.4017 | 9.475 |  |  |  | 8.1 | 9.175 |  |  |  | 9.325 | 10.35 |  |  |  | 12.975 | 14.0624 |  |  |  | 19.9 | 20.9333
20.5 | [8 empty] | 8.04 | 9.025 |  |  |  | 9.2 | 10.26 |  |  |  | 12.7833 | 13.8856 |  |  |  | 19.1 | 20.3
20.6 | [8 empty] | 7.9 | 8.9 |  |  |  | 9.1 | 10.2 |  |  |  | 12.5 | 13.719 |  |  |  | 18.54 | 19.68
20.7 | [8 empty] | 7.825 | 8.8 |  |  |  | 9.075 | 10.1 |  |  |  | 12.4 | 13.5356 |  |  |  | 18.0667 | 19.05
20.8 | [13 empty] | 8.9 | 10 |  |  |  | 12.2167 | 13.3422 |  |  |  | 17.66 | 18.575
20.9 | [13 empty] | 8.9 | 9.875 |  |  |  | 12.0333 | 13.2022 |  |  |  | 17.25 | 18.3143
21 | [13 empty] | 8.8 | 9.8 |  |  |  | 11.9 | 13.0422 |  |  |  | 16.88 | 17.9
21.1 | [13 empty] | 8.7 | 9.7 |  |  |  | 11.8 | 12.9021 |  |  |  | 16.6 | 17.64
21.2 | [13 empty] | 8.6 | 9.6 |  |  |  | 11.7 | 12.7521 |  |  |  | 16.24 | 17.32
21.3 | [13 empty] | 8.5 | 9.5 |  |  |  | 11.5 | 12.6021 |  |  |  | 15.92 | 17.04
21.4 | [13 empty] | 8.4 | 9.45 |  |  |  | 11.4333 | 12.4221 |  |  |  | 15.7 | 16.85
21.5 | [18 empty] | 11.24 | 12.3021 |  |  |  | 15.5 | 16.5667
21.6 | [18 empty] | 11.1571 | 12.202 |  |  |  | 15.3 | 16.3667
21.7 | [18 empty] | 11 | 12.0687 |  |  |  | 15.1 | 16.1167
21.8 | [18 empty] | 10.88 | 11.9163 |  |  |  | 14.875 | 15.8857
21.9 | [18 empty] | 10.7857 | 11.852 |  |  |  | 14.7 | 15.6667
22 | [18 empty] | 10.62 | 11.682 |  |  |  | 14.5333 | 15.48
22.1 | [18 empty] | 10.55 | 11.5819 |  |  |  | 14.2 | 15.3
22.2 | [18 empty] | 10.48 | 11.4876 |  |  |  | 14.025 | 15.1333
22.3 | [18 empty] | 10.3 | 11.4019 |  |  |  | 13.875 | 14.94
22.4 | [18 empty] | 10.2 | 11.2819 |  |  |  | 13.7 | 14.74
22.5 | [18 empty] | 10.08 | 11.1019 |  |  |  | 13.575 | 14.6
22.6 | [18 empty] | 9.9143 | 10.9818 |  |  |  | 13.4 | 14.46
22.7 | [18 empty] | 9.84 | 10.9018 |  |  |  | 13.22 | 14.4
22.8 | [18 empty] | 9.7333 | 10.8018 |  |  |  | 13.1 | 14.2333
22.9 | [18 empty] | 9.66 | 10.7018 |  |  |  | 12.9 | 14.02

## Wedge 30 Longitudinal Profiles
Longitudinal Profiles
Wedge 30
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:3 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 12.725
-30.8 |  |  |  |  | 12.85
-30.7 |  |  |  |  | 12.975
-30.6 |  |  |  |  | 13.1
-30.5 |  |  |  |  | 13.225
-30.4 |  |  |  |  | 13.4375
-30.3 |  |  |  |  | 13.625
-30.2 |  |  |  |  | 13.7
-30.1 |  |  |  |  | 13.8286
-30 |  |  |  |  | 13.9714
-29.9 |  |  |  |  | 14.1333
-29.8 |  |  |  |  | 14.2857
-29.7 |  |  |  |  | 14.48
-29.6 |  |  |  |  | 14.6
-29.5 |  |  |  |  | 14.7857
-29.4 |  |  |  |  | 14.9
-29.3 |  |  |  |  | 15.05
-29.2 |  |  |  |  | 15.25
-29.1 |  |  |  |  | 15.4
-29 |  |  |  |  | 15.66
-28.9 |  |  |  | 11.2 | 15.825
-28.8 |  |  |  | 11.2571 | 16
-28.7 |  |  |  | 11.3875 | 16.1857
-28.6 |  |  |  | 11.6125 | 16.45
-28.5 |  |  |  | 11.7375 | 16.5667
-28.4 |  |  |  | 11.8 | 16.7857
-28.3 |  |  |  | 11.8875 | 17.025
-28.2 |  |  |  | 12.0286 | 17.25
-28.1 |  |  |  | 12.1714 | 17.4286
-28 |  |  |  | 12.3143 | 17.7
-27.9 |  |  |  | 12.45 | 17.9714
-27.8 |  |  |  | 12.5 | 18.2571
-27.7 |  |  |  | 12.725 | 18.5
-27.6 |  |  |  | 12.85 | 18.8667
-27.5 |  |  |  | 12.9857 | 19.1667
-27.4 |  |  |  | 13.1375 | 19.6
-27.3 |  |  |  | 13.2833 | 20.0714
-27.2 |  |  |  | 13.425 | 20.4
-27.1 |  |  |  | 13.6375 | 21.02
-27 |  |  |  | 13.7625 | 21.7143
-26.9 |  |  | 9.6 | 13.975 | 22.4286
-26.8 |  |  | 9.7 | 14.2 | 23.3143
-26.7 |  |  | 9.8 | 14.325 | 24.48
-26.6 |  |  | 9.8833 | 14.5143 | 25.7625
-26.5 |  |  | 10 | 14.7222 | 27.5
-26.4 |  |  | 10.1 | 14.95 | 29.6875
-26.3 |  |  | 10.2 | 15.2 | 32.3
-26.2 |  |  | 10.2857 | 15.3857 | 35.8
-26.1 |  |  | 10.4 | 15.64 | 39.9
-26 |  |  | 10.4833 | 15.8 | 44.7
-25.9 |  | 8.8875 | 10.6 | 16.1444 | 50
-25.8 |  | 8.9667 | 10.7 | 16.375 | 55.44
-25.7 |  | 9.0286 | 10.8 | 16.625 | 60.7571
-25.6 |  | 9.1 | 10.9 | 16.9625 | 64.9833
-25.5 | 9.4 | 9.2 | 11.0571 | 17.3667 | 69.0778
-25.4 | 9.4833 | 9.2714 | 11.2 | 17.7 | 71.8125
-25.3 | 9.5 | 9.3167 | 11.3 | 18.1714 | 74.1429
-25.2 | 9.5667 | 9.4714 | 11.4 | 18.7 | 75.7571
-25.1 | 9.7 | 9.5286 | 11.5429 | 19.2667 | 77.34
-25 | 9.7833 | 9.6 | 11.7125 | 19.9 | 78.45
-24.9 | 9.85 | 9.7286 | 11.85 | 20.8 | 79.4
-24.8 | 10 | 9.8 | 12.0333 | 21.7143 | 80.2
-24.7 | 10.0833 | 9.9 | 12.1857 | 22.95 | 80.8
-24.6 | 10.1375 | 10 | 12.3333 | 24.4 | 81.475
-24.5 | 10.2833 | 10.06 | 12.55 | 26.3333 | 82
-24.4 | 10.3286 | 10.14 | 12.7 | 28.9 | 82.5286
-24.3 | 10.4 | 10.3 | 12.8667 | 32.1 | 82.9
-24.2 | 10.5333 | 10.3714 | 13.0857 | 36.3875 | 83.36
-24.1 | 10.6 | 10.5 | 13.2667 | 41.2875 | 83.78
-24 | 10.7 | 10.6 | 13.4571 | 46.8875 | 84.1375
-23.9 | 10.8286 | 10.7 | 13.7 | 52.8375 | 84.46
-23.8 | 10.9 | 10.85 | 13.975 | 59.1857 | 84.8667
-23.7 | 11 | 10.9333 | 14.2429 | 64.6571 | 85.15
-23.6 | 11.1286 | 11.1 | 14.44 | 69.3714 | 85.4286
-23.5 | 11.2833 | 11.2667 | 14.8571 | 73.2 | 85.75
-23.4 | 11.4 | 11.4 | 15.15 | 76.4571 | 86.05
-23.3 | 11.4714 | 11.5714 | 15.65 | 78.8 | 86.2167
-23.2 | 11.6125 | 11.7125 | 16.2 | 80.5 | 86.4
-23.1 | 11.76 | 11.85 | 16.7857 | 81.8 | 86.7
-23 | 11.9333 | 12.1 | 17.4 | 82.8 | 86.8875
-22.9 | 12.1 | 12.3 | 18.25 | 83.775 | 87.05
-22.8 | 12.2833 | 12.5429 | 19.3 | 84.525 | 87.2
-22.7 | 12.45 | 12.7286 | 20.5 | 85.15 | 87.4333
-22.6 | 12.6143 | 13.0143 | 22.12 | 85.8375 | 87.5857
-22.5 | 12.7571 | 13.3 | 24.2286 | 86.2714 | 87.76
-22.4 | 13.0125 | 13.6286 | 26.9 | 86.7571 | 87.9
-22.3 | 13.1857 | 14 | 30.7667 | 87.1667 | 88.05
-22.2 | 13.4 | 14.5286 | 35.9667 | 87.6333 | 88.25
-22.1 | 13.68 | 15.0429 | 41.6 | 88.05 | 88.4
-22 | 14 | 15.7571 | 48.56 | 88.4333 | 88.7125
-21.9 | 14.3286 | 16.625 | 55.7 | 88.65 | 88.8375
-21.8 | 14.8 | 17.7333 | 62.9429 | 88.95 | 88.9833
-21.7 | 15.3143 | 18.9714 | 69.6 | 89.2333 | 89.2333
-21.6 | 15.9571 | 20.6 | 75.325 | 89.6 | 89.4
-21.5 | 16.7714 | 23.0714 | 79.5 | 89.85 | 89.6667
-21.4 | 17.8 | 26.25 | 82.5286 | 90 | 89.82
-21.3 | 19 | 30.5 | 84.7 | 90.175 | 89.9
-21.2 | 20.6286 | 36.0571 | 86.3571 | 90.3625 | 90.0571
-21.1 | 22.9143 | 42.8571 | 87.65 | 90.5 | 90.3
-21 | 26.02 | 50.1143 | 88.7 | 90.6571 | 90.45
-20.9 | 30.5 | 57.7 | 89.575 | 90.7875 | 90.7
-20.8 | 36.14 | 65.9167 | 90.28 | 90.8857 | 90.875
-20.7 | 42.8167 | 73.1 | 90.9571 | 91.1 | 91.1
-20.6 | 50.8375 | 78.8143 | 91.4714 | 91.25 | 91.2571
-20.5 | 59.1 | 83.225 | 91.94 | 91.3857 | 91.4625
-20.4 | 67.4667 | 86.15 | 92.4 | 91.55 | 91.7
-20.3 | 74.3667 | 88.5333 | 92.8667 | 91.75 | 91.8143
-20.2 | 80.45 | 89.9857 | 93.2 | 91.95 | 91.9875
-20.1 | 84.5571 | 91.2857 | 93.45 | 92.075 | 92.1
-20 | 87.5 | 92.1429 | 93.6571 | 92.1889 | 92.2
-19.9 | 89.4857 | 92.925 | 93.9167 | 92.4286 | 92.42
-19.8 | 90.9333 | 93.55 | 94.1429 | 92.5 | 92.6
-19.7 | 92.0286 | 94 | 94.3333 | 92.75 | 92.7714
-19.6 | 92.9 | 94.5714 | 94.5 | 92.9 | 92.9429
-19.5 | 93.675 | 95.0667 | 94.6667 | 93.1429 | 93.075
-19.4 | 94.3429 | 95.3714 | 94.8143 | 93.3286 | 93.2
-19.3 | 94.7143 | 95.7 | 94.9667 | 93.5571 | 93.35
-19.2 | 95.22 | 96.0333 | 95.1 | 93.6 | 93.5
-19.1 | 95.5333 | 96.3 | 95.2 | 93.85 | 93.6571
-19 | 95.7 | 96.48 | 95.3 | 94 | 93.8
-18.9 | 95.96 | 96.6286 | 95.3714 | 94.1 | 93.9
-18.8 | 96.24 | 96.8 | 95.52 | 94.3125 | 93.9778
-18.7 | 96.4333 | 96.88 | 95.7 | 94.45 | 94.1
-18.6 | 96.6167 | 96.9333 | 95.7833 | 94.5 | 94.15
-18.5 | 96.7714 | 97.0857 | 95.85 | 94.6667 | 94.3
-18.4 | 96.8 | 97.1 | 96.0167 | 94.8 | 94.45
-18.3 | 96.9 | 97.1857 | 96.1833 | 95 | 94.5
-18.2 | 97 | 97.2167 | 96.2167 | 95.075 | 94.5833
-18.1 | 97.025 | 97.3714 | 96.4429 | 95.2333 | 94.8
-18 | 97.1 | 97.42 | 96.5333 | 95.3857 | 94.9
-17.9 | 97.1 | 97.5 | 96.6857 | 95.5333 | 94.9375
-17.8 | 97.18 | 97.5 | 96.84 | 95.675 | 95.0714
-17.7 | 97.2 | 97.6 | 96.9 | 95.8125 | 95.1333
-17.6 | 97.3 | 97.6667 | 97.06 | 95.9 | 95.275
-17.5 | 97.3833 | 97.8857 | 97.16 | 95.975 | 95.3889
-17.4 | 97.4167 | 98 | 97.4 | 96.1 | 95.5
-17.3 | 97.5 | 98.05 | 97.5286 | 96.2 | 95.55
-17.2 | 97.6 | 98.2167 | 97.6833 | 96.2571 | 95.7125
-17.1 | 97.7143 | 98.3167 | 97.75 | 96.4 | 95.8429
-17 | 97.8 | 98.5 | 97.9 | 96.5 | 95.975
-16.9 | 97.9333 | 98.6 | 97.9714 | 96.5833 | 96.1
-16.8 | 98.02 | 98.7 | 98.1 | 96.6 | 96.1667
-16.7 | 98.1 | 98.8 | 98.1667 | 96.6222 | 96.3
-16.6 | 98.25 | 98.8857 | 98.25 | 96.7375 | 96.38
-16.5 | 98.3875 | 99 | 98.3286 | 96.8625 | 96.4375
-16.4 | 98.5 | 99.1167 | 98.4 | 96.9143 | 96.625
-16.3 | 98.6 | 99.3 | 98.5 | 97.0571 | 96.7875
-16.2 | 98.6667 | 99.4 | 98.6 | 97.1875 | 96.9125
-16.1 | 98.85 | 99.4 | 98.7 | 97.3 | 97
-16 | 99 | 99.5222 | 98.75 | 97.3571 | 97.0857
-15.9 | 99.05 | 99.6 | 98.8143 | 97.4 | 97.2
-15.8 | 99.1 | 99.6 | 98.9 | 97.6 | 97.2286
-15.7 | 99.25 | 99.7 | 99 | 97.6714 | 97.3
-15.6 | 99.3 | 99.8 | 99 | 97.7 | 97.5
-15.5 | 99.4 | 99.8 | 99 | 97.85 | 97.5667
-15.4 | 99.5 | 99.9 | 99.1 | 98 | 97.7
-15.3 | 99.5 | 99.9 | 99.1 | 98.05 | 97.8
-15.2 | 99.5167 | 99.9 | 99.2 | 98.2 | 97.85
-15.1 | 99.6 | 99.9 | 99.2429 | 98.3222 | 97.975
-15 | 99.6 | 100 | 99.3 | 98.4 | 98
-14.9 | 99.6571 | 100.075 | 99.3833 | 98.5 | 98.1
-14.8 | 99.7 | 100.1167 | 99.5 | 98.5714 | 98.1571
-14.7 | 99.8 | 100.2 | 99.6 | 98.7 | 98.3125
-14.6 | 99.9 | 100.2125 | 99.6833 | 98.8 | 98.4
-14.5 | 99.9 | 100.3 | 99.7 | 98.8429 | 98.4625
-14.4 | 99.9 | 100.3 | 99.8 | 98.9 | 98.675
-14.3 | 99.9 | 100.3625 | 99.88 | 99.0714 | 98.7
-14.2 | 99.9 | 100.4 | 99.9 | 99.1 | 98.8
-14.1 | 99.9833 | 100.4429 | 100 | 99.2 | 98.8857
-14 | 100.0571 | 100.5 | 100.06 | 99.2 | 98.98
-13.9 | 100.1125 | 100.6 | 100.2 | 99.3 | 99.0333
-13.8 | 100.2 | 100.6 | 100.3 | 99.3 | 99.1
-13.7 | 100.24 | 100.7 | 100.3 | 99.4 | 99.1
-13.6 | 100.3 | 100.8 | 100.4 | 99.5 | 99.1429
-13.5 | 100.3833 | 100.85 | 100.5 | 99.6 | 99.2
-13.4 | 100.4333 | 100.9167 | 100.5571 | 99.6571 | 99.2889
-13.3 | 100.5 | 101 | 100.6 | 99.8 | 99.3
-13.2 | 100.5714 | 101.1 | 100.6833 | 99.8714 | 99.3875
-13.1 | 100.6333 | 101.1 | 100.7 | 99.9125 | 99.4
-13 | 100.775 | 101.2375 | 100.8 | 100 | 99.3
-12.9 | 100.8 | 101.3 | 100.82 | 100 | 99.4
-12.8 | 100.8857 | 101.3 | 100.9 | 100.1 | 99.4
-12.7 | 100.925 | 101.4 | 101 | 100.2 | 99.4
-12.6 | 101 | 101.4 | 101 | 100.2 | 99.4
-12.5 | 101.06 | 101.5 | 101.0667 | 100.2 | 99.4
-12.4 | 101.1 | 101.5 | 101.1 | 100.3 | 99.4
-12.3 | 101.1286 | 101.52 | 101.2 | 100.3 | 99.5
-12.2 | 101.2 | 101.6 | 101.2857 | 100.3 | 99.5833
-12.1 | 101.3 | 101.6 | 101.3429 | 100.3 | 99.6333
-12 | 101.3 | 101.7 | 101.4 | 100.3 | 99.8
-11.9 | 101.3714 | 101.7 | 101.4 | 100.3 | 99.8
-11.8 | 101.4 | 101.8 | 101.4 | 100.3 | 99.8
-11.7 | 101.4 | 101.8 | 101.5 | 100.3 | 99.9
-11.6 | 101.4571 | 101.8333 | 101.5 | 100.3 | 99.9556
-11.5 | 101.5 | 101.9 | 101.5 | 100.4 | 100.0857
-11.4 | 101.6 | 101.9 | 101.5 | 100.4 | 100.1375
-11.3 | 101.6 | 102 | 101.5 | 100.5 | 100.2
-11.2 | 101.6 | 102 | 101.5 | 100.5571 | 100.2
-11.1 | 101.7 | 102 | 101.4 | 100.6 | 100.3
-11 | 101.7 | 102 | 101.4 | 100.6 | 100.3
-10.9 | 101.7 | 102 | 101.4 | 100.6 | 100.4
-10.8 | 101.7 | 102 | 101.5 | 100.7 | 100.4286
-10.7 | 101.7 | 101.95 | 101.5 | 100.7 | 100.5
-10.6 | 101.7 | 101.9 | 101.5 | 100.7286 | 100.5
-10.5 | 101.7 | 101.9 | 101.5 | 100.8 | 100.54
-10.4 | 101.6 | 101.9 | 101.5 | 100.9 | 100.6
-10.3 | 101.6 | 101.9 | 101.6 | 101 | 100.6143
-10.2 | 101.6 | 101.9 | 101.6 | 101.0571 | 100.7
-10.1 | 101.6 | 101.9 | 101.6 | 101.1 | 100.7
-10 | 101.6 | 101.9 | 101.7 | 101.1286 | 100.7
-9.9 | 101.6 | 101.9 | 101.7 | 101.2 | 100.8
-9.8 | 101.6 | 101.9 | 101.7143 | 101.2 | 100.8
-9.7 | 101.6286 | 101.975 | 101.8 | 101.2571 | 100.8
-9.6 | 101.7 | 102 | 101.9 | 101.3 | 100.9
-9.5 | 101.8 | 102 | 101.9 | 101.3167 | 100.9
-9.4 | 101.8 | 102.1 | 102 | 101.4 | 100.9
-9.3 | 101.9 | 102.2 | 102 | 101.4 | 101
-9.2 | 101.9 | 102.3 | 102 | 101.4 | 101
-9.1 | 102 | 102.3 | 102 | 101.4 | 101
-9 | 102 | 102.4 | 102.1 | 101.5 | 101
-8.9 | 102.1 | 102.4 | 102.1 | 101.5 | 101.1333
-8.8 | 102.1 | 102.4 | 102.1 | 101.5 | 101.2
-8.7 | 102.2 | 102.4 | 102.2 | 101.5556 | 101.2
-8.6 | 102.2333 | 102.5 | 102.2 | 101.6 | 101.2875
-8.5 | 102.3 | 102.5 | 102.2 | 101.6 | 101.3125
-8.4 | 102.3 | 102.5 | 102.2 | 101.65 | 101.3571
-8.3 | 102.3 | 102.5 | 102.2 | 101.7 | 101.3857
-8.2 | 102.3 | 102.5 | 102.2 | 101.7 | 101.4
-8.1 | 102.3 | 102.6 | 102.2 | 101.7 | 101.4
-8 | 102.4 | 102.6 | 102.3 | 101.7 | 101.4
-7.9 | 102.4 | 102.6 | 102.3 | 101.7 | 101.5
-7.8 | 102.4 | 102.6 | 102.3 | 101.8 | 101.6
-7.7 | 102.4 | 102.6 | 102.4 | 101.8 | 101.6
-7.6 | 102.4 | 102.6 | 102.4 | 101.8 | 101.6
-7.5 | 102.4 | 102.6 | 102.4 | 101.9 | 101.6
-7.4 | 102.4 | 102.65 | 102.5 | 101.9 | 101.6
-7.3 | 102.5 | 102.7 | 102.5 | 101.925 | 101.6667
-7.2 | 102.5 | 102.75 | 102.5 | 102 | 101.7
-7.1 | 102.5 | 102.8 | 102.6 | 102 | 101.7
-7 | 102.5833 | 102.8 | 102.6 | 102.1 | 101.7
-6.9 | 102.6667 | 102.8 | 102.6 | 102.1 | 101.74
-6.8 | 102.7 | 102.84 | 102.6 | 102.1 | 101.8
-6.7 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
-6.6 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
-6.5 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
-6.4 | 102.7 | 102.9 | 102.6 | 102.2 | 101.9
-6.3 | 102.7 | 102.9 | 102.65 | 102.1 | 101.9
-6.2 | 102.8 | 102.9 | 102.7 | 102.1 | 101.9
-6.1 | 102.8 | 102.9143 | 102.7 | 102.1 | 101.9
-6 | 102.8 | 102.92 | 102.7 | 102.1 | 101.8
-5.9 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
-5.8 | 102.8 | 102.9571 | 102.7 | 102.1 | 101.8
-5.7 | 102.8 | 103 | 102.7 | 102.1 | 101.8
-5.6 | 102.8 | 103 | 102.7 | 102.1 | 101.8
-5.5 | 102.8 | 102.9667 | 102.7 | 102.1 | 101.8
-5.4 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
-5.3 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
-5.2 | 102.8 | 102.9 | 102.7 | 102.0833 | 101.8
-5.1 | 102.8 | 102.9 | 102.6 | 102 | 101.8
-5 | 102.8 | 102.9 | 102.6 | 102 | 101.7125
-4.9 | 102.8 | 102.8 | 102.5 | 102 | 101.7
-4.8 | 102.75 | 102.8 | 102.5 | 102 | 101.7
-4.7 | 102.7 | 102.8 | 102.5 | 102 | 101.7
-4.6 | 102.7 | 102.8 | 102.5 | 102 | 101.7
-4.5 | 102.6 | 102.7 | 102.5 | 102 | 101.7
-4.4 | 102.6 | 102.7 | 102.5 | 102 | 101.7
-4.3 | 102.6 | 102.6 | 102.5 | 101.9 | 101.7
-4.2 | 102.6 | 102.6 | 102.4 | 101.9 | 101.6
-4.1 | 102.5143 | 102.5 | 102.4 | 101.9 | 101.6
-4 | 102.5 | 102.5 | 102.32 | 101.8571 | 101.6
-3.9 | 102.4143 | 102.5 | 102.3 | 101.8 | 101.6
-3.8 | 102.4 | 102.5 | 102.3 | 101.8 | 101.5
-3.7 | 102.3429 | 102.4 | 102.2 | 101.8 | 101.5
-3.6 | 102.3 | 102.4 | 102.2 | 101.8 | 101.5
-3.5 | 102.3 | 102.4 | 102.12 | 101.8 | 101.4
-3.4 | 102.2 | 102.3 | 102.1 | 101.7 | 101.4
-3.3 | 102.2 | 102.2143 | 102.1 | 101.7 | 101.3
-3.2 | 102.2 | 102.2 | 102.0571 | 101.6 | 101.2429
-3.1 | 102.1 | 102.1 | 102 | 101.5833 | 101.2
-3 | 102.1 | 102.08 | 101.9 | 101.5 | 101.2
-2.9 | 102 | 102 | 101.8 | 101.4125 | 101.12
-2.8 | 101.9 | 101.9 | 101.6833 | 101.32 | 101.0571
-2.7 | 101.8 | 101.8667 | 101.6 | 101.2625 | 101
-2.6 | 101.7 | 101.8 | 101.5 | 101.2 | 100.9625
-2.5 | 101.5571 | 101.65 | 101.4167 | 101.1 | 100.8
-2.4 | 101.5 | 101.5 | 101.35 | 101.0429 | 100.8
-2.3 | 101.4 | 101.4571 | 101.2125 | 101 | 100.8
-2.2 | 101.2667 | 101.3 | 101.0714 | 100.9286 | 100.8
-2.1 | 101.2 | 101.2333 | 101 | 100.8833 | 100.7
-2 | 101.1333 | 101.1 | 100.95 | 100.8 | 100.7
-1.9 | 101.06 | 101.0143 | 100.9 | 100.7 | 100.6
-1.8 | 100.9 | 100.9 | 100.85 | 100.6 | 100.575
-1.7 | 100.9 | 100.8333 | 100.8 | 100.6 | 100.5
-1.6 | 100.8 | 100.7667 | 100.7333 | 100.5 | 100.425
-1.5 | 100.64 | 100.7 | 100.6571 | 100.4833 | 100.4
-1.4 | 100.6 | 100.6 | 100.6 | 100.4 | 100.4
-1.3 | 100.5 | 100.6 | 100.6 | 100.4 | 100.3125
-1.2 | 100.5 | 100.5 | 100.5 | 100.3 | 100.3
-1.1 | 100.4125 | 100.4 | 100.45 | 100.3 | 100.3
-1 | 100.3875 | 100.3 | 100.3857 | 100.2571 | 100.2111
-0.9 | 100.3 | 100.2 | 100.3 | 100.2 | 100.2
-0.8 | 100.2 | 100.2 | 100.2 | 100.0857 | 100.15
-0.7 | 100.1 | 100.1571 | 100.2 | 100 | 100.1
-0.6 | 100.1 | 100.1 | 100.1 | 100 | 100
-0.5 | 100.1 | 100 | 100.1 | 100 | 100
-0.4 | 100 | 100 | 100.0857 | 100 | 100
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 99.9 | 100
0.3 | 100 | 100 | 100.0143 | 100 | 100
0.4 | 100 | 100 | 100.1 | 99.9571 | 100
0.5 | 100.1 | 100 | 100.1 | 100 | 100
0.6 | 100.1 | 100.1 | 100.1 | 100 | 100
0.7 | 100.1 | 100.1571 | 100.16 | 100 | 100.1
0.8 | 100.2 | 100.2 | 100.2 | 100.1 | 100.15
0.9 | 100.3 | 100.2 | 100.3 | 100.1714 | 100.2
1 | 100.3875 | 100.3 | 100.4 | 100.3 | 100.2111
1.1 | 100.4125 | 100.4 | 100.5 | 100.3 | 100.3
1.2 | 100.5 | 100.5 | 100.5 | 100.3333 | 100.3
1.3 | 100.5 | 100.6 | 100.5667 | 100.4 | 100.3125
1.4 | 100.6 | 100.6 | 100.6 | 100.4 | 100.4
1.5 | 100.64 | 100.7 | 100.6333 | 100.5 | 100.4
1.6 | 100.8 | 100.7667 | 100.7167 | 100.5 | 100.425
1.7 | 100.9 | 100.8333 | 100.8 | 100.6 | 100.5
1.8 | 100.9 | 100.9 | 100.825 | 100.62 | 100.575
1.9 | 101.06 | 101.0143 | 100.9 | 100.7 | 100.6
2 | 101.1333 | 101.1 | 101 | 100.76 | 100.7
2.1 | 101.2 | 101.2333 | 101 | 100.8286 | 100.7
2.2 | 101.2667 | 101.3 | 101.15 | 100.9 | 100.8
2.3 | 101.4 | 101.4571 | 101.2 | 101 | 100.8
2.4 | 101.5 | 101.5 | 101.3167 | 101 | 100.8
2.5 | 101.5571 | 101.65 | 101.5 | 101.1 | 100.8
2.6 | 101.7 | 101.8 | 101.5429 | 101.175 | 100.9625
2.7 | 101.8 | 101.8667 | 101.6 | 101.2143 | 101
2.8 | 101.9 | 101.9 | 101.7 | 101.3 | 101.0571
2.9 | 102 | 102 | 101.8143 | 101.4 | 101.12
3 | 102.1 | 102.08 | 101.9 | 101.4875 | 101.2
3.1 | 102.1 | 102.1 | 102 | 101.6 | 101.2
3.2 | 102.2 | 102.2 | 102.0286 | 101.6 | 101.2429
3.3 | 102.2 | 102.2143 | 102.1 | 101.65 | 101.3
3.4 | 102.2 | 102.3 | 102.1 | 101.7167 | 101.4
3.5 | 102.3 | 102.4 | 102.1 | 101.8 | 101.4
3.6 | 102.3 | 102.4 | 102.2 | 101.8 | 101.5
3.7 | 102.3429 | 102.4 | 102.2 | 101.8 | 101.5
3.8 | 102.4 | 102.5 | 102.2875 | 101.8 | 101.5
3.9 | 102.4143 | 102.5 | 102.3 | 101.8 | 101.6
4 | 102.5 | 102.5 | 102.4 | 101.9 | 101.6
4.1 | 102.5143 | 102.5 | 102.4 | 101.9 | 101.6
4.2 | 102.6 | 102.6 | 102.4 | 101.9 | 101.6
4.3 | 102.6 | 102.6 | 102.5 | 101.9 | 101.7
4.4 | 102.6 | 102.7 | 102.5 | 102 | 101.7
4.5 | 102.6 | 102.7 | 102.5 | 102 | 101.7
4.6 | 102.7 | 102.8 | 102.5 | 102 | 101.7
4.7 | 102.7 | 102.8 | 102.5 | 102 | 101.7
4.8 | 102.75 | 102.8 | 102.5 | 102 | 101.7
4.9 | 102.8 | 102.8 | 102.5714 | 102 | 101.7
5 | 102.8 | 102.9 | 102.6 | 102 | 101.7125
5.1 | 102.8 | 102.9 | 102.6 | 102 | 101.8
5.2 | 102.8 | 102.9 | 102.7 | 102.0333 | 101.8
5.3 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
5.4 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
5.5 | 102.8 | 102.9667 | 102.7 | 102.1 | 101.8
5.6 | 102.8 | 103 | 102.7 | 102.1 | 101.8
5.7 | 102.8 | 103 | 102.7 | 102.1 | 101.8
5.8 | 102.8 | 102.9571 | 102.7 | 102.1 | 101.8
5.9 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
6 | 102.8 | 102.92 | 102.7 | 102.1 | 101.8
6.1 | 102.8 | 102.9143 | 102.7 | 102.1 | 101.9
6.2 | 102.8 | 102.9 | 102.7 | 102.1 | 101.9
6.3 | 102.7 | 102.9 | 102.675 | 102.15 | 101.9
6.4 | 102.7 | 102.9 | 102.6 | 102.2 | 101.9
6.5 | 102.7 | 102.9 | 102.6 | 102.1375 | 101.8
6.6 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
6.7 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
6.8 | 102.7 | 102.84 | 102.6 | 102.1 | 101.8
6.9 | 102.6667 | 102.8 | 102.6 | 102.1 | 101.74
7 | 102.5833 | 102.8 | 102.6 | 102.1 | 101.7
7.1 | 102.5 | 102.8 | 102.6 | 102.0429 | 101.7
7.2 | 102.5 | 102.75 | 102.52 | 102 | 101.7
7.3 | 102.5 | 102.7 | 102.5 | 101.9714 | 101.6667
7.4 | 102.4 | 102.65 | 102.4571 | 101.9 | 101.6
7.5 | 102.4 | 102.6 | 102.4 | 101.85 | 101.6
7.6 | 102.4 | 102.6 | 102.4 | 101.8 | 101.6
7.7 | 102.4 | 102.6 | 102.4 | 101.8 | 101.6
7.8 | 102.4 | 102.6 | 102.3 | 101.8 | 101.6
7.9 | 102.4 | 102.6 | 102.3 | 101.7125 | 101.5
8 | 102.4 | 102.6 | 102.2833 | 101.7 | 101.4
8.1 | 102.3 | 102.6 | 102.2 | 101.7 | 101.4
8.2 | 102.3 | 102.5 | 102.2 | 101.7 | 101.4
8.3 | 102.3 | 102.5 | 102.2 | 101.7 | 101.3857
8.4 | 102.3 | 102.5 | 102.2 | 101.6 | 101.3571
8.5 | 102.3 | 102.5 | 102.2 | 101.6 | 101.3125
8.6 | 102.2333 | 102.5 | 102.2 | 101.6 | 101.2875
8.7 | 102.2 | 102.4 | 102.2 | 101.6 | 101.2
8.8 | 102.1 | 102.4 | 102.1 | 101.5 | 101.2
8.9 | 102.1 | 102.4 | 102.1 | 101.5 | 101.1333
9 | 102 | 102.4 | 102.1 | 101.5 | 101
9.1 | 102 | 102.3 | 102 | 101.4 | 101
9.2 | 101.9 | 102.3 | 102 | 101.4 | 101
9.3 | 101.9 | 102.2 | 102 | 101.4 | 101
9.4 | 101.8 | 102.1 | 102 | 101.4 | 100.9
9.5 | 101.8 | 102 | 101.9 | 101.3 | 100.9
9.6 | 101.7 | 102 | 101.8286 | 101.3 | 100.9
9.7 | 101.6286 | 101.975 | 101.775 | 101.3 | 100.8
9.8 | 101.6 | 101.9 | 101.7 | 101.2 | 100.8
9.9 | 101.6 | 101.9 | 101.7 | 101.2 | 100.8
10 | 101.6 | 101.9 | 101.7 | 101.1 | 100.7
10.1 | 101.6 | 101.9 | 101.6 | 101.1 | 100.7
10.2 | 101.6 | 101.9 | 101.6 | 101.1 | 100.7
10.3 | 101.6 | 101.9 | 101.6 | 101 | 100.6143
10.4 | 101.6 | 101.9 | 101.5 | 100.9 | 100.6
10.5 | 101.7 | 101.9 | 101.5 | 100.85 | 100.54
10.6 | 101.7 | 101.9 | 101.5 | 100.7833 | 100.5
10.7 | 101.7 | 101.95 | 101.5 | 100.7 | 100.5
10.8 | 101.7 | 102 | 101.5 | 100.7 | 100.4286
10.9 | 101.7 | 102 | 101.4125 | 100.6 | 100.4
11 | 101.7 | 102 | 101.4 | 100.6 | 100.3
11.1 | 101.7 | 102 | 101.4 | 100.6 | 100.3
11.2 | 101.6 | 102 | 101.5 | 100.5143 | 100.2
11.3 | 101.6 | 102 | 101.5 | 100.5 | 100.2
11.4 | 101.6 | 101.9 | 101.5 | 100.45 | 100.1375
11.5 | 101.5 | 101.9 | 101.5 | 100.4 | 100.0857
11.6 | 101.4571 | 101.8333 | 101.5 | 100.3429 | 99.9556
11.7 | 101.4 | 101.8 | 101.5 | 100.3 | 99.9
11.8 | 101.4 | 101.8 | 101.425 | 100.3 | 99.8
11.9 | 101.3714 | 101.7 | 101.4 | 100.3 | 99.8
12 | 101.3 | 101.7 | 101.4 | 100.3 | 99.8
12.1 | 101.3 | 101.6 | 101.3714 | 100.3 | 99.6333
12.2 | 101.2 | 101.6 | 101.3 | 100.3 | 99.5833
12.3 | 101.1286 | 101.52 | 101.2 | 100.3 | 99.5
12.4 | 101.1 | 101.5 | 101.1125 | 100.2833 | 99.4
12.5 | 101.06 | 101.5 | 101.0889 | 100.2 | 99.4
12.6 | 101 | 101.4 | 101 | 100.2 | 99.4
12.7 | 100.925 | 101.4 | 101 | 100.1833 | 99.4
12.8 | 100.8857 | 101.3 | 100.9 | 100.1 | 99.4
12.9 | 100.8 | 101.3 | 100.84 | 100.0429 | 99.4
13 | 100.775 | 101.2375 | 100.8 | 100 | 99.3
13.1 | 100.6333 | 101.1 | 100.7 | 99.9 | 99.4
13.2 | 100.5714 | 101.1 | 100.7 | 99.8167 | 99.3875
13.3 | 100.5 | 101 | 100.6 | 99.78 | 99.3
13.4 | 100.4333 | 100.9167 | 100.5 | 99.7 | 99.2889
13.5 | 100.3833 | 100.85 | 100.5 | 99.575 | 99.2
13.6 | 100.3 | 100.8 | 100.4 | 99.45 | 99.1429
13.7 | 100.24 | 100.7 | 100.3 | 99.4 | 99.1
13.8 | 100.2 | 100.6 | 100.22 | 99.36 | 99.1
13.9 | 100.1125 | 100.6 | 100.2 | 99.3 | 99.0333
14 | 100.0571 | 100.5 | 100.08 | 99.2 | 98.98
14.1 | 99.9833 | 100.4429 | 100 | 99.2 | 98.8857
14.2 | 99.9 | 100.4 | 99.9 | 99.15 | 98.8
14.3 | 99.9 | 100.3625 | 99.8 | 99.025 | 98.7
14.4 | 99.9 | 100.3 | 99.8 | 98.9 | 98.675
14.5 | 99.9 | 100.3 | 99.7 | 98.8 | 98.4625
14.6 | 99.9 | 100.2125 | 99.6167 | 98.8 | 98.4
14.7 | 99.8 | 100.2 | 99.6 | 98.7 | 98.3125
14.8 | 99.7 | 100.1167 | 99.46 | 98.625 | 98.1571
14.9 | 99.6571 | 100.075 | 99.4 | 98.5 | 98.1
15 | 99.6 | 100 | 99.3 | 98.4 | 98
15.1 | 99.6 | 99.9 | 99.2714 | 98.3 | 97.975
15.2 | 99.5167 | 99.9 | 99.2 | 98.2429 | 97.85
15.3 | 99.5 | 99.9 | 99.1 | 98.1 | 97.8
15.4 | 99.5 | 99.9 | 99.0714 | 97.9714 | 97.7
15.5 | 99.4 | 99.8 | 99 | 97.9 | 97.5667
15.6 | 99.3 | 99.8 | 99 | 97.7571 | 97.5
15.7 | 99.25 | 99.7 | 99 | 97.7 | 97.3
15.8 | 99.1 | 99.6 | 98.8857 | 97.575 | 97.2286
15.9 | 99.05 | 99.6 | 98.8 | 97.45 | 97.2
16 | 99 | 99.5222 | 98.7833 | 97.3143 | 97.0857
16.1 | 98.85 | 99.4 | 98.7 | 97.275 | 97
16.2 | 98.6667 | 99.4 | 98.6286 | 97.2 | 96.9125
16.3 | 98.6 | 99.3 | 98.5 | 97.1125 | 96.7875
16.4 | 98.5 | 99.1167 | 98.4 | 96.9571 | 96.625
16.5 | 98.3875 | 99 | 98.35 | 96.9 | 96.4375
16.6 | 98.25 | 98.8857 | 98.2 | 96.7889 | 96.38
16.7 | 98.1 | 98.8 | 98.2 | 96.6667 | 96.3
16.8 | 98.02 | 98.7 | 98.0375 | 96.6 | 96.1667
16.9 | 97.9333 | 98.6 | 98 | 96.6 | 96.1
17 | 97.8 | 98.5 | 97.9 | 96.475 | 95.975
17.1 | 97.7143 | 98.3167 | 97.7714 | 96.35 | 95.8429
17.2 | 97.6 | 98.2167 | 97.6167 | 96.3 | 95.7125
17.3 | 97.5 | 98.05 | 97.5571 | 96.175 | 95.55
17.4 | 97.4167 | 98 | 97.425 | 96.05 | 95.5
17.5 | 97.3833 | 97.8857 | 97.1 | 95.925 | 95.3889
17.6 | 97.3 | 97.6667 | 97.0833 | 95.8875 | 95.275
17.7 | 97.2 | 97.6 | 96.9286 | 95.7625 | 95.1333
17.8 | 97.18 | 97.5 | 96.8667 | 95.6286 | 95.0714
17.9 | 97.1 | 97.5 | 96.7 | 95.5833 | 94.9375
18 | 97.1 | 97.42 | 96.55 | 95.375 | 94.9
18.1 | 97.025 | 97.3714 | 96.4 | 95.2 | 94.8
18.2 | 97 | 97.2167 | 96.2429 | 95.1286 | 94.5833
18.3 | 96.9 | 97.1857 | 96.2 | 94.9857 | 94.5
18.4 | 96.8 | 97.1 | 96 | 94.85 | 94.45
18.5 | 96.7714 | 97.0857 | 95.8833 | 94.7143 | 94.3
18.6 | 96.6167 | 96.9333 | 95.7167 | 94.5571 | 94.15
18.7 | 96.4333 | 96.88 | 95.6429 | 94.4 | 94.1
18.8 | 96.24 | 96.8 | 95.5 | 94.2571 | 93.9778
18.9 | 95.96 | 96.6286 | 95.4 | 94.125 | 93.9
19 | 95.7 | 96.48 | 95.3 | 94 | 93.8
19.1 | 95.5333 | 96.3 | 95.2 | 93.8 | 93.6571
19.2 | 95.22 | 96.0333 | 95.1286 | 93.66 | 93.5
19.3 | 94.7143 | 95.7 | 94.9 | 93.5125 | 93.35
19.4 | 94.3429 | 95.3714 | 94.8 | 93.3571 | 93.2
19.5 | 93.675 | 95.0667 | 94.7 | 93.0875 | 93.075
19.6 | 92.9 | 94.5714 | 94.4667 | 92.9143 | 92.9429
19.7 | 92.0286 | 94 | 94.3 | 92.7143 | 92.7714
19.8 | 90.9333 | 93.55 | 94.1 | 92.5429 | 92.6
19.9 | 89.4857 | 92.925 | 93.9429 | 92.3714 | 92.42
20 | 87.5 | 92.1429 | 93.6714 | 92.2444 | 92.2
20.1 | 84.5571 | 91.2857 | 93.4 | 92.125 | 92.1
20.2 | 80.45 | 89.9857 | 93.15 | 91.9 | 91.9875
20.3 | 74.3667 | 88.5333 | 92.82 | 91.7714 | 91.8143
20.4 | 67.4667 | 86.15 | 92.4 | 91.6 | 91.7
20.5 | 59.1 | 83.225 | 91.98 | 91.4333 | 91.4625
20.6 | 50.8375 | 78.8143 | 91.4571 | 91.3 | 91.2571
20.7 | 42.8167 | 73.1 | 90.9714 | 91.15 | 91.1
20.8 | 36.14 | 65.9167 | 90.25 | 90.8857 | 90.875
20.9 | 30.5 | 57.7 | 89.5 | 90.7429 | 90.7
21 | 26.02 | 50.1143 | 88.7 | 90.6125 | 90.45
21.1 | 22.9143 | 42.8571 | 87.68 | 90.4429 | 90.3
21.2 | 20.6286 | 36.0571 | 86.4286 | 90.3125 | 90.0571
21.3 | 19 | 30.5 | 84.6857 | 90.1714 | 89.9
21.4 | 17.8 | 26.25 | 82.6286 | 89.9429 | 89.82
21.5 | 16.7714 | 23.0714 | 79.45 | 89.8 | 89.6667
21.6 | 15.9571 | 20.6 | 75.25 | 89.6 | 89.4
21.7 | 15.3143 | 18.9714 | 69.625 | 89.1833 | 89.2333
21.8 | 14.8 | 17.7333 | 63.1571 | 89 | 88.9833
21.9 | 14.3286 | 16.625 | 55.75 | 88.6571 | 88.8375
22 | 14 | 15.7571 | 48.4 | 88.3667 | 88.7125
22.1 | 13.68 | 15.0429 | 41.2833 | 88.0571 | 88.4
22.2 | 13.4 | 14.5286 | 35.7 | 87.58 | 88.25
22.3 | 13.1857 | 14 | 30.8 | 87.0857 | 88.05
22.4 | 13.0125 | 13.6286 | 27.0571 | 86.7714 | 87.9
22.5 | 12.7571 | 13.3 | 24.2 | 86.2625 | 87.76
22.6 | 12.6143 | 13.0143 | 22.1 | 85.7857 | 87.5857
22.7 | 12.45 | 12.7286 | 20.6 | 85.1571 | 87.4333
22.8 | 12.2833 | 12.5429 | 19.32 | 84.525 | 87.2
22.9 | 12.1 | 12.3 | 18.2571 | 83.7375 | 87.05
23 | 11.9333 | 12.1 | 17.4714 | 82.8 | 86.8875
23.1 | 11.76 | 11.85 | 16.775 | 81.8 | 86.7
23.2 | 11.6125 | 11.7125 | 16.16 | 80.4 | 86.4
23.3 | 11.4714 | 11.5714 | 15.6 | 78.7333 | 86.2167
23.4 | 11.4 | 11.4 | 15.25 | 76.6 | 86.05
23.5 | 11.2833 | 11.2667 | 14.8333 | 73.15 | 85.75
23.6 | 11.1286 | 11.1 | 14.5 | 69.1429 | 85.4286
23.7 | 11 | 10.9333 | 14.2 | 64.4286 | 85.15
23.8 | 10.9 | 10.85 | 13.925 | 59 | 84.8667
23.9 | 10.8286 | 10.7 | 13.7167 | 52.8 | 84.46
24 | 10.7 | 10.6 | 13.5143 | 46.875 | 84.1375
24.1 | 10.6 | 10.5 | 13.2833 | 41.3125 | 83.78
24.2 | 10.5333 | 10.3714 | 13.1 | 36.2875 | 83.36
24.3 | 10.4 | 10.3 | 12.8571 | 32.3667 | 82.9
24.4 | 10.3286 | 10.14 | 12.725 | 28.9 | 82.5286
24.5 | 10.2833 | 10.06 | 12.5 | 26.42 | 82
24.6 | 10.1375 | 10 | 12.3667 | 24.45 | 81.475
24.7 | 10.0833 | 9.9 | 12.2143 | 22.9 | 80.8
24.8 | 10 | 9.8 | 12.06 | 21.7 | 80.2
24.9 | 9.85 | 9.7286 | 11.7714 | 20.8 | 79.4
25 | 9.7833 | 9.6 | 11.6286 | 19.98 | 78.45
25.1 | 9.7 | 9.5286 | 11.5625 | 19.2429 | 77.34
25.2 | 9.5667 | 9.4714 | 11.4286 | 18.7 | 75.7571
25.3 | 9.5 | 9.3167 | 11.2714 | 18.1571 | 74.1429
25.4 | 9.4833 | 9.2714 | 11.2 | 17.75 | 71.8125
25.5 | 9.4 | 9.2 | 11.0833 | 17.3667 | 69.0778
25.6 |  | 9.1 | 10.9167 | 17.0125 | 64.9833
25.7 |  | 9.0286 | 10.8 | 16.6375 | 60.7571
25.8 |  | 8.9667 | 10.7 | 16.3778 | 55.44
25.9 |  | 8.8875 | 10.6 | 16.125 | 50
26 |  |  | 10.4167 | 15.8429 | 44.7
26.1 |  |  | 10.4 | 15.6 | 39.9
26.2 |  |  | 10.3 | 15.3857 | 35.8
26.3 |  |  | 10.2 | 15.2 | 32.3
26.4 |  |  | 10.12 | 14.9556 | 29.6875
26.5 |  |  | 10 | 14.7143 | 27.5
26.6 |  |  | 9.9 | 14.5143 | 25.7625
26.7 |  |  | 9.75 | 14.2857 | 24.48
26.8 |  |  | 9.7 | 14.1556 | 23.3143
26.9 |  |  | 9.625 | 13.975 | 22.4286
27 |  |  |  | 13.8125 | 21.7143
27.1 |  |  |  | 13.5875 | 21.02
27.2 |  |  |  | 13.4625 | 20.4
27.3 |  |  |  | 13.3286 | 20.0714
27.4 |  |  |  | 13.0857 | 19.6
27.5 |  |  |  | 12.9333 | 19.1667
27.6 |  |  |  | 12.9 | 18.8667
27.7 |  |  |  | 12.75 | 18.5
27.8 |  |  |  | 12.5429 | 18.2571
27.9 |  |  |  | 12.4125 | 17.9714
28 |  |  |  | 12.2625 | 17.7
28.1 |  |  |  | 12.2 | 17.4286
28.2 |  |  |  | 12.0833 | 17.25
28.3 |  |  |  | 11.9375 | 17.025
28.4 |  |  |  | 11.8125 | 16.7857
28.5 |  |  |  | 11.6875 | 16.5667
28.6 |  |  |  | 11.5625 | 16.45
28.7 |  |  |  | 11.4375 | 16.1857
28.8 |  |  |  | 11.3125 | 16
28.9 |  |  |  | 11.2 | 15.825
29 |  |  |  |  | 15.66
29.1 |  |  |  |  | 15.4
29.2 |  |  |  |  | 15.25
29.3 |  |  |  |  | 15.05
29.4 |  |  |  |  | 14.9
29.5 |  |  |  |  | 14.7857
29.6 |  |  |  |  | 14.6
29.7 |  |  |  |  | 14.48
29.8 |  |  |  |  | 14.2857
29.9 |  |  |  |  | 14.1333
30 |  |  |  |  | 13.9714
30.1 |  |  |  |  | 13.8286
30.2 |  |  |  |  | 13.7
30.3 |  |  |  |  | 13.625
30.4 |  |  |  |  | 13.4375
30.5 |  |  |  |  | 13.225
30.6 |  |  |  |  | 13.1
30.7 |  |  |  |  | 12.975
30.8 |  |  |  |  | 12.85
30.9 |  |  |  |  | 12.725

## Wedge 30 Output Factors
Output factors
Wedge 30 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.866 | 0.884 | 0.892 | 0.901 | 0.907 | 0.913 | 0.917 | 0.921
| 4 | 0.886 | 0.906 | 0.916 | 0.928 | 0.936 | 0.943 | 0.948 | 0.954
| 5 | 0.896 | 0.918 | 0.93 | 0.944 | 0.953 | 0.962 | 0.969 | 0.976
| 7 | 0.907 | 0.933 | 0.947 | 0.964 | 0.977 | 0.989 | 0.997 | 1.007
| 10 | 0.917 | 0.946 | 0.963 | 0.983 | 1 | 1.017 | 1.028 | 1.041
| 15 | 0.927 | 0.958 | 0.977 | 1.001 | 1.023 | 1.045 | 1.061 | 1.08
| 20 | 0.933 | 0.964 | 0.985 | 1.011 | 1.036 | 1.063 | 1.082 | 1.106
| 30 | 0.939 | 0.973 | 0.995 | 1.025 | 1.054 | 1.086 | 1.111 | 1.142
| 40 | 0.944 | 0.979 | 1.001 | 1.033 | 1.065 | 1.101 | 1.13 | 1.164

## Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 24 | 30.5 | 43.7 | 52.5
0.1 | 27.6 | 34.2667 | 47.4 | 56.175
0.2 | 33.38 | 40 | 53.3 | 61.64
0.3 | 42.1 | 48.15 | 60.475 | 68.2
0.4 | 50.25 | 55.44 | 67.24 | 73.8
0.5 | 57.1 | 61.9 | 72.68 | 78.8667
0.6 | 63.1 | 67.8 | 77.32 | 82.8
0.7 | 68.5 | 72.4 | 80.86 | 86.125
0.8 | 72.925 | 76.54 | 84.2 | 88.6
0.9 | 76.9 | 80.2 | 87 | 90.9667
1 | 79.925 | 83.2 | 89.48 | 92.925
1.1 | 83.3 | 85.675 | 91.3667 | 94.525
1.2 | 85.7 | 87.9 | 92.9 | 95.8
1.3 | 88.2 | 89.9 | 94.45 | 96.7
1.4 | 89.85 | 91.55 | 95.6 | 97.6
1.5 | 91.8 | 93.075 | 96.45 | 98.25
1.6 | 93.1 | 94.3 | 97.2833 | 98.925
1.7 | 94.2 | 95.325 | 98.0333 | 99.35
1.8 | 95.52 | 96.5 | 98.5 | 99.55
1.9 | 96.6 | 97.12 | 99.04 | 99.8
2 | 97.3 | 97.7 | 99.38 | 99.9
2.1 | 97.95 | 98.4 | 99.68 | 100
2.2 | 98.48 | 98.7 | 99.8 | 100
2.3 | 98.88 | 99.1 | 99.9 | 100
2.4 | 99.25 | 99.35 | 100 | 99.9
2.5 | 99.44 | 99.65 | 100 | 99.825
2.6 | 99.7 | 99.8 | 99.9 | 99.7
2.7 | 99.85 | 99.8833 | 99.9 | 99.6
2.8 | 99.9 | 100 | 99.74 | 99.375
2.9 | 99.9833 | 100 | 99.6 | 99.16
3 | 99.9 | 99.9 | 99.5 | 98.86
3.1 | 99.9333 | 99.8 | 99.3 | 98.55
3.2 | 99.8667 | 99.6 | 99.1333 | 98.38
3.3 | 99.8 | 99.5 | 98.75 | 98.05
3.4 | 99.64 | 99.34 | 98.54 | 97.7333
3.5 | 99.44 | 99.2 | 98.2667 | 97.5
3.6 | 99.2 | 98.9 | 98.02 | 97.15
3.7 | 98.95 | 98.7667 | 97.72 | 96.9
3.8 | 98.6 | 98.4667 | 97.35 | 96.52
3.9 | 98.3 | 98.2167 | 97.1 | 96.2
4 | 97.9333 | 97.92 | 96.8 | 95.925
4.1 | 97.7 | 97.58 | 96.475 | 95.55
4.2 | 97.35 | 97.28 | 96.1333 | 95.15
4.3 | 97 | 96.96 | 95.88 | 94.78
4.4 | 96.8 | 96.6333 | 95.56 | 94.525
4.5 | 96.18 | 96.175 | 95.1667 | 94.2
4.6 | 95.94 | 95.92 | 94.8 | 93.86
4.7 | 95.6 | 95.64 | 94.5 | 93.6
4.8 | 95.125 | 95.32 | 94.175 | 93.24
4.9 | 94.8 | 94.92 | 93.775 | 92.8
5 | 94.425 | 94.575 | 93.525 | 92.5667
5.1 | 94.0333 | 94.3667 | 93.1 | 92.28
5.2 | 93.7 | 94 | 92.8 | 91.9333
5.3 | 93.3 | 93.5 | 92.375 | 91.5
5.4 | 92.8 | 93.06 | 92.15 | 91.2667
5.5 | 92.4333 | 92.66 | 91.68 | 90.85
5.6 | 91.975 | 92.2333 | 91.3 | 90.5
5.7 | 91.68 | 91.95 | 90.9667 | 90.22
5.8 | 91.2 | 91.58 | 90.68 | 89.925
5.9 | 90.72 | 91.2 | 90.4 | 89.6
6 | 90.38 | 90.8 | 90.1 | 89.3
6.1 | 89.88 | 90.6 | 89.85 | 88.98
6.2 | 89.5 | 90.3 | 89.48 | 88.65
6.3 | 89.1857 | 89.94 | 89.1 | 88.3
6.4 | 88.7 | 89.5333 | 88.8 | 87.95
6.5 | 88.26 | 89.1667 | 88.425 | 87.58
6.6 | 87.8 | 88.9143 | 88.2 | 87.2333
6.7 | 87.48 | 88.4 | 87.9 | 87.0333
6.8 | 87.2 | 88.12 | 87.4 | 86.7
6.9 | 86.74 | 87.72 | 87.1667 | 86.28
7 | 86.4 | 87.4 | 86.6667 | 85.95
7.1 | 85.9667 | 87 | 86.4 | 85.58
7.2 | 85.6 | 86.72 | 86.05 | 85.4
7.3 | 85.2 | 86.3 | 85.7 | 85.1
7.4 | 84.9 | 85.95 | 85.4 | 84.68
7.5 | 84.5 | 85.625 | 85.0667 | 84.4
7.6 | 84.225 | 85.3 | 84.68 | 84.1
7.7 | 83.7 | 84.95 | 84.44 | 83.75
7.8 | 83.4 | 84.5 | 84.1667 | 83.425
7.9 | 83.075 | 84.16 | 83.74 | 83.1
8 | 82.55 | 83.8833 | 83.5 | 82.75
8.1 | 82.1 | 83.7 | 83.22 | 82.5
8.2 | 81.56 | 83.15 | 82.9 | 82.1
8.3 | 81.3 | 82.92 | 82.5667 | 81.85
8.4 | 81 | 82.44 | 82.2333 | 81.525
8.5 | 80.68 | 82.0571 | 82 | 81.2
8.6 | 80.24 | 81.66 | 81.65 | 80.8667
8.7 | 79.8833 | 81.45 | 81.26 | 80.62
8.8 | 79.55 | 81.2 | 80.98 | 80.36
8.9 | 79.12 | 80.66 | 80.65 | 80.06
9 | 78.68 | 80.44 | 80.32 | 79.675
9.1 | 78.35 | 80.1667 | 80.04 | 79.4
9.2 | 78.1 | 79.8333 | 79.8 | 79.08
9.3 | 77.7 | 79.6 | 79.4 | 78.75
9.4 | 77.4143 | 79.16 | 79.14 | 78.44
9.5 | 77.05 | 78.76 | 78.8667 | 78.1667
9.6 | 76.62 | 78.5 | 78.5 | 77.8333
9.7 | 76.25 | 78.1 | 78.26 | 77.6
9.8 | 75.8 | 77.84 | 77.96 | 77.2667
9.9 | 75.375 | 77.44 | 77.6667 | 76.96
10 | 75.08 | 77.1667 | 77.4 | 76.6
10.1 | 74.7333 | 76.72 | 77.06 | 76.3667
10.2 | 74.4 | 76.4667 | 76.6667 | 76
10.3 | 74.06 | 76.16 | 76.42 | 75.775
10.4 | 73.775 | 75.86 | 76.125 | 75.45
10.5 | 73.38 | 75.5333 | 75.7 | 75.2
10.6 | 72.98 | 75.22 | 75.3667 | 74.9167
10.7 | 72.72 | 74.88 | 75.2 | 74.68
10.8 | 72.42 | 74.45 | 75 | 74.4667
10.9 | 72.1 | 74.18 | 74.58 | 74.175
11 | 71.76 | 73.95 | 74.35 | 73.8
11.1 | 71.45 | 73.625 | 73.9 | 73.45
11.2 | 70.9667 | 73.36 | 73.7 | 73.225
11.3 | 70.7 | 72.96 | 73.4 | 72.98
11.4 | 70.3571 | 72.725 | 73.15 | 72.75
11.5 | 70 | 72.5 | 72.7667 | 72.45
11.6 | 69.75 | 72.2 | 72.5 | 72.2
11.7 | 69.42 | 71.85 | 72.26 | 71.7667
11.8 | 69.0667 | 71.55 | 72.05 | 71.525
11.9 | 68.825 | 71.3 | 71.82 | 71.3
12 | 68.5 | 70.98 | 71.46 | 71
12.1 | 68.25 | 70.58 | 71.08 | 70.8
12.2 | 67.85 | 70.3 | 70.85 | 70.46
12.3 | 67.45 | 69.9667 | 70.62 | 70.2
12.4 | 67.2 | 69.7 | 70.28 | 69.92
12.5 | 66.94 | 69.475 | 70.06 | 69.62
12.6 | 66.55 | 69.125 | 69.85 | 69.42
12.7 | 66.2 | 68.8833 | 69.55 | 69.175
12.8 | 66 | 68.7167 | 69.2333 | 68.85
12.9 | 65.625 | 68.34 | 68.98 | 68.625
13 | 65.3 | 68 | 68.6667 | 68.3667
13.1 | 64.96 | 67.64 | 68.42 | 68.1
13.2 | 64.7 | 67.38 | 68.16 | 67.8
13.3 | 64.4333 | 67.175 | 67.86 | 67.6
13.4 | 64.18 | 66.85 | 67.5833 | 67.3
13.5 | 63.84 | 66.5667 | 67.275 | 67
13.6 | 63.475 | 66.28 | 67.08 | 66.7429
13.7 | 63.2 | 66 | 66.8 | 66.525
13.8 | 62.96 | 65.7 | 66.5667 | 66.2333
13.9 | 62.6 | 65.375 | 66.32 | 66.05
14 | 62.36 | 65.14 | 66.05 | 65.75
14.1 | 61.98 | 64.85 | 65.74 | 65.5
14.2 | 61.7 | 64.625 | 65.46 | 65.22
14.3 | 61.42 | 64.3667 | 65.1833 | 64.96
14.4 | 61.14 | 64.1 | 65 | 64.7
14.5 | 60.8 | 63.7833 | 64.74 | 64.475
14.6 | 60.55 | 63.6 | 64.5 | 64.1833
14.7 | 60.225 | 63.25 | 64.2 | 63.9
14.8 | 60 | 62.95 | 64 | 63.6
14.9 | 59.7667 | 62.56 | 63.68 | 63.42
15 | 59.56 | 62.36 | 63.45 | 63.22
15.1 | 59.325 | 62.0833 | 63.15 | 62.95
15.2 | 59 | 61.875 | 62.94 | 62.68
15.3 | 58.7167 | 61.65 | 62.7 | 62.4286
15.4 | 58.3667 | 61.2714 | 62.5 | 62.2333
15.5 | 58.075 | 60.94 | 62.175 | 61.95
15.6 | 57.8 | 60.7 | 61.9833 | 61.8
15.7 | 57.54 | 60.55 | 61.7333 | 61.5
15.8 | 57.325 | 60.24 | 61.575 | 61.3
15.9 | 57.08 | 60 | 61.3 | 61.1167
16 | 56.8 | 59.8167 | 60.9833 | 60.85
16.1 | 56.5 | 59.54 | 60.8 | 60.6
16.2 | 56.26 | 59.32 | 60.575 | 60.35
16.3 | 56.1 | 59 | 60.34 | 60.08
16.4 | 55.86 | 58.7333 | 60.04 | 59.7857
16.5 | 55.55 | 58.475 | 59.7 | 59.6
16.6 | 55.3 | 58.2333 | 59.5 | 59.42
16.7 | 55.02 | 57.9833 | 59.3 | 59.2429
16.8 | 54.8 | 57.8167 | 59.0667 | 59.05
16.9 | 54.4667 | 57.475 | 58.8 | 58.7333
17 | 54.2 | 57.2 | 58.5833 | 58.5
17.1 | 53.92 | 57 | 58.3571 | 58.25
17.2 | 53.6667 | 56.75 | 58.12 | 58.1
17.3 | 53.4 | 56.425 | 57.8333 | 57.85
17.4 | 53.22 | 56.2 | 57.66 | 57.55
17.5 | 53.05 | 56.0333 | 57.38 | 57.32
17.6 | 52.775 | 55.8 | 57.22 | 57.05
17.7 | 52.5 | 55.56 | 56.94 | 56.85
17.8 | 52.3 | 55.36 | 56.75 | 56.6667
17.9 | 52 | 55.0833 | 56.5833 | 56.4167
18 | 51.76 | 54.86 | 56.4 | 56.22
18.1 | 51.56 | 54.6333 | 56.08 | 56.1
18.2 | 51.3 | 54.4 | 55.88 | 55.8
18.3 | 51 | 54.15 | 55.66 | 55.7
18.4 | 50.78 | 53.9 | 55.36 | 55.4333
18.5 | 50.575 | 53.7333 | 55.225 | 55.08
18.6 | 50.36 | 53.56 | 54.9667 | 54.9
18.7 | 50.1 | 53.1667 | 54.775 | 54.78
18.8 | 49.85 | 52.9 | 54.6 | 54.46
18.9 | 49.6 | 52.675 | 54.3333 | 54.24
19 | 49.28 | 52.54 | 54.0833 | 54.025
19.1 | 49.08 | 52.3 | 53.8333 | 53.85
19.2 | 48.88 | 52.12 | 53.6 | 53.64
19.3 | 48.7 | 51.92 | 53.42 | 53.5
19.4 | 48.48 | 51.5833 | 53.25 | 53.225
19.5 | 48.26 | 51.4167 | 53.0333 | 53
19.6 | 48.1 | 51.2 | 52.75 | 52.8333
19.7 | 47.84 | 50.975 | 52.56 | 52.56
19.8 | 47.65 | 50.76 | 52.35 | 52.4
19.9 | 47.4333 | 50.5 | 52.1667 | 52.3
20 | 47.15 | 50.28 | 51.95 | 51.975
20.1 | 47 | 50.1 | 51.8 | 51.725
20.2 | 46.7286 | 49.9667 | 51.575 | 51.5
20.3 | 46.5333 | 49.7 | 51.3 | 51.38
20.4 | 46.4 | 49.5 | 51.1 | 51.2
20.5 | 46.1286 | 49.18 | 50.9833 | 50.9
20.6 | 45.9 | 49 | 50.78 | 50.76
20.7 | 45.74 | 48.85 | 50.6 | 50.6
20.8 | 45.5333 | 48.5667 | 50.4 | 50.475
20.9 | 45.36 | 48.4 | 50.1 | 50.1667
21 | 45.15 | 48.225 | 50 | 50
21.1 | 44.85 | 47.94 | 49.775 | 49.84
21.2 | 44.5667 | 47.7667 | 49.54 | 49.64
21.3 | 44.34 | 47.58 | 49.4 | 49.4
21.4 | 44.2 | 47.38 | 49.2 | 49.175
21.5 | 43.9 | 47.1667 | 49.06 | 49.05
21.6 | 43.74 | 46.9 | 48.7667 | 48.9
21.7 | 43.52 | 46.7667 | 48.6 | 48.75
21.8 | 43.4 | 46.55 | 48.4 | 48.525
21.9 | 43.22 | 46.34 | 48.2667 | 48.3
22 | 43 | 46.14 | 48.1 | 48.1
22.1 | 42.9 | 46 | 47.9 | 47.8857
22.2 | 42.7 | 45.8 | 47.66 | 47.7
22.3 | 42.5 | 45.6 | 47.3833 | 47.525
22.4 | 42.3333 | 45.4143 | 47.28 | 47.4
22.5 | 42.125 | 45.3 | 46.98 | 47.2667
22.6 | 41.88 | 45.05 | 46.8 | 47.05
22.7 | 41.7 | 44.825 | 46.72 | 46.85
22.8 | 41.5333 | 44.6 | 46.55 | 46.6167
22.9 | 41.36 | 44.44 | 46.4 | 46.425
23 | 41.225 | 44.3 | 46.225 | 46.22
23.1 | 41.0167 | 44 | 46 | 46.1
23.2 | 40.9 | 43.9 | 45.775 | 45.9
23.3 | 40.675 | 43.68 | 45.56 | 45.7333
23.4 | 40.45 | 43.5 | 45.4 | 45.56
23.5 | 40.24 | 43.4 | 45.24 | 45.3
23.6 | 40.05 | 43.16 | 45.04 | 45.1
23.7 | 39.82 | 42.9667 | 44.82 | 45
23.8 | 39.74 | 42.7 | 44.6 | 44.875
23.9 | 39.54 | 42.6 | 44.5 | 44.65
24 | 39.375 | 42.4 | 44.275 | 44.4
24.1 | 39.175 | 42.18 | 44.1 | 44.26
24.2 | 39 | 41.975 | 43.9 | 44.15
24.3 | 38.8667 | 41.74 | 43.7667 | 43.96
24.4 | 38.7 | 41.65 | 43.56 | 43.78
24.5 | 38.4333 | 41.5 | 43.4 | 43.5833
24.6 | 38.25 | 41.325 | 43.3 | 43.4167
24.7 | 38.025 | 41.1 | 43.12 | 43.225
24.8 | 37.92 | 40.95 | 42.9167 | 43
24.9 | 37.75 | 40.8 | 42.7 | 42.9
25 | 37.6 | 40.62 | 42.5833 | 42.8
25.1 | 37.475 | 40.425 | 42.4 | 42.6
25.2 | 37.3 | 40.26 | 42.26 | 42.45
25.3 | 37.05 | 40.05 | 42.125 | 42.3
25.4 | 36.9 | 39.9 | 41.975 | 42.1
25.5 | 36.7 | 39.78 | 41.75 | 42
25.6 | 36.6 | 39.66 | 41.6 | 41.7333
25.7 | 36.425 | 39.4 | 41.4833 | 41.5667
25.8 | 36.3 | 39.2667 | 41.28 | 41.4
25.9 | 36.15 | 39.1 | 41.08 | 41.26
26 | 35.9 | 38.9167 | 40.9 | 41.16
26.1 | 35.8 | 38.72 | 40.76 | 41.04
26.2 | 35.6667 | 38.55 | 40.58 | 40.825
26.3 | 35.45 | 38.45 | 40.35 | 40.65
26.4 | 35.3 | 38.22 | 40.2 | 40.48
26.5 | 35.1 | 38.04 | 40.1 | 40.3
26.6 | 35.02 | 37.8667 | 40 | 40.15
26.7 | 34.9 | 37.8 | 39.8 | 40.02
26.8 | 34.7 | 37.6167 | 39.65 | 39.8167
26.9 | 34.54 | 37.44 | 39.5 | 39.7167
27 | 34.34 | 37.22 | 39.35 | 39.52
27.1 | 34.2 | 37.14 | 39.2 | 39.38
27.2 | 34.0833 | 36.95 | 38.9833 | 39.2
27.3 | 33.9167 | 36.74 | 38.9 | 39.0333
27.4 | 33.8333 | 36.64 | 38.7 | 38.94
27.5 | 33.7 | 36.5 | 38.5 | 38.825
27.6 | 33.525 | 36.375 | 38.36 | 38.65
27.7 | 33.3833 | 36.25 | 38.2 | 38.45
27.8 | 33.2167 | 36.0333 | 38.025 | 38.325
27.9 | 33.04 | 35.9 | 37.9 | 38.2
28 | 32.94 | 35.68 | 37.74 | 38.1
28.1 | 32.76 | 35.58 | 37.625 | 37.9333
28.2 | 32.6 | 35.3667 | 37.4 | 37.8
28.3 | 32.5333 | 35.3 | 37.28 | 37.6
28.4 | 32.36 | 35.12 | 37.2 | 37.375
28.5 | 32.15 | 34.95 | 37.02 | 37.28
28.6 | 32 | 34.8 | 36.92 | 37.1667
28.7 | 31.8833 | 34.65 | 36.7333 | 37.075
28.8 | 31.7167 | 34.5 | 36.56 | 36.8667
28.9 | 31.58 | 34.38 | 36.44 | 36.775
29 | 31.44 | 34.25 | 36.3 | 36.575
29.1 | 31.2667 | 34.16 | 36.18 | 36.4
29.2 | 31.175 | 33.95 | 36.05 | 36.3
29.3 | 30.95 | 33.8 | 35.9 | 36.15
29.4 | 30.8167 | 33.6 | 35.7333 | 36
29.5 | 30.75 | 33.5 | 35.5333 | 35.9
29.6 | 30.56 | 33.38 | 35.44 | 35.7
29.7 | 30.38 | 33.28 | 35.35 | 35.5833
29.8 | 30.2667 | 33.075 | 35.2 | 35.5
29.9 | 30.18 | 32.925 | 35 | 35.32
30 | 30.0143 | 32.8 | 34.9 | 35.14

## SSD80 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 25.4 | 34.8 | 53.2 | 64.3
0.1 | 29.2 | 38.4 | 57.2 | 67.85
0.2 | 35.4 | 44.82 | 62.45 | 73
0.3 | 44.6 | 52.6 | 69.7 | 78.52
0.4 | 52.2 | 60.2 | 75.2 | 83.66
0.5 | 59.45 | 66.65 | 80.3 | 87.5333
0.6 | 65.6 | 71.9 | 84.4667 | 90.5
0.7 | 70.75 | 76.5667 | 87.6 | 93
0.8 | 75.3 | 80.5 | 90.3 | 94.9
0.9 | 79.06 | 83.5375 | 92.5 | 96.35
1 | 82.05 | 86.35 | 94.1571 | 97.5
1.1 | 85.125 | 88.7625 | 95.6 | 98.35
1.2 | 87.475 | 90.775 | 96.7714 | 98.95
1.3 | 89.7 | 92.4875 | 97.675 | 99.45
1.4 | 91.5 | 93.9875 | 98.425 | 99.775
1.5 | 93.275 | 95.2556 | 98.95 | 99.95
1.6 | 94.575 | 96.25 | 99.45 | 100
1.7 | 95.7 | 97.075 | 99.72 | 100
1.8 | 96.575 | 97.8857 | 99.9 | 99.9667
1.9 | 97.6333 | 98.5857 | 99.92 | 99.8
2 | 98.3333 | 98.925 | 100 | 99.64
2.1 | 98.8 | 99.35 | 99.98 | 99.4
2.2 | 99.12 | 99.66 | 99.85 | 99.15
2.3 | 99.46 | 99.9 | 99.7333 | 98.9
2.4 | 99.7 | 99.975 | 99.575 | 98.48
2.5 | 99.8 | 100 | 99.325 | 98.1
2.6 | 99.9 | 100 | 99.075 | 97.8
2.7 | 100 | 100 | 98.8 | 97.4
2.8 | 100 | 99.9143 | 98.54 | 97
2.9 | 99.9333 | 99.85 | 98.2333 | 96.6857
3 | 99.875 | 99.7 | 97.9 | 96.34
3.1 | 99.55 | 99.45 | 97.55 | 95.95
3.2 | 99.375 | 99.2 | 97.2 | 95.45
3.3 | 99.125 | 98.95 | 96.82 | 95.1667
3.4 | 98.95 | 98.65 | 96.45 | 94.8
3.5 | 98.5714 | 98.375 | 96.02 | 94.35
3.6 | 98.35 | 98.05 | 95.7 | 93.9
3.7 | 97.95 | 97.78 | 95.325 | 93.5
3.8 | 97.7 | 97.5 | 94.95 | 93.05
3.9 | 97.28 | 97.1 | 94.5333 | 92.6571
4 | 96.8 | 96.75 | 94.05 | 92.34
4.1 | 96.4333 | 96.3 | 93.725 | 91.95
4.2 | 96.0333 | 95.875 | 93.3 | 91.55
4.3 | 95.65 | 95.6 | 92.875 | 91.1333
4.4 | 95.2 | 95.1 | 92.5 | 90.8
4.5 | 94.75 | 94.675 | 92.1667 | 90.3667
4.6 | 94.35 | 94.44 | 91.6667 | 89.92
4.7 | 93.9429 | 94 | 91.275 | 89.58
4.8 | 93.3 | 93.6 | 90.95 | 89.1
4.9 | 93 | 93.2 | 90.54 | 88.875
5 | 92.66 | 92.65 | 90.2 | 88.4333
5.1 | 92.18 | 92.3 | 89.65 | 88.025
5.2 | 91.675 | 91.88 | 89.28 | 87.65
5.3 | 91.325 | 91.5 | 88.9 | 87.325
5.4 | 90.65 | 91 | 88.58 | 86.8
5.5 | 90.25 | 90.6 | 88.2 | 86.4
5.6 | 89.9 | 90.25 | 87.72 | 86.05
5.7 | 89.4 | 89.8 | 87.425 | 85.7
5.8 | 89.12 | 89.35 | 86.975 | 85.35
5.9 | 88.56 | 89 | 86.625 | 84.95
6 | 88.15 | 88.6333 | 86.2667 | 84.575
6.1 | 87.625 | 88.15 | 85.9667 | 84.225
6.2 | 87.225 | 87.825 | 85.5 | 83.8333
6.3 | 86.85 | 87.35 | 85.2 | 83.5
6.4 | 86.35 | 86.95 | 84.7 | 83.2
6.5 | 85.9 | 86.55 | 84.4 | 82.8
6.6 | 85.45 | 86.1333 | 84.05 | 82.45
6.7 | 85 | 85.8 | 83.6667 | 82.125
6.8 | 84.7 | 85.4667 | 83.2 | 81.75
6.9 | 84.2 | 84.95 | 82.85 | 81.35
7 | 83.7857 | 84.525 | 82.55 | 80.975
7.1 | 83.4333 | 84.2 | 82.2 | 80.6375
7.2 | 83.0667 | 83.75 | 81.8 | 80.275
7.3 | 82.65 | 83.4 | 81.4 | 80
7.4 | 82.175 | 83 | 81.1 | 79.65
7.5 | 81.62 | 82.6 | 80.7667 | 79.35
7.6 | 81.225 | 82.35 | 80.4 | 79
7.7 | 80.95 | 81.8333 | 80.1 | 78.66
7.8 | 80.45 | 81.4 | 79.825 | 78.25
7.9 | 80.15 | 81.06 | 79.375 | 77.95
8 | 79.68 | 80.6 | 79.08 | 77.625
8.1 | 79.28 | 80.35 | 78.66 | 77.275
8.2 | 78.9 | 79.85 | 78.3 | 76.88
8.3 | 78.45 | 79.5 | 78.02 | 76.6
8.4 | 78.1 | 79.25 | 77.65 | 76.25
8.5 | 77.66 | 78.75 | 77.25 | 75.9
8.6 | 77.325 | 78.4625 | 76.85 | 75.55
8.7 | 76.975 | 78.08 | 76.6 | 75.34
8.8 | 76.55 | 77.7 | 76.35 | 74.925
8.9 | 76.04 | 77.35 | 75.9 | 74.65
9 | 75.6 | 77 | 75.5667 | 74.25
9.1 | 75.3 | 76.7 | 75.2 | 73.975
9.2 | 74.85 | 76.25 | 74.8 | 73.7
9.3 | 74.4 | 75.9 | 74.6 | 73.3
9.4 | 74.15 | 75.5 | 74.175 | 73.05
9.5 | 73.7 | 75.22 | 73.925 | 72.7
9.6 | 73.3667 | 74.75 | 73.575 | 72.325
9.7 | 73 | 74.375 | 73.225 | 72.05
9.8 | 72.5 | 74.05 | 72.85 | 71.76
9.9 | 72.2 | 73.675 | 72.525 | 71.45
10 | 71.8 | 73.325 | 72.2 | 71.075
10.1 | 71.475 | 72.98 | 71.95 | 70.7875
10.2 | 71.15 | 72.65 | 71.6 | 70.475
10.3 | 70.775 | 72.225 | 71.25 | 70.1375
10.4 | 70.425 | 71.9 | 70.92 | 69.78
10.5 | 69.9667 | 71.65 | 70.65 | 69.475
10.6 | 69.75 | 71.25 | 70.3 | 69.15
10.7 | 69.375 | 70.8571 | 69.96 | 68.9
10.8 | 69.025 | 70.6222 | 69.6 | 68.575
10.9 | 68.675 | 70.25 | 69.35 | 68.3
11 | 68.35 | 69.88 | 69.04 | 67.975
11.1 | 67.95 | 69.58 | 68.775 | 67.7
11.2 | 67.64 | 69.2429 | 68.45 | 67.475
11.3 | 67.2 | 68.96 | 68.15 | 67.125
11.4 | 66.9 | 68.6 | 67.825 | 66.85
11.5 | 66.6 | 68.3 | 67.475 | 66.6
11.6 | 66.05 | 68 | 67.24 | 66.25
11.7 | 65.9 | 67.6 | 66.9 | 65.925
11.8 | 65.6 | 67.25 | 66.6333 | 65.575
11.9 | 65.325 | 66.875 | 66.2667 | 65.325
12 | 64.95 | 66.625 | 66.025 | 65.075
12.1 | 64.5 | 66.36 | 65.68 | 64.825
12.2 | 64.35 | 66.05 | 65.45 | 64.575
12.3 | 63.96 | 65.7 | 65.1 | 64.28
12.4 | 63.65 | 65.35 | 64.95 | 63.94
12.5 | 63.25 | 65 | 64.6 | 63.7
12.6 | 62.975 | 64.75 | 64.275 | 63.3667
12.7 | 62.65 | 64.5 | 64.075 | 63.0333
12.8 | 62.375 | 64.15 | 63.6833 | 62.8
12.9 | 62.0333 | 63.825 | 63.4667 | 62.54
13 | 61.8 | 63.5625 | 63.175 | 62.3
13.1 | 61.4 | 63.225 | 62.9429 | 62.05
13.2 | 61.15 | 62.9333 | 62.6125 | 61.8
13.3 | 60.7 | 62.65 | 62.3 | 61.5
13.4 | 60.4 | 62.3 | 61.95 | 61.24
13.5 | 60.2 | 62.05 | 61.7 | 61.02
13.6 | 59.9 | 61.8111 | 61.5 | 60.675
13.7 | 59.5667 | 61.575 | 61.175 | 60.4
13.8 | 59.25 | 61.2375 | 60.9 | 60.125
13.9 | 58.875 | 60.875 | 60.725 | 59.8667
14 | 58.75 | 60.5875 | 60.46 | 59.6
14.1 | 58.325 | 60.275 | 60.1286 | 59.3
14.2 | 58.0857 | 60.025 | 59.8833 | 59.1
14.3 | 57.78 | 59.7714 | 59.675 | 58.8333
14.4 | 57.525 | 59.5 | 59.3375 | 58.55
14.5 | 57.125 | 59.15 | 59.075 | 58.3333
14.6 | 56.9 | 58.9 | 58.8 | 58.1
14.7 | 56.7 | 58.6667 | 58.4857 | 57.8
14.8 | 56.3375 | 58.3 | 58.225 | 57.6
14.9 | 55.96 | 58 | 57.98 | 57.35
15 | 55.7429 | 57.72 | 57.675 | 57.02
15.1 | 55.475 | 57.4 | 57.4333 | 56.8
15.2 | 55.2 | 57.175 | 57.15 | 56.475
15.3 | 54.85 | 56.85 | 56.9 | 56.3
15.4 | 54.6 | 56.6 | 56.7 | 56.05
15.5 | 54.3667 | 56.35 | 56.45 | 55.8
15.6 | 54.025 | 56.1 | 56.175 | 55.55
15.7 | 53.775 | 55.8 | 55.925 | 55.275
15.8 | 53.525 | 55.6 | 55.7 | 55.025
15.9 | 53.25 | 55.35 | 55.45 | 54.8
16 | 53 | 55.0667 | 55.2 | 54.55
16.1 | 52.775 | 54.82 | 54.95 | 54.4
16.2 | 52.5 | 54.6 | 54.75 | 54.2
16.3 | 52.3 | 54.34 | 54.475 | 53.85
16.4 | 51.84 | 54.025 | 54.225 | 53.675
16.5 | 51.7 | 53.775 | 53.975 | 53.4333
16.6 | 51.45 | 53.48 | 53.825 | 53.25
16.7 | 51.1 | 53.25 | 53.5 | 53
16.8 | 50.92 | 53 | 53.3 | 52.75
16.9 | 50.7 | 52.7 | 53.05 | 52.5
17 | 50.38 | 52.5 | 52.8 | 52.375
17.1 | 50.125 | 52.275 | 52.55 | 52.125
17.2 | 49.9 | 52 | 52.3 | 51.85
17.3 | 49.55 | 51.75 | 52.1 | 51.625
17.4 | 49.4 | 51.5 | 51.9143 | 51.3125
17.5 | 49.2 | 51.25 | 51.66 | 51.1222
17.6 | 49 | 51 | 51.42 | 51
17.7 | 48.7 | 50.8 | 51.2 | 50.75
17.8 | 48.5 | 50.625 | 50.925 | 50.5
17.9 | 48.275 | 50.3 | 50.775 | 50.3
18 | 47.92 | 50.05 | 50.54 | 50.1
18.1 | 47.775 | 49.85 | 50.3 | 49.825
18.2 | 47.525 | 49.65 | 50.05 | 49.6
18.3 | 47.275 | 49.4222 | 49.8 | 49.35
18.4 | 47.1 | 49.125 | 49.66 | 49.1
18.5 | 46.775 | 48.925 | 49.425 | 48.95
18.6 | 46.525 | 48.68 | 49.175 | 48.7
18.7 | 46.3333 | 48.45 | 48.95 | 48.5
18.8 | 46.05 | 48.2 | 48.8 | 48.2
18.9 | 45.775 | 47.975 | 48.55 | 48.05
19 | 45.6 | 47.725 | 48.3 | 47.8
19.1 | 45.3 | 47.575 | 48.05 | 47.65
19.2 | 45.15 | 47.325 | 47.9 | 47.5
19.3 | 44.9 | 47.1 | 47.66 | 47.275
19.4 | 44.7 | 46.85 | 47.425 | 47.025
19.5 | 44.4375 | 46.7 | 47.175 | 46.9
19.6 | 44.275 | 46.45 | 46.94 | 46.7
19.7 | 44.04 | 46.2 | 46.8 | 46.5
19.8 | 43.8333 | 45.96 | 46.55 | 46.25
19.9 | 43.6 | 45.725 | 46.4 | 46
20 | 43.4 | 45.475 | 46.15 | 45.85
20.1 | 43.2 | 45.3 | 46.025 | 45.6556
20.2 | 43 | 45.1 | 45.7 | 45.475
20.3 | 42.8 | 44.925 | 45.575 | 45.225
20.4 | 42.5667 | 44.675 | 45.325 | 44.9875
20.5 | 42.3 | 44.5 | 45.175 | 44.8625
20.6 | 42.125 | 44.175 | 45 | 44.675
20.7 | 41.9 | 44.025 | 44.775 | 44.4444
20.8 | 41.7 | 43.875 | 44.525 | 44.3
20.9 | 41.5 | 43.675 | 44.4 | 44.05
21 | 41.2 | 43.425 | 44.1667 | 43.8
21.1 | 41.16 | 43.3 | 43.98 | 43.65
21.2 | 40.9333 | 43.05 | 43.85 | 43.5
21.3 | 40.7 | 42.82 | 43.6 | 43.325
21.4 | 40.62 | 42.6667 | 43.4 | 43.15
21.5 | 40.3 | 42.475 | 43.2 | 42.925
21.6 | 40.2 | 42.2 | 43.05 | 42.775
21.7 | 40 | 42 | 42.78 | 42.575
21.8 | 39.725 | 41.825 | 42.65 | 42.45
21.9 | 39.5 | 41.575 | 42.4 | 42.25
22 | 39.34 | 41.425 | 42.25 | 42.1
22.1 | 39.2 | 41.275 | 42.1 | 41.86
22.2 | 38.95 | 41.075 | 41.95 | 41.7
22.3 | 38.8 | 40.9 | 41.75 | 41.475
22.4 | 38.58 | 40.7 | 41.5 | 41.35
22.5 | 38.4 | 40.5 | 41.3333 | 41.125
22.6 | 38.2 | 40.3 | 41.2 | 41
22.7 | 38.025 | 40.125 | 41 | 40.7333
22.8 | 37.9 | 39.9 | 40.8 | 40.6
22.9 | 37.7 | 39.8 | 40.62 | 40.45
23 | 37.525 | 39.52 | 40.5 | 40.2
23.1 | 37.3 | 39.4 | 40.3333 | 40.1
23.2 | 37.125 | 39.225 | 40.1 | 39.9
23.3 | 36.8875 | 38.975 | 39.925 | 39.75
23.4 | 36.8 | 38.7625 | 39.8 | 39.5
23.5 | 36.55 | 38.6375 | 39.575 | 39.35
23.6 | 36.35 | 38.425 | 39.4111 | 39.2
23.7 | 36.1875 | 38.275 | 39.175 | 39.1
23.8 | 36.025 | 38.025 | 39 | 38.85
23.9 | 35.8444 | 37.9 | 38.875 | 38.7
24 | 35.725 | 37.725 | 38.7 | 38.5667
24.1 | 35.6 | 37.575 | 38.6 | 38.4
24.2 | 35.35 | 37.4 | 38.35 | 38.275
24.3 | 35.15 | 37.2 | 38.2 | 38.025
24.4 | 35.025 | 37.025 | 38 | 37.9
24.5 | 34.9 | 36.88 | 37.9 | 37.75
24.6 | 34.7667 | 36.7 | 37.7 | 37.6
24.7 | 34.55 | 36.5 | 37.55 | 37.4
24.8 | 34.3667 | 36.4 | 37.4 | 37.2
24.9 | 34.2 | 36.25 | 37.2 | 37.1
25 | 34.075 | 36.1 | 37.1 | 36.95
25.1 | 33.9 | 35.9 | 36.9 | 36.8
25.2 | 33.66 | 35.65 | 36.6667 | 36.6143
25.3 | 33.5 | 35.5 | 36.5 | 36.4833
25.4 | 33.3 | 35.3 | 36.4 | 36.28
25.5 | 33.25 | 35.125 | 36.25 | 36.1
25.6 | 33 | 35 | 36.0333 | 36
25.7 | 32.9 | 34.825 | 35.9 | 35.8
25.8 | 32.7 | 34.7 | 35.76 | 35.7
25.9 | 32.6 | 34.55 | 35.6 | 35.5
26 | 32.5 | 34.4 | 35.5 | 35.4
26.1 | 32.375 | 34.225 | 35.3167 | 35.3
26.2 | 32.2 | 34.1 | 35.175 | 35.1
26.3 | 31.975 | 33.95 | 34.925 | 34.925
26.4 | 31.8 | 33.75 | 34.875 | 34.8
26.5 | 31.7 | 33.675 | 34.725 | 34.6
26.6 | 31.55 | 33.5 | 34.55 | 34.4857
26.7 | 31.4 | 33.28 | 34.35 | 34.32
26.8 | 31.25 | 33.15 | 34.2125 | 34.18
26.9 | 31.1 | 33 | 34.0875 | 34.05
27 | 30.96 | 32.85 | 33.9625 | 33.8667
27.1 | 30.825 | 32.7 | 33.775 | 33.7
27.2 | 30.675 | 32.5 | 33.6125 | 33.6
27.3 | 30.525 | 32.4 | 33.5 | 33.425
27.4 | 30.3 | 32.25 | 33.3 | 33.3
27.5 | 30.275 | 32.1 | 33.2 | 33.2
27.6 | 30.025 | 31.975 | 33.1 | 33
27.7 | 29.9 | 31.85 | 32.85 | 32.9
27.8 | 29.8 | 31.7 | 32.8 | 32.8
27.9 | 29.675 | 31.5 | 32.6 | 32.65
28 | 29.425 | 31.325 | 32.5 | 32.5
28.1 | 29.3 | 31.175 | 32.3 | 32.35
28.2 | 29.2 | 31.025 | 32.2 | 32.2
28.3 | 29.1 | 30.9 | 32.1 | 32.06
28.4 | 28.95 | 30.8 | 31.9 | 32
28.5 | 28.8 | 30.6 | 31.8 | 31.8
28.6 | 28.7 | 30.525 | 31.65 | 31.625
28.7 | 28.6 | 30.375 | 31.5 | 31.5
28.8 | 28.35 | 30.225 | 31.3 | 31.4
28.9 | 28.25 | 30.1 | 31.175 | 31.2
29 | 28.1 | 29.9 | 31.1 | 31.1
29.1 | 28 | 29.85 | 31 | 31
29.2 | 27.9 | 29.7 | 30.8 | 30.9
29.3 | 27.725 | 29.55 | 30.7 | 30.7
29.4 | 27.625 | 29.4 | 30.5 | 30.65
29.5 | 27.475 | 29.25 | 30.4 | 30.5
29.6 | 27.325 | 29.2 | 30.3 | 30.35
29.7 | 27.1333 | 29 | 30.125 | 30.2
29.8 | 27.1 | 28.9 | 30.04 | 30.05
29.9 | 26.95 | 28.7333 | 29.9 | 29.9625
30 | 26.9 | 28.6 | 29.8 | 29.7875

## SSD120 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 23.5 | 29.6 | 40.4 | 47.7
0.1 | 27.175 | 33.275 | 43.95 | 51.35
0.2 | 33.125 | 39.5 | 49.425 | 56.9
0.3 | 41.875 | 47.95 | 57.1 | 63.9
0.4 | 49.75 | 54.9 | 63.75 | 69.7667
0.5 | 56.875 | 61.475 | 69.475 | 74.6
0.6 | 62.725 | 66.56 | 74.3 | 78.95
0.7 | 67.3 | 71.2 | 77.92 | 82.4
0.8 | 72 | 75.3286 | 81.5 | 85.6
0.9 | 75.65 | 78.9 | 84.4 | 88
1 | 79.2 | 81.95 | 86.95 | 90.4
1.1 | 81.98 | 84.6 | 89 | 92.1
1.2 | 84.7 | 87.1 | 91 | 93.75
1.3 | 86.7333 | 89 | 92.6 | 94.9
1.4 | 88.8 | 90.5833 | 94.1 | 96
1.5 | 90.6 | 92.06 | 95.2 | 96.925
1.6 | 92.275 | 93.5571 | 96.175 | 97.75
1.7 | 93.525 | 94.6625 | 97 | 98.35
1.8 | 95.025 | 95.5375 | 97.6 | 98.7857
1.9 | 95.75 | 96.5 | 98.225 | 99.1667
2 | 96.65 | 97.1143 | 98.65 | 99.4333
2.1 | 97.1 | 97.96 | 99.125 | 99.7
2.2 | 97.86 | 98.34 | 99.375 | 99.9
2.3 | 98.475 | 98.8429 | 99.6 | 99.9667
2.4 | 98.7667 | 99.24 | 99.775 | 100
2.5 | 99.0714 | 99.5 | 99.825 | 100
2.6 | 99.6 | 99.6 | 99.9 | 100
2.7 | 99.725 | 99.775 | 100 | 99.9
2.8 | 99.9 | 99.9 | 100 | 99.8125
2.9 | 99.9625 | 99.9 | 99.9 | 99.65
3 | 100 | 100 | 99.9 | 99.5125
3.1 | 100 | 100 | 99.8 | 99.2833
3.2 | 99.825 | 99.9 | 99.6 | 99.2
3.3 | 99.7167 | 99.8 | 99.4 | 98.975
3.4 | 99.6 | 99.7 | 99.26 | 98.725
3.5 | 99.35 | 99.5 | 98.9 | 98.4571
3.6 | 99.125 | 99.3 | 98.75 | 98.1
3.7 | 99 | 99.16 | 98.5 | 97.95
3.8 | 98.8714 | 98.85 | 98.25 | 97.7
3.9 | 98.6125 | 98.6 | 97.925 | 97.46
4 | 98.325 | 98.46 | 97.7 | 97.225
4.1 | 98 | 98.225 | 97.525 | 96.875
4.2 | 97.68 | 97.925 | 97.25 | 96.625
4.3 | 97.5 | 97.65 | 96.925 | 96.3
4.4 | 97.2 | 97.3 | 96.575 | 95.925
4.5 | 96.85 | 96.9857 | 96.425 | 95.6
4.6 | 96.525 | 96.74 | 96.0333 | 95.35
4.7 | 96.1 | 96.45 | 95.7 | 95
4.8 | 95.6 | 96.125 | 95.3667 | 94.6111
4.9 | 95.3 | 95.85 | 95.2 | 94.3125
5 | 95.02 | 95.58 | 94.95 | 94.025
5.1 | 94.7 | 95.2 | 94.625 | 93.86
5.2 | 94.375 | 94.85 | 94.18 | 93.5
5.3 | 94 | 94.4 | 93.925 | 93.2
5.4 | 93.65 | 94.175 | 93.6 | 93
5.5 | 93.1 | 93.8 | 93.2778 | 92.7
5.6 | 92.75 | 93.5333 | 92.9429 | 92.375
5.7 | 92.4 | 93.25 | 92.7 | 92.075
5.8 | 92.1333 | 92.8333 | 92.45 | 91.75
5.9 | 91.775 | 92.5333 | 92.05 | 91.45
6 | 91.35 | 92.0333 | 91.775 | 91.14
6.1 | 91.125 | 91.6 | 91.48 | 90.7
6.2 | 90.65 | 91.4 | 91.2 | 90.5
6.3 | 90.1333 | 91.2 | 90.9 | 90.3
6.4 | 89.8667 | 90.7333 | 90.5 | 89.875
6.5 | 89.5 | 90.42 | 90.25 | 89.625
6.6 | 89.275 | 90.175 | 89.9 | 89.38
6.7 | 88.925 | 89.84 | 89.65 | 89.05
6.8 | 88.425 | 89.5 | 89.3 | 88.8
6.9 | 87.75 | 89.3 | 89.05 | 88.4
7 | 87.475 | 88.9 | 88.725 | 88
7.1 | 87.2 | 88.6 | 88.45 | 87.875
7.2 | 86.9 | 88.125 | 88.1 | 87.55
7.3 | 86.5 | 87.74 | 87.8 | 87.175
7.4 | 86.15 | 87.4571 | 87.5 | 86.925
7.5 | 85.82 | 87.2222 | 87.2 | 86.7
7.6 | 85.3667 | 86.8625 | 87 | 86.45
7.7 | 85.1667 | 86.5625 | 86.7 | 86.1
7.8 | 84.8375 | 86.3857 | 86.4 | 85.675
7.9 | 84.4667 | 86.05 | 86.125 | 85.425
8 | 84.15 | 85.625 | 85.7333 | 85.18
8.1 | 83.725 | 85.3 | 85.475 | 84.95
8.2 | 83.4333 | 85.05 | 85.1333 | 84.6
8.3 | 83 | 84.7 | 84.85 | 84.35
8.4 | 82.6 | 84.35 | 84.54 | 84
8.5 | 82.4 | 84 | 84.2 | 83.65
8.6 | 82.06 | 83.66 | 83.94 | 83.42
8.7 | 81.75 | 83.425 | 83.7 | 83.175
8.8 | 81.275 | 83.1167 | 83.45 | 82.925
8.9 | 80.88 | 82.775 | 83.15 | 82.56
9 | 80.7 | 82.66 | 82.775 | 82.2333
9.1 | 80.4 | 82.1625 | 82.525 | 82.1
9.2 | 79.9 | 81.975 | 82.2 | 81.8
9.3 | 79.35 | 81.65 | 81.9667 | 81.475
9.4 | 79.025 | 81.3167 | 81.6 | 81.2
9.5 | 78.8 | 81.1 | 81.325 | 80.9
9.6 | 78.55 | 80.7 | 81.075 | 80.6333
9.7 | 78.05 | 80.3 | 80.725 | 80.375
9.8 | 77.65 | 80.05 | 80.475 | 80.025
9.9 | 77.38 | 79.75 | 80.225 | 79.6667
10 | 77.15 | 79.275 | 80 | 79.5
10.1 | 76.8 | 79.025 | 79.75 | 79.2667
10.2 | 76.5 | 78.86 | 79.4 | 78.95
10.3 | 76.3 | 78.5 | 79.15 | 78.575
10.4 | 75.9 | 78.2 | 78.8 | 78.425
10.5 | 75.55 | 77.95 | 78.46 | 78
10.6 | 75.2 | 77.6 | 78.1 | 77.9
10.7 | 74.85 | 77.35 | 77.9667 | 77.5667
10.8 | 74.5167 | 77.075 | 77.65 | 77.325
10.9 | 74.2625 | 76.75 | 77.375 | 77.075
11 | 73.925 | 76.2 | 77.1667 | 76.825
11.1 | 73.675 | 76.06 | 76.9 | 76.575
11.2 | 73.3375 | 75.725 | 76.5 | 76.225
11.3 | 72.95 | 75.45 | 76.325 | 75.975
11.4 | 72.55 | 75.175 | 76.15 | 75.65
11.5 | 72.2 | 74.925 | 75.725 | 75.5
11.6 | 71.94 | 74.675 | 75.5 | 75.15
11.7 | 71.675 | 74.425 | 75.3 | 74.9
11.8 | 71.3333 | 74.15 | 74.85 | 74.65
11.9 | 70.9 | 73.75 | 74.54 | 74.4
12 | 70.75 | 73.5 | 74.35 | 74.2667
12.1 | 70.6 | 73.225 | 74.025 | 73.9
12.2 | 70.2 | 72.9 | 73.8857 | 73.6333
12.3 | 69.84 | 72.6 | 73.72 | 73.3667
12.4 | 69.5 | 72.2 | 73.38 | 73.2
12.5 | 69.15 | 71.95 | 73.2 | 72.9
12.6 | 68.925 | 71.8 | 72.98 | 72.6
12.7 | 68.675 | 71.5667 | 72.75 | 72.425
12.8 | 68.35 | 71.3 | 72.5 | 72.25
12.9 | 67.95 | 71.05 | 72.1 | 71.925
13 | 67.7 | 70.625 | 71.7714 | 71.6
13.1 | 67.4 | 70.375 | 71.525 | 71.45
13.2 | 67.1 | 70.025 | 71.4 | 71.2
13.3 | 66.9 | 69.775 | 71.0833 | 70.85
13.4 | 66.5 | 69.65 | 70.7667 | 70.6
13.5 | 66.0667 | 69.3 | 70.475 | 70.3
13.6 | 65.8 | 69.05 | 70.225 | 70.15
13.7 | 65.5667 | 68.8 | 69.975 | 69.9
13.8 | 65.4 | 68.55 | 69.725 | 69.65
13.9 | 65.2 | 68.225 | 69.575 | 69.4
14 | 64.9 | 67.95 | 69.34 | 69.12
14.1 | 64.65 | 67.6 | 69 | 68.8
14.2 | 64.25 | 67.4 | 68.85 | 68.56
14.3 | 64 | 67.1333 | 68.5 | 68.3
14.4 | 63.75 | 66.86 | 68.25 | 68.1286
14.5 | 63.375 | 66.5 | 68.0667 | 67.9
14.6 | 63.1 | 66.4 | 67.8 | 67.5556
14.7 | 62.825 | 66.15 | 67.54 | 67.3143
14.8 | 62.575 | 65.825 | 67.3 | 67.125
14.9 | 62.325 | 65.575 | 66.95 | 66.9
15 | 62.0333 | 65.325 | 66.7 | 66.725
15.1 | 61.75 | 65.1 | 66.55 | 66.5
15.2 | 61.425 | 64.75 | 66.32 | 66.15
15.3 | 61.28 | 64.5 | 66.075 | 65.875
15.4 | 61.05 | 64.35 | 65.825 | 65.7333
15.5 | 60.8333 | 64.05 | 65.575 | 65.4
15.6 | 60.5 | 63.675 | 65.325 | 65.1667
15.7 | 60.1667 | 63.3833 | 65.075 | 64.95
15.8 | 59.9667 | 63.175 | 64.925 | 64.75
15.9 | 59.725 | 63.0125 | 64.6 | 64.5429
16 | 59.4333 | 62.675 | 64.35 | 64.3
16.1 | 59.15 | 62.425 | 64.1 | 64.075
16.2 | 58.9 | 62.275 | 63.8667 | 63.925
16.3 | 58.6 | 62.025 | 63.7 | 63.65
16.4 | 58.44 | 61.8 | 63.5 | 63.34
16.5 | 58.1 | 61.65 | 63.275 | 63.2
16.6 | 57.85 | 61.35 | 63 | 63
16.7 | 57.6 | 61.025 | 62.85 | 62.8
16.8 | 57.45 | 60.85 | 62.6 | 62.575
16.9 | 57.125 | 60.5333 | 62.275 | 62.325
17 | 56.975 | 60.3 | 62.125 | 62
17.1 | 56.625 | 60.2 | 61.9 | 61.9
17.2 | 56.475 | 60 | 61.725 | 61.7
17.3 | 56.15 | 59.65 | 61.5 | 61.52
17.4 | 55.88 | 59.4 | 61.225 | 61.2
17.5 | 55.7 | 59.16 | 61 | 60.9667
17.6 | 55.5 | 58.85 | 60.75 | 60.75
17.7 | 55.26 | 58.625 | 60.525 | 60.5
17.8 | 54.95 | 58.475 | 60.375 | 60.25
17.9 | 54.7 | 58.125 | 60.2 | 60.1
18 | 54.45 | 57.95 | 60 | 59.95
18.1 | 54.2 | 57.65 | 59.65 | 59.7
18.2 | 53.85 | 57.4667 | 59.5 | 59.45
18.3 | 53.7 | 57.375 | 59.275 | 59.32
18.4 | 53.6 | 57.125 | 59 | 59.06
18.5 | 53.45 | 56.9 | 58.7 | 58.825
18.6 | 53.2 | 56.775 | 58.5 | 58.6
18.7 | 52.86 | 56.4875 | 58.275 | 58.425
18.8 | 52.6 | 56.1571 | 58.14 | 58.175
18.9 | 52.375 | 55.925 | 57.95 | 57.925
19 | 52.125 | 55.675 | 57.8 | 57.775
19.1 | 51.975 | 55.525 | 57.55 | 57.625
19.2 | 51.74 | 55.3 | 57.4 | 57.36
19.3 | 51.5 | 55.1 | 57.1 | 57.125
19.4 | 51.3 | 54.96 | 56.9 | 56.875
19.5 | 50.9 | 54.65 | 56.6 | 56.7
19.6 | 50.7333 | 54.375 | 56.425 | 56.54
19.7 | 50.575 | 54.05 | 56.24 | 56.3333
19.8 | 50.425 | 53.875 | 56.1 | 56.1
19.9 | 50.175 | 53.7 | 55.85 | 55.9
20 | 49.925 | 53.5 | 55.6 | 55.7
20.1 | 49.675 | 53.35 | 55.4 | 55.5
20.2 | 49.5125 | 53.1 | 55.2333 | 55.275
20.3 | 49.3714 | 52.8667 | 54.95 | 55.125
20.4 | 49.1 | 52.7 | 54.8 | 54.925
20.5 | 48.875 | 52.475 | 54.6667 | 54.7
20.6 | 48.7222 | 52.3 | 54.55 | 54.525
20.7 | 48.5 | 52.075 | 54.3 | 54.4
20.8 | 48.35 | 51.825 | 54.125 | 54.15
20.9 | 48.05 | 51.58 | 53.8 | 53.9
21 | 47.8333 | 51.35 | 53.625 | 53.8
21.1 | 47.7125 | 51.1 | 53.4 | 53.55
21.2 | 47.3833 | 50.95 | 53.2714 | 53.35
21.3 | 47.175 | 50.8 | 53.05 | 53.125
21.4 | 46.925 | 50.475 | 52.9 | 52.975
21.5 | 46.7 | 50.3556 | 52.575 | 52.825
21.6 | 46.7 | 50.175 | 52.4 | 52.675
21.7 | 46.45 | 49.9 | 52.275 | 52.44
21.8 | 46.225 | 49.76 | 52.1 | 52.2
21.9 | 45.975 | 49.45 | 51.95 | 52
22 | 45.8 | 49.3 | 51.7 | 51.9
22.1 | 45.6333 | 49.125 | 51.5 | 51.75
22.2 | 45.425 | 49 | 51.32 | 51.6
22.3 | 45.175 | 48.825 | 51.125 | 51.35
22.4 | 45 | 48.5667 | 50.975 | 51.1
22.5 | 44.9 | 48.325 | 50.725 | 50.85
22.6 | 44.75 | 48.175 | 50.575 | 50.7
22.7 | 44.6 | 48.025 | 50.4 | 50.6
22.8 | 44.3 | 47.88 | 50.225 | 50.4
22.9 | 44.1 | 47.6667 | 49.975 | 50.25
23 | 43.95 | 47.38 | 49.725 | 50
23.1 | 43.7 | 47.2 | 49.5 | 49.86
23.2 | 43.5 | 47.1 | 49.45 | 49.6
23.3 | 43.3 | 46.95 | 49.3 | 49.4
23.4 | 43.175 | 46.7 | 49.05 | 49.3
23.5 | 43.025 | 46.55 | 48.9 | 49.1
23.6 | 42.9167 | 46.3 | 48.675 | 48.925
23.7 | 42.575 | 46.2 | 48.52 | 48.675
23.8 | 42.3667 | 46 | 48.3 | 48.5
23.9 | 42.2 | 45.86 | 48.1 | 48.375
24 | 42.0143 | 45.625 | 48 | 48.24
24.1 | 41.8 | 45.475 | 47.825 | 48.1
24.2 | 41.65 | 45.225 | 47.575 | 47.95
24.3 | 41.4 | 45.075 | 47.425 | 47.7
24.4 | 41.3 | 44.95 | 47.175 | 47.55
24.5 | 41.12 | 44.775 | 47.0125 | 47.4
24.6 | 40.9857 | 44.6429 | 46.9 | 47.26
24.7 | 40.9 | 44.38 | 46.725 | 47
24.8 | 40.7 | 44.15 | 46.475 | 46.875
24.9 | 40.525 | 43.98 | 46.3125 | 46.7
25 | 40.3667 | 43.85 | 46.1857 | 46.5
25.1 | 40.2 | 43.7 | 46.02 | 46.3
25.2 | 39.9667 | 43.5 | 45.8889 | 46.2
25.3 | 39.7 | 43.3 | 45.7667 | 45.95
25.4 | 39.6333 | 43.15 | 45.48 | 45.8
25.5 | 39.55 | 42.9 | 45.36 | 45.675
25.6 | 39.35 | 42.76 | 45.1667 | 45.425
25.7 | 39.1 | 42.6 | 45 | 45.3
25.8 | 38.975 | 42.3667 | 44.825 | 45.2111
25.9 | 38.725 | 42.225 | 44.7 | 45.0375
26 | 38.6 | 42.1 | 44.6 | 44.825
26.1 | 38.5 | 41.875 | 44.3 | 44.5875
26.2 | 38.3 | 41.7 | 44.1 | 44.4625
26.3 | 38.05 | 41.48 | 44 | 44.275
26.4 | 37.9 | 41.35 | 43.875 | 44.15
26.5 | 37.75 | 41.3 | 43.7 | 44
26.6 | 37.675 | 41.05 | 43.5 | 43.9
26.7 | 37.4333 | 40.9 | 43.3 | 43.7
26.8 | 37.3 | 40.775 | 43.2 | 43.575
26.9 | 37.24 | 40.6 | 43 | 43.4
27 | 37 | 40.475 | 42.875 | 43.275
27.1 | 36.85 | 40.325 | 42.7 | 43.1
27.2 | 36.675 | 40.08 | 42.58 | 42.94
27.3 | 36.55 | 39.9 | 42.45 | 42.8
27.4 | 36.425 | 39.8 | 42.3 | 42.65
27.5 | 36.3 | 39.64 | 42.15 | 42.4
27.6 | 36.0667 | 39.4 | 42 | 42.2875
27.7 | 35.8 | 39.3 | 41.825 | 42.1625
27.8 | 35.6667 | 39.2 | 41.6286 | 42.0375
27.9 | 35.575 | 39 | 41.4 | 41.9125
28 | 35.475 | 38.9 | 41.3 | 41.6875
28.1 | 35.34 | 38.8 | 41.1 | 41.5333
28.2 | 35.2 | 38.625 | 41 | 41.35
28.3 | 35.05 | 38.4 | 40.9 | 41.225
28.4 | 34.9 | 38.3 | 40.7 | 41.1
28.5 | 34.75 | 38.1 | 40.6 | 40.95
28.6 | 34.6 | 38 | 40.4167 | 40.75
28.7 | 34.5 | 37.9 | 40.2875 | 40.625
28.8 | 34.3 | 37.7 | 40.15 | 40.5
28.9 | 34.225 | 37.5 | 40.025 | 40.375
29 | 34.075 | 37.35 | 39.8 | 40.225
29.1 | 33.9 | 37.2 | 39.65 | 40.1
29.2 | 33.8 | 37.1 | 39.5 | 39.9
29.3 | 33.6 | 36.95 | 39.4 | 39.8
29.4 | 33.5 | 36.8 | 39.2 | 39.6
29.5 | 33.4 | 36.65 | 39.1 | 39.5
29.6 | 33.25 | 36.5 | 39 | 39.375
29.7 | 33.0667 | 36.3 | 38.875 | 39.25
29.8 | 32.9125 | 36.25 | 38.7 | 39.025
29.9 | 32.7857 | 36 | 38.5 | 38.8
30 | 32.6333 | 35.9 | 38.4 | 38.7

## Wedge 45 Lateral Profiles
Lateral Profiles
Wedge 45
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:3 cm FieldSize:4x4 cm2 | Depth:3 cm FieldSize:10x10 cm2 | Depth:3 cm FieldSize:20x20 cm2 | Depth:3 cm FieldSize:40x20 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:40x20 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:40x20 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:40x20 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:40x20 cm2
-17.9 | [18 empty] | 12.2 | 15.3167
-17.8 | [18 empty] | 12.42 | 15.4714
-17.7 | [18 empty] | 12.6167 | 15.6333
-17.6 | [18 empty] | 12.8 | 15.8
-17.5 | [18 empty] | 12.9333 | 16.0667
-17.4 | [18 empty] | 13.2 | 16.3333
-17.3 | [18 empty] | 13.42 | 16.52
-17.2 | [18 empty] | 13.6 | 16.75
-17.1 | [18 empty] | 13.84 | 16.92
-17 | [18 empty] | 14.05 | 17.14
-16.9 | [14 empty] | 9.9 | 12.7 |  |  | 14.3 | 17.34
-16.8 | [14 empty] | 10.0667 | 12.8333 |  |  | 14.5667 | 17.55
-16.7 | [14 empty] | 10.3 | 13.025 |  |  | 14.76 | 17.85
-16.6 | [14 empty] | 10.5 | 13.26 |  |  | 14.975 | 18.0667
-16.5 | [14 empty] | 10.66 | 13.44 |  |  | 15.15 | 18.34
-16.4 | [14 empty] | 10.825 | 13.675 |  |  | 15.52 | 18.575
-16.3 | [14 empty] | 11.0333 | 13.8 |  |  | 15.8 | 18.9
-16.2 | [14 empty] | 11.1857 | 14.0167 |  |  | 15.95 | 19.2333
-16.1 | [14 empty] | 11.4 | 14.2167 |  |  | 16.3 | 19.4
-16 | [14 empty] | 11.5714 | 14.3833 |  |  | 16.64 | 19.6333
-15.9 | [10 empty] | 7.5 | 9.9808 |  |  | 11.775 | 14.6 |  |  | 16.95 | 19.9667
-15.8 | [10 empty] | 7.6 | 10.0725 |  |  | 12 | 14.85 |  |  | 17.15 | 20.3
-15.7 | [10 empty] | 7.7667 | 10.1975 |  |  | 12.225 | 15.06 |  |  | 17.5 | 20.6333
-15.6 | [10 empty] | 7.9 | 10.3724 |  |  | 12.55 | 15.3 |  |  | 17.8 | 20.9667
-15.5 | [10 empty] | 8.025 | 10.4641 |  |  | 12.7333 | 15.5143 |  |  | 18.1 | 21.2667
-15.4 | [6 empty] | 6.4 | 8.8 |  |  | 8.175 | 10.6307 |  |  | 12.9 | 15.7333 |  |  | 18.44 | 21.66
-15.3 | [6 empty] | 6.6 | 8.9 |  |  | 8.3 | 10.8473 |  |  | 13.2 | 15.98 |  |  | 18.8 | 22
-15.2 |  |  | 7.1667 | 9.8 |  |  | 6.7 | 9 |  |  | 8.44 | 10.9773 |  |  | 13.425 | 16.18 |  |  | 19.3 | 22.3333
-15.1 |  |  | 7.275 | 9.88 |  |  | 6.8 | 9.15 |  |  | 8.62 | 11.0972 |  |  | 13.64 | 16.4667 |  |  | 19.6 | 22.7333
-15 |  |  | 7.4 | 9.98 |  |  | 6.9 | 9.2333 |  |  | 8.8 | 11.2972 |  |  | 13.9 | 16.75 |  |  | 19.96 | 23.05
-14.9 |  |  | 7.5 | 10.1667 |  |  | 7 | 9.4 |  |  | 8.9 | 11.4472 |  |  | 14.175 | 17.05 |  |  | 20.44 | 23.4667
-14.8 |  |  | 7.7 | 10.3 |  |  | 7.1 | 9.52 |  |  | 9.1 | 11.5721 |  |  | 14.5 | 17.3333 |  |  | 20.85 | 23.95
-14.7 |  |  | 7.8 | 10.4667 |  |  | 7.2667 | 9.7 |  |  | 9.2833 | 11.8399 |  |  | 14.7667 | 17.6667 |  |  | 21.4 | 24.36
-14.6 |  |  | 7.9 | 10.54 |  |  | 7.4667 | 9.82 |  |  | 9.475 | 11.997 |  |  | 15.06 | 17.98 |  |  | 21.92 | 24.98
-14.5 |  |  | 8 | 10.7 |  |  | 7.6 | 10 |  |  | 9.6 | 12.237 |  |  | 15.3 | 18.2 |  |  | 22.4667 | 25.56
-14.4 |  |  | 8.1 | 10.8333 |  |  | 7.7 | 10.1 |  |  | 9.84 | 12.4469 |  |  | 15.7 | 18.54 |  |  | 23 | 26.34
-14.3 |  |  | 8.2667 | 11 |  |  | 7.9 | 10.24 |  |  | 10.0667 | 12.6568 |  |  | 16 | 18.8 |  |  | 23.875 | 26.92
-14.2 |  |  | 8.4 | 11.14 |  |  | 8 | 10.4 |  |  | 10.25 | 12.8968 |  |  | 16.34 | 19.34 |  |  | 24.7 | 27.6667
-14.1 |  |  | 8.6 | 11.3167 |  |  | 8.1667 | 10.54 |  |  | 10.46 | 13.0967 |  |  | 16.75 | 19.7 |  |  | 25.6333 | 28.68
-14 |  |  | 8.7 | 11.4 |  |  | 8.3 | 10.76 |  |  | 10.675 | 13.2967 |  |  | 17.1333 | 20.1333 |  |  | 26.8333 | 29.5
-13.9 |  |  | 8.9 | 11.625 |  |  | 8.5 | 10.9 |  |  | 10.92 | 13.5566 |  |  | 17.64 | 20.5667 |  |  | 27.9857 | 30.8
-13.8 |  |  | 9 | 11.72 |  |  | 8.625 | 11.1 |  |  | 11.1667 | 13.7566 |  |  | 18.04 | 21 |  |  | 29.55 | 32.4
-13.7 |  |  | 9.2 | 11.925 |  |  | 8.825 | 11.2167 |  |  | 11.4 | 13.9565 |  |  | 18.5 | 21.5 |  |  | 31.4 | 34.12
-13.6 |  |  | 9.35 | 12.1167 |  |  | 9 | 11.5 |  |  | 11.7 | 14.3164 |  |  | 19 | 21.96 |  |  | 33.8 | 36.4333
-13.5 |  |  | 9.5 | 12.3 |  |  | 9.125 | 11.7 |  |  | 11.96 | 14.5564 |  |  | 19.6 | 22.6 |  |  | 36.9 | 39.1833
-13.4 |  |  | 9.7 | 12.5 |  |  | 9.3 | 11.825 |  |  | 12.2667 | 14.7963 |  |  | 20.3 | 23.14 |  |  | 41.15 | 43.15
-13.3 |  |  | 9.9 | 12.675 |  |  | 9.5 | 12.06 |  |  | 12.4833 | 15.1212 |  |  | 21.075 | 23.85 |  |  | 46.58 | 48.9
-13.2 |  |  | 10.1 | 12.85 |  |  | 9.7 | 12.28 |  |  | 12.9 | 15.4561 |  |  | 21.9 | 24.6667 |  |  | 54.775 | 55.9
-13.1 |  |  | 10.3 | 13.05 |  |  | 9.9333 | 12.48 |  |  | 13.22 | 15.7961 |  |  | 22.9 | 25.6 |  |  | 64.4 | 65.3286
-13 |  |  | 10.45 | 13.3 |  |  | 10.225 | 12.68 |  |  | 13.52 | 16.156 |  |  | 24.025 | 26.7 |  |  | 75.7 | 75.5333
-12.9 |  |  | 10.65 | 13.5 |  |  | 10.4 | 13 |  |  | 13.94 | 16.5958 |  |  | 25.325 | 28 |  |  | 86.88 | 86.5
-12.8 |  |  | 10.8667 | 13.72 |  |  | 10.675 | 13.2167 |  |  | 14.36 | 16.9757 |  |  | 26.8 | 29.66 |  |  | 98.2333 | 96.5333
-12.7 |  |  | 11.1 | 13.92 |  |  | 11.02 | 13.48 |  |  | 14.76 | 17.3957 |  |  | 28.9 | 31.4 |  |  | 107.6667 | 105.7
-12.6 |  |  | 11.3 | 14.2 |  |  | 11.275 | 13.775 |  |  | 15.34 | 17.8456 |  |  | 31.5 | 33.875 |  |  | 115.1 | 112.8
-12.5 |  |  | 11.5 | 14.425 |  |  | 11.55 | 14.075 |  |  | 15.85 | 18.3954 |  |  | 34.74 | 37.14 |  |  | 120.2 | 117.7
-12.4 |  |  | 11.8667 | 14.675 |  |  | 11.8667 | 14.44 |  |  | 16.4667 | 19.0286 |  |  | 39.525 | 41.54 |  |  | 123.3833 | 121.1667
-12.3 |  |  | 12.075 | 14.95 |  |  | 12.25 | 14.8333 |  |  | 17.2 | 19.7151 |  |  | 46.16 | 47.98 |  |  | 125.82 | 123.6667
-12.2 |  |  | 12.375 | 15.2667 |  |  | 12.6 | 15.2 |  |  | 18 | 20.5948 |  |  | 55.1 | 56.82 |  |  | 127.84 | 125.4
-12.1 |  |  | 12.725 | 15.56 |  |  | 13.1 | 15.62 |  |  | 19 | 21.4146 |  |  | 66.1333 | 66.88 |  |  | 129.2 | 126.9
-12 |  |  | 13 | 15.96 |  |  | 13.5333 | 16.12 |  |  | 20.15 | 22.5444 |  |  | 78.8833 | 78.55 |  |  | 130.15 | 127.88
-11.9 |  |  | 13.5333 | 16.34 |  |  | 14.18 | 16.76 |  |  | 21.46 | 23.8341 |  |  | 91.35 | 90.94 |  |  | 131.1 | 128.75
-11.8 |  |  | 13.95 | 16.72 |  |  | 14.85 | 17.4 |  |  | 23.1 | 25.5537 |  |  | 102.1 | 102.075 |  |  | 131.68 | 129.3
-11.7 |  |  | 14.375 | 17.3 |  |  | 15.6 | 18.12 |  |  | 25.2 | 27.4598 |  |  | 112.2 | 111.32 |  |  | 132.1714 | 129.8333
-11.6 |  |  | 15.0333 | 17.85 |  |  | 16.5 | 19 |  |  | 27.92 | 30.0925 |  |  | 120.06 | 118.7833 |  |  | 132.5 | 130.2
-11.5 |  |  | 15.6 | 18.45 |  |  | 17.65 | 20.2 |  |  | 31.54 | 33.4666 |  |  | 125.1 | 123.3 |  |  | 132.8667 | 130.425
-11.4 |  |  | 16.5 | 19.275 |  |  | 18.9333 | 21.36 |  |  | 36.54 | 38.0705 |  |  | 128.3333 | 126.4 |  | 6.5 | 133.1 | 130.64
-11.3 |  |  | 17.4 | 20.1 |  |  | 20.5 | 23.0333 |  |  | 43.6 | 45.5886 |  |  | 130.975 | 128.7667 |  | 6.575 | 133.28 | 130.75
-11.2 |  |  | 18.7 | 21.3 |  |  | 22.5 | 25 |  |  | 53.8 | 55.2662 |  |  | 132.68 | 130.48 |  | 6.7 | 133.3 | 130.9
-11.1 |  |  | 20.2 | 22.8 |  |  | 25.4333 | 27.86 |  |  | 67.24 | 67.1632 |  |  | 134.0571 | 131.5 |  | 6.8333 | 133.4 | 131
-11 |  |  | 22.1667 | 24.65 |  |  | 29.1 | 31.325 |  |  | 81.4 | 81.1797 |  |  | 135.15 | 132.6 |  | 7 | 133.4 | 131.1
-10.9 |  |  | 24.8667 | 27.18 |  |  | 34.9 | 36.575 |  |  | 95.72 | 94.4014 |  | 5.0987 | 135.85 | 133.32 |  | 7.175 | 133.4 | 131.1
-10.8 |  |  | 28.5333 | 30.775 |  |  | 42.7 | 45.1 |  |  | 108.6 | 106.9067 |  | 5.1987 | 136.3 | 133.875 |  | 7.3 | 133.3 | 131.1
-10.7 |  |  | 34.2667 | 35.7 |  |  | 55.1 | 56.15 |  |  | 118.875 | 117.5704 |  | 5.2987 | 136.7429 | 134.36 |  | 7.4 | 133.3 | 130.9833
-10.6 |  |  | 42.45 | 43.82 |  |  | 68.4 | 69.4 |  |  | 127.22 | 125.5284 |  | 5.3653 | 137.08 | 134.66 |  | 7.5 | 133.2 | 130.8
-10.5 |  |  | 54.65 | 55.08 |  |  | 83.3 | 83.9 |  |  | 132.4 | 130.2675 |  | 5.3986 | 137.25 | 134.88 |  | 7.725 | 133.125 | 130.7
-10.4 |  |  | 68.7667 | 69.3 |  |  | 99.2333 | 98.875 |  | 3.5667 | 135.5 | 133.467 |  | 5.4986 | 137.3333 | 134.9 |  | 7.8667 | 133 | 130.6
-10.3 |  |  | 84.8 | 84.3 |  |  | 113.1667 | 111.85 |  | 3.6 | 137.8 | 135.866 |  | 5.5986 | 137.4 | 135 |  | 8 | 132.75 | 130.5
-10.2 |  |  | 100.6 | 98.84 |  | 2.925 | 123.9 | 123.1 |  | 3.68 | 139.6 | 137.466 |  | 5.7786 | 137.3333 | 135 |  | 8.125 | 132.54 | 130.4
-10.1 |  | 3.2 | 114.9 | 112.04 |  | 3 | 132.05 | 130.7333 |  | 3.8 | 140.62 | 138.605 |  | 5.8985 | 137.3 | 134.9 |  | 8.35 | 132.3333 | 130.25
-10 |  | 3.3 | 126 | 124.08 |  | 3.1 | 137.2667 | 135.3 |  | 3.875 | 141.475 | 139.265 |  | 5.9652 | 137.2333 | 134.8 |  | 8.4 | 132.1 | 130.04
-9.9 |  | 3.3 | 134.1 | 131.76 |  | 3.1 | 140.35 | 138.38 |  | 3.9 | 141.96 | 139.765 |  | 6.0985 | 137.16 | 134.7 |  | 8.5 | 131.8 | 129.7667
-9.8 |  | 3.4 | 138.8 | 136.5 |  | 3.2 | 142.45 | 140.4 |  | 4 | 142.325 | 140.245 |  | 6.2984 | 136.9833 | 134.7 |  | 8.74 | 131.66 | 129.4333
-9.7 |  | 3.5 | 141.95 | 139.24 |  | 3.2 | 144.05 | 141.8 |  | 4.1 | 142.575 | 140.465 |  | 6.3984 | 136.8167 | 134.38 |  | 8.975 | 131.4667 | 129.28
-9.6 |  | 3.5 | 143.725 | 141.14 |  | 3.3 | 144.9 | 142.7 |  | 4.2 | 142.7 | 140.565 |  | 6.4984 | 136.625 | 134.2 |  | 9.125 | 131.2167 | 129.15
-9.5 |  | 3.54 | 145.025 | 142.475 |  | 3.4 | 145.55 | 143.4 |  | 4.2 | 142.8 | 140.565 |  | 6.6983 | 136.3 | 134.1 |  | 9.3 | 131.05 | 128.9667
-9.4 |  | 3.6 | 145.825 | 143.4 |  | 3.4 | 145.88 | 143.8 |  | 4.35 | 142.72 | 140.5483 |  | 6.7983 | 136.15 | 133.8667 |  | 9.475 | 130.85 | 128.8
-9.3 |  | 3.7 | 146.3 | 143.7 |  | 3.5 | 146.2 | 144.05 |  | 4.45 | 142.6 | 140.465 |  | 6.9982 | 135.88 | 133.625 |  | 9.7 | 130.4 | 128.5
-9.2 |  | 3.7 | 146.4 | 143.9333 |  | 3.5 | 146.3 | 144.1 |  | 4.5667 | 142.4 | 140.285 |  | 7.1982 | 135.5333 | 133.4 |  | 9.85 | 130.1667 | 128.32
-9.1 |  | 3.8 | 146.4 | 144 |  | 3.6 | 146.3 | 144.1 |  | 4.7 | 142.075 | 140.185 |  | 7.3482 | 135.2 | 132.9833 |  | 10 | 130.02 | 128.1
-9 |  | 3.9 | 146.4 | 144 |  | 3.6667 | 146.3 | 144.1 |  | 4.7 | 141.9 | 139.9317 |  | 7.4981 | 134.7 | 132.775 |  | 10.3 | 129.8 | 127.74
-8.9 |  | 3.9 | 146.2 | 143.9 |  | 3.8 | 146.1 | 143.9 |  | 4.8 | 141.7 | 139.645 |  | 7.6231 | 134.48 | 132.45 |  | 10.6 | 129.56 | 127.4667
-8.8 |  | 4 | 146.1 | 143.68 |  | 3.825 | 145.8333 | 143.6667 |  | 4.95 | 141.35 | 139.3317 |  | 7.923 | 134.3 | 132.175 |  | 10.92 | 129.25 | 127.25
-8.7 |  | 4.1 | 145.8 | 143.4 |  | 3.9 | 145.525 | 143.3 |  | 5.1 | 140.775 | 139.065 |  | 8.098 | 134.04 | 131.875 |  | 11.25 | 128.875 | 127.1667
-8.6 |  | 4.1667 | 145.4 | 143.075 |  | 4 | 145.175 | 142.82 |  | 5.3 | 140.38 | 138.515 |  | 8.3229 | 133.62 | 131.55 |  | 11.5 | 128.475 | 126.9
-8.5 |  | 4.2667 | 145 | 142.75 |  | 4.125 | 144.75 | 142.45 |  | 5.4 | 140 | 138.205 |  | 8.5729 | 133.25 | 131.2 |  | 11.8 | 128.22 | 126.6
-8.4 |  | 4.4 | 144.5333 | 142.5 |  | 4.2 | 144.2 | 142.1 |  | 5.5667 | 139.7 | 137.866 |  | 8.7978 | 133 | 130.82 |  | 12.15 | 128.025 | 126.34
-8.3 |  | 4.5 | 144.1 | 142 |  | 4.3 | 143.75 | 141.625 |  | 5.7 | 139.24 | 137.416 |  | 8.9977 | 132.75 | 130.525 |  | 12.45 | 127.8667 | 126.0571
-8.2 |  | 4.6 | 143.5 | 141.2667 |  | 4.4 | 143.2 | 141.05 |  | 5.86 | 138.8 | 137.006 |  | 9.2227 | 132.22 | 130.24 |  | 12.8 | 127.6 | 125.76
-8.1 |  | 4.7 | 142.875 | 140.8 |  | 4.5667 | 142.7 | 140.54 |  | 6.025 | 138.325 | 136.516 |  | 9.5976 | 131.925 | 129.95 |  | 13.4 | 127.34 | 125.42
-8 |  | 4.8 | 142.2 | 140.25 |  | 4.7 | 142.2 | 140.18 |  | 6.2333 | 137.82 | 136.126 |  | 9.8475 | 131.6 | 129.65 |  | 13.7667 | 126.96 | 125.1667
-7.9 |  | 4.9 | 141.8 | 139.7 |  | 4.8 | 141.6 | 139.5 |  | 6.4 | 137.38 | 135.666 |  | 10.1974 | 131.2667 | 129.2333 |  | 14.2667 | 126.65 | 124.95
-7.8 |  | 5 | 141.3333 | 139.2 |  | 4.9667 | 140.9 | 138.9 |  | 6.6 | 136.92 | 135.166 |  | 10.5641 | 130.9333 | 129 |  | 14.9 | 126.325 | 124.75
-7.7 |  | 5.2 | 140.64 | 138.64 |  | 5.1 | 140.4 | 138.45 |  | 6.825 | 136.4714 | 134.666 |  | 10.8973 | 130.45 | 128.5667 |  | 15.6 | 126.1 | 124.5
-7.6 |  | 5.3 | 140.1 | 138.08 |  | 5.2 | 139.8 | 138.0667 |  | 7.1 | 135.8667 | 134.266 |  | 11.3639 | 129.98 | 128.18 | 2.1 | 16.375 | 125.82 | 124.2
-7.5 |  | 5.475 | 139.5 | 137.55 |  | 5.4333 | 139.15 | 137.4 |  | 7.375 | 135.3 | 133.717 |  | 11.6971 | 129.58 | 127.8 | 2.4 | 17.2 | 125.525 | 123.92
-7.4 |  | 5.6 | 138.9 | 136.9667 |  | 5.7 | 138.8 | 136.9333 |  | 7.5333 | 134.8 | 133.367 |  | 12.147 | 129.18 | 127.64 | 2.5 | 18.4 | 125.1333 | 123.7
-7.3 |  | 5.7 | 138.4667 | 136.44 |  | 5.9 | 138.0667 | 136.25 |  | 7.9 | 134.4 | 132.987 | 1.9 | 12.8218 | 128.8 | 127.2 | 2.6 | 19.6 | 124.84 | 123.3
-7.2 |  | 5.95 | 137.775 | 135.88 |  | 6.025 | 137.6667 | 135.8 |  | 8.2333 | 133.82 | 132.442 | 1.9 | 13.5216 | 128.44 | 126.8 | 2.7 | 21.2333 | 124.54 | 123
-7.1 |  | 6.15 | 137.25 | 135.38 |  | 6.275 | 137.0667 | 135.3 | 1.3 | 8.6 | 133.35 | 132.117 | 1.9 | 14.3964 | 128.22 | 126.5 | 2.75 | 23.2 | 124.2143 | 122.55
-7 | 1.2 | 6.4 | 136.7 | 134.875 | 1.1 | 6.6 | 136.35 | 134.65 | 1.4 | 9 | 133 | 131.627 | 2 | 15.2712 | 127.7667 | 126.1667 | 2.8 | 25.8667 | 123.9 | 122.4
-6.9 | 1.2 | 6.6 | 135.9667 | 134.44 | 1.1 | 6.9 | 135.8667 | 134.3 | 1.4 | 9.5 | 132.45 | 131.127 | 2 | 16.2959 | 127.35 | 125.84 | 2.9 | 29.3 | 123.42 | 122.12
-6.8 | 1.2 | 6.9 | 135.4 | 133.64 | 1.1 | 7.3 | 135.4 | 133.72 | 1.4 | 10.1 | 131.92 | 130.767 | 2.1 | 17.6956 | 127.0167 | 125.4 | 2.9 | 34.4 | 123.2333 | 121.8143
-6.7 | 1.2 | 7.075 | 134.7667 | 133.05 | 1.1 | 7.6 | 134.9 | 133.16 | 1.4 | 10.6667 | 131.4667 | 130.167 | 2.1 | 19.2452 | 126.52 | 124.96 | 2.9 | 41.4 | 122.85 | 121.5
-6.6 | 1.2 | 7.475 | 134.2 | 132.55 | 1.1 | 8.0667 | 134.1667 | 132.65 | 1.4 | 11.425 | 131.04 | 129.718 | 2.2 | 21.2281 | 126.18 | 124.6 | 3 | 49.9 | 122.44 | 121.3
-6.5 | 1.2 | 7.875 | 133.5667 | 132.0333 | 1.2 | 8.7 | 133.6 | 132.08 | 1.5 | 12.32 | 130.6 | 129.148 | 2.2 | 24.0607 | 125.7 | 124.3 | 3.1 | 59.2333 | 122.075 | 121
-6.4 | 1.3 | 8.3667 | 132.975 | 131.6 | 1.2 | 9.2667 | 133 | 131.5333 | 1.5 | 13.4 | 129.85 | 128.588 | 2.3 | 27.893 | 125.35 | 123.9 | 3.1 | 70 | 121.8 | 120.72
-6.3 | 1.3 | 8.9333 | 132.4 | 130.94 | 1.2 | 10.1 | 132.5 | 131.04 | 1.5 | 14.8 | 129.425 | 128.168 | 2.3 | 33.1417 | 124.9667 | 123.62 | 3.2 | 79.9 | 121.4667 | 120.275
-6.2 | 1.3 | 9.6 | 131.85 | 130.38 | 1.2 | 11 | 131.7333 | 130.36 | 1.6 | 16.4 | 129.1 | 127.7251 | 2.3 | 40.9898 | 124.6 | 123.26 | 3.3 | 88.5333 | 121.2 | 120.025
-6.1 | 1.3 | 10.425 | 131.4 | 129.95 | 1.3 | 12.1 | 131.325 | 129.84 | 1.65 | 18.65 | 128.5 | 127.3347 | 2.4 | 49.6543 | 124.3333 | 122.76 | 3.4 | 95.425 | 120.75 | 119.6
-6 | 1.3 | 11.4 | 130.7 | 129.3 | 1.3 | 13.6 | 130.6667 | 129.2 | 1.7 | 21.4 | 128 | 126.708 | 2.4 | 60.5848 | 123.8 | 122.4 | 3.4 | 99.925 | 120.4667 | 119.4
-5.9 | 1.3 | 12.8667 | 130.1 | 128.8 | 1.3 | 15.6 | 130.15 | 128.7 | 1.7 | 25.3667 | 127.42 | 126.1017 | 2.5 | 71.5321 | 123.4 | 121.9 | 3.5 | 103.5 | 120.2167 | 119.0667
-5.8 | 1.3 | 14.6 | 129.5667 | 128.3 | 1.3 | 18.2 | 129.4667 | 128.0667 | 1.8 | 31.9333 | 126.9 | 125.719 | 2.6 | 82.0794 | 123.04 | 121.6 | 3.6 | 105.875 | 119.85 | 118.7714
-5.7 | 1.4 | 17.2333 | 129.1333 | 127.575 | 1.4 | 21.7 | 129.0333 | 127.625 | 1.8 | 40.1 | 126.38 | 125.219 | 2.6 | 90.844 | 122.6 | 121.36 | 3.7 | 107.7 | 119.4833 | 118.5
-5.6 | 1.4 | 20.5 | 128.5333 | 127.06 | 1.4 | 26.75 | 128.3333 | 127.08 | 1.9 | 50.925 | 125.85 | 124.8023 | 2.7 | 97.6756 | 122.15 | 120.95 | 3.8 | 108.75 | 119.15 | 118.2
-5.5 | 1.5 | 25.7 | 127.8333 | 126.56 | 1.4 | 34.4 | 127.7333 | 126.4 | 1.9 | 62.5 | 125.4 | 124.249 | 2.8 | 102.374 | 121.72 | 120.45 | 3.9 | 109.6 | 118.84 | 117.85
-5.4 | 1.5 | 34.7 | 127.2667 | 126.1 | 1.5 | 45.1333 | 127.2 | 126 | 2 | 73.8 | 124.8 | 123.819 | 2.9 | 105.5934 | 121.2 | 120.1 | 3.95 | 110.3333 | 118.5333 | 117.5833
-5.3 | 1.5 | 45.0667 | 126.65 | 125.6 | 1.5 | 57.35 | 126.8 | 125.4 | 2 | 85 | 124.28 | 123.419 | 3 | 107.873 | 120.8667 | 119.86 | 4 | 111.1 | 118.2 | 117.1667
-5.2 | 1.6 | 57.3 | 126.1 | 124.9 | 1.6 | 69.5 | 126.2 | 124.95 | 2.1 | 94.2333 | 123.8 | 122.829 | 3.1 | 109.423 | 120.6 | 119.36 | 4.15 | 111.4 | 117.8 | 116.82
-5.1 | 1.6 | 70.9 | 125.7 | 124.3333 | 1.6 | 82.1 | 125.7333 | 124.42 | 2.2 | 101.8 | 123.3 | 122.469 | 3.2 | 110.672 | 120.1333 | 119 | 4.3 | 111.725 | 117.4143 | 116.575
-5 | 1.6 | 82.725 | 124.9667 | 123.8 | 1.6 | 92.2 | 125 | 123.8 | 2.2 | 105.6 | 122.82 | 121.97 | 3.3 | 111.497 | 119.8 | 118.6 | 4.4 | 111.9333 | 117.1 | 116.3
-4.9 | 1.7 | 94.4 | 124.6 | 123.22 | 1.7 | 100.8 | 124.55 | 123.25 | 2.3 | 108.8 | 122.2667 | 121.6033 | 3.4 | 112.172 | 119.4 | 118.35 | 4.5 | 112.1 | 116.7333 | 116
-4.8 | 1.7 | 102.65 | 123.9 | 122.78 | 1.7 | 106.2 | 123.9333 | 122.82 | 2.3 | 110.9 | 121.775 | 121.1414 | 3.5 | 112.472 | 118.9 | 117.9 | 4.7 | 112.1 | 116.35 | 115.6
-4.7 | 1.8 | 108.2333 | 123.3667 | 122.25 | 1.8 | 109.9 | 123.3667 | 122.32 | 2.4 | 112.4 | 121.2333 | 120.61 | 3.7 | 112.772 | 118.56 | 117.52 | 4.8 | 112.2 | 116.0833 | 115.3
-4.6 | 1.8 | 111.5667 | 122.8 | 121.75 | 1.9 | 112.3 | 122.8 | 121.825 | 2.5 | 113.25 | 120.84 | 120.17 | 3.8 | 112.972 | 118.3 | 117.1 | 5.1 | 112.2 | 115.8333 | 115.0429
-4.5 | 1.9 | 113.75 | 122.2667 | 121.3 | 2 | 113.5667 | 122.4 | 121.15 | 2.6 | 113.8667 | 120.45 | 119.67 | 4 | 113.072 | 117.775 | 116.7 | 5.3 | 112.2 | 115.6 | 114.8333
-4.4 | 2 | 115 | 121.7667 | 120.72 | 2.05 | 114.8 | 121.9333 | 120.65 | 2.75 | 114.3333 | 119.875 | 119.17 | 4.2 | 113.172 | 117.4 | 116.4 | 5.6 | 112.175 | 115.3333 | 114.5286
-4.3 | 2.1 | 115.85 | 121.35 | 120.16 | 2.1 | 115.4 | 121.2 | 120.34 | 2.9 | 114.575 | 119.35 | 118.77 | 4.3 | 113.222 | 117.04 | 116.12 | 5.8 | 112.1 | 114.98 | 114.2
-4.2 | 2.1 | 116.5 | 120.7333 | 119.75 | 2.2 | 115.85 | 120.6 | 119.75 | 3.1 | 114.7 | 118.95 | 118.37 | 4.6 | 113.272 | 116.4333 | 115.82 | 6.1 | 112 | 114.7 | 113.8833
-4.1 | 2.2 | 116.8 | 120.0333 | 119.3 | 2.3 | 116.1333 | 120.1667 | 119.1667 | 3.2 | 114.8 | 118.52 | 117.871 | 4.8 | 113.172 | 116.1667 | 115.36 | 6.4 | 111.85 | 114.2 | 113.55
-4 | 2.3 | 116.9 | 119.5 | 118.775 | 2.5 | 116.3 | 119.7 | 118.7 | 3.4 | 114.8 | 118.075 | 117.331 | 5.1 | 113.0387 | 115.775 | 115.02 | 6.9 | 111.7 | 113.92 | 113.3
-3.9 | 2.4 | 116.9 | 118.95 | 118.1 | 2.6 | 116.3 | 119.25 | 118.3667 | 3.6 | 114.7 | 117.58 | 116.971 | 5.3 | 112.872 | 115.54 | 114.7333 | 7.25 | 111.55 | 113.58 | 112.8
-3.8 | 2.5 | 116.8333 | 118.475 | 117.575 | 2.8 | 116.3 | 118.65 | 117.675 | 3.85 | 114.6667 | 117.06 | 116.5139 | 5.65 | 112.672 | 115.2 | 114.36 | 7.65 | 111.4 | 113.36 | 112.5333
-3.7 | 2.7 | 116.675 | 117.9333 | 117.26 | 3 | 116.2 | 118.15 | 117.325 | 4.2 | 114.5 | 116.54 | 115.971 | 6.1 | 112.472 | 114.75 | 113.9333 | 8.2 | 111.3 | 112.86 | 112.2333
-3.6 | 2.8667 | 116.5 | 117.5333 | 116.775 | 3.3 | 115.9667 | 117.5333 | 116.85 | 4.5 | 114.4 | 116.0667 | 115.621 | 6.6 | 112.272 | 114.3 | 113.55 | 9.1 | 111.1 | 112.5833 | 111.875
-3.5 | 3.1 | 116.1333 | 117.1 | 116.3333 | 3.5 | 115.7 | 117.1333 | 116.32 | 4.9 | 114.05 | 115.75 | 115.271 | 7.2 | 112.072 | 113.9 | 113.2 | 10.1 | 110.7 | 112.3333 | 111.625
-3.4 | 3.4 | 115.9 | 116.5 | 115.8 | 3.9 | 115.5 | 116.7 | 115.9 | 5.35 | 113.8 | 115.1 | 114.731 | 7.9 | 111.872 | 113.54 | 112.7333 | 11.2 | 110.4 | 111.95 | 111.3
-3.3 | 3.7 | 115.35 | 116.05 | 115.4 | 4.25 | 115.25 | 116.2 | 115.35 | 5.8 | 113.525 | 114.68 | 114.411 | 8.7 | 111.647 | 113.1 | 112.4 | 12.3 | 110.0333 | 111.54 | 111
-3.2 | 4.1 | 115.075 | 115.525 | 114.8 | 4.75 | 114.85 | 115.7 | 114.8 | 6.6 | 113.2667 | 114.3333 | 113.932 | 9.9 | 111.397 | 112.6 | 112 | 14.2 | 109.8667 | 111.05 | 110.68
-3.1 | 4.6 | 114.875 | 114.9 | 114.26 | 5.4 | 114.4333 | 115.2 | 114.425 | 7.5 | 113 | 113.85 | 113.492 | 11.3 | 111.0387 | 112.15 | 111.55 | 16.15 | 109.6333 | 110.725 | 110.3
-3 | 5.3 | 114.5333 | 114.55 | 113.82 | 6.3 | 113.98 | 114.5667 | 113.84 | 8.6 | 112.625 | 113.36 | 112.892 | 13 | 110.697 | 111.76 | 111.18 | 19.35 | 109.1 | 110.4 | 109.82
-2.9 | 6 | 114 | 114.05 | 113.425 | 7.2 | 113.68 | 114.125 | 113.32 | 9.8 | 112.2 | 112.9 | 112.492 | 15.3 | 110.297 | 111.4286 | 110.675 | 23.8 | 108.9333 | 110.08 | 109.54
-2.8 | 7.3 | 113.3 | 113.5 | 113 | 8.6 | 113.1667 | 113.4667 | 113 | 11.5 | 111.9333 | 112.3 | 112.022 | 18.5 | 109.973 | 110.88 | 110.28 | 30.1 | 108.7667 | 109.56 | 109.2
-2.7 | 8.7 | 112.8 | 112.8333 | 112.425 | 10.5 | 112.7 | 113.1333 | 112.4 | 14.1 | 111.35 | 111.7 | 111.397 | 23.3 | 109.5397 | 110.475 | 109.875 | 38 | 108.4667 | 109.1333 | 108.95
-2.6 | 10.9 | 112.5 | 112.3 | 111.86 | 12.7 | 112.3 | 112.5 | 111.8 | 17.6 | 111 | 111.2333 | 110.997 | 30 | 109.273 | 110.05 | 109.625 | 48 | 108.2 | 108.7 | 108.45
-2.5 | 13.6 | 112.1 | 111.9 | 111.35 | 16.4 | 111.7 | 111.9 | 111.2 | 22.1 | 110.7 | 110.8 | 110.532 | 37.8 | 108.8063 | 109.54 | 109.2 | 57.9 | 107.875 | 108.3667 | 108.0333
-2.4 | 17.85 | 111.4 | 111.2 | 110.74 | 21.55 | 111.3667 | 111.3333 | 110.82 | 29.7 | 110.3 | 110.34 | 109.9924 | 48.8 | 108.573 | 109.05 | 108.7 | 67.8667 | 107.6 | 108.02 | 107.7
-2.3 | 24.6 | 110.9 | 110.7 | 110.2 | 28.9 | 110.7 | 110.8667 | 110.3 | 38.7 | 109.7 | 109.82 | 109.523 | 59 | 108.223 | 108.5667 | 108.2 | 76.8 | 107.3 | 107.625 | 107.4
-2.2 | 34.7 | 110.4 | 110.05 | 109.7 | 39 | 110.1333 | 110.3 | 109.7333 | 49.6 | 109.25 | 109.44 | 109.053 | 69.65 | 107.823 | 108.175 | 107.8667 | 84.85 | 106.95 | 107.2857 | 107
-2.1 | 46.2 | 109.7 | 109.65 | 109.2 | 49.7 | 109.6 | 109.775 | 109.1 | 61.35 | 108.7 | 108.88 | 108.5444 | 79.2667 | 107.448 | 107.85 | 107.4 | 90.4 | 106.6 | 106.96 | 106.6667
-2 | 57.65 | 109.3 | 109.0667 | 108.66 | 62.35 | 109.0667 | 109.0667 | 108.66 | 71.8 | 108.2333 | 108.4 | 108.073 | 86.7 | 107.073 | 107.44 | 107.14 | 93.6 | 106.25 | 106.55 | 106.4
-1.9 | 69.2 | 108.8 | 108.5 | 108.2333 | 73.2 | 108.5667 | 108.5 | 108.2333 | 81.4 | 107.9 | 107.86 | 107.653 | 91.8 | 106.653 | 107.05 | 106.7 | 96.35 | 105.9 | 106.2 | 106.0333
-1.8 | 79.8 | 108.25 | 108.0667 | 107.7333 | 81.9 | 108.1 | 108.2 | 107.62 | 88.3 | 107.4 | 107.45 | 107.123 | 95.3 | 106.3397 | 106.8 | 106.3 | 98.2 | 105.7333 | 105.9333 | 105.6
-1.7 | 88.0333 | 107.8 | 107.45 | 107.18 | 90 | 107.6667 | 107.65 | 107.15 | 93.1 | 107.05 | 106.85 | 106.7397 | 97.5 | 105.874 | 106.3 | 105.9333 | 99.65 | 105.4 | 105.65 | 105.4
-1.6 | 93.9 | 107.225 | 107.0333 | 106.65 | 94.45 | 107.2 | 107.0667 | 106.7333 | 96.55 | 106.65 | 106.4 | 106.3397 | 98.9 | 105.624 | 105.8286 | 105.6 | 100.6 | 105.2 | 105.24 | 105.0714
-1.5 | 97.4 | 106.8 | 106.5333 | 106.2 | 97.2 | 106.55 | 106.625 | 106.28 | 98.7 | 105.9 | 105.9333 | 105.9238 | 100.2 | 105.274 | 105.4333 | 105.25 | 101.35 | 104.8 | 104.92 | 104.66
-1.4 | 99.4 | 106.25 | 106.05 | 105.75 | 99.3 | 106.3 | 106.2 | 105.8 | 99.6 | 105.5 | 105.64 | 105.434 | 100.9 | 104.924 | 105.0143 | 104.8333 | 101.8 | 104.4333 | 104.54 | 104.2667
-1.3 | 100.7 | 105.8333 | 105.65 | 105.3 | 100.3 | 105.9 | 105.675 | 105.3333 | 100.45 | 105.1333 | 105.4 | 105.094 | 101.1 | 104.574 | 104.7167 | 104.5 | 102 | 104.15 | 104.2 | 103.9333
-1.2 | 101.55 | 105.4667 | 105.2 | 104.9 | 101.3 | 105.3 | 105.2 | 104.975 | 101.3 | 104.75 | 104.78 | 104.724 | 101.7 | 104.274 | 104.4 | 104.06 | 102.2 | 103.9 | 103.7667 | 103.6
-1.1 | 102 | 105.0333 | 104.7667 | 104.55 | 101.7 | 104.9 | 104.9 | 104.6 | 101.7 | 104.375 | 104.375 | 104.249 | 101.9 | 103.974 | 104.05 | 103.7 | 102.4 | 103.625 | 103.4333 | 103.3667
-1 | 102.1 | 104.4 | 104.3333 | 104.1 | 101.9 | 104.4 | 104.3333 | 104.2 | 101.8 | 104.05 | 103.95 | 103.849 | 102.1 | 103.524 | 103.7 | 103.4 | 102.4 | 103.25 | 103.125 | 103.0667
-0.9 | 102.2 | 103.7 | 103.8 | 103.55 | 102 | 103.9 | 103.9 | 103.6667 | 101.9 | 103.7 | 103.6 | 103.474 | 102.1 | 103.224 | 103.3571 | 103.0667 | 102.2 | 102.75 | 102.8 | 102.72
-0.8 | 102.2 | 103.5 | 103.3 | 103.1 | 102 | 103.4333 | 103.5667 | 103.225 | 101.9 | 103.1667 | 103.1 | 103.074 | 102.1 | 102.824 | 102.86 | 102.7 | 102.1 | 102.45 | 102.48 | 102.5
-0.7 | 102 | 102.9 | 102.9 | 102.75 | 102 | 103.025 | 103.0333 | 102.66 | 101.7 | 102.675 | 102.7 | 102.614 | 101.8 | 102.4073 | 102.54 | 102.4667 | 101.8 | 102.2 | 102.1429 | 102.2
-0.6 | 101.8 | 102.45 | 102.4333 | 102.325 | 101.8 | 102.65 | 102.5333 | 102.4 | 101.5 | 102.2 | 102.175 | 102.299 | 101.6 | 101.975 | 102.12 | 102.0429 | 101.6 | 101.9 | 101.725 | 101.88
-0.5 | 101.6 | 102 | 102.15 | 101.96 | 101.7 | 102.2667 | 102.1667 | 101.95 | 101.3 | 101.8 | 101.85 | 101.9148 | 101.4 | 101.675 | 101.74 | 101.65 | 101.4 | 101.7 | 101.55 | 101.5
-0.4 | 101.4 | 101.7 | 101.6333 | 101.55 | 101.4 | 101.75 | 101.7 | 101.5 | 101.2 | 101.6 | 101.4333 | 101.475 | 101.3 | 101.425 | 101.3 | 101.3 | 101.05 | 101.2667 | 101.25 | 101.18
-0.3 | 101 | 101.1333 | 101.2 | 101.02 | 101.1 | 101.3 | 101.2333 | 101.08 | 100.95 | 101.0667 | 101.05 | 101.0417 | 101 | 100.95 | 100.92 | 100.8833 | 100.7 | 100.875 | 100.8714 | 100.9
-0.2 | 100.6 | 100.675 | 100.9 | 100.7333 | 100.7 | 100.9 | 100.85 | 100.75 | 100.6 | 100.5667 | 100.6 | 100.695 | 100.7 | 100.675 | 100.6857 | 100.6 | 100.55 | 100.6 | 100.5667 | 100.5333
-0.1 | 100.3 | 100.35 | 100.42 | 100.3 | 100.4 | 100.3333 | 100.5 | 100.45 | 100.3 | 100.3 | 100.275 | 100.225 | 100.3 | 100.275 | 100.36 | 100.2333 | 100.3 | 100.2333 | 100.3333 | 100.35
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.5 | 99.625 | 99.5 | 99.58 | 99.6 | 99.65 | 99.7333 | 99.62 | 99.7 | 99.6667 | 99.56 | 99.6751 | 99.7 | 99.7001 | 99.8 | 99.75 | 99.7 | 99.7 | 99.5333 | 99.68
0.2 | 99 | 99.225 | 99.175 | 99.3 | 99.2 | 99.4 | 99.325 | 99.2857 | 99.3 | 99.3 | 99.225 | 99.3609 | 99.5 | 99.4751 | 99.46 | 99.36 | 99.5 | 99.5 | 99.2667 | 99.46
0.3 | 98.45 | 98.8 | 98.875 | 99 | 98.8 | 99.0667 | 99 | 98.98 | 98.9 | 98.9 | 98.94 | 99.0086 | 99.25 | 99.0752 | 99.05 | 99.1 | 99.25 | 99.3 | 98.95 | 99.2
0.4 | 98.2 | 98.5333 | 98.6 | 98.7 | 98.3 | 98.7 | 98.7 | 98.76 | 98.5 | 98.6 | 98.6 | 98.5754 | 98.75 | 98.7753 | 98.85 | 98.9 | 98.8 | 98.9667 | 98.8 | 99
0.5 | 97.85 | 98.25 | 98.2 | 98.4333 | 97.9 | 98.3667 | 98.38 | 98.45 | 98.1 | 98.4 | 98.3667 | 98.2421 | 98.4 | 98.5004 | 98.56 | 98.575 | 98.6 | 98.7 | 98.6 | 98.7667
0.6 | 97.15 | 97.9 | 97.9 | 98.1333 | 97.4 | 97.975 | 98.05 | 98.1667 | 97.8 | 98.1 | 98.1167 | 98.0422 | 98 | 98.3087 | 98.275 | 98.275 | 98.3667 | 98.4 | 98.36 | 98.5
0.7 | 96.65 | 97.6 | 97.6 | 97.925 | 96.7 | 97.6 | 97.8 | 97.88 | 97.1 | 97.8 | 97.8 | 97.8755 | 97.6333 | 97.9005 | 98.06 | 98.14 | 97.9 | 98.2667 | 98.05 | 98.3
0.8 | 96 | 97.2 | 97.4 | 97.65 | 96.4 | 97.425 | 97.5 | 97.5667 | 96.5 | 97.3667 | 97.46 | 97.5756 | 97.1 | 97.6756 | 97.82 | 97.8333 | 97.35 | 98 | 97.82 | 98
0.9 | 95.5 | 96.9333 | 97.1 | 97.2 | 95.7 | 97.1 | 97.3 | 97.275 | 96 | 97.2667 | 97.2 | 97.309 | 96.6 | 97.3757 | 97.5 | 97.55 | 96.9 | 97.72 | 97.6 | 97.875
1 | 94.8 | 96.65 | 96.8667 | 96.9 | 95.1 | 96.825 | 97 | 97.0333 | 95.3 | 96.9 | 96.9 | 97.1257 | 96 | 97.0757 | 97.3 | 97.325 | 96.6 | 97.425 | 97.26 | 97.625
1.1 | 93.9 | 96.35 | 96.6 | 96.7667 | 94.1 | 96.425 | 96.7 | 96.7 | 94.4 | 96.6 | 96.66 | 96.8258 | 95.6 | 96.8258 | 97.1 | 97.12 | 96.2 | 97.175 | 97.075 | 97.42
1.2 | 92.8 | 96.05 | 96.2 | 96.5 | 92.9 | 96.1 | 96.475 | 96.46 | 93.6 | 96.3667 | 96.4 | 96.5559 | 94.6 | 96.5426 | 96.875 | 96.95 | 95.6 | 96.9 | 96.925 | 97.2167
1.3 | 91.6 | 95.725 | 95.925 | 96.2333 | 91.7 | 95.85 | 96.2333 | 96.26 | 92.4 | 96.1 | 96.14 | 96.3093 | 94.2333 | 96.176 | 96.5 | 96.74 | 95.05 | 96.6333 | 96.7 | 96.92
1.4 | 90 | 95.4333 | 95.7 | 95.9 | 90.2 | 95.6 | 95.875 | 95.95 | 91.05 | 95.7 | 95.9 | 96.026 | 93.1 | 95.976 | 96.2 | 96.52 | 94.45 | 96.4 | 96.56 | 96.8
1.5 | 87.85 | 95.0667 | 95.5 | 95.7167 | 88.3 | 95.2667 | 95.6 | 95.675 | 89.8 | 95.3333 | 95.6667 | 95.8161 | 91.9 | 95.6761 | 95.92 | 96.25 | 93.5 | 96.2 | 96.28 | 96.56
1.6 | 84.5 | 94.8 | 95.2 | 95.48 | 85 | 95 | 95.3667 | 95.45 | 87.4 | 94.95 | 95.3333 | 95.6011 | 90.8 | 95.5761 | 95.7 | 95.9333 | 92.7 | 96 | 95.9667 | 96.3
1.7 | 79.5 | 94.55 | 94.95 | 95.15 | 81.4 | 94.85 | 95.1 | 95.1667 | 84.9 | 94.8 | 95.08 | 95.2762 | 89.2 | 95.3262 | 95.54 | 95.675 | 91.7 | 95.6 | 95.875 | 96.1
1.8 | 72.7 | 94.25 | 94.6 | 94.92 | 75.1 | 94.6 | 94.9 | 94.96 | 80.7 | 94.5333 | 94.9 | 94.9596 | 87.15 | 94.9763 | 95.34 | 95.525 | 90.2 | 95.4333 | 95.7 | 95.9333
1.9 | 63.6 | 94 | 94.4 | 94.6667 | 67.4 | 94.2 | 94.65 | 94.68 | 75 | 94.4 | 94.5333 | 94.7763 | 84.1 | 94.6763 | 95.08 | 95.175 | 88.8 | 95.0667 | 95.4 | 95.65
2 | 53.5 | 93.7 | 94.1667 | 94.3857 | 57.2 | 93.8 | 94.3667 | 94.4 | 67.3 | 94 | 94.275 | 94.5364 | 80.3 | 94.4014 | 94.84 | 94.9667 | 86.8 | 94.8667 | 95.18 | 95.4333
2.1 | 42.7 | 93.5 | 93.9 | 94.1 | 47.85 | 93.5333 | 94 | 94.18 | 58.8 | 93.7 | 94.1 | 94.2764 | 74.8 | 94.0515 | 94.625 | 94.65 | 83.5 | 94.7 | 94.92 | 95.1
2.2 | 32.75 | 93.2 | 93.7 | 93.88 | 37.7 | 93.2667 | 93.8667 | 93.95 | 47.7 | 93.4 | 93.925 | 94.1565 | 67.9 | 93.7266 | 94.44 | 94.52 | 80.5 | 94.3 | 94.55 | 94.96
2.3 | 24.1 | 92.9667 | 93.5 | 93.6 | 28 | 92.925 | 93.65 | 93.74 | 38.8 | 93.0333 | 93.7 | 93.8765 | 59.2 | 93.3567 | 94.2 | 94.3 | 74.9 | 94 | 94.4 | 94.74
2.4 | 17.3 | 92.6333 | 93.225 | 93.45 | 20.8 | 92.625 | 93.4667 | 93.5 | 29.8 | 92.7667 | 93.2833 | 93.5433 | 49.9 | 93.1767 | 93.9167 | 94.1 | 67.7 | 93.7 | 94.14 | 94.5
2.5 | 13 | 92.25 | 93.0667 | 93.2 | 15.9 | 92.4333 | 93.2 | 93.2 | 22.4 | 92.4 | 93.1 | 93.4017 | 40 | 92.8518 | 93.65 | 93.75 | 59.5 | 93.5 | 93.9 | 94.3333
2.6 | 10 | 91.9 | 92.825 | 92.9 | 12.4 | 92.15 | 92.9 | 92.98 | 17.1 | 92.2 | 92.9 | 93.1517 | 31.5 | 92.5769 | 93.35 | 93.55 | 50.65 | 93.26 | 93.7833 | 93.98
2.7 | 8.1 | 91.65 | 92.425 | 92.6667 | 10 | 91.8667 | 92.6 | 92.72 | 13.8 | 91.9 | 92.6667 | 92.8601 | 24.6 | 92.2769 | 93.1333 | 93.35 | 42.1 | 92.9333 | 93.6167 | 93.9
2.8 | 6.6 | 91.375 | 92.3 | 92.5 | 8.2 | 91.45 | 92.4 | 92.48 | 11.3 | 91.675 | 92.45 | 92.6768 | 19.3667 | 91.927 | 93 | 93.12 | 33.5 | 92.7 | 93.34 | 93.7333
2.9 | 5.6 | 91 | 92.0333 | 92.22 | 7 | 91 | 92.2 | 92.35 | 9.4333 | 91.34 | 92.2 | 92.3769 | 15.7 | 91.6771 | 92.8714 | 92.925 | 26.4 | 92.45 | 93.1 | 93.5
3 | 4.7 | 90.65 | 91.7 | 91.9333 | 5.95 | 90.65 | 91.85 | 92.0333 | 8.05 | 90.95 | 91.8143 | 92.052 | 13.3 | 91.3772 | 92.6333 | 92.8 | 21.3 | 92.2 | 92.8667 | 93.28
3.1 | 4.15 | 90.2667 | 91.3333 | 91.7 | 5.1 | 90.4333 | 91.5667 | 91.64 | 6.95 | 90.5 | 91.5833 | 91.9103 | 11.6 | 91.0522 | 92.4 | 92.5 | 17.3 | 91.9 | 92.6 | 93.025
3.2 | 3.65 | 89.9333 | 91.1333 | 91.4667 | 4.5 | 90.1 | 91.4 | 91.4 | 6.2 | 90.1 | 91.2857 | 91.6771 | 9.9 | 90.6523 | 92.1 | 92.2 | 14.75 | 91.5 | 92.3667 | 92.8
3.3 | 3.3333 | 89.55 | 90.9 | 91.1 | 3.95 | 89.6667 | 91.1 | 91.2 | 5.55 | 89.7 | 91.025 | 91.4172 | 8.7 | 90.4107 | 91.8 | 92.0333 | 12.9 | 91.15 | 92.1714 | 92.6
3.4 | 3 | 89.1 | 90.6 | 90.8333 | 3.55 | 89.175 | 90.85 | 90.85 | 5 | 89.4 | 90.86 | 91.0772 | 7.9 | 90.0775 | 91.675 | 91.7 | 11.3 | 90.9 | 92 | 92.4
3.5 | 2.7 | 88.7 | 90.2 | 90.5 | 3.2 | 88.6667 | 90.55 | 90.62 | 4.5 | 89 | 90.575 | 90.8373 | 7 | 89.7776 | 91.35 | 91.5333 | 9.8 | 90.6333 | 91.8 | 92.1667
3.6 | 2.6 | 88 | 89.9333 | 90.28 | 2.9 | 88.2667 | 90.2333 | 90.4 | 4.1 | 88.5667 | 90.325 | 90.5274 | 6.4 | 89.4027 | 90.975 | 91.34 | 8.8 | 90.3 | 91.58 | 91.95
3.7 | 2.4 | 87.5 | 89.65 | 89.94 | 2.8 | 87.6 | 89.95 | 90.0667 | 3.8 | 88 | 90.1 | 90.3274 | 6 | 88.9111 | 90.725 | 91.06 | 8.3333 | 89.9 | 91.3 | 91.6333
3.8 | 2.3 | 87.1667 | 89.275 | 89.68 | 2.5 | 87.2 | 89.6333 | 89.75 | 3.5 | 87.58 | 89.75 | 89.9775 | 5.6 | 88.5112 | 90.55 | 90.7429 | 7.6 | 89.6333 | 91.05 | 91.4
3.9 | 2.1 | 86.575 | 89 | 89.38 | 2.4 | 86.5 | 89.2333 | 89.4333 | 3.3 | 86.9 | 89.4 | 89.6443 | 5.2333 | 88.103 | 90.28 | 90.48 | 7.05 | 89.3 | 90.8 | 91.2
4 | 2 | 85.9 | 88.8 | 88.98 | 2.2 | 85.9 | 89 | 89.175 | 3.1 | 86.35 | 89.1 | 89.311 | 4.85 | 87.5781 | 90.0667 | 90.3 | 6.6 | 88.95 | 90.675 | 90.9667
4.1 | 1.9 | 85.2 | 88.4 | 88.7 | 2.1 | 85.3 | 88.6 | 88.78 | 2.95 | 85.8667 | 88.85 | 89.1027 | 4.6 | 87.3115 | 89.68 | 89.86 | 6.2 | 88.475 | 90.3 | 90.7
4.2 | 1.8 | 84.3667 | 88.1 | 88.45 | 2 | 84.2 | 88.175 | 88.52 | 2.85 | 85.1 | 88.54 | 88.8578 | 4.3 | 86.5784 | 89.4 | 89.52 | 5.8 | 88.1667 | 90.15 | 90.475
4.3 | 1.7 | 83.4 | 87.7667 | 88.1 | 1.9 | 83.5667 | 87.9 | 88.1833 | 2.65 | 84.5 | 88.2333 | 88.6778 | 4.1 | 86.1285 | 89.08 | 89.26 | 5.5 | 87.925 | 89.82 | 90.18
4.4 | 1.7 | 82.2 | 87.5333 | 87.86 | 1.8 | 82.4333 | 87.6 | 87.9667 | 2.5 | 83.7 | 87.92 | 88.3179 | 3.9 | 85.6453 | 88.8667 | 89.04 | 5.3 | 87.2333 | 89.5 | 90
4.5 | 1.6 | 80.5667 | 87.225 | 87.55 | 1.7 | 81.2667 | 87.475 | 87.58 | 2.4 | 82.8 | 87.6667 | 87.978 | 3.8 | 85.0788 | 88.6167 | 88.8 | 5.1 | 86.8 | 89.3333 | 89.6571
4.6 | 1.5 | 78.7 | 86.9 | 87.175 | 1.6 | 79.5 | 87.175 | 87.3667 | 2.3 | 82.0667 | 87.4167 | 87.5781 | 3.6 | 84.6039 | 88.28 | 88.45 | 5 | 86.4333 | 89.0333 | 89.42
4.7 | 1.5 | 75.875 | 86.55 | 86.94 | 1.6 | 77.7 | 86.8667 | 87.15 | 2.2 | 80.7 | 87.06 | 87.3532 | 3.4 | 83.829 | 87.95 | 88.2 | 4.8 | 85.95 | 88.78 | 89.14
4.8 | 1.4 | 71.7333 | 86.3 | 86.64 | 1.5 | 74.8 | 86.6 | 86.7667 | 2.1 | 79.075 | 86.74 | 87.0615 | 3.3 | 82.946 | 87.74 | 87.95 | 4.6 | 85.3333 | 88.375 | 88.96
4.9 | 1.4 | 65.8667 | 85.95 | 86.3 | 1.5 | 70.95 | 86.2 | 86.52 | 2 | 77.35 | 86.4 | 86.8116 | 3.2 | 82.4127 | 87.44 | 87.66 | 4.4 | 84.9 | 88.15 | 88.65
5 | 1.4 | 58.55 | 85.6667 | 86.04 | 1.4 | 65.2 | 85.9667 | 86.1 | 1.9 | 75.2667 | 86.1 | 86.4984 | 3.1 | 81.5397 | 87.08 | 87.4333 | 4.3 | 84.325 | 87.9667 | 88.3667
5.1 | 1.3 | 49.9 | 85.3667 | 85.7 | 1.4 | 57.7 | 85.6333 | 85.8 | 1.8 | 71.86 | 85.8667 | 86.2534 | 3 | 80.5799 | 86.8 | 87.1 | 4.2 | 83.6 | 87.7167 | 88.0333
5.2 | 1.3 | 40.5 | 85.0667 | 85.45 | 1.3 | 49.6 | 85.3 | 85.45 | 1.8 | 67.525 | 85.6 | 85.9285 | 2.9 | 79.4801 | 86.55 | 86.7667 | 4 | 82.85 | 87.325 | 87.8
5.3 | 1.3 | 32.3 | 84.7333 | 85.12 | 1.3 | 40.8667 | 85 | 85.2 | 1.7 | 61.2 | 85.2 | 85.6036 | 2.8 | 77.7806 | 86.2333 | 86.525 | 3.9 | 82.225 | 87.1 | 87.5667
5.4 | 1.3 | 25.1667 | 84.4 | 84.72 | 1.2 | 33 | 84.575 | 84.82 | 1.7 | 54.2 | 84.9667 | 85.3187 | 2.7 | 76.181 | 85.95 | 86.26 | 3.8 | 81.62 | 86.85 | 87.28
5.5 | 1.2 | 19.1333 | 84.0667 | 84.4 | 1.2 | 25.4667 | 84.3 | 84.5 | 1.7 | 45.9 | 84.7 | 85.0787 | 2.6 | 73.9815 | 85.72 | 85.9667 | 3.7 | 80.7 | 86.6333 | 86.95
5.6 | 1.2 | 15.3 | 83.7 | 84.05 | 1.2 | 20.05 | 84.1 | 84.2167 | 1.6 | 37.9 | 84.4 | 84.7122 | 2.5 | 71.1156 | 85.4333 | 85.6333 | 3.6 | 79.6 | 86.28 | 86.64
5.7 | 1.2 | 12.8333 | 83.4333 | 83.78 | 1.2 | 15.9 | 83.5667 | 84.02 | 1.55 | 30.35 | 83.95 | 84.3289 | 2.5 | 67.1632 | 85.0667 | 85.3 | 3.5 | 78.3 | 86.1 | 86.3667
5.8 | 1.1 | 11.05 | 83.0333 | 83.46 | 1.1 | 13.6 | 83.3 | 83.6 | 1.5 | 24.5 | 83.65 | 84.059 | 2.4 | 61.3847 | 84.74 | 85.05 | 3.3 | 77.2 | 85.8667 | 86.1
5.9 | 1.1 | 9.725 | 82.7333 | 83.0667 | 1.1 | 11.7 | 83.1 | 83.3167 | 1.4 | 19.6 | 83.35 | 83.799 | 2.4 | 54.553 | 84.62 | 84.76 | 3.3 | 75.3 | 85.525 | 85.85
6 | 1.1 | 8.7 | 82.3 | 82.7571 | 1.1 | 10.25 | 82.65 | 82.9 | 1.4 | 16.6 | 83.1 | 83.5291 | 2.3 | 47.5382 | 84.26 | 84.46 | 3.2 | 73.3 | 85.1 | 85.55
6.1 | 1.1 | 7.9667 | 82 | 82.4 | 1.05 | 9.2 | 82.4 | 82.66 | 1.4 | 14.3667 | 82.775 | 83.1192 | 2.2 | 39.99 | 83.95 | 84.08 | 3.1 | 70.75 | 84.95 | 85.3333
6.2 | 1.1 | 7.3667 | 81.6667 | 82.0667 | 1 | 8.3 | 82 | 82.32 | 1.4 | 12.7667 | 82.4 | 82.7543 | 2.1 | 33.2167 | 83.62 | 83.8333 | 3 | 66.775 | 84.68 | 85
6.3 | 1 | 6.85 | 81.35 | 81.725 | 1 | 7.6667 | 81.6333 | 81.9667 | 1.3 | 11.4 | 82.1 | 82.4294 | 2.1 | 27.1683 | 83.275 | 83.6 | 3 | 61.4 | 84.4 | 84.75
6.4 | 1 | 6.4333 | 81 | 81.4333 | 1 | 7.1 | 81.4 | 81.5 | 1.3 | 10.4 | 81.84 | 82.0595 | 2 | 22.6444 | 83.04 | 83.25 | 2.9 | 55.5 | 84.05 | 84.45
6.5 | 1 | 6.1333 | 80.6 | 81.125 | 1 | 6.6333 | 80.95 | 81.2 | 1.3 | 9.65 | 81.35 | 81.8046 | 2 | 19.6951 | 82.56 | 82.9333 | 2.8 | 48.525 | 83.7 | 84.1333
6.6 | 1 | 5.8667 | 80.3 | 80.6667 | 1 | 6.2333 | 80.6333 | 80.85 | 1.3 | 8.9667 | 81.0667 | 81.3397 | 1.9 | 17.1957 | 82.325 | 82.66 | 2.8 | 41.475 | 83.35 | 83.78
6.7 | 1 | 5.5 | 79.9333 | 80.3 | 1 | 5.8 | 80.3333 | 80.525 | 1.2 | 8.35 | 80.7 | 80.9548 | 1.9 | 15.5961 | 82 | 82.25 | 2.7 | 35.4333 | 83.1167 | 83.64
6.8 | 0.9 | 5.325 | 79.55 | 80.08 | 0.9 | 5.5667 | 79.975 | 80.2 | 1.2 | 7.8667 | 80.4 | 80.6798 | 1.8 | 14.2464 | 81.6714 | 81.9667 | 2.65 | 29.35 | 82.9 | 83.34
6.9 | 0.9 | 5.1 | 79.2 | 79.68 | 0.9 | 5.275 | 79.6 | 79.92 | 1.2 | 7.4667 | 80 | 80.3632 | 1.8 | 13.1468 | 81.34 | 81.7 | 2.6 | 24.9 | 82.7 | 83.05
7 | 0.9 | 4.9333 | 78.825 | 79.28 | 0.9 | 5.075 | 79.25 | 79.56 | 1.2 | 7.125 | 79.6333 | 80.06 | 1.8 | 12.247 | 80.94 | 81.3333 | 2.5 | 21.75 | 82.4 | 82.74
7.1 |  | 4.8 | 78.55 | 78.95 | 0.9 | 4.875 | 78.9 | 79.22 | 1.1 | 6.8 | 79.425 | 79.6801 | 1.8 | 11.5971 | 80.74 | 81 | 2.5 | 19.55 | 81.95 | 82.44
7.2 |  | 4.7 | 78.2 | 78.62 |  | 4.7 | 78.6 | 78.9 | 1.1 | 6.4667 | 79 | 79.4401 | 1.7 | 10.9306 | 80.5 | 80.6833 | 2.5 | 17.7 | 81.5667 | 82.0667
7.3 |  | 4.5 | 77.85 | 78.24 |  | 4.5333 | 78.25 | 78.52 |  | 6.1333 | 78.66 | 79.0802 | 1.7 | 10.2974 | 80.08 | 80.5167 | 2.4 | 16.3 | 81.4 | 81.7333
7.4 |  | 4.4 | 77.5 | 77.925 |  | 4.4 | 77.775 | 78.22 |  | 5.9 | 78.375 | 78.747 | 1.7 | 9.8975 | 79.85 | 80.14 | 2.4 | 15.3 | 81.125 | 81.5
7.5 |  | 4.3 | 77.1333 | 77.5333 |  | 4.2 | 77.5 | 77.75 |  | 5.68 | 78.05 | 78.4304 |  | 9.5476 | 79.525 | 79.72 | 2.4 | 14.3333 | 80.7333 | 81.275
7.6 |  | 4.1333 | 76.75 | 77.2 |  | 4.1 | 77.25 | 77.5 |  | 5.5333 | 77.74 | 78.1055 |  | 9.1727 | 79.2167 | 79.45 | 2.2 | 13.8 | 80.45 | 80.85
7.7 |  | 4.0667 | 76.5 | 76.96 |  | 4 | 76.8667 | 77.22 |  | 5.4 | 77.34 | 77.7472 |  | 8.8478 | 78.84 | 79.2 |  | 13 | 80.1833 | 80.52
7.8 |  | 4 | 76.1667 | 76.55 |  | 3.9 | 76.5 | 76.8 |  | 5.2 | 77.0667 | 77.3607 |  | 8.5978 | 78.52 | 78.8 |  | 12.4333 | 79.9333 | 80.22
7.9 |  | 3.9 | 75.7 | 76.15 |  | 3.8 | 76.12 | 76.45 |  | 5 | 76.6667 | 76.9474 |  | 8.2479 | 78.05 | 78.4 |  | 11.8667 | 79.66 | 79.94
8 |  | 3.8 | 75.45 | 75.78 |  | 3.7 | 75.8667 | 76.05 |  | 4.9 | 76.3 | 76.7058 |  | 8.0647 | 77.7667 | 78.0333 |  | 11.55 | 79.32 | 79.6
8.1 |  | 3.7 | 75.05 | 75.5 |  | 3.6 | 75.4333 | 75.725 |  | 4.775 | 75.9667 | 76.3142 |  | 7.798 | 77.4333 | 77.7 |  | 11.225 | 78.9 | 79.4
8.2 |  | 3.6 | 74.6667 | 75.05 |  | 3.5 | 74.95 | 75.375 |  | 4.625 | 75.62 | 76.031 |  | 7.5231 | 77.1 | 77.4667 |  | 10.8333 | 78.5667 | 79.05
8.3 |  | 3.5333 | 74.3 | 74.72 |  | 3.4 | 74.7 | 74.925 |  | 4.5 | 75.18 | 75.6611 |  | 7.3982 | 76.7667 | 77.0857 |  | 10.35 | 78.2 | 78.7
8.4 |  | 3.5 | 73.9 | 74.38 |  | 3.3 | 74.3 | 74.5 |  | 4.4 | 74.84 | 75.3312 |  | 7.1982 | 76.38 | 76.775 |  | 10.1667 | 77.9333 | 78.45
8.5 |  | 3.4 | 73.5667 | 73.95 |  | 3.225 | 73.8333 | 74.12 |  | 4.3 | 74.45 | 74.9063 |  | 6.9649 | 75.88 | 76.35 |  | 9.8 | 77.66 | 78.1
8.6 |  | 3.3 | 73.225 | 73.55 |  | 3.2 | 73.5 | 73.75 |  | 4.2 | 74.05 | 74.5814 |  | 6.7483 | 75.65 | 75.95 |  | 9.5667 | 77.4 | 77.8
8.7 |  | 3.3 | 72.6 | 73.15 |  | 3.1 | 73 | 73.4 |  | 4.1 | 73.58 | 74.0815 |  | 6.5983 | 75.3667 | 75.64 |  | 9.3 | 77.0667 | 77.34
8.8 |  | 3.22 | 72.2 | 72.725 |  | 3 | 72.65 | 72.975 |  | 4 | 73.25 | 73.7816 |  | 6.4317 | 75 | 75.35 |  | 9.14 | 76.6667 | 77.1
8.9 |  | 3.1667 | 71.8 | 72.2333 |  | 3 | 72.0667 | 72.46 |  | 3.9 | 72.9 | 73.4816 |  | 6.2984 | 74.56 | 74.98 |  | 8.9667 | 76.3429 | 76.82
9 |  | 3.1 | 71.1 | 71.6 |  | 2.9 | 71.7 | 72.02 |  | 3.8 | 72.56 | 73.0817 |  | 6.1651 | 74.275 | 74.625 |  | 8.7333 | 76 | 76.45
9.1 |  | 3 | 70.6667 | 71.18 |  | 2.9 | 71.2 | 71.5 |  | 3.7 | 72.0833 | 72.7018 |  | 6.0185 | 74 | 74.35 |  | 8.5 | 75.65 | 76.1667
9.2 |  | 3 | 70 | 70.575 |  | 2.8 | 70.4667 | 71.1 |  | 3.6 | 71.64 | 72.202 |  | 5.7985 | 73.45 | 73.8 |  | 8.4 | 75.34 | 75.9167
9.3 |  | 3 | 69.4 | 69.8 |  | 2.8 | 69.95 | 70.35 |  | 3.525 | 71.15 | 71.7321 |  | 5.7236 | 73.15 | 73.525 |  | 8.3 | 74.95 | 75.45
9.4 |  | 2.9 | 68.7667 | 69.12 |  | 2.7 | 69.3667 | 69.675 |  | 3.475 | 70.7 | 71.1322 |  | 5.5986 | 72.86 | 73.06 |  | 8.1 | 74.6 | 75.0667
9.5 |  | 2.9 | 67.6 | 68.3 |  | 2.7 | 68.6667 | 69 |  | 3.4 | 70.25 | 70.7157 |  | 5.4986 | 72.45 | 72.72 |  | 7.8 | 74.25 | 74.75
9.6 |  | 2.8 | 66.625 | 67.175 |  | 2.6 | 67.85 | 68.12 |  | 3.3 | 69.7 | 70.2424 |  | 5.3986 | 72.04 | 72.4 |  | 7.7 | 73.9 | 74.5
9.7 |  | 2.8 | 65.2667 | 66 |  | 2.56 | 66.8 | 67.38 |  | 3.225 | 69.125 | 69.7326 |  | 5.1987 | 71.7667 | 72.06 |  | 7.55 | 73.7 | 74.05
9.8 |  | 2.7 | 63.6333 | 64.3 |  | 2.5 | 65.7 | 66.18 |  | 3.2 | 68.475 | 69.2227 |  | 5.1987 | 71.26 | 71.7 |  | 7.4 | 73.275 | 73.75
9.9 |  | 2.7 | 61.075 | 62.2 |  | 2.5 | 64.25 | 64.92 |  | 3.125 | 67.88 | 68.3829 |  | 4.9988 | 70.9667 | 71.3 |  | 7.275 | 72.95 | 73.5
10 |  | 2.6 | 57.4667 | 58.92 |  | 2.5 | 62.425 | 63.0167 |  | 3.1 | 67.26 | 67.6831 |  | 4.8988 | 70.54 | 70.78 |  | 7.125 | 72.62 | 73.2667
10.1 |  | 2.6 | 52.7 | 54.4 |  | 2.4 | 60.3 | 60.7 |  | 3.05 | 66.2 | 66.8633 |  | 4.8788 | 70.1 | 70.4 |  | 7 | 72.28 | 72.85
10.2 |  |  | 46.9 | 48.4 |  | 2.4 | 56.7 | 57.35 |  | 3 | 65.4 | 65.8835 |  | 4.7488 | 69.7 | 69.9667 |  | 6.8 | 71.9 | 72.5143
10.3 |  |  | 40.3333 | 42.45 |  |  | 52.1 | 52.8 |  | 2.95 | 64.14 | 64.8838 |  | 4.6988 | 69.18 | 69.6 |  | 6.6 | 71.6 | 72.15
10.4 |  |  | 34 | 36.3 |  |  | 46.5333 | 47.18 |  | 2.9 | 62.8333 | 63.6091 |  | 4.5988 | 68.7 | 69.15 |  | 6.5 | 71.26 | 71.7667
10.5 |  |  | 27.5 | 30.4 |  |  | 39.925 | 41.32 |  |  | 61.2 | 62.0245 |  | 4.4989 | 68.15 | 68.56 |  | 6.4 | 70.825 | 71.325
10.6 |  |  | 22.2 | 25.35 |  |  | 33.825 | 35.3 |  |  | 58.8 | 59.9017 |  | 4.3656 | 67.65 | 68.18 |  | 6.2667 | 70.45 | 71.02
10.7 |  |  | 18.4333 | 21.24 |  |  | 27.3 | 29.2 |  |  | 55.775 | 56.7458 |  | 4.2989 | 67.15 | 67.56 |  | 6.1333 | 70.1 | 70.6
10.8 |  |  | 16.1 | 18.75 |  |  | 22.6 | 24.4 |  |  | 51.48 | 52.5619 |  | 4.1989 | 66.55 | 67.1 |  | 6 | 69.68 | 70.2667
10.9 |  |  | 14.3 | 17.075 |  |  | 18.625 | 20.9 |  |  | 46.08 | 47.4881 |  | 4.124 | 65.95 | 66.2667 |  | 5.9 | 69.38 | 69.9333
11 |  |  | 13.1 | 15.9 |  |  | 16.0667 | 18.3 |  |  | 40.6 | 41.8895 |  |  | 65.16 | 65.6 |  | 5.8 | 68.9 | 69.62
11.1 |  |  | 12.2333 | 14.96 |  |  | 14.3 | 16.625 |  |  | 34.35 | 36.3243 |  |  | 64.35 | 64.8 |  | 5.7 | 68.525 | 69.02
11.2 |  |  | 11.55 | 14.25 |  |  | 13.1 | 15.3333 |  |  | 28.525 | 30.9923 |  |  | 63.5 | 63.925 |  | 5.6 | 68.025 | 68.65
11.3 |  |  | 11 | 13.7 |  |  | 12.02 | 14.38 |  |  | 24.2 | 26.1185 |  |  | 62.425 | 62.9333 |  | 5.5 | 67.65 | 68.3
11.4 |  |  | 10.5333 | 13.25 |  |  | 11.3 | 13.66 |  |  | 20.2667 | 22.5694 |  |  | 61.28 | 61.7 |  | 5.5 | 67.15 | 67.76
11.5 |  |  | 10.2 | 12.9 |  |  | 10.6667 | 13.06 |  |  | 17.7 | 20.095 |  |  | 59.5667 | 60.2 |  |  | 66.64 | 67.3
11.6 |  |  | 9.9 | 12.65 |  |  | 10.2333 | 12.58 |  |  | 15.86 | 18.3555 |  |  | 57.55 | 58.02 |  |  | 66.14 | 66.76
11.7 |  |  | 9.6 | 12.36 |  |  | 9.8 | 12.2 |  |  | 14.54 | 16.9957 |  |  | 54.56 | 55.3 |  |  | 65.5333 | 66.18
11.8 |  |  | 9.4667 | 12.2 |  |  | 9.45 | 11.85 |  |  | 13.5 | 16.046 |  |  | 50.875 | 51.6 |  |  | 65.08 | 65.62
11.9 |  |  | 9.3 | 11.94 |  |  | 9.1 | 11.525 |  |  | 12.725 | 15.2462 |  |  | 46.025 | 47.2 |  |  | 64.5286 | 65
12 |  |  | 9.125 | 11.825 |  |  | 8.975 | 11.4 |  |  | 12.1 | 14.6213 |  |  | 41 | 42.3 |  |  | 63.8 | 64.42
12.1 |  |  | 9.0333 | 11.7 |  |  | 8.7 | 11.1667 |  |  | 11.6 | 14.1165 |  |  | 35.4 | 37.3 |  |  | 62.9 | 63.65
12.2 |  |  | 8.95 | 11.6 |  |  | 8.6 | 11 |  |  | 11.1 | 13.6466 |  |  | 30.4 | 32.3 |  |  | 62 | 62.88
12.3 |  |  | 8.9 | 11.5333 |  |  | 8.4 | 10.8 |  |  | 10.7333 | 13.2767 |  |  | 26.15 | 28.14 |  |  | 61.15 | 61.72
12.4 |  |  | 8.8 | 11.45 |  |  | 8.3 | 10.7 |  |  | 10.5 | 12.9568 |  |  | 22.64 | 24.9 |  |  | 60 | 60.6167
12.5 |  |  | 8.7 | 11.3 |  |  | 8.2 | 10.6 |  |  | 10.15 | 12.6769 |  |  | 20.2 | 22.64 |  |  | 58.4 | 59.2333
12.6 |  |  | 8.6 | 11.3 |  |  | 8.1 | 10.48 |  |  | 9.9 | 12.4302 |  |  | 18.54 | 21 |  |  | 56.5 | 57.4
12.7 |  |  | 8.6 | 11.2 |  |  | 8 | 10.4 |  |  | 9.725 | 12.172 |  |  | 17.3 | 19.625 |  |  | 53.6167 | 55.05
12.8 |  |  | 8.5 | 11.1 |  |  | 7.9 | 10.3 |  |  | 9.5 | 11.997 |  |  | 16.2 | 18.7 |  |  | 50.425 | 51.8
12.9 |  |  | 8.45 | 11.04 |  |  | 7.8 | 10.2 |  |  | 9.3333 | 11.847 |  |  | 15.4 | 17.94 |  |  | 46.16 | 47.8
13 |  |  | 8.4 | 10.925 |  |  | 7.8 | 10.18 |  |  | 9.2 | 11.6971 |  |  | 14.7 | 17.3 |  |  | 41.65 | 43.5333
13.1 |  |  | 8.3333 | 10.9 |  |  | 7.7 | 10.1 |  |  | 9.1 | 11.4721 |  |  | 14.2 | 16.76 |  |  | 36.78 | 39
13.2 |  |  | 8.2333 | 10.8 |  |  | 7.6 | 10 |  |  | 8.92 | 11.3972 |  |  | 13.6 | 16.35 |  |  | 32.125 | 34.6
13.3 |  |  | 8.175 | 10.7 |  |  | 7.5 | 9.92 |  |  | 8.8 | 11.2972 |  |  | 13.24 | 15.9 |  |  | 27.9667 | 30.66
13.4 |  |  | 8.1 | 10.6 |  |  | 7.5 | 9.8 |  |  | 8.7 | 11.1805 |  |  | 12.875 | 15.48 |  |  | 24.9 | 27.45
13.5 |  |  | 8 | 10.5 |  |  | 7.4 | 9.7 |  |  | 8.6 | 10.9973 |  |  | 12.58 | 15.2 |  |  | 22.46 | 25.0667
13.6 |  |  | 7.9 | 10.46 |  |  | 7.3 | 9.68 |  |  | 8.5 | 10.9223 |  |  | 12.25 | 14.8667 |  |  | 20.58 | 23.35
13.7 |  |  | 7.8 | 10.34 |  |  | 7.2667 | 9.6 |  |  | 8.4 | 10.7973 |  |  | 11.94 | 14.6571 |  |  | 19.3667 | 22.1
13.8 |  |  | 7.7 | 10.24 |  |  | 7.2 | 9.475 |  |  | 8.3 | 10.6773 |  |  | 11.725 | 14.4 |  |  | 18.2 | 21.16
13.9 |  |  | 7.6 | 10.15 |  |  | 7.0667 | 9.36 |  |  | 8.175 | 10.5974 |  |  | 11.5 | 14.275 |  |  | 17.4 | 20.4333
14 |  |  | 7.55 | 10.05 |  |  | 7 | 9.3 |  |  | 8.1 | 10.4974 |  |  | 11.325 | 14.025 |  |  | 16.78 | 19.7
14.1 |  |  | 7.4667 | 10 |  |  | 6.9 | 9.2 |  |  | 8 | 10.3974 |  |  | 11.2 | 13.8333 |  |  | 16.16 | 19.15
14.2 |  |  | 7.4 | 9.9 |  |  | 6.8 | 9.1 |  |  | 7.975 | 10.2974 |  |  | 11 | 13.66 |  |  | 15.66 | 18.6
14.3 |  |  | 7.3 | 9.7429 |  |  | 6.8 | 9 |  |  | 7.9 | 10.2225 |  |  | 10.8667 | 13.4667 |  |  | 15.3 | 18.2
14.4 |  |  | 7.2 | 9.675 |  |  | 6.675 | 8.9 |  |  | 7.72 | 10.1975 |  |  | 10.75 | 13.275 |  |  | 14.8 | 17.8333
14.5 |  |  | 7.1 | 9.5 |  |  | 6.5667 | 8.8 |  |  | 7.625 | 9.9975 |  |  | 10.5 | 13.175 |  |  | 14.5333 | 17.5
14.6 |  |  | 7 | 9.48 |  |  | 6.5 | 8.7 |  |  | 7.5333 | 9.8975 |  |  | 10.3667 | 13 |  |  | 14.15 | 17.2333
14.7 |  |  | 6.825 | 9.3 |  |  | 6.3667 | 8.6 |  |  | 7.46 | 9.7642 |  |  | 10.2 | 12.8286 |  |  | 13.92 | 16.875
14.8 |  |  | 6.7 | 9.2 |  |  | 6.3 | 8.5 |  |  | 7.3 | 9.6976 |  |  | 10.1 | 12.7 |  |  | 13.7 | 16.6
14.9 |  |  | 6.7 | 9.1 |  |  | 6.2 | 8.425 |  |  | 7.2667 | 9.5976 |  |  | 9.9667 | 12.5833 |  |  | 13.4333 | 16.28
15 |  |  | 6.6 | 9 |  |  | 6.1 | 8.3 |  |  | 7.2 | 9.4976 |  |  | 9.85 | 12.4 |  |  | 13.2 | 16.1
15.1 |  |  | 6.5 | 8.96 |  |  | 6 | 8.275 |  |  | 7.1 | 9.3977 |  |  | 9.74 | 12.3 |  |  | 13 | 15.84
15.2 |  |  | 6.4 | 8.8571 |  |  | 5.98 | 8.15 |  |  | 7 | 9.331 |  |  | 9.675 | 12.15 |  |  | 12.8 | 15.7
15.3 | [6 empty] | 5.9 | 8.1 |  |  | 6.9 | 9.231 |  |  | 9.5 | 12.025 |  |  | 12.65 | 15.5
15.4 | [6 empty] | 5.8 | 7.9167 |  |  | 6.8 | 9.0477 |  |  | 9.375 | 11.9667 |  |  | 12.5 | 15.3333
15.5 | [10 empty] | 6.7167 | 8.9977 |  |  | 9.25 | 11.8 |  |  | 12.28 | 15.2
15.6 | [10 empty] | 6.64 | 8.8978 |  |  | 9.14 | 11.7 |  |  | 12.1 | 15
15.7 | [10 empty] | 6.5 | 8.7978 |  |  | 9.05 | 11.6 |  |  | 12 | 14.8286
15.8 | [10 empty] | 6.4 | 8.6978 |  |  | 8.925 | 11.5 |  |  | 11.78 | 14.7
15.9 | [10 empty] | 6.28 | 8.5978 |  |  | 8.8 | 11.3 |  |  | 11.6 | 14.54
16 | [14 empty] | 8.675 | 11.2 |  |  | 11.55 | 14.24
16.1 | [14 empty] | 8.6 | 11.1 |  |  | 11.34 | 14.2
16.2 | [14 empty] | 8.475 | 11 |  |  | 11.2 | 14.0333
16.3 | [14 empty] | 8.375 | 10.8 |  |  | 11.1 | 13.9
16.4 | [14 empty] | 8.25 | 10.7333 |  |  | 11 | 13.8
16.5 | [14 empty] | 8.1 | 10.6333 |  |  | 10.82 | 13.675
16.6 | [14 empty] | 8.025 | 10.5 |  |  | 10.7 | 13.44
16.7 | [14 empty] | 7.9333 | 10.4167 |  |  | 10.6 | 13.3
16.8 | [14 empty] | 7.8333 | 10.2571 |  |  | 10.4667 | 13.2
16.9 | [14 empty] | 7.75 | 10.1 |  |  | 10.3 | 13.08
17 | [18 empty] | 10.1333 | 12.975
17.1 | [18 empty] | 10 | 12.8
17.2 | [18 empty] | 9.95 | 12.7
17.3 | [18 empty] | 9.8 | 12.6
17.4 | [18 empty] | 9.58 | 12.475
17.5 | [18 empty] | 9.5 | 12.34
17.6 | [18 empty] | 9.48 | 12.2
17.7 | [18 empty] | 9.4 | 12.1
17.8 | [18 empty] | 9.3 | 11.98
17.9 | [18 empty] | 9.2 | 11.8

## Wedge 45 Longitudinal Profiles
Longitudinal Profiles
Wedge 45
SSD=100cm
Field Size=40x20cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:3 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 11.2556
-30.8 |  |  |  |  | 11.4
-30.7 |  |  |  |  | 11.4429
-30.6 |  |  |  |  | 11.575
-30.5 |  |  |  |  | 11.72
-30.4 |  |  |  |  | 11.875
-30.3 |  |  |  |  | 12
-30.2 |  |  |  |  | 12.0857
-30.1 |  |  |  |  | 12.2143
-30 |  |  |  |  | 12.3571
-29.9 |  |  |  |  | 12.5167
-29.8 |  |  |  |  | 12.6833
-29.7 |  |  |  |  | 12.825
-29.6 |  |  |  |  | 13.0333
-29.5 |  |  |  |  | 13.2
-29.4 |  |  |  |  | 13.4
-29.3 |  |  |  |  | 13.525
-29.2 |  |  |  |  | 13.7
-29.1 |  |  |  |  | 13.8667
-29 |  |  |  |  | 14.2
-28.9 |  |  |  | 9.8 | 14.325
-28.8 |  |  |  | 9.9 | 14.5375
-28.7 |  |  |  | 10 | 14.725
-28.6 |  |  |  | 10.1286 | 14.8875
-28.5 |  |  |  | 10.2 | 15.1125
-28.4 |  |  |  | 10.3 | 15.275
-28.3 |  |  |  | 10.4375 | 15.4625
-28.2 |  |  |  | 10.5714 | 15.7
-28.1 |  |  |  | 10.725 | 15.8667
-28 |  |  |  | 10.85 | 16.2125
-27.9 |  |  |  | 10.9 | 16.4125
-27.8 |  |  |  | 11.0667 | 16.6714
-27.7 |  |  |  | 11.2333 | 16.9857
-27.6 |  |  |  | 11.3143 | 17.3286
-27.5 |  |  |  | 11.4571 | 17.6333
-27.4 |  |  |  | 11.625 | 17.9875
-27.3 |  |  |  | 11.8375 | 18.3625
-27.2 |  |  |  | 11.9625 | 18.95
-27.1 |  |  |  | 12.175 | 19.4875
-27 |  |  |  | 12.3333 | 20.1889
-26.9 |  |  | 8.3 | 12.5 | 20.875
-26.8 |  |  | 8.3571 | 12.625 | 21.9
-26.7 |  |  | 8.4143 | 12.8333 | 22.9333
-26.6 |  |  | 8.55 | 13.0333 | 24.45
-26.5 |  |  | 8.6 | 13.25 | 26.225
-26.4 |  |  | 8.7 | 13.4286 | 28.48
-26.3 |  |  | 8.7833 | 13.5625 | 31.4
-26.2 |  |  | 8.8429 | 13.775 | 35.2
-26.1 |  |  | 8.9 | 14.0571 | 39.3875
-26 |  |  | 9.0667 | 14.3429 | 44.3875
-25.9 |  | 7.5 | 9.1286 | 14.625 | 49.85
-25.8 |  | 7.52 | 9.2 | 14.8667 | 55.6125
-25.7 |  | 7.6 | 9.3143 | 15.1333 | 61
-25.6 |  | 7.7 | 9.4 | 15.4 | 65.76
-25.5 | 8 | 7.7 | 9.4875 | 15.8143 | 69.6571
-25.4 | 8.1 | 7.8 | 9.6333 | 16.15 | 72.8857
-25.3 | 8.1 | 7.9 | 9.7857 | 16.6143 | 75.0833
-25.2 | 8.2 | 7.9875 | 9.88 | 17.0571 | 76.7857
-25.1 | 8.2 | 8 | 10.05 | 17.6667 | 78.28
-25 | 8.3 | 8.0333 | 10.2 | 18.3571 | 79.42
-24.9 | 8.4 | 8.1857 | 10.3 | 19.2167 | 80.3286
-24.8 | 8.45 | 8.2 | 10.4143 | 20.2375 | 81.3143
-24.7 | 8.5 | 8.2625 | 10.5571 | 21.375 | 81.9
-24.6 | 8.55 | 8.3333 | 10.8 | 22.7875 | 82.55
-24.5 | 8.7 | 8.4 | 10.9667 | 24.88 | 83.2143
-24.4 | 8.7 | 8.5571 | 11.1125 | 27.3333 | 83.6667
-24.3 | 8.7857 | 8.6 | 11.25 | 30.8 | 84.1
-24.2 | 8.8286 | 8.7 | 11.5 | 35.14 | 84.425
-24.1 | 8.9625 | 8.8 | 11.7429 | 40.1 | 84.75
-24 | 9.0143 | 8.8857 | 11.9 | 46.05 | 85.1333
-23.9 | 9.1 | 9.025 | 12.15 | 52.4 | 85.5
-23.8 | 9.12 | 9.1667 | 12.3333 | 58.75 | 85.86
-23.7 | 9.3 | 9.3143 | 12.5714 | 64.9 | 86.2333
-23.6 | 9.3714 | 9.4 | 12.9 | 69.7 | 86.55
-23.5 | 9.4375 | 9.52 | 13.2125 | 73.85 | 86.75
-23.4 | 9.5833 | 9.7 | 13.5875 | 77.1 | 87.025
-23.3 | 9.6667 | 9.85 | 13.9625 | 79.5333 | 87.3
-23.2 | 9.8333 | 10.0143 | 14.5286 | 81.4 | 87.55
-23.1 | 10 | 10.1571 | 15.1286 | 82.7 | 87.775
-23 | 10.08 | 10.2875 | 15.7714 | 83.8 | 88.025
-22.9 | 10.25 | 10.52 | 16.7 | 84.6875 | 88.2375
-22.8 | 10.42 | 10.7 | 17.6 | 85.4286 | 88.3625
-22.7 | 10.5857 | 10.9 | 19 | 85.9 | 88.4
-22.6 | 10.6571 | 11.2429 | 20.6667 | 86.6143 | 88.625
-22.5 | 10.875 | 11.5286 | 22.7714 | 87.1 | 88.8
-22.4 | 11.0667 | 11.8143 | 25.9 | 87.75 | 88.975
-22.3 | 11.25 | 12.2 | 29.9333 | 88.1 | 89.2
-22.2 | 11.525 | 12.6333 | 34.875 | 88.5667 | 89.36
-22.1 | 11.7857 | 13.2333 | 41.25 | 88.925 | 89.5429
-22 | 12.0714 | 13.8667 | 48.2714 | 89.2571 | 89.75
-21.9 | 12.4 | 14.7 | 55.62 | 89.5875 | 89.9222
-21.8 | 12.8714 | 15.7429 | 63.1333 | 89.9167 | 90.0429
-21.7 | 13.3667 | 17.0667 | 70.05 | 90.1875 | 90.2
-21.6 | 14 | 18.8167 | 75.9375 | 90.4 | 90.375
-21.5 | 14.85 | 21 | 80.4 | 90.65 | 90.625
-21.4 | 15.8667 | 24.3 | 83.3167 | 90.9333 | 90.7875
-21.3 | 17.1714 | 28.95 | 85.6 | 91.1 | 91.0125
-21.2 | 19 | 34.6 | 87.2 | 91.325 | 91.1375
-21.1 | 21.2 | 41.5 | 88.55 | 91.4571 | 91.2714
-21 | 24.7 | 48.7857 | 89.6667 | 91.6 | 91.5286
-20.9 | 29.5 | 57.2167 | 90.3714 | 91.7286 | 91.6833
-20.8 | 35.65 | 65.7167 | 91.12 | 91.9 | 91.9667
-20.7 | 42.9 | 72.9286 | 91.7875 | 92.05 | 92.16
-20.6 | 51.04 | 79.1857 | 92.4125 | 92.2 | 92.2429
-20.5 | 59.525 | 83.5286 | 92.8429 | 92.3375 | 92.45
-20.4 | 67.7333 | 86.8714 | 93.2714 | 92.525 | 92.6
-20.3 | 75.2 | 88.8667 | 93.6333 | 92.7 | 92.75
-20.2 | 81.5 | 90.3714 | 93.9667 | 92.8571 | 92.95
-20.1 | 85.5714 | 91.7 | 94.3333 | 93.075 | 93.0333
-20 | 88.6667 | 92.7571 | 94.6167 | 93.2125 | 93.2
-19.9 | 90.6 | 93.55 | 94.8429 | 93.3429 | 93.3
-19.8 | 92.12 | 94.1833 | 95 | 93.4857 | 93.45
-19.7 | 93.2833 | 94.8143 | 95.2286 | 93.6286 | 93.6571
-19.6 | 94.1286 | 95.3 | 95.3714 | 93.8667 | 93.8
-19.5 | 94.86 | 95.7 | 95.525 | 94 | 94
-19.4 | 95.3571 | 96.06 | 95.6 | 94.1571 | 94.0375
-19.3 | 95.8 | 96.26 | 95.7286 | 94.3833 | 94.1625
-19.2 | 96.1857 | 96.625 | 95.8714 | 94.5143 | 94.375
-19.1 | 96.5429 | 96.8375 | 95.9125 | 94.68 | 94.4286
-19 | 96.8667 | 96.9714 | 96.05 | 94.8667 | 94.5625
-18.9 | 97.12 | 97.1333 | 96.1 | 95 | 94.6
-18.8 | 97.2857 | 97.3714 | 96.2833 | 95.0667 | 94.7
-18.7 | 97.5667 | 97.5 | 96.4143 | 95.15 | 94.8333
-18.6 | 97.7 | 97.6 | 96.5667 | 95.4143 | 94.9
-18.5 | 97.775 | 97.7 | 96.62 | 95.5571 | 95.0429
-18.4 | 97.8125 | 97.7 | 96.7857 | 95.7125 | 95.175
-18.3 | 97.9 | 97.7667 | 96.9429 | 95.86 | 95.3
-18.2 | 97.9143 | 97.8 | 97 | 96.0286 | 95.45
-18.1 | 98 | 97.8714 | 97.2444 | 96.1 | 95.5
-18 | 98.1 | 97.95 | 97.3 | 96.2667 | 95.7
-17.9 | 98.1 | 98 | 97.4143 | 96.3286 | 95.75
-17.8 | 98.1875 | 98.1375 | 97.5667 | 96.4833 | 95.8875
-17.7 | 98.2 | 98.2 | 97.74 | 96.5429 | 95.9143
-17.6 | 98.275 | 98.35 | 97.8 | 96.6 | 96
-17.5 | 98.3143 | 98.44 | 97.9667 | 96.6375 | 96.1286
-17.4 | 98.4 | 98.5 | 98.0286 | 96.7714 | 96.2833
-17.3 | 98.46 | 98.5667 | 98.1833 | 96.9 | 96.4
-17.2 | 98.6 | 98.7286 | 98.3 | 97 | 96.4
-17.1 | 98.7 | 98.8714 | 98.4 | 97.05 | 96.5375
-17 | 98.8 | 99 | 98.5 | 97.2 | 96.6
-16.9 | 98.9 | 99.1 | 98.6 | 97.2571 | 96.7667
-16.8 | 98.9 | 99.1833 | 98.6857 | 97.425 | 96.9429
-16.7 | 98.9833 | 99.25 | 98.7 | 97.5 | 97.1
-16.6 | 99.1375 | 99.4167 | 98.8125 | 97.6 | 97.2
-16.5 | 99.2833 | 99.5 | 98.9 | 97.6333 | 97.2167
-16.4 | 99.4 | 99.6 | 98.9875 | 97.7857 | 97.3
-16.3 | 99.5 | 99.7143 | 99.1 | 97.8 | 97.4
-16.2 | 99.54 | 99.8 | 99.1667 | 97.85 | 97.45
-16.1 | 99.7 | 99.9 | 99.3 | 97.9 | 97.6
-16 | 99.7667 | 99.9 | 99.4 | 98.0375 | 97.75
-15.9 | 99.9 | 100 | 99.4333 | 98.1 | 97.9
-15.8 | 100 | 100.1 | 99.5 | 98.1875 | 98.0375
-15.7 | 100.0667 | 100.2 | 99.5 | 98.3429 | 98.1
-15.6 | 100.1 | 100.2 | 99.575 | 98.4 | 98.1778
-15.5 | 100.2 | 100.2714 | 99.625 | 98.55 | 98.2
-15.4 | 100.2 | 100.3 | 99.7 | 98.6 | 98.3286
-15.3 | 100.2875 | 100.4 | 99.7167 | 98.6571 | 98.4
-15.2 | 100.3143 | 100.4 | 99.8 | 98.7167 | 98.42
-15.1 | 100.4 | 100.4 | 99.8 | 98.8 | 98.5
-15 | 100.4 | 100.4 | 99.8 | 98.9 | 98.6
-14.9 | 100.4833 | 100.4667 | 99.8714 | 99.0429 | 98.675
-14.8 | 100.5 | 100.5 | 99.9 | 99.1 | 98.72
-14.7 | 100.5 | 100.5857 | 99.9625 | 99.2 | 98.8
-14.6 | 100.5571 | 100.6 | 100.1 | 99.2833 | 98.9667
-14.5 | 100.62 | 100.6 | 100.2 | 99.4 | 99.0125
-14.4 | 100.7 | 100.6 | 100.2714 | 99.45 | 99.1
-14.3 | 100.7 | 100.6714 | 100.3286 | 99.5 | 99.1625
-14.2 | 100.7 | 100.7167 | 100.4 | 99.6 | 99.2778
-14.1 | 100.7857 | 100.8 | 100.5 | 99.6333 | 99.3
-14 | 100.8 | 100.8 | 100.58 | 99.7 | 99.3714
-13.9 | 100.8167 | 100.86 | 100.64 | 99.8 | 99.4167
-13.8 | 100.9 | 100.9 | 100.7 | 99.9 | 99.5
-13.7 | 100.9 | 101 | 100.8 | 99.9857 | 99.6
-13.6 | 100.9833 | 101 | 100.8 | 100.1 | 99.6
-13.5 | 101 | 101.1 | 100.9 | 100.1 | 99.6
-13.4 | 101.1 | 101.2 | 100.96 | 100.1714 | 99.6
-13.3 | 101.2 | 101.3 | 101 | 100.2 | 99.6
-13.2 | 101.2 | 101.3625 | 101 | 100.2 | 99.6167
-13.1 | 101.28 | 101.4 | 101.1 | 100.3 | 99.7
-13 | 101.3 | 101.5 | 101.1 | 100.3429 | 99.7
-12.9 | 101.35 | 101.525 | 101.15 | 100.4 | 99.7
-12.8 | 101.4 | 101.6 | 101.2 | 100.5 | 99.7
-12.7 | 101.4 | 101.7 | 101.3 | 100.5 | 99.8
-12.6 | 101.4667 | 101.7 | 101.3333 | 100.5 | 99.8
-12.5 | 101.5 | 101.8 | 101.4 | 100.5875 | 99.8571
-12.4 | 101.5 | 101.8 | 101.4571 | 100.5429 | 99.9
-12.3 | 101.6 | 101.9 | 101.5 | 100.6 | 100
-12.2 | 101.6 | 101.9 | 101.5143 | 100.6 | 100
-12.1 | 101.6 | 101.95 | 101.6 | 100.6 | 100
-12 | 101.7 | 102 | 101.6 | 100.6 | 100
-11.9 | 101.7 | 102.0667 | 101.6 | 100.6 | 100.2
-11.8 | 101.8 | 102.1 | 101.6 | 100.6 | 100.2571
-11.7 | 101.8 | 102.1 | 101.7 | 100.6875 | 100.3
-11.6 | 101.8167 | 102.1 | 101.7 | 100.6429 | 100.3125
-11.5 | 101.9 | 102.1 | 101.7 | 100.7 | 100.4
-11.4 | 102 | 102.1 | 101.7 | 100.7 | 100.4
-11.3 | 102 | 102.1 | 101.7 | 100.75 | 100.4
-11.2 | 102 | 102.1 | 101.7 | 100.8 | 100.5
-11.1 | 101.9714 | 102.1 | 101.7 | 100.8286 | 100.525
-11 | 101.9833 | 102.1 | 101.6143 | 100.9 | 100.6
-10.9 | 102 | 102.1 | 101.6 | 100.9 | 100.6
-10.8 | 102 | 102.1 | 101.6 | 101 | 100.7
-10.7 | 102 | 102.1 | 101.6 | 101 | 100.7833
-10.6 | 102 | 102.1 | 101.6 | 101.1 | 100.8
-10.5 | 102 | 102.1 | 101.6 | 101.1833 | 100.8
-10.4 | 101.9429 | 102 | 101.65 | 101.25 | 100.9143
-10.3 | 101.9 | 102 | 101.7 | 101.3 | 101
-10.2 | 101.9 | 102 | 101.7286 | 101.3 | 101
-10.1 | 101.9 | 102 | 101.8 | 101.4 | 101
-10 | 101.9 | 102 | 101.8 | 101.4 | 101
-9.9 | 101.9 | 102 | 101.9 | 101.4 | 101.0429
-9.8 | 101.9 | 102.0286 | 101.9714 | 101.5 | 101.1
-9.7 | 101.9 | 102.1 | 102 | 101.5 | 101.1
-9.6 | 101.9286 | 102.2 | 102.1 | 101.5 | 101.1
-9.5 | 102 | 102.2714 | 102.1 | 101.6 | 101.1
-9.4 | 102.1 | 102.3 | 102.1143 | 101.6375 | 101.1
-9.3 | 102.2 | 102.4 | 102.2 | 101.7 | 101.2
-9.2 | 102.3 | 102.4167 | 102.2 | 101.7 | 101.3
-9.1 | 102.3 | 102.5 | 102.3 | 101.8 | 101.3571
-9 | 102.3 | 102.5 | 102.35 | 101.7625 | 101.4
-8.9 | 102.3833 | 102.5857 | 102.4 | 101.7625 | 101.4
-8.8 | 102.46 | 102.6 | 102.4 | 101.8 | 101.4
-8.7 | 102.5333 | 102.6 | 102.4 | 101.8714 | 101.5
-8.6 | 102.6 | 102.6 | 102.4 | 101.9 | 101.4167
-8.5 | 102.6 | 102.6167 | 102.4 | 101.9 | 101.46
-8.4 | 102.7 | 102.7 | 102.4 | 101.9 | 101.6
-8.3 | 102.6714 | 102.7 | 102.4 | 101.9 | 101.65
-8.2 | 102.6 | 102.7 | 102.4286 | 101.9 | 101.7
-8.1 | 102.65 | 102.8 | 102.5 | 101.9 | 101.7
-8 | 102.7 | 102.8 | 102.5 | 101.9833 | 101.8
-7.9 | 102.7 | 102.8 | 102.5833 | 102 | 101.8
-7.8 | 102.7 | 102.8833 | 102.6 | 102 | 101.8
-7.7 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
-7.6 | 102.75 | 102.9 | 102.6 | 102.1 | 101.9
-7.5 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8
-7.4 | 102.8 | 102.9 | 102.7 | 102.1857 | 101.8571
-7.3 | 102.85 | 102.9 | 102.8 | 102.2143 | 101.9
-7.2 | 102.9 | 102.9833 | 102.8 | 102.3 | 101.9
-7.1 | 102.9 | 103 | 102.8 | 102.3 | 101.9
-7 | 102.94 | 103 | 102.8 | 102.3 | 101.9
-6.9 | 103 | 103 | 102.8 | 102.3 | 101.9429
-6.8 | 103 | 103 | 102.8571 | 102.3 | 102
-6.7 | 103 | 103 | 102.9 | 102.3 | 102
-6.6 | 103 | 103.0833 | 102.9 | 102.3 | 102
-6.5 | 103.0667 | 103.1 | 102.9 | 102.4 | 102
-6.4 | 103 | 103.1 | 102.9 | 102.4 | 102
-6.3 | 103 | 103.2 | 102.9 | 102.4 | 102
-6.2 | 103.1 | 103.2 | 102.9 | 102.4 | 102
-6.1 | 103.1 | 103.2 | 102.9 | 102.4 | 102
-6 | 103.1 | 103.2 | 102.9 | 102.4 | 102
-5.9 | 103.1 | 103.2 | 102.9 | 102.4 | 102.1
-5.8 | 103.1 | 103.15 | 102.9 | 102.4 | 102
-5.7 | 103.1 | 103.1833 | 102.9 | 102.3625 | 102
-5.6 | 103.2 | 103.1 | 102.8667 | 102.3 | 102
-5.5 | 103.2 | 103.1 | 102.8 | 102.3 | 101.9667
-5.4 | 103.2 | 103.1 | 102.8667 | 102.3 | 101.9
-5.3 | 103.1 | 103.1 | 102.8 | 102.3 | 101.9
-5.2 | 103.1 | 103.1 | 102.8 | 102.3 | 101.9
-5.1 | 103.1 | 103.1 | 102.8 | 102.2 | 101.9
-5 | 103 | 103.0333 | 102.8 | 102.16 | 101.95
-4.9 | 103 | 103 | 102.8 | 102.1875 | 101.9
-4.8 | 103 | 103 | 102.8 | 102.2 | 101.9
-4.7 | 103 | 103 | 102.7 | 102.1889 | 101.9
-4.6 | 103 | 102.9167 | 102.7 | 102.1 | 101.9
-4.5 | 102.9 | 102.9 | 102.6857 | 102.1 | 101.9
-4.4 | 102.9 | 102.8857 | 102.55 | 102.1 | 101.9
-4.3 | 102.8 | 102.8 | 102.5 | 102.0125 | 101.9857
-4.2 | 102.8 | 102.7 | 102.5 | 102 | 101.9
-4.1 | 102.8 | 102.7 | 102.5 | 102 | 101.9
-4 | 102.8 | 102.7 | 102.5 | 102 | 101.8
-3.9 | 102.7 | 102.6 | 102.5 | 102 | 101.8
-3.8 | 102.6429 | 102.6 | 102.5 | 102 | 101.7375
-3.7 | 102.6 | 102.5 | 102.5 | 101.9167 | 101.7
-3.6 | 102.5 | 102.5 | 102.4 | 102 | 101.7
-3.5 | 102.5 | 102.4714 | 102.3167 | 101.9 | 101.6
-3.4 | 102.45 | 102.4 | 102.2429 | 101.9 | 101.6
-3.3 | 102.38 | 102.35 | 102.1 | 101.8 | 101.5
-3.2 | 102.3 | 102.3 | 102.1 | 101.7125 | 101.5
-3.1 | 102.3 | 102.25 | 102 | 101.6 | 101.5
-3 | 102.25 | 102.2 | 102 | 101.6 | 101.4
-2.9 | 102.2 | 102.0429 | 101.9429 | 101.4857 | 101.3
-2.8 | 102.1 | 101.9857 | 101.9 | 101.3333 | 101.2375
-2.7 | 102 | 101.8333 | 101.8 | 101.3 | 101.2
-2.6 | 102 | 101.8 | 101.7 | 101.2286 | 101.04
-2.5 | 101.7667 | 101.7 | 101.5833 | 101.1833 | 101
-2.4 | 101.6 | 101.6429 | 101.4167 | 101.1 | 101
-2.3 | 101.5333 | 101.5125 | 101.3833 | 101 | 100.9
-2.2 | 101.4333 | 101.35 | 101.3 | 100.9 | 100.8833
-2.1 | 101.3667 | 101.2 | 101.1429 | 100.8 | 100.8
-2 | 101.2 | 101.1 | 101.0125 | 100.8 | 100.75
-1.9 | 101.1 | 101 | 100.9833 | 100.7 | 100.7
-1.8 | 100.9857 | 100.9 | 100.9 | 100.7 | 100.6571
-1.7 | 100.9 | 100.8167 | 100.8 | 100.6 | 100.6
-1.6 | 100.8 | 100.7 | 100.7 | 100.5571 | 100.5
-1.5 | 100.7 | 100.6 | 100.7 | 100.5 | 100.5
-1.4 | 100.6 | 100.6 | 100.6857 | 100.4286 | 100.4286
-1.3 | 100.5 | 100.5 | 100.6 | 100.4 | 100.35
-1.2 | 100.5 | 100.5 | 100.4667 | 100.4 | 100.3
-1.1 | 100.4333 | 100.4 | 100.4 | 100.4 | 100.3
-1 | 100.3667 | 100.3 | 100.36 | 100.3 | 100.2111
-0.9 | 100.3 | 100.2667 | 100.3 | 100.2 | 100.1
-0.8 | 100.2286 | 100.2 | 100.2429 | 100.2 | 100.1
-0.7 | 100.2 | 100.1333 | 100.2667 | 100.2 | 100.0286
-0.6 | 100.2 | 100.1 | 100.2 | 100.1 | 100
-0.5 | 100.1 | 100 | 100.2 | 100 | 100
-0.4 | 100.1 | 100 | 100.1 | 100 | 100
-0.3 | 100 | 100 | 100.025 | 100 | 99.9333
-0.2 | 100 | 100 | 100 | 99.9125 | 99.9333
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 99.9125 | 99.9
0.3 | 100.0167 | 100 | 100.0429 | 100 | 99.9
0.4 | 100.1 | 100 | 100.1 | 100 | 100
0.5 | 100.1 | 100 | 100.2 | 100 | 100
0.6 | 100.2 | 100.1 | 100.2 | 100.1 | 100
0.7 | 100.2 | 100.1333 | 100.3 | 100.2 | 100
0.8 | 100.2571 | 100.2 | 100.3 | 100.2 | 100.1
0.9 | 100.3 | 100.2667 | 100.3 | 100.2 | 100.1
1 | 100.4 | 100.3 | 100.3 | 100.3 | 100.2
1.1 | 100.4 | 100.4 | 100.4 | 100.4 | 100.2857
1.2 | 100.5 | 100.5 | 100.4833 | 100.4 | 100.3
1.3 | 100.5 | 100.5 | 100.6 | 100.4 | 100.4
1.4 | 100.6 | 100.6 | 100.6125 | 100.4286 | 100.4714
1.5 | 100.7167 | 100.6 | 100.7 | 100.5 | 100.5
1.6 | 100.8 | 100.7 | 100.7 | 100.5571 | 100.55
1.7 | 100.9 | 100.8167 | 100.8 | 100.6 | 100.6
1.8 | 101 | 100.9 | 100.9 | 100.7 | 100.7
1.9 | 101.1 | 101 | 100.9 | 100.7 | 100.7
2 | 101.1571 | 101.1 | 101.0375 | 100.8 | 100.7
2.1 | 101.3 | 101.2 | 101.1714 | 100.8 | 100.8
2.2 | 101.4667 | 101.35 | 101.3 | 100.9 | 100.9
2.3 | 101.56 | 101.5125 | 101.4 | 101 | 100.9
2.4 | 101.6286 | 101.6429 | 101.4429 | 101.1 | 101
2.5 | 101.7833 | 101.7 | 101.6 | 101.1833 | 101
2.6 | 101.9333 | 101.8 | 101.6714 | 101.2286 | 101.0833
2.7 | 102 | 101.8333 | 101.8 | 101.3 | 101.15
2.8 | 102.0833 | 101.9857 | 101.9 | 101.3333 | 101.2875
2.9 | 102.2 | 102.0429 | 101.9667 | 101.4857 | 101.3
3 | 102.2 | 102.2 | 102 | 101.6 | 101.4
3.1 | 102.3 | 102.25 | 102 | 101.6 | 101.5
3.2 | 102.3 | 102.3 | 102.05 | 101.7125 | 101.5
3.3 | 102.4 | 102.35 | 102.1 | 101.8 | 101.5143
3.4 | 102.4667 | 102.4 | 102.2571 | 101.9 | 101.6
3.5 | 102.5 | 102.4714 | 102.3 | 101.9 | 101.6
3.6 | 102.5 | 102.5 | 102.4 | 102 | 101.68
3.7 | 102.6 | 102.5 | 102.48 | 101.9167 | 101.7
3.8 | 102.6714 | 102.6 | 102.5 | 102 | 101.7875
3.9 | 102.7333 | 102.6 | 102.5 | 102 | 101.8
4 | 102.8 | 102.7 | 102.5 | 102 | 101.8429
4.1 | 102.8 | 102.7 | 102.5 | 102 | 101.9
4.2 | 102.8 | 102.7 | 102.5 | 102 | 101.9
4.3 | 102.8 | 102.8 | 102.5 | 102.0125 | 101.9333
4.4 | 102.8857 | 102.8857 | 102.575 | 102.1 | 101.925
4.5 | 102.9 | 102.9 | 102.7 | 102.1 | 101.9
4.6 | 103 | 102.9167 | 102.7 | 102.1 | 101.9
4.7 | 103 | 103 | 102.7 | 102.1889 | 101.9
4.8 | 103 | 103 | 102.8 | 102.2 | 101.9
4.9 | 103 | 103 | 102.8 | 102.1875 | 101.9
5 | 103 | 103.0333 | 102.8 | 102.16 | 101.9
5.1 | 103.1 | 103.1 | 102.8 | 102.2 | 101.9
5.2 | 103.1 | 103.1 | 102.8 | 102.3 | 101.9
5.3 | 103.1 | 103.1 | 102.8 | 102.3 | 101.9
5.4 | 103.2 | 103.1 | 102.8333 | 102.3 | 101.9
5.5 | 103.2 | 103.1 | 102.8 | 102.3 | 101.9143
5.6 | 103.2 | 103.1 | 102.8 | 102.3 | 102
5.7 | 103.1 | 103.1833 | 102.9 | 102.3625 | 102
5.8 | 103.1 | 103.15 | 102.9 | 102.4 | 102.05
5.9 | 103.1 | 103.2 | 102.9 | 102.4 | 102.1
6 | 103.1 | 103.2 | 102.9 | 102.4 | 102
6.1 | 103.1 | 103.2 | 102.9 | 102.4 | 102
6.2 | 103.1 | 103.2 | 102.9 | 102.4 | 102
6.3 | 103.06 | 103.2 | 102.9 | 102.4 | 102
6.4 | 103 | 103.1 | 102.9 | 102.4 | 102
6.5 | 103.1 | 103.1 | 102.9 | 102.4 | 102
6.6 | 103 | 103.0833 | 102.9 | 102.3 | 102
6.7 | 103 | 103 | 102.9 | 102.3 | 102
6.8 | 103 | 103 | 102.8286 | 102.3 | 101.9571
6.9 | 103 | 103 | 102.8 | 102.3 | 101.9
7 | 102.9167 | 103 | 102.8 | 102.3 | 101.94
7.1 | 102.9 | 103 | 102.8 | 102.3 | 101.9
7.2 | 102.9 | 102.9833 | 102.8 | 102.3 | 101.9
7.3 | 102.8333 | 102.9 | 102.8 | 102.2143 | 101.9
7.4 | 102.8 | 102.9 | 102.7 | 102.1857 | 101.9
7.5 | 102.8 | 102.9 | 102.7 | 102.1 | 101.8333
7.6 | 102.72 | 102.9 | 102.6857 | 102.1 | 101.8833
7.7 | 102.7 | 102.9 | 102.6 | 102.1 | 101.8
7.8 | 102.7 | 102.8833 | 102.6 | 102 | 101.8
7.9 | 102.7 | 102.8 | 102.56 | 102 | 101.8
8 | 102.7 | 102.8 | 102.5 | 101.9833 | 101.8
8.1 | 102.6333 | 102.8 | 102.5 | 101.9 | 101.7375
8.2 | 102.6 | 102.7 | 102.4 | 101.9 | 101.7
8.3 | 102.7 | 102.7 | 102.4 | 101.9 | 101.7
8.4 | 102.7 | 102.7 | 102.4 | 101.9 | 101.5714
8.5 | 102.6 | 102.6167 | 102.4 | 101.9 | 101.5
8.6 | 102.6 | 102.6 | 102.4 | 101.9 | 101.4
8.7 | 102.6 | 102.6 | 102.4 | 101.8714 | 101.4714
8.8 | 102.4333 | 102.6 | 102.4 | 101.8 | 101.4
8.9 | 102.36 | 102.5857 | 102.4 | 101.7625 | 101.4
9 | 102.3 | 102.5 | 102.4 | 101.7625 | 101.4
9.1 | 102.3 | 102.5 | 102.2833 | 101.8 | 101.4
9.2 | 102.2833 | 102.4167 | 102.2 | 101.7 | 101.2571
9.3 | 102.2 | 102.4 | 102.2 | 101.7 | 101.2
9.4 | 102.1429 | 102.3 | 102.2 | 101.6375 | 101.1429
9.5 | 102 | 102.2714 | 102.1 | 101.6 | 101.1
9.6 | 102 | 102.2 | 102.1 | 101.5 | 101.1
9.7 | 101.9 | 102.1 | 102.0286 | 101.5 | 101.1
9.8 | 101.9 | 102.0286 | 101.95 | 101.5 | 101.1
9.9 | 101.9 | 102 | 101.9 | 101.4 | 101
10 | 101.9 | 102 | 101.8333 | 101.4 | 101
10.1 | 101.9 | 102 | 101.8 | 101.4 | 101
10.2 | 101.9 | 102 | 101.7 | 101.3 | 101
10.3 | 101.9 | 102 | 101.7 | 101.3 | 101
10.4 | 101.9667 | 102 | 101.625 | 101.25 | 100.9
10.5 | 102 | 102.1 | 101.6 | 101.1833 | 100.8286
10.6 | 102 | 102.1 | 101.6 | 101.1 | 100.8
10.7 | 102 | 102.1 | 101.6 | 101 | 100.7333
10.8 | 102 | 102.1 | 101.6 | 101 | 100.7
10.9 | 102 | 102.1 | 101.6 | 100.9 | 100.6333
11 | 101.9571 | 102.1 | 101.6375 | 100.9 | 100.6
11.1 | 101.9 | 102.1 | 101.7 | 100.8286 | 100.5
11.2 | 102 | 102.1 | 101.7 | 100.8 | 100.5
11.3 | 102 | 102.1 | 101.7 | 100.75 | 100.4286
11.4 | 102 | 102.1 | 101.7 | 100.7 | 100.4
11.5 | 101.9 | 102.1 | 101.7 | 100.7 | 100.4
11.6 | 101.8 | 102.1 | 101.7 | 100.6429 | 100.3625
11.7 | 101.8 | 102.1 | 101.7 | 100.6875 | 100.3
11.8 | 101.8 | 102.1 | 101.6 | 100.6 | 100.2
11.9 | 101.7 | 102.0667 | 101.6 | 100.6 | 100.1625
12 | 101.7 | 102 | 101.6 | 100.6 | 100.0375
12.1 | 101.6 | 101.95 | 101.6 | 100.6 | 100.0857
12.2 | 101.6 | 101.9 | 101.5 | 100.6 | 100
12.3 | 101.6 | 101.9 | 101.5 | 100.6 | 100
12.4 | 101.5 | 101.8 | 101.5 | 100.5429 | 99.95
12.5 | 101.5 | 101.8 | 101.4 | 100.5875 | 99.9
12.6 | 101.4429 | 101.7 | 101.3167 | 100.5 | 99.8
12.7 | 101.4 | 101.7 | 101.3 | 100.5 | 99.7429
12.8 | 101.4 | 101.6 | 101.2 | 100.5 | 99.7
12.9 | 101.3167 | 101.525 | 101.2 | 100.4 | 99.7
13 | 101.3 | 101.5 | 101.1 | 100.3429 | 99.7
13.1 | 101.25 | 101.4 | 101.1 | 100.3 | 99.7
13.2 | 101.2 | 101.3625 | 101 | 100.2 | 99.6667
13.3 | 101.2 | 101.3 | 101 | 100.2 | 99.6
13.4 | 101.1167 | 101.2 | 100.9333 | 100.1714 | 99.6
13.5 | 101.04 | 101.1 | 100.8857 | 100.1 | 99.6
13.6 | 100.95 | 101 | 100.8 | 100.1 | 99.6
13.7 | 100.9 | 101 | 100.7875 | 99.9857 | 99.6
13.8 | 100.9 | 100.9 | 100.7 | 99.9 | 99.54
13.9 | 100.8 | 100.86 | 100.6167 | 99.8 | 99.4667
14 | 100.8 | 100.8 | 100.55 | 99.7 | 99.325
14.1 | 100.7571 | 100.8 | 100.5 | 99.6333 | 99.3
14.2 | 100.7 | 100.7167 | 100.4571 | 99.6 | 99.2375
14.3 | 100.7 | 100.6714 | 100.4 | 99.5 | 99.1125
14.4 | 100.7 | 100.6 | 100.25 | 99.45 | 99.0875
14.5 | 100.6 | 100.6 | 100.2 | 99.4 | 99
14.6 | 100.5286 | 100.6 | 100.0714 | 99.2833 | 98.9167
14.7 | 100.5 | 100.5857 | 100 | 99.2 | 98.8333
14.8 | 100.5 | 100.5 | 99.9 | 99.1 | 98.7833
14.9 | 100.5 | 100.4667 | 99.8429 | 99.0429 | 98.6286
15 | 100.4 | 100.4 | 99.8 | 98.9 | 98.6
15.1 | 100.4 | 100.4 | 99.8 | 98.8 | 98.5
15.2 | 100.3 | 100.4 | 99.8 | 98.7167 | 98.4
15.3 | 100.2625 | 100.4 | 99.7833 | 98.6571 | 98.4
15.4 | 100.2 | 100.3 | 99.7 | 98.6 | 98.3
15.5 | 100.1833 | 100.2714 | 99.6 | 98.55 | 98.25
15.6 | 100.1 | 100.2 | 99.6 | 98.4 | 98.2
15.7 | 100.04 | 100.2 | 99.5143 | 98.3429 | 98.1111
15.8 | 100 | 100.1 | 99.5 | 98.1875 | 98.0875
15.9 | 99.9333 | 100 | 99.5 | 98.1 | 97.9143
16 | 99.8 | 99.9 | 99.4 | 98.0375 | 97.8
16.1 | 99.6833 | 99.9 | 99.3 | 97.9 | 97.6
16.2 | 99.5167 | 99.8 | 99.18 | 97.85 | 97.4
16.3 | 99.5 | 99.7143 | 99.0875 | 97.8 | 97.3571
16.4 | 99.45 | 99.6 | 98.9625 | 97.7857 | 97.3
16.5 | 99.36 | 99.5 | 98.9 | 97.6333 | 97.2571
16.6 | 99.125 | 99.4167 | 98.8 | 97.6 | 97.2
16.7 | 99 | 99.25 | 98.7333 | 97.5 | 97.05
16.8 | 98.9125 | 99.1833 | 98.6667 | 97.425 | 96.9
16.9 | 98.8714 | 99.1 | 98.6 | 97.2571 | 96.8167
17 | 98.8 | 99 | 98.4667 | 97.2 | 96.65
17.1 | 98.6714 | 98.8714 | 98.4 | 97.05 | 96.4875
17.2 | 98.6 | 98.7286 | 98.3143 | 97 | 96.4
17.3 | 98.5 | 98.5667 | 98.16 | 96.9 | 96.3714
17.4 | 98.4 | 98.5 | 98.1 | 96.7714 | 96.2286
17.5 | 98.3 | 98.44 | 97.9333 | 96.6375 | 96.1833
17.6 | 98.25 | 98.35 | 97.8857 | 96.6 | 96.0167
17.7 | 98.2 | 98.2 | 97.72 | 96.5429 | 95.9625
17.8 | 98.1625 | 98.1375 | 97.6429 | 96.4833 | 95.8444
17.9 | 98.1 | 98 | 97.4857 | 96.3286 | 95.7
18 | 98.0857 | 97.95 | 97.3429 | 96.2667 | 95.7
18.1 | 98 | 97.8714 | 97.2222 | 96.1 | 95.5333
18.2 | 97.9 | 97.8 | 97.0714 | 96.0286 | 95.4
18.3 | 97.9 | 97.7667 | 96.9167 | 95.86 | 95.2571
18.4 | 97.8875 | 97.7 | 96.7667 | 95.7125 | 95.125
18.5 | 97.8 | 97.7 | 96.7 | 95.5571 | 95
18.6 | 97.7143 | 97.6 | 96.5429 | 95.4143 | 94.9429
18.7 | 97.5333 | 97.5 | 96.3875 | 95.15 | 94.7833
18.8 | 97.3333 | 97.3714 | 96.25 | 95.0667 | 94.7
18.9 | 97.1 | 97.1333 | 96.1833 | 95 | 94.65
19 | 96.9 | 96.9714 | 96.0286 | 94.8667 | 94.5125
19.1 | 96.5286 | 96.8375 | 95.9 | 94.68 | 94.4
19.2 | 96.2 | 96.625 | 95.8429 | 94.5143 | 94.3375
19.3 | 95.8 | 96.26 | 95.7 | 94.3833 | 94.2125
19.4 | 95.3875 | 96.06 | 95.6333 | 94.1571 | 94.0875
19.5 | 94.9 | 95.7 | 95.5 | 94 | 93.9667
19.6 | 94.15 | 95.3 | 95.35 | 93.8667 | 93.85
19.7 | 93.3 | 94.8143 | 95.2 | 93.6286 | 93.7
19.8 | 92.1 | 94.1833 | 95.05 | 93.4857 | 93.475
19.9 | 90.7 | 93.55 | 94.85 | 93.3429 | 93.3333
20 | 88.6 | 92.7571 | 94.6 | 93.2125 | 93.16
20.1 | 85.7 | 91.7 | 94.3429 | 93.075 | 93.0833
20.2 | 81.3 | 90.3714 | 94.0333 | 92.8571 | 92.9286
20.3 | 75.12 | 88.8667 | 93.7 | 92.7 | 92.7
20.4 | 67.72 | 86.8714 | 93.2857 | 92.525 | 92.6
20.5 | 59.5 | 83.5286 | 92.8571 | 92.3375 | 92.4444
20.6 | 50.9333 | 79.1857 | 92.3714 | 92.2 | 92.3
20.7 | 42.96 | 72.9286 | 91.85 | 92.05 | 92.2
20.8 | 35.4286 | 65.7167 | 91.1 | 91.9 | 92.0167
20.9 | 29.7333 | 57.2167 | 90.35 | 91.7286 | 91.68
21 | 24.7429 | 48.7857 | 89.7667 | 91.6 | 91.4714
21.1 | 21.2 | 41.5 | 88.54 | 91.4571 | 91.3286
21.2 | 18.8667 | 34.6 | 87.1857 | 91.325 | 91.175
21.3 | 17.15 | 28.95 | 85.5667 | 91.1 | 90.9625
21.4 | 15.8 | 24.3 | 83.4667 | 90.9333 | 90.775
21.5 | 14.8 | 21 | 80.3 | 90.65 | 90.6125
21.6 | 14.025 | 18.8167 | 75.8625 | 90.4 | 90.3857
21.7 | 13.35 | 17.0667 | 69.9875 | 90.1875 | 90.2429
21.8 | 12.8857 | 15.7429 | 63.5 | 89.9167 | 90.0889
21.9 | 12.4 | 14.7 | 55.8 | 89.5875 | 89.9429
22 | 12.1 | 13.8667 | 48.1 | 89.2571 | 89.7286
22.1 | 11.75 | 13.2333 | 41.2 | 88.925 | 89.6
22.2 | 11.5714 | 12.6333 | 35.0429 | 88.5667 | 89.3167
22.3 | 11.24 | 12.2 | 29.6 | 88.1 | 89.15
22.4 | 11.04 | 11.8143 | 25.6667 | 87.75 | 88.95
22.5 | 10.9429 | 11.5286 | 22.8667 | 87.1 | 88.8
22.6 | 10.7286 | 11.2429 | 20.6 | 86.6143 | 88.65
22.7 | 10.5571 | 10.9 | 19.0286 | 85.9 | 88.45
22.8 | 10.4 | 10.7 | 17.6833 | 85.4286 | 88.3143
22.9 | 10.22 | 10.52 | 16.6625 | 84.6875 | 88.2
23 | 10.1 | 10.2875 | 15.7 | 83.8 | 88.0875
23.1 | 9.9833 | 10.1571 | 15.0857 | 82.7 | 87.7889
23.2 | 9.9 | 10.0143 | 14.5429 | 81.4 | 87.5
23.3 | 9.64 | 9.85 | 13.9571 | 79.5333 | 87.3
23.4 | 9.56 | 9.7 | 13.5286 | 77.1 | 87.05
23.5 | 9.5 | 9.52 | 13.1375 | 73.85 | 86.7
23.6 | 9.35 | 9.4 | 12.925 | 69.7 | 86.5429
23.7 | 9.28 | 9.3143 | 12.5857 | 64.9 | 86.2333
23.8 | 9.2 | 9.1667 | 12.4 | 58.75 | 85.8667
23.9 | 9.1 | 9.025 | 12.1333 | 52.4 | 85.5143
24 | 9 | 8.8857 | 11.95 | 46.05 | 85.1333
24.1 | 8.9429 | 8.8 | 11.7429 | 40.1 | 84.8
24.2 | 8.9 | 8.7 | 11.4833 | 35.14 | 84.425
24.3 | 8.8 | 8.6 | 11.3286 | 30.8 | 84.05
24.4 | 8.7 | 8.5571 | 11.0857 | 27.3333 | 83.6667
24.5 | 8.7 | 8.4 | 10.9333 | 24.88 | 83.1857
24.6 | 8.6 | 8.3333 | 10.75 | 22.7875 | 82.5
24.7 | 8.5 | 8.2625 | 10.5333 | 21.375 | 81.9143
24.8 | 8.425 | 8.2 | 10.4833 | 20.2375 | 81.4
24.9 | 8.3833 | 8.1857 | 10.3 | 19.2167 | 80.3571
25 | 8.3 | 8.0333 | 10.1714 | 18.3571 | 79.48
25.1 | 8.2571 | 8 | 10.0286 | 17.6667 | 78.3
25.2 | 8.2 | 7.9875 | 9.85 | 17.0571 | 76.8571
25.3 | 8.1 | 7.9 | 9.7667 | 16.6143 | 75.1
25.4 | 8.08 | 7.8 | 9.6 | 16.15 | 72.8
25.5 | 8 | 7.7 | 9.4625 | 15.8143 | 69.7571
25.6 |  | 7.7 | 9.4 | 15.4 | 65.88
25.7 |  | 7.6 | 9.3 | 15.1333 | 60.7
25.8 |  | 7.52 | 9.2429 | 14.8667 | 55.675
25.9 |  | 7.5 | 9.1 | 14.625 | 49.925
26 |  |  | 9.04 | 14.3429 | 44.35
26.1 |  |  | 8.9 | 14.0571 | 39.3111
26.2 |  |  | 8.8167 | 13.775 | 35.375
26.3 |  |  | 8.7667 | 13.5625 | 31.45
26.4 |  |  | 8.6857 | 13.4286 | 28.5429
26.5 |  |  | 8.6 | 13.25 | 26.3143
26.6 |  |  | 8.525 | 13.0333 | 24.3
26.7 |  |  | 8.4 | 12.8333 | 23.0667
26.8 |  |  | 8.3333 | 12.625 | 21.8571
26.9 |  |  | 8.3 | 12.5 | 20.875
27 |  |  |  | 12.3333 | 20.2
27.1 |  |  |  | 12.175 | 19.4375
27.2 |  |  |  | 11.9625 | 18.8875
27.3 |  |  |  | 11.8375 | 18.3875
27.4 |  |  |  | 11.625 | 18.025
27.5 |  |  |  | 11.4571 | 17.7
27.6 |  |  |  | 11.3143 | 17.32
27.7 |  |  |  | 11.2333 | 17
27.8 |  |  |  | 11.0667 | 16.6571
27.9 |  |  |  | 10.9 | 16.3875
28 |  |  |  | 10.85 | 16.1875
28.1 |  |  |  | 10.725 | 15.9167
28.2 |  |  |  | 10.5714 | 15.7
28.3 |  |  |  | 10.4375 | 15.5125
28.4 |  |  |  | 10.3 | 15.275
28.5 |  |  |  | 10.2 | 15.0625
28.6 |  |  |  | 10.1286 | 14.875
28.7 |  |  |  | 10 | 14.7125
28.8 |  |  |  | 9.9 | 14.4875
28.9 |  |  |  | 9.8 | 14.325
29 |  |  |  |  | 14.1375
29.1 |  |  |  |  | 13.9444
29.2 |  |  |  |  | 13.7
29.3 |  |  |  |  | 13.575
29.4 |  |  |  |  | 13.4429
29.5 |  |  |  |  | 13.2
29.6 |  |  |  |  | 13.0556
29.7 |  |  |  |  | 12.8625
29.8 |  |  |  |  | 12.7286
29.9 |  |  |  |  | 12.56
30 |  |  |  |  | 12.325
30.1 |  |  |  |  | 12.1714
30.2 |  |  |  |  | 12.0286
30.3 |  |  |  |  | 11.9833
30.4 |  |  |  |  | 11.8286
30.5 |  |  |  |  | 11.7833
30.6 |  |  |  |  | 11.6286
30.7 |  |  |  |  | 11.4
30.8 |  |  |  |  | 11.35
30.9 |  |  |  |  | 11.3

## Wedge 45 Output Factors
Output factors
Wedge 45 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20
| 3 | 0.868 | 0.884 | 0.893 | 0.902 | 0.908 | 0.914 | 0.918
| 4 | 0.886 | 0.906 | 0.916 | 0.928 | 0.936 | 0.943 | 0.948
| 5 | 0.896 | 0.918 | 0.93 | 0.943 | 0.953 | 0.962 | 0.968
| 7 | 0.908 | 0.933 | 0.947 | 0.964 | 0.977 | 0.989 | 0.997
| 10 | 0.919 | 0.946 | 0.963 | 0.983 | 1 | 1.016 | 1.027
| 15 | 0.928 | 0.958 | 0.977 | 1.001 | 1.022 | 1.044 | 1.06
| 20 | 0.933 | 0.965 | 0.985 | 1.011 | 1.036 | 1.061 | 1.08
| 30 | 0.94 | 0.973 | 0.995 | 1.025 | 1.054 | 1.085 | 1.109
| 40 | 0.945 | 0.979 | 1.001 | 1.033 | 1.064 | 1.099 | 1.125

## Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 24 | 31.1 | 38.2 | 49
0.1 | 27.3 | 34.7 | 41.925 | 52.55
0.2 | 33.05 | 40.325 | 46.925 | 57.8
0.3 | 41.7333 | 48.775 | 54.8 | 64.8667
0.4 | 50.1 | 56.24 | 62 | 70.85
0.5 | 57.25 | 62.4 | 68.2 | 75.8
0.6 | 63.4 | 67.9 | 73.125 | 80.3
0.7 | 68.45 | 72.85 | 77.4 | 83.8
0.8 | 72.82 | 76.7 | 81.0333 | 86.6
0.9 | 76.875 | 80.4667 | 83.95 | 89.3
1 | 80.3 | 83.325 | 86.65 | 91.275
1.1 | 83.275 | 85.95 | 89 | 93.05
1.2 | 85.8 | 88.275 | 90.9333 | 94.55
1.3 | 88 | 90.175 | 92.7 | 95.725
1.4 | 89.9 | 91.675 | 94 | 96.85
1.5 | 91.7 | 93.4 | 95.2 | 97.5
1.6 | 93.2 | 94.55 | 96.2 | 98.2
1.7 | 94.45 | 95.8 | 97.1 | 98.8
1.8 | 95.375 | 96.7 | 97.9 | 99.3667
1.9 | 96.76 | 97.3 | 98.35 | 99.55
2 | 97.35 | 98 | 98.8667 | 99.8
2.1 | 97.85 | 98.5 | 99.25 | 99.9
2.2 | 98.675 | 98.95 | 99.5 | 100
2.3 | 99.025 | 99.2 | 99.7 | 100
2.4 | 99.4333 | 99.525 | 99.82 | 100
2.5 | 99.6 | 99.7 | 99.9 | 100
2.6 | 99.8 | 99.8 | 100 | 99.88
2.7 | 99.85 | 99.9 | 100 | 99.8
2.8 | 99.9333 | 100 | 99.9 | 99.6
2.9 | 100 | 100 | 99.8 | 99.4
3 | 99.9 | 99.9 | 99.7 | 99.2
3.1 | 99.9 | 99.825 | 99.525 | 99
3.2 | 99.8 | 99.775 | 99.375 | 98.65
3.3 | 99.675 | 99.625 | 99.15 | 98.42
3.4 | 99.5 | 99.4 | 99 | 98.175
3.5 | 99.3 | 99.2 | 98.825 | 97.925
3.6 | 99.15 | 98.94 | 98.5333 | 97.675
3.7 | 98.9 | 98.68 | 98.3 | 97.25
3.8 | 98.5 | 98.45 | 98 | 96.975
3.9 | 98.15 | 98.1 | 97.65 | 96.7
4 | 97.9 | 97.8 | 97.45 | 96.4
4.1 | 97.7 | 97.45 | 97.1333 | 96.06
4.2 | 97.35 | 97.22 | 96.75 | 95.7
4.3 | 97 | 96.9333 | 96.35 | 95.3
4.4 | 96.55 | 96.55 | 96.125 | 95
4.5 | 96.05 | 96.2333 | 95.75 | 94.7
4.6 | 95.65 | 95.8333 | 95.4667 | 94.4
4.7 | 95.3 | 95.6333 | 95.0667 | 94.0333
4.8 | 95 | 95.3 | 94.75 | 93.7333
4.9 | 94.65 | 94.85 | 94.5 | 93.425
5 | 94.3 | 94.5 | 94.0667 | 93.05
5.1 | 93.6667 | 94.1333 | 93.7 | 92.54
5.2 | 93.32 | 93.775 | 93.2667 | 92.2
5.3 | 92.95 | 93.425 | 92.98 | 91.9
5.4 | 92.5 | 93.1 | 92.65 | 91.625
5.5 | 92.12 | 92.825 | 92.3 | 91.275
5.6 | 91.75 | 92.5 | 91.9333 | 91
5.7 | 91.3 | 92 | 91.6 | 90.6333
5.8 | 91.1 | 91.7 | 91.375 | 90.3667
5.9 | 90.84 | 91.3 | 91.025 | 89.9
6 | 90.4 | 90.9333 | 90.55 | 89.55
6.1 | 89.78 | 90.525 | 90.225 | 89.3
6.2 | 89.5 | 90.1 | 89.9667 | 88.875
6.3 | 89.075 | 89.75 | 89.6 | 88.6
6.4 | 88.725 | 89.35 | 89.25 | 88.2
6.5 | 88.275 | 89.12 | 88.8 | 87.95
6.6 | 87.9 | 88.75 | 88.5 | 87.625
6.7 | 87.4667 | 88.375 | 88.225 | 87.275
6.8 | 86.9 | 88.025 | 87.86 | 86.95
6.9 | 86.6 | 87.8 | 87.55 | 86.75
7 | 86.05 | 87.5 | 87.1 | 86.3
7.1 | 85.675 | 87 | 86.8667 | 86
7.2 | 85.45 | 86.6 | 86.6 | 85.6667
7.3 | 84.9 | 86.3 | 86.15 | 85.325
7.4 | 84.7 | 85.875 | 85.775 | 85.0333
7.5 | 84.3 | 85.5 | 85.55 | 84.7
7.6 | 84.075 | 85.18 | 85.2 | 84.425
7.7 | 83.6 | 84.9 | 84.95 | 84.16
7.8 | 83.125 | 84.5 | 84.7 | 83.75
7.9 | 82.625 | 84.225 | 84.28 | 83.4
8 | 82.3 | 83.8667 | 83.825 | 83.2
8.1 | 81.9 | 83.55 | 83.55 | 82.75
8.2 | 81.65 | 83.2 | 83.18 | 82.45
8.3 | 81.3 | 82.9 | 82.72 | 82.125
8.4 | 80.9 | 82.6 | 82.5 | 81.875
8.5 | 80.4333 | 82.1 | 82.2333 | 81.48
8.6 | 80 | 81.8 | 81.925 | 81.1667
8.7 | 79.65 | 81.52 | 81.6 | 80.85
8.8 | 79.3 | 81.225 | 81.25 | 80.62
8.9 | 79 | 80.8 | 80.925 | 80.3
9 | 78.65 | 80.5 | 80.6 | 80
9.1 | 78.1 | 80.05 | 80.4 | 79.775
9.2 | 77.6667 | 79.625 | 80.075 | 79.35
9.3 | 77.4 | 79.35 | 79.725 | 79.1
9.4 | 77.1333 | 79.15 | 79.275 | 78.85
9.5 | 76.6667 | 78.74 | 79.05 | 78.4
9.6 | 76.3 | 78.5 | 78.8 | 78.075
9.7 | 76.05 | 78.0333 | 78.6333 | 77.725
9.8 | 75.7333 | 77.6 | 78.1333 | 77.4
9.9 | 75.3667 | 77.35 | 77.7333 | 77.075
10 | 74.975 | 77.0667 | 77.425 | 76.8
10.1 | 74.74 | 76.7333 | 77.1333 | 76.6
10.2 | 74.3 | 76.38 | 76.9 | 76.4
10.3 | 73.95 | 76.1 | 76.55 | 76.1
10.4 | 73.65 | 75.86 | 76.28 | 75.65
10.5 | 73.35 | 75.55 | 76.05 | 75.425
10.6 | 73.1 | 75.2 | 75.8 | 75.15
10.7 | 72.6667 | 74.9 | 75.52 | 74.925
10.8 | 72.225 | 74.5333 | 75.125 | 74.475
10.9 | 71.9 | 74.2333 | 74.75 | 74.24
11 | 71.55 | 73.875 | 74.475 | 73.95
11.1 | 71.2 | 73.64 | 74.175 | 73.7333
11.2 | 70.9667 | 73.3 | 73.85 | 73.5
11.3 | 70.6 | 73.025 | 73.56 | 73.2
11.4 | 70.275 | 72.7 | 73.25 | 72.8
11.5 | 70.04 | 72.45 | 72.95 | 72.575
11.6 | 69.7 | 72.1 | 72.75 | 72.325
11.7 | 69.4333 | 71.8 | 72.325 | 72.0667
11.8 | 69.1 | 71.45 | 72.025 | 71.8
11.9 | 68.66 | 71.075 | 71.825 | 71.4667
12 | 68.25 | 70.825 | 71.6 | 71.1
12.1 | 68 | 70.575 | 71.25 | 70.9
12.2 | 67.7333 | 70.35 | 70.9667 | 70.7
12.3 | 67.3 | 69.9 | 70.675 | 70.325
12.4 | 67.1 | 69.65 | 70.4 | 69.975
12.5 | 66.8333 | 69.4 | 70.075 | 69.8
12.6 | 66.5667 | 69.05 | 69.825 | 69.6333
12.7 | 66.2 | 68.8 | 69.575 | 69.3
12.8 | 65.875 | 68.5333 | 69.25 | 69.1
12.9 | 65.625 | 68.2 | 68.9333 | 68.6667
13 | 65.28 | 67.8 | 68.66 | 68.4333
13.1 | 65.05 | 67.6 | 68.4 | 68.1
13.2 | 64.7667 | 67.4 | 68.1 | 67.8
13.3 | 64.3667 | 67.1 | 67.7 | 67.6
13.4 | 64 | 66.875 | 67.5333 | 67.3
13.5 | 63.72 | 66.5 | 67.35 | 67.15
13.6 | 63.4 | 66.2 | 67.05 | 66.8
13.7 | 63.225 | 65.925 | 66.8333 | 66.6333
13.8 | 63.04 | 65.65 | 66.6 | 66.3667
13.9 | 62.7333 | 65.225 | 66.3 | 66.175
14 | 62.4 | 65 | 65.975 | 65.75
14.1 | 62.1 | 64.8 | 65.6667 | 65.475
14.2 | 61.675 | 64.5 | 65.5 | 65.3333
14.3 | 61.4 | 64.325 | 65.2667 | 65
14.4 | 61.3 | 63.95 | 65.0333 | 64.75
14.5 | 60.7667 | 63.625 | 64.725 | 64.5
14.6 | 60.425 | 63.28 | 64.48 | 64.3
14.7 | 60.1667 | 63.1 | 64.3 | 63.9333
14.8 | 60 | 62.9 | 63.975 | 63.775
14.9 | 59.7 | 62.675 | 63.65 | 63.58
15 | 59.5 | 62.325 | 63.375 | 63.35
15.1 | 59.3 | 62.075 | 63.05 | 63
15.2 | 58.95 | 61.825 | 62.78 | 62.75
15.3 | 58.75 | 61.5 | 62.55 | 62.52
15.4 | 58.4 | 61.2 | 62.4 | 62.2333
15.5 | 58.16 | 60.9 | 62.15 | 62.02
15.6 | 57.875 | 60.7 | 62 | 61.7
15.7 | 57.575 | 60.525 | 61.7333 | 61.525
15.8 | 57.2 | 60.1 | 61.25 | 61.3
15.9 | 56.9 | 59.9 | 61.075 | 60.95
16 | 56.7 | 59.7 | 60.85 | 60.8
16.1 | 56.45 | 59.5 | 60.7 | 60.55
16.2 | 56.275 | 59.3 | 60.3667 | 60.4
16.3 | 55.925 | 59 | 60.1 | 60.225
16.4 | 55.675 | 58.7 | 59.9 | 59.85
16.5 | 55.5 | 58.4667 | 59.7 | 59.525
16.6 | 55.28 | 58.225 | 59.38 | 59.4
16.7 | 55.05 | 57.875 | 59.15 | 59.18
16.8 | 54.65 | 57.7 | 58.875 | 58.9
16.9 | 54.3 | 57.4667 | 58.65 | 58.7
17 | 54.16 | 57.175 | 58.4 | 58.45
17.1 | 53.925 | 57 | 58.1667 | 58.2
17.2 | 53.65 | 56.7 | 57.9 | 58
17.3 | 53.375 | 56.5 | 57.6667 | 57.78
17.4 | 53.14 | 56.3 | 57.425 | 57.5
17.5 | 52.8667 | 56.15 | 57.175 | 57.2
17.6 | 52.6333 | 55.875 | 57 | 57.075
17.7 | 52.4 | 55.625 | 56.775 | 56.825
17.8 | 52.2 | 55.375 | 56.5 | 56.6
17.9 | 52 | 55.125 | 56.3 | 56.4
18 | 51.8 | 54.9 | 56 | 56.2
18.1 | 51.525 | 54.65 | 55.875 | 56.025
18.2 | 51.3 | 54.4 | 55.5667 | 55.775
18.3 | 51 | 54.24 | 55.4 | 55.5
18.4 | 50.725 | 54 | 55.2 | 55.3
18.5 | 50.5333 | 53.8 | 55 | 55.1
18.6 | 50.4 | 53.62 | 54.775 | 54.86
18.7 | 50.1 | 53.25 | 54.5 | 54.65
18.8 | 49.85 | 53 | 54.3 | 54.525
18.9 | 49.62 | 52.775 | 54.075 | 54.2
19 | 49.375 | 52.5 | 53.8 | 53.98
19.1 | 49.125 | 52.2 | 53.6333 | 53.7333
19.2 | 48.9 | 52.05 | 53.4 | 53.4
19.3 | 48.75 | 51.82 | 53.25 | 53.3
19.4 | 48.375 | 51.575 | 53 | 53.1333
19.5 | 48.175 | 51.4 | 52.8 | 52.9667
19.6 | 48.0333 | 51.1667 | 52.6 | 52.725
19.7 | 47.7 | 50.9667 | 52.35 | 52.475
19.8 | 47.575 | 50.7 | 52.1333 | 52.24
19.9 | 47.4 | 50.4667 | 51.875 | 52.025
20 | 47.2667 | 50.3 | 51.625 | 51.8
20.1 | 47.1 | 50.025 | 51.475 | 51.65
20.2 | 46.85 | 49.9 | 51.225 | 51.5
20.3 | 46.675 | 49.675 | 51.1 | 51.25
20.4 | 46.45 | 49.575 | 50.9 | 51.0667
20.5 | 46.14 | 49.325 | 50.7 | 50.9
20.6 | 45.875 | 49.2 | 50.5 | 50.7
20.7 | 45.6 | 48.85 | 50.3 | 50.5
20.8 | 45.4333 | 48.65 | 50.075 | 50.3
20.9 | 45.28 | 48.4 | 49.85 | 50.06
21 | 45.0667 | 48.2 | 49.675 | 49.925
21.1 | 44.9 | 48 | 49.375 | 49.7
21.2 | 44.75 | 47.75 | 49.2667 | 49.525
21.3 | 44.5 | 47.525 | 49 | 49.3
21.4 | 44.4 | 47.375 | 48.775 | 49.2
21.5 | 44.1 | 47.24 | 48.625 | 48.9667
21.6 | 43.9 | 47.1 | 48.4 | 48.7
21.7 | 43.7 | 46.85 | 48.3 | 48.575
21.8 | 43.5 | 46.7 | 48.1 | 48.325
21.9 | 43.3 | 46.5 | 47.85 | 48.06
22 | 43.1 | 46.275 | 47.7 | 47.925
22.1 | 42.9 | 46 | 47.4 | 47.775
22.2 | 42.775 | 45.82 | 47.24 | 47.65
22.3 | 42.5 | 45.5667 | 47.0667 | 47.4
22.4 | 42.275 | 45.35 | 46.8 | 47.2
22.5 | 42.1 | 45.2 | 46.6 | 47
22.6 | 41.9 | 45.05 | 46.5 | 46.825
22.7 | 41.75 | 44.825 | 46.35 | 46.6
22.8 | 41.6 | 44.5667 | 46.1 | 46.4
22.9 | 41.4 | 44.4 | 45.95 | 46.3
23 | 41.25 | 44.3 | 45.7 | 46.125
23.1 | 41.1 | 44.1 | 45.5667 | 45.875
23.2 | 40.8 | 43.925 | 45.3 | 45.7
23.3 | 40.7 | 43.8 | 45.1 | 45.5333
23.4 | 40.55 | 43.65 | 44.9 | 45.375
23.5 | 40.275 | 43.4 | 44.7667 | 45.2
23.6 | 40.125 | 43.1 | 44.6 | 44.98
23.7 | 39.9 | 42.975 | 44.4333 | 44.8667
23.8 | 39.7 | 42.825 | 44.3 | 44.6
23.9 | 39.5 | 42.575 | 44.1 | 44.45
24 | 39.3333 | 42.4 | 43.85 | 44.25
24.1 | 39.175 | 42.2 | 43.7 | 44.1
24.2 | 39 | 42.025 | 43.575 | 43.9
24.3 | 38.8 | 41.8667 | 43.4 | 43.8
24.4 | 38.6 | 41.6 | 43.3 | 43.625
24.5 | 38.4 | 41.5 | 43.1 | 43.3
24.6 | 38.3 | 41.3333 | 42.8333 | 43.2
24.7 | 38.15 | 41.2 | 42.625 | 42.9
24.8 | 38.025 | 41 | 42.475 | 42.8
24.9 | 37.875 | 40.8 | 42.325 | 42.625
25 | 37.7 | 40.64 | 42.2 | 42.56
25.1 | 37.575 | 40.4333 | 42 | 42.4
25.2 | 37.425 | 40.35 | 41.8 | 42.3
25.3 | 37.2 | 40.1667 | 41.65 | 42.1
25.4 | 36.95 | 40 | 41.5 | 41.85
25.5 | 36.8 | 39.8 | 41.3 | 41.7
25.6 | 36.6333 | 39.6333 | 41.125 | 41.55
25.7 | 36.4 | 39.4333 | 40.9667 | 41.42
25.8 | 36.225 | 39.2 | 40.8 | 41.2
25.9 | 36.175 | 39.1 | 40.6 | 41.0333
26 | 36.025 | 38.9333 | 40.45 | 40.9
26.1 | 35.875 | 38.8 | 40.3 | 40.65
26.2 | 35.64 | 38.6 | 40.1 | 40.5
26.3 | 35.6 | 38.5333 | 39.95 | 40.4
26.4 | 35.4 | 38.3 | 39.8 | 40.25
26.5 | 35.2 | 38.15 | 39.625 | 40.05
26.6 | 34.975 | 38 | 39.48 | 39.9
26.7 | 34.8 | 37.8 | 39.3 | 39.8
26.8 | 34.7 | 37.625 | 39.1 | 39.6
26.9 | 34.6 | 37.475 | 39 | 39.45
27 | 34.4 | 37.3 | 38.825 | 39.3
27.1 | 34.3 | 37.2 | 38.7 | 39.1
27.2 | 34.1 | 37.1 | 38.5 | 39
27.3 | 33.9333 | 36.85 | 38.3 | 38.8333
27.4 | 33.8 | 36.7 | 38.2 | 38.675
27.5 | 33.6667 | 36.575 | 38.1 | 38.4667
27.6 | 33.4 | 36.4 | 37.9 | 38.3
27.7 | 33.3 | 36.3 | 37.75 | 38.2
27.8 | 33.2 | 36.15 | 37.6333 | 38.1
27.9 | 33.1 | 36 | 37.5 | 37.9
28 | 32.9 | 35.85 | 37.275 | 37.8
28.1 | 32.72 | 35.7 | 37.125 | 37.6
28.2 | 32.575 | 35.475 | 37 | 37.425
28.3 | 32.425 | 35.4 | 36.825 | 37.2667
28.4 | 32.3 | 35.175 | 36.7 | 37.1
28.5 | 32.15 | 35 | 36.5 | 37
28.6 | 32 | 34.9 | 36.4 | 36.9
28.7 | 31.8333 | 34.7 | 36.3 | 36.7
28.8 | 31.8 | 34.6 | 36.1 | 36.6
28.9 | 31.575 | 34.45 | 36 | 36.4
29 | 31.4 | 34.3667 | 35.8667 | 36.3
29.1 | 31.2 | 34.2 | 35.675 | 36.1
29.2 | 31 | 34.1 | 35.5 | 36
29.3 | 30.9 | 33.9 | 35.45 | 35.825
29.4 | 30.9 | 33.78 | 35.275 | 35.7
29.5 | 30.8 | 33.6 | 35.1 | 35.6
29.6 | 30.6 | 33.4667 | 34.975 | 35.5
29.7 | 30.5 | 33.32 | 34.8 | 35.325
29.8 | 30.275 | 33.1 | 34.7 | 35.2
29.9 | 30.1667 | 33 | 34.5333 | 35.1
30 | 30.075 | 32.875 | 34.4 | 34.9

## SSD80 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 25.1 | 35.2 | 45.4 | 60.3
0.1 | 29 | 38.75 | 49.3333 | 63.9
0.2 | 35.55 | 44.75 | 54.7333 | 69.2
0.3 | 44.3 | 53.15 | 62.1 | 75.2
0.4 | 52.025 | 60.5 | 68.8 | 80.3
0.5 | 59.3333 | 67 | 74.4875 | 84.8
0.6 | 65.225 | 72.3143 | 79.25 | 88.15
0.7 | 70.5 | 76.825 | 83.05 | 90.7333
0.8 | 75.075 | 80.6 | 86.3625 | 93.1
0.9 | 78.9 | 83.8 | 88.9667 | 94.78
1 | 82.375 | 86.6 | 91.2 | 96.2
1.1 | 85.12 | 89 | 93 | 97.35
1.2 | 87.4 | 90.95 | 94.5 | 98.2167
1.3 | 89.7 | 92.7 | 95.8 | 98.85
1.4 | 91.55 | 94.15 | 96.775 | 99.24
1.5 | 93.1667 | 95.4 | 97.65 | 99.65
1.6 | 94.5 | 96.4 | 98.375 | 99.92
1.7 | 95.75 | 97.3 | 98.95 | 100
1.8 | 96.8 | 98.05 | 99.425 | 100
1.9 | 97.45 | 98.6 | 99.75 | 100
2 | 98.2 | 99.05 | 99.9 | 100
2.1 | 98.66 | 99.525 | 99.95 | 99.85
2.2 | 99.15 | 99.7 | 100 | 99.6
2.3 | 99.5 | 99.82 | 100 | 99.42
2.4 | 99.725 | 99.975 | 99.925 | 99.2
2.5 | 99.9 | 100 | 99.9 | 98.875
2.6 | 100 | 100 | 99.8 | 98.625
2.7 | 100 | 99.9 | 99.6 | 98.25
2.8 | 100 | 99.9 | 99.3 | 97.9333
2.9 | 99.9 | 99.7667 | 98.9833 | 97.575
3 | 99.8 | 99.55 | 98.775 | 97.15
3.1 | 99.55 | 99.4 | 98.45 | 96.85
3.2 | 99.4 | 99.1 | 98.24 | 96.375
3.3 | 99.15 | 98.775 | 97.94 | 96.0667
3.4 | 98.8 | 98.55 | 97.6 | 95.7
3.5 | 98.55 | 98.2625 | 97.175 | 95.2667
3.6 | 98.3 | 97.95 | 96.8 | 94.85
3.7 | 97.95 | 97.6 | 96.45 | 94.4667
3.8 | 97.5 | 97.2143 | 96.12 | 94.1333
3.9 | 97.175 | 96.9 | 95.7571 | 93.7
4 | 96.85 | 96.475 | 95.4 | 93.325
4.1 | 96.475 | 96.1 | 95 | 92.9
4.2 | 95.95 | 95.775 | 94.5444 | 92.5667
4.3 | 95.675 | 95.44 | 94.08 | 92.125
4.4 | 95.35 | 94.9714 | 93.725 | 91.7
4.5 | 94.88 | 94.6 | 93.325 | 91.36
4.6 | 94.4667 | 94.15 | 92.9 | 90.925
4.7 | 93.9429 | 93.7333 | 92.5333 | 90.55
4.8 | 93.375 | 93.4 | 92.125 | 90.15
4.9 | 92.95 | 92.8667 | 91.675 | 89.775
5 | 92.5 | 92.5333 | 91.325 | 89.35
5.1 | 92 | 92.0625 | 90.86 | 89
5.2 | 91.6 | 91.6875 | 90.425 | 88.4875
5.3 | 91.25 | 91.25 | 90.1 | 88.1
5.4 | 90.65 | 90.75 | 89.6667 | 87.8
5.5 | 90.275 | 90.3556 | 89.34 | 87.3333
5.6 | 89.875 | 89.925 | 88.95 | 86.925
5.7 | 89.46 | 89.55 | 88.45 | 86.5667
5.8 | 89 | 89.175 | 88.075 | 86.2125
5.9 | 88.3 | 88.775 | 87.58 | 85.7833
6 | 87.9333 | 88.42 | 87.2 | 85.5125
6.1 | 87.6 | 88.025 | 87 | 85.1375
6.2 | 87.15 | 87.56 | 86.5 | 84.8
6.3 | 86.9 | 87.1667 | 86.16 | 84.425
6.4 | 86.4 | 86.75 | 85.775 | 84.05
6.5 | 85.9667 | 86.35 | 85.375 | 83.675
6.6 | 85.3667 | 85.9 | 84.95 | 83.3
6.7 | 85.04 | 85.55 | 84.65 | 82.9625
6.8 | 84.575 | 85.2 | 84.35 | 82.5875
6.9 | 84.3333 | 84.75 | 83.9 | 82.2125
7 | 83.85 | 84.325 | 83.45 | 81.8
7.1 | 83.275 | 83.875 | 83.1 | 81.4625
7.2 | 82.925 | 83.625 | 82.7667 | 81.1
7.3 | 82.38 | 83.15 | 82.34 | 80.725
7.4 | 82.1 | 82.75 | 81.975 | 80.4
7.5 | 81.6 | 82.4 | 81.65 | 80.05
7.6 | 81.15 | 82.05 | 81.3 | 79.6667
7.7 | 80.82 | 81.7 | 80.9 | 79.275
7.8 | 80.45 | 81.3 | 80.5 | 78.95
7.9 | 80.05 | 80.9333 | 80.175 | 78.575
8 | 79.55 | 80.5875 | 79.825 | 78.3
8.1 | 79.14 | 80.15 | 79.55 | 77.9
8.2 | 78.7 | 79.7333 | 79.14 | 77.7
8.3 | 78.35 | 79.375 | 78.7 | 77.25
8.4 | 78.1 | 78.98 | 78.35 | 77
8.5 | 77.55 | 78.6 | 78 | 76.625
8.6 | 77.22 | 78.2 | 77.625 | 76.225
8.7 | 76.82 | 77.9 | 77.32 | 75.85
8.8 | 76.45 | 77.55 | 77 | 75.525
8.9 | 76.0333 | 77.1 | 76.65 | 75.3
9 | 75.76 | 76.7 | 76.3 | 74.95
9.1 | 75.3 | 76.4 | 75.95 | 74.5
9.2 | 74.8 | 75.96 | 75.6 | 74.25
9.3 | 74.35 | 75.7 | 75.325 | 73.8
9.4 | 74.1 | 75.35 | 74.9625 | 73.55
9.5 | 73.76 | 74.94 | 74.5667 | 73.2
9.6 | 73.525 | 74.625 | 74.2 | 72.875
9.7 | 72.95 | 74.225 | 73.95 | 72.5333
9.8 | 72.625 | 73.875 | 73.5 | 72.225
9.9 | 72.15 | 73.625 | 73.1667 | 71.9
10 | 71.75 | 73.175 | 72.86 | 71.6
10.1 | 71.5333 | 72.825 | 72.475 | 71.25
10.2 | 71.1 | 72.55 | 72.2 | 70.94
10.3 | 70.65 | 72.125 | 71.8429 | 70.65
10.4 | 70.275 | 71.9 | 71.5333 | 70.3
10.5 | 69.9 | 71.55 | 71.175 | 70
10.6 | 69.6571 | 71.22 | 70.85 | 69.6
10.7 | 69.2667 | 70.8 | 70.475 | 69.35
10.8 | 68.9 | 70.4333 | 70.15 | 69.025
10.9 | 68.625 | 69.975 | 69.9 | 68.78
11 | 68.2 | 69.7 | 69.55 | 68.4333
11.1 | 67.925 | 69.4 | 69.2 | 68.1
11.2 | 67.5 | 69.05 | 68.95 | 67.85
11.3 | 67.1 | 68.7 | 68.65 | 67.5
11.4 | 66.7 | 68.325 | 68.3 | 67.26
11.5 | 66.475 | 68.1 | 68.025 | 66.925
11.6 | 66.15 | 67.85 | 67.675 | 66.675
11.7 | 65.775 | 67.525 | 67.425 | 66.325
11.8 | 65.45 | 67.18 | 67.075 | 66.1
11.9 | 65.25 | 66.9 | 66.725 | 65.725
12 | 65 | 66.5 | 66.5 | 65.425
12.1 | 64.66 | 66.16 | 66.15 | 65.1333
12.2 | 64.225 | 65.86 | 65.8 | 64.8333
12.3 | 63.875 | 65.525 | 65.5667 | 64.6
12.4 | 63.625 | 65.35 | 65.22 | 64.3333
12.5 | 63.275 | 65.025 | 64.9 | 64
12.6 | 62.95 | 64.6143 | 64.7 | 63.65
12.7 | 62.6 | 64.325 | 64.4 | 63.4
12.8 | 62.325 | 64 | 64.0667 | 63.175
12.9 | 62.1 | 63.72 | 63.65 | 62.9
13 | 61.725 | 63.4 | 63.45 | 62.575
13.1 | 61.35 | 63.15 | 63.1333 | 62.325
13.2 | 61.05 | 62.775 | 62.9 | 62
13.3 | 60.7 | 62.525 | 62.675 | 61.76
13.4 | 60.4125 | 62.2 | 62.35 | 61.425
13.5 | 60.125 | 61.95 | 62.0667 | 61.175
13.6 | 59.8625 | 61.625 | 61.8 | 60.925
13.7 | 59.525 | 61.3 | 61.4333 | 60.68
13.8 | 59.2 | 61.05 | 61.275 | 60.4
13.9 | 58.95 | 60.8 | 60.95 | 60.1
14 | 58.7 | 60.45 | 60.6667 | 59.85
14.1 | 58.25 | 60.0857 | 60.4 | 59.6
14.2 | 58 | 59.9 | 60.05 | 59.3
14.3 | 57.675 | 59.5333 | 59.8 | 59.04
14.4 | 57.525 | 59.325 | 59.52 | 58.775
14.5 | 57.2125 | 59 | 59.225 | 58.525
14.6 | 56.875 | 58.65 | 59.0167 | 58.275
14.7 | 56.625 | 58.4429 | 58.82 | 58.02
14.8 | 56.275 | 58.125 | 58.5 | 57.775
14.9 | 55.95 | 57.875 | 58.15 | 57.525
15 | 55.675 | 57.55 | 57.9 | 57.275
15.1 | 55.35 | 57.24 | 57.7 | 56.925
15.2 | 55.05 | 56.975 | 57.35 | 56.675
15.3 | 54.8 | 56.74 | 57.1 | 56.425
15.4 | 54.5667 | 56.4 | 56.82 | 56.2
15.5 | 54.24 | 56.15 | 56.525 | 55.95
15.6 | 54.075 | 56 | 56.275 | 55.7
15.7 | 53.825 | 55.65 | 56 | 55.5
15.8 | 53.475 | 55.5 | 55.875 | 55.2
15.9 | 53.225 | 55.26 | 55.525 | 54.95
16 | 53 | 54.95 | 55.3 | 54.7
16.1 | 52.76 | 54.675 | 55.125 | 54.475
16.2 | 52.5333 | 54.425 | 54.85 | 54.325
16.3 | 52.2 | 54.1 | 54.6 | 53.975
16.4 | 51.9667 | 53.94 | 54.375 | 53.725
16.5 | 51.76 | 53.7 | 54.05 | 53.475
16.6 | 51.4333 | 53.4 | 53.8 | 53.225
16.7 | 51.2143 | 53.1 | 53.6 | 53
16.8 | 50.92 | 52.9 | 53.3 | 52.75
16.9 | 50.625 | 52.6 | 53.025 | 52.6
17 | 50.3778 | 52.4 | 52.86 | 52.35
17.1 | 50.15 | 52.2 | 52.6 | 52.02
17.2 | 49.9 | 51.85 | 52.2 | 51.7875
17.3 | 49.56 | 51.7 | 52 | 51.625
17.4 | 49.425 | 51.36 | 51.85 | 51.375
17.5 | 49.15 | 51.1 | 51.6667 | 51.2125
17.6 | 48.8 | 50.875 | 51.4667 | 50.9778
17.7 | 48.675 | 50.6 | 51.2333 | 50.7714
17.8 | 48.425 | 50.375 | 51 | 50.5333
17.9 | 48.075 | 50.225 | 50.8 | 50.3
18 | 47.84 | 50 | 50.5 | 50.1333
18.1 | 47.7 | 49.7333 | 50.2333 | 49.85
18.2 | 47.45 | 49.5 | 50.1 | 49.62
18.3 | 47.2 | 49.275 | 49.8 | 49.4286
18.4 | 46.95 | 49.1 | 49.6 | 49.2
18.5 | 46.7 | 48.775 | 49.3 | 48.975
18.6 | 46.575 | 48.525 | 49.1 | 48.725
18.7 | 46.325 | 48.3 | 48.875 | 48.5
18.8 | 46 | 48.1 | 48.7 | 48.3
18.9 | 45.7778 | 47.9 | 48.475 | 48.05
19 | 45.55 | 47.65 | 48.225 | 47.9
19.1 | 45.4 | 47.4 | 48.1 | 47.625
19.2 | 45.15 | 47.275 | 47.76 | 47.475
19.3 | 44.9 | 47.05 | 47.6125 | 47.225
19.4 | 44.675 | 46.7 | 47.4 | 47.025
19.5 | 44.4625 | 46.525 | 47.175 | 46.775
19.6 | 44.275 | 46.3 | 46.94 | 46.6
19.7 | 44.06 | 46.05 | 46.7 | 46.525
19.8 | 43.8 | 45.8 | 46.525 | 46.25
19.9 | 43.575 | 45.64 | 46.275 | 46
20 | 43.375 | 45.35 | 46.05 | 45.8
20.1 | 43.14 | 45.2 | 45.88 | 45.6
20.2 | 43 | 45 | 45.65 | 45.4
20.3 | 42.75 | 44.75 | 45.4333 | 45.2
20.4 | 42.5 | 44.55 | 45.2333 | 45.0667
20.5 | 42.35 | 44.4143 | 45.05 | 44.8333
20.6 | 42.12 | 44.1 | 44.8 | 44.5667
20.7 | 41.95 | 44 | 44.55 | 44.4
20.8 | 41.725 | 43.675 | 44.4 | 44.16
20.9 | 41.6 | 43.425 | 44.2 | 43.94
21 | 41.3 | 43.3 | 44 | 43.8
21.1 | 41.0667 | 43.125 | 43.8 | 43.625
21.2 | 40.95 | 42.9 | 43.625 | 43.45
21.3 | 40.7333 | 42.65 | 43.48 | 43.26
21.4 | 40.5 | 42.525 | 43.2429 | 43.04
21.5 | 40.32 | 42.38 | 43.0375 | 42.9
21.6 | 40.125 | 42.1 | 42.8444 | 42.6625
21.7 | 39.975 | 41.925 | 42.7 | 42.5375
21.8 | 39.7 | 41.675 | 42.44 | 42.325
21.9 | 39.575 | 41.525 | 42.3 | 42.1
22 | 39.425 | 41.375 | 42.125 | 41.9333
22.1 | 39.1 | 41.1 | 41.8 | 41.6667
22.2 | 38.9667 | 40.9 | 41.6333 | 41.5857
22.3 | 38.7333 | 40.75 | 41.5 | 41.45
22.4 | 38.5 | 40.5 | 41.2778 | 41.22
22.5 | 38.4 | 40.35 | 41.1625 | 41
22.6 | 38.175 | 40.125 | 40.975 | 40.725
22.7 | 38 | 39.975 | 40.74 | 40.6
22.8 | 37.8 | 39.75 | 40.6 | 40.425
22.9 | 37.55 | 39.6167 | 40.35 | 40.3
23 | 37.4 | 39.375 | 40.2 | 40.0667
23.1 | 37.25 | 39.14 | 40 | 39.9
23.2 | 37.05 | 39 | 39.8 | 39.75
23.3 | 36.85 | 38.85 | 39.7 | 39.6
23.4 | 36.7 | 38.65 | 39.475 | 39.475
23.5 | 36.55 | 38.525 | 39.3 | 39.225
23.6 | 36.3556 | 38.4 | 39.175 | 39.1
23.7 | 36.2375 | 38.175 | 39.025 | 38.9
23.8 | 36 | 38 | 38.825 | 38.75
23.9 | 35.8714 | 37.8333 | 38.6 | 38.55
24 | 35.7 | 37.6 | 38.4 | 38.4
24.1 | 35.55 | 37.4714 | 38.2 | 38.2667
24.2 | 35.3 | 37.3 | 38.05 | 38.075
24.3 | 35.1 | 37.0833 | 37.9 | 37.9
24.4 | 35 | 36.9167 | 37.7 | 37.725
24.5 | 34.85 | 36.7875 | 37.575 | 37.5
24.6 | 34.7 | 36.65 | 37.425 | 37.3625
24.7 | 34.5 | 36.425 | 37.3 | 37.2375
24.8 | 34.325 | 36.2 | 37 | 37.025
24.9 | 34.2 | 36.05 | 36.9 | 36.95
25 | 34.025 | 35.95 | 36.74 | 36.725
25.1 | 34 | 35.8 | 36.5 | 36.55
25.2 | 33.75 | 35.6 | 36.4 | 36.4
25.3 | 33.5 | 35.4 | 36.225 | 36.2333
25.4 | 33.35 | 35.2333 | 36.0667 | 36.1
25.5 | 33.2 | 35.1 | 35.9571 | 35.9
25.6 | 33 | 34.925 | 35.775 | 35.8
25.7 | 32.9 | 34.775 | 35.54 | 35.6
25.8 | 32.7 | 34.55 | 35.4 | 35.475
25.9 | 32.525 | 34.475 | 35.25 | 35.325
26 | 32.4 | 34.25 | 35.1 | 35.2
26.1 | 32.3 | 34.1 | 34.96 | 35.05
26.2 | 32.1 | 34 | 34.75 | 34.9
26.3 | 31.9625 | 33.8 | 34.675 | 34.75
26.4 | 31.7889 | 33.6833 | 34.5 | 34.575
26.5 | 31.625 | 33.5 | 34.275 | 34.4
26.6 | 31.5 | 33.375 | 34.125 | 34.24
26.7 | 31.3333 | 33.1625 | 34 | 34.1
26.8 | 31.2 | 33.0375 | 33.85 | 33.9375
26.9 | 31.1 | 32.875 | 33.7 | 33.8222
27 | 30.9 | 32.74 | 33.55 | 33.7
27.1 | 30.8 | 32.6 | 33.5 | 33.475
27.2 | 30.6 | 32.425 | 33.3 | 33.35
27.3 | 30.45 | 32.3 | 33.1667 | 33.225
27.4 | 30.3 | 32.15 | 32.975 | 33.05
27.5 | 30.15 | 32 | 32.825 | 33
27.6 | 30 | 31.85 | 32.7167 | 32.75
27.7 | 29.9 | 31.7 | 32.4875 | 32.6
27.8 | 29.7 | 31.55 | 32.3667 | 32.55
27.9 | 29.6 | 31.4 | 32.25 | 32.3
28 | 29.5 | 31.2 | 32.125 | 32.275
28.1 | 29.3 | 31.1 | 32.0111 | 32.1
28.2 | 29.2 | 31 | 31.875 | 31.95
28.3 | 29.0667 | 30.82 | 31.6625 | 31.8
28.4 | 28.9 | 30.675 | 31.5375 | 31.7
28.5 | 28.8 | 30.525 | 31.4125 | 31.525
28.6 | 28.65 | 30.4 | 31.225 | 31.4
28.7 | 28.4556 | 30.3 | 31.175 | 31.275
28.8 | 28.3375 | 30.175 | 31 | 31.125
28.9 | 28.2125 | 29.925 | 30.9 | 30.98
29 | 28.0875 | 29.9 | 30.75 | 30.85
29.1 | 27.9625 | 29.75 | 30.6 | 30.7
29.2 | 27.8375 | 29.6 | 30.45 | 30.56
29.3 | 27.7125 | 29.45 | 30.3 | 30.4
29.4 | 27.5875 | 29.3 | 30.175 | 30.3
29.5 | 27.4625 | 29.2 | 30.025 | 30.2
29.6 | 27.3 | 29.1 | 29.975 | 30.075
29.7 | 27.2 | 28.975 | 29.8 | 30
29.8 | 27.04 | 28.8 | 29.68 | 29.875
29.9 | 26.9333 | 28.675 | 29.55 | 29.7
30 | 26.8 | 28.54 | 29.3667 | 29.575

## SSD120 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 23.7 | 30 | 36.1 | 45.1
0.1 | 27.1 | 33.5 | 39.7125 | 48.6
0.2 | 33.1 | 39.25 | 45.5875 | 54.12
0.3 | 41.55 | 47.7 | 53.15 | 61.3
0.4 | 49.5333 | 55.35 | 59.6778 | 67.6
0.5 | 56.5 | 61.5 | 66.2 | 72.8
0.6 | 62.2 | 67 | 71.02 | 77.075
0.7 | 67.6 | 71.6 | 75.2 | 80.75
0.8 | 71.8 | 75.475 | 78.775 | 84
0.9 | 76.05 | 78.8 | 82.05 | 86.7
1 | 78.98 | 82.25 | 85.025 | 88.98
1.1 | 82.1 | 84.675 | 87.3 | 90.9667
1.2 | 84.9 | 87.125 | 89.45 | 92.7
1.3 | 87.05 | 89.1 | 91.15 | 94
1.4 | 89.3 | 90.875 | 92.4857 | 95.34
1.5 | 91 | 92.625 | 93.9167 | 96.325
1.6 | 92.625 | 93.8 | 95.1 | 97
1.7 | 93.675 | 95.1667 | 96 | 97.825
1.8 | 94.6 | 95.88 | 96.8333 | 98.3
1.9 | 95.7 | 96.8 | 97.475 | 98.875
2 | 96.65 | 97.4 | 98.02 | 99.22
2.1 | 97.4 | 97.95 | 98.5 | 99.4333
2.2 | 98.1 | 98.5 | 98.9 | 99.7
2.3 | 98.6 | 99.2 | 99.35 | 99.8
2.4 | 98.7 | 99.5 | 99.5 | 99.9
2.5 | 99.15 | 99.6 | 99.7 | 100
2.6 | 99.3 | 99.7 | 99.875 | 100
2.7 | 99.65 | 99.8 | 100 | 99.9
2.8 | 99.8 | 99.9 | 100 | 99.9
2.9 | 99.9 | 99.9 | 100 | 99.7
3 | 100 | 100 | 99.9 | 99.6
3.1 | 100 | 100 | 99.9 | 99.425
3.2 | 100 | 99.9 | 99.75 | 99.3
3.3 | 99.925 | 99.8 | 99.6 | 99.2
3.4 | 99.7 | 99.675 | 99.4333 | 99.02
3.5 | 99.5 | 99.5 | 99.2 | 98.725
3.6 | 99.2333 | 99.34 | 98.95 | 98.4833
3.7 | 99.1 | 99.1 | 98.8 | 98.275
3.8 | 98.825 | 98.9 | 98.55 | 98.04
3.9 | 98.6333 | 98.7 | 98.2 | 97.55
4 | 98.35 | 98.4 | 97.95 | 97.3
4.1 | 98.025 | 98.05 | 97.8 | 97.15
4.2 | 97.85 | 97.9 | 97.46 | 96.9
4.3 | 97.45 | 97.65 | 97.1 | 96.5
4.4 | 97.075 | 97.15 | 96.8 | 96.25
4.5 | 96.7667 | 96.92 | 96.5 | 96
4.6 | 96.4 | 96.65 | 96.275 | 95.6
4.7 | 96 | 96.5 | 96 | 95.34
4.8 | 95.5 | 96.24 | 95.7 | 95.1
4.9 | 95.25 | 95.95 | 95.35 | 94.75
5 | 95.1 | 95.64 | 95.075 | 94.52
5.1 | 94.775 | 95.3 | 94.8625 | 94.175
5.2 | 94.375 | 94.95 | 94.4875 | 93.825
5.3 | 93.95 | 94.6 | 94.175 | 93.575
5.4 | 93.54 | 94.32 | 93.925 | 93.275
5.5 | 93.3 | 94 | 93.575 | 92.8
5.6 | 92.8 | 93.72 | 93.2875 | 92.45
5.7 | 92.4 | 93.25 | 92.9571 | 92.2
5.8 | 91.9 | 92.925 | 92.55 | 91.95
5.9 | 91.5 | 92.525 | 92.2857 | 91.72
6 | 91.3 | 92.1 | 91.9667 | 91.36
6.1 | 91.025 | 91.7 | 91.7 | 91.05
6.2 | 90.625 | 91.3 | 91.45 | 90.675
6.3 | 90.25 | 90.9667 | 91.2 | 90.35
6.4 | 89.925 | 90.75 | 90.9286 | 90.1
6.5 | 89.6 | 90.5333 | 90.45 | 89.85
6.6 | 89.15 | 90.2 | 90.225 | 89.6
6.7 | 88.68 | 89.98 | 89.9667 | 89.1429
6.8 | 88.4 | 89.675 | 89.6 | 88.9375
6.9 | 88.1 | 89.3714 | 89.1667 | 88.6667
7 | 87.725 | 89.05 | 88.9 | 88.3
7.1 | 87.28 | 88.6375 | 88.475 | 87.975
7.2 | 86.925 | 88.4 | 88.15 | 87.725
7.3 | 86.475 | 88.15 | 87.8 | 87.475
7.4 | 86.225 | 87.8 | 87.6 | 87.15
7.5 | 85.88 | 87.55 | 87 | 86.85
7.6 | 85.45 | 87.2667 | 86.85 | 86.64
7.7 | 84.9 | 86.825 | 86.7 | 86.3
7.8 | 84.7 | 86.3 | 86.35 | 85.95
7.9 | 84.5 | 86.175 | 86.1 | 85.62
8 | 84.15 | 85.85 | 85.75 | 85.5
8.1 | 83.7 | 85.46 | 85.425 | 85.1667
8.2 | 83.4 | 85.15 | 85.2 | 84.875
8.3 | 83.2 | 84.9 | 85.04 | 84.58
8.4 | 82.7667 | 84.55 | 84.5714 | 84.1
8.5 | 82.3 | 84.2 | 84.2333 | 83.8
8.6 | 82.0667 | 83.9 | 84.025 | 83.65
8.7 | 81.7333 | 83.625 | 83.75 | 83.275
8.8 | 81.4 | 83.25 | 83.5 | 82.925
8.9 | 80.95 | 82.9 | 83 | 82.7
9 | 80.575 | 82.625 | 82.725 | 82.5
9.1 | 80.325 | 82.3 | 82.4 | 82.1
9.2 | 80.04 | 82.05 | 82.15 | 81.85
9.3 | 79.5 | 81.75 | 81.85 | 81.6
9.4 | 79.2 | 81.44 | 81.525 | 81.26
9.5 | 78.95 | 81.2 | 81.275 | 80.925
9.6 | 78.725 | 80.8 | 80.9 | 80.675
9.7 | 78.35 | 80.5 | 80.75 | 80.525
9.8 | 78.025 | 80.2333 | 80.3333 | 80.25
9.9 | 77.675 | 79.975 | 80.12 | 79.85
10 | 77.3375 | 79.8 | 79.875 | 79.7
10.1 | 77 | 79.3 | 79.7 | 79.45
10.2 | 76.725 | 78.85 | 79.35 | 79.15
10.3 | 76.4 | 78.7 | 78.95 | 78.85
10.4 | 76 | 78.35 | 78.5 | 78.5
10.5 | 75.4 | 78.1 | 78.35 | 78.1556
10.6 | 75.1429 | 77.76 | 78.2 | 77.88
10.7 | 74.925 | 77.4286 | 77.8714 | 77.675
10.8 | 74.575 | 77.2 | 77.62 | 77.3875
10.9 | 74.2875 | 76.85 | 77.3833 | 77.1375
11 | 73.9571 | 76.4833 | 77.1 | 76.925
11.1 | 73.54 | 76.275 | 76.8625 | 76.6667
11.2 | 73.3 | 75.95 | 76.4875 | 76.35
11.3 | 72.95 | 75.7 | 76.2375 | 76.1
11.4 | 72.6 | 75.375 | 75.9375 | 75.775
11.5 | 72.35 | 75.1 | 75.6 | 75.5
11.6 | 72.04 | 74.875 | 75.25 | 75.26
11.7 | 71.65 | 74.525 | 75 | 75.1125
11.8 | 71.325 | 74.275 | 74.85 | 74.7
11.9 | 71.075 | 74.025 | 74.625 | 74.44
12 | 70.8333 | 73.75 | 74.275 | 74.2
12.1 | 70.46 | 73.35 | 74.025 | 73.95
12.2 | 70.15 | 73.075 | 73.7 | 73.7
12.3 | 69.8 | 72.825 | 73.575 | 73.36
12.4 | 69.45 | 72.58 | 73.3 | 73.125
12.5 | 69.2 | 72.325 | 73 | 72.875
12.6 | 69 | 72.1 | 72.75 | 72.625
12.7 | 68.62 | 71.75 | 72.4 | 72.4
12.8 | 68.375 | 71.3 | 72.275 | 72.2
12.9 | 68.025 | 71.15 | 71.925 | 71.9
13 | 67.775 | 71 | 71.675 | 71.6
13.1 | 67.425 | 70.675 | 71.4625 | 71.45
13.2 | 67.175 | 70.425 | 71.3286 | 71.2
13.3 | 66.875 | 70.075 | 71 | 70.875
13.4 | 66.64 | 69.78 | 70.6333 | 70.6222
13.5 | 66.3 | 69.5 | 70.35 | 70.4625
13.6 | 66.125 | 69.2 | 70 | 70.2
13.7 | 65.75 | 68.95 | 69.825 | 69.875
13.8 | 65.4 | 68.7 | 69.4667 | 69.58
13.9 | 65.06 | 68.425 | 69.2375 | 69.36
14 | 64.825 | 68.0857 | 69.025 | 69.15
14.1 | 64.5429 | 67.825 | 68.85 | 69
14.2 | 64.275 | 67.5 | 68.5 | 68.6
14.3 | 63.9 | 67.275 | 68.2 | 68.275
14.4 | 63.6 | 67.025 | 67.9714 | 68.025
14.5 | 63.4 | 66.875 | 67.775 | 67.775
14.6 | 63.05 | 66.68 | 67.5 | 67.625
14.7 | 62.75 | 66.3 | 67.1 | 67.4
14.8 | 62.45 | 66.1 | 66.925 | 67.05
14.9 | 62.225 | 65.8 | 66.65 | 66.7
15 | 61.95 | 65.5 | 66.34 | 66.55
15.1 | 61.625 | 65.25 | 66.2 | 66.3
15.2 | 61.4 | 65 | 65.875 | 66.05
15.3 | 61.225 | 64.8 | 65.6 | 65.9
15.4 | 60.9 | 64.475 | 65.48 | 65.675
15.5 | 60.6 | 64.25 | 65.25 | 65.5
15.6 | 60.5 | 63.98 | 65.1 | 65.275
15.7 | 60.25 | 63.75 | 64.75 | 64.95
15.8 | 60 | 63.52 | 64.4 | 64.7
15.9 | 59.75 | 63.2667 | 64.25 | 64.475
16 | 59.525 | 62.975 | 64.1 | 64.2571
16.1 | 59.175 | 62.8 | 63.9 | 64.05
16.2 | 58.8 | 62.5333 | 63.7 | 63.7
16.3 | 58.5714 | 62.25 | 63.475 | 63.5286
16.4 | 58.4 | 62.02 | 63.2 | 63.3222
16.5 | 58.2375 | 61.875 | 63.05 | 63.075
16.6 | 58.025 | 61.625 | 62.825 | 62.825
16.7 | 57.775 | 61.375 | 62.55 | 62.58
16.8 | 57.475 | 61.1 | 62.2 | 62.4
16.9 | 57.15 | 60.88 | 61.875 | 62.3
17 | 56.875 | 60.54 | 61.62 | 61.95
17.1 | 56.6333 | 60.3 | 61.375 | 61.7
17.2 | 56.4 | 60.15 | 61.125 | 61.55
17.3 | 56.125 | 59.8 | 60.975 | 61.3
17.4 | 55.875 | 59.55 | 60.725 | 61.075
17.5 | 55.675 | 59.4 | 60.525 | 60.925
17.6 | 55.4 | 59.15 | 60.3625 | 60.7333
17.7 | 55.275 | 58.95 | 60.1125 | 60.44
17.8 | 55 | 58.65 | 59.84 | 60.25
17.9 | 54.775 | 58.5 | 59.7 | 59.9
18 | 54.5 | 58.2 | 59.45 | 59.675
18.1 | 54.2 | 58.05 | 59.3 | 59.5125
18.2 | 54.05 | 57.8 | 59.15 | 59.3857
18.3 | 53.8 | 57.675 | 58.94 | 59.14
18.4 | 53.6 | 57.425 | 58.575 | 58.875
18.5 | 53.325 | 57.075 | 58.3 | 58.6333
18.6 | 53.05 | 56.825 | 58.175 | 58.4
18.7 | 52.9 | 56.5 | 57.9 | 58.175
18.8 | 52.675 | 56.325 | 57.68 | 58
18.9 | 52.425 | 56.2 | 57.42 | 57.825
19 | 52.275 | 55.9 | 57.15 | 57.6571
19.1 | 52.075 | 55.7333 | 57 | 57.475
19.2 | 51.84 | 55.4333 | 56.7286 | 57.225
19.3 | 51.6 | 55.2 | 56.5167 | 56.95
19.4 | 51.3 | 55 | 56.36 | 56.7
19.5 | 51.1 | 54.85 | 56.2 | 56.55
19.6 | 50.8667 | 54.72 | 56.0667 | 56.4
19.7 | 50.6 | 54.4 | 55.7 | 56.15
19.8 | 50.475 | 54.15 | 55.5667 | 56
19.9 | 50.225 | 54 | 55.325 | 55.7667
20 | 50 | 53.8 | 55.1 | 55.4
20.1 | 49.775 | 53.55 | 54.95 | 55.3
20.2 | 49.5333 | 53.3 | 54.65 | 55.175
20.3 | 49.35 | 53.075 | 54.425 | 54.94
20.4 | 49.125 | 52.92 | 54.2 | 54.8
20.5 | 48.85 | 52.6714 | 54.05 | 54.6125
20.6 | 48.65 | 52.3857 | 53.8 | 54.3444
20.7 | 48.5 | 52.125 | 53.5875 | 54.2
20.8 | 48.4 | 52 | 53.425 | 53.95
20.9 | 48.2 | 51.8 | 53.2167 | 53.75
21 | 48.025 | 51.7 | 53.1 | 53.6
21.1 | 47.675 | 51.5 | 52.94 | 53.3
21.2 | 47.425 | 51.25 | 52.8 | 53.2
21.3 | 47.2 | 51 | 52.55 | 53
21.4 | 47.0333 | 50.8 | 52.32 | 52.8
21.5 | 46.85 | 50.56 | 52.175 | 52.56
21.6 | 46.6 | 50.4 | 51.925 | 52.4
21.7 | 46.36 | 50.2 | 51.7 | 52.3
21.8 | 46.225 | 50 | 51.525 | 52.025
21.9 | 46.1 | 49.8 | 51.375 | 51.8
22 | 45.8 | 49.6 | 51.2 | 51.625
22.1 | 45.7 | 49.4 | 50.9333 | 51.36
22.2 | 45.45 | 49.3 | 50.8 | 51.225
22.3 | 45.3 | 49 | 50.6 | 51.075
22.4 | 45.1 | 48.825 | 50.4 | 50.925
22.5 | 44.95 | 48.575 | 50.2 | 50.76
22.6 | 44.7 | 48.425 | 50 | 50.55
22.7 | 44.45 | 48.175 | 49.825 | 50.3
22.8 | 44.3 | 48 | 49.6 | 50.15
22.9 | 44.05 | 47.8 | 49.5 | 50
23 | 43.9 | 47.65 | 49.1 | 49.75
23.1 | 43.75 | 47.4 | 48.9 | 49.6
23.2 | 43.6 | 47.35 | 48.775 | 49.475
23.3 | 43.5 | 47.2 | 48.625 | 49.25
23.4 | 43.3 | 46.9 | 48.475 | 48.975
23.5 | 43.1 | 46.7 | 48.3 | 48.8
23.6 | 42.9 | 46.6 | 48.2 | 48.675
23.7 | 42.65 | 46.475 | 48 | 48.525
23.8 | 42.4667 | 46.275 | 47.725 | 48.375
23.9 | 42.4 | 46.025 | 47.6 | 48.125
24 | 42.075 | 45.875 | 47.45 | 47.875
24.1 | 41.9429 | 45.7 | 47.3 | 47.7333
24.2 | 41.725 | 45.5833 | 47.05 | 47.6
24.3 | 41.55 | 45.375 | 46.9 | 47.35
24.4 | 41.3667 | 45.1 | 46.65 | 47.25
24.5 | 41.2 | 44.975 | 46.5 | 47.125
24.6 | 41.1 | 44.8 | 46.275 | 46.9444
24.7 | 40.94 | 44.6 | 46.125 | 46.75
24.8 | 40.7 | 44.45 | 45.9 | 46.6
24.9 | 40.55 | 44.3 | 45.8 | 46.375
25 | 40.3 | 44.1286 | 45.675 | 46.2
25.1 | 40.15 | 43.9 | 45.475 | 46.025
25.2 | 39.9 | 43.8 | 45.225 | 45.9
25.3 | 39.8 | 43.7 | 45.08 | 45.775
25.4 | 39.7 | 43.45 | 44.95 | 45.5667
25.5 | 39.56 | 43.2 | 44.8 | 45.4
25.6 | 39.325 | 43 | 44.55 | 45.1833
25.7 | 39.075 | 42.825 | 44.4 | 45.1
25.8 | 38.925 | 42.675 | 44.26 | 44.9
25.9 | 38.8 | 42.58 | 44 | 44.725
26 | 38.5 | 42.35 | 43.9 | 44.6
26.1 | 38.4 | 42.2 | 43.725 | 44.45
26.2 | 38.25 | 42.1 | 43.6 | 44.2
26.3 | 38.1 | 41.8571 | 43.4 | 44.06
26.4 | 37.95 | 41.7111 | 43.3 | 43.9
26.5 | 37.75 | 41.4875 | 43.1 | 43.8
26.6 | 37.7 | 41.3625 | 43 | 43.6
26.7 | 37.575 | 41.2375 | 42.7333 | 43.5
26.8 | 37.425 | 41.025 | 42.575 | 43.275
26.9 | 37.175 | 40.8 | 42.4 | 43.1
27 | 37.025 | 40.6667 | 42.2 | 43
27.1 | 36.875 | 40.5 | 42.1 | 42.8
27.2 | 36.725 | 40.4 | 41.9625 | 42.7
27.3 | 36.5 | 40.2 | 41.8 | 42.5556
27.4 | 36.3333 | 40.075 | 41.65 | 42.375
27.5 | 36.175 | 39.925 | 41.4 | 42.2125
27.6 | 36.025 | 39.775 | 41.35 | 42.0875
27.7 | 35.8 | 39.625 | 41.15 | 41.9625
27.8 | 35.7 | 39.4 | 41.025 | 41.775
27.9 | 35.575 | 39.225 | 40.9 | 41.6125
28 | 35.45 | 39 | 40.6875 | 41.4875
28.1 | 35.3 | 38.9 | 40.5625 | 41.325
28.2 | 35.15 | 38.7 | 40.475 | 41.175
28.3 | 35 | 38.6333 | 40.3 | 40.925
28.4 | 34.95 | 38.5 | 40.2 | 40.8
28.5 | 34.8 | 38.4 | 39.95 | 40.7
28.6 | 34.6 | 38.3 | 39.7 | 40.475
28.7 | 34.525 | 38.175 | 39.575 | 40.325
28.8 | 34.3 | 37.8333 | 39.425 | 40.175
28.9 | 34.2 | 37.7 | 39.375 | 40.025
29 | 34.1 | 37.6 | 39.225 | 39.9
29.1 | 34 | 37.4667 | 39 | 39.75
29.2 | 33.85 | 37.35 | 38.9 | 39.6
29.3 | 33.7 | 37.225 | 38.8 | 39.45
29.4 | 33.55 | 37.0857 | 38.6 | 39.4
29.5 | 33.4 | 36.9333 | 38.5 | 39.175
29.6 | 33.16 | 36.8 | 38.3 | 39.025
29.7 | 33.025 | 36.6 | 38.15 | 38.9
29.8 | 32.8833 | 36.5 | 38 | 38.8125
29.9 | 32.7375 | 36.35 | 37.85 | 38.625
30 | 32.675 | 36.22 | 37.7 | 38.475

## Wedge 60 Lateral Profiles
Lateral Profiles
Wedge 60
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:3 cm FieldSize:4x4 cm2 | Depth:3 cm FieldSize:10x10 cm2 | Depth:3 cm FieldSize:15x15 cm2 | Depth:3 cm FieldSize:40x15 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:15x15 cm2 | Depth:5 cm FieldSize:40x15 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:15x15 cm2 | Depth:10 cm FieldSize:40x15 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:15x15 cm2 | Depth:20 cm FieldSize:40x15 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:15x15 cm2 | Depth:30 cm FieldSize:40x15 cm2
-14.7 | [18 empty] | 11.2138 | 15.7
-14.6 | [18 empty] | 11.4385 | 15.98
-14.5 | [18 empty] | 11.6883 | 16.2333
-14.4 | [18 empty] | 11.8881 | 16.425
-14.3 | [18 empty] | 12.0879 | 16.6
-14.2 | [18 empty] | 12.3876 | 16.8
-14.1 | [18 empty] | 12.5541 | 17.1
-14 | [18 empty] | 12.8871 | 17.3
-13.9 | [14 empty] | 9.2 | 13.2 |  |  | 13.1535 | 17.5
-13.8 | [14 empty] | 9.4667 | 13.425 |  |  | 13.2867 | 17.675
-13.7 | [14 empty] | 9.625 | 13.6333 |  |  | 13.4116 | 17.9
-13.6 | [14 empty] | 9.75 | 13.8 |  |  | 13.7862 | 18.225
-13.5 | [14 empty] | 10 | 14 |  |  | 13.986 | 18.55
-13.4 | [14 empty] | 10.175 | 14.15 |  |  | 14.1608 | 18.8
-13.3 | [14 empty] | 10.3333 | 14.325 |  |  | 14.5105 | 19.075
-13.2 | [10 empty] | 6.9 | 10.48 |  |  | 10.6 | 14.6 |  |  | 14.7852 | 19.3
-13.1 | [10 empty] | 7.05 | 10.6667 |  |  | 10.75 | 14.7333 |  |  | 14.985 | 19.6
-13 | [10 empty] | 7.2 | 10.775 |  |  | 10.9 | 14.95 |  |  | 15.1848 | 19.9
-12.9 | [10 empty] | 7.3 | 10.9667 |  |  | 11.1667 | 15.2 |  |  | 15.5644 | 20.25
-12.8 | [6 empty] | 6 | 9.4 |  |  | 7.4 | 11.05 |  |  | 11.4 | 15.5 |  |  | 15.8092 | 20.48
-12.7 |  |  | 6.6 | 10.4 |  |  | 6.1 | 9.5 |  |  | 7.6 | 11.2 |  |  | 11.6 | 15.7333 |  |  | 15.984 | 20.825
-12.6 |  |  | 6.7 | 10.5667 |  |  | 6.2 | 9.6 |  |  | 7.7 | 11.4 |  |  | 11.775 | 15.9333 |  |  | 16.4336 | 21.1333
-12.5 |  |  | 6.75 | 10.7 |  |  | 6.3 | 9.7 |  |  | 7.9 | 11.6 |  |  | 12.05 | 16.3 |  |  | 16.6633 | 21.525
-12.4 |  |  | 6.84 | 10.8 |  |  | 6.4 | 9.9 |  |  | 8.075 | 11.7667 |  |  | 12.25 | 16.4667 |  |  | 17.0496 | 21.8333
-12.3 |  |  | 7 | 10.9 |  |  | 6.5 | 10 |  |  | 8.2 | 11.9333 |  |  | 12.575 | 16.725 |  |  | 17.4076 | 22.2
-12.2 |  |  | 7.1 | 11.0667 |  |  | 6.7 | 10.125 |  |  | 8.3667 | 12.1 |  |  | 12.8333 | 17.025 |  |  | 17.7572 | 22.5
-12.1 |  |  | 7.3 | 11.25 |  |  | 6.8 | 10.2333 |  |  | 8.6 | 12.275 |  |  | 13.0667 | 17.3 |  |  | 18.1069 | 22.8667
-12 |  |  | 7.375 | 11.425 |  |  | 7 | 10.5 |  |  | 8.7667 | 12.425 |  |  | 13.4 | 17.65 |  |  | 18.5314 | 23.3
-11.9 |  |  | 7.525 | 11.52 |  |  | 7.1 | 10.6 |  |  | 8.9 | 12.675 |  |  | 13.65 | 17.9 |  |  | 18.9211 | 23.6333
-11.8 |  |  | 7.625 | 11.725 |  |  | 7.1 | 10.7667 |  |  | 9.08 | 12.875 |  |  | 13.9 | 18.2 |  |  | 19.3307 | 24.1667
-11.7 |  |  | 7.7667 | 11.85 |  |  | 7.3 | 10.9 |  |  | 9.2667 | 13.1333 |  |  | 14.25 | 18.5 |  |  | 19.9134 | 24.5
-11.6 |  |  | 7.9 | 12 |  |  | 7.45 | 11 |  |  | 9.475 | 13.4 |  |  | 14.55 | 18.9 |  |  | 20.4295 | 25.1333
-11.5 |  |  | 8.1 | 12.2 |  |  | 7.6 | 11.25 |  |  | 9.625 | 13.575 |  |  | 14.9 | 19.2667 |  |  | 20.8791 | 25.6333
-11.4 |  |  | 8.15 | 12.4 |  |  | 7.8 | 11.4 |  |  | 9.9 | 13.7 |  |  | 15.3 | 19.625 |  | 7.4 | 21.4785 | 26.2667
-11.3 |  |  | 8.4 | 12.5 |  |  | 7.9 | 11.65 |  |  | 10.2 | 14 |  |  | 15.625 | 19.95 |  | 7.575 | 22.0779 | 26.8
-11.2 |  |  | 8.5 | 12.7 |  |  | 8.0667 | 11.8 |  |  | 10.4 | 14.3 |  |  | 15.975 | 20.36 |  | 7.6667 | 22.7173 | 27.3
-11.1 |  |  | 8.6667 | 12.9 |  |  | 8.3 | 11.925 |  |  | 10.6 | 14.5333 |  |  | 16.48 | 20.775 |  | 7.9 | 23.6963 | 28.2
-11 |  |  | 8.85 | 13.1 |  |  | 8.5 | 12.175 |  |  | 10.8 | 14.8 |  |  | 16.85 | 21.25 |  | 8.0667 | 24.4755 | 29.0333
-10.9 |  |  | 9.1 | 13.2667 |  |  | 8.675 | 12.4 |  |  | 11.15 | 15.1 |  | 5.9 | 17.45 | 21.6667 |  | 8.1667 | 25.5244 | 30.075
-10.8 |  |  | 9.2 | 13.475 |  |  | 8.9 | 12.625 |  |  | 11.5 | 15.35 |  | 6 | 17.85 | 22.25 |  | 8.35 | 26.5734 | 31.25
-10.7 |  |  | 9.3 | 13.7 |  |  | 9.025 | 12.82 |  |  | 11.75 | 15.575 |  | 6.125 | 18.46 | 22.7667 |  | 8.5 | 28.0719 | 32.5
-10.6 |  |  | 9.6 | 13.9 |  |  | 9.2 | 13.0667 |  |  | 11.975 | 15.9667 |  | 6.3 | 19.05 | 23.3 |  | 8.65 | 29.8951 | 34.1667
-10.5 |  |  | 9.8 | 14.1667 |  |  | 9.475 | 13.3 |  |  | 12.3667 | 16.2667 |  | 6.375 | 19.75 | 23.9 |  | 8.82 | 31.7682 | 36.125
-10.4 |  |  | 10.025 | 14.4 |  |  | 9.7 | 13.55 |  | 4.3 | 12.825 | 16.675 |  | 6.42 | 20.55 | 24.7 |  | 9 | 34.2657 | 38.425
-10.3 |  |  | 10.26 | 14.6 |  |  | 9.95 | 13.8 |  | 4.4 | 13.15 | 17.075 |  | 6.575 | 21.3 | 25.5 |  | 9.2 | 37.7122 | 41.475
-10.2 |  |  | 10.4 | 14.75 |  | 3.7012 | 10.2667 | 14.075 |  | 4.5 | 13.525 | 17.55 |  | 6.6333 | 22.2 | 26.4 |  | 9.3 | 41.958 | 45.9
-10.1 |  | 4 | 10.7 | 15.025 |  | 3.8013 | 10.5 | 14.375 |  | 4.54 | 13.975 | 17.875 |  | 6.8 | 23.3333 | 27.32 |  | 9.4 | 48.3017 | 50.6
-10 |  | 4 | 10.9 | 15.275 |  | 3.8013 | 10.8333 | 14.725 |  | 4.675 | 14.5 | 18.45 |  | 6.9 | 24.5667 | 28.7 |  | 9.65 | 56.7433 | 58.75
-9.9 |  | 4.1 | 11.3333 | 15.625 |  | 3.9013 | 11.15 | 15.05 |  | 4.7 | 14.95 | 18.9 |  | 7 | 26.2 | 30.125 |  | 9.8 | 67.3826 | 68.9
-9.8 |  | 4.1667 | 11.5667 | 15.825 |  | 3.9013 | 11.5333 | 15.425 |  | 4.8 | 15.5 | 19.55 |  | 7.2 | 28.075 | 32 |  | 10.05 | 80.3796 | 81.45
-9.7 |  | 4.2 | 11.75 | 16.225 |  | 4.0013 | 11.9 | 15.775 |  | 4.9 | 16.25 | 20.1 |  | 7.3 | 30.375 | 34.275 |  | 10.2 | 93.8811 | 94.3
-9.6 |  | 4.3 | 12.15 | 16.55 |  | 4.0013 | 12.35 | 16.2 |  | 5 | 17 | 20.9 |  | 7.6 | 33.45 | 37.05 |  | 10.48 | 107.7675 | 107.15
-9.5 |  | 4.3333 | 12.5 | 16.9 |  | 4.1014 | 12.9 | 16.7 |  | 5.1 | 17.9 | 21.8 |  | 7.7 | 37.375 | 40.8 |  | 10.675 | 120.255 | 117.65
-9.4 |  | 4.4667 | 12.8667 | 17.3 |  | 4.2014 | 13.4 | 17.2 |  | 5.2 | 18.9 | 22.75 |  | 7.9 | 42.85 | 45.925 |  | 10.825 | 130.7942 | 127.55
-9.3 |  | 4.5 | 13.25 | 17.7 |  | 4.3014 | 14.05 | 17.85 |  | 5.3 | 20.1 | 23.9 |  | 8.2 | 51.025 | 53.3 |  | 11.05 | 138.262 | 134.8333
-9.2 |  | 4.6 | 13.8 | 18.25 |  | 4.4015 | 14.8 | 18.5 |  | 5.5 | 21.54 | 25.25 |  | 8.3 | 61.8 | 63.5333 |  | 11.325 | 143.4315 | 139.45
-9.1 |  | 4.6667 | 14.35 | 18.8 |  | 4.5015 | 15.6 | 19.3333 |  | 5.6 | 23.2 | 26.925 |  | 8.475 | 74.4 | 75.2 |  | 11.575 | 146.254 | 142.75
-9 |  | 4.725 | 15.075 | 19.45 |  | 4.5348 | 16.6 | 20.2667 |  | 5.7 | 25.35 | 29.025 |  | 8.7 | 89.55 | 89.35 |  | 11.85 | 148.6515 | 144.6
-8.9 |  | 4.8 | 15.8 | 20.3 |  | 4.6682 | 17.6667 | 21.42 |  | 5.7333 | 27.925 | 31.8 |  | 8.95 | 105 | 103.8 |  | 12.1 | 150.25 | 146.2
-8.8 |  | 4.9 | 16.8 | 21.2 |  | 4.8016 | 19.1667 | 22.75 |  | 5.9 | 31.925 | 35.3 |  | 9.1 | 119.8 | 117.8 |  | 12.4 | 151.4485 | 147.25
-8.7 |  | 5 | 17.9667 | 22.35 |  | 4.8016 | 21.06 | 24.66 |  | 6.075 | 37 | 39.9 |  | 9.3 | 131.7 | 128.1 |  | 12.7 | 152.148 | 147.9333
-8.6 |  | 5.1 | 19.3 | 23.65 |  | 4.9016 | 23.25 | 26.825 |  | 6.225 | 44.05 | 47 |  | 9.5 | 140.65 | 137.025 |  | 13.0667 | 152.514 | 148.375
-8.5 |  | 5.2 | 21.4 | 25.425 |  | 5.0017 | 26.2 | 29.55 |  | 6.375 | 54.7333 | 57.15 |  | 9.8 | 146.475 | 142.75 |  | 13.3667 | 152.687 | 148.525
-8.4 |  | 5.3 | 23.8 | 27.925 |  | 5.2017 | 30.6333 | 33.2 |  | 6.5 | 68.8 | 71.4 |  | 10.14 | 150.75 | 146.45 |  | 13.9333 | 152.707 | 148.7
-8.3 |  | 5.4 | 27.0333 | 31.25 |  | 5.3018 | 36 | 38.875 |  | 6.7 | 84.5 | 86.7 |  | 10.375 | 153.25 | 149.05 |  | 14.15 | 152.672 | 148.7
-8.2 |  | 5.5 | 31.7 | 35.9 |  | 5.4018 | 44.6 | 47.1333 |  | 6.8 | 102.7 | 103.4 |  | 10.6667 | 154.975 | 150.625 |  | 14.575 | 152.547 | 148.5
-8.1 |  | 5.6 | 38.8 | 42.25 |  | 5.5018 | 57.9 | 58.8333 |  | 7.1 | 119.1 | 117.45 |  | 10.9333 | 155.9 | 151.8 |  | 15.05 | 152.223 | 148.3667
-8 |  | 5.725 | 49.7333 | 52.6 |  | 5.6269 | 74.15 | 74.3 |  | 7.325 | 133.36 | 131.325 |  | 11.35 | 156.55 | 152.4333 |  | 15.55 | 151.948 | 148
-7.9 |  | 5.875 | 65.7 | 67.3 |  | 5.8269 | 92.2 | 91.45 |  | 7.56 | 144.54 | 141.775 |  | 11.675 | 156.825 | 152.7 |  | 16.15 | 151.648 | 147.6
-7.8 |  | 6 | 82.3 | 84.3 |  | 6.022 | 110.5333 | 109.3 |  | 7.8 | 151.9 | 149.1333 |  | 12.05 | 157 | 152.8 |  | 16.8 | 151.109 | 147.3667
-7.7 |  | 6.1333 | 102.1 | 102.1667 |  | 6.1521 | 127.2667 | 125.3 |  | 8.05 | 156.65 | 153.225 |  | 12.55 | 156.975 | 152.8 |  | 17.74 | 150.549 | 147
-7.6 |  | 6.3 | 120.1667 | 120.15 |  | 6.3771 | 142 | 138.4 |  | 8.28 | 159.425 | 155.8 |  | 12.925 | 156.8333 | 152.7 | 2.7986 | 18.6 | 150.05 | 146.5
-7.5 |  | 6.475 | 136.575 | 134.4 |  | 6.5422 | 152.05 | 148.5 |  | 8.55 | 161.3 | 157.475 |  | 13.4 | 156.5 | 152.35 | 2.7986 | 19.75 | 149.5005 | 146
-7.4 |  | 6.6333 | 149.825 | 146.475 |  | 6.7523 | 158.675 | 155.06 |  | 8.8333 | 162.275 | 158.425 | 2.1 | 13.95 | 156.25 | 151.9 | 2.8986 | 20.85 | 149.0177 | 145.4
-7.3 |  | 6.875 | 158.425 | 154.75 |  | 7.0023 | 162.75 | 158.7333 |  | 9.25 | 162.85 | 158.9 | 2.2 | 14.7 | 155.7333 | 151.6667 | 2.8986 | 22.55 | 148.518 | 144.8
-7.2 |  | 7.1 | 163.125 | 158.95 |  | 7.2024 | 164.875 | 161.025 |  | 9.7 | 163 | 159.1 | 2.2 | 15.45 | 155.1333 | 151.3 | 2.9985 | 24.35 | 148.0187 | 144.3
-7.1 |  | 7.35 | 165.775 | 161.675 | 1.3 | 7.4775 | 166.2 | 162.275 | 1.6005 | 10.05 | 163 | 159.1 | 2.2 | 16.3 | 154.3 | 150.7 | 3.0985 | 26.7 | 146.9197 | 143.7
-7 | 1.4007 | 7.5 | 167.225 | 162.925 | 1.35 | 7.9026 | 166.9 | 162.9 | 1.6005 | 10.55 | 162.8 | 158.95 | 2.3 | 17.45 | 153.75 | 150.1667 | 3.0985 | 29.9 | 146.304 | 143.1
-6.9 | 1.4007 | 7.8 | 167.8 | 163.42 | 1.4 | 8.2527 | 167 | 163.1 | 1.6005 | 11.1667 | 162.2667 | 158.5 | 2.3 | 18.6 | 153.1333 | 149.5 | 3.0985 | 34.0667 | 145.654 | 142.5
-6.8 | 1.4007 | 8.1333 | 167.9 | 163.5 | 1.4 | 8.6529 | 167 | 163.1 | 1.6005 | 11.75 | 161.8 | 158.0667 | 2.4 | 20.18 | 152.4 | 148.9 | 3.0985 | 40.1 | 145.105 | 141.8
-6.7 | 1.4007 | 8.4 | 167.85 | 163.5 | 1.4 | 9.053 | 166.8 | 162.9 | 1.6005 | 12.4333 | 161.2 | 157.5667 | 2.4 | 22.2 | 151.7667 | 148.225 | 3.0985 | 48.3 | 144.605 | 141.4
-6.6 | 1.4007 | 8.8333 | 167.5 | 163.3 | 1.4 | 9.6032 | 166.2 | 162.6 | 1.7006 | 13.3 | 160.7 | 156.8667 | 2.4 | 24.525 | 151.025 | 147.4 | 3.1984 | 58 | 143.6813 | 140.7
-6.5 | 1.4007 | 9.325 | 167.15 | 162.7 | 1.4 | 10.2034 | 165.5 | 162 | 1.7006 | 14.4 | 159.84 | 156.225 | 2.5 | 27.675 | 150.25 | 146.575 | 3.365 | 70.7 | 143.207 | 140.0667
-6.4 | 1.5008 | 9.85 | 166.225 | 162 | 1.4 | 11.0037 | 164.6 | 161.4 | 1.8006 | 15.6667 | 158.925 | 155.5 | 2.5 | 32.3667 | 149.28 | 145.925 | 3.3983 | 81.8 | 142.657 | 139.4667
-6.3 | 1.4007 | 10.5 | 165.325 | 161 | 1.5 | 11.8539 | 163.9 | 160.475 | 1.8006 | 17.3 | 158 | 154.3333 | 2.6 | 38.775 | 148.575 | 145.175 | 3.4983 | 94 | 141.808 | 138.8
-6.2 | 1.5008 | 11.3 | 164.325 | 160.28 | 1.5 | 13.0043 | 162.875 | 159.4333 | 1.8006 | 19.1 | 157.3 | 153.4333 | 2.7 | 47.7667 | 147.65 | 144.7 | 3.4983 | 104.1333 | 140.759 | 138
-6.1 | 1.5008 | 12.3 | 163.275 | 159.1 | 1.5 | 14.4048 | 161.875 | 158.625 | 1.8006 | 21.775 | 156.3333 | 152.9 | 2.7 | 59 | 147.1 | 143.8 | 3.5982 | 111.8 | 140.36 | 137.3
-6 | 1.5008 | 13.4333 | 162.2 | 158.2 | 1.5 | 16.0453 | 160.9 | 157.775 | 1.9006 | 25.1 | 155.075 | 152 | 2.8 | 71.6 | 145.9 | 143 | 3.6982 | 117.6 | 139.4105 | 136.9
-5.9 | 1.5008 | 15 | 161.2 | 156.975 | 1.5 | 18.2861 | 159.8 | 156.475 | 1.9006 | 29.85 | 154.12 | 150.8333 | 2.9 | 84.2 | 145.2 | 142.2 | 3.7981 | 120.9 | 138.9277 | 136.125
-5.8 | 1.6008 | 17.15 | 159.72 | 156 | 1.6 | 21.107 | 158.7 | 155.5333 | 1.934 | 37.05 | 153.2 | 150 | 3 | 96.7333 | 144.2 | 141.3 | 3.8981 | 123.7 | 138.162 | 135.5
-5.7 | 1.6008 | 20.2 | 158.7 | 154.55 | 1.6 | 25.3751 | 157.5 | 154.7 | 2.0007 | 47.525 | 152.2 | 148.9667 | 3 | 106.4 | 143.25 | 140.4 | 3.998 | 125.0667 | 137.1297 | 134.75
-5.6 | 1.6008 | 24.4 | 157.35 | 153.7333 | 1.6 | 31.6439 | 156.4 | 153.4 | 2.1007 | 60.275 | 151.05 | 148 | 3.1 | 114.75 | 142.6 | 139.65 | 4.0979 | 126.1667 | 136.6297 | 134.175
-5.5 | 1.6008 | 30.4 | 155.9 | 152.7 | 1.7 | 41.1637 | 155 | 152.3 | 2.1007 | 74.475 | 149.8 | 147 | 3.1 | 120.3 | 141.65 | 138.9667 | 4.1979 | 126.8333 | 135.989 | 133.525
-5.4 | 1.7008 | 40.12 | 154.6 | 151.5333 | 1.7 | 53.4178 | 153.8 | 151.0333 | 2.1007 | 87.8 | 148.95 | 146 | 3.2 | 123.3333 | 140.9 | 137.95 | 4.2979 | 127.1 | 135.19 | 132.86
-5.3 | 1.7008 | 52.375 | 153.58 | 150.4 | 1.7 | 67.4225 | 152.8 | 150.1 | 2.2007 | 100.3 | 147.9 | 144.9 | 3.3 | 125.7 | 139.8 | 137.25 | 4.3978 | 127.3667 | 134.4408 | 131.9
-5.2 | 1.7008 | 67.6667 | 152.575 | 149.1 | 1.8 | 81.8273 | 151.6 | 148.85 | 2.2674 | 111.9 | 146.7667 | 143.9 | 3.4 | 127.04 | 139.1667 | 136.6 | 4.4977 | 127.4 | 133.267 | 131.4
-5.1 | 1.8009 | 83.35 | 151.35 | 147.9 | 1.8 | 96.5322 | 150.4 | 147.9 | 2.3008 | 120.15 | 145.7333 | 142.975 | 3.5 | 127.95 | 138.25 | 135.65 | 4.5977 | 127.4 | 132.717 | 130.4
-5 | 1.8009 | 97.925 | 149.775 | 146.8 | 1.8 | 109.436 | 149 | 146.55 | 2.4008 | 124.56 | 144.625 | 141.825 | 3.5 | 128.425 | 137.4 | 134.9 | 4.8642 | 127.3 | 131.868 | 129.85
-4.9 | 1.9009 | 110.725 | 148.6 | 145.6 | 1.9 | 119.3895 | 147.875 | 145.1 | 2.4008 | 127.75 | 143.5 | 140.875 | 3.7333 | 128.6 | 136.4667 | 134 | 4.9975 | 127.1 | 131.169 | 129.25
-4.8 | 1.9009 | 121 | 147.3 | 144.275 | 1.9 | 125.642 | 146.5667 | 144.1 | 2.5008 | 129.75 | 142.4 | 139.82 | 3.9 | 128.6 | 135.5 | 133.1667 | 5.1474 | 126.85 | 130.569 | 128.5
-4.7 | 2.001 | 127.7 | 146.2 | 143.1667 | 2 | 129.393 | 145.2667 | 143.05 | 2.6009 | 130.875 | 141.3333 | 138.775 | 4.1 | 128.6 | 134.625 | 132.35 | 5.2973 | 126.6 | 129.92 | 127.9
-4.6 | 2.001 | 130.86 | 144.7 | 142.3 | 2.1 | 131.269 | 144.3333 | 141.975 | 2.7509 | 131.35 | 139.9 | 137.725 | 4.2 | 128.4 | 133.45 | 131.5 | 5.3973 | 126.0333 | 129.046 | 127.15
-4.5 | 2.101 | 132.8 | 143.6 | 140.6 | 2.15 | 132.694 | 142.8 | 140.7 | 2.901 | 131.7 | 138.9667 | 136.7667 | 4.3 | 128.15 | 132.8 | 130.6 | 5.5972 | 125.4667 | 128.272 | 126.6
-4.4 | 2.1677 | 134 | 142.25 | 139.7 | 2.2 | 133.294 | 141.48 | 139.55 | 3.001 | 131.775 | 138.1 | 135.7 | 4.5 | 127.8 | 132 | 129.75 | 5.8971 | 125.1 | 127.5973 | 125.85
-4.3 | 2.2011 | 134.4 | 141 | 138.4 | 2.3 | 133.545 | 140.45 | 138.2 | 3.101 | 131.64 | 136.85 | 134.65 | 4.7 | 127.275 | 131.2 | 129.1 | 6.0969 | 124.7 | 127.0063 | 125.3
-4.2 | 2.3011 | 134.4 | 139.8 | 137.3 | 2.4 | 133.645 | 139.3 | 137.24 | 3.3011 | 131.4 | 135.95 | 133.95 | 4.9 | 126.925 | 130.25 | 128.5 | 6.2969 | 124.18 | 126.174 | 124.65
-4.1 | 2.4012 | 134.25 | 138.65 | 136 | 2.6 | 133.545 | 138 | 136.2 | 3.5012 | 130.9 | 134.7 | 132.8 | 5.1 | 126.375 | 129.4 | 127.475 | 6.8299 | 123.65 | 125.624 | 124.225
-4 | 2.5013 | 133.9333 | 137.5 | 135 | 2.7 | 133.3107 | 136.825 | 135 | 3.7012 | 130.6 | 133.9 | 131.8333 | 5.4 | 125.95 | 128.7 | 126.825 | 7.2964 | 123.15 | 125.1083 | 123.6
-3.9 | 2.6013 | 133.5 | 136.3333 | 133.9 | 2.85 | 132.694 | 135.675 | 133.8 | 3.9013 | 130.15 | 132.8 | 130.8 | 5.65 | 125.36 | 127.9 | 126 | 7.6962 | 122.65 | 124.201 | 122.85
-3.8 | 2.8014 | 132.8667 | 135.1 | 132.7333 | 3.1 | 131.944 | 134.425 | 132.925 | 4.068 | 129.4 | 131.825 | 129.9 | 6.05 | 124.85 | 126.9333 | 125.2 | 8.096 | 122.3 | 123.6097 | 122.2
-3.7 | 2.9015 | 132.1 | 134.1 | 131.7 | 3.2 | 131.4107 | 133.275 | 131.8 | 4.4015 | 128.75 | 130.8 | 129.2 | 6.45 | 124.4 | 126.3 | 124.625 | 8.8956 | 121.8 | 122.902 | 121.6333
-3.6 | 3.1016 | 131.1 | 132.7 | 130.65 | 3.5 | 130.443 | 132.125 | 130.78 | 4.7016 | 128.15 | 129.85 | 127.9 | 7.1 | 123.7 | 125.5 | 123.8333 | 9.4953 | 121.25 | 122.378 | 121.0667
-3.5 | 3.3517 | 130.4 | 131.6 | 129.4 | 3.8 | 129.6097 | 131.2 | 129.7333 | 5.2017 | 127.4 | 128.925 | 127.225 | 7.7 | 123.05 | 125 | 123.2 | 10.5947 | 120.9333 | 121.828 | 120.45
-3.4 | 3.6018 | 129.4667 | 130.7 | 128.6667 | 4.2 | 128.918 | 130.1333 | 128.825 | 5.6519 | 126.75 | 127.86 | 126.4 | 8.3 | 122.5 | 124.2 | 122.5333 | 11.6442 | 120 | 121.129 | 119.8667
-3.3 | 3.9019 | 128.775 | 129.7 | 127.7 | 4.5 | 127.9097 | 129.15 | 127.8 | 6.2021 | 126.025 | 127.1 | 125.575 | 9.4 | 121.8 | 123.4667 | 121.85 | 13.2934 | 119.3 | 120.38 | 119.325
-3.2 | 4.4022 | 128.1 | 128.7 | 126.6 | 5.05 | 127.442 | 128.1667 | 126.8 | 7.0023 | 125.025 | 126.15 | 124.8 | 10.4 | 121.35 | 122.75 | 121.3 | 14.9925 | 118.9333 | 119.78 | 118.85
-3.1 | 4.9525 | 127 | 127.475 | 125.7 | 5.7 | 126.542 | 127.4 | 125.9 | 7.936 | 124.275 | 125.3667 | 123.9 | 12 | 120.8 | 121.95 | 120.5667 | 17.2914 | 118.3 | 119.1477 | 118.2
-3 | 5.6028 | 126 | 126.6 | 124.82 | 6.5 | 125.6753 | 126.3333 | 125.175 | 9.053 | 123.9 | 124.5 | 123 | 13.65 | 120.3 | 121.2 | 119.9 | 20.6897 | 117.9 | 118.5568 | 117.4
-2.9 | 6.4532 | 125.25 | 125.35 | 123.95 | 7.7 | 124.822 | 125.1333 | 124.15 | 10.6035 | 123.1667 | 123.5 | 122.3 | 16.05 | 119.4 | 120.4 | 119.35 | 25.4873 | 117.38 | 118.032 | 116.8
-2.8 | 7.7039 | 124.3 | 124.46 | 123.1 | 9.2 | 124.116 | 124.175 | 123.2333 | 12.3041 | 122.375 | 122.6667 | 121.3 | 19.45 | 118.6 | 119.55 | 118.3 | 32.1839 | 116.7 | 117.258 | 116.1
-2.7 | 9.3046 | 123.5 | 123.64 | 122.1 | 10.9 | 123.216 | 123.3 | 122.3 | 14.8049 | 121.4 | 121.5667 | 120.675 | 24.6 | 117.9 | 118.86 | 117.65 | 40.8796 | 116.025 | 116.4503 | 115.575
-2.6 | 11.6392 | 122.55 | 122.7 | 121.1 | 13.45 | 122.266 | 122.2667 | 121.45 | 18.4061 | 120.5 | 120.9333 | 119.7 | 31.1 | 117.325 | 117.96 | 116.925 | 51.0745 | 115.375 | 115.5843 | 115.15
-2.5 | 14.5573 | 121.7 | 121.8 | 120.25 | 16.9 | 121.265 | 121.5 | 120.3 | 23.6078 | 119.775 | 119.9 | 118.725 | 39.9 | 116.6333 | 117.2 | 116.275 | 61.6192 | 114.825 | 114.9517 | 114.3
-2.4 | 18.9095 | 120.6 | 120.9 | 119.3 | 22 | 120.365 | 120.45 | 119.3 | 31.5105 | 118.9667 | 118.9 | 118.025 | 51 | 116 | 116.4 | 115.525 | 72.6137 | 114.2 | 114.4527 | 113.725
-2.3 | 26.8135 | 119.7 | 119.7 | 118.5 | 30 | 119.44 | 119.5 | 118.6 | 41.6139 | 118.3 | 118.1 | 117.1 | 61.8 | 115.3 | 115.85 | 114.775 | 82.1089 | 113.7 | 113.686 | 113.15
-2.2 | 36.3182 | 118.6 | 118.7 | 117.5333 | 40.6 | 118.5397 | 118.5 | 117.65 | 52.9676 | 117.45 | 117.125 | 116.3 | 73.3 | 114.5 | 115 | 114 | 89.9051 | 113 | 113.187 | 112.7
-2.1 | 49.0746 | 117.675 | 117.75 | 116.725 | 52.85 | 117.489 | 117.4 | 116.6 | 65.3218 | 116.6 | 116.375 | 115.4 | 83.9 | 113.8 | 114.18 | 113.4 | 95.8521 | 112.5667 | 112.5378 | 112.05
-2 | 60.8304 | 117 | 116.8 | 115.7333 | 64.6333 | 116.639 | 116.5 | 115.8 | 75.9253 | 115.425 | 115.375 | 114.5 | 91.1 | 113 | 113.625 | 112.8 | 99.2504 | 111.65 | 112.038 | 111.5
-1.9 | 73.4367 | 116 | 115.925 | 114.7 | 76.2 | 115.939 | 115.7 | 114.8667 | 85.8786 | 114.5 | 114.7 | 113.8 | 97 | 112.4667 | 113.075 | 112 | 101.2157 | 111 | 111.2557 | 110.78
-1.8 | 84.8424 | 114.9 | 115.1 | 113.9 | 86.05 | 114.838 | 114.7667 | 114 | 93.4811 | 113.75 | 113.8 | 113.0333 | 99.95 | 111.8333 | 112.075 | 111.3667 | 103.248 | 110.44 | 110.589 | 110.25
-1.7 | 93.3467 | 114.1 | 114.125 | 113.1 | 94.05 | 113.938 | 114.125 | 113.1 | 98.0327 | 113.125 | 113 | 112.2 | 102.1 | 111.1667 | 111.35 | 110.7 | 104.148 | 109.85 | 109.84 | 109.625
-1.6 | 99.2496 | 113.2 | 113.175 | 112.225 | 98.6667 | 113.138 | 113.175 | 112.35 | 100.934 | 112.3 | 112.1 | 111.375 | 103.65 | 110.375 | 110.75 | 110.2 | 104.848 | 109.075 | 109.291 | 108.95
-1.5 | 102.351 | 112.5 | 112.075 | 111.42 | 101.6 | 112.137 | 112.05 | 111.4 | 103.134 | 111.5333 | 111.2 | 110.6 | 104.65 | 109.7 | 110 | 109.4667 | 105.247 | 108.5667 | 108.7913 | 108.475
-1.4 | 104.452 | 111.5 | 111.2 | 110.5 | 103.7 | 111.437 | 110.975 | 110.7 | 104.335 | 110.75 | 110.625 | 109.9333 | 105.1333 | 108.9 | 109.1667 | 108.775 | 105.497 | 108.25 | 108.067 | 107.875
-1.3 | 105.253 | 110.65 | 110.35 | 109.7 | 104.55 | 110.537 | 110.3 | 109.8 | 104.935 | 109.9333 | 109.7 | 109.1333 | 105.3 | 108.2667 | 108.6333 | 108.0667 | 105.547 | 107.4 | 107.592 | 107.26
-1.2 | 105.553 | 109.5667 | 109.7 | 108.8 | 104.9 | 109.7703 | 109.5 | 108.925 | 105.135 | 109.2 | 108.95 | 108.4 | 105.4 | 107.725 | 107.875 | 107.45 | 105.547 | 106.9 | 106.8597 | 106.7
-1.1 | 105.653 | 108.55 | 108.9 | 108 | 105.1 | 108.836 | 108.7 | 108.25 | 105.235 | 108.45 | 108.175 | 107.6 | 105.4 | 106.875 | 107.225 | 106.9 | 105.347 | 106.4 | 106.244 | 106
-1 | 105.653 | 107.825 | 107.9 | 107.3 | 105.1 | 108.0693 | 107.8 | 107.5 | 105.135 | 107.6333 | 107.5 | 106.875 | 105.2 | 106.3 | 106.4 | 106.1667 | 104.898 | 105.7667 | 105.5945 | 105.5
-0.9 | 105.553 | 107.1 | 107.2 | 106.35 | 105 | 107.036 | 106.9 | 106.6 | 104.935 | 106.875 | 106.7667 | 106.3 | 104.9 | 105.5 | 105.9 | 105.4667 | 104.548 | 105.2 | 104.995 | 104.9
-0.8 | 105.353 | 106.25 | 106.4 | 105.725 | 104.7 | 106.285 | 106.2 | 105.75 | 104.735 | 106.2 | 105.9667 | 105.525 | 104.6 | 105 | 105.2 | 104.85 | 104.298 | 104.525 | 104.595 | 104.3
-0.7 | 104.652 | 105.45 | 105.5333 | 105.15 | 104.4333 | 105.535 | 105.325 | 105.1 | 104.435 | 105.4 | 105.2 | 104.6667 | 104.0333 | 104.35 | 104.5 | 104.2 | 103.848 | 103.9333 | 103.9627 | 104.05
-0.6 | 104.252 | 104.6 | 104.5667 | 104.2333 | 103.9 | 104.735 | 104.6 | 104.3 | 103.835 | 104.6667 | 104.5 | 103.925 | 103.55 | 103.7 | 103.7 | 103.6 | 103.148 | 103.4333 | 103.397 | 103.55
-0.5 | 103.652 | 103.8 | 103.9 | 103.6 | 103.2 | 103.735 | 103.925 | 103.5333 | 103.234 | 103.9 | 103.55 | 103.25 | 103.05 | 103.12 | 103.2 | 103 | 102.649 | 102.8 | 102.7303 | 102.9667
-0.4 | 102.751 | 103.1333 | 103.1 | 102.8 | 102.55 | 103.034 | 103.025 | 102.9 | 102.534 | 103.15 | 103 | 102.6 | 102.45 | 102.475 | 102.325 | 102.4 | 102.149 | 102.2 | 102.098 | 102.1
-0.3 | 102.401 | 102.2 | 102.2 | 102.0667 | 101.9 | 102.234 | 102.2 | 102 | 101.884 | 102.3 | 102 | 101.9 | 101.7 | 101.725 | 101.65 | 101.8 | 101.849 | 101.8 | 101.4987 | 101.7
-0.2 | 101.351 | 101.35 | 101.6 | 101.325 | 101.35 | 101.584 | 101.42 | 101.3333 | 101.234 | 101.625 | 101.4667 | 101.1 | 101.05 | 101.05 | 101.05 | 101.1 | 101.149 | 101.175 | 100.949 | 101.1
-0.1 | 100.75 | 100.6 | 100.7 | 100.525 | 100.7 | 101.034 | 100.6333 | 100.7 | 100.834 | 100.925 | 100.8 | 100.6 | 100.7 | 100.55 | 100.46 | 100.55 | 100.45 | 100.65 | 100.45 | 100.5
0 | 100 | 100 | 100 | 100 | 100 | 99.9998 | 100 | 100 | 99.9998 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.3497 | 99.2 | 99.2667 | 99.3 | 99.1 | 99.2831 | 99.3 | 99.3333 | 99.4331 | 99.5 | 99.25 | 99.275 | 99.35 | 99.325 | 99.4 | 99.35 | 99.3504 | 99.375 | 99.2008 | 99.575
0.2 | 98.5493 | 98.55 | 98.575 | 98.55 | 98.5 | 98.6329 | 98.5 | 98.675 | 98.8329 | 99.05 | 98.825 | 98.6667 | 98.8 | 98.85 | 98.7 | 98.725 | 99.0005 | 98.725 | 98.7762 | 99.225
0.3 | 97.6988 | 97.975 | 97.9667 | 97.9 | 97.7 | 97.7659 | 97.9 | 98 | 98.0326 | 98.4 | 98 | 98 | 98.2 | 98.3 | 98.3 | 98.3 | 98.3508 | 98.08 | 98.2268 | 98.6
0.4 | 97.0485 | 97.125 | 97.1 | 97.25 | 97.1 | 97.2657 | 97.325 | 97.45 | 97.3324 | 97.65 | 97.475 | 97.4 | 97.7 | 97.4667 | 97.6 | 97.8 | 97.7511 | 97.65 | 97.5358 | 98.2
0.5 | 96.3482 | 96.45 | 96.7 | 96.7 | 96.3 | 96.5822 | 96.6333 | 96.875 | 96.6822 | 96.8 | 97 | 96.9667 | 97.2 | 96.875 | 97.2 | 97.2667 | 97.2514 | 97.1 | 97.2028 | 97.7
0.6 | 95.5478 | 95.8 | 95.95 | 96.225 | 95.3 | 95.9987 | 96.1 | 96.4 | 95.882 | 96.3 | 96.3 | 96.3 | 96.25 | 96.4 | 96.6333 | 96.7 | 96.6517 | 96.725 | 96.7532 | 97.2
0.7 | 94.7474 | 95.4 | 95.3 | 95.75 | 94.4 | 95.5318 | 95.6 | 95.7 | 95.0816 | 95.9333 | 95.75 | 95.7333 | 95.7 | 95.9 | 96.1333 | 96.2667 | 96.2519 | 96.2 | 96.2371 | 96.7
0.8 | 93.8469 | 94.5667 | 94.72 | 94.95 | 93.7 | 94.8316 | 94.85 | 95 | 94.0313 | 95.1 | 95.1 | 95.1667 | 95.05 | 95.4 | 95.6 | 95.8 | 95.3523 | 95.78 | 95.6377 | 96.2
0.9 | 92.9465 | 94 | 94.25 | 94.45 | 92.9 | 94.0813 | 94.4 | 94.4 | 93.2311 | 94.6 | 94.45 | 94.7 | 94.25 | 94.85 | 95.0333 | 95.3 | 94.8526 | 95.2 | 95.2548 | 95.7667
1 | 91.7459 | 93.3667 | 93.8333 | 93.8333 | 91.75 | 93.6312 | 93.62 | 93.9 | 92.4308 | 93.9 | 93.9 | 94.1 | 93.4 | 94.5 | 94.525 | 94.8 | 94.3029 | 94.8 | 94.8052 | 95.2667
1.1 | 90.6453 | 92.55 | 93.125 | 93.4 | 90.6 | 92.981 | 93.175 | 93.4333 | 91.2304 | 93.35 | 93.45 | 93.7 | 92.65 | 93.95 | 94 | 94.3 | 93.4033 | 94.3 | 94.4056 | 94.75
1.2 | 89.3947 | 92.16 | 92.45 | 92.7667 | 89.25 | 92.4308 | 92.5 | 93.025 | 90.13 | 93.025 | 92.9 | 93.16 | 91.5 | 93.375 | 93.3667 | 93.7667 | 92.8036 | 93.78 | 93.8062 | 94.525
1.3 | 87.8939 | 91.5667 | 91.875 | 92.25 | 87.9 | 91.8306 | 91.8667 | 92.25 | 88.5796 | 92.3 | 92.25 | 92.575 | 90.6 | 92.775 | 93 | 93.275 | 91.904 | 93.3 | 93.34 | 93.925
1.4 | 85.9429 | 91.1 | 91.45 | 91.6 | 86.2 | 91.2054 | 91.5 | 91.75 | 87.129 | 91.6667 | 91.8667 | 91.95 | 89.6 | 92.15 | 92.5333 | 92.8 | 91.1544 | 92.925 | 92.8571 | 93.525
1.5 | 83.1415 | 90.3667 | 90.9 | 91.0667 | 84.1 | 90.6052 | 90.84 | 91.15 | 85.4285 | 91.025 | 91.2667 | 91.45 | 88.3 | 91.6 | 91.925 | 92.4 | 89.8551 | 92.35 | 92.4825 | 92.975
1.6 | 79.99 | 89.7 | 90.1 | 90.5 | 81.1 | 90.18 | 90.3 | 90.5 | 83.2277 | 90.45 | 90.7 | 90.95 | 87.1 | 91.1 | 91.45 | 91.825 | 88.9888 | 91.75 | 92.0879 | 92.64
1.7 | 75.271 | 89.2 | 89.6 | 89.95 | 77.3 | 89.3548 | 89.75 | 90.1333 | 80.4268 | 90 | 90.2 | 90.275 | 85.1 | 90.5 | 91 | 91.4333 | 87.5562 | 91.35 | 91.5684 | 92.3
1.8 | 68.1841 | 88.7 | 89.0667 | 89.5 | 71.6 | 88.7045 | 89.1667 | 89.6333 | 76.2587 | 89.55 | 89.625 | 89.8333 | 82.8 | 89.95 | 90.4 | 91 | 85.7571 | 90.9 | 91.1089 | 91.9333
1.9 | 59.5298 | 88.15 | 88.7 | 88.9 | 63.4 | 88.4295 | 88.4 | 89.1 | 70.9236 | 88.92 | 89.2333 | 89.4 | 80.2 | 89.4 | 90 | 90.65 | 84.2579 | 90.6 | 90.7093 | 91.4
2 | 50.2251 | 87.45 | 88.0333 | 88.4 | 54.7 | 87.6792 | 87.9667 | 88.6 | 63.4211 | 88.275 | 88.4667 | 88.9667 | 76.2 | 88.9333 | 89.56 | 90 | 81.8591 | 89.9 | 90.2098 | 91
2.1 | 39.92 | 86.95 | 87.54 | 87.98 | 45.2 | 87.2291 | 87.5 | 87.95 | 54.8683 | 87.675 | 88 | 88.4667 | 71.4 | 88.375 | 89.06 | 89.55 | 79.3603 | 89.4333 | 89.7103 | 90.54
2.2 | 30.6153 | 86.35 | 87 | 87.4 | 36 | 86.7289 | 87.1333 | 87.4 | 45.2151 | 87.2 | 87.55 | 87.9667 | 63.8 | 87.95 | 88.55 | 89.2 | 75.5289 | 88.9 | 89.3706 | 90.15
2.3 | 22.7781 | 85.8 | 86.6 | 87.05 | 27.1 | 86.1287 | 86.475 | 86.9667 | 36.0787 | 86.525 | 87 | 87.6 | 56.3 | 87.275 | 88.05 | 88.725 | 70.6647 | 88.475 | 88.7779 | 89.65
2.4 | 16.5583 | 85.3 | 86 | 86.575 | 19.85 | 85.4284 | 86.0333 | 86.6 | 27.7092 | 86.05 | 86.6333 | 87 | 47.55 | 86.8667 | 87.65 | 88.3 | 63.5682 | 88.2 | 88.5115 | 89.35
2.5 | 12.4062 | 84.85 | 85.4 | 86.15 | 15.25 | 84.9683 | 85.6 | 86.0333 | 21.4071 | 85.5333 | 86 | 86.55 | 38.4 | 86.45 | 87 | 87.6667 | 55.922 | 87.5 | 88.0786 | 88.85
2.6 | 9.5548 | 84.4 | 84.95 | 85.5 | 11.9 | 84.4781 | 85.05 | 85.625 | 16.3054 | 84.85 | 85.4333 | 86.1 | 30.85 | 85.775 | 86.475 | 87.2333 | 47.2764 | 87 | 87.5458 | 88.5
2.7 | 7.6538 | 83.7 | 84.6 | 85.15 | 9.7 | 83.9947 | 84.6 | 85.0667 | 13.1044 | 84.55 | 84.975 | 85.6 | 23.8 | 85.36 | 86.225 | 86.9 | 38.6807 | 86.5 | 87.1129 | 88.1
2.8 | 6.2031 | 83.125 | 84.175 | 84.7 | 7.8 | 83.3945 | 84.1 | 84.75 | 10.8036 | 84.0667 | 84.6 | 85.2667 | 18.95 | 85 | 85.675 | 86.3667 | 31.6842 | 85.95 | 86.7799 | 87.6
2.9 | 5.4027 | 82.65 | 83.66 | 84.3 | 6.7 | 82.9526 | 83.7 | 84.225 | 9.003 | 83.5333 | 84.1 | 84.75 | 15.6 | 84.5 | 85.2333 | 86 | 24.9875 | 85.625 | 86.347 | 87.2
3 | 4.5023 | 82.1333 | 83.15 | 83.7667 | 5.7 | 82.4275 | 83.3333 | 83.7 | 7.8026 | 82.94 | 83.825 | 84.275 | 12.9 | 84 | 84.875 | 85.5667 | 19.99 | 85.18 | 85.8891 | 86.8667
3.1 | 4.002 | 81.8 | 82.7 | 83.275 | 4.9 | 82.0607 | 82.9333 | 83.4 | 6.7022 | 82.5 | 83.2 | 83.85 | 11 | 83.5 | 84.55 | 85.2 | 16.4918 | 84.65 | 85.6144 | 86.3
3.2 | 3.5017 | 81.1667 | 82.25 | 82.8667 | 4.3 | 81.5272 | 82.45 | 82.9 | 5.902 | 81.8 | 82.8333 | 83.375 | 9.65 | 83.0333 | 84.1 | 84.82 | 14.1929 | 84.25 | 85.1815 | 86
3.3 | 3.2016 | 80.7 | 81.8667 | 82.4333 | 3.9 | 80.977 | 82 | 82.5 | 5.4018 | 81.48 | 82.35 | 83.075 | 8.55 | 82.525 | 83.6667 | 84.4 | 11.8941 | 83.75 | 84.6154 | 85.55
3.4 | 2.9348 | 80.1333 | 81.3 | 82.0333 | 3.45 | 80.2267 | 81.55 | 82.0667 | 4.8016 | 80.95 | 82.05 | 82.65 | 7.7 | 81.8667 | 83 | 83.9 | 10.8446 | 83.06 | 84.2657 | 85.175
3.5 | 2.7014 | 79.3333 | 80.9 | 81.5667 | 3.1667 | 79.6266 | 81 | 81.6667 | 4.4015 | 80.45 | 81.5333 | 82.2333 | 6.9 | 81.5 | 82.7333 | 83.6 | 9.7451 | 82.725 | 83.8162 | 84.95
3.6 | 2.5013 | 78.825 | 80.4 | 81.1 | 2.85 | 79.0263 | 80.475 | 81.2 | 4.0013 | 79.775 | 81.1 | 81.7333 | 6.3 | 80.92 | 82.3 | 83.275 | 8.5957 | 82.4 | 83.4566 | 84.6
3.7 | 2.3011 | 78.275 | 79.9 | 80.6333 | 2.65 | 78.4261 | 80.05 | 80.6667 | 3.7012 | 79.05 | 80.6 | 81.3667 | 5.95 | 80.54 | 81.95 | 82.75 | 7.946 | 81.95 | 83.1918 | 84.15
3.8 | 2.2011 | 77.6667 | 79.45 | 80.125 | 2.4667 | 77.7926 | 79.575 | 80.3333 | 3.5012 | 78.65 | 80.2 | 80.85 | 5.5 | 80.0667 | 81.45 | 82.2 | 7.4296 | 81.5 | 82.8672 | 83.8
3.9 | 2.101 | 76.9 | 79 | 79.75 | 2.3 | 77.0924 | 79.0333 | 79.7667 | 3.2511 | 78 | 79.75 | 80.5 | 5.2 | 79.375 | 81.1667 | 81.925 | 6.8966 | 81 | 82.2511 | 83.375
4 | 2.001 | 76.1 | 78.45 | 79.175 | 2.2 | 76.3255 | 78.65 | 79.4 | 3.001 | 77.25 | 79.3 | 79.98 | 5 | 78.8 | 80.575 | 81.425 | 6.4468 | 80.5 | 81.9431 | 82.9667
4.1 | 1.9009 | 75.35 | 78.1 | 78.9 | 2 | 75.4251 | 78.3 | 78.95 | 2.901 | 76.5 | 78.75 | 79.58 | 4.5333 | 78.1667 | 80.25 | 80.925 | 5.997 | 80.25 | 81.3936 | 82.55
4.2 | 1.8009 | 74.4 | 77.7333 | 78.45 | 1.9 | 74.6499 | 77.7 | 78.45 | 2.6509 | 75.975 | 78.425 | 79.2 | 4.35 | 77.8 | 79.94 | 80.65 | 5.6971 | 79.75 | 81.1522 | 82.15
4.3 | 1.7008 | 73.45 | 77.175 | 78 | 1.9 | 73.7996 | 77.4 | 77.95 | 2.5008 | 75.15 | 78.025 | 78.6333 | 4.1 | 76.95 | 79.4 | 80.24 | 5.3973 | 79.1 | 80.7193 | 81.675
4.4 | 1.6508 | 72.3 | 76.68 | 77.5 | 1.8 | 72.5492 | 76.9 | 77.625 | 2.4008 | 74.3 | 77.4333 | 78.12 | 4 | 76.4 | 79 | 79.8 | 5.1974 | 78.65 | 80.1948 | 81.38
4.5 | 1.6008 | 70.65 | 76.2667 | 77.05 | 1.7 | 71.2237 | 76.5 | 77.2 | 2.4008 | 73.5 | 77.06 | 77.6667 | 3.8 | 75.85 | 78.5333 | 79.4 | 4.9975 | 78.05 | 79.8202 | 81
4.6 | 1.5008 | 68.78 | 75.8 | 76.58 | 1.7 | 69.99 | 75.9667 | 76.7 | 2.3008 | 72.2 | 76.65 | 77.325 | 3.6 | 75.1 | 78.0667 | 78.9333 | 4.8975 | 77.65 | 79.5954 | 80.7
4.7 | 1.5008 | 66.425 | 75.4 | 76.2 | 1.6 | 67.9226 | 75.475 | 76.175 | 2.2007 | 71.15 | 76.1 | 76.9 | 3.5 | 74.6 | 77.6 | 78.6 | 4.6977 | 76.9 | 79.3207 | 80.2
4.8 | 1.5008 | 62.725 | 74.95 | 75.7667 | 1.5 | 65.6219 | 75.0333 | 75.85 | 2.1007 | 69.8 | 75.7 | 76.35 | 3.4 | 73.7 | 77.2333 | 78.2 | 4.5977 | 76.3 | 78.8212 | 79.9
4.9 | 1.4007 | 57.55 | 74.4 | 75.2 | 1.5 | 61.8706 | 74.6 | 75.35 | 2.0007 | 67.8 | 75.15 | 75.88 | 3.3 | 72.875 | 76.725 | 77.75 | 4.3978 | 75.75 | 78.2218 | 79.38
5 | 1.4007 | 51.05 | 73.95 | 74.7667 | 1.45 | 56.919 | 74.1333 | 74.9333 | 1.9006 | 65.96 | 74.85 | 75.55 | 3.2 | 72.075 | 76.325 | 77.2 | 4.1979 | 75.2 | 77.9221 | 78.95
5.1 | 1.3007 | 43.65 | 73.4667 | 74.4 | 1.4 | 50.4168 | 73.6667 | 74.45 | 1.9006 | 63.3 | 74.4 | 75.1 | 3.1 | 71 | 75.9333 | 76.9 | 4.0646 | 74.7 | 77.5725 | 78.7333
5.2 | 1.3007 | 36.0333 | 73.1 | 73.9333 | 1.3 | 43.3144 | 73.2667 | 74 | 1.8006 | 59.025 | 73.9 | 74.7 | 3 | 69.8 | 75.45 | 76.5333 | 3.8981 | 73.9 | 77.356 | 78.1
5.3 | 1.3007 | 29 | 72.6 | 73.45 | 1.3 | 36.1721 | 72.75 | 73.5333 | 1.7006 | 53.55 | 73.6 | 74.2 | 2.9 | 68.5 | 74.95 | 76.02 | 3.7981 | 73.02 | 76.9731 | 77.8
5.4 | 1.3007 | 22.075 | 72.0667 | 72.9333 | 1.3 | 29.3348 | 72.3 | 73.12 | 1.7006 | 47.0667 | 73.0667 | 73.7 | 2.8 | 66.925 | 74.4667 | 75.5 | 3.7981 | 72.175 | 76.3986 | 77.4333
5.5 | 1.2006 | 17.6 | 71.75 | 72.45 | 1.2 | 23.2077 | 71.8 | 72.625 | 1.7006 | 40.3 | 72.6 | 73.35 | 2.7 | 64.75 | 74.15 | 75.05 | 3.5982 | 71.325 | 75.9241 | 76.9
5.6 | 1.2006 | 14.15 | 71.15 | 71.95 | 1.2 | 18.031 | 71.25 | 72.025 | 1.6005 | 33.4 | 72.05 | 72.875 | 2.6 | 61.9667 | 73.6333 | 74.6667 | 3.3983 | 70.425 | 75.4579 | 76.5
5.7 | 1.2006 | 11.85 | 70.55 | 71.4667 | 1.2 | 14.7299 | 70.9 | 71.56 | 1.6005 | 27.15 | 71.725 | 72.525 | 2.6 | 58.7 | 73.2667 | 74.26 | 3.3983 | 69.3 | 74.975 | 76.125
5.8 | 1.1005 | 10.2 | 69.9667 | 71.0667 | 1.1 | 12.4042 | 70.32 | 71.18 | 1.5005 | 21.8 | 71.125 | 71.975 | 2.5 | 53.8 | 72.8 | 73.85 | 3.1984 | 67.775 | 74.5755 | 75.7333
5.9 | 1.1005 | 9.1 | 69.6 | 70.575 | 1.1 | 10.9036 | 69.85 | 70.7 | 1.4505 | 18 | 70.65 | 71.625 | 2.4 | 48.1333 | 72.4 | 73.35 | 3.1984 | 66.4 | 74.2508 | 75.3
6 | 1.1005 | 8.3 | 69.1 | 70 | 1.1 | 9.7032 | 69.3333 | 70.2 | 1.4005 | 15.1 | 70.1 | 70.9 | 2.4 | 41.5667 | 71.925 | 72.9 | 3.0985 | 64.24 | 73.7763 | 75
6.1 | 1.1005 | 7.575 | 68.5 | 69.45 | 1.1 | 8.6279 | 68.65 | 69.6 | 1.4005 | 13.2 | 69.55 | 70.4 | 2.3 | 35.2 | 71.5 | 72.6333 | 3.0985 | 61.85 | 73.3766 | 74.45
6.2 | 1.1005 | 7.05 | 67.9 | 68.9 | 1.1 | 7.9026 | 68.125 | 69.2333 | 1.3004 | 11.8 | 69.1 | 69.85 | 2.2 | 29.36 | 71 | 72 | 3.0985 | 58.6 | 73.027 | 74.2
6.3 | 1.1005 | 6.625 | 67.4 | 68.4333 | 1 | 7.3024 | 67.575 | 68.5 | 1.3004 | 10.7 | 68.5333 | 69.44 | 2.2 | 24.625 | 70.7 | 71.6 | 2.9985 | 53.8 | 72.4276 | 73.75
6.4 | 1.0505 | 6.275 | 66.88 | 67.9 | 1 | 6.8023 | 66.975 | 68.025 | 1.3004 | 9.8333 | 68 | 68.9333 | 2.2 | 20.6 | 70.2 | 71.2 | 2.8986 | 48.6667 | 72.1279 | 73.4
6.5 | 1.0005 | 6.025 | 66.175 | 67.2 | 1 | 6.4021 | 66.4 | 67.375 | 1.3004 | 9.05 | 67.5 | 68.4 | 2.2 | 17.8 | 69.8 | 70.8 | 2.8986 | 42.6 | 71.8282 | 73.1667
6.6 | 1.0005 | 5.725 | 65.6 | 66.4667 | 1 | 6.002 | 65.8 | 66.725 | 1.3004 | 8.55 | 67.0333 | 67.8 | 2.1 | 15.8 | 69.22 | 70.25 | 2.7986 | 36.4 | 71.3287 | 72.64
6.7 | 1.0005 | 5.5 | 64.775 | 65.9 | 1 | 5.7519 | 65.1 | 66.075 | 1.2004 | 7.85 | 66.4 | 67.3 | 2.1 | 14.425 | 68.7667 | 69.7333 | 2.7986 | 31.4 | 70.9791 | 72.125
6.8 | 1.0005 | 5.3 | 64.05 | 65.1 | 1 | 5.5018 | 64.4333 | 65.3333 | 1.2004 | 7.4667 | 65.9 | 66.7 | 2 | 13.1 | 68.3 | 69.2 | 2.6987 | 26.5 | 70.6044 | 71.7333
6.9 | 0.90045 | 5.125 | 63.1 | 64.2 | 0.9 | 5.2017 | 63.55 | 64.525 | 1.1004 | 7.2 | 65.175 | 66.15 | 1.9667 | 12.25 | 67.7667 | 68.7 | 2.4987 | 22.625 | 70.03 | 71.3667
7 | 0.90045 | 5 | 62.025 | 63.2 | 0.9 | 5.1017 | 62.6 | 63.7 | 1.1004 | 6.85 | 64.525 | 65.475 | 1.9 | 11.6 | 67.1667 | 68.225 | 2.4987 | 19.7 | 69.4806 | 70.8667
7.1 |  | 4.8333 | 60.8 | 61.9333 |  | 4.9016 | 61.7 | 62.675 | 1.1004 | 6.6 | 63.775 | 64.9 | 1.8 | 10.9 | 66.75 | 67.75 | 2.3988 | 17.775 | 69.011 | 70.4667
7.2 |  | 4.75 | 59.3 | 60.35 |  | 4.7349 | 60.3 | 61.5 | 1.1004 | 6.4 | 62.95 | 64.075 | 1.8 | 10.4 | 66.25 | 67.175 | 2.2988 | 16.3 | 68.5315 | 70
7.3 |  | 4.6 | 57.2 | 58.225 |  | 4.6015 | 58.9 | 60.05 |  | 6.1 | 62.2 | 63.12 | 1.7 | 9.9667 | 65.65 | 66.675 | 2.2988 | 15.1 | 68.2318 | 69.425
7.4 |  | 4.4667 | 54 | 55.2 |  | 4.4015 | 57.1 | 58.3 |  | 5.9 | 61.4 | 62.4 | 1.7 | 9.5 | 65.075 | 66.15 | 2.2988 | 14.2 | 67.8322 | 69
7.5 |  | 4.3333 | 49.4 | 50.95 |  | 4.3014 | 54.4333 | 56 |  | 5.7667 | 60.3 | 61.4 |  | 9.1667 | 64.4 | 65.7 | 2.2988 | 13.4 | 67.1662 | 68.5
7.6 |  | 4.3 | 44.1 | 46.15 |  | 4.1514 | 50.9667 | 52.5 |  | 5.5333 | 59.2 | 60.1 |  | 8.85 | 63.875 | 65.1 | 2.2988 | 12.6 | 66.6667 | 68.2
7.7 |  | 4.2 | 37.85 | 40.25 |  | 4.0514 | 46.2333 | 48 |  | 5.4 | 57.55 | 58.5 |  | 8.5 | 63.26 | 64.3 |  | 12.1 | 66.3337 | 67.6
7.8 |  | 4.1667 | 32.025 | 34.225 |  | 4.0013 | 40.4667 | 42.8 |  | 5.2 | 55.6 | 56.8 |  | 8.2 | 62.7 | 63.85 |  | 11.7 | 65.8742 | 67.2
7.9 |  | 4.025 | 26.35 | 29 |  | 3.9013 | 34.3667 | 36.9667 |  | 5.1 | 53.1 | 54.1 |  | 7.98 | 61.95 | 63.2333 |  | 11.3333 | 65.2848 | 66.65
8 |  | 4 | 21.1667 | 24.1 |  | 3.8013 | 28.8 | 31.7 |  | 5 | 49.15 | 50.5 |  | 7.74 | 61.25 | 62.4 |  | 11 | 64.702 | 66
8.1 |  | 3.9 | 17.1 | 20.6 |  | 3.7012 | 23.45 | 26.4 |  | 4.825 | 44.7667 | 46.05 |  | 7.6 | 60.35 | 61.7 |  | 10.65 | 64.2158 | 65.45
8.2 |  | 3.8333 | 14.6 | 18.15 |  | 3.7012 | 19 | 22.3 |  | 4.7 | 39.425 | 40.7 |  | 7.3 | 59.4 | 60.8 |  | 10.34 | 63.8362 | 65.1333
8.3 |  | 3.8 | 12.85 | 16.55 |  | 3.6012 | 15.85 | 19.1667 |  | 4.6 | 33.4333 | 35.575 |  | 7.15 | 58.2 | 59.6 |  | 10.04 | 63.1702 | 64.5
8.4 |  | 3.8 | 11.6 | 15.4 |  | 3.6012 | 13.6667 | 17 |  | 4.5 | 27.9 | 30.54 |  | 7.06 | 56.9 | 58.5 |  | 9.8 | 62.6374 | 64.0333
8.5 |  | 3.8 | 10.82 | 14.5667 |  | 3.5012 | 12.2 | 15.575 |  | 4.4 | 23.3 | 26.075 |  | 6.825 | 55.35 | 56.85 |  | 9.5 | 61.9714 | 63.45
8.6 |  | 3.7 | 10.0667 | 13.95 |  | 3.5012 | 11.1667 | 14.56 |  | 4.3 | 19.5667 | 22.4 |  | 6.65 | 53.2 | 54.85 |  | 9.3 | 61.3387 | 62.725
8.7 |  | 3.7 | 9.5667 | 13.4333 |  | 3.5012 | 10.2667 | 13.8 |  | 4.3 | 16.5 | 19.7 |  | 6.5 | 50.375 | 52.12 |  | 8.92 | 60.6144 | 62.0333
8.8 |  | 3.7 | 9.1667 | 13.0667 |  | 3.5012 | 9.6667 | 13.1667 |  | 4.3 | 14.55 | 17.95 |  | 6.4 | 46.5 | 48.5667 |  | 8.7 | 59.8735 | 61.3
8.9 |  | 3.6333 | 8.85 | 12.7333 |  | 3.4261 | 9.1333 | 12.6333 |  | 4.2 | 13.3 | 16.625 |  | 6.2667 | 42.1 | 44.6 |  | 8.6 | 58.9411 | 60.6333
9 |  | 3.6 | 8.6 | 12.4 |  | 3.4011 | 8.725 | 12.1667 |  | 4.2 | 12.3333 | 15.7 |  | 6.125 | 37.025 | 39.7 |  | 8.4 | 57.992 | 59.5333
9.1 |  | 3.6 | 8.4 | 12.2 |  | 3.4011 | 8.4333 | 11.9 |  | 4.1 | 11.525 | 14.95 |  | 6 | 32.2 | 35 |  | 8.3 | 57.1096 | 58.4
9.2 |  | 3.5 | 8.2333 | 12.1 |  | 3.3511 | 8.175 | 11.625 |  | 4 | 10.8333 | 14.3667 |  | 5.825 | 27.6 | 30.5 |  | 8.1333 | 55.4196 | 57.15
9.3 |  | 3.5 | 8.1 | 11.9 |  | 3.3011 | 7.9 | 11.4 |  | 4 | 10.4 | 13.8667 |  | 5.7 | 23.4 | 27.1 |  | 7.95 | 53.7963 | 55.3
9.4 |  | 3.5 | 8 | 11.8 |  | 3.3011 | 7.7 | 11.1333 |  | 4 | 9.9333 | 13.3667 |  | 5.6 | 20.1 | 23.95 |  | 7.7 | 51.3487 | 53.4
9.5 |  | 3.4 | 7.9 | 11.725 |  | 3.2011 | 7.55 | 11 |  | 3.9 | 9.6 | 13 |  | 5.5 | 18.1 | 21.8333 |  | 7.55 | 48.4515 | 50.5
9.6 |  | 3.4 | 7.9 | 11.7 |  | 3.2011 | 7.4 | 10.8667 |  | 3.9 | 9.2667 | 12.7667 |  | 5.45 | 16.5 | 20.24 |  | 7.4 | 44.6553 | 46.975
9.7 |  | 3.375 | 7.8 | 11.6 |  | 3.1511 | 7.3 | 10.7 |  | 3.8 | 9 | 12.45 |  | 5.4 | 15.3333 | 19.1 |  | 7.4 | 40.4095 | 43.05
9.8 |  | 3.3 | 7.8 | 11.575 |  | 3.101 | 7.3 | 10.6333 |  | 3.7 | 8.7 | 12.2 |  | 5.3 | 14.375 | 18.275 |  | 7.275 | 35.631 | 38.675
9.9 |  | 3.3 | 7.7 | 11.5 |  | 3.001 | 7.2 | 10.6 |  | 3.6667 | 8.6 | 12.025 |  | 5.3 | 13.65 | 17.525 |  | 7.025 | 31.2188 | 34.4667
10 |  | 3.25 | 7.7 | 11.5 |  | 3.001 | 7.1667 | 10.5 |  | 3.6 | 8.4 | 11.875 |  | 5.225 | 13.05 | 16.925 |  | 6.9 | 27.1728 | 30.55
10.1 |  | 3.2 | 7.7 | 11.4 |  | 3.001 | 7.1 | 10.5 |  | 3.6 | 8.3 | 11.725 |  | 5.2 | 12.5333 | 16.5333 |  | 6.7333 | 23.5564 | 27.4667
10.2 |  |  | 7.6 | 11.4 |  | 2.901 | 7 | 10.475 |  | 3.5 | 8.1 | 11.525 |  | 5.1 | 12.0333 | 16.1 |  | 6.6 | 21.1588 | 25.1
10.3 |  |  | 7.6 | 11.3 |  |  | 7 | 10.4 |  | 3.5 | 8 | 11.475 |  | 5 | 11.6667 | 15.65 |  | 6.5 | 19.3307 | 23.1
10.4 |  |  | 7.5 | 11.3 |  |  | 6.9 | 10.4 |  | 3.4 | 7.9 | 11.35 |  | 5 | 11.3333 | 15.3 |  | 6.4667 | 17.982 | 21.96
10.5 |  |  | 7.5 | 11.2 |  |  | 6.875 | 10.3 |  |  | 7.8 | 11.3 |  | 4.8 | 11.0333 | 15 |  | 6.3 | 16.9081 | 20.94
10.6 |  |  | 7.4 | 11.1 |  |  | 6.8 | 10.2 |  |  | 7.75 | 11.15 |  | 4.8 | 10.7 | 14.75 |  | 6.24 | 15.8508 | 20.15
10.7 |  |  | 7.375 | 11.05 |  |  | 6.7333 | 10.2 |  |  | 7.6 | 11.1 |  | 4.8 | 10.5667 | 14.5 |  | 6.2 | 15.2348 | 19.5
10.8 |  |  | 7.275 | 10.95 |  |  | 6.7 | 10.1 |  |  | 7.6 | 11 |  | 4.7 | 10.4 | 14.25 |  | 6.1 | 14.6853 | 18.9
10.9 |  |  | 7.1667 | 10.9 |  |  | 6.6 | 10 |  |  | 7.5 | 10.9 |  | 4.6 | 10.2 | 14.025 |  | 6.1 | 14.2524 | 18.44
11 |  |  | 7.1 | 10.8 |  |  | 6.6 | 9.9 |  |  | 7.4 | 10.8 |  |  | 10 | 13.9 |  | 6 | 13.8528 | 18.06
11.1 |  |  | 7 | 10.7 |  |  | 6.5 | 9.875 |  |  | 7.4 | 10.76 |  |  | 9.875 | 13.7 |  | 5.9 | 13.4366 | 17.725
11.2 |  |  | 7 | 10.6 |  |  | 6.4 | 9.7 |  |  | 7.3 | 10.7 |  |  | 9.7 | 13.5 |  | 5.9 | 13.0869 | 17.375
11.3 |  |  | 6.9 | 10.5 |  |  | 6.35 | 9.7 |  |  | 7.2 | 10.6 |  |  | 9.6 | 13.3667 |  | 5.9 | 12.8538 | 17.05
11.4 |  |  | 6.8 | 10.4 |  |  | 6.3 | 9.575 |  |  | 7.1 | 10.5 |  |  | 9.4 | 13.2 |  | 5.8 | 12.6374 | 16.7
11.5 |  |  | 6.7 | 10.3 |  |  | 6.2 | 9.5 |  |  | 7 | 10.4 |  |  | 9.3 | 13.1 |  |  | 12.2877 | 16.5
11.6 |  |  | 6.6 | 10.2 |  |  | 6.1 | 9.4 |  |  | 7 | 10.3 |  |  | 9.2 | 12.925 |  |  | 12.0629 | 16.2333
11.7 |  |  | 6.5 | 10.1 |  |  | 6 | 9.3 |  |  | 6.9 | 10.2 |  |  | 9.1 | 12.9 |  |  | 11.8548 | 16.1
11.8 |  |  | 6.425 | 10 |  |  | 5.9 | 9.2 |  |  | 6.8 | 10.15 |  |  | 8.975 | 12.8 |  |  | 11.6383 | 15.85
11.9 |  |  | 6.3 | 9.9 |  |  | 5.875 | 9.1 |  |  | 6.7 | 10 |  |  | 8.825 | 12.7 |  |  | 11.3886 | 15.65
12 |  |  | 6.25 | 9.8 |  |  | 5.775 | 9 |  |  | 6.6 | 9.9 |  |  | 8.7667 | 12.5333 |  |  | 11.2887 | 15.45
12.1 |  |  | 6.1 | 9.7 |  |  | 5.7 | 8.9 |  |  | 6.5 | 9.8 |  |  | 8.6 | 12.4 |  |  | 11.0556 | 15.3333
12.2 |  |  | 6 | 9.525 |  |  | 5.6 | 8.8 |  |  | 6.4333 | 9.8 |  |  | 8.525 | 12.3 |  |  | 10.8891 | 15.175
12.3 |  |  | 5.975 | 9.4 |  |  | 5.5 | 8.7 |  |  | 6.3 | 9.6 |  |  | 8.4667 | 12.1 |  |  | 10.7892 | 14.875
12.4 |  |  | 5.9 | 9.375 |  |  | 5.4 | 8.575 |  |  | 6.2 | 9.5 |  |  | 8.4 | 12 |  |  | 10.656 | 14.775
12.5 |  |  | 5.8 | 9.2 |  |  | 5.35 | 8.5 |  |  | 6.1 | 9.4 |  |  | 8.2 | 11.9 |  |  | 10.5894 | 14.5667
12.6 |  |  | 5.7 | 9.2 |  |  | 5.3 | 8.4 |  |  | 6.06 | 9.3 |  |  | 8.1 | 11.8 |  |  | 10.4562 | 14.4333
12.7 |  |  | 5.7 | 9 |  |  | 5.2 | 8.3 |  |  | 6 | 9.225 |  |  | 8.1 | 11.7 |  |  | 10.2897 | 14.3
12.8 | [6 empty] | 5.1 | 8.225 |  |  | 5.9 | 9.1 |  |  | 7.9 | 11.6 |  |  | 10.1898 | 14.2
12.9 | [10 empty] | 5.85 | 9 |  |  | 7.85 | 11.5 |  |  | 9.99 | 14.05
13 | [10 empty] | 5.8 | 8.9 |  |  | 7.8 | 11.4 |  |  | 9.8901 | 14
13.1 | [10 empty] | 5.7 | 8.8 |  |  | 7.7 | 11.2667 |  |  | 9.7902 | 13.76
13.2 | [10 empty] | 5.6 | 8.7 |  |  | 7.5333 | 11.175 |  |  | 9.6903 | 13.7
13.3 | [14 empty] | 7.4 | 11.025 |  |  | 9.5904 | 13.5
13.4 | [14 empty] | 7.3 | 10.9 |  |  | 9.3906 | 13.5
13.5 | [14 empty] | 7.2 | 10.8 |  |  | 9.2907 | 13.3
13.6 | [14 empty] | 7.2 | 10.6 |  |  | 9.1908 | 13.2
13.7 | [14 empty] | 7.1 | 10.5 |  |  | 9.0909 | 13.0667
13.8 | [14 empty] | 7 | 10.4333 |  |  | 9.0909 | 12.9
13.9 | [14 empty] | 6.9 | 10.3 |  |  | 8.991 | 12.75
14 | [18 empty] | 8.8911 | 12.65
14.1 | [18 empty] | 8.7912 | 12.6
14.2 | [18 empty] | 8.5914 | 12.4
14.3 | [18 empty] | 8.4915 | 12.3
14.4 | [18 empty] | 8.3916 | 12.2
14.5 | [18 empty] | 8.2917 | 12.1
14.6 | [18 empty] | 8.1918 | 12
14.7 | [18 empty] | 8.0919 | 11.825

## Wedge 60 Longitudinal Profiles
Longitudinal Profiles
Wedge 60
SSD=100cm
Field Size=40x15cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:3 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-30.9 |  |  |  |  | 10.46
-30.8 |  |  |  |  | 10.6
-30.7 |  |  |  |  | 10.7
-30.6 |  |  |  |  | 10.7571
-30.5 |  |  |  |  | 10.9
-30.4 |  |  |  |  | 10.95
-30.3 |  |  |  |  | 11.1
-30.2 |  |  |  |  | 11.2
-30.1 |  |  |  |  | 11.3
-30 |  |  |  |  | 11.5333
-29.9 |  |  |  |  | 11.6
-29.8 |  |  |  |  | 11.78
-29.7 |  |  |  |  | 11.86
-29.6 |  |  |  |  | 12.04
-29.5 |  |  |  |  | 12.2
-29.4 |  |  |  |  | 12.3429
-29.3 |  |  |  |  | 12.5167
-29.2 |  |  |  |  | 12.7
-29.1 |  |  |  |  | 12.8667
-29 |  |  |  |  | 13.0167
-28.9 |  |  |  | 9.1143 | 13.2
-28.8 |  |  |  | 9.2444 | 13.4
-28.7 |  |  |  | 9.3625 | 13.55
-28.6 |  |  |  | 9.4 | 13.8286
-28.5 |  |  |  | 9.5667 | 14.0667
-28.4 |  |  |  | 9.6333 | 14.24
-28.3 |  |  |  | 9.775 | 14.4286
-28.2 |  |  |  | 9.9 | 14.6571
-28.1 |  |  |  | 9.9429 | 14.9
-28 |  |  |  | 10.0857 | 15.15
-27.9 |  |  |  | 10.16 | 15.3889
-27.8 |  |  |  | 10.3 | 15.625
-27.7 |  |  |  | 10.48 | 15.875
-27.6 |  |  |  | 10.6 | 16.2375
-27.5 |  |  |  | 10.7 | 16.575
-27.4 |  |  |  | 10.85 | 16.8875
-27.3 |  |  |  | 11.0667 | 17.4375
-27.2 |  |  |  | 11.2125 | 18.0375
-27.1 |  |  |  | 11.3429 | 18.5
-27 |  |  |  | 11.5 | 19.2
-26.9 |  |  | 7.9 | 11.68 | 19.95
-26.8 |  |  | 7.95 | 11.8571 | 21
-26.7 |  |  | 8 | 12.1 | 22.05
-26.6 |  |  | 8.1 | 12.2667 | 23.5
-26.5 |  |  | 8.15 | 12.4222 | 25.2667
-26.4 |  |  | 8.2 | 12.6571 | 27.4625
-26.3 |  |  | 8.2333 | 12.8833 | 30.5625
-26.2 |  |  | 8.3 | 13.05 | 34.375
-26.1 |  |  | 8.4167 | 13.2333 | 38.875
-26 |  |  | 8.5 | 13.4333 | 43.975
-25.9 |  | 7.4 | 8.5875 | 13.7667 | 49.5
-25.8 |  | 7.4 | 8.7 | 13.95 | 55.3875
-25.7 |  | 7.5 | 8.8 | 14.2333 | 60.825
-25.6 |  | 7.6 | 8.8429 | 14.6333 | 65.4625
-25.5 | 7.7167 | 7.6 | 8.9 | 14.925 | 69.6375
-25.4 | 7.8 | 7.7 | 9.05 | 15.35 | 72.7333
-25.3 | 7.8333 | 7.7 | 9.1143 | 15.8 | 74.9
-25.2 | 7.9 | 7.7 | 9.2 | 16.2714 | 76.8
-25.1 | 8 | 7.8 | 9.3167 | 16.94 | 78.3
-25 | 8 | 7.8125 | 9.4833 | 17.54 | 79.5
-24.9 | 8.0143 | 7.9 | 9.65 | 18.45 | 80.4
-24.8 | 8.1 | 7.9 | 9.725 | 19.5 | 81.225
-24.7 | 8.1167 | 8 | 9.9375 | 20.725 | 82
-24.6 | 8.2 | 8.05 | 10.0833 | 22.4 | 82.5
-24.5 | 8.3 | 8.1 | 10.2125 | 24.45 | 83.0286
-24.4 | 8.3 | 8.1857 | 10.35 | 26.9444 | 83.6
-24.3 | 8.4 | 8.2429 | 10.575 | 30.525 | 84
-24.2 | 8.4 | 8.3 | 10.7286 | 35.0375 | 84.5
-24.1 | 8.5 | 8.4 | 11 | 40.35 | 84.84
-24 | 8.5286 | 8.5 | 11.15 | 46.375 | 85.2
-23.9 | 8.6 | 8.5714 | 11.4143 | 52.75 | 85.575
-23.8 | 8.68 | 8.7 | 11.6333 | 58.8333 | 85.95
-23.7 | 8.7429 | 8.7714 | 11.925 | 64.7 | 86.2714
-23.6 | 8.8 | 8.96 | 12.2 | 69.7333 | 86.5286
-23.5 | 8.9 | 9.0429 | 12.5125 | 73.8286 | 86.8333
-23.4 | 9 | 9.2143 | 12.9 | 77.2 | 87.0571
-23.3 | 9.1 | 9.3667 | 13.3 | 79.7 | 87.3429
-23.2 | 9.2286 | 9.52 | 13.8 | 81.4778 | 87.5571
-23.1 | 9.3 | 9.7 | 14.4 | 82.9143 | 87.8
-23 | 9.45 | 9.88 | 15 | 83.9875 | 88.05
-22.9 | 9.5333 | 10.05 | 15.9571 | 84.8625 | 88.2667
-22.8 | 9.7 | 10.3 | 16.975 | 85.55 | 88.4167
-22.7 | 9.86 | 10.4667 | 18.22 | 86.3125 | 88.5833
-22.6 | 10.04 | 10.7 | 19.9 | 86.86 | 88.7375
-22.5 | 10.24 | 10.94 | 22.125 | 87.4429 | 88.925
-22.4 | 10.4167 | 11.3571 | 25.22 | 87.85 | 89.0778
-22.3 | 10.6667 | 11.75 | 29.1 | 88.2625 | 89.225
-22.2 | 10.8667 | 12.15 | 34.2714 | 88.7857 | 89.3667
-22.1 | 11.0667 | 12.7333 | 40.4429 | 89.125 | 89.5
-22 | 11.3429 | 13.425 | 47.5 | 89.3429 | 89.7125
-21.9 | 11.76 | 14.3 | 55.2 | 89.7714 | 89.875
-21.8 | 12.1857 | 15.3143 | 62.525 | 90.1167 | 90.125
-21.7 | 12.65 | 16.6667 | 69.6 | 90.3429 | 90.2875
-21.6 | 13.25 | 18.45 | 75.3 | 90.65 | 90.4333
-21.5 | 14.1429 | 20.9333 | 79.625 | 90.8667 | 90.725
-21.4 | 15.0625 | 24.4143 | 82.9 | 91.075 | 90.9
-21.3 | 16.45 | 28.9 | 85.1714 | 91.2625 | 91.0857
-21.2 | 18.1 | 34.5 | 86.8 | 91.5 | 91.25
-21.1 | 20.5333 | 41.4 | 88.3 | 91.5714 | 91.45
-21 | 23.92 | 49.1571 | 89.3 | 91.7143 | 91.5667
-20.9 | 28.5375 | 57.3 | 90.175 | 91.9143 | 91.7
-20.8 | 34.65 | 65.6429 | 90.8714 | 92.0333 | 91.9667
-20.7 | 41.9143 | 73.0429 | 91.4 | 92.1333 | 92.225
-20.6 | 50.4 | 78.9833 | 92.0429 | 92.3 | 92.3444
-20.5 | 58.6667 | 83.3286 | 92.7 | 92.4429 | 92.525
-20.4 | 66.7333 | 86.5 | 93.0857 | 92.6714 | 92.6875
-20.3 | 74.4 | 88.8429 | 93.3 | 92.8429 | 92.8143
-20.2 | 80.4714 | 90.3667 | 93.6429 | 93.05 | 92.9667
-20.1 | 85.2 | 91.6 | 94.0333 | 93.2 | 93.1143
-20 | 88.2667 | 92.6 | 94.3 | 93.35 | 93.3
-19.9 | 90.2857 | 93.4 | 94.55 | 93.6 | 93.475
-19.8 | 91.7857 | 93.94 | 94.7333 | 93.7857 | 93.6
-19.7 | 92.84 | 94.5667 | 94.9 | 93.925 | 93.725
-19.6 | 93.78 | 95.04 | 95.18 | 94.1429 | 93.85
-19.5 | 94.4429 | 95.4571 | 95.3 | 94.2857 | 93.9667
-19.4 | 95.1286 | 95.8375 | 95.4 | 94.5286 | 94.175
-19.3 | 95.6 | 96.175 | 95.5 | 94.6 | 94.3125
-19.2 | 95.88 | 96.4667 | 95.55 | 94.6875 | 94.475
-19.1 | 96.25 | 96.7571 | 95.675 | 94.9111 | 94.6556
-19 | 96.62 | 96.9167 | 95.8 | 95.0286 | 94.7857
-18.9 | 96.775 | 97 | 95.9333 | 95.1833 | 94.8
-18.8 | 96.9571 | 97.1 | 96 | 95.325 | 94.8571
-18.7 | 97.1833 | 97.2667 | 96.1333 | 95.4571 | 95
-18.6 | 97.2571 | 97.3333 | 96.2 | 95.6 | 95.1286
-18.5 | 97.5 | 97.4 | 96.3375 | 95.7 | 95.2
-18.4 | 97.5 | 97.45 | 96.4833 | 95.8 | 95.3667
-18.3 | 97.6 | 97.5143 | 96.66 | 95.9286 | 95.4333
-18.2 | 97.7143 | 97.6 | 96.7125 | 96.0833 | 95.5
-18.1 | 97.8 | 97.7 | 96.85 | 96.3125 | 95.6667
-18 | 97.8167 | 97.7833 | 97.0143 | 96.4 | 95.7286
-17.9 | 97.9 | 97.8375 | 97.1571 | 96.5714 | 95.8625
-17.8 | 98 | 97.9714 | 97.3143 | 96.6333 | 95.9
-17.7 | 98 | 98.0333 | 97.4571 | 96.7 | 96
-17.6 | 98.1 | 98.2 | 97.6 | 96.8 | 96.0625
-17.5 | 98.2 | 98.2571 | 97.6857 | 96.9 | 96.1875
-17.4 | 98.2 | 98.4 | 97.9 | 97 | 96.4
-17.3 | 98.3 | 98.5 | 97.9857 | 97.1714 | 96.425
-17.2 | 98.4 | 98.5857 | 98.1286 | 97.2 | 96.5667
-17.1 | 98.5 | 98.6286 | 98.2714 | 97.225 | 96.6333
-17 | 98.6333 | 98.7714 | 98.32 | 97.3 | 96.7857
-16.9 | 98.7 | 98.9 | 98.4 | 97.3857 | 96.8429
-16.8 | 98.8 | 99 | 98.4833 | 97.5 | 96.9
-16.7 | 98.9 | 99.15 | 98.6571 | 97.5833 | 97
-16.6 | 98.9 | 99.2 | 98.8167 | 97.7167 | 97.1
-16.5 | 99.1 | 99.3429 | 98.9 | 97.8 | 97.1625
-16.4 | 99.2 | 99.4 | 99 | 97.9 | 97.2875
-16.3 | 99.3667 | 99.5571 | 99 | 97.9 | 97.32
-16.2 | 99.4333 | 99.6286 | 99.1 | 98 | 97.4857
-16.1 | 99.6 | 99.7 | 99.1571 | 98 | 97.65
-16 | 99.6 | 99.8 | 99.2167 | 98.1 | 97.7125
-15.9 | 99.6 | 99.8 | 99.3833 | 98.2 | 97.8375
-15.8 | 99.7 | 99.9286 | 99.4 | 98.3833 | 97.9833
-15.7 | 99.7857 | 100 | 99.4143 | 98.5 | 98.1375
-15.6 | 99.8 | 100.1 | 99.5 | 98.6833 | 98.2556
-15.5 | 99.9 | 100.1 | 99.5 | 98.8 | 98.3
-15.4 | 99.9714 | 100.2 | 99.5 | 98.8 | 98.4
-15.3 | 100.0429 | 100.2 | 99.6 | 98.85 | 98.425
-15.2 | 100.175 | 100.2 | 99.6 | 98.9 | 98.5444
-15.1 | 100.2 | 100.2 | 99.6 | 99.05 | 98.6833
-15 | 100.3 | 100.3 | 99.6571 | 99.175 | 98.7375
-14.9 | 100.3 | 100.3 | 99.8 | 99.22 | 98.8556
-14.8 | 100.3 | 100.45 | 99.8 | 99.3857 | 98.9
-14.7 | 100.4 | 100.5 | 99.9667 | 99.5 | 99
-14.6 | 100.4 | 100.5 | 100 | 99.5 | 99.0375
-14.5 | 100.4 | 100.5875 | 100.1 | 99.6 | 99.1714
-14.4 | 100.5 | 100.6 | 100.1833 | 99.6 | 99.3
-14.3 | 100.5 | 100.6 | 100.3 | 99.7 | 99.3
-14.2 | 100.5 | 100.7 | 100.3 | 99.7667 | 99.3222
-14.1 | 100.5 | 100.7 | 100.4 | 99.8 | 99.4
-14 | 100.6 | 100.7 | 100.4857 | 99.9 | 99.4
-13.9 | 100.65 | 100.75 | 100.55 | 100 | 99.5
-13.8 | 100.7 | 100.8 | 100.65 | 100.0375 | 99.5
-13.7 | 100.7 | 100.8667 | 100.7167 | 100.1 | 99.6
-13.6 | 100.8 | 100.96 | 100.8 | 100.2 | 99.6
-13.5 | 100.9 | 101.0333 | 100.9167 | 100.2 | 99.7
-13.4 | 100.9 | 101.1 | 101 | 100.3 | 99.7
-13.3 | 100.9 | 101.16 | 101.0143 | 100.3375 | 99.7
-13.2 | 100.9167 | 101.2167 | 101.1 | 100.4625 | 99.8
-13.1 | 101.0714 | 101.3 | 101.1 | 100.5 | 99.8
-13 | 101.14 | 101.3333 | 101.1 | 100.5 | 99.8
-12.9 | 101.3 | 101.4 | 101.1 | 100.5833 | 99.8286
-12.8 | 101.3 | 101.5 | 101.2 | 100.6 | 99.9
-12.7 | 101.4 | 101.5 | 101.2 | 100.6 | 99.9
-12.6 | 101.4 | 101.6 | 101.2 | 100.7 | 99.9
-12.5 | 101.4 | 101.6 | 101.2857 | 100.7 | 100
-12.4 | 101.4 | 101.7 | 101.3 | 100.7 | 100
-12.3 | 101.5 | 101.7 | 101.4 | 100.7 | 100
-12.2 | 101.5 | 101.8 | 101.4 | 100.7 | 100
-12.1 | 101.5 | 101.8 | 101.4 | 100.7 | 100.1
-12 | 101.5714 | 101.8 | 101.4 | 100.7 | 100.1
-11.9 | 101.6 | 101.8 | 101.4 | 100.7 | 100.1
-11.8 | 101.6 | 101.8 | 101.475 | 100.7375 | 100.2
-11.7 | 101.6 | 101.8 | 101.6 | 100.8 | 100.2833
-11.6 | 101.7 | 101.9 | 101.6333 | 100.8 | 100.3
-11.5 | 101.7 | 101.9333 | 101.6 | 100.8 | 100.3833
-11.4 | 101.7 | 102 | 101.6 | 100.8 | 100.4222
-11.3 | 101.7 | 102.0714 | 101.6 | 100.8833 | 100.5
-11.2 | 101.8 | 102.1 | 101.6 | 100.9 | 100.5
-11.1 | 101.9333 | 102.1 | 101.6 | 100.9 | 100.6
-11 | 102 | 102.1 | 101.6 | 101 | 100.6
-10.9 | 102 | 102.1 | 101.6 | 101 | 100.7
-10.8 | 102 | 102.1 | 101.6 | 101.0429 | 100.8
-10.7 | 101.95 | 102.1 | 101.6 | 101.1 | 100.8429
-10.6 | 101.9 | 102 | 101.6 | 101.2 | 100.9167
-10.5 | 101.9 | 101.95 | 101.6 | 101.3 | 101
-10.4 | 101.8 | 101.9 | 101.6 | 101.3125 | 101
-10.3 | 101.8 | 101.9 | 101.7 | 101.4 | 101
-10.2 | 101.8 | 101.9 | 101.7 | 101.4625 | 101
-10.1 | 101.8 | 101.9 | 101.7 | 101.5 | 101
-10 | 101.8 | 101.95 | 101.7714 | 101.5 | 101
-9.9 | 101.82 | 102 | 101.8889 | 101.5 | 101
-9.8 | 101.9 | 102 | 101.9 | 101.6 | 101.0286
-9.7 | 101.9 | 102.1 | 102 | 101.6143 | 101.1
-9.6 | 101.9167 | 102.1 | 102 | 101.7 | 101.1571
-9.5 | 102 | 102.2 | 102.1 | 101.7 | 101.2
-9.4 | 102.0571 | 102.2625 | 102.1 | 101.7 | 101.2143
-9.3 | 102.1 | 102.3 | 102.2 | 101.7 | 101.3
-9.2 | 102.1833 | 102.3 | 102.2 | 101.7 | 101.3167
-9.1 | 102.2 | 102.35 | 102.2429 | 101.78 | 101.4
-9 | 102.3 | 102.4 | 102.3 | 101.8 | 101.5
-8.9 | 102.3 | 102.4429 | 102.3 | 101.8875 | 101.5
-8.8 | 102.4 | 102.5857 | 102.3 | 101.95 | 101.5
-8.7 | 102.4 | 102.6 | 102.3167 | 102 | 101.5
-8.6 | 102.4667 | 102.6 | 102.4 | 102 | 101.5
-8.5 | 102.5 | 102.7 | 102.4143 | 102 | 101.5
-8.4 | 102.5 | 102.7 | 102.5 | 102 | 101.5875
-8.3 | 102.5 | 102.7 | 102.5 | 102 | 101.6
-8.2 | 102.5667 | 102.7 | 102.5 | 102 | 101.6833
-8.1 | 102.6 | 102.7 | 102.5 | 102.1 | 101.7
-8 | 102.6 | 102.7 | 102.5 | 102.1 | 101.7
-7.9 | 102.6 | 102.7 | 102.5 | 102.1 | 101.75
-7.8 | 102.6 | 102.7 | 102.5 | 102.2 | 101.8
-7.7 | 102.7 | 102.8 | 102.5 | 102.2714 | 101.7333
-7.6 | 102.7 | 102.8 | 102.56 | 102.3 | 101.8
-7.5 | 102.7 | 102.8 | 102.6 | 102.3 | 101.8
-7.4 | 102.8429 | 102.8429 | 102.62 | 102.3 | 101.8
-7.3 | 102.9 | 102.9 | 102.7 | 102.3 | 101.8
-7.2 | 102.9 | 102.9 | 102.8 | 102.325 | 101.85
-7.1 | 102.9 | 102.9333 | 102.8 | 102.4 | 101.9
-7 | 102.9 | 103 | 102.8 | 102.4 | 101.9
-6.9 | 102.9 | 102.94 | 102.8 | 102.4 | 101.9
-6.8 | 102.9 | 103 | 102.8 | 102.4 | 102
-6.7 | 102.9 | 103 | 102.9 | 102.4 | 102
-6.6 | 102.9571 | 103 | 102.9 | 102.4 | 102
-6.5 | 103.02 | 103.1 | 102.9 | 102.4 | 102.0333
-6.4 | 103.1 | 103.1 | 102.9833 | 102.4 | 102.05
-6.3 | 103.1 | 103.1 | 102.9714 | 102.5 | 102.0857
-6.2 | 103.1 | 103.1 | 102.9 | 102.5 | 102.1
-6.1 | 103.1 | 103.1 | 103 | 102.5 | 102.1
-6 | 103.1 | 103.1 | 103 | 102.6 | 102.1
-5.9 | 103.1 | 103.1 | 103 | 102.5429 | 102.1
-5.8 | 103.1 | 103.1 | 103 | 102.5 | 102.1
-5.7 | 103.1 | 103.1 | 103 | 102.4714 | 102.1
-5.6 | 103.0167 | 103.1 | 102.9167 | 102.4 | 102.025
-5.5 | 103 | 103.1 | 102.9 | 102.4 | 101.9
-5.4 | 103 | 103 | 102.9 | 102.4 | 101.95
-5.3 | 103 | 103 | 102.875 | 102.4 | 102
-5.2 | 103 | 103 | 102.8 | 102.4 | 101.9
-5.1 | 103.1 | 103 | 102.8 | 102.4 | 101.9
-5 | 103.08 | 103 | 102.8 | 102.4 | 101.9125
-4.9 | 103 | 103 | 102.8 | 102.4 | 102
-4.8 | 102.9833 | 103 | 102.7667 | 102.3167 | 102
-4.7 | 102.9 | 102.9 | 102.6857 | 102.3 | 102
-4.6 | 102.9 | 102.8286 | 102.6 | 102.25 | 102
-4.5 | 102.8333 | 102.8 | 102.6 | 102.2 | 102
-4.4 | 102.7833 | 102.7857 | 102.6 | 102.2 | 101.9286
-4.3 | 102.7 | 102.7 | 102.6 | 102.2 | 101.9
-4.2 | 102.7 | 102.7 | 102.6 | 102.2 | 101.9
-4.1 | 102.6 | 102.7 | 102.5667 | 102.2 | 101.9
-4 | 102.6 | 102.7 | 102.5 | 102.1286 | 101.9
-3.9 | 102.6 | 102.6286 | 102.5 | 102.1143 | 101.7
-3.8 | 102.6 | 102.5667 | 102.5 | 102.1429 | 101.7
-3.7 | 102.5857 | 102.5 | 102.5 | 102.1 | 101.6125
-3.6 | 102.5 | 102.5 | 102.4 | 102.0429 | 101.6
-3.5 | 102.5 | 102.5 | 102.4 | 102 | 101.6
-3.4 | 102.4 | 102.4 | 102.3 | 101.9167 | 101.6
-3.3 | 102.3 | 102.4 | 102.3 | 101.8571 | 101.5667
-3.2 | 102.2 | 102.3857 | 102.2 | 101.8 | 101.5
-3.1 | 102.2 | 102.2429 | 102.1 | 101.7 | 101.4
-3 | 102.1 | 102.2 | 102.0429 | 101.6 | 101.35
-2.9 | 102.1 | 102.1 | 102 | 101.6 | 101.3
-2.8 | 102 | 102 | 101.9167 | 101.55 | 101.175
-2.7 | 102 | 101.9167 | 101.8571 | 101.4857 | 101.1
-2.6 | 101.8667 | 101.8 | 101.68 | 101.3333 | 101.0571
-2.5 | 101.8 | 101.7286 | 101.5 | 101.1667 | 100.9833
-2.4 | 101.7 | 101.5857 | 101.4167 | 101.1 | 100.9
-2.3 | 101.6 | 101.4333 | 101.3429 | 101 | 100.9
-2.2 | 101.4667 | 101.3 | 101.3 | 101 | 100.7833
-2.1 | 101.3 | 101.2833 | 101.15 | 100.9 | 100.7
-2 | 101.1571 | 101.1 | 101.08 | 100.9857 | 100.7
-1.9 | 101.0143 | 101.02 | 101 | 100.9 | 100.7
-1.8 | 100.9 | 100.95 | 100.85 | 100.8429 | 100.7
-1.7 | 100.9 | 100.8 | 100.8 | 100.7857 | 100.6143
-1.6 | 100.8 | 100.7333 | 100.74 | 100.6333 | 100.6
-1.5 | 100.7667 | 100.7 | 100.7 | 100.6 | 100.5857
-1.4 | 100.6833 | 100.6 | 100.6167 | 100.5 | 100.5
-1.3 | 100.5167 | 100.5286 | 100.6 | 100.5 | 100.46
-1.2 | 100.4714 | 100.5 | 100.5667 | 100.5 | 100.4
-1.1 | 100.4 | 100.5 | 100.5 | 100.4143 | 100.3
-1 | 100.4 | 100.3333 | 100.4 | 100.4 | 100.2
-0.9 | 100.3857 | 100.3 | 100.26 | 100.3 | 100.1167
-0.8 | 100.3 | 100.3 | 100.2 | 100.22 | 100.1
-0.7 | 100.2 | 100.2 | 100.2 | 100.16 | 100.1
-0.6 | 100.1333 | 100.1714 | 100.1667 | 100.1 | 100.1
-0.5 | 100.0571 | 100.1 | 100.1 | 100.1 | 100.1
-0.4 | 100 | 100 | 100.0667 | 100 | 100.1
-0.3 | 100 | 100 | 100 | 100 | 100.1
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100 | 100.075
0.4 | 100 | 100.0286 | 100.0667 | 100.0429 | 100.1
0.5 | 100.0833 | 100.1 | 100.1 | 100.1 | 100.1
0.6 | 100.16 | 100.1571 | 100.1667 | 100.1 | 100.1
0.7 | 100.2286 | 100.2 | 100.2 | 100.2 | 100.1
0.8 | 100.3 | 100.2833 | 100.2 | 100.2714 | 100.1
0.9 | 100.4 | 100.3 | 100.26 | 100.3 | 100.1571
1 | 100.4 | 100.3 | 100.4 | 100.4 | 100.2
1.1 | 100.4 | 100.4833 | 100.5 | 100.4667 | 100.3
1.2 | 100.4857 | 100.5 | 100.5667 | 100.5 | 100.4
1.3 | 100.54 | 100.5 | 100.6 | 100.5 | 100.4
1.4 | 100.7 | 100.6 | 100.6167 | 100.5333 | 100.5
1.5 | 100.8 | 100.7 | 100.7 | 100.6 | 100.6
1.6 | 100.8 | 100.7 | 100.74 | 100.6833 | 100.6
1.7 | 100.9 | 100.8714 | 100.8 | 100.7429 | 100.6
1.8 | 100.9 | 100.92 | 100.85 | 100.8857 | 100.7
1.9 | 101.0333 | 101 | 101 | 100.9 | 100.7
2 | 101.1857 | 101.1 | 101.08 | 100.9 | 100.7
2.1 | 101.3 | 101.26 | 101.15 | 100.9 | 100.7
2.2 | 101.4286 | 101.3 | 101.3 | 100.9714 | 100.8
2.3 | 101.6 | 101.5143 | 101.3429 | 101 | 100.8833
2.4 | 101.7 | 101.6 | 101.4167 | 101.0714 | 100.9
2.5 | 101.8 | 101.7 | 101.5 | 101.1333 | 100.9333
2.6 | 101.8857 | 101.85 | 101.68 | 101.3833 | 101
2.7 | 102 | 101.9 | 101.8571 | 101.5 | 101.1
2.8 | 102.0167 | 102.0571 | 101.9167 | 101.5 | 101.1286
2.9 | 102.1 | 102.1 | 102 | 101.6 | 101.2714
3 | 102.1143 | 102.1714 | 102.0429 | 101.6429 | 101.3
3.1 | 102.2 | 102.3 | 102.1 | 101.7143 | 101.4
3.2 | 102.22 | 102.3571 | 102.2 | 101.8 | 101.4833
3.3 | 102.3 | 102.4 | 102.3 | 101.9 | 101.6
3.4 | 102.4 | 102.4 | 102.3 | 101.9667 | 101.6
3.5 | 102.5 | 102.5 | 102.4 | 102 | 101.6
3.6 | 102.5 | 102.5 | 102.4 | 102 | 101.6
3.7 | 102.6 | 102.5 | 102.5 | 102.1 | 101.6625
3.8 | 102.6 | 102.6 | 102.5 | 102.1875 | 101.7
3.9 | 102.6 | 102.7 | 102.5 | 102.1 | 101.725
4 | 102.6 | 102.7 | 102.5 | 102.175 | 101.8571
4.1 | 102.625 | 102.7 | 102.5667 | 102.2 | 101.9
4.2 | 102.7 | 102.7 | 102.6 | 102.2 | 101.9
4.3 | 102.7 | 102.7 | 102.6 | 102.2 | 101.9
4.4 | 102.7 | 102.7571 | 102.6 | 102.2 | 101.9
4.5 | 102.8 | 102.8 | 102.6 | 102.2 | 101.9833
4.6 | 102.9 | 102.8 | 102.6 | 102.2 | 102
4.7 | 102.9 | 102.9 | 102.6857 | 102.3 | 102
4.8 | 103 | 103 | 102.7667 | 102.3714 | 102
4.9 | 103 | 103 | 102.8 | 102.4 | 102
5 | 103.05 | 103 | 102.8 | 102.4 | 101.9625
5.1 | 103.0667 | 103.0286 | 102.8 | 102.4 | 101.9
5.2 | 103.1 | 103.0444 | 102.8 | 102.4 | 101.9
5.3 | 103 | 103 | 102.875 | 102.4 | 101.95
5.4 | 103 | 103 | 102.9 | 102.4 | 102
5.5 | 103 | 103.1 | 102.9 | 102.4 | 101.9
5.6 | 103 | 103.1 | 102.9167 | 102.4 | 101.9857
5.7 | 103.1 | 103.1 | 103 | 102.5 | 102.1
5.8 | 103.1 | 103.1 | 103 | 102.5 | 102.1
5.9 | 103.1 | 103.1 | 103 | 102.5 | 102.1
6 | 103.1 | 103.1 | 103 | 102.5625 | 102.1
6.1 | 103.1 | 103.1 | 103 | 102.5 | 102.1
6.2 | 103.1 | 103.1 | 102.9 | 102.5 | 102.1
6.3 | 103.1 | 103.1 | 102.9714 | 102.46 | 102.025
6.4 | 103.1 | 103.1 | 102.9833 | 102.4 | 102.1
6.5 | 103 | 103.0333 | 102.9 | 102.4 | 102
6.6 | 102.9333 | 103 | 102.9 | 102.4 | 102
6.7 | 102.9 | 103 | 102.9 | 102.4 | 102
6.8 | 102.9 | 103 | 102.8 | 102.4 | 102
6.9 | 102.9 | 102.9 | 102.8 | 102.4 | 101.9286
7 | 102.9 | 103 | 102.8 | 102.4 | 101.9
7.1 | 102.9 | 102.9 | 102.8 | 102.4 | 101.9
7.2 | 102.9 | 102.9 | 102.8 | 102.3 | 101.9
7.3 | 102.9 | 102.9 | 102.7 | 102.3 | 101.8
7.4 | 102.8167 | 102.8714 | 102.62 | 102.3 | 101.8
7.5 | 102.7667 | 102.8 | 102.6 | 102.3 | 101.8
7.6 | 102.7 | 102.8 | 102.56 | 102.2833 | 101.8
7.7 | 102.7 | 102.8 | 102.5 | 102.2 | 101.7125
7.8 | 102.6 | 102.7167 | 102.5 | 102.2 | 101.8
7.9 | 102.6 | 102.7 | 102.5 | 102.1429 | 101.7
8 | 102.6 | 102.7 | 102.5 | 102.1 | 101.7
8.1 | 102.6 | 102.7 | 102.5 | 102.0833 | 101.7
8.2 | 102.5333 | 102.7 | 102.5 | 102 | 101.6286
8.3 | 102.5 | 102.7 | 102.5 | 102 | 101.6
8.4 | 102.5 | 102.7 | 102.5 | 102 | 101.6
8.5 | 102.5 | 102.7 | 102.4143 | 102 | 101.5
8.6 | 102.5 | 102.6 | 102.4 | 102 | 101.5
8.7 | 102.4 | 102.6 | 102.3167 | 102 | 101.5
8.8 | 102.4 | 102.6 | 102.3 | 102 | 101.5
8.9 | 102.3 | 102.4714 | 102.3 | 101.9 | 101.5
9 | 102.2714 | 102.4 | 102.3 | 101.8 | 101.4625
9.1 | 102.2 | 102.3833 | 102.2429 | 101.7286 | 101.4
9.2 | 102.1571 | 102.3 | 102.2 | 101.7 | 101.3667
9.3 | 102.1 | 102.3 | 102.2 | 101.7 | 101.3
9.4 | 102.1 | 102.2875 | 102.1 | 101.7 | 101.2625
9.5 | 102 | 102.2 | 102.1 | 101.7 | 101.2
9.6 | 101.9 | 102.1 | 102 | 101.7 | 101.1125
9.7 | 101.9 | 102.025 | 102 | 101.6 | 101.1
9.8 | 101.9 | 102 | 101.9 | 101.6 | 101.075
9.9 | 101.8 | 102 | 101.8889 | 101.5571 | 101
10 | 101.8 | 101.9 | 101.7714 | 101.5 | 101
10.1 | 101.8 | 101.9 | 101.7 | 101.5 | 101
10.2 | 101.8 | 101.9 | 101.7 | 101.5 | 101
10.3 | 101.8 | 101.9 | 101.7 | 101.4 | 101
10.4 | 101.8 | 101.9 | 101.6 | 101.3625 | 101
10.5 | 101.9 | 101.9286 | 101.6 | 101.3 | 101
10.6 | 101.9 | 102 | 101.6 | 101.1833 | 100.96
10.7 | 101.9667 | 102.0833 | 101.6 | 101.1 | 100.8875
10.8 | 102 | 102.1 | 101.6 | 101 | 100.7667
10.9 | 102 | 102.1 | 101.6 | 101 | 100.7
11 | 102 | 102.1 | 101.6 | 101 | 100.6
11.1 | 101.9167 | 102.1 | 101.6 | 100.9143 | 100.5714
11.2 | 101.8333 | 102.1 | 101.6 | 100.9 | 100.5
11.3 | 101.7667 | 102 | 101.6 | 100.9 | 100.4889
11.4 | 101.7 | 102 | 101.6 | 100.8 | 100.4
11.5 | 101.7 | 101.9667 | 101.6 | 100.8 | 100.4
11.6 | 101.6857 | 101.9 | 101.6333 | 100.8857 | 100.3
11.7 | 101.6 | 101.8 | 101.6 | 100.8 | 100.2286
11.8 | 101.6 | 101.8 | 101.475 | 100.7 | 100.2
11.9 | 101.6 | 101.8 | 101.4 | 100.7 | 100.1
12 | 101.6 | 101.8 | 101.4 | 100.7 | 100.1
12.1 | 101.5 | 101.8 | 101.4 | 100.7 | 100.075
12.2 | 101.5 | 101.8 | 101.4 | 100.7 | 100
12.3 | 101.4778 | 101.7 | 101.4 | 100.7 | 100
12.4 | 101.4 | 101.7 | 101.3 | 100.7 | 100
12.5 | 101.4 | 101.6 | 101.2857 | 100.7 | 99.9714
12.6 | 101.4 | 101.6 | 101.2 | 100.7 | 99.9
12.7 | 101.3875 | 101.5167 | 101.2 | 100.6 | 99.9
12.8 | 101.3 | 101.5 | 101.2 | 100.6 | 99.9
12.9 | 101.2833 | 101.4143 | 101.1 | 100.6 | 99.8
13 | 101.2 | 101.36 | 101.1 | 100.5 | 99.8
13.1 | 101.05 | 101.3 | 101.1 | 100.5 | 99.8
13.2 | 101 | 101.2429 | 101.1 | 100.4125 | 99.7625
13.3 | 100.9 | 101.1833 | 101.0143 | 100.3 | 99.7
13.4 | 100.9 | 101.1 | 101 | 100.2625 | 99.7
13.5 | 100.8833 | 101 | 100.9167 | 100.2 | 99.7
13.6 | 100.8 | 101 | 100.8 | 100.2 | 99.6
13.7 | 100.7 | 100.8 | 100.7167 | 100.1 | 99.6
13.8 | 100.7 | 100.8 | 100.65 | 100 | 99.5
13.9 | 100.6286 | 100.775 | 100.55 | 99.95 | 99.4875
14 | 100.5833 | 100.7 | 100.4857 | 99.875 | 99.4
14.1 | 100.5 | 100.7 | 100.4 | 99.8 | 99.4
14.2 | 100.5 | 100.66 | 100.3 | 99.7143 | 99.3714
14.3 | 100.5 | 100.6 | 100.3 | 99.7 | 99.3
14.4 | 100.5 | 100.6 | 100.1833 | 99.6143 | 99.2833
14.5 | 100.4 | 100.5167 | 100.1 | 99.6 | 99.2
14.6 | 100.4 | 100.5 | 100 | 99.5143 | 99
14.7 | 100.4 | 100.5 | 99.9667 | 99.4857 | 99
14.8 | 100.3 | 100.4833 | 99.8 | 99.4 | 98.9333
14.9 | 100.3 | 100.3286 | 99.8 | 99.2667 | 98.8125
15 | 100.2833 | 100.3 | 99.6571 | 99.125 | 98.7857
15.1 | 100.2 | 100.2143 | 99.6 | 99.1 | 98.6333
15.2 | 100.1571 | 100.2 | 99.6 | 98.9571 | 98.5
15.3 | 100.0143 | 100.2 | 99.6 | 98.8 | 98.4714
15.4 | 99.95 | 100.2 | 99.5 | 98.8 | 98.4
15.5 | 99.8833 | 100.1 | 99.5 | 98.7833 | 98.3333
15.6 | 99.8 | 100.1 | 99.5 | 98.6286 | 98.2111
15.7 | 99.7625 | 100 | 99.4143 | 98.5143 | 98.1857
15.8 | 99.7 | 99.9571 | 99.4 | 98.4 | 98.02
15.9 | 99.6571 | 99.8 | 99.3833 | 98.2 | 97.8875
16 | 99.6 | 99.7429 | 99.2167 | 98.0833 | 97.7571
16.1 | 99.5833 | 99.7 | 99.1571 | 98 | 97.6143
16.2 | 99.5 | 99.6 | 99.1 | 97.95 | 97.5
16.3 | 99.4 | 99.5 | 99 | 97.9 | 97.3833
16.4 | 99.2714 | 99.4286 | 99 | 97.8875 | 97.2375
16.5 | 99.1 | 99.3 | 98.9 | 97.8 | 97.2
16.6 | 98.95 | 99.2143 | 98.8167 | 97.7 | 97.1
16.7 | 98.8833 | 99.0714 | 98.6571 | 97.6286 | 97.0167
16.8 | 98.8 | 99 | 98.4833 | 97.4714 | 96.9429
16.9 | 98.7 | 98.8333 | 98.4 | 97.4 | 96.9
17 | 98.6 | 98.8 | 98.32 | 97.3 | 96.8
17.1 | 98.5 | 98.65 | 98.2714 | 97.275 | 96.6833
17.2 | 98.3667 | 98.6 | 98.1286 | 97.2 | 96.5167
17.3 | 98.3 | 98.5 | 97.9857 | 97.125 | 96.475
17.4 | 98.2571 | 98.4 | 97.9 | 97.06 | 96.35
17.5 | 98.2 | 98.2857 | 97.6857 | 96.8667 | 96.1667
17.6 | 98.1 | 98.2 | 97.6 | 96.8 | 96.1
17.7 | 98 | 98.06 | 97.4571 | 96.7 | 95.9857
17.8 | 98 | 98 | 97.3143 | 96.6833 | 95.9
17.9 | 97.9667 | 97.8571 | 97.1571 | 96.5375 | 95.8125
18 | 97.9 | 97.7 | 97.0143 | 96.3875 | 95.7
18.1 | 97.8 | 97.7 | 96.85 | 96.2625 | 95.6167
18.2 | 97.7 | 97.6 | 96.7125 | 96.1286 | 95.55
18.3 | 97.6375 | 97.5429 | 96.66 | 95.9833 | 95.4833
18.4 | 97.5667 | 97.4 | 96.4833 | 95.8167 | 95.4
18.5 | 97.4833 | 97.4 | 96.3375 | 95.7 | 95.2333
18.6 | 97.3375 | 97.3571 | 96.2 | 95.5667 | 95.075
18.7 | 97.1667 | 97.3 | 96.1333 | 95.5 | 94.95
18.8 | 97 | 97.1333 | 96 | 95.375 | 94.9
18.9 | 96.7571 | 97 | 95.9333 | 95.2333 | 94.8
19 | 96.6 | 96.9 | 95.8 | 95.05 | 94.7286
19.1 | 96.2857 | 96.7714 | 95.675 | 94.8556 | 94.7
19.2 | 95.86 | 96.44 | 95.55 | 94.7375 | 94.4875
19.3 | 95.65 | 96.125 | 95.5 | 94.6125 | 94.325
19.4 | 95.15 | 95.8125 | 95.4 | 94.4714 | 94.1375
19.5 | 94.4857 | 95.4375 | 95.3 | 94.3375 | 93.925
19.6 | 93.8 | 95.05 | 95.18 | 94.2 | 93.9
19.7 | 92.8833 | 94.6 | 94.9 | 93.925 | 93.775
19.8 | 91.75 | 93.95 | 94.7333 | 93.7375 | 93.65
19.9 | 90.3375 | 93.36 | 94.55 | 93.6 | 93.5333
20 | 88.2429 | 92.6 | 94.3 | 93.375 | 93.3
20.1 | 85.2 | 91.6 | 94.0333 | 93.25 | 93.1429
20.2 | 80.5 | 90.4286 | 93.6429 | 93.05 | 92.9167
20.3 | 74.4571 | 88.6571 | 93.3 | 92.9 | 92.825
20.4 | 67.3 | 86.52 | 93.0857 | 92.6857 | 92.6375
20.5 | 58.2667 | 83.6 | 92.7 | 92.4857 | 92.5125
20.6 | 50.1 | 78.8143 | 92.0429 | 92.3429 | 92.3
20.7 | 41.7 | 73.1125 | 91.4 | 92.16 | 92.1714
20.8 | 34.6 | 65.6 | 90.8714 | 92.08 | 92.0167
20.9 | 28.8143 | 57.6571 | 90.175 | 91.9143 | 91.7
21 | 23.9333 | 48.95 | 89.3 | 91.7 | 91.5222
21.1 | 20.7 | 41.6 | 88.3 | 91.6167 | 91.5
21.2 | 18.12 | 34.7 | 86.8 | 91.4667 | 91.275
21.3 | 16.4571 | 28.85 | 85.1714 | 91.3143 | 91.0667
21.4 | 15.0857 | 24.2667 | 82.9 | 91.0667 | 90.9
21.5 | 14.1 | 20.96 | 79.625 | 90.8167 | 90.675
21.6 | 13.3143 | 18.4 | 75.3 | 90.6429 | 90.4667
21.7 | 12.6667 | 16.5571 | 69.6 | 90.3571 | 90.2375
21.8 | 12.2 | 15.3 | 62.525 | 90.1333 | 90.1125
21.9 | 11.75 | 14.3125 | 55.2 | 89.8 | 89.8875
22 | 11.3667 | 13.475 | 47.5 | 89.4333 | 89.725
22.1 | 11.1 | 12.7667 | 40.4429 | 89.1286 | 89.5444
22.2 | 10.84 | 12.18 | 34.2714 | 88.8143 | 89.4143
22.3 | 10.6667 | 11.7 | 29.1 | 88.2875 | 89.1714
22.4 | 10.3833 | 11.3286 | 25.22 | 87.8 | 89.0286
22.5 | 10.2167 | 11 | 22.125 | 87.4286 | 88.9111
22.6 | 10.02 | 10.7333 | 19.9 | 86.8143 | 88.7714
22.7 | 9.84 | 10.4833 | 18.22 | 86.2875 | 88.5333
22.8 | 9.6714 | 10.2571 | 16.975 | 85.575 | 88.4667
22.9 | 9.6 | 10.0667 | 15.9571 | 84.9 | 88.2286
23 | 9.4333 | 9.8 | 15 | 83.9 | 88.05
23.1 | 9.36 | 9.6667 | 14.4 | 82.9 | 87.7857
23.2 | 9.2 | 9.55 | 13.8 | 81.6 | 87.5833
23.3 | 9.1 | 9.3 | 13.3 | 79.675 | 87.3571
23.4 | 8.9833 | 9.1667 | 12.9 | 77.2 | 87.0429
23.5 | 8.9 | 9.0667 | 12.5125 | 73.9 | 86.8167
23.6 | 8.8 | 8.9 | 12.2 | 69.7667 | 86.525
23.7 | 8.8 | 8.7875 | 11.925 | 64.6 | 86.275
23.8 | 8.65 | 8.65 | 11.6333 | 59.0556 | 86
23.9 | 8.6 | 8.6 | 11.4143 | 52.5667 | 85.525
24 | 8.5 | 8.4571 | 11.15 | 46.2333 | 85.1857
24.1 | 8.5 | 8.4 | 11 | 40.3 | 84.8333
24.2 | 8.4 | 8.3333 | 10.7286 | 35.025 | 84.55
24.3 | 8.4 | 8.2714 | 10.575 | 30.5 | 84
24.4 | 8.3375 | 8.2 | 10.35 | 27.1111 | 83.6143
24.5 | 8.3 | 8.1 | 10.2125 | 24.425 | 83.0714
24.6 | 8.2 | 8 | 10.0833 | 22.4 | 82.55
24.7 | 8.1857 | 8 | 9.9375 | 20.725 | 82.0333
24.8 | 8.1 | 7.9 | 9.725 | 19.5 | 81.2714
24.9 | 8 | 7.9 | 9.65 | 18.4 | 80.4
25 | 8 | 7.8286 | 9.4833 | 17.56 | 79.475
25.1 | 7.9833 | 7.8 | 9.3167 | 16.975 | 78.3
25.2 | 7.9 | 7.7143 | 9.2 | 16.25 | 76.75
25.3 | 7.8167 | 7.7 | 9.1143 | 15.8333 | 74.9
25.4 | 7.8 | 7.7 | 9.05 | 15.3125 | 72.6444
25.5 | 7.7 | 7.6 | 8.9 | 14.9375 | 69.575
25.6 |  | 7.6 | 8.8429 | 14.6333 | 65.5125
25.7 |  | 7.5 | 8.8 | 14.2429 | 60.3
25.8 |  | 7.4 | 8.7 | 14.025 | 55.4125
25.9 |  | 7.4 | 8.5875 | 13.7333 | 49.5375
26 |  |  | 8.5 | 13.4667 | 43.9125
26.1 |  |  | 8.4167 | 13.1833 | 38.8875
26.2 |  |  | 8.3 | 13 | 34.3125
26.3 |  |  | 8.2333 | 12.8667 | 30.5875
26.4 |  |  | 8.2 | 12.65 | 27.525
26.5 |  |  | 8.15 | 12.4111 | 25.0375
26.6 |  |  | 8.1 | 12.2167 | 23.5
26.7 |  |  | 8 | 12.05 | 22.075
26.8 |  |  | 7.95 | 11.9125 | 20.95
26.9 |  |  | 7.9 | 11.7333 | 19.975
27 |  |  |  | 11.54 | 19.2
27.1 |  |  |  | 11.375 | 18.5667
27.2 |  |  |  | 11.1625 | 17.9875
27.3 |  |  |  | 11.0167 | 17.4875
27.4 |  |  |  | 10.875 | 16.95
27.5 |  |  |  | 10.75 | 16.5625
27.6 |  |  |  | 10.6143 | 16.225
27.7 |  |  |  | 10.42 | 15.9125
27.8 |  |  |  | 10.2571 | 15.625
27.9 |  |  |  | 10.1143 | 15.4
28 |  |  |  | 10.0429 | 15.15
28.1 |  |  |  | 9.9875 | 14.9
28.2 |  |  |  | 9.9 | 14.7
28.3 |  |  |  | 9.8286 | 14.4333
28.4 |  |  |  | 9.6833 | 14.1833
28.5 |  |  |  | 9.5167 | 14.0167
28.6 |  |  |  | 9.45 | 13.8143
28.7 |  |  |  | 9.3125 | 13.6
28.8 |  |  |  | 9.2 | 13.4167
28.9 |  |  |  | 9.1667 | 13.25
29 |  |  |  |  | 13.0667
29.1 |  |  |  |  | 12.8286
29.2 |  |  |  |  | 12.68
29.3 |  |  |  |  | 12.4667
29.4 |  |  |  |  | 12.325
29.5 |  |  |  |  | 12.2
29.6 |  |  |  |  | 12
29.7 |  |  |  |  | 11.9
29.8 |  |  |  |  | 11.72
29.9 |  |  |  |  | 11.65
30 |  |  |  |  | 11.48
30.1 |  |  |  |  | 11.3143
30.2 |  |  |  |  | 11.2
30.3 |  |  |  |  | 11.1
30.4 |  |  |  |  | 10.9
30.5 |  |  |  |  | 10.8429
30.6 |  |  |  |  | 10.8
30.7 |  |  |  |  | 10.675
30.8 |  |  |  |  | 10.55
30.9 |  |  |  |  | 10.4

## Wedge 60 Output Factors
Output factors
Wedge 60 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15
| 3 | 0.863 | 0.881 | 0.889 | 0.898 | 0.905 | 0.911
| 4 | 0.882 | 0.902 | 0.913 | 0.924 | 0.933 | 0.941
| 5 | 0.892 | 0.914 | 0.926 | 0.94 | 0.951 | 0.961
| 7 | 0.905 | 0.929 | 0.944 | 0.961 | 0.976 | 0.989
| 10 | 0.916 | 0.944 | 0.961 | 0.983 | 1 | 1.018
| 15 | 0.926 | 0.957 | 0.976 | 1.002 | 1.025 | 1.048
| 20 | 0.931 | 0.964 | 0.985 | 1.013 | 1.04 | 1.068
| 30 | 0.94 | 0.974 | 0.997 | 1.029 | 1.061 | 1.096
| 40 | 0.944 | 0.98 | 1.003 | 1.037 | 1.072 | 1.11