## Title
Varian TrueBeam
4MV Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables |  | 2. Plots
| Open field |  | Open Field
| Open Field Depth Doses |  | Wedge 15
| Open Field Profiles at 1.2cm |  | Wedge 30 Plots
| Open Field Profiles at 5cm |  | Wedge 45 Plots
| Open Field Profiles at 10cm |  | Wedge 60 Plots
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
Wedge 15
Wedge 15 Depth Doses
SSD80 Wedge 15 Depth Doses
SSD120 Wedge 15 Depth Doses
Wedge 15 Lateral Profiles
Wedge 15 Longitudinal Profiles
Wedge 15 Dosimetric Data
Wedge 30
Wedge 30 Depth Doses
SSD80 Wedge 30 Depth Doses
SSD120 Wedge 30 Depth Doses
Wedge 30 Lateral Profiles
Wedge 30 Longitudinal Profiles
Wedge 30 Dosimetric Data
Wedge 45
Wedge 45 Depth Doses
SSD80 Wedge 45 Depth Doses
Wedge 45 Longitudinal Profiles
Wedge 45 Lateral Profiles
Wedge 45 Longitudinal Profiles
Wedge 45 Dosimetric Data
Wedge 60
Wedge 60 Depth Doses
SSD80 Wedge 60 Depth Doses
SSD120 Wedge 60 Depth Doses
Wedge 60 Lateral Profiles
Wedge 60 Longitudinal Profiles
Wedge 60 Dosimetric Data

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 59.8 | 60.2 | 61.9 | 62.8 | 64.9 | 71.8 | 76.7 | 79.6
0.1 | 64.5667 | 65.2 | 66.3429 | 67.825 | 69.44 | 75.7333 | 80.125 | 83
0.2 | 72.9 | 73.56 | 74.6 | 75.4 | 76.25 | 82.02 | 85.6 | 87.75
0.3 | 82.025 | 82 | 82.8286 | 83.6667 | 84.4 | 88.4 | 91.1143 | 92.44
0.4 | 88.72 | 88.9 | 89.6 | 89.9 | 90.36 | 93.14 | 95 | 95.7
0.5 | 93.3714 | 93.425 | 93.7 | 94 | 94.5 | 96.0667 | 97.25 | 97.84
0.6 | 96.4 | 96.3 | 96.6 | 96.575 | 96.8556 | 98 | 98.75 | 99
0.7 | 98.3 | 98.05 | 98.1667 | 98.325 | 98.4333 | 99.1667 | 99.5714 | 99.7
0.8 | 99.3 | 99.2 | 99.4 | 99.3 | 99.36 | 99.78 | 99.94 | 99.975
0.9 | 99.9143 | 99.7714 | 99.9 | 99.82 | 99.84 | 100 | 100 | 100
1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
1.1 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9667 | 99.925
1.2 | 99.9 | 99.94 | 99.975 | 99.94 | 99.9667 | 99.75 | 99.76 | 99.725
1.3 | 99.8 | 99.6857 | 99.75 | 99.7 | 99.75 | 99.4 | 99.5 | 99.6
1.4 | 99.12 | 99.3 | 99.425 | 99.35 | 99.34 | 99.2167 | 99.075 | 99.24
1.5 | 98.8 | 98.95 | 98.95 | 98.9 | 99.02 | 98.9 | 98.8 | 98.9
1.6 | 98.1667 | 98.4 | 98.48 | 98.45 | 98.5333 | 98.575 | 98.48 | 98.55
1.7 | 97.6 | 97.8429 | 98.1 | 98.1 | 98.15 | 98.16 | 98.2 | 98.3
1.8 | 97.06 | 97.26 | 97.5333 | 97.5667 | 97.68 | 97.7333 | 97.8 | 97.95
1.9 | 96.6 | 96.82 | 97.1333 | 97.1 | 97.35 | 97.5333 | 97.52 | 97.72
2 | 96.05 | 96.2667 | 96.7 | 96.75 | 96.78 | 97.2 | 97.12 | 97.35
2.1 | 95.325 | 95.725 | 96.2 | 96.16 | 96.4 | 96.66 | 96.8 | 97
2.2 | 94.8 | 95.0667 | 95.6333 | 95.7667 | 96 | 96.2333 | 96.4333 | 96.66
2.3 | 94.3 | 94.5 | 95.08 | 95.3167 | 95.5556 | 95.8667 | 96.1 | 96.3
2.4 | 93.75 | 93.8889 | 94.68 | 94.7 | 95.1 | 95.625 | 95.7333 | 95.925
2.5 | 92.9 | 93.36 | 94.15 | 94.35 | 94.7143 | 95.1 | 95.4 | 95.6
2.6 | 92.38 | 92.9333 | 93.7125 | 93.925 | 94.2 | 94.74 | 94.95 | 95.24
2.7 | 91.8 | 92.38 | 93 | 93.5 | 93.78 | 94.2667 | 94.6333 | 94.86
2.8 | 91.26 | 91.7333 | 92.56 | 92.94 | 93.3 | 93.9 | 94.25 | 94.5
2.9 | 90.6 | 91.16 | 92.22 | 92.5 | 92.9 | 93.6286 | 93.875 | 94.1333
3 | 90.02 | 90.7333 | 91.6 | 92.05 | 92.45 | 93.3 | 93.7 | 93.8
3.1 | 89.5 | 90.2 | 91.1 | 91.6 | 91.9429 | 92.7667 | 93.28 | 93.46
3.2 | 88.85 | 89.5667 | 90.6 | 91.05 | 91.5857 | 92.4 | 92.8667 | 93.22
3.3 | 88.225 | 89.1 | 90.1 | 90.74 | 91.04 | 92.0333 | 92.56 | 92.875
3.4 | 87.65 | 88.4 | 89.6 | 90.22 | 90.58 | 91.6 | 92.2143 | 92.36
3.5 | 87.075 | 87.85 | 89.0667 | 89.66 | 90.25 | 91.2333 | 91.7 | 92.08
3.6 | 86.4222 | 87.3429 | 88.54 | 89.25 | 89.8 | 91.02 | 91.35 | 91.78
3.7 | 85.84 | 86.8 | 88.15 | 88.75 | 89.3333 | 90.54 | 90.98 | 91.4
3.8 | 85.3 | 86.2875 | 87.55 | 88.275 | 88.8571 | 90.05 | 90.64 | 91.05
3.9 | 84.72 | 85.68 | 87.05 | 87.8 | 88.4333 | 89.8 | 90.3 | 90.6
4 | 84.05 | 85.1429 | 86.55 | 87.3429 | 88 | 89.28 | 89.9333 | 90.38
4.1 | 83.5667 | 84.64 | 86 | 86.8286 | 87.5429 | 88.975 | 89.62 | 90.05
4.2 | 82.9571 | 84.1 | 85.5 | 86.35 | 87.0857 | 88.54 | 89.18 | 89.65
4.3 | 82.32 | 83.5 | 85.1 | 85.8 | 86.5571 | 88.1 | 88.8 | 89.3857
4.4 | 81.825 | 82.9 | 84.55 | 85.3 | 86.1 | 87.6 | 88.5 | 89
4.5 | 81.25 | 82.325 | 84.1 | 84.9333 | 85.6 | 87.35 | 88.1 | 88.64
4.6 | 80.775 | 81.9 | 83.55 | 84.5 | 85.25 | 86.88 | 87.6 | 88.2
4.7 | 80.2 | 81.3 | 83.125 | 84.05 | 84.72 | 86.425 | 87.375 | 87.88
4.8 | 79.7 | 80.74 | 82.5857 | 83.5 | 84.25 | 86.1 | 87.075 | 87.55
4.9 | 79.18 | 80.25 | 82.06 | 82.975 | 83.875 | 85.7333 | 86.725 | 87.175
5 | 78.6 | 79.75 | 81.575 | 82.54 | 83.45 | 85.3333 | 86.25 | 86.8
5.1 | 78.04 | 79.3 | 81.04 | 82.1 | 82.95 | 84.9 | 85.8429 | 86.5
5.2 | 77.4 | 78.78 | 80.5286 | 81.6286 | 82.5333 | 84.6 | 85.6 | 86.14
5.3 | 76.8 | 78.2286 | 80.2 | 81.15 | 82.1 | 84.15 | 85.175 | 85.8286
5.4 | 76.3 | 77.72 | 79.6375 | 80.65 | 81.6111 | 83.8 | 84.82 | 85.6
5.5 | 75.9 | 77.2 | 79.08 | 80.16 | 81.24 | 83.46 | 84.4 | 85.15
5.6 | 75.25 | 76.7 | 78.5714 | 79.775 | 80.7429 | 83 | 83.95 | 84.8
5.7 | 74.85 | 76.06 | 78.1 | 79.38 | 80.4 | 82.64 | 83.6125 | 84.36
5.8 | 74.26 | 75.58 | 77.6444 | 78.8571 | 79.8625 | 82.3222 | 83.26 | 84
5.9 | 73.7667 | 75 | 77.1 | 78.3667 | 79.48 | 81.86 | 83.02 | 83.675
6 | 73.1 | 74.5667 | 76.68 | 77.92 | 79.05 | 81.5 | 82.5667 | 83.325
6.1 | 72.5333 | 74.05 | 76.1857 | 77.5 | 78.5 | 81.1 | 82.3 | 82.98
6.2 | 72.15 | 73.55 | 75.7 | 77.05 | 78.1 | 80.65 | 81.9 | 82.7
6.3 | 71.56 | 73.14 | 75.2333 | 76.45 | 77.62 | 80.3 | 81.6 | 82.2
6.4 | 71.1 | 72.5333 | 74.8 | 75.9714 | 77.1571 | 79.9286 | 81.15 | 81.9
6.5 | 70.7 | 72.0667 | 74.2667 | 75.64 | 76.8429 | 79.56 | 80.8 | 81.5
6.6 | 70.1 | 71.55 | 73.78 | 75.2 | 76.4 | 79.2 | 80.46 | 81.2333
6.7 | 69.78 | 71.02 | 73.26 | 74.74 | 75.95 | 78.84 | 80.04 | 80.8
6.8 | 69.1333 | 70.46 | 72.8 | 74.3333 | 75.58 | 78.3571 | 79.68 | 80.425
6.9 | 68.8 | 70 | 72.4 | 73.95 | 75.0714 | 78.05 | 79.4 | 80.15
7 | 68.2 | 69.525 | 71.85 | 73.4 | 74.6 | 77.62 | 79.0333 | 79.6667
7.1 | 67.7 | 69.15 | 71.4 | 72.925 | 74.15 | 77.16 | 78.65 | 79.375
7.2 | 67.2333 | 68.6 | 70.875 | 72.46 | 73.8 | 76.86 | 78.25 | 79.1
7.3 | 66.8333 | 68.1333 | 70.45 | 72.1 | 73.2 | 76.46 | 77.925 | 78.8333
7.4 | 66.2143 | 67.65 | 70.04 | 71.68 | 72.92 | 76.02 | 77.5 | 78.3571
7.5 | 65.8 | 67.15 | 69.55 | 71.2 | 72.4333 | 75.64 | 77.1 | 78
7.6 | 65.2875 | 66.75 | 69.06 | 70.65 | 72.1 | 75.24 | 76.7857 | 77.75
7.7 | 64.76 | 66.2571 | 68.7286 | 70.3 | 71.575 | 74.88 | 76.45 | 77.4
7.8 | 64.25 | 65.84 | 68.15 | 69.9 | 71.25 | 74.55 | 76.075 | 77
7.9 | 63.75 | 65.3778 | 67.6875 | 69.4 | 70.7 | 74.2143 | 75.7333 | 76.6
8 | 63.45 | 64.82 | 67.3 | 69.08 | 70.4 | 73.76 | 75.425 | 76.2333
8.1 | 62.975 | 64.45 | 66.8714 | 68.55 | 69.875 | 73.45 | 75.1 | 75.82
8.2 | 62.5 | 63.9667 | 66.4714 | 68.1333 | 69.525 | 73.14 | 74.7714 | 75.625
8.3 | 62.24 | 63.55 | 66.1143 | 67.8333 | 69.1333 | 72.68 | 74.275 | 75.25
8.4 | 61.75 | 63.25 | 65.68 | 67.35 | 68.7667 | 72.34 | 74.025 | 74.975
8.5 | 61.28 | 62.8 | 65.22 | 66.95 | 68.4 | 71.92 | 73.56 | 74.675
8.6 | 60.7333 | 62.2667 | 64.775 | 66.5 | 67.9333 | 71.55 | 73.325 | 74.25
8.7 | 60.2 | 61.85 | 64.35 | 66.1 | 67.55 | 71.22 | 73.06 | 73.825
8.8 | 59.95 | 61.3857 | 63.94 | 65.7 | 67.16 | 70.92 | 72.675 | 73.58
8.9 | 59.4 | 60.95 | 63.45 | 65.25 | 66.775 | 70.46 | 72.25 | 73.3
9 | 59.04 | 60.5 | 63.08 | 64.84 | 66.35 | 70.2 | 71.95 | 73.1
9.1 | 58.55 | 60.1571 | 62.66 | 64.4667 | 65.94 | 69.7714 | 71.525 | 72.7
9.2 | 58.26 | 59.7286 | 62.28 | 64.15 | 65.5667 | 69.44 | 71.25 | 72.35
9.3 | 57.75 | 59.38 | 61.84 | 63.6429 | 65.15 | 69.15 | 70.86 | 72.06
9.4 | 57.28 | 58.9444 | 61.48 | 63.26 | 64.7 | 68.7 | 70.55 | 71.65
9.5 | 56.9875 | 58.6 | 60.95 | 62.8 | 64.425 | 68.45 | 70.2 | 71.3
9.6 | 56.6 | 58.05 | 60.6143 | 62.44 | 64.1 | 68.1 | 69.9286 | 71
9.7 | 56.1429 | 57.72 | 60.26 | 62.0857 | 63.5625 | 67.7 | 69.6 | 70.7
9.8 | 55.76 | 57.2714 | 59.8143 | 61.7 | 63.24 | 67.3 | 69.2125 | 70.35
9.9 | 55.4875 | 56.82 | 59.36 | 61.275 | 62.925 | 66.95 | 68.95 | 70.04
10 | 55 | 56.3714 | 58.975 | 60.94 | 62.48 | 66.72 | 68.5 | 69.75
10.1 | 54.55 | 56.075 | 58.68 | 60.55 | 62.08 | 66.25 | 68.26 | 69.3333
10.2 | 54.2 | 55.7286 | 58.2429 | 60.2 | 61.7 | 65.95 | 67.9143 | 69.075
10.3 | 53.8714 | 55.35 | 58 | 59.8125 | 61.25 | 65.625 | 67.6286 | 68.775
10.4 | 53.45 | 54.9333 | 57.5333 | 59.45 | 61 | 65.35 | 67.4 | 68.45
10.5 | 53 | 54.52 | 57 | 59.05 | 60.64 | 64.875 | 66.94 | 68.1
10.6 | 52.65 | 54.18 | 56.6571 | 58.6 | 60.2 | 64.575 | 66.6 | 67.825
10.7 | 52.325 | 53.82 | 56.3 | 58.2571 | 59.94 | 64.225 | 66.26 | 67.4333
10.8 | 51.85 | 53.5 | 55.9625 | 57.92 | 59.6 | 63.95 | 66 | 67.2
10.9 | 51.575 | 53 | 55.58 | 57.5333 | 59.2 | 63.66 | 65.7 | 66.9
11 | 51.15 | 52.775 | 55.2714 | 57.22 | 58.9 | 63.18 | 65.3 | 66.55
11.1 | 50.75 | 52.28 | 54.8429 | 56.7625 | 58.5 | 62.8 | 64.9667 | 66.2
11.2 | 50.36 | 51.95 | 54.5 | 56.425 | 58 | 62.55 | 64.6 | 65.95
11.3 | 50.1 | 51.6 | 54.12 | 56.0667 | 57.72 | 62.25 | 64.4 | 65.62
11.4 | 49.75 | 51.25 | 53.7143 | 55.65 | 57.375 | 62 | 64.15 | 65.425
11.5 | 49.54 | 50.84 | 53.4286 | 55.3667 | 57.025 | 61.6 | 63.75 | 64.9667
11.6 | 49.05 | 50.375 | 53.1 | 54.95 | 56.7333 | 61.3 | 63.42 | 64.725
11.7 | 48.74 | 50.14 | 52.72 | 54.6667 | 56.35 | 60.9 | 63.16 | 64.4
11.8 | 48.28 | 49.84 | 52.3333 | 54.3 | 55.92 | 60.65 | 62.925 | 64.0667
11.9 | 48 | 49.54 | 52.0286 | 54 | 55.625 | 60.2 | 62.575 | 63.925
12 | 47.7 | 49.15 | 51.72 | 53.58 | 55.32 | 59.95 | 62.18 | 63.5333
12.1 | 47.38 | 48.72 | 51.38 | 53.35 | 54.96 | 59.6 | 61.9 | 63.175
12.2 | 47 | 48.38 | 51.05 | 53 | 54.65 | 59.25 | 61.6 | 62.88
12.3 | 46.7 | 48.0333 | 50.6 | 52.575 | 54.2571 | 58.94 | 61.2 | 62.54
12.4 | 46.36 | 47.76 | 50.25 | 52.175 | 53.92 | 58.675 | 60.95 | 62.2667
12.5 | 46.0143 | 47.4286 | 49.875 | 51.85 | 53.64 | 58.2 | 60.65 | 61.9
12.6 | 45.6429 | 47 | 49.6 | 51.6 | 53.26 | 57.975 | 60.3 | 61.7
12.7 | 45.4 | 46.75 | 49.3 | 51.25 | 52.96 | 57.725 | 60 | 61.38
12.8 | 45.0333 | 46.4 | 48.96 | 50.9 | 52.575 | 57.375 | 59.62 | 61
12.9 | 44.775 | 46.0667 | 48.6 | 50.55 | 52.26 | 57.1 | 59.3 | 60.74
13 | 44.4714 | 45.75 | 48.3 | 50.26 | 51.9667 | 56.75 | 59.16 | 60.4
13.1 | 44.14 | 45.5 | 48.0143 | 49.9 | 51.6 | 56.45 | 58.85 | 60.15
13.2 | 43.775 | 45.16 | 47.65 | 49.6 | 51.25 | 56.12 | 58.475 | 59.96
13.3 | 43.56 | 44.825 | 47.34 | 49.3333 | 51 | 55.8 | 58.325 | 59.65
13.4 | 43.18 | 44.45 | 46.95 | 48.98 | 50.66 | 55.5667 | 57.975 | 59.375
13.5 | 42.875 | 44.2 | 46.7 | 48.65 | 50.46 | 55.2667 | 57.7333 | 59.125
13.6 | 42.5714 | 43.95 | 46.42 | 48.42 | 50.06 | 54.9125 | 57.4 | 58.82
13.7 | 42.24 | 43.575 | 46.1 | 48.0286 | 49.8 | 54.65 | 57.125 | 58.5286
13.8 | 41.975 | 43.34 | 45.8333 | 47.7 | 49.48 | 54.3714 | 56.86 | 58.24
13.9 | 41.725 | 43.04 | 45.5 | 47.375 | 49.125 | 54.04 | 56.6 | 57.9667
14 | 41.38 | 42.675 | 45.175 | 47.14 | 48.8333 | 53.775 | 56.18 | 57.64
14.1 | 41.025 | 42.425 | 44.8 | 46.8429 | 48.475 | 53.46 | 55.9333 | 57.45
14.2 | 40.78 | 42.18 | 44.575 | 46.6 | 48.225 | 53.1857 | 55.6 | 57.1
14.3 | 40.45 | 41.96 | 44.26 | 46.125 | 47.88 | 52.94 | 55.45 | 56.8
14.4 | 40.2333 | 41.58 | 43.95 | 45.875 | 47.5333 | 52.6286 | 55.1 | 56.5
14.5 | 40 | 41.3667 | 43.72 | 45.6625 | 47.175 | 52.32 | 54.7667 | 56.325
14.6 | 39.75 | 41 | 43.4 | 45.34 | 46.95 | 52.075 | 54.6 | 56.0667
14.7 | 39.4 | 40.6286 | 43.075 | 45.1 | 46.675 | 51.775 | 54.3 | 55.75
14.8 | 39.125 | 40.34 | 42.8333 | 44.7667 | 46.45 | 51.5333 | 54.0429 | 55.5
14.9 | 38.9 | 40.1 | 42.575 | 44.45 | 46.2 | 51.2333 | 53.75 | 55.175
15 | 38.6 | 39.9 | 42.2375 | 44.2 | 45.9 | 50.975 | 53.5333 | 54.9
15.1 | 38.36 | 39.65 | 42.02 | 43.94 | 45.5667 | 50.6333 | 53.18 | 54.625
15.2 | 38.125 | 39.4 | 41.68 | 43.6714 | 45.3 | 50.4 | 52.925 | 54.375
15.3 | 37.82 | 39.125 | 41.42 | 43.4 | 45.1 | 50.15 | 52.6667 | 54.125
15.4 | 37.6 | 38.86 | 41.08 | 43.075 | 44.7 | 49.8 | 52.375 | 53.8
15.5 | 37.275 | 38.6 | 40.825 | 42.7 | 44.5 | 49.55 | 52.125 | 53.625
15.6 | 37.02 | 38.3 | 40.58 | 42.5286 | 44.15 | 49.2 | 51.775 | 53.325
15.7 | 36.725 | 38.025 | 40.28 | 42.3 | 43.8667 | 48.9667 | 51.55 | 53.075
15.8 | 36.6 | 37.775 | 40.05 | 41.9286 | 43.6 | 48.8 | 51.3 | 52.8
15.9 | 36.3143 | 37.525 | 39.85 | 41.7 | 43.35 | 48.4667 | 51.1 | 52.575
16 | 35.96 | 37.225 | 39.5 | 41.4125 | 43.12 | 48.225 | 50.76 | 52.34
16.1 | 35.775 | 36.975 | 39.3 | 41.16 | 42.78 | 47.95 | 50.46 | 52
16.2 | 35.44 | 36.675 | 38.95 | 40.8714 | 42.575 | 47.65 | 50.3 | 51.78
16.3 | 35.275 | 36.425 | 38.7 | 40.62 | 42.28 | 47.4 | 50 | 51.5
16.4 | 34.975 | 36.175 | 38.5167 | 40.4286 | 42.075 | 47.1333 | 49.7 | 51.2
16.5 | 34.8 | 35.9429 | 38.2 | 40.12 | 41.8 | 46.9 | 49.52 | 51
16.6 | 34.55 | 35.7 | 37.9 | 39.88 | 41.5833 | 46.7 | 49.275 | 50.8
16.7 | 34.275 | 35.4875 | 37.725 | 39.575 | 41.3 | 46.3857 | 49.025 | 50.5
16.8 | 34.08 | 35.34 | 37.475 | 39.3 | 41.025 | 46.04 | 48.775 | 50.3
16.9 | 33.9 | 35.0667 | 37.225 | 39.1 | 40.7 | 45.8 | 48.525 | 50.05
17 | 33.6 | 34.8 | 36.975 | 38.8429 | 40.45 | 45.66 | 48.28 | 49.8
17.1 | 33.4 | 34.6 | 36.7333 | 38.54 | 40.3 | 45.3375 | 47.98 | 49.5571
17.2 | 33.14 | 34.275 | 36.475 | 38.32 | 40.04 | 45.1 | 47.75 | 49.2429
17.3 | 32.9 | 34.1 | 36.28 | 38.15 | 39.7 | 44.8333 | 47.48 | 48.94
17.4 | 32.65 | 33.85 | 36.02 | 37.825 | 39.5 | 44.6 | 47.25 | 48.76
17.5 | 32.5 | 33.54 | 35.78 | 37.6 | 39.22 | 44.325 | 47 | 48.44
17.6 | 32.2667 | 33.325 | 35.5 | 37.3333 | 38.9333 | 43.975 | 46.65 | 48.25
17.7 | 32.025 | 33.0667 | 35.2 | 37.06 | 38.72 | 43.8 | 46.4667 | 48
17.8 | 31.7 | 32.875 | 35.0667 | 36.8333 | 38.375 | 43.575 | 46.225 | 47.75
17.9 | 31.6 | 32.65 | 34.8 | 36.54 | 38.175 | 43.35 | 46 | 47.46
18 | 31.4 | 32.46 | 34.55 | 36.3167 | 37.9714 | 43.075 | 45.7 | 47.225
18.1 | 31.14 | 32.15 | 34.3429 | 36.16 | 37.72 | 42.825 | 45.475 | 47.1
18.2 | 30.9 | 32.1 | 34.1333 | 35.85 | 37.5 | 42.5 | 45.325 | 46.8667
18.3 | 30.725 | 31.85 | 33.8857 | 35.6429 | 37.275 | 42.3 | 44.975 | 46.6
18.4 | 30.5 | 31.6 | 33.7 | 35.45 | 37.1 | 42.15 | 44.8 | 46.35
18.5 | 30.2 | 31.3333 | 33.4333 | 35.15 | 36.875 | 41.9 | 44.575 | 46.1
18.6 | 30.0143 | 31.1667 | 33.2 | 35 | 36.58 | 41.6 | 44.325 | 45.9
18.7 | 29.84 | 30.9286 | 32.9857 | 34.825 | 36.4 | 41.4 | 44.08 | 45.725
18.8 | 29.7 | 30.7444 | 32.86 | 34.62 | 36.1714 | 41.2 | 43.86 | 45.475
18.9 | 29.5 | 30.575 | 32.575 | 34.4143 | 35.98 | 41 | 43.6 | 45.225
19 | 29.3 | 30.3333 | 32.46 | 34.16 | 35.7 | 40.75 | 43.4 | 44.975
19.1 | 29.12 | 30.1 | 32.18 | 33.9 | 35.44 | 40.5 | 43.2 | 44.725
19.2 | 28.8667 | 29.9 | 31.96 | 33.72 | 35.2625 | 40.275 | 42.975 | 44.62
19.3 | 28.725 | 29.8 | 31.74 | 33.5 | 35 | 40.05 | 42.74 | 44.375
19.4 | 28.5 | 29.52 | 31.52 | 33.275 | 34.8667 | 39.925 | 42.4857 | 44.1667
19.5 | 28.3 | 29.3 | 31.25 | 33.025 | 34.6 | 39.575 | 42.3 | 43.9
19.6 | 28.15 | 29.15 | 31.04 | 32.85 | 34.3667 | 39.425 | 42.0778 | 43.6833
19.7 | 27.94 | 29 | 30.9143 | 32.6 | 34.1375 | 39.2 | 41.85 | 43.4
19.8 | 27.7857 | 28.7667 | 30.675 | 32.44 | 33.975 | 39.025 | 41.65 | 43.225
19.9 | 27.44 | 28.55 | 30.5 | 32.2 | 33.8 | 38.7667 | 41.45 | 42.94
20 | 27.32 | 28.3667 | 30.32 | 32.04 | 33.625 | 38.5 | 41.24 | 42.85
20.1 | 27.2 | 28.15 | 30.1 | 31.8 | 33.35 | 38.3 | 41 | 42.5
20.2 | 27 | 28.02 | 29.9 | 31.6 | 33.08 | 38.2 | 40.8 | 42.3
20.3 | 26.8 | 27.8667 | 29.7 | 31.4 | 32.8667 | 37.96 | 40.55 | 42.14
20.4 | 26.6 | 27.625 | 29.5 | 31.2 | 32.68 | 37.72 | 40.35 | 41.9444
20.5 | 26.5 | 27.45 | 29.25 | 30.95 | 32.4 | 37.46 | 40.2 | 41.7286
20.6 | 26.3 | 27.225 | 29.1 | 30.8 | 32.34 | 37.3 | 39.95 | 41.5
20.7 | 26.16 | 27 | 28.95 | 30.65 | 32.0857 | 37.025 | 39.7 | 41.32
20.8 | 25.925 | 26.875 | 28.8 | 30.42 | 31.84 | 36.8333 | 39.5 | 41.1
20.9 | 25.7 | 26.725 | 28.5667 | 30.175 | 31.7111 | 36.65 | 39.3 | 40.925
21 | 25.6 | 26.5 | 28.425 | 30 | 31.5 | 36.475 | 39.06 | 40.7
21.1 | 25.375 | 26.325 | 28.1667 | 29.9 | 31.35 | 36.225 | 38.8667 | 40.46
21.2 | 25.225 | 26.1 | 28.1 | 29.625 | 31.1 | 36.05 | 38.7667 | 40.2714
21.3 | 25 | 26 | 27.875 | 29.5 | 30.9 | 35.74 | 38.4667 | 40.0571
21.4 | 24.9 | 25.86 | 27.625 | 29.24 | 30.7 | 35.62 | 38.28 | 39.9
21.5 | 24.625 | 25.6571 | 27.475 | 29.1 | 30.475 | 35.425 | 38.0429 | 39.74
21.6 | 24.575 | 25.45 | 27.3333 | 28.84 | 30.3 | 35.225 | 37.8286 | 39.5
21.7 | 24.325 | 25.2571 | 27.1 | 28.725 | 30.125 | 35 | 37.62 | 39.275
21.8 | 24.2333 | 25.1 | 26.94 | 28.5 | 29.975 | 34.88 | 37.44 | 39.1
21.9 | 24.1 | 24.9667 | 26.7 | 28.3 | 29.8 | 34.65 | 37.22 | 38.8
22 | 23.9333 | 24.8 | 26.55 | 28.2 | 29.625 | 34.44 | 37.08 | 38.7
22.1 | 23.8 | 24.6333 | 26.4 | 28 | 29.4 | 34.3 | 36.8833 | 38.5
22.2 | 23.5 | 24.46 | 26.275 | 27.86 | 29.2 | 34.05 | 36.6571 | 38.28
22.3 | 23.4 | 24.3 | 26.025 | 27.6 | 29 | 33.825 | 36.425 | 38
22.4 | 23.25 | 24.16 | 25.9 | 27.425 | 28.85 | 33.65 | 36.3 | 37.8667
22.5 | 23.075 | 24.0222 | 25.725 | 27.26 | 28.7 | 33.44 | 36.0833 | 37.5857
22.6 | 22.9 | 23.82 | 25.62 | 27.025 | 28.425 | 33.2 | 35.9 | 37.5
22.7 | 22.7 | 23.6 | 25.4 | 26.9 | 28.375 | 33.1333 | 35.7 | 37.2667
22.8 | 22.7 | 23.42 | 25.2167 | 26.7333 | 28.125 | 32.9 | 35.5 | 37.1
22.9 | 22.475 | 23.3 | 25.05 | 26.6 | 27.975 | 32.675 | 35.2857 | 36.85
23 | 22.3 | 23.1 | 24.8571 | 26.425 | 27.8 | 32.525 | 35.1 | 36.6333
23.1 | 22.2 | 23 | 24.75 | 26.2333 | 27.6333 | 32.32 | 34.95 | 36.5167
23.2 | 22 | 22.8667 | 24.54 | 26.05 | 27.475 | 32.2 | 34.775 | 36.35
23.3 | 21.84 | 22.7 | 24.4 | 25.95 | 27.32 | 32.02 | 34.625 | 36.2167
23.4 | 21.8 | 22.55 | 24.25 | 25.72 | 27.1 | 31.75 | 34.3571 | 35.975
23.5 | 21.5889 | 22.4 | 24.12 | 25.55 | 27.025 | 31.6 | 34.2 | 35.78
23.6 | 21.46 | 22.2 | 23.9286 | 25.4 | 26.775 | 31.4 | 34.1 | 35.625
23.7 | 21.2571 | 22.1 | 23.8 | 25.275 | 26.625 | 31.25 | 33.82 | 35.375
23.8 | 21.1143 | 22 | 23.6286 | 25.1 | 26.5 | 31.1 | 33.6571 | 35.3
23.9 | 21 | 21.9 | 23.5 | 24.925 | 26.25 | 30.9 | 33.5 | 35.08
24 | 20.8 | 21.7 | 23.3 | 24.8 | 26.1 | 30.75 | 33.325 | 34.85
24.1 | 20.7 | 21.525 | 23.2 | 24.625 | 26 | 30.6 | 33.12 | 34.7333
24.2 | 20.65 | 21.4 | 23 | 24.4667 | 25.85 | 30.38 | 32.9571 | 34.5
24.3 | 20.5 | 21.3 | 22.9 | 24.3 | 25.7 | 30.24 | 32.8 | 34.38
24.4 | 20.35 | 21.1 | 22.68 | 24.2 | 25.5 | 30.08 | 32.6444 | 34.2
24.5 | 20.2 | 20.98 | 22.6 | 24.05 | 25.3 | 29.9 | 32.44 | 34
24.6 | 20.1 | 20.825 | 22.5 | 23.8 | 25.2 | 29.7 | 32.2 | 33.9
24.7 | 19.95 | 20.75 | 22.3 | 23.7 | 25 | 29.55 | 32.1 | 33.68
24.8 | 19.7571 | 20.6 | 22.2 | 23.55 | 24.9 | 29.35 | 31.95 | 33.525
24.9 | 19.7 | 20.5 | 22 | 23.5 | 24.76 | 29.2 | 31.8 | 33.325
25 | 19.5 | 20.3 | 21.86 | 23.25 | 24.5667 | 29 | 31.56 | 33.1429
25.1 | 19.4 | 20.1 | 21.75 | 23.1 | 24.4 | 28.9 | 31.44 | 33
25.2 | 19.225 | 20 | 21.6143 | 23 | 24.2667 | 28.7 | 31.2714 | 32.825
25.3 | 19.15 | 19.9 | 21.4 | 22.8 | 24.1667 | 28.52 | 31.1 | 32.64
25.4 | 19 | 19.8 | 21.2571 | 22.7 | 24.02 | 28.4 | 30.9 | 32.46
25.5 | 18.9 | 19.7 | 21.1143 | 22.5 | 23.88 | 28.26 | 30.8 | 32.3
25.6 | 18.8 | 19.5333 | 21 | 22.4 | 23.6714 | 28.0571 | 30.6375 | 32.16
25.7 | 18.65 | 19.4 | 20.9 | 22.275 | 23.5 | 27.9 | 30.35 | 32
25.8 | 18.525 | 19.3 | 20.7667 | 22.1 | 23.375 | 27.75 | 30.3 | 31.825
25.9 | 18.4 | 19.1 | 20.6143 | 22 | 23.225 | 27.6 | 30.1 | 31.7
26 | 18.3 | 19 | 20.5 | 21.8833 | 23.1 | 27.3667 | 29.9167 | 31.475
26.1 | 18.2 | 18.88 | 20.4 | 21.7 | 22.95 | 27.275 | 29.76 | 31.25
26.2 | 18.0857 | 18.7 | 20.2 | 21.575 | 22.84 | 27.1 | 29.625 | 31.15
26.3 | 17.9429 | 18.58 | 20.1 | 21.425 | 22.66 | 26.925 | 29.46 | 30.96
26.4 | 17.825 | 18.5 | 20 | 21.275 | 22.525 | 26.82 | 29.325 | 30.78
26.5 | 17.7 | 18.4 | 19.85 | 21.1 | 22.4 | 26.7 | 29.175 | 30.66
26.6 | 17.6 | 18.3 | 19.7 | 21.04 | 22.24 | 26.5 | 29.025 | 30.5
26.7 | 17.5 | 18.1 | 19.6 | 20.96 | 22.1 | 26.4 | 28.875 | 30.4
26.8 | 17.35 | 18 | 19.5 | 20.8 | 22 | 26.225 | 28.7 | 30.2
26.9 | 17.2 | 17.9 | 19.4 | 20.675 | 21.8 | 26.1 | 28.58 | 30.025
27 | 17.06 | 17.8 | 19.2 | 20.55 | 21.7 | 25.9 | 28.3 | 29.95
27.1 | 17 | 17.7 | 19.1 | 20.4 | 21.6 | 25.8 | 28.2 | 29.7
27.2 | 16.9 | 17.55 | 18.94 | 20.2286 | 21.5 | 25.675 | 28.04 | 29.55
27.3 | 16.8 | 17.46 | 18.9 | 20.175 | 21.3 | 25.475 | 27.95 | 29.42
27.4 | 16.66 | 17.35 | 18.8 | 20.05 | 21.15 | 25.3 | 27.7833 | 29.3
27.5 | 16.6 | 17.2 | 18.6333 | 19.875 | 21.05 | 25.2 | 27.6 | 29.14
27.6 | 16.5 | 17.175 | 18.5 | 19.72 | 20.9 | 25 | 27.4286 | 29
27.7 | 16.4 | 17.02 | 18.4 | 19.6 | 20.8 | 24.9 | 27.3 | 28.84
27.8 | 16.275 | 16.9333 | 18.3 | 19.5 | 20.7 | 24.8 | 27.125 | 28.6833
27.9 | 16.2 | 16.82 | 18.18 | 19.4 | 20.5 | 24.6 | 27.06 | 28.5
28 | 16.075 | 16.7 | 18 | 19.3 | 20.4 | 24.5 | 26.925 | 28.3833
28.1 | 15.925 | 16.6 | 17.9125 | 19.1 | 20.3 | 24.3 | 26.7667 | 28.3
28.2 | 15.825 | 16.42 | 17.8 | 19 | 20.175 | 24.2 | 26.6167 | 28.0333
28.3 | 15.7 | 16.4 | 17.7286 | 18.85 | 20.025 | 24.1 | 26.46 | 27.96
28.4 | 15.625 | 16.28 | 17.6 | 18.75 | 19.9 | 23.95 | 26.3714 | 27.85
28.5 | 15.5 | 16.2 | 17.5 | 18.7 | 19.7857 | 23.8 | 26.2 | 27.66
28.6 | 15.425 | 16.1 | 17.32 | 18.55 | 19.62 | 23.7 | 26.05 | 27.5
28.7 | 15.3333 | 15.9 | 17.2333 | 18.4 | 19.5429 | 23.6 | 25.925 | 27.375
28.8 | 15.24 | 15.8 | 17.1 | 18.35 | 19.4 | 23.4 | 25.76 | 27.25
28.9 | 15.15 | 15.775 | 17.05 | 18.2 | 19.3 | 23.3 | 25.7 | 27
29 | 15.05 | 15.6 | 16.9 | 18.1 | 19.2 | 23.175 | 25.525 | 26.96
29.1 | 15 | 15.5 | 16.8 | 17.95 | 19.0833 | 23 | 25.375 | 26.8
29.2 | 14.8333 | 15.4 | 16.7 | 17.8571 | 19 | 22.86 | 25.22 | 26.7
29.3 | 14.7857 | 15.35 | 16.6 | 17.7143 | 18.8833 | 22.7 | 25.075 | 26.575
29.4 | 14.7 | 15.3 | 16.5 | 17.7 | 18.7 | 22.6 | 24.92 | 26.4
29.5 | 14.5857 | 15.1857 | 16.4 | 17.56 | 18.6 | 22.5 | 24.775 | 26.2
29.6 | 14.5 | 15.1 | 16.25 | 17.4667 | 18.5 | 22.3333 | 24.625 | 26.1
29.7 | 14.3375 | 14.975 | 16.2 | 17.3 | 18.4 | 22.28 | 24.5667 | 25.9667
29.8 | 14.25 | 14.85 | 16.1 | 17.2 | 18.3 | 22.1 | 24.425 | 25.88
29.9 | 14.2 | 14.775 | 15.925 | 17.1 | 18.2 | 21.975 | 24.3 | 25.725
30 | 14.1 | 14.62 | 15.9 | 16.9667 | 18.1 | 21.825 | 24.2 | 25.6

## Open Field Profiles at 1.2cm
Lateral Profiles
Open field
SSD=100cm
Depth=1.2cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.2 | [7 empty] | 3.1
-28.1 | [7 empty] | 3.1
-28 | [7 empty] | 3.1143
-27.9 | [7 empty] | 3.2
-27.8 | [7 empty] | 3.2
-27.7 | [7 empty] | 3.28
-27.6 | [7 empty] | 3.3
-27.5 | [7 empty] | 3.3
-27.4 | [7 empty] | 3.4
-27.3 | [7 empty] | 3.4
-27.2 | [7 empty] | 3.4143
-27.1 | [7 empty] | 3.5
-27 | [7 empty] | 3.5857
-26.9 | [7 empty] | 3.6
-26.8 | [7 empty] | 3.65
-26.7 | [7 empty] | 3.7
-26.6 | [7 empty] | 3.7125
-26.5 | [7 empty] | 3.8
-26.4 | [7 empty] | 3.8
-26.3 | [7 empty] | 3.9
-26.2 | [7 empty] | 3.9
-26.1 | [7 empty] | 4
-26 | [7 empty] | 4.0333
-25.9 | [7 empty] | 4.1
-25.8 | [7 empty] | 4.2
-25.7 | [7 empty] | 4.2
-25.6 | [7 empty] | 4.3
-25.5 | [7 empty] | 4.3833
-25.4 | [7 empty] | 4.4
-25.3 | [7 empty] | 4.5
-25.2 | [7 empty] | 4.5571
-25.1 | [7 empty] | 4.6
-25 | [7 empty] | 4.7
-24.9 | [7 empty] | 4.8
-24.8 | [7 empty] | 4.9
-24.7 | [7 empty] | 4.95
-24.6 | [7 empty] | 5.1
-24.5 | [7 empty] | 5.1286
-24.4 | [7 empty] | 5.2625
-24.3 | [7 empty] | 5.3
-24.2 | [7 empty] | 5.4143
-24.1 | [7 empty] | 5.55
-24 | [7 empty] | 5.6125
-23.9 | [7 empty] | 5.7429
-23.8 | [7 empty] | 5.9
-23.7 | [7 empty] | 5.9429
-23.6 | [7 empty] | 6.075
-23.5 | [7 empty] | 6.2143
-23.4 | [7 empty] | 6.38
-23.3 | [7 empty] | 6.45
-23.2 | [7 empty] | 6.6
-23.1 | [7 empty] | 6.7714
-23 | [7 empty] | 6.9
-22.9 | [7 empty] | 7.0375
-22.8 | [7 empty] | 7.1714
-22.7 | [7 empty] | 7.3
-22.6 | [7 empty] | 7.5429
-22.5 | [7 empty] | 7.7
-22.4 | [7 empty] | 7.86
-22.3 | [7 empty] | 8
-22.2 | [7 empty] | 8.2
-22.1 | [7 empty] | 8.375
-22 | [7 empty] | 8.5833
-21.9 | [7 empty] | 8.75
-21.8 | [7 empty] | 8.9333
-21.7 | [7 empty] | 9.1833
-21.6 | [7 empty] | 9.4
-21.5 | [7 empty] | 9.58
-21.4 | [7 empty] | 9.86
-21.3 | [7 empty] | 10.14
-21.2 | [7 empty] | 10.48
-21.1 | [7 empty] | 10.8625
-21 | [7 empty] | 11.3429
-20.9 | [7 empty] | 12.075
-20.8 | [7 empty] | 13.18
-20.7 | [7 empty] | 15.2333
-20.6 | [7 empty] | 19.5
-20.5 | [7 empty] | 26.92
-20.4 | [7 empty] | 37.2
-20.3 | [7 empty] | 48.7
-20.2 | [7 empty] | 61.4
-20.1 | [6 empty] | 4.3 | 74.9
-20 | [6 empty] | 4.4 | 85.8
-19.9 | [6 empty] | 4.5 | 94.5667
-19.8 | [6 empty] | 4.52 | 99.96
-19.7 | [6 empty] | 4.6 | 102.9375
-19.6 | [6 empty] | 4.7 | 104.2
-19.5 | [6 empty] | 4.8 | 104.88
-19.4 | [6 empty] | 4.9 | 105.2667
-19.3 | [6 empty] | 4.92 | 105.48
-19.2 | [6 empty] | 5 | 105.7125
-19.1 | [6 empty] | 5.1667 | 105.86
-19 | [6 empty] | 5.24 | 106.0125
-18.9 | [6 empty] | 5.3333 | 106.1
-18.8 | [6 empty] | 5.4 | 106.12
-18.7 | [6 empty] | 5.55 | 106.22
-18.6 | [6 empty] | 5.6 | 106.3857
-18.5 | [6 empty] | 5.7333 | 106.4
-18.4 | [6 empty] | 5.825 | 106.4
-18.3 | [6 empty] | 5.95 | 106.5
-18.2 | [6 empty] | 6.1 | 106.5375
-18.1 | [6 empty] | 6.2 | 106.6
-18 | [6 empty] | 6.35 | 106.7
-17.9 | [6 empty] | 6.4167 | 106.7
-17.8 | [6 empty] | 6.5833 | 106.7
-17.7 | [6 empty] | 6.66 | 106.8
-17.6 | [6 empty] | 6.88 | 106.8
-17.5 | [6 empty] | 6.94 | 106.9
-17.4 | [6 empty] | 7.15 | 106.9
-17.3 | [6 empty] | 7.3 | 106.9
-17.2 | [6 empty] | 7.42 | 106.9375
-17.1 | [6 empty] | 7.6 | 107
-17 | [6 empty] | 7.7667 | 107
-16.9 | [6 empty] | 7.94 | 107.1
-16.8 | [6 empty] | 8.12 | 107.1
-16.7 | [6 empty] | 8.2857 | 107.1
-16.6 | [6 empty] | 8.5 | 107.2
-16.5 | [6 empty] | 8.72 | 107.2
-16.4 | [6 empty] | 8.8857 | 107.2
-16.3 | [6 empty] | 9.1667 | 107.2
-16.2 | [6 empty] | 9.46 | 107.2875
-16.1 | [6 empty] | 9.72 | 107.3
-16 | [6 empty] | 10.1 | 107.3
-15.9 | [6 empty] | 10.5667 | 107.4
-15.8 | [6 empty] | 11.4625 | 107.4
-15.7 | [6 empty] | 12.8818 | 107.4
-15.6 | [6 empty] | 15.14 | 107.5
-15.5 | [6 empty] | 19.7 | 107.5
-15.4 | [6 empty] | 28.65 | 107.5
-15.3 | [6 empty] | 40.1 | 107.5
-15.2 | [6 empty] | 52.5625 | 107.5
-15.1 | [6 empty] | 65.25 | 107.5
-15 | [5 empty] | 3.175 | 78.0833 | 107.5
-14.9 | [5 empty] | 3.2 | 88.32 | 107.5
-14.8 | [5 empty] | 3.225 | 96.1 | 107.5
-14.7 | [5 empty] | 3.3 | 100.3 | 107.5
-14.6 | [5 empty] | 3.4 | 102.22 | 107.5
-14.5 | [5 empty] | 3.4 | 103.26 | 107.5
-14.4 | [5 empty] | 3.5125 | 103.7375 | 107.5
-14.3 | [5 empty] | 3.6 | 104.1 | 107.45
-14.2 | [5 empty] | 3.625 | 104.3 | 107.4
-14.1 | [5 empty] | 3.7 | 104.5 | 107.375
-14 | [5 empty] | 3.8 | 104.5667 | 107.3
-13.9 | [5 empty] | 3.9 | 104.7 | 107.225
-13.8 | [5 empty] | 3.9 | 104.7 | 107.2
-13.7 | [5 empty] | 4.02 | 104.8 | 107.1
-13.6 | [5 empty] | 4.1 | 104.8 | 107.1
-13.5 | [5 empty] | 4.2 | 104.8571 | 107.0111
-13.4 | [5 empty] | 4.3 | 104.9 | 106.9
-13.3 | [5 empty] | 4.4 | 104.9 | 106.9
-13.2 | [5 empty] | 4.5 | 104.9 | 106.8
-13.1 | [5 empty] | 4.6 | 104.9 | 106.7714
-13 | [5 empty] | 4.7 | 104.9 | 106.7
-12.9 | [5 empty] | 4.8 | 104.9 | 106.6
-12.8 | [5 empty] | 4.9 | 104.9 | 106.6
-12.7 | [5 empty] | 5 | 104.9 | 106.5625
-12.6 | [5 empty] | 5.2 | 104.9 | 106.5
-12.5 | [5 empty] | 5.3 | 104.9 | 106.4
-12.4 | [5 empty] | 5.4 | 104.9 | 106.4
-12.3 | [5 empty] | 5.5 | 104.9 | 106.3714
-12.2 | [5 empty] | 5.7 | 104.9 | 106.3
-12.1 | [5 empty] | 5.875 | 104.9 | 106.3
-12 | [5 empty] | 6 | 104.9 | 106.2333
-11.9 | [5 empty] | 6.175 | 104.9429 | 106.2
-11.8 | [5 empty] | 6.3 | 105 | 106.2
-11.7 | [5 empty] | 6.5 | 105 | 106.2
-11.6 | [5 empty] | 6.6667 | 104.9 | 106.1
-11.5 | [5 empty] | 6.825 | 104.9 | 106.1
-11.4 | [5 empty] | 7 | 104.9 | 106.1
-11.3 | [5 empty] | 7.3 | 104.9 | 106.0125
-11.2 | [5 empty] | 7.5 | 104.9 | 106
-11.1 | [5 empty] | 7.7333 | 104.9 | 105.9429
-11 | [5 empty] | 8.05 | 104.9 | 105.9
-10.9 | [5 empty] | 8.48 | 104.9 | 105.9
-10.8 | [5 empty] | 9.025 | 104.9 | 105.9
-10.7 | [5 empty] | 9.9 | 104.9 | 105.8
-10.6 | [5 empty] | 11.425 | 104.84 | 105.8
-10.5 | [5 empty] | 14.4 | 104.8 | 105.7
-10.4 | [5 empty] | 20.6429 | 104.8 | 105.6429
-10.3 | [5 empty] | 30.3286 | 104.8 | 105.6
-10.2 | [5 empty] | 41.9 | 104.7 | 105.5875
-10.1 | [5 empty] | 54.9 | 104.7 | 105.5
-10 | [5 empty] | 67.7 | 104.675 | 105.4
-9.9 |  |  |  |  | 1.7 | 79.3333 | 104.6 | 105.3
-9.8 |  |  |  |  | 1.7 | 89.3 | 104.6 | 105.3
-9.7 |  |  |  |  | 1.7 | 95.35 | 104.5 | 105.2571
-9.6 |  |  |  |  | 1.7 | 98.5 | 104.5 | 105.2
-9.5 |  |  |  |  | 1.8 | 99.9 | 104.5 | 105.1875
-9.4 |  |  |  |  | 1.8 | 100.6 | 104.41 | 105.0571
-9.3 |  |  |  |  | 1.85 | 101.12 | 104.4 | 105
-9.2 |  |  |  |  | 1.9 | 101.4 | 104.3417 | 104.9
-9.1 |  |  |  |  | 1.9 | 101.6 | 104.3 | 104.8
-9 |  |  |  | 1.2 | 2 | 101.9 | 104.3 | 104.8
-8.9 |  |  |  | 1.3 | 2 | 101.925 | 104.3 | 104.78
-8.8 |  |  |  | 1.3 | 2.05 | 102.0571 | 104.3 | 104.7
-8.7 |  |  |  | 1.3 | 2.1 | 102.175 | 104.2 | 104.6
-8.6 |  |  |  | 1.3 | 2.14 | 102.275 | 104.2 | 104.6
-8.5 |  |  |  | 1.4 | 2.2 | 102.3833 | 104.2 | 104.5
-8.4 |  |  |  | 1.4 | 2.3 | 102.5 | 104.2 | 104.5
-8.3 |  |  |  | 1.4 | 2.3 | 102.5 | 104.2 | 104.5
-8.2 |  |  |  | 1.4 | 2.4 | 102.6 | 104.1 | 104.475
-8.1 |  |  |  | 1.5 | 2.45 | 102.625 | 104.1 | 104.4
-8 |  |  |  | 1.5 | 2.5 | 102.7 | 104.1 | 104.3125
-7.9 |  |  | 0.9 | 1.5 | 2.6 | 102.725 | 104.05 | 104.3
-7.8 |  |  | 0.9 | 1.6 | 2.7 | 102.8 | 104 | 104.3
-7.7 |  |  | 0.9 | 1.6 | 2.7 | 102.8 | 103.9167 | 104.3
-7.6 |  |  | 0.975 | 1.7 | 2.8 | 102.8 | 103.9 | 104.2857
-7.5 |  |  | 1 | 1.7 | 2.9571 | 102.8 | 103.9 | 104.2
-7.4 |  |  | 1 | 1.7 | 3 | 102.8 | 103.9 | 104.2
-7.3 |  |  | 1 | 1.8 | 3.1167 | 102.9 | 103.8 | 104.0857
-7.2 |  |  | 1 | 1.8 | 3.225 | 102.9 | 103.8 | 104
-7.1 |  |  | 1.1 | 1.9 | 3.3125 | 102.9 | 103.8 | 104
-7 |  |  | 1.1 | 2 | 3.4429 | 102.9 | 103.7 | 104
-6.9 |  | 0.6 | 1.1 | 2 | 3.6 | 102.9 | 103.7 | 103.95
-6.8 |  | 0.6 | 1.1 | 2.1 | 3.675 | 102.9 | 103.7 | 103.9
-6.7 |  | 0.6 | 1.2 | 2.1667 | 3.8 | 102.8 | 103.6 | 103.9
-6.6 |  | 0.6 | 1.2 | 2.2 | 4 | 102.8 | 103.6 | 103.8
-6.5 |  | 0.6 | 1.2 | 2.3 | 4.1 | 102.8 | 103.55 | 103.7625
-6.4 | 0.4 | 0.6 | 1.3 | 2.4 | 4.3 | 102.8 | 103.5 | 103.7
-6.3 | 0.4 | 0.65 | 1.3 | 2.5 | 4.5 | 102.8 | 103.4167 | 103.7
-6.2 | 0.475 | 0.7 | 1.4 | 2.6 | 4.7 | 102.8 | 103.4 | 103.6625
-6.1 | 0.5 | 0.7 | 1.4 | 2.7 | 4.9 | 102.8 | 103.4 | 103.6
-6 | 0.5 | 0.7 | 1.4 | 2.8 | 5.08 | 102.8 | 103.4 | 103.5
-5.9 | 0.5 | 0.7 | 1.5 | 2.9 | 5.4143 | 102.8 | 103.4 | 103.5
-5.8 | 0.5 | 0.7 | 1.5 | 3 | 5.8143 | 102.7 | 103.2571 | 103.4
-5.7 | 0.5 | 0.7 | 1.6 | 3.1 | 6.475 | 102.7 | 103.2 | 103.4
-5.6 | 0.5 | 0.8 | 1.6 | 3.2 | 7.55 | 102.7 | 103.2 | 103.3
-5.5 | 0.5 | 0.8 | 1.7 | 3.4 | 9.25 | 102.7 | 103.1 | 103.2714
-5.4 | 0.5 | 0.8 | 1.8 | 3.5 | 13 | 102.6 | 103.1 | 103.2
-5.3 | 0.5 | 0.8 | 1.8 | 3.7 | 20.2 | 102.6 | 103.0143 | 103.1
-5.2 | 0.5 | 0.8 | 1.9 | 3.9 | 30.9 | 102.575 | 103 | 103.1
-5.1 | 0.6 | 0.9 | 2 | 4.0333 | 42.8 | 102.5 | 102.9 | 103.075
-5 | 0.6 | 0.9 | 2.1 | 4.2 | 56.05 | 102.5 | 102.9 | 103
-4.9 | 0.6 | 0.9 | 2.1667 | 4.5333 | 68.975 | 102.5 | 102.8857 | 102.9125
-4.8 | 0.6 | 1 | 2.2333 | 4.9 | 80.25 | 102.4 | 102.72 | 102.875
-4.7 | 0.6 | 1 | 2.3 | 5.4333 | 89.2 | 102.3 | 102.7 | 102.8
-4.6 | 0.6 | 1 | 2.4 | 6.3 | 94.825 | 102.3 | 102.6 | 102.8
-4.5 | 0.65 | 1.1 | 2.5333 | 7.9 | 97.4 | 102.2714 | 102.5667 | 102.7
-4.4 | 0.7 | 1.1 | 2.7 | 10.8 | 98.6667 | 102.2 | 102.5 | 102.6
-4.3 | 0.7 | 1.15 | 2.8 | 17.3 | 99.275 | 102.2 | 102.4286 | 102.55
-4.2 | 0.7 | 1.2 | 2.95 | 26.8 | 99.775 | 102.1 | 102.4 | 102.4143
-4.1 | 0.7 | 1.3 | 3.1 | 39.4 | 100 | 102 | 102.3 | 102.3571
-4 | 0.8 | 1.3 | 3.3 | 52.625 | 100.2 | 101.975 | 102.3 | 102.3
-3.9 | 0.8 | 1.4 | 3.55 | 65.8 | 100.3 | 101.9 | 102.2833 | 102.2
-3.8 | 0.8 | 1.4 | 3.8 | 77.6 | 100.4 | 101.8 | 102.2 | 102.15
-3.7 | 0.9 | 1.5 | 4.3 | 87 | 100.425 | 101.8 | 102.05 | 102.1
-3.6 | 0.9 | 1.6 | 5.1 | 93.7 | 100.5 | 101.7 | 102 | 102
-3.5 | 1 | 1.7 | 6.3 | 96.8 | 100.6 | 101.7 | 101.8727 | 101.95
-3.4 | 1 | 1.8 | 9.0333 | 98.1 | 100.7 | 101.6 | 101.8 | 101.9
-3.3 | 1 | 1.9 | 14.5 | 98.85 | 100.7 | 101.6 | 101.8 | 101.8
-3.2 | 1.1 | 2 | 23.6 | 99.3 | 100.8 | 101.575 | 101.7091 | 101.72
-3.1 | 1.2 | 2.1 | 35.5 | 99.5 | 100.8 | 101.5 | 101.7 | 101.7
-3 | 1.2 | 2.3 | 48.1 | 99.7 | 100.9 | 101.4 | 101.6 | 101.6143
-2.9 | 1.3 | 2.475 | 61.8 | 99.8 | 100.9 | 101.4 | 101.5286 | 101.5857
-2.8 | 1.4 | 2.7333 | 74.7 | 99.9 | 100.9 | 101.3 | 101.5 | 101.4429
-2.7 | 1.5 | 3.1 | 86 | 100.0667 | 100.9 | 101.3 | 101.4 | 101.4
-2.6 | 1.6 | 3.8 | 92.55 | 100.2 | 100.9 | 101.2 | 101.4 | 101.375
-2.5 | 1.7 | 5.05 | 96.35 | 100.2 | 100.9 | 101.175 | 101.3667 | 101.3
-2.4 | 1.9 | 7.3 | 97.9333 | 100.2 | 100.875 | 101.1 | 101.3 | 101.3
-2.3 | 2.1 | 12.8 | 98.74 | 100.3 | 100.8667 | 101 | 101.2636 | 101.26
-2.2 | 2.5333 | 21.3 | 99.2 | 100.3 | 100.875 | 101 | 101.1727 | 101.1867
-2.1 | 3.2 | 33.1 | 99.5 | 100.4 | 100.8 | 101 | 101.0778 | 101.0533
-2 | 4.3 | 46.0667 | 99.7 | 100.4 | 100.8 | 100.9 | 100.975 | 101
-1.9 | 6.9 | 60.4 | 99.8 | 100.4 | 100.8 | 100.9 | 100.9 | 100.9909
-1.8 | 12.1 | 73.9 | 99.9 | 100.4 | 100.7 | 100.9 | 100.9 | 100.9
-1.7 | 21.4 | 84.6 | 99.9 | 100.4 | 100.7 | 100.8 | 100.8 | 100.8625
-1.6 | 33 | 92.1667 | 100 | 100.4 | 100.7 | 100.7 | 100.8 | 100.8
-1.5 | 45.8 | 96.3 | 100.1 | 100.3 | 100.6 | 100.675 | 100.7 | 100.7
-1.4 | 59.9 | 98.0667 | 100.1 | 100.3 | 100.6 | 100.6 | 100.68 | 100.7
-1.3 | 73.5 | 98.95 | 100.1 | 100.3 | 100.55 | 100.5 | 100.6 | 100.6455
-1.2 | 84.6 | 99.35 | 100.1 | 100.3 | 100.5 | 100.4667 | 100.54 | 100.5333
-1.1 | 92.35 | 99.7 | 100.1 | 100.3 | 100.4 | 100.4 | 100.45 | 100.4
-1 | 96.2 | 99.8333 | 100.1 | 100.2 | 100.4 | 100.3 | 100.4 | 100.4
-0.9 | 98.1 | 99.9 | 100.1 | 100.2 | 100.3 | 100.3 | 100.4 | 100.3
-0.8 | 99.0333 | 100 | 100.1 | 100.1 | 100.2 | 100.3 | 100.3 | 100.3
-0.7 | 99.4 | 100 | 100.1 | 100.1 | 100.2 | 100.2 | 100.3 | 100.3
-0.6 | 99.65 | 100 | 100 | 100.0667 | 100.2 | 100.1 | 100.2333 | 100.2
-0.5 | 99.8 | 100 | 100 | 100 | 100.2 | 100 | 100.1571 | 100.1625
-0.4 | 99.8333 | 100 | 100 | 100 | 100.1 | 100 | 100.1 | 100.1
-0.3 | 99.9 | 100 | 100 | 100 | 100.1 | 100 | 100.1 | 100.0889
-0.2 | 100 | 100 | 100 | 100 | 100.04 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100.05 | 100 | 100 | 100.0111
0.3 | 99.9 | 100 | 100 | 100 | 100.1 | 100 | 100.1 | 100.1
0.4 | 99.8333 | 100 | 100 | 100 | 100.1 | 100 | 100.1 | 100.1
0.5 | 99.8 | 100 | 100 | 100 | 100.15 | 100 | 100.1571 | 100.2
0.6 | 99.6333 | 100 | 100 | 100.1 | 100.2 | 100.1 | 100.2333 | 100.2
0.7 | 99.375 | 100 | 100.1 | 100.1 | 100.2 | 100.2 | 100.3 | 100.2625
0.8 | 99.0333 | 100 | 100.1 | 100.1 | 100.2 | 100.3 | 100.3 | 100.3
0.9 | 98.2 | 99.9 | 100.1 | 100.2 | 100.3 | 100.3 | 100.4 | 100.3429
1 | 96.2 | 99.8333 | 100.1 | 100.2 | 100.4 | 100.32 | 100.4 | 100.4
1.1 | 92.35 | 99.7 | 100.1 | 100.3 | 100.4 | 100.4 | 100.45 | 100.4333
1.2 | 84.3667 | 99.35 | 100.1 | 100.3 | 100.5 | 100.4667 | 100.54 | 100.5
1.3 | 73.5 | 98.95 | 100.1 | 100.3 | 100.55 | 100.5 | 100.6 | 100.5769
1.4 | 59.9 | 98.0667 | 100.1 | 100.3 | 100.6 | 100.6 | 100.68 | 100.6875
1.5 | 45.8 | 96.3 | 100.1 | 100.3 | 100.6 | 100.7 | 100.7 | 100.7286
1.6 | 33 | 92.2333 | 100 | 100.4 | 100.7 | 100.7 | 100.8 | 100.8
1.7 | 21.4 | 84.6 | 99.9 | 100.4 | 100.7 | 100.8 | 100.8 | 100.8
1.8 | 12.15 | 72.9 | 99.9 | 100.4 | 100.7 | 100.9 | 100.9 | 100.9
1.9 | 7 | 60.4 | 99.8 | 100.4 | 100.7167 | 100.9 | 100.9 | 100.9267
2 | 4.3 | 46.8667 | 99.7 | 100.4 | 100.8 | 100.9 | 100.9769 | 100.9933
2.1 | 3.2 | 33.1 | 99.4667 | 100.4 | 100.8 | 101 | 101.0778 | 101.06
2.2 | 2.5333 | 21.5 | 99.2 | 100.3 | 100.875 | 101 | 101.1727 | 101.1267
2.3 | 2.1 | 12.8 | 98.7 | 100.3 | 100.8571 | 101 | 101.2 | 101.1933
2.4 | 1.9 | 7.3 | 97.9333 | 100.2 | 100.875 | 101.1 | 101.3 | 101.2692
2.5 | 1.7 | 5.05 | 96.35 | 100.2 | 100.9 | 101.175 | 101.3667 | 101.3
2.6 | 1.6 | 3.8 | 92.4667 | 100.15 | 100.9 | 101.2 | 101.4 | 101.3125
2.7 | 1.5 | 3.1 | 85.3 | 100.05 | 100.9 | 101.3 | 101.4 | 101.4
2.8 | 1.4 | 2.7 | 74.7 | 99.95 | 100.9 | 101.3 | 101.5 | 101.475
2.9 | 1.3 | 2.4667 | 62.5 | 99.8 | 100.9 | 101.4 | 101.5286 | 101.6
3 | 1.2 | 2.3 | 48.1 | 99.7 | 100.9 | 101.4 | 101.6 | 101.6
3.1 | 1.2 | 2.1 | 36 | 99.5 | 100.8 | 101.5 | 101.7 | 101.7
3.2 | 1.1 | 2 | 23.6 | 99.3 | 100.8 | 101.6 | 101.71 | 101.7667
3.3 | 1 | 1.9 | 14.8 | 98.85 | 100.7 | 101.6 | 101.8 | 101.8143
3.4 | 1 | 1.8 | 9.1333 | 98.1 | 100.7 | 101.6 | 101.8 | 101.9
3.5 | 1 | 1.7 | 6.4 | 96.8 | 100.6 | 101.7 | 101.8727 | 101.9
3.6 | 0.9 | 1.6 | 5.1 | 93.8 | 100.5 | 101.7 | 102 | 102.0143
3.7 | 0.9 | 1.5 | 4.3 | 87.1 | 100.425 | 101.8 | 102.05 | 102.1
3.8 | 0.8 | 1.4 | 3.8 | 77.3 | 100.4 | 101.8 | 102.2 | 102.1857
3.9 | 0.8 | 1.4 | 3.55 | 65.8667 | 100.3 | 101.9 | 102.2833 | 102.2143
4 | 0.8 | 1.3 | 3.3 | 51.8 | 100.2 | 101.975 | 102.3 | 102.3
4.1 | 0.7 | 1.3 | 3.1 | 39.15 | 100 | 102 | 102.3 | 102.3875
4.2 | 0.7 | 1.2 | 2.95 | 26.8 | 99.775 | 102.1 | 102.4 | 102.45
4.3 | 0.7 | 1.15 | 2.8 | 17.3 | 99.3 | 102.2 | 102.4333 | 102.5
4.4 | 0.7 | 1.1 | 2.7 | 10.8 | 98.75 | 102.2 | 102.5 | 102.6
4.5 | 0.65 | 1.1 | 2.6 | 7.9 | 97.4 | 102.2714 | 102.5667 | 102.72
4.6 | 0.6 | 1 | 2.4 | 6.3 | 94.975 | 102.3 | 102.6 | 102.8
4.7 | 0.6 | 1 | 2.3 | 5.4333 | 89.35 | 102.3 | 102.7 | 102.8
4.8 | 0.6 | 1 | 2.2 | 4.9 | 80.25 | 102.4 | 102.72 | 102.9
4.9 | 0.6 | 0.9 | 2.1667 | 4.5 | 68.975 | 102.5 | 102.8857 | 102.9429
5 | 0.6 | 0.9 | 2.0667 | 4.2 | 56.05 | 102.5 | 102.9 | 103
5.1 | 0.6 | 0.9 | 2 | 4.0333 | 42.9667 | 102.5 | 102.9 | 103.1
5.2 | 0.5 | 0.8 | 1.9 | 3.85 | 30.3167 | 102.6 | 103 | 103.1
5.3 | 0.5 | 0.8 | 1.8 | 3.7 | 20.6 | 102.6 | 103.0167 | 103.1375
5.4 | 0.5 | 0.8 | 1.8 | 3.5 | 12.75 | 102.6 | 103.1 | 103.2
5.5 | 0.5 | 0.8 | 1.7 | 3.4 | 9.25 | 102.6 | 103.1 | 103.2
5.6 | 0.5 | 0.8 | 1.6 | 3.2 | 7.5333 | 102.7 | 103.2 | 103.3
5.7 | 0.5 | 0.7 | 1.6 | 3.1 | 6.475 | 102.7 | 103.2 | 103.3833
5.8 | 0.5 | 0.7 | 1.5 | 3 | 5.8143 | 102.7 | 103.2667 | 103.4
5.9 | 0.5 | 0.7 | 1.5 | 2.9 | 5.4714 | 102.8 | 103.4 | 103.4286
6 | 0.5 | 0.7 | 1.4 | 2.8 | 5.0667 | 102.8 | 103.4 | 103.5
6.1 | 0.5 | 0.7 | 1.4 | 2.7 | 4.875 | 102.8 | 103.4 | 103.6
6.2 | 0.46667 | 0.7 | 1.4 | 2.6 | 4.7 | 102.8 | 103.4 | 103.7
6.3 | 0.4 | 0.65 | 1.3 | 2.5 | 4.5 | 102.8 | 103.4 | 103.7
6.4 | 0.4 | 0.6 | 1.3 | 2.4 | 4.3 | 102.8 | 103.4833 | 103.7
6.5 |  | 0.6 | 1.2 | 2.3 | 4.1 | 102.8 | 103.55 | 103.7
6.6 |  | 0.6 | 1.2 | 2.2 | 4 | 102.8 | 103.6 | 103.84
6.7 |  | 0.6 | 1.2 | 2.1667 | 3.85 | 102.8 | 103.6 | 103.9
6.8 |  | 0.6 | 1.1 | 2.1 | 3.675 | 102.9 | 103.7 | 103.9
6.9 |  | 0.6 | 1.1 | 2 | 3.6 | 102.9 | 103.7 | 103.9833
7 |  |  | 1.1 | 2 | 3.4333 | 102.9 | 103.7 | 104
7.1 |  |  | 1.1 | 1.9 | 3.3111 | 102.9 | 103.8 | 104
7.2 |  |  | 1 | 1.8 | 3.225 | 102.9 | 103.8 | 104
7.3 |  |  | 1 | 1.8 | 3.1167 | 102.9 | 103.8 | 104.1
7.4 |  |  | 1 | 1.7 | 3 | 102.8 | 103.9 | 104.15
7.5 |  |  | 1 | 1.7 | 2.9429 | 102.8 | 103.9 | 104.2
7.6 |  |  | 0.96 | 1.7 | 2.8 | 102.8 | 103.9 | 104.3
7.7 |  |  | 0.9 | 1.6 | 2.7 | 102.8 | 103.9167 | 104.3
7.8 |  |  | 0.9 | 1.6 | 2.7 | 102.8 | 104 | 104.3
7.9 |  |  | 0.9 | 1.5 | 2.6 | 102.725 | 104.05 | 104.3
8 |  |  |  | 1.5 | 2.5 | 102.7 | 104.1 | 104.3
8.1 |  |  |  | 1.5 | 2.4333 | 102.625 | 104.1 | 104.375
8.2 |  |  |  | 1.4 | 2.4 | 102.6 | 104.1 | 104.4143
8.3 |  |  |  | 1.4 | 2.3 | 102.5 | 104.2 | 104.5
8.4 |  |  |  | 1.4 | 2.3 | 102.5 | 104.2 | 104.5
8.5 |  |  |  | 1.4 | 2.2 | 102.3714 | 104.2 | 104.5
8.6 |  |  |  | 1.3 | 2.14 | 102.275 | 104.2 | 104.5714
8.7 |  |  |  | 1.3 | 2.1 | 102.16 | 104.2 | 104.6
8.8 |  |  |  | 1.3 | 2.025 | 102.05 | 104.3 | 104.7
8.9 |  |  |  | 1.3 | 2 | 101.925 | 104.3 | 104.7
9 |  |  |  | 1.2 | 2 | 101.9 | 104.3 | 104.8
9.1 |  |  |  |  | 1.9 | 101.6 | 104.3 | 104.8
9.2 |  |  |  |  | 1.9 | 101.4 | 104.3455 | 104.9
9.3 |  |  |  |  | 1.85 | 101.12 | 104.4 | 105
9.4 |  |  |  |  | 1.8 | 100.6 | 104.41 | 105.0875
9.5 |  |  |  |  | 1.8 | 99.9 | 104.5 | 105.125
9.6 |  |  |  |  | 1.7 | 98.4 | 104.5 | 105.2
9.7 |  |  |  |  | 1.7 | 95.3 | 104.5 | 105.2
9.8 |  |  |  |  | 1.7 | 89.3 | 104.6 | 105.3
9.9 |  |  |  |  | 1.7 | 79.4 | 104.6 | 105.3
10 | [5 empty] | 67.05 | 104.68 | 105.4
10.1 | [5 empty] | 54.2 | 104.7 | 105.4857
10.2 | [5 empty] | 42.14 | 104.7 | 105.6
10.3 | [5 empty] | 30.5714 | 104.8 | 105.6
10.4 | [5 empty] | 20.7714 | 104.8 | 105.6875
10.5 | [5 empty] | 14.42 | 104.8 | 105.7143
10.6 | [5 empty] | 11.425 | 104.85 | 105.8
10.7 | [5 empty] | 9.9 | 104.9 | 105.8
10.8 | [5 empty] | 9.025 | 104.9 | 105.9
10.9 | [5 empty] | 8.45 | 104.9 | 105.9
11 | [5 empty] | 8.075 | 104.9 | 105.9
11.1 | [5 empty] | 7.7333 | 104.9 | 105.9
11.2 | [5 empty] | 7.5 | 104.9 | 106
11.3 | [5 empty] | 7.3 | 104.9 | 106.0429
11.4 | [5 empty] | 7 | 104.9 | 106.1
11.5 | [5 empty] | 6.82 | 104.9 | 106.1
11.6 | [5 empty] | 6.6667 | 104.9 | 106.1
11.7 | [5 empty] | 6.5 | 105 | 106.1667
11.8 | [5 empty] | 6.3 | 105 | 106.2
11.9 | [5 empty] | 6.175 | 104.9429 | 106.2
12 | [5 empty] | 6 | 104.9 | 106.2714
12.1 | [5 empty] | 5.875 | 104.9 | 106.3
12.2 | [5 empty] | 5.7 | 104.9 | 106.3
12.3 | [5 empty] | 5.5 | 104.9 | 106.4
12.4 | [5 empty] | 5.4 | 104.9 | 106.4
12.5 | [5 empty] | 5.3 | 104.9 | 106.4167
12.6 | [5 empty] | 5.2 | 104.9 | 106.5
12.7 | [5 empty] | 5 | 104.9 | 106.5
12.8 | [5 empty] | 4.9 | 104.9 | 106.6
12.9 | [5 empty] | 4.8 | 104.9 | 106.6125
13 | [5 empty] | 4.7 | 104.9 | 106.7
13.1 | [5 empty] | 4.6 | 104.9 | 106.7
13.2 | [5 empty] | 4.5 | 104.9 | 106.8
13.3 | [5 empty] | 4.4 | 104.9 | 106.8714
13.4 | [5 empty] | 4.3 | 104.9 | 106.9125
13.5 | [5 empty] | 4.2 | 104.8571 | 107
13.6 | [5 empty] | 4.1 | 104.8 | 107.0462
13.7 | [5 empty] | 4 | 104.8 | 107.1
13.8 | [5 empty] | 3.96 | 104.7 | 107.1375
13.9 | [5 empty] | 3.9 | 104.7 | 107.2625
14 | [5 empty] | 3.8 | 104.5667 | 107.3
14.1 | [5 empty] | 3.7 | 104.4833 | 107.3143
14.2 | [5 empty] | 3.625 | 104.3 | 107.3857
14.3 | [5 empty] | 3.6 | 104.0857 | 107.4
14.4 | [5 empty] | 3.5125 | 103.7 | 107.5
14.5 | [5 empty] | 3.4 | 103.2 | 107.5
14.6 | [5 empty] | 3.4 | 102.22 | 107.5
14.7 | [5 empty] | 3.3 | 100.3 | 107.5
14.8 | [5 empty] | 3.225 | 96.1 | 107.5
14.9 | [5 empty] | 3.2 | 88.52 | 107.5
15 | [5 empty] | 3.175 | 78.1333 | 107.5
15.1 | [6 empty] | 65.55 | 107.5
15.2 | [6 empty] | 52.5571 | 107.5
15.3 | [6 empty] | 39.4333 | 107.5
15.4 | [6 empty] | 28.8 | 107.5
15.5 | [6 empty] | 19.8 | 107.5
15.6 | [6 empty] | 15.14 | 107.4875
15.7 | [6 empty] | 12.8818 | 107.4
15.8 | [6 empty] | 11.4625 | 107.4
15.9 | [6 empty] | 10.65 | 107.3857
16 | [6 empty] | 10.1 | 107.3
16.1 | [6 empty] | 9.72 | 107.3
16.2 | [6 empty] | 9.46 | 107.3
16.3 | [6 empty] | 9.16 | 107.225
16.4 | [6 empty] | 8.9667 | 107.2
16.5 | [6 empty] | 8.7167 | 107.2
16.6 | [6 empty] | 8.55 | 107.1875
16.7 | [6 empty] | 8.2833 | 107.1
16.8 | [6 empty] | 8.1167 | 107.1
16.9 | [6 empty] | 7.94 | 107.1
17 | [6 empty] | 7.76 | 107
17.1 | [6 empty] | 7.6 | 107
17.2 | [6 empty] | 7.42 | 107
17.3 | [6 empty] | 7.3 | 106.9
17.4 | [6 empty] | 7.15 | 106.9
17.5 | [6 empty] | 7.0333 | 106.8625
17.6 | [6 empty] | 6.8667 | 106.8
17.7 | [6 empty] | 6.66 | 106.7625
17.8 | [6 empty] | 6.58 | 106.7
17.9 | [6 empty] | 6.4143 | 106.7
18 | [6 empty] | 6.35 | 106.6625
18.1 | [6 empty] | 6.2 | 106.6
18.2 | [6 empty] | 6.1 | 106.5
18.3 | [6 empty] | 5.95 | 106.4714
18.4 | [6 empty] | 5.825 | 106.4
18.5 | [6 empty] | 5.72 | 106.4
18.6 | [6 empty] | 5.6 | 106.36
18.7 | [6 empty] | 5.55 | 106.2
18.8 | [6 empty] | 5.4 | 106.1
18.9 | [6 empty] | 5.3333 | 106.1
19 | [6 empty] | 5.24 | 105.975
19.1 | [6 empty] | 5.1667 | 105.9
19.2 | [6 empty] | 5 | 105.6714
19.3 | [6 empty] | 4.9167 | 105.5
19.4 | [6 empty] | 4.9 | 105.3
19.5 | [6 empty] | 4.8 | 104.9
19.6 | [6 empty] | 4.7 | 104.2667
19.7 | [6 empty] | 4.6 | 102.775
19.8 | [6 empty] | 4.5167 | 100.1
19.9 | [6 empty] | 4.5 | 94.8
20 | [6 empty] | 4.4 | 86.04
20.1 | [6 empty] | 4.3 | 74.58
20.2 | [7 empty] | 61.5667
20.3 | [7 empty] | 49.44
20.4 | [7 empty] | 36.6857
20.5 | [7 empty] | 26.66
20.6 | [7 empty] | 19.6
20.7 | [7 empty] | 15.3714
20.8 | [7 empty] | 13.2333
20.9 | [7 empty] | 12.1
21 | [7 empty] | 11.35
21.1 | [7 empty] | 10.8875
21.2 | [7 empty] | 10.45
21.3 | [7 empty] | 10.14
21.4 | [7 empty] | 9.84
21.5 | [7 empty] | 9.65
21.6 | [7 empty] | 9.36
21.7 | [7 empty] | 9.16
21.8 | [7 empty] | 8.9833
21.9 | [7 empty] | 8.7333
22 | [7 empty] | 8.55
22.1 | [7 empty] | 8.4
22.2 | [7 empty] | 8.16
22.3 | [7 empty] | 7.9714
22.4 | [7 empty] | 7.8167
22.5 | [7 empty] | 7.6714
22.6 | [7 empty] | 7.5
22.7 | [7 empty] | 7.3714
22.8 | [7 empty] | 7.2375
22.9 | [7 empty] | 7
23 | [7 empty] | 6.8714
23.1 | [7 empty] | 6.7375
23.2 | [7 empty] | 6.5833
23.3 | [7 empty] | 6.5
23.4 | [7 empty] | 6.34
23.5 | [7 empty] | 6.25
23.6 | [7 empty] | 6.1
23.7 | [7 empty] | 6
23.8 | [7 empty] | 5.8714
23.9 | [7 empty] | 5.7167
24 | [7 empty] | 5.6667
24.1 | [7 empty] | 5.5143
24.2 | [7 empty] | 5.3857
24.3 | [7 empty] | 5.3
24.4 | [7 empty] | 5.225
24.5 | [7 empty] | 5.1
24.6 | [7 empty] | 5.0571
24.7 | [7 empty] | 5
24.8 | [7 empty] | 4.8667
24.9 | [7 empty] | 4.8
25 | [7 empty] | 4.7
25.1 | [7 empty] | 4.6
25.2 | [7 empty] | 4.525
25.3 | [7 empty] | 4.5
25.4 | [7 empty] | 4.4
25.5 | [7 empty] | 4.34
25.6 | [7 empty] | 4.2857
25.7 | [7 empty] | 4.2
25.8 | [7 empty] | 4.16
25.9 | [7 empty] | 4.1
26 | [7 empty] | 4
26.1 | [7 empty] | 4
26.2 | [7 empty] | 3.9429
26.3 | [7 empty] | 3.9
26.4 | [7 empty] | 3.8429
26.5 | [7 empty] | 3.8
26.6 | [7 empty] | 3.7
26.7 | [7 empty] | 3.7
26.8 | [7 empty] | 3.6143
26.9 | [7 empty] | 3.6
27 | [7 empty] | 3.55
27.1 | [7 empty] | 3.5
27.2 | [7 empty] | 3.4857
27.3 | [7 empty] | 3.4
27.4 | [7 empty] | 3.4
27.5 | [7 empty] | 3.3
27.6 | [7 empty] | 3.3
27.7 | [7 empty] | 3.24
27.8 | [7 empty] | 3.2
27.9 | [7 empty] | 3.2
28 | [7 empty] | 3.1
28.1 | [7 empty] | 3.1

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 4.3
-28.8 | [7 empty] | 4.3
-28.7 | [7 empty] | 4.4
-28.6 | [7 empty] | 4.4
-28.5 | [7 empty] | 4.4875
-28.4 | [7 empty] | 4.6
-28.3 | [7 empty] | 4.6
-28.2 | [7 empty] | 4.6857
-28.1 | [7 empty] | 4.7143
-28 | [7 empty] | 4.8
-27.9 | [7 empty] | 4.875
-27.8 | [7 empty] | 4.9222
-27.7 | [7 empty] | 5
-27.6 | [7 empty] | 5.075
-27.5 | [7 empty] | 5.1
-27.4 | [7 empty] | 5.1375
-27.3 | [7 empty] | 5.2556
-27.2 | [7 empty] | 5.35
-27.1 | [7 empty] | 5.4267
-27 | [7 empty] | 5.4933
-26.9 | [7 empty] | 5.6
-26.8 | [7 empty] | 5.6333
-26.7 | [7 empty] | 5.7
-26.6 | [7 empty] | 5.8
-26.5 | [7 empty] | 5.9
-26.4 | [7 empty] | 6
-26.3 | [7 empty] | 6.0571
-26.2 | [7 empty] | 6.2
-26.1 | [7 empty] | 6.225
-26 | [7 empty] | 6.38
-25.9 | [7 empty] | 6.5
-25.8 | [7 empty] | 6.6
-25.7 | [7 empty] | 6.7
-25.6 | [7 empty] | 6.8
-25.5 | [7 empty] | 6.9125
-25.4 | [7 empty] | 7.05
-25.3 | [7 empty] | 7.225
-25.2 | [7 empty] | 7.3571
-25.1 | [7 empty] | 7.4
-25 | [7 empty] | 7.6
-24.9 | [7 empty] | 7.7
-24.8 | [7 empty] | 7.8857
-24.7 | [7 empty] | 8.025
-24.6 | [7 empty] | 8.1571
-24.5 | [7 empty] | 8.2875
-24.4 | [7 empty] | 8.5125
-24.3 | [7 empty] | 8.6429
-24.2 | [7 empty] | 8.85
-24.1 | [7 empty] | 9.025
-24 | [7 empty] | 9.2429
-23.9 | [7 empty] | 9.45
-23.8 | [7 empty] | 9.6
-23.7 | [7 empty] | 9.7667
-23.6 | [7 empty] | 10
-23.5 | [7 empty] | 10.2375
-23.4 | [7 empty] | 10.4429
-23.3 | [7 empty] | 10.7
-23.2 | [7 empty] | 10.9571
-23.1 | [7 empty] | 11.1625
-23 | [7 empty] | 11.4625
-22.9 | [7 empty] | 11.7
-22.8 | [7 empty] | 11.975
-22.7 | [7 empty] | 12.3143
-22.6 | [7 empty] | 12.575
-22.5 | [7 empty] | 12.875
-22.4 | [7 empty] | 13.2714
-22.3 | [7 empty] | 13.575
-22.2 | [7 empty] | 13.98
-22.1 | [7 empty] | 14.4
-22 | [7 empty] | 14.8
-21.9 | [7 empty] | 15.3286
-21.8 | [7 empty] | 15.9
-21.7 | [7 empty] | 16.675
-21.6 | [7 empty] | 17.7714
-21.5 | [7 empty] | 19.3375
-21.4 | [7 empty] | 22.4
-21.3 | [7 empty] | 27.95
-21.2 | [7 empty] | 35.4714
-21.1 | [7 empty] | 45.1
-21 | [7 empty] | 55.85
-20.9 | [7 empty] | 66.525
-20.8 | [7 empty] | 77.2714
-20.7 | [7 empty] | 85.2
-20.6 | [6 empty] | 5.94 | 91.1571
-20.5 | [6 empty] | 6.06 | 94.7143
-20.4 | [6 empty] | 6.15 | 96.55
-20.3 | [6 empty] | 6.2667 | 97.6333
-20.2 | [6 empty] | 6.34 | 98.4
-20.1 | [6 empty] | 6.5 | 99.04
-20 | [6 empty] | 6.6714 | 99.5067
-19.9 | [6 empty] | 6.75 | 99.8182
-19.8 | [6 empty] | 6.8333 | 100.2
-19.7 | [6 empty] | 7 | 100.3571
-19.6 | [6 empty] | 7.1667 | 100.6875
-19.5 | [6 empty] | 7.2333 | 100.9
-19.4 | [6 empty] | 7.4 | 101.1667
-19.3 | [6 empty] | 7.5 | 101.375
-19.2 | [6 empty] | 7.6667 | 101.7143
-19.1 | [6 empty] | 7.8333 | 101.8667
-19 | [6 empty] | 7.9857 | 102.0267
-18.9 | [6 empty] | 8.1667 | 102.16
-18.8 | [6 empty] | 8.35 | 102.2778
-18.7 | [6 empty] | 8.5 | 102.4333
-18.6 | [6 empty] | 8.625 | 102.6286
-18.5 | [6 empty] | 8.82 | 102.8133
-18.4 | [6 empty] | 9.05 | 102.88
-18.3 | [6 empty] | 9.2 | 102.9933
-18.2 | [6 empty] | 9.4 | 103.125
-18.1 | [6 empty] | 9.62 | 103.25
-18 | [6 empty] | 9.8333 | 103.34
-17.9 | [6 empty] | 10.025 | 103.42
-17.8 | [6 empty] | 10.2667 | 103.62
-17.7 | [6 empty] | 10.6 | 103.78
-17.6 | [6 empty] | 10.8 | 103.9067
-17.5 | [6 empty] | 11.12 | 103.9733
-17.4 | [6 empty] | 11.3333 | 104.04
-17.3 | [6 empty] | 11.6667 | 104.1125
-17.2 | [6 empty] | 11.9 | 104.2
-17.1 | [6 empty] | 12.24 | 104.2071
-17 | [6 empty] | 12.6333 | 104.2786
-16.9 | [6 empty] | 12.94 | 104.3778
-16.8 | [6 empty] | 13.3667 | 104.48
-16.7 | [6 empty] | 13.8 | 104.5286
-16.6 | [6 empty] | 14.3 | 104.6833
-16.5 | [6 empty] | 14.92 | 104.7143
-16.4 | [6 empty] | 15.82 | 104.8
-16.3 | [6 empty] | 17 | 104.875
-16.2 | [6 empty] | 18.9 | 104.9
-16.1 | [6 empty] | 22.8714 | 105
-16 | [6 empty] | 29.6 | 105
-15.9 | [6 empty] | 38.1 | 105.1
-15.8 | [6 empty] | 48.8 | 105.1375
-15.7 | [6 empty] | 60.0333 | 105.2
-15.6 | [6 empty] | 70.7 | 105.2
-15.5 | [6 empty] | 80.8 | 105.2375
-15.4 | [5 empty] | 4.3714 | 88.3 | 105.2286
-15.3 | [5 empty] | 4.5 | 92.76 | 105.3
-15.2 | [5 empty] | 4.55 | 95.375 | 105.3
-15.1 | [5 empty] | 4.6375 | 96.86 | 105.3
-15 | [5 empty] | 4.7 | 97.7 | 105.3
-14.9 | [5 empty] | 4.8 | 98.34 | 105.3
-14.8 | [5 empty] | 4.9 | 98.8 | 105.3
-14.7 | [5 empty] | 5 | 99.2 | 105.3
-14.6 | [5 empty] | 5.1 | 99.6 | 105.3
-14.5 | [5 empty] | 5.275 | 99.85 | 105.2375
-14.4 | [5 empty] | 5.3 | 100.1 | 105.2
-14.3 | [5 empty] | 5.4 | 100.2167 | 105.2
-14.2 | [5 empty] | 5.6 | 100.4667 | 105.2
-14.1 | [5 empty] | 5.7 | 100.7 | 105.2
-14 | [5 empty] | 5.775 | 100.8833 | 105.2
-13.9 | [5 empty] | 5.9 | 100.975 | 105.2
-13.8 | [5 empty] | 6.1 | 101.06 | 105.1375
-13.7 | [5 empty] | 6.1625 | 101.3 | 105.1
-13.6 | [5 empty] | 6.375 | 101.4 | 105
-13.5 | [5 empty] | 6.475 | 101.5 | 105
-13.4 | [5 empty] | 6.675 | 101.5714 | 105
-13.3 | [5 empty] | 6.78 | 101.6 | 104.9714
-13.2 | [5 empty] | 6.925 | 101.7667 | 104.9
-13.1 | [5 empty] | 7.125 | 101.8333 | 104.9
-13 | [5 empty] | 7.4 | 101.9 | 104.9
-12.9 | [5 empty] | 7.525 | 102 | 104.9
-12.8 | [5 empty] | 7.7667 | 102.1 | 104.9
-12.7 | [5 empty] | 7.9 | 102.18 | 104.9
-12.6 | [5 empty] | 8.175 | 102.2 | 104.9
-12.5 | [5 empty] | 8.4 | 102.26 | 104.8
-12.4 | [5 empty] | 8.5875 | 102.35 | 104.8
-12.3 | [5 empty] | 8.8286 | 102.5 | 104.8
-12.2 | [5 empty] | 9.0889 | 102.58 | 104.8
-12.1 | [5 empty] | 9.4 | 102.6 | 104.8
-12 | [5 empty] | 9.66 | 102.6 | 104.7857
-11.9 | [5 empty] | 9.9286 | 102.6 | 104.7
-11.8 | [5 empty] | 10.3 | 102.7 | 104.7
-11.7 | [5 empty] | 10.65 | 102.72 | 104.6
-11.6 | [5 empty] | 11 | 102.8 | 104.6
-11.5 | [5 empty] | 11.35 | 102.8 | 104.6
-11.4 | [5 empty] | 11.82 | 102.8667 | 104.6
-11.3 | [5 empty] | 12.3 | 102.9 | 104.6
-11.2 | [5 empty] | 13.075 | 102.9 | 104.6
-11.1 | [5 empty] | 14.08 | 102.9 | 104.6
-11 | [5 empty] | 15.4667 | 102.95 | 104.6
-10.9 | [5 empty] | 18.4 | 103 | 104.5625
-10.8 | [5 empty] | 23.3 | 103 | 104.5
-10.7 | [5 empty] | 31.38 | 103 | 104.4
-10.6 | [5 empty] | 41.025 | 103 | 104.4
-10.5 | [5 empty] | 52.66 | 103 | 104.4
-10.4 | [5 empty] | 63.725 | 103 | 104.4
-10.3 | [5 empty] | 74.2 | 103 | 104.3
-10.2 | [5 empty] | 83.32 | 103 | 104.3
-10.1 |  |  |  |  | 2.4 | 89.28 | 103 | 104.3
-10 |  |  |  |  | 2.4 | 93.1 | 103 | 104.2
-9.9 |  |  |  |  | 2.5 | 94.8 | 103 | 104.1375
-9.8 |  |  |  |  | 2.5 | 95.95 | 103 | 104.1
-9.7 |  |  |  |  | 2.6 | 96.65 | 103 | 104
-9.6 |  |  |  |  | 2.65 | 97.2 | 103 | 104
-9.5 |  |  |  |  | 2.7 | 97.6 | 103 | 104
-9.4 |  |  |  |  | 2.8 | 98.02 | 103 | 103.9
-9.3 |  |  |  |  | 2.825 | 98.4 | 103 | 103.9
-9.2 |  |  |  |  | 2.9 | 98.65 | 103 | 103.8857
-9.1 |  |  |  | 1.8 | 3 | 98.8 | 103 | 103.8
-9 |  |  |  | 1.9 | 3 | 99.1 | 102.9 | 103.8
-8.9 |  |  |  | 1.9 | 3.125 | 99.35 | 102.9 | 103.8
-8.8 |  |  |  | 2 | 3.2 | 99.5 | 102.9 | 103.75
-8.7 |  |  |  | 2 | 3.3 | 99.75 | 102.9 | 103.7
-8.6 |  |  |  | 2 | 3.4 | 99.9 | 102.9 | 103.7
-8.5 |  |  |  | 2.1 | 3.5 | 100.1143 | 102.9 | 103.6
-8.4 |  |  |  | 2.1 | 3.6 | 100.3 | 102.9 | 103.6
-8.3 |  |  |  | 2.2 | 3.7 | 100.45 | 102.9 | 103.6
-8.2 |  |  |  | 2.3 | 3.8 | 100.5 | 102.9 | 103.6
-8.1 |  |  |  | 2.3 | 3.9 | 100.7 | 102.9 | 103.6
-8 |  |  | 1.4 | 2.4 | 4.06 | 100.8 | 102.9 | 103.6
-7.9 |  |  | 1.4 | 2.45 | 4.2 | 100.9 | 102.9 | 103.55
-7.8 |  |  | 1.4 | 2.5 | 4.3 | 100.975 | 102.9 | 103.5
-7.7 |  |  | 1.4333 | 2.6 | 4.4667 | 101 | 102.9 | 103.5
-7.6 |  |  | 1.5 | 2.7 | 4.6333 | 101.1 | 102.9 | 103.5
-7.5 |  |  | 1.5 | 2.8 | 4.7667 | 101.2 | 102.8 | 103.46
-7.4 |  |  | 1.6 | 2.85 | 5 | 101.2 | 102.8 | 103.4
-7.3 |  |  | 1.6 | 2.9 | 5.175 | 101.3 | 102.8 | 103.4
-7.2 |  |  | 1.7 | 3 | 5.375 | 101.4 | 102.8 | 103.3571
-7.1 |  |  | 1.7 | 3.1 | 5.525 | 101.4 | 102.8 | 103.3
-7 |  | 0.9 | 1.7 | 3.25 | 5.775 | 101.4 | 102.8 | 103.3
-6.9 |  | 0.9 | 1.8 | 3.3333 | 6 | 101.5 | 102.8 | 103.2556
-6.8 |  | 0.9 | 1.9 | 3.5 | 6.2667 | 101.6 | 102.7 | 103.1667
-6.7 |  | 0.9 | 1.9 | 3.6 | 6.5 | 101.6 | 102.7 | 103.1
-6.6 |  | 0.96667 | 2 | 3.7 | 6.775 | 101.6 | 102.7 | 103.1
-6.5 | 0.6 | 1 | 2 | 3.9 | 7.0333 | 101.7 | 102.7 | 103.1
-6.4 | 0.6 | 1 | 2.1 | 4 | 7.375 | 101.7 | 102.6 | 103.1
-6.3 | 0.7 | 1 | 2.2 | 4.2 | 7.75 | 101.7 | 102.6 | 103.0091
-6.2 | 0.7 | 1.0667 | 2.25 | 4.4 | 8.1 | 101.7 | 102.6 | 103
-6.1 | 0.7 | 1.1 | 2.3 | 4.6 | 8.55 | 101.8 | 102.6 | 102.9429
-6 | 0.7 | 1.1 | 2.4 | 4.75 | 9.1 | 101.8 | 102.5 | 102.9
-5.9 | 0.7 | 1.1 | 2.5 | 4.9667 | 9.95 | 101.8 | 102.5 | 102.8
-5.8 | 0.7 | 1.2 | 2.6 | 5.2 | 11.1333 | 101.8 | 102.5 | 102.8
-5.7 | 0.8 | 1.2 | 2.7 | 5.4 | 13 | 101.8 | 102.4 | 102.7357
-5.6 | 0.8 | 1.2 | 2.8 | 5.65 | 16.4 | 101.8 | 102.4 | 102.7
-5.5 | 0.8 | 1.3 | 2.9 | 5.95 | 22.8 | 101.8 | 102.4 | 102.6
-5.4 | 0.8 | 1.3 | 3.05 | 6.2 | 32.25 | 101.8 | 102.35 | 102.6
-5.3 | 0.8 | 1.4 | 3.2 | 6.5 | 43.2 | 101.8 | 102.3 | 102.6
-5.2 | 0.9 | 1.4 | 3.3 | 6.9 | 55 | 101.8 | 102.2167 | 102.6
-5.1 | 0.9 | 1.5 | 3.5 | 7.25 | 66.3 | 101.7 | 102.2 | 102.5429
-5 | 0.9 | 1.5 | 3.7 | 7.8 | 77 | 101.7 | 102.1875 | 102.4733
-4.9 | 1 | 1.6 | 3.8 | 8.4 | 85.25 | 101.7 | 102.1 | 102.4067
-4.8 | 1 | 1.7 | 4 | 9.2 | 90.9 | 101.7 | 102.1 | 102.34
-4.7 | 1 | 1.8 | 4.2 | 10.5 | 93.8 | 101.7 | 102.025 | 102.3
-4.6 | 1.1 | 1.8 | 4.4 | 12.85 | 95.3 | 101.7 | 101.9385 | 102.27
-4.5 | 1.1 | 1.9 | 4.6 | 17.05 | 96.05 | 101.6 | 101.9 | 102.2
-4.4 | 1.2 | 2 | 4.8 | 24.85 | 96.78 | 101.6 | 101.8333 | 102.2
-4.3 | 1.2 | 2.1 | 5.1 | 35 | 97.3 | 101.6 | 101.8 | 102.1
-4.2 | 1.2333 | 2.2 | 5.4333 | 46.5 | 97.7 | 101.5333 | 101.75 | 102
-4.1 | 1.3 | 2.3 | 5.7667 | 58.75 | 98 | 101.5 | 101.7 | 102
-4 | 1.4 | 2.4 | 6.2 | 69.85 | 98.25 | 101.475 | 101.6143 | 101.8923
-3.9 | 1.4 | 2.5 | 6.7 | 80.7 | 98.5 | 101.4 | 101.6 | 101.8154
-3.8 | 1.5 | 2.7 | 7.35 | 88.1 | 98.8 | 101.3 | 101.55 | 101.8
-3.7 | 1.6 | 2.8 | 8.4 | 92.35 | 98.95 | 101.25 | 101.5 | 101.78
-3.6 | 1.7 | 3 | 10.05 | 94.65 | 99.18 | 101.2 | 101.5 | 101.7133
-3.5 | 1.8 | 3.2 | 13.28 | 95.9 | 99.42 | 101.2 | 101.5 | 101.6467
-3.4 | 1.9 | 3.4 | 18.6333 | 96.6 | 99.5667 | 101.2 | 101.4 | 101.6
-3.3 | 1.9 | 3.6 | 28.15 | 97.15 | 99.7 | 101.2 | 101.4 | 101.5636
-3.2 | 2.1 | 3.8 | 38.8667 | 97.6 | 99.775 | 101.2 | 101.3 | 101.4625
-3.1 | 2.2 | 4.0333 | 51.55 | 98 | 99.825 | 101.15 | 101.3 | 101.4
-3 | 2.35 | 4.4 | 64.1 | 98.18 | 99.96 | 101.1 | 101.2167 | 101.3167
-2.9 | 2.5 | 4.75 | 74.7 | 98.3667 | 100.0375 | 101.1 | 101.2 | 101.3
-2.8 | 2.6 | 5.3 | 84.3 | 98.7 | 100.1 | 101.1 | 101.175 | 101.2
-2.7 | 2.85 | 6.1 | 90.8 | 98.9 | 100.1375 | 101.1 | 101.1 | 101.2
-2.6 | 3.05 | 7.3 | 94.2 | 99 | 100.2 | 101 | 101 | 101.2
-2.5 | 3.3 | 9.6 | 95.8 | 99.2 | 100.2 | 100.975 | 101 | 101.11
-2.4 | 3.6 | 13.825 | 96.9 | 99.3 | 100.25 | 100.9 | 100.9 | 101.0857
-2.3 | 4.1 | 21.3667 | 97.5 | 99.4 | 100.3 | 100.875 | 100.9 | 101
-2.2 | 4.7667 | 32.3 | 97.9 | 99.5 | 100.3 | 100.8 | 100.8429 | 100.9
-2.1 | 5.8 | 44.6 | 98.4 | 99.6 | 100.4 | 100.8 | 100.8 | 100.9
-2 | 7.6 | 57.8 | 98.6 | 99.7 | 100.4 | 100.7 | 100.7167 | 100.825
-1.9 | 11.5 | 69.7 | 98.8 | 99.8 | 100.4 | 100.7 | 100.7 | 100.8
-1.8 | 18.2 | 80.1333 | 99 | 99.9 | 100.4 | 100.7 | 100.7 | 100.7
-1.7 | 28.8333 | 88.8 | 99.2 | 100 | 100.4 | 100.6 | 100.6571 | 100.7
-1.6 | 40.4333 | 93.9333 | 99.3 | 100 | 100.4 | 100.525 | 100.6 | 100.62
-1.5 | 53.7 | 96.375 | 99.5 | 100 | 100.4 | 100.5 | 100.6 | 100.6
-1.4 | 66.95 | 97.7 | 99.6 | 100.1 | 100.4 | 100.4 | 100.5 | 100.5333
-1.3 | 78.1 | 98.3 | 99.7 | 100.1 | 100.3 | 100.4 | 100.4 | 100.5
-1.2 | 87.6 | 98.8667 | 99.7 | 100.1 | 100.3 | 100.4 | 100.375 | 100.4625
-1.1 | 93.4 | 99.1 | 99.8 | 100.1 | 100.3 | 100.3 | 100.3 | 100.4
-1 | 96.5 | 99.3333 | 99.9 | 100.1 | 100.2 | 100.3 | 100.2 | 100.3
-0.9 | 98 | 99.6 | 99.95 | 100.1 | 100.2 | 100.2 | 100.2 | 100.3
-0.8 | 98.8 | 99.7 | 100 | 100 | 100.2 | 100.2 | 100.1 | 100.275
-0.7 | 99.3 | 99.8 | 100 | 100 | 100.2 | 100.2 | 100.1 | 100.2
-0.6 | 99.5 | 99.9 | 100 | 100 | 100.12 | 100.2 | 100.1 | 100.2
-0.5 | 99.7 | 100 | 100 | 100 | 100.1 | 100.2 | 100.1 | 100.1
-0.4 | 99.8 | 100 | 100.1 | 100 | 100.1 | 100.1 | 100 | 100.1
-0.3 | 99.9 | 100 | 100.1 | 100 | 100.025 | 100.1 | 100 | 100.0125
-0.2 | 99.9333 | 100 | 100.0333 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100.05 | 100 | 100 | 100 | 100 | 100
0.3 | 99.9 | 100 | 100.1 | 100 | 100.05 | 100.1 | 100 | 100
0.4 | 99.8 | 100 | 100.1 | 100 | 100.1 | 100.1 | 100 | 100.1
0.5 | 99.7 | 100 | 100 | 100 | 100.1 | 100.2 | 100.1 | 100.1429
0.6 | 99.5 | 99.9333 | 100 | 100 | 100.1 | 100.2 | 100.1 | 100.2
0.7 | 99.3 | 99.8 | 100 | 100 | 100.2 | 100.2 | 100.1 | 100.2
0.8 | 98.8 | 99.7 | 100 | 100 | 100.2 | 100.2 | 100.1 | 100.25
0.9 | 98 | 99.6 | 99.95 | 100.1 | 100.2 | 100.2 | 100.2 | 100.3
1 | 96.5 | 99.3333 | 99.9 | 100.1 | 100.2 | 100.3 | 100.2 | 100.3
1.1 | 93.4 | 99.1 | 99.8 | 100.1 | 100.3 | 100.3 | 100.3 | 100.4
1.2 | 87.7 | 98.85 | 99.7 | 100.1 | 100.3 | 100.4 | 100.4 | 100.4286
1.3 | 78.1 | 98.3 | 99.7 | 100.1 | 100.3 | 100.4 | 100.4 | 100.5
1.4 | 66.95 | 97.7 | 99.6 | 100.1 | 100.4 | 100.4 | 100.5 | 100.5
1.5 | 53.65 | 96.4 | 99.5 | 100 | 100.4 | 100.5 | 100.6 | 100.6
1.6 | 40.25 | 93.85 | 99.3 | 100 | 100.4 | 100.525 | 100.6 | 100.6
1.7 | 28.9 | 88.8 | 99.2 | 100 | 100.4 | 100.6 | 100.6 | 100.7
1.8 | 18.2333 | 80.15 | 99 | 99.9 | 100.4 | 100.7 | 100.7 | 100.7
1.9 | 11.22 | 69.7 | 98.8 | 99.8 | 100.4 | 100.7 | 100.7 | 100.7571
2 | 7.6 | 57.8 | 98.6 | 99.7667 | 100.4 | 100.7 | 100.7 | 100.8
2.1 | 5.8 | 44.6 | 98.3667 | 99.6 | 100.4 | 100.8 | 100.8 | 100.9
2.2 | 4.8 | 32.3 | 97.9 | 99.5 | 100.3 | 100.8 | 100.8 | 100.9
2.3 | 4.1 | 21.3667 | 97.5 | 99.4 | 100.3 | 100.875 | 100.9 | 101
2.4 | 3.65 | 13.8 | 96.9 | 99.3 | 100.25 | 100.9 | 100.9 | 101.1
2.5 | 3.3 | 9.5 | 95.8 | 99.2 | 100.2 | 101 | 101 | 101.1
2.6 | 3.05 | 7.4 | 94.2 | 99.05 | 100.2 | 101 | 101 | 101.1889
2.7 | 2.8333 | 6.1 | 90.8 | 98.9 | 100.1375 | 101.1 | 101.0833 | 101.2
2.8 | 2.7 | 5.3 | 84.5 | 98.7 | 100.1 | 101.1 | 101.12 | 101.2571
2.9 | 2.5 | 4.75 | 74.7 | 98.35 | 100.0375 | 101.1 | 101.2 | 101.3
3 | 2.35 | 4.4 | 64.1 | 98.175 | 99.96 | 101.1 | 101.2429 | 101.3
3.1 | 2.2 | 4.0333 | 51.3 | 98 | 99.825 | 101.1667 | 101.3 | 101.4
3.2 | 2.1 | 3.8 | 38.8667 | 97.5667 | 99.775 | 101.2 | 101.3 | 101.43
3.3 | 1.9667 | 3.6 | 27.8 | 97.18 | 99.7 | 101.2 | 101.4 | 101.5333
3.4 | 1.9 | 3.4 | 18.6333 | 96.6 | 99.55 | 101.2 | 101.4 | 101.6
3.5 | 1.7667 | 3.2 | 13.1 | 95.9 | 99.42 | 101.2 | 101.45 | 101.6133
3.6 | 1.7 | 3 | 10.05 | 94.6 | 99.175 | 101.2 | 101.5 | 101.68
3.7 | 1.6 | 2.8 | 8.4 | 92.45 | 98.94 | 101.25 | 101.5 | 101.7467
3.8 | 1.5 | 2.7 | 7.35 | 88.1 | 98.8 | 101.3 | 101.5 | 101.8133
3.9 | 1.4 | 2.5 | 6.65 | 80.7 | 98.55 | 101.4 | 101.6 | 101.88
4 | 1.4 | 2.4 | 6.2 | 70.05 | 98.25 | 101.475 | 101.65 | 101.97
4.1 | 1.3 | 2.3 | 5.7667 | 58.8 | 98 | 101.5 | 101.7 | 102
4.2 | 1.3 | 2.2 | 5.4 | 46.6 | 97.7 | 101.5333 | 101.7833 | 102.0375
4.3 | 1.2 | 2.1 | 5.1 | 35.15 | 97.3 | 101.6 | 101.8 | 102.1
4.4 | 1.2 | 2 | 4.8 | 25.1 | 96.78 | 101.6 | 101.8667 | 102.16
4.5 | 1.1 | 1.9 | 4.6 | 17.15 | 96.1 | 101.6 | 101.9 | 102.2
4.6 | 1.1 | 1.8 | 4.425 | 12.9 | 95.3 | 101.7 | 101.975 | 102.2333
4.7 | 1 | 1.8 | 4.2 | 10.55 | 93.8 | 101.7 | 102.07 | 102.3308
4.8 | 1 | 1.7 | 4 | 9.2 | 90.9 | 101.7 | 102.1 | 102.4067
4.9 | 1 | 1.6 | 3.8 | 8.4 | 85.25 | 101.7 | 102.1 | 102.4733
5 | 0.9 | 1.5 | 3.7 | 7.75 | 76.25 | 101.7 | 102.1286 | 102.5
5.1 | 0.9 | 1.5 | 3.5 | 7.3 | 66.4333 | 101.7 | 102.2 | 102.5077
5.2 | 0.9 | 1.4 | 3.3 | 6.9 | 54.525 | 101.8 | 102.2 | 102.5846
5.3 | 0.8 | 1.4 | 3.2 | 6.5 | 42.65 | 101.8 | 102.3 | 102.6
5.4 | 0.8 | 1.3 | 3.0333 | 6.2 | 31.825 | 101.8 | 102.3 | 102.6
5.5 | 0.8 | 1.3 | 2.9 | 5.95 | 23 | 101.8 | 102.4 | 102.6714
5.6 | 0.8 | 1.2 | 2.8 | 5.6 | 16.35 | 101.8 | 102.4 | 102.7091
5.7 | 0.8 | 1.2 | 2.7 | 5.4 | 13 | 101.8 | 102.45 | 102.8
5.8 | 0.7 | 1.2 | 2.6 | 5.2 | 11.1 | 101.8 | 102.5 | 102.8
5.9 | 0.7 | 1.1 | 2.5 | 4.9667 | 10 | 101.8 | 102.5 | 102.8
6 | 0.7 | 1.1 | 2.4 | 4.7 | 9.1 | 101.8 | 102.5 | 102.8714
6.1 | 0.7 | 1.1 | 2.3 | 4.6 | 8.55 | 101.8 | 102.55 | 102.9125
6.2 | 0.7 | 1.0667 | 2.2 | 4.4 | 8.1 | 101.7 | 102.6 | 103
6.3 | 0.7 | 1 | 2.2 | 4.2 | 7.75 | 101.7 | 102.6 | 103
6.4 | 0.6 | 1 | 2.1 | 4 | 7.3667 | 101.7 | 102.6 | 103.0615
6.5 | 0.6 | 1 | 2 | 3.9 | 7.0333 | 101.7 | 102.66 | 103.1
6.6 |  | 0.96667 | 2 | 3.7 | 6.75 | 101.6 | 102.7 | 103.1
6.7 |  | 0.9 | 1.9 | 3.6 | 6.5 | 101.6 | 102.7 | 103.1667
6.8 |  | 0.9 | 1.9 | 3.5 | 6.25 | 101.6 | 102.7 | 103.2385
6.9 |  | 0.9 | 1.8 | 3.3 | 6 | 101.5 | 102.8 | 103.3
7 |  | 0.9 | 1.7 | 3.25 | 5.775 | 101.4 | 102.8 | 103.3
7.1 |  |  | 1.7 | 3.1 | 5.55 | 101.4 | 102.8 | 103.3
7.2 |  |  | 1.7 | 3 | 5.375 | 101.4 | 102.8 | 103.325
7.3 |  |  | 1.6 | 2.9 | 5.175 | 101.3 | 102.8 | 103.4
7.4 |  |  | 1.6 | 2.8333 | 5 | 101.2 | 102.8 | 103.4
7.5 |  |  | 1.5 | 2.8 | 4.7667 | 101.2 | 102.8 | 103.4308
7.6 |  |  | 1.5 | 2.7 | 4.625 | 101.1 | 102.9 | 103.5
7.7 |  |  | 1.4 | 2.6 | 4.4667 | 101 | 102.9 | 103.5
7.8 |  |  | 1.4 | 2.5 | 4.3 | 100.975 | 102.9 | 103.5
7.9 |  |  | 1.4 | 2.45 | 4.2 | 100.9 | 102.9 | 103.5143
8 |  |  | 1.4 | 2.4 | 4.05 | 100.8 | 102.9 | 103.6
8.1 |  |  | 1.3 | 2.3 | 3.9 | 100.7 | 102.9 | 103.6
8.2 |  |  |  | 2.3 | 3.8 | 100.5 | 102.9 | 103.6
8.3 |  |  |  | 2.2 | 3.7 | 100.4333 | 102.9 | 103.6
8.4 |  |  |  | 2.1 | 3.6 | 100.3 | 102.9 | 103.6
8.5 |  |  |  | 2.1 | 3.5 | 100.1143 | 102.9 | 103.6
8.6 |  |  |  | 2 | 3.4 | 99.9 | 102.9 | 103.7
8.7 |  |  |  | 2 | 3.3 | 99.74 | 102.9 | 103.7
8.8 |  |  |  | 2 | 3.2 | 99.5 | 102.9 | 103.8
8.9 |  |  |  | 1.9 | 3.125 | 99.35 | 102.9 | 103.8
9 |  |  |  | 1.9 | 3 | 99.1 | 102.9 | 103.8
9.1 |  |  |  | 1.8 | 3 | 98.8 | 102.9286 | 103.8
9.2 |  |  |  |  | 2.9 | 98.65 | 103 | 103.85
9.3 |  |  |  |  | 2.825 | 98.4 | 103 | 103.9
9.4 |  |  |  |  | 2.8 | 98 | 103 | 103.9
9.5 |  |  |  |  | 2.7 | 97.6 | 103 | 103.9714
9.6 |  |  |  |  | 2.6333 | 97.2 | 103 | 104
9.7 |  |  |  |  | 2.6 | 96.7 | 103 | 104.0375
9.8 |  |  |  |  | 2.5 | 95.95 | 103 | 104.1
9.9 |  |  |  |  | 2.5 | 94.8 | 103 | 104.1
10 |  |  |  |  | 2.4 | 92.9 | 103 | 104.2
10.1 |  |  |  |  | 2.4 | 89.2 | 103 | 104.2714
10.2 | [5 empty] | 82.9 | 103 | 104.3
10.3 | [5 empty] | 74.5 | 103 | 104.3
10.4 | [5 empty] | 63.95 | 103 | 104.3625
10.5 | [5 empty] | 52.9 | 103 | 104.4
10.6 | [5 empty] | 41.575 | 103 | 104.4
10.7 | [5 empty] | 31 | 103 | 104.4625
10.8 | [5 empty] | 23.3667 | 103 | 104.5
10.9 | [5 empty] | 18.4 | 103 | 104.6
11 | [5 empty] | 15.5 | 103 | 104.6
11.1 | [5 empty] | 14.08 | 102.9 | 104.6
11.2 | [5 empty] | 13.075 | 102.9 | 104.6
11.3 | [5 empty] | 12.3 | 102.9 | 104.6
11.4 | [5 empty] | 11.82 | 102.8333 | 104.6
11.5 | [5 empty] | 11.35 | 102.8 | 104.6
11.6 | [5 empty] | 11 | 102.8 | 104.6
11.7 | [5 empty] | 10.65 | 102.7 | 104.65
11.8 | [5 empty] | 10.3 | 102.7 | 104.7
11.9 | [5 empty] | 10.0286 | 102.6 | 104.7
12 | [5 empty] | 9.65 | 102.6 | 104.75
12.1 | [5 empty] | 9.4 | 102.6 | 104.8
12.2 | [5 empty] | 9.075 | 102.5333 | 104.8
12.3 | [5 empty] | 8.825 | 102.5 | 104.8
12.4 | [5 empty] | 8.5857 | 102.4167 | 104.8
12.5 | [5 empty] | 8.4 | 102.22 | 104.8
12.6 | [5 empty] | 8.15 | 102.2 | 104.9
12.7 | [5 empty] | 7.9 | 102.1333 | 104.9
12.8 | [5 empty] | 7.7667 | 102.1 | 104.9
12.9 | [5 empty] | 7.55 | 102 | 104.9
13 | [5 empty] | 7.4 | 101.9 | 104.9
13.1 | [5 empty] | 7.125 | 101.8 | 104.9
13.2 | [5 empty] | 6.925 | 101.7333 | 104.9
13.3 | [5 empty] | 6.8 | 101.6 | 105
13.4 | [5 empty] | 6.675 | 101.5333 | 105
13.5 | [5 empty] | 6.475 | 101.5 | 105
13.6 | [5 empty] | 6.36 | 101.375 | 105
13.7 | [5 empty] | 6.1625 | 101.25 | 105.0714
13.8 | [5 empty] | 6.1 | 101.1 | 105.2
13.9 | [5 empty] | 5.9 | 100.92 | 105.2
14 | [5 empty] | 5.775 | 100.8429 | 105.2
14.1 | [5 empty] | 5.7 | 100.66 | 105.2
14.2 | [5 empty] | 5.6 | 100.46 | 105.2
14.3 | [5 empty] | 5.4 | 100.2833 | 105.2
14.4 | [5 empty] | 5.3 | 100.1167 | 105.2
14.5 | [5 empty] | 5.275 | 99.8 | 105.3
14.6 | [5 empty] | 5.1 | 99.56 | 105.3
14.7 | [5 empty] | 5 | 99.22 | 105.3
14.8 | [5 empty] | 4.9 | 98.7714 | 105.3
14.9 | [5 empty] | 4.8 | 98.34 | 105.3
15 | [5 empty] | 4.7 | 97.6333 | 105.3
15.1 | [5 empty] | 4.6375 | 96.9 | 105.3
15.2 | [5 empty] | 4.55 | 95.4 | 105.3
15.3 | [5 empty] | 4.5 | 92.85 | 105.2375
15.4 | [5 empty] | 4.3571 | 88.3 | 105.2
15.5 | [6 empty] | 80.8333 | 105.275
15.6 | [6 empty] | 71.06 | 105.2
15.7 | [6 empty] | 60 | 105.2
15.8 | [6 empty] | 48.8667 | 105.1
15.9 | [6 empty] | 38.4333 | 105.1
16 | [6 empty] | 29.2 | 105
16.1 | [6 empty] | 23 | 104.975
16.2 | [6 empty] | 19.12 | 104.9
16.3 | [6 empty] | 16.9 | 104.8125
16.4 | [6 empty] | 15.8 | 104.7857
16.5 | [6 empty] | 15 | 104.7
16.6 | [6 empty] | 14.35 | 104.7
16.7 | [6 empty] | 13.75 | 104.5667
16.8 | [6 empty] | 13.35 | 104.425
16.9 | [6 empty] | 12.96 | 104.4
17 | [6 empty] | 12.66 | 104.31
17.1 | [6 empty] | 12.2667 | 104.24
17.2 | [6 empty] | 11.9333 | 104.2
17.3 | [6 empty] | 11.65 | 104.1091
17.4 | [6 empty] | 11.3667 | 104
17.5 | [6 empty] | 11.1286 | 104
17.6 | [6 empty] | 10.82 | 103.88
17.7 | [6 empty] | 10.5667 | 103.7467
17.8 | [6 empty] | 10.3 | 103.6133
17.9 | [6 empty] | 10.1 | 103.48
18 | [6 empty] | 9.8 | 103.375
18.1 | [6 empty] | 9.66 | 103.3125
18.2 | [6 empty] | 9.36 | 103.1933
18.3 | [6 empty] | 9.25 | 103.06
18.4 | [6 empty] | 9.02 | 102.9267
18.5 | [6 empty] | 8.88 | 102.7933
18.6 | [6 empty] | 8.7 | 102.6357
18.7 | [6 empty] | 8.5 | 102.4214
18.8 | [6 empty] | 8.32 | 102.31
18.9 | [6 empty] | 8.1333 | 102.2182
19 | [6 empty] | 7.9667 | 102.04
19.1 | [6 empty] | 7.7833 | 101.8667
19.2 | [6 empty] | 7.7 | 101.6857
19.3 | [6 empty] | 7.525 | 101.4
19.4 | [6 empty] | 7.35 | 101.1714
19.5 | [6 empty] | 7.2857 | 100.9571
19.6 | [6 empty] | 7.1333 | 100.7
19.7 | [6 empty] | 6.9667 | 100.3625
19.8 | [6 empty] | 6.9 | 100.1571
19.9 | [6 empty] | 6.7 | 99.9
20 | [6 empty] | 6.6333 | 99.5154
20.1 | [6 empty] | 6.4667 | 99.0267
20.2 | [6 empty] | 6.4 | 98.4462
20.3 | [6 empty] | 6.2333 | 97.6769
20.4 | [6 empty] | 6.2 | 96.6429
20.5 | [6 empty] | 6.1 | 94.7
20.6 | [6 empty] | 6 | 91.1286
20.7 | [7 empty] | 85.4
20.8 | [7 empty] | 76.8857
20.9 | [7 empty] | 66.8375
21 | [7 empty] | 55.8857
21.1 | [7 empty] | 45.2
21.2 | [7 empty] | 35.2857
21.3 | [7 empty] | 27.9571
21.4 | [7 empty] | 22.5
21.5 | [7 empty] | 19.55
21.6 | [7 empty] | 17.7
21.7 | [7 empty] | 16.7
21.8 | [7 empty] | 15.925
21.9 | [7 empty] | 15.4
22 | [7 empty] | 14.85
22.1 | [7 empty] | 14.3857
22.2 | [7 empty] | 13.96
22.3 | [7 empty] | 13.575
22.4 | [7 empty] | 13.2625
22.5 | [7 empty] | 12.9143
22.6 | [7 empty] | 12.575
22.7 | [7 empty] | 12.275
22.8 | [7 empty] | 12.0143
22.9 | [7 empty] | 11.72
23 | [7 empty] | 11.45
23.1 | [7 empty] | 11.175
23.2 | [7 empty] | 10.925
23.3 | [7 empty] | 10.675
23.4 | [7 empty] | 10.5
23.5 | [7 empty] | 10.2429
23.6 | [7 empty] | 10.0375
23.7 | [7 empty] | 9.8
23.8 | [7 empty] | 9.6429
23.9 | [7 empty] | 9.425
24 | [7 empty] | 9.25
24.1 | [7 empty] | 9.0429
24.2 | [7 empty] | 8.825
24.3 | [7 empty] | 8.675
24.4 | [7 empty] | 8.5
24.5 | [7 empty] | 8.325
24.6 | [7 empty] | 8.1875
24.7 | [7 empty] | 8.0143
24.8 | [7 empty] | 7.9
24.9 | [7 empty] | 7.6875
25 | [7 empty] | 7.5571
25.1 | [7 empty] | 7.425
25.2 | [7 empty] | 7.2875
25.3 | [7 empty] | 7.15
25.4 | [7 empty] | 7.0857
25.5 | [7 empty] | 6.9333
25.6 | [7 empty] | 6.84
25.7 | [7 empty] | 6.7
25.8 | [7 empty] | 6.6
25.9 | [7 empty] | 6.4857
26 | [7 empty] | 6.4
26.1 | [7 empty] | 6.2625
26.2 | [7 empty] | 6.2
26.3 | [7 empty] | 6.1
26.4 | [7 empty] | 5.9625
26.5 | [7 empty] | 5.9
26.6 | [7 empty] | 5.8
26.7 | [7 empty] | 5.7
26.8 | [7 empty] | 5.6625
26.9 | [7 empty] | 5.5286
27 | [7 empty] | 5.4333
27.1 | [7 empty] | 5.36
27.2 | [7 empty] | 5.3
27.3 | [7 empty] | 5.28
27.4 | [7 empty] | 5.175
27.5 | [7 empty] | 5.1
27.6 | [7 empty] | 5.1
27.7 | [7 empty] | 4.975
27.8 | [7 empty] | 4.9
27.9 | [7 empty] | 4.8143
28 | [7 empty] | 4.7875
28.1 | [7 empty] | 4.7
28.2 | [7 empty] | 4.7
28.3 | [7 empty] | 4.6
28.4 | [7 empty] | 4.5571
28.5 | [7 empty] | 4.5
28.6 | [7 empty] | 4.4
28.7 | [7 empty] | 4.4
28.8 | [7 empty] | 4.3167
28.9 | [7 empty] | 4.3

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 6.125
-29.7 | [7 empty] | 6.2
-29.6 | [7 empty] | 6.3
-29.5 | [7 empty] | 6.4
-29.4 | [7 empty] | 6.4714
-29.3 | [7 empty] | 6.52
-29.2 | [7 empty] | 6.6143
-29.1 | [7 empty] | 6.7571
-29 | [7 empty] | 6.8
-28.9 | [7 empty] | 6.9
-28.8 | [7 empty] | 6.9714
-28.7 | [7 empty] | 7
-28.6 | [7 empty] | 7.2
-28.5 | [7 empty] | 7.2375
-28.4 | [7 empty] | 7.3714
-28.3 | [7 empty] | 7.48
-28.2 | [7 empty] | 7.58
-28.1 | [7 empty] | 7.6889
-28 | [7 empty] | 7.7692
-27.9 | [7 empty] | 7.9091
-27.8 | [7 empty] | 8.0667
-27.7 | [7 empty] | 8.2
-27.6 | [7 empty] | 8.2667
-27.5 | [7 empty] | 8.3667
-27.4 | [7 empty] | 8.5
-27.3 | [7 empty] | 8.6538
-27.2 | [7 empty] | 8.7778
-27.1 | [7 empty] | 8.9778
-27 | [7 empty] | 9.0222
-26.9 | [7 empty] | 9.2
-26.8 | [7 empty] | 9.3875
-26.7 | [7 empty] | 9.5143
-26.6 | [7 empty] | 9.65
-26.5 | [7 empty] | 9.85
-26.4 | [7 empty] | 10.0143
-26.3 | [7 empty] | 10.1571
-26.2 | [7 empty] | 10.3167
-26.1 | [7 empty] | 10.5071
-26 | [7 empty] | 10.7182
-25.9 | [7 empty] | 10.9
-25.8 | [7 empty] | 11.0875
-25.7 | [7 empty] | 11.28
-25.6 | [7 empty] | 11.48
-25.5 | [7 empty] | 11.7125
-25.4 | [7 empty] | 11.8667
-25.3 | [7 empty] | 12.08
-25.2 | [7 empty] | 12.28
-25.1 | [7 empty] | 12.54
-25 | [7 empty] | 12.7923
-24.9 | [7 empty] | 13.0375
-24.8 | [7 empty] | 13.3429
-24.7 | [7 empty] | 13.5636
-24.6 | [7 empty] | 13.7909
-24.5 | [7 empty] | 14
-24.4 | [7 empty] | 14.325
-24.3 | [7 empty] | 14.7
-24.2 | [7 empty] | 14.9286
-24.1 | [7 empty] | 15.2
-24 | [7 empty] | 15.5
-23.9 | [7 empty] | 15.8333
-23.8 | [7 empty] | 16.1667
-23.7 | [7 empty] | 16.5375
-23.6 | [7 empty] | 16.9
-23.5 | [7 empty] | 17.2429
-23.4 | [7 empty] | 17.65
-23.3 | [7 empty] | 18.0571
-23.2 | [7 empty] | 18.4429
-23.1 | [7 empty] | 18.8714
-23 | [7 empty] | 19.3733
-22.9 | [7 empty] | 19.92
-22.8 | [7 empty] | 20.5467
-22.7 | [7 empty] | 21.28
-22.6 | [7 empty] | 22.3077
-22.5 | [7 empty] | 23.9143
-22.4 | [7 empty] | 26.4375
-22.3 | [7 empty] | 31.05
-22.2 | [7 empty] | 37.3364
-22.1 | [7 empty] | 45.1571
-22 | [7 empty] | 54.025
-21.9 | [7 empty] | 63
-21.8 | [7 empty] | 71.3333
-21.7 | [7 empty] | 77.8533
-21.6 | [7 empty] | 83.24
-21.5 | [7 empty] | 86.44
-21.4 | [6 empty] | 8.35 | 88.4818
-21.3 | [6 empty] | 8.475 | 89.7
-21.2 | [6 empty] | 8.65 | 90.5
-21.1 | [6 empty] | 8.8364 | 91.0286
-21 | [6 empty] | 8.9429 | 91.55
-20.9 | [6 empty] | 9.1167 | 91.9429
-20.8 | [6 empty] | 9.2833 | 92.4
-20.7 | [6 empty] | 9.3833 | 92.8571
-20.6 | [6 empty] | 9.5429 | 93.225
-20.5 | [6 empty] | 9.7 | 93.5125
-20.4 | [6 empty] | 9.9444 | 93.8786
-20.3 | [6 empty] | 10.1429 | 94.2357
-20.2 | [6 empty] | 10.2889 | 94.5929
-20.1 | [6 empty] | 10.48 | 94.8556
-20 | [6 empty] | 10.7143 | 95.14
-19.9 | [6 empty] | 10.8667 | 95.44
-19.8 | [6 empty] | 11.0667 | 95.66
-19.7 | [6 empty] | 11.28 | 95.8333
-19.6 | [6 empty] | 11.5333 | 96.1
-19.5 | [6 empty] | 11.7 | 96.4182
-19.4 | [6 empty] | 11.9667 | 96.6
-19.3 | [6 empty] | 12.2167 | 96.8375
-19.2 | [6 empty] | 12.425 | 97.0429
-19.1 | [6 empty] | 12.7 | 97.3
-19 | [6 empty] | 12.98 | 97.5375
-18.9 | [6 empty] | 13.2 | 97.6714
-18.8 | [6 empty] | 13.5 | 97.8
-18.7 | [6 empty] | 13.7667 | 98.0286
-18.6 | [6 empty] | 14.1 | 98.2429
-18.5 | [6 empty] | 14.4143 | 98.4
-18.4 | [6 empty] | 14.7 | 98.5429
-18.3 | [6 empty] | 15 | 98.75
-18.2 | [6 empty] | 15.3667 | 98.9125
-18.1 | [6 empty] | 15.68 | 99.05
-18 | [6 empty] | 16 | 99.2556
-17.9 | [6 empty] | 16.3857 | 99.3222
-17.8 | [6 empty] | 16.7875 | 99.475
-17.7 | [6 empty] | 17.2 | 99.6556
-17.6 | [6 empty] | 17.65 | 99.775
-17.5 | [6 empty] | 18.05 | 99.9778
-17.4 | [6 empty] | 18.65 | 100.1
-17.3 | [6 empty] | 19.25 | 100.2571
-17.2 | [6 empty] | 20.0167 | 100.4111
-17.1 | [6 empty] | 21 | 100.5286
-17 | [6 empty] | 22.42 | 100.6
-16.9 | [6 empty] | 25.05 | 100.7
-16.8 | [6 empty] | 29.4667 | 100.8286
-16.7 | [6 empty] | 35.96 | 100.9625
-16.6 | [6 empty] | 44.45 | 101.1
-16.5 | [6 empty] | 53.55 | 101.1286
-16.4 | [6 empty] | 62.7286 | 101.2556
-16.3 | [6 empty] | 71.65 | 101.3
-16.2 | [6 empty] | 79.3167 | 101.4
-16.1 | [6 empty] | 84.65 | 101.4875
-16 | [6 empty] | 87.7833 | 101.6
-15.9 | [5 empty] | 6.5 | 89.5333 | 101.6667
-15.8 | [5 empty] | 6.6625 | 90.5667 | 101.7143
-15.7 | [5 empty] | 6.775 | 91.3 | 101.8
-15.6 | [5 empty] | 6.9 | 92.0571 | 101.8
-15.5 | [5 empty] | 7 | 92.52 | 101.9
-15.4 | [5 empty] | 7.2 | 92.9667 | 101.9
-15.3 | [5 empty] | 7.2857 | 93.42 | 101.9
-15.2 | [5 empty] | 7.45 | 93.7 | 102
-15.1 | [5 empty] | 7.6 | 94.12 | 102
-15 | [5 empty] | 7.775 | 94.375 | 102
-14.9 | [5 empty] | 7.9 | 94.7143 | 102.1
-14.8 | [5 empty] | 8.075 | 95.0167 | 102.1
-14.7 | [5 empty] | 8.225 | 95.1833 | 102.1
-14.6 | [5 empty] | 8.375 | 95.5 | 102.1
-14.5 | [5 empty] | 8.625 | 95.8 | 102.1
-14.4 | [5 empty] | 8.7667 | 96.05 | 102.1429
-14.3 | [5 empty] | 8.925 | 96.2 | 102.2
-14.2 | [5 empty] | 9.1 | 96.4444 | 102.2
-14.1 | [5 empty] | 9.325 | 96.66 | 102.2
-14 | [5 empty] | 9.5375 | 96.83 | 102.2
-13.9 | [5 empty] | 9.725 | 97 | 102.3
-13.8 | [5 empty] | 9.975 | 97.2667 | 102.3
-13.7 | [5 empty] | 10.2222 | 97.4333 | 102.3
-13.6 | [5 empty] | 10.4286 | 97.575 | 102.3
-13.5 | [5 empty] | 10.6667 | 97.7333 | 102.4
-13.4 | [5 empty] | 10.9 | 97.8857 | 102.4
-13.3 | [5 empty] | 11.225 | 98 | 102.4
-13.2 | [5 empty] | 11.4333 | 98.225 | 102.4
-13.1 | [5 empty] | 11.775 | 98.36 | 102.4
-13 | [5 empty] | 12.0333 | 98.5 | 102.4
-12.9 | [5 empty] | 12.3333 | 98.6429 | 102.5
-12.8 | [5 empty] | 12.6667 | 98.8 | 102.5
-12.7 | [5 empty] | 12.98 | 98.9 | 102.5
-12.6 | [5 empty] | 13.28 | 99 | 102.6
-12.5 | [5 empty] | 13.65 | 99.16 | 102.6
-12.4 | [5 empty] | 13.98 | 99.2667 | 102.6
-12.3 | [5 empty] | 14.4 | 99.4143 | 102.6
-12.2 | [5 empty] | 14.8 | 99.5 | 102.625
-12.1 | [5 empty] | 15.2286 | 99.6 | 102.7
-12 | [5 empty] | 15.75 | 99.7833 | 102.7
-11.9 | [5 empty] | 16.16 | 99.8444 | 102.7
-11.8 | [5 empty] | 16.7667 | 99.9625 | 102.7
-11.7 | [5 empty] | 17.5 | 100.0167 | 102.7
-11.6 | [5 empty] | 18.6 | 100.1 | 102.7
-11.5 | [5 empty] | 20.05 | 100.15 | 102.7
-11.4 | [5 empty] | 22.625 | 100.2167 | 102.7
-11.3 | [5 empty] | 26.85 | 100.3 | 102.8
-11.2 | [5 empty] | 34.0667 | 100.4 | 102.8
-11.1 | [5 empty] | 42.6667 | 100.4 | 102.8
-11 | [5 empty] | 51.9333 | 100.5 | 102.8
-10.9 | [5 empty] | 61.8667 | 100.5833 | 102.78
-10.8 | [5 empty] | 71.3 | 100.6 | 102.7
-10.7 | [5 empty] | 78.6 | 100.7 | 102.7
-10.6 | [5 empty] | 84.3 | 100.7 | 102.6
-10.5 | [5 empty] | 87.6857 | 100.8 | 102.6
-10.4 |  |  |  |  | 3.7 | 89.62 | 100.8 | 102.6
-10.3 |  |  |  |  | 3.8 | 90.7 | 100.8 | 102.6
-10.2 |  |  |  |  | 3.9 | 91.6 | 100.8462 | 102.6
-10.1 |  |  |  |  | 4 | 92.1 | 100.9 | 102.5
-10 |  |  |  |  | 4.1 | 92.65 | 100.9 | 102.5
-9.9 |  |  |  |  | 4.2 | 93.2 | 101 | 102.5
-9.8 |  |  |  |  | 4.3 | 93.58 | 101 | 102.5
-9.7 |  |  |  |  | 4.4 | 93.95 | 101.1 | 102.5
-9.6 |  |  |  |  | 4.5 | 94.3 | 101.2 | 102.5
-9.5 |  |  |  |  | 4.6 | 94.7333 | 101.2 | 102.5
-9.4 |  |  |  |  | 4.7 | 94.95 | 101.2 | 102.5
-9.3 |  |  |  | 3 | 4.8 | 95.2 | 101.2167 | 102.5
-9.2 |  |  |  | 3.1 | 4.95 | 95.475 | 101.3 | 102.5
-9.1 |  |  |  | 3.2 | 5.1 | 95.76 | 101.3 | 102.5
-9 |  |  |  | 3.2 | 5.2 | 96.1 | 101.3 | 102.5
-8.9 |  |  |  | 3.3 | 5.4 | 96.42 | 101.3 | 102.5
-8.8 |  |  |  | 3.4 | 5.5 | 96.6667 | 101.4 | 102.5
-8.7 |  |  |  | 3.5 | 5.7 | 96.825 | 101.4 | 102.5
-8.6 |  |  |  | 3.5 | 5.8 | 97.075 | 101.4667 | 102.5
-8.5 |  |  |  | 3.6 | 6 | 97.325 | 101.5 | 102.4
-8.4 |  |  |  | 3.8 | 6.15 | 97.475 | 101.5 | 102.4
-8.3 |  |  |  | 3.9 | 6.35 | 97.7 | 101.6 | 102.4
-8.2 |  |  | 2.2 | 3.9 | 6.5 | 97.875 | 101.6 | 102.4
-8.1 |  |  | 2.3 | 4.0667 | 6.7 | 98.025 | 101.6 | 102.4
-8 |  |  | 2.4 | 4.2 | 6.9 | 98.175 | 101.6 | 102.4
-7.9 |  |  | 2.4 | 4.3 | 7.1667 | 98.3 | 101.6 | 102.4
-7.8 |  |  | 2.5 | 4.4 | 7.4 | 98.45 | 101.6 | 102.4
-7.7 |  |  | 2.5 | 4.5 | 7.55 | 98.7 | 101.6167 | 102.4
-7.6 |  |  | 2.6 | 4.7 | 7.9 | 98.8 | 101.7 | 102.4
-7.5 |  |  | 2.7 | 4.8 | 8.1 | 98.85 | 101.7 | 102.4
-7.4 |  |  | 2.7 | 5 | 8.4 | 99 | 101.7 | 102.3875
-7.3 |  |  | 2.8 | 5.1 | 8.7 | 99.1 | 101.7 | 102.3
-7.2 |  |  | 2.9 | 5.2 | 9 | 99.2333 | 101.7 | 102.3
-7.1 |  | 1.4 | 3 | 5.5 | 9.3 | 99.4 | 101.7 | 102.3
-7 |  | 1.5 | 3.1 | 5.6667 | 9.6 | 99.475 | 101.7 | 102.3
-6.9 |  | 1.5 | 3.2 | 5.8667 | 9.9 | 99.5 | 101.7 | 102.2
-6.8 |  | 1.6 | 3.2667 | 6 | 10.25 | 99.5 | 101.7 | 102.2
-6.7 |  | 1.6 | 3.4 | 6.2 | 10.6333 | 99.6 | 101.7 | 102.2
-6.6 |  | 1.6 | 3.5 | 6.4 | 11 | 99.8 | 101.7 | 102.2
-6.5 | 1.1 | 1.7 | 3.6 | 6.65 | 11.4667 | 99.9 | 101.7 | 102.2
-6.4 | 1.1 | 1.7 | 3.7 | 6.9 | 12 | 99.96 | 101.7 | 102.2
-6.3 | 1.1 | 1.8 | 3.8 | 7.1667 | 12.5667 | 100 | 101.7 | 102.2
-6.2 | 1.1 | 1.8 | 3.9 | 7.5 | 13.3 | 100.075 | 101.7 | 102.15
-6.1 | 1.2 | 1.9 | 4.1 | 7.6667 | 14.2333 | 100.125 | 101.7 | 102.1
-6 | 1.2 | 1.9 | 4.25 | 8 | 15.75 | 100.2 | 101.7 | 102.1
-5.9 | 1.2 | 2 | 4.4 | 8.2667 | 18.325 | 100.225 | 101.7 | 102.1
-5.8 | 1.3 | 2.05 | 4.5 | 8.625 | 22.75 | 100.3 | 101.7 | 102.0143
-5.7 | 1.3 | 2.1 | 4.7 | 8.92 | 30.15 | 100.375 | 101.7 | 102
-5.6 | 1.3 | 2.2 | 4.85 | 9.3 | 39.375 | 100.4 | 101.7 | 102
-5.5 | 1.4 | 2.3 | 5.075 | 9.6333 | 49.9 | 100.4 | 101.6 | 102
-5.4 | 1.4 | 2.35 | 5.225 | 10 | 60.7 | 100.475 | 101.6 | 101.9
-5.3 | 1.5 | 2.4 | 5.475 | 10.55 | 69.8 | 100.5 | 101.5 | 101.9
-5.2 | 1.5 | 2.5 | 5.7 | 11.075 | 78.8333 | 100.5 | 101.5 | 101.9
-5.1 | 1.6 | 2.6 | 5.875 | 11.8 | 85.6 | 100.5125 | 101.5 | 101.9
-5 | 1.6 | 2.7 | 6.125 | 12.7 | 89.5 | 100.6 | 101.5 | 101.8625
-4.9 | 1.65 | 2.8 | 6.375 | 14.075 | 91.7 | 100.6 | 101.5 | 101.8
-4.8 | 1.7 | 2.9 | 6.65 | 16.4 | 92.9667 | 100.6 | 101.4667 | 101.7
-4.7 | 1.8 | 3 | 6.9 | 20.8 | 93.8 | 100.6 | 101.4 | 101.6625
-4.6 | 1.9 | 3.1 | 7.2 | 27.6 | 94.6 | 100.6 | 101.4 | 101.6
-4.5 | 1.9 | 3.2667 | 7.5 | 36.6 | 95.15 | 100.6 | 101.4 | 101.5
-4.4 | 2 | 3.4 | 7.9 | 47.8 | 95.5 | 100.6 | 101.4 | 101.5
-4.3 | 2.1 | 3.55 | 8.2 | 58.5 | 95.875 | 100.6 | 101.4 | 101.5
-4.2 | 2.1 | 3.7 | 8.7 | 69.6 | 96.1 | 100.6 | 101.325 | 101.4
-4.1 | 2.2 | 3.9 | 9.2 | 78.6 | 96.5 | 100.6 | 101.3 | 101.3
-4 | 2.3 | 4.1 | 9.85 | 85.9 | 96.8 | 100.7 | 101.25 | 101.3
-3.9 | 2.4 | 4.2333 | 10.7667 | 90.2 | 97.1667 | 100.7 | 101.2 | 101.3
-3.8 | 2.5 | 4.45 | 12.15 | 92.4 | 97.4 | 100.6 | 101.15 | 101.3
-3.7 | 2.6 | 4.625 | 14.4 | 94 | 97.6 | 100.6 | 101.1 | 101.2429
-3.6 | 2.7667 | 4.8 | 18.7 | 94.85 | 97.75 | 100.6 | 101.1 | 101.2
-3.5 | 2.9 | 5.1 | 25.825 | 95.4 | 98 | 100.6 | 101.1 | 101.175
-3.4 | 3.05 | 5.35 | 35.275 | 95.9333 | 98.2 | 100.6 | 101.0143 | 101.1
-3.3 | 3.2 | 5.625 | 45.94 | 96.4 | 98.4 | 100.6 | 101 | 101.1
-3.2 | 3.3667 | 5.9 | 58 | 96.7333 | 98.6 | 100.6 | 101 | 101.1
-3.1 | 3.5 | 6.3333 | 68.8667 | 97 | 98.7 | 100.6 | 101 | 101.1
-3 | 3.7 | 6.8 | 78.775 | 97.3 | 98.9 | 100.6 | 100.9 | 101
-2.9 | 3.9 | 7.3667 | 86.5 | 97.5 | 99 | 100.5 | 100.8857 | 101
-2.8 | 4.1 | 8.225 | 91.2 | 97.8 | 99.2 | 100.5 | 100.8 | 100.925
-2.7 | 4.3 | 9.5667 | 93.7667 | 98 | 99.25 | 100.5 | 100.8 | 100.9
-2.6 | 4.65 | 11.9 | 95.2 | 98.4 | 99.4 | 100.4 | 100.8 | 100.8625
-2.5 | 5 | 16.2 | 96.1333 | 98.5 | 99.5 | 100.4 | 100.8 | 100.8
-2.4 | 5.5 | 23.9 | 96.65 | 98.7 | 99.6 | 100.4 | 100.7 | 100.7875
-2.3 | 6.1 | 33.5 | 97.1333 | 98.9 | 99.6 | 100.4 | 100.7 | 100.7
-2.2 | 7.15 | 44.9 | 97.5667 | 99 | 99.7 | 100.4 | 100.7 | 100.7
-2.1 | 9 | 58 | 97.8333 | 99.1 | 99.7 | 100.4 | 100.6 | 100.65
-2 | 12.25 | 69.5667 | 98.1 | 99.3 | 99.8 | 100.4 | 100.6 | 100.6
-1.9 | 17.8 | 79.925 | 98.3 | 99.4 | 99.9 | 100.4 | 100.6 | 100.6
-1.8 | 26.675 | 87.8333 | 98.6 | 99.5 | 99.9 | 100.3333 | 100.5857 | 100.5
-1.7 | 38.4 | 92.75 | 98.8 | 99.6 | 99.9 | 100.4 | 100.5 | 100.5
-1.6 | 51 | 95.45 | 99 | 99.7 | 100 | 100.3 | 100.5 | 100.4
-1.5 | 63.6 | 96.9 | 99.1 | 99.8 | 100 | 100.3 | 100.4 | 100.4
-1.4 | 75.2 | 97.75 | 99.3 | 99.8 | 100 | 100.3 | 100.4 | 100.4
-1.3 | 84.85 | 98.3333 | 99.4 | 99.9 | 100.1 | 100.2 | 100.4 | 100.3625
-1.2 | 91.7 | 98.85 | 99.5 | 99.9 | 100.1 | 100.2 | 100.3 | 100.2286
-1.1 | 95.3 | 99 | 99.6 | 100 | 100.1 | 100.2 | 100.3 | 100.2
-1 | 97.25 | 99.15 | 99.7 | 100 | 100.1 | 100.1 | 100.3 | 100.1
-0.9 | 98.35 | 99.35 | 99.8 | 100 | 100.1 | 100.1 | 100.2 | 100.1
-0.8 | 98.9 | 99.55 | 99.9 | 100 | 100.1 | 100 | 100.2 | 100.1
-0.7 | 99.25 | 99.6 | 99.9 | 100 | 100 | 100 | 100.2 | 100.1
-0.6 | 99.6 | 99.7 | 100 | 100 | 100 | 100 | 100.1167 | 100.1
-0.5 | 99.7 | 99.8 | 99.9 | 100.1 | 100 | 100 | 100.1 | 100.1
-0.4 | 99.8 | 99.8 | 100 | 100.1 | 100 | 100 | 100.1 | 100
-0.3 | 99.9 | 99.9 | 100 | 100.1 | 100 | 100 | 100.1 | 100
-0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0333 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100.05 | 100
0.3 | 99.9 | 99.9 | 100 | 100.1 | 100 | 100 | 100.1 | 100
0.4 | 99.8 | 99.8 | 100 | 100.1 | 100 | 100 | 100.1 | 100
0.5 | 99.7 | 99.75 | 99.9 | 100.1 | 100 | 100 | 100.1 | 100.0286
0.6 | 99.6 | 99.7 | 100 | 100 | 100 | 100 | 100.1167 | 100.1
0.7 | 99.25 | 99.6 | 99.9 | 100 | 100 | 100 | 100.2 | 100.1
0.8 | 98.9 | 99.55 | 99.9 | 100 | 100.1 | 100 | 100.2 | 100.1
0.9 | 98.35 | 99.35 | 99.8 | 100 | 100.1 | 100.1 | 100.2 | 100.1
1 | 97.2 | 99.15 | 99.7 | 100 | 100.1 | 100.1 | 100.3 | 100.1
1.1 | 95.4 | 99 | 99.6 | 100 | 100.1 | 100.2 | 100.3 | 100.1286
1.2 | 91.7 | 98.8333 | 99.5 | 99.9 | 100.1 | 100.2 | 100.3 | 100.2625
1.3 | 84.4 | 98.35 | 99.4 | 99.9 | 100.1 | 100.2 | 100.4 | 100.4
1.4 | 75.2 | 97.75 | 99.3 | 99.8 | 100 | 100.3 | 100.4 | 100.4
1.5 | 63.5 | 96.9 | 99.1 | 99.8 | 100 | 100.3 | 100.4 | 100.4
1.6 | 51 | 95.45 | 99 | 99.7 | 100 | 100.3 | 100.4667 | 100.4429
1.7 | 38.4 | 92.8 | 98.8 | 99.6 | 99.9667 | 100.4 | 100.5 | 100.5
1.8 | 26.6333 | 87.7 | 98.5667 | 99.5 | 99.9 | 100.3 | 100.5857 | 100.5
1.9 | 17.8 | 79.7 | 98.3 | 99.4 | 99.85 | 100.4 | 100.6 | 100.6
2 | 12.1 | 69.5667 | 98.1 | 99.3 | 99.8 | 100.4 | 100.6 | 100.6
2.1 | 9 | 58 | 97.8333 | 99.1 | 99.7 | 100.4 | 100.6 | 100.68
2.2 | 7.15 | 44.9 | 97.5667 | 99 | 99.7 | 100.4 | 100.7 | 100.7
2.3 | 6.1 | 33.8 | 97.125 | 98.85 | 99.6 | 100.4 | 100.7 | 100.7
2.4 | 5.5 | 23.9 | 96.75 | 98.7 | 99.6 | 100.4 | 100.7 | 100.725
2.5 | 5 | 16.2 | 96.1667 | 98.5 | 99.5 | 100.4 | 100.8 | 100.8
2.6 | 4.65 | 11.9 | 95.2 | 98.35 | 99.4 | 100.4 | 100.8 | 100.9
2.7 | 4.3 | 9.5 | 93.7 | 98 | 99.25 | 100.5 | 100.8 | 100.9
2.8 | 4.1333 | 8.3 | 91.2333 | 97.8 | 99.2 | 100.5 | 100.8 | 100.95
2.9 | 3.9 | 7.3667 | 86.5 | 97.5 | 99 | 100.5 | 100.8875 | 101
3 | 3.7 | 6.8 | 78.4 | 97.3 | 98.9 | 100.6 | 100.9 | 101
3.1 | 3.5 | 6.3333 | 68.75 | 97 | 98.7 | 100.6 | 101 | 101.08
3.2 | 3.3667 | 5.9 | 57.3 | 96.7333 | 98.6 | 100.6 | 101 | 101.1
3.3 | 3.2 | 5.625 | 45.8 | 96.4 | 98.4 | 100.6 | 101 | 101.1
3.4 | 3.05 | 5.35 | 34.7 | 95.9 | 98.2 | 100.6 | 101.0167 | 101.1
3.5 | 2.9 | 5.1 | 25.4 | 95.4 | 98 | 100.6 | 101.1 | 101.2
3.6 | 2.7667 | 4.8 | 18.68 | 94.8 | 97.8 | 100.6 | 101.1 | 101.2
3.7 | 2.6 | 4.625 | 14.5 | 94 | 97.6 | 100.6 | 101.1 | 101.275
3.8 | 2.5 | 4.4667 | 12.175 | 92.5 | 97.4 | 100.6 | 101.15 | 101.3
3.9 | 2.4 | 4.2333 | 10.7 | 90.25 | 97.15 | 100.7 | 101.2 | 101.3
4 | 2.3 | 4.1 | 9.9 | 85.7 | 96.8 | 100.7 | 101.25 | 101.3
4.1 | 2.2 | 3.9 | 9.2333 | 78.6333 | 96.4667 | 100.6 | 101.3 | 101.3
4.2 | 2.1 | 3.7 | 8.7 | 69 | 96.1 | 100.6 | 101.325 | 101.4
4.3 | 2.1 | 3.5333 | 8.2 | 58.5 | 95.86 | 100.6 | 101.4 | 101.4714
4.4 | 2 | 3.4 | 7.9 | 47.9 | 95.5 | 100.6 | 101.4 | 101.5
4.5 | 1.9 | 3.25 | 7.5 | 36.9 | 95.15 | 100.6 | 101.4 | 101.5286
4.6 | 1.9 | 3.1 | 7.2 | 27.7 | 94.6 | 100.6 | 101.4 | 101.6
4.7 | 1.8 | 3 | 6.9 | 20.9 | 93.8 | 100.6 | 101.4 | 101.6
4.8 | 1.7 | 2.9 | 6.725 | 16.5 | 92.95 | 100.6 | 101.4667 | 101.7286
4.9 | 1.6333 | 2.8 | 6.3667 | 14.0667 | 91.8 | 100.6 | 101.5 | 101.8
5 | 1.6 | 2.7 | 6.1 | 12.7 | 89.475 | 100.6 | 101.5 | 101.9
5.1 | 1.6 | 2.6 | 5.875 | 11.8 | 85.8 | 100.6 | 101.5 | 101.9
5.2 | 1.5 | 2.5 | 5.7 | 11.1 | 78.9667 | 100.5 | 101.5 | 101.9
5.3 | 1.5 | 2.4 | 5.475 | 10.55 | 70.2 | 100.5 | 101.5 | 101.9
5.4 | 1.4 | 2.3333 | 5.225 | 10 | 60.7 | 100.475 | 101.6 | 101.925
5.5 | 1.4 | 2.3 | 5.075 | 9.6333 | 49.9333 | 100.4 | 101.6 | 102
5.6 | 1.3 | 2.2 | 4.825 | 9.3 | 39.45 | 100.4 | 101.7 | 102
5.7 | 1.3 | 2.1 | 4.7 | 8.95 | 30.25 | 100.375 | 101.7 | 102
5.8 | 1.3 | 2.05 | 4.525 | 8.6 | 22.925 | 100.3 | 101.7 | 102.05
5.9 | 1.2 | 2 | 4.4 | 8.25 | 18.325 | 100.225 | 101.7 | 102.1
6 | 1.2 | 1.9 | 4.2333 | 8 | 15.825 | 100.2 | 101.7 | 102.1
6.1 | 1.2 | 1.9 | 4.075 | 7.6667 | 14.3 | 100.1 | 101.7 | 102.1
6.2 | 1.1 | 1.8 | 3.9 | 7.5 | 13.3 | 100.06 | 101.7 | 102.1857
6.3 | 1.1 | 1.8 | 3.8 | 7.15 | 12.5667 | 100 | 101.7 | 102.2
6.4 | 1.1 | 1.7 | 3.7 | 6.9 | 12 | 99.95 | 101.7 | 102.2
6.5 | 1.05 | 1.7 | 3.6 | 6.65 | 11.45 | 99.9 | 101.7 | 102.2
6.6 | 1 | 1.6 | 3.5 | 6.4 | 11 | 99.775 | 101.7 | 102.2
6.7 |  | 1.6 | 3.4 | 6.2 | 10.6333 | 99.6 | 101.7 | 102.2
6.8 |  | 1.6 | 3.2667 | 6 | 10.25 | 99.5 | 101.7 | 102.2
6.9 |  | 1.5 | 3.2 | 5.8667 | 9.9 | 99.5 | 101.7 | 102.2143
7 |  | 1.5 | 3.1 | 5.6667 | 9.6 | 99.475 | 101.7 | 102.3
7.1 |  | 1.4 | 3 | 5.5 | 9.3 | 99.4 | 101.7 | 102.3
7.2 |  |  | 2.9 | 5.3 | 9 | 99.2333 | 101.7 | 102.3
7.3 |  |  | 2.8 | 5.1 | 8.7 | 99.1 | 101.7 | 102.3
7.4 |  |  | 2.7 | 5 | 8.4 | 98.98 | 101.7 | 102.3286
7.5 |  |  | 2.7 | 4.8 | 8.1333 | 98.85 | 101.7 | 102.4
7.6 |  |  | 2.6 | 4.7 | 7.9 | 98.8 | 101.7 | 102.4
7.7 |  |  | 2.5 | 4.5 | 7.6 | 98.7 | 101.6143 | 102.4
7.8 |  |  | 2.5 | 4.4 | 7.35 | 98.45 | 101.6 | 102.4
7.9 |  |  | 2.4 | 4.3 | 7.15 | 98.3 | 101.6 | 102.4
8 |  |  | 2.4 | 4.2 | 7 | 98.175 | 101.6 | 102.4
8.1 |  |  | 2.3 | 4.0667 | 6.75 | 98.025 | 101.6 | 102.4
8.2 |  |  | 2.2 | 3.9 | 6.5 | 97.875 | 101.6 | 102.4
8.3 |  |  |  | 3.9 | 6.35 | 97.725 | 101.6 | 102.4
8.4 |  |  |  | 3.8 | 6.15 | 97.4667 | 101.5 | 102.4
8.5 |  |  |  | 3.6 | 6 | 97.32 | 101.5 | 102.4125
8.6 |  |  |  | 3.5 | 5.8 | 97.075 | 101.4667 | 102.5
8.7 |  |  |  | 3.5 | 5.7 | 96.8 | 101.4 | 102.5
8.8 |  |  |  | 3.4 | 5.5 | 96.6667 | 101.4 | 102.5
8.9 |  |  |  | 3.3 | 5.4 | 96.4 | 101.3 | 102.5
9 |  |  |  | 3.2 | 5.2 | 96.1 | 101.3 | 102.5
9.1 |  |  |  | 3.2 | 5.1 | 95.75 | 101.3 | 102.5
9.2 |  |  |  | 3.1 | 4.9333 | 95.4667 | 101.3 | 102.5
9.3 |  |  |  | 3 | 4.8 | 95.2 | 101.2167 | 102.5
9.4 |  |  |  |  | 4.7 | 95.025 | 101.2 | 102.5
9.5 |  |  |  |  | 4.6 | 94.7333 | 101.2 | 102.5
9.6 |  |  |  |  | 4.5 | 94.2667 | 101.2 | 102.5
9.7 |  |  |  |  | 4.4 | 93.9333 | 101.1 | 102.5
9.8 |  |  |  |  | 4.3 | 93.625 | 101 | 102.5
9.9 |  |  |  |  | 4.2 | 93.2 | 101 | 102.5
10 |  |  |  |  | 4.1 | 92.65 | 100.9 | 102.5
10.1 |  |  |  |  | 4 | 92.2 | 100.9 | 102.525
10.2 |  |  |  |  | 3.9 | 91.6 | 100.8417 | 102.6
10.3 |  |  |  |  | 3.8 | 90.7 | 100.8 | 102.6
10.4 |  |  |  |  | 3.7 | 89.5 | 100.8 | 102.6
10.5 | [5 empty] | 87.7333 | 100.8 | 102.6
10.6 | [5 empty] | 84.3 | 100.7 | 102.6
10.7 | [5 empty] | 78.6 | 100.7 | 102.7
10.8 | [5 empty] | 70.8 | 100.6 | 102.7
10.9 | [5 empty] | 61.9 | 100.5833 | 102.72
11 | [5 empty] | 52 | 100.5 | 102.8
11.1 | [5 empty] | 42.8 | 100.4 | 102.8
11.2 | [5 empty] | 34.0667 | 100.4 | 102.8
11.3 | [5 empty] | 26.85 | 100.3 | 102.7857
11.4 | [5 empty] | 22.6 | 100.2167 | 102.7
11.5 | [5 empty] | 20.1 | 100.15 | 102.7
11.6 | [5 empty] | 18.6 | 100.1 | 102.7
11.7 | [5 empty] | 17.5 | 100.0167 | 102.7
11.8 | [5 empty] | 16.7667 | 99.9625 | 102.7
11.9 | [5 empty] | 16.175 | 99.8375 | 102.7
12 | [5 empty] | 15.72 | 99.7833 | 102.7
12.1 | [5 empty] | 15.2857 | 99.6 | 102.7
12.2 | [5 empty] | 14.78 | 99.5 | 102.6
12.3 | [5 empty] | 14.45 | 99.4 | 102.6
12.4 | [5 empty] | 14.02 | 99.2667 | 102.6
12.5 | [5 empty] | 13.65 | 99.16 | 102.6
12.6 | [5 empty] | 13.36 | 99 | 102.5571
12.7 | [5 empty] | 12.98 | 98.9 | 102.5
12.8 | [5 empty] | 12.6667 | 98.8 | 102.5
12.9 | [5 empty] | 12.325 | 98.6429 | 102.4625
13 | [5 empty] | 12.025 | 98.5 | 102.4
13.1 | [5 empty] | 11.775 | 98.36 | 102.4
13.2 | [5 empty] | 11.4333 | 98.225 | 102.4
13.3 | [5 empty] | 11.22 | 98 | 102.4
13.4 | [5 empty] | 10.9 | 97.8857 | 102.4
13.5 | [5 empty] | 10.66 | 97.7333 | 102.4
13.6 | [5 empty] | 10.425 | 97.575 | 102.3
13.7 | [5 empty] | 10.2222 | 97.4333 | 102.3
13.8 | [5 empty] | 9.95 | 97.2667 | 102.3
13.9 | [5 empty] | 9.7556 | 97.0667 | 102.2714
14 | [5 empty] | 9.5571 | 96.8444 | 102.2
14.1 | [5 empty] | 9.325 | 96.6545 | 102.2
14.2 | [5 empty] | 9.1 | 96.425 | 102.2
14.3 | [5 empty] | 8.925 | 96.2 | 102.2
14.4 | [5 empty] | 8.75 | 96.025 | 102.1125
14.5 | [5 empty] | 8.62 | 95.775 | 102.1
14.6 | [5 empty] | 8.36 | 95.5 | 102.1
14.7 | [5 empty] | 8.225 | 95.1833 | 102.1
14.8 | [5 empty] | 8.075 | 95.0167 | 102.1
14.9 | [5 empty] | 7.9 | 94.7143 | 102.075
15 | [5 empty] | 7.775 | 94.4286 | 102
15.1 | [5 empty] | 7.6 | 94.1 | 102
15.2 | [5 empty] | 7.45 | 93.7571 | 102
15.3 | [5 empty] | 7.275 | 93.42 | 101.9
15.4 | [5 empty] | 7.2 | 92.94 | 101.9
15.5 | [5 empty] | 7.075 | 92.5 | 101.8857
15.6 | [5 empty] | 6.9 | 92.0667 | 101.8
15.7 | [5 empty] | 6.775 | 91.3 | 101.8
15.8 | [5 empty] | 6.6625 | 90.5667 | 101.7
15.9 | [5 empty] | 6.5 | 89.4143 | 101.62
16 | [6 empty] | 87.7833 | 101.6
16.1 | [6 empty] | 84.65 | 101.4571
16.2 | [6 empty] | 79.3333 | 101.4
16.3 | [6 empty] | 71.75 | 101.3
16.4 | [6 empty] | 62.9429 | 101.225
16.5 | [6 empty] | 53.55 | 101.2
16.6 | [6 empty] | 44.6 | 101.0625
16.7 | [6 empty] | 35.96 | 100.9286
16.8 | [6 empty] | 29.6 | 100.8
16.9 | [6 empty] | 25.1 | 100.7
17 | [6 empty] | 22.5 | 100.6286
17.1 | [6 empty] | 21 | 100.5
17.2 | [6 empty] | 20.0167 | 100.3714
17.3 | [6 empty] | 19.25 | 100.2286
17.4 | [6 empty] | 18.65 | 100.15
17.5 | [6 empty] | 18.05 | 99.92
17.6 | [6 empty] | 17.65 | 99.7429
17.7 | [6 empty] | 17.2 | 99.6125
17.8 | [6 empty] | 16.8222 | 99.51
17.9 | [6 empty] | 16.3857 | 99.3857
18 | [6 empty] | 16 | 99.25
18.1 | [6 empty] | 15.6667 | 99.0286
18.2 | [6 empty] | 15.3571 | 98.875
18.3 | [6 empty] | 15.1 | 98.7429
18.4 | [6 empty] | 14.7 | 98.6
18.5 | [6 empty] | 14.4143 | 98.4429
18.6 | [6 empty] | 14.1 | 98.2375
18.7 | [6 empty] | 13.7667 | 98
18.8 | [6 empty] | 13.54 | 97.8429
18.9 | [6 empty] | 13.25 | 97.6375
19 | [6 empty] | 12.98 | 97.5
19.1 | [6 empty] | 12.7 | 97.2429
19.2 | [6 empty] | 12.4091 | 96.975
19.3 | [6 empty] | 12.2143 | 96.8
19.4 | [6 empty] | 11.9667 | 96.65
19.5 | [6 empty] | 11.7 | 96.3933
19.6 | [6 empty] | 11.5286 | 96.1267
19.7 | [6 empty] | 11.2667 | 95.8923
19.8 | [6 empty] | 11.06 | 95.6909
19.9 | [6 empty] | 10.8667 | 95.45
20 | [6 empty] | 10.7143 | 95.15
20.1 | [6 empty] | 10.4636 | 94.875
20.2 | [6 empty] | 10.2875 | 94.5846
20.3 | [6 empty] | 10.125 | 94.24
20.4 | [6 empty] | 9.9375 | 93.84
20.5 | [6 empty] | 9.7 | 93.5
20.6 | [6 empty] | 9.6333 | 93.1714
20.7 | [6 empty] | 9.46 | 92.85
20.8 | [6 empty] | 9.2833 | 92.4429
20.9 | [6 empty] | 9.1167 | 91.9571
21 | [6 empty] | 8.9333 | 91.5
21.1 | [6 empty] | 8.8273 | 91
21.2 | [6 empty] | 8.6444 | 90.4286
21.3 | [6 empty] | 8.475 | 89.65
21.4 | [6 empty] | 8.35 | 88.475
21.5 | [7 empty] | 86.28
21.6 | [7 empty] | 83.1
21.7 | [7 empty] | 78.1
21.8 | [7 empty] | 71.3
21.9 | [7 empty] | 62.775
22 | [7 empty] | 53.9857
22.1 | [7 empty] | 45.3
22.2 | [7 empty] | 37.6
22.3 | [7 empty] | 31.4615
22.4 | [7 empty] | 26.55
22.5 | [7 empty] | 23.9
22.6 | [7 empty] | 22.3
22.7 | [7 empty] | 21.3067
22.8 | [7 empty] | 20.525
22.9 | [7 empty] | 19.9385
23 | [7 empty] | 19.3231
23.1 | [7 empty] | 18.88
23.2 | [7 empty] | 18.4667
23.3 | [7 empty] | 18.05
23.4 | [7 empty] | 17.6429
23.5 | [7 empty] | 17.2556
23.6 | [7 empty] | 16.8273
23.7 | [7 empty] | 16.4714
23.8 | [7 empty] | 16.15
23.9 | [7 empty] | 15.8364
24 | [7 empty] | 15.5125
24.1 | [7 empty] | 15.2444
24.2 | [7 empty] | 14.875
24.3 | [7 empty] | 14.6333
24.4 | [7 empty] | 14.3286
24.5 | [7 empty] | 14.025
24.6 | [7 empty] | 13.82
24.7 | [7 empty] | 13.575
24.8 | [7 empty] | 13.325
24.9 | [7 empty] | 13.0455
25 | [7 empty] | 12.7733
25.1 | [7 empty] | 12.5067
25.2 | [7 empty] | 12.28
25.3 | [7 empty] | 12.0833
25.4 | [7 empty] | 11.9167
25.5 | [7 empty] | 11.675
25.6 | [7 empty] | 11.5077
25.7 | [7 empty] | 11.2812
25.8 | [7 empty] | 11.0937
25.9 | [7 empty] | 10.8846
26 | [7 empty] | 10.6667
26.1 | [7 empty] | 10.5
26.2 | [7 empty] | 10.3571
26.3 | [7 empty] | 10.15
26.4 | [7 empty] | 9.975
26.5 | [7 empty] | 9.85
26.6 | [7 empty] | 9.65
26.7 | [7 empty] | 9.4778
26.8 | [7 empty] | 9.3571
26.9 | [7 empty] | 9.2143
27 | [7 empty] | 9.0875
27.1 | [7 empty] | 8.9625
27.2 | [7 empty] | 8.7889
27.3 | [7 empty] | 8.6455
27.4 | [7 empty] | 8.5333
27.5 | [7 empty] | 8.4
27.6 | [7 empty] | 8.2667
27.7 | [7 empty] | 8.1333
27.8 | [7 empty] | 8
27.9 | [7 empty] | 7.9231
28 | [7 empty] | 7.7923
28.1 | [7 empty] | 7.6556
28.2 | [7 empty] | 7.5444
28.3 | [7 empty] | 7.4455
28.4 | [7 empty] | 7.4
28.5 | [7 empty] | 7.3
28.6 | [7 empty] | 7.1778
28.7 | [7 empty] | 7.0667
28.8 | [7 empty] | 7
28.9 | [7 empty] | 6.9
29 | [7 empty] | 6.8
29.1 | [7 empty] | 6.725
29.2 | [7 empty] | 6.6
29.3 | [7 empty] | 6.5
29.4 | [7 empty] | 6.4375
29.5 | [7 empty] | 6.4
29.6 | [7 empty] | 6.3
29.7 | [7 empty] | 6.2286
29.8 | [7 empty] | 6.1

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.7 | [7 empty] | 12.4571
-29.6 | [7 empty] | 12.64
-29.5 | [7 empty] | 12.8143
-29.4 | [7 empty] | 13.0143
-29.3 | [7 empty] | 13.1875
-29.2 | [7 empty] | 13.325
-29.1 | [7 empty] | 13.4571
-29 | [7 empty] | 13.675
-28.9 | [7 empty] | 13.85
-28.8 | [7 empty] | 14.05
-28.7 | [7 empty] | 14.2714
-28.6 | [7 empty] | 14.4286
-28.5 | [7 empty] | 14.65
-28.4 | [7 empty] | 14.7857
-28.3 | [7 empty] | 15.0143
-28.2 | [7 empty] | 15.18
-28.1 | [7 empty] | 15.48
-28 | [7 empty] | 15.68
-27.9 | [7 empty] | 15.88
-27.8 | [7 empty] | 16.1
-27.7 | [7 empty] | 16.32
-27.6 | [7 empty] | 16.5
-27.5 | [7 empty] | 16.8
-27.4 | [7 empty] | 17.0667
-27.3 | [7 empty] | 17.2667
-27.2 | [7 empty] | 17.54
-27.1 | [7 empty] | 17.8167
-27 | [7 empty] | 18
-26.9 | [7 empty] | 18.3
-26.8 | [7 empty] | 18.6
-26.7 | [7 empty] | 18.9
-26.6 | [7 empty] | 19.16
-26.5 | [7 empty] | 19.4
-26.4 | [7 empty] | 19.7125
-26.3 | [7 empty] | 20.0875
-26.2 | [7 empty] | 20.4
-26.1 | [7 empty] | 20.6571
-26 | [7 empty] | 20.925
-25.9 | [7 empty] | 21.3
-25.8 | [7 empty] | 21.6857
-25.7 | [7 empty] | 22.025
-25.6 | [7 empty] | 22.4
-25.5 | [7 empty] | 22.6857
-25.4 | [7 empty] | 23.0875
-25.3 | [7 empty] | 23.6
-25.2 | [7 empty] | 23.9571
-25.1 | [7 empty] | 24.35
-25 | [7 empty] | 24.85
-24.9 | [7 empty] | 25.3
-24.8 | [7 empty] | 25.8857
-24.7 | [7 empty] | 26.625
-24.6 | [7 empty] | 27.6
-24.5 | [7 empty] | 29
-24.4 | [7 empty] | 31.175
-24.3 | [7 empty] | 34.5571
-24.2 | [7 empty] | 39.5
-24.1 | [7 empty] | 45.3857
-24 | [7 empty] | 52.02
-23.9 | [7 empty] | 58.6
-23.8 | [7 empty] | 64.24
-23.7 | [7 empty] | 69.7
-23.6 | [7 empty] | 73.66
-23.5 | [7 empty] | 76.26
-23.4 | [7 empty] | 77.86
-23.3 | [7 empty] | 78.94
-23.2 | [7 empty] | 79.64
-23.1 | [7 empty] | 80.28
-23 | [7 empty] | 80.86
-22.9 | [6 empty] | 12.7 | 81.32
-22.8 | [6 empty] | 12.8333 | 81.82
-22.7 | [6 empty] | 13.0833 | 82.22
-22.6 | [6 empty] | 13.325 | 82.65
-22.5 | [6 empty] | 13.45 | 83
-22.4 | [6 empty] | 13.65 | 83.25
-22.3 | [6 empty] | 13.9143 | 83.78
-22.2 | [6 empty] | 14.08 | 84.0333
-22.1 | [6 empty] | 14.3 | 84.2714
-22 | [6 empty] | 14.48 | 84.6714
-21.9 | [6 empty] | 14.6833 | 84.94
-21.8 | [6 empty] | 14.96 | 85.3429
-21.7 | [6 empty] | 15.15 | 85.76
-21.6 | [6 empty] | 15.3333 | 85.9667
-21.5 | [6 empty] | 15.6667 | 86.32
-21.4 | [6 empty] | 15.8833 | 86.68
-21.3 | [6 empty] | 16.1333 | 86.94
-21.2 | [6 empty] | 16.4333 | 87.14
-21.1 | [6 empty] | 16.625 | 87.44
-21 | [6 empty] | 16.8667 | 87.675
-20.9 | [6 empty] | 17.1714 | 88.0714
-20.8 | [6 empty] | 17.4143 | 88.2875
-20.7 | [6 empty] | 17.7286 | 88.475
-20.6 | [6 empty] | 18.0143 | 88.8143
-20.5 | [6 empty] | 18.3333 | 89.04
-20.4 | [6 empty] | 18.5833 | 89.1875
-20.3 | [6 empty] | 18.9 | 89.45
-20.2 | [6 empty] | 19.24 | 89.7143
-20.1 | [6 empty] | 19.5667 | 89.8875
-20 | [6 empty] | 19.875 | 90.25
-19.9 | [6 empty] | 20.2 | 90.4571
-19.8 | [6 empty] | 20.5571 | 90.675
-19.7 | [6 empty] | 20.8667 | 90.925
-19.6 | [6 empty] | 21.25 | 91.1556
-19.5 | [6 empty] | 21.7 | 91.4182
-19.4 | [6 empty] | 22 | 91.5778
-19.3 | [6 empty] | 22.4364 | 91.8
-19.2 | [6 empty] | 22.8 | 91.9333
-19.1 | [6 empty] | 23.3286 | 92.1333
-19 | [6 empty] | 23.7286 | 92.4
-18.9 | [6 empty] | 24.2 | 92.5333
-18.8 | [6 empty] | 24.8125 | 92.7333
-18.7 | [6 empty] | 25.4667 | 93
-18.6 | [6 empty] | 26.4143 | 93.1818
-18.5 | [6 empty] | 27.8333 | 93.3286
-18.4 | [6 empty] | 30.2 | 93.4667
-18.3 | [6 empty] | 33.3333 | 93.6
-18.2 | [6 empty] | 38.3667 | 93.7333
-18.1 | [6 empty] | 44.3857 | 93.925
-18 | [6 empty] | 51.65 | 94.1556
-17.9 | [6 empty] | 58.85 | 94.36
-17.8 | [6 empty] | 65.4444 | 94.48
-17.7 | [6 empty] | 71.5 | 94.5143
-17.6 | [6 empty] | 75.8857 | 94.7
-17.5 | [6 empty] | 78.6 | 94.875
-17.4 | [6 empty] | 80.4286 | 95.0125
-17.3 | [6 empty] | 81.6 | 95.16
-17.2 | [6 empty] | 82.45 | 95.3125
-17.1 | [6 empty] | 83.0143 | 95.4375
-17 | [6 empty] | 83.7 | 95.5
-16.9 | [5 empty] | 10.68 | 84.1143 | 95.6143
-16.8 | [5 empty] | 10.8667 | 84.6 | 95.7667
-16.7 | [5 empty] | 11.0333 | 84.95 | 95.8125
-16.6 | [5 empty] | 11.2 | 85.4 | 95.9375
-16.5 | [5 empty] | 11.3667 | 85.7286 | 96.0714
-16.4 | [5 empty] | 11.6 | 86.2 | 96.2
-16.3 | [5 empty] | 11.76 | 86.4818 | 96.3
-16.2 | [5 empty] | 11.9 | 86.86 | 96.3833
-16.1 | [5 empty] | 12.2 | 87.1 | 96.4286
-16 | [5 empty] | 12.4 | 87.4 | 96.5
-15.9 | [5 empty] | 12.6 | 87.7333 | 96.6167
-15.8 | [5 empty] | 12.8 | 88.14 | 96.7
-15.7 | [5 empty] | 13 | 88.3857 | 96.8143
-15.6 | [5 empty] | 13.1667 | 88.7 | 96.9571
-15.5 | [5 empty] | 13.5 | 88.9 | 97.0875
-15.4 | [5 empty] | 13.675 | 89.2167 | 97.1143
-15.3 | [5 empty] | 13.925 | 89.5 | 97.25
-15.2 | [5 empty] | 14.15 | 89.6667 | 97.3
-15.1 | [5 empty] | 14.4 | 89.96 | 97.4
-15 | [5 empty] | 14.6 | 90.2143 | 97.4833
-14.9 | [5 empty] | 14.9 | 90.45 | 97.55
-14.8 | [5 empty] | 15.16 | 90.74 | 97.6857
-14.7 | [5 empty] | 15.45 | 90.86 | 97.7143
-14.6 | [5 empty] | 15.8 | 91.0667 | 97.8
-14.5 | [5 empty] | 16.05 | 91.44 | 97.9
-14.4 | [5 empty] | 16.3 | 91.64 | 97.9
-14.3 | [5 empty] | 16.6 | 91.8222 | 98
-14.2 | [5 empty] | 16.95 | 92.04 | 98.05
-14.1 | [5 empty] | 17.3 | 92.225 | 98.2
-14 | [5 empty] | 17.6 | 92.4333 | 98.2
-13.9 | [5 empty] | 17.95 | 92.7167 | 98.3
-13.8 | [5 empty] | 18.3 | 92.9 | 98.3571
-13.7 | [5 empty] | 18.65 | 93.05 | 98.5
-13.6 | [5 empty] | 19 | 93.275 | 98.5875
-13.5 | [5 empty] | 19.4667 | 93.45 | 98.625
-13.4 | [5 empty] | 19.7 | 93.7167 | 98.7
-13.3 | [5 empty] | 20.1333 | 93.9 | 98.7
-13.2 | [5 empty] | 20.6 | 94.08 | 98.8
-13.1 | [5 empty] | 20.9667 | 94.2667 | 98.8571
-13 | [5 empty] | 21.5 | 94.4333 | 98.9
-12.9 | [5 empty] | 22 | 94.5 | 98.9333
-12.8 | [5 empty] | 22.7 | 94.6667 | 99.02
-12.7 | [5 empty] | 23.4 | 94.75 | 99.1
-12.6 | [5 empty] | 24.4 | 94.9 | 99.2
-12.5 | [5 empty] | 25.8 | 95.1 | 99.2
-12.4 | [5 empty] | 28 | 95.3 | 99.3
-12.3 | [5 empty] | 31.875 | 95.4 | 99.3
-12.2 | [5 empty] | 37.35 | 95.5286 | 99.3
-12.1 | [5 empty] | 44.4 | 95.6833 | 99.3
-12 | [5 empty] | 52.3 | 95.8 | 99.4
-11.9 | [5 empty] | 59.9 | 95.92 | 99.4429
-11.8 | [5 empty] | 67.3 | 96.1 | 99.5
-11.7 | [5 empty] | 73.5333 | 96.2 | 99.5125
-11.6 | [5 empty] | 78.4375 | 96.3 | 99.6
-11.5 | [5 empty] | 81.6167 | 96.4333 | 99.6125
-11.4 | [5 empty] | 83.55 | 96.5 | 99.7
-11.3 | [5 empty] | 84.8 | 96.6 | 99.7
-11.2 | [5 empty] | 85.7125 | 96.7167 | 99.7143
-11.1 | [5 empty] | 86.35 | 96.9 | 99.8
-11 | [5 empty] | 86.95 | 97 | 99.8
-10.9 |  |  |  |  | 6.76 | 87.4 | 97.1 | 99.8625
-10.8 |  |  |  |  | 6.8 | 87.9 | 97.2 | 99.9
-10.7 |  |  |  |  | 6.9667 | 88.325 | 97.3667 | 99.9375
-10.6 |  |  |  |  | 7.1125 | 88.725 | 97.5 | 100
-10.5 |  |  |  |  | 7.2375 | 89.075 | 97.5333 | 100
-10.4 |  |  |  |  | 7.3625 | 89.45 | 97.6 | 100
-10.3 |  |  |  |  | 7.575 | 89.875 | 97.68 | 100
-10.2 |  |  |  |  | 7.7 | 90.1333 | 97.7 | 100
-10.1 |  |  |  |  | 7.85 | 90.55 | 97.8 | 100.1
-10 |  |  |  |  | 8.05 | 90.875 | 98 | 100.1
-9.9 |  |  |  |  | 8.175 | 91.125 | 98.0833 | 100.2
-9.8 |  |  |  |  | 8.3 | 91.475 | 98.2 | 100.25
-9.7 |  |  |  | 5.5 | 8.525 | 91.825 | 98.2833 | 100.3
-9.6 |  |  |  | 5.625 | 8.7 | 92.075 | 98.35 | 100.325
-9.5 |  |  |  | 5.7 | 8.85 | 92.325 | 98.4 | 100.4
-9.4 |  |  |  | 5.85 | 9.1 | 92.5 | 98.4714 | 100.4
-9.3 |  |  |  | 6 | 9.325 | 92.85 | 98.6 | 100.5
-9.2 |  |  |  | 6.1 | 9.54 | 93.2 | 98.6 | 100.5
-9.1 |  |  |  | 6.2 | 9.675 | 93.45 | 98.7 | 100.5
-9 |  |  |  | 6.325 | 9.95 | 93.6 | 98.7 | 100.5
-8.9 |  |  |  | 6.5 | 10.175 | 93.925 | 98.8 | 100.5
-8.8 |  |  |  | 6.6 | 10.325 | 94.2 | 98.9 | 100.6
-8.7 |  |  |  | 6.8 | 10.575 | 94.425 | 98.9 | 100.6
-8.6 |  |  |  | 6.925 | 10.825 | 94.5875 | 99 | 100.6
-8.5 |  |  | 4.2 | 7.1 | 11.075 | 94.8 | 99.1 | 100.7
-8.4 |  |  | 4.2667 | 7.25 | 11.325 | 95.075 | 99.1571 | 100.7
-8.3 |  |  | 4.4 | 7.425 | 11.65 | 95.2286 | 99.2 | 100.7
-8.2 |  |  | 4.4667 | 7.6 | 11.9 | 95.5 | 99.3 | 100.8
-8.1 |  |  | 4.5 | 7.8 | 12.25 | 95.725 | 99.3 | 100.8
-8 |  |  | 4.6 | 8 | 12.5 | 95.85 | 99.4 | 100.7
-7.9 |  |  | 4.8 | 8.1 | 12.9 | 96 | 99.4857 | 100.7286
-7.8 |  |  | 4.8 | 8.3333 | 13.1667 | 96.1 | 99.55 | 100.8
-7.7 |  |  | 5 | 8.5 | 13.5 | 96.2 | 99.6 | 100.8
-7.6 |  |  | 5.05 | 8.8 | 13.85 | 96.45 | 99.6833 | 100.8
-7.5 |  |  | 5.2 | 8.975 | 14.175 | 96.6 | 99.7 | 100.8
-7.4 |  |  | 5.3 | 9.2 | 14.55 | 96.82 | 99.7 | 100.8625
-7.3 |  | 2.7 | 5.45 | 9.45 | 14.95 | 96.95 | 99.7714 | 100.9
-7.2 |  | 2.8 | 5.6 | 9.7 | 15.45 | 97.1 | 99.8 | 100.9
-7.1 |  | 2.8 | 5.7 | 9.875 | 15.85 | 97.225 | 99.8 | 100.9
-7 |  | 2.8 | 5.8 | 10.2 | 16.25 | 97.375 | 99.9 | 100.85
-6.9 |  | 2.9 | 6 | 10.5 | 16.85 | 97.5 | 99.94 | 100.8
-6.8 |  | 3 | 6.15 | 10.7667 | 17.38 | 97.6 | 100 | 100.8
-6.7 | 1.9 | 3.1 | 6.3 | 11 | 18.2 | 97.7 | 100.0833 | 100.8533
-6.6 | 1.9333 | 3.1 | 6.5 | 11.3 | 19.2 | 97.84 | 100.1 | 100.9
-6.5 | 2 | 3.2 | 6.65 | 11.65 | 20.8 | 97.975 | 100.1 | 100.9
-6.4 | 2.0333 | 3.3 | 6.8 | 12 | 23.5 | 98.1 | 100.2 | 100.9
-6.3 | 2.1 | 3.3333 | 7.05 | 12.35 | 27.7667 | 98.2 | 100.2 | 100.9
-6.2 | 2.2 | 3.4 | 7.2 | 12.6333 | 34.2667 | 98.3 | 100.2 | 100.9
-6.1 | 2.2 | 3.6 | 7.4 | 12.9667 | 42.4 | 98.3 | 100.2 | 100.9
-6 | 2.2 | 3.6 | 7.6 | 13.35 | 50.9333 | 98.4 | 100.2 | 100.9
-5.9 | 2.275 | 3.7 | 7.75 | 13.76 | 60.175 | 98.5 | 100.2 | 100.9
-5.8 | 2.3 | 3.8 | 7.98 | 14.3 | 69.5 | 98.74 | 100.3 | 100.9
-5.7 | 2.4 | 3.9 | 8.3 | 14.7667 | 77.4 | 98.8 | 100.3 | 100.9
-5.6 | 2.5 | 4 | 8.475 | 15.4 | 83.4333 | 98.9 | 100.35 | 100.9
-5.5 | 2.5 | 4.1 | 8.725 | 16.1 | 86.9 | 98.9 | 100.4 | 100.9
-5.4 | 2.6 | 4.2 | 9 | 17.175 | 89.175 | 99 | 100.3429 | 100.9
-5.3 | 2.675 | 4.35 | 9.325 | 18.5333 | 90.5 | 99.1 | 100.3 | 100.9
-5.2 | 2.7 | 4.5 | 9.575 | 20.925 | 91.475 | 99.2 | 100.3 | 100.9
-5.1 | 2.8 | 4.6 | 9.8 | 25.275 | 92.15 | 99.2 | 100.3 | 100.8
-5 | 2.8 | 4.7333 | 10.1 | 31.05 | 92.85 | 99.3 | 100.3 | 100.8
-4.9 | 2.9 | 4.9 | 10.5 | 39.6 | 93.25 | 99.3167 | 100.3 | 100.8
-4.8 | 3 | 5.1 | 10.8 | 49.0667 | 93.64 | 99.4 | 100.3 | 100.8
-4.7 | 3.1 | 5.2 | 11.2 | 58.6 | 94.08 | 99.4 | 100.4 | 100.8
-4.6 | 3.2 | 5.4 | 11.6 | 68.9 | 94.475 | 99.5 | 100.4 | 100.8
-4.5 | 3.3 | 5.575 | 12.1 | 76.7 | 94.8 | 99.5 | 100.4 | 100.8
-4.4 | 3.4 | 5.7 | 12.6 | 83.55 | 95.2 | 99.5 | 100.4 | 100.8
-4.3 | 3.5 | 5.9 | 13.35 | 87.8667 | 95.5333 | 99.6 | 100.4 | 100.8
-4.2 | 3.6 | 6.1 | 14.3 | 90.4 | 95.825 | 99.7 | 100.4 | 100.8
-4.1 | 3.7 | 6.2667 | 15.8 | 91.95 | 96.05 | 99.7 | 100.4 | 100.8
-4 | 3.9 | 6.5 | 18.3 | 93 | 96.2 | 99.7 | 100.4 | 100.7556
-3.9 | 4 | 6.8 | 22 | 93.8 | 96.5 | 99.8 | 100.4 | 100.7
-3.8 | 4.1333 | 7 | 28.6 | 94.26 | 96.825 | 99.8 | 100.4 | 100.7
-3.7 | 4.3 | 7.25 | 37.5 | 94.7 | 97.1 | 99.9 | 100.4 | 100.675
-3.6 | 4.4333 | 7.5 | 47.8 | 95.18 | 97.2667 | 99.9 | 100.4 | 100.6
-3.5 | 4.6 | 7.85 | 58.5333 | 95.45 | 97.4333 | 99.9 | 100.4 | 100.6
-3.4 | 4.725 | 8.15 | 68.3 | 95.875 | 97.6333 | 99.95 | 100.4 | 100.6
-3.3 | 4.975 | 8.55 | 77.8667 | 96.125 | 97.8 | 100 | 100.4 | 100.6
-3.2 | 5.125 | 9.05 | 84.8 | 96.375 | 98.05 | 100 | 100.4 | 100.6
-3.1 | 5.3 | 9.7333 | 89.5667 | 96.725 | 98.2 | 100.05 | 100.35 | 100.6
-3 | 5.525 | 10.6 | 92.5 | 97.02 | 98.425 | 100.1 | 100.4 | 100.6
-2.9 | 5.7667 | 12.1 | 93.9667 | 97.2 | 98.575 | 100.1 | 100.4 | 100.6
-2.8 | 6 | 14.4 | 95 | 97.4 | 98.7 | 100.1 | 100.4 | 100.6
-2.7 | 6.4 | 18.7 | 95.9 | 97.6 | 98.8333 | 100.1 | 100.3 | 100.5
-2.6 | 6.7667 | 25.35 | 96.4 | 97.85 | 98.9 | 100.1 | 100.3 | 100.5
-2.5 | 7.4 | 34.9 | 96.7 | 98.1 | 99.0333 | 100.1 | 100.3 | 100.5
-2.4 | 8.4 | 46.3 | 97.1 | 98.25 | 99.2 | 100.1 | 100.3 | 100.5
-2.3 | 9.7 | 57.1 | 97.4 | 98.4 | 99.3 | 100.1 | 100.3 | 100.5
-2.2 | 12.2 | 69.1 | 97.55 | 98.5 | 99.4 | 100.1 | 100.3 | 100.5
-2.1 | 16.4 | 78.2 | 97.9333 | 98.7 | 99.4 | 100.1 | 100.3 | 100.4875
-2 | 23.4 | 86.1 | 98.2 | 98.84 | 99.5 | 100.1 | 100.3 | 100.4
-1.9 | 33.5 | 91.65 | 98.4 | 99.075 | 99.6 | 100.1 | 100.3 | 100.4
-1.8 | 44.4 | 94.4 | 98.6 | 99.2 | 99.6 | 100.1 | 100.3 | 100.3
-1.7 | 57 | 96.2 | 98.8 | 99.26 | 99.7 | 100.1 | 100.2 | 100.3
-1.6 | 68.9 | 97.2 | 98.9 | 99.3 | 99.7 | 100.1 | 100.2 | 100.3
-1.5 | 78.8 | 97.9667 | 99 | 99.4 | 99.8 | 100.1 | 100.2833 | 100.2
-1.4 | 86.9 | 98.3333 | 99.2 | 99.5 | 99.9 | 100.1 | 100.275 | 100.2
-1.3 | 92.9 | 98.65 | 99.4 | 99.55 | 99.9 | 100.1 | 100.2333 | 100.2
-1.2 | 96 | 99 | 99.5 | 99.6 | 100 | 100.1 | 100.2 | 100.2
-1.1 | 97.5 | 99.2 | 99.6 | 99.6 | 100 | 100.1 | 100.1 | 100.2
-1 | 98.2 | 99.3 | 99.7 | 99.7 | 100 | 100.1 | 100.0727 | 100.175
-0.9 | 98.8667 | 99.425 | 99.7 | 99.7 | 100 | 100.1 | 100 | 100.1
-0.8 | 99.2 | 99.5667 | 99.8 | 99.8 | 100.025 | 100.1 | 100 | 100.1
-0.7 | 99.5 | 99.6667 | 99.8 | 99.8 | 100.075 | 100 | 100 | 100.1
-0.6 | 99.7 | 99.7 | 99.9 | 99.9 | 100.1 | 100 | 100 | 100.1
-0.5 | 99.9 | 99.8 | 100 | 99.9 | 100 | 100 | 100.0857 | 100.1
-0.4 | 100 | 99.8 | 100 | 100 | 100 | 100 | 100.1 | 100.075
-0.3 | 100 | 99.875 | 100 | 100 | 100 | 100 | 100.0167 | 100
-0.2 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 99.86 | 100 | 100 | 100 | 100 | 100.06 | 100
0.4 | 100 | 99.8 | 100 | 100 | 100 | 100 | 100.1 | 100.075
0.5 | 99.9 | 99.8 | 100 | 99.9 | 100 | 100 | 100.0571 | 100.1
0.6 | 99.7 | 99.7 | 99.9 | 99.9 | 100.1 | 100 | 100 | 100.1
0.7 | 99.5 | 99.6667 | 99.8 | 99.825 | 100.06 | 100 | 100 | 100.1
0.8 | 99.2 | 99.5667 | 99.8 | 99.8 | 100.025 | 100.1 | 100 | 100.1
0.9 | 98.85 | 99.4 | 99.7 | 99.7 | 100 | 100.1 | 100.025 | 100.1
1 | 98.2 | 99.3 | 99.7 | 99.7 | 100 | 100.1 | 100.1 | 100.175
1.1 | 97.5 | 99.2 | 99.6 | 99.6 | 100 | 100.1 | 100.1 | 100.2
1.2 | 95.85 | 99 | 99.5 | 99.6 | 100 | 100.1 | 100.2 | 100.2
1.3 | 92.9 | 98.65 | 99.4 | 99.55 | 99.9 | 100.1 | 100.2 | 100.2
1.4 | 86.9 | 98.3333 | 99.2 | 99.5 | 99.9 | 100.1 | 100.25 | 100.2
1.5 | 79.3 | 98 | 99.05 | 99.4 | 99.8 | 100.1 | 100.25 | 100.2
1.6 | 68.75 | 97.2 | 98.9 | 99.3 | 99.7 | 100.1 | 100.2 | 100.3
1.7 | 57.1 | 96.2 | 98.7667 | 99.25 | 99.7 | 100.1 | 100.2 | 100.3
1.8 | 44.4 | 94.6 | 98.6 | 99.2 | 99.6 | 100.1 | 100.275 | 100.3
1.9 | 33.3 | 91.5 | 98.4 | 99.075 | 99.6 | 100.1 | 100.3 | 100.4
2 | 23.4 | 86.1 | 98.2 | 98.9 | 99.5 | 100.1 | 100.3 | 100.4
2.1 | 16.6 | 78.2 | 97.9 | 98.7 | 99.4 | 100.1 | 100.3 | 100.4875
2.2 | 12.3 | 69.1 | 97.5333 | 98.5 | 99.4 | 100.1 | 100.3 | 100.5
2.3 | 9.7 | 57.1 | 97.4 | 98.4 | 99.3 | 100.1 | 100.3 | 100.5
2.4 | 8.4 | 46.3 | 97.1 | 98.25 | 99.2 | 100.1 | 100.3 | 100.5
2.5 | 7.4 | 34.7 | 96.7 | 98.1 | 99.0333 | 100.1 | 100.3 | 100.5
2.6 | 6.8 | 25.95 | 96.4 | 97.85 | 98.9 | 100.1 | 100.3 | 100.5
2.7 | 6.4 | 18.7 | 95.9 | 97.6 | 98.825 | 100.1 | 100.3 | 100.5
2.8 | 6.1 | 14.4 | 95.025 | 97.4 | 98.7 | 100.1 | 100.3714 | 100.6
2.9 | 5.8 | 12.1 | 93.9667 | 97.2 | 98.575 | 100.1 | 100.4 | 100.6
3 | 5.6 | 10.6 | 92.4 | 97.02 | 98.425 | 100.1 | 100.4 | 100.6
3.1 | 5.3 | 9.7333 | 89.5667 | 96.725 | 98.2 | 100.05 | 100.3833 | 100.6
3.2 | 5.12 | 9 | 85.1 | 96.375 | 98.025 | 100 | 100.4 | 100.6
3.3 | 4.975 | 8.5667 | 77.7 | 96.125 | 97.775 | 100 | 100.4 | 100.6
3.4 | 4.725 | 8.1 | 68.1 | 95.875 | 97.625 | 99.95 | 100.4 | 100.6
3.5 | 4.6 | 7.8333 | 58.35 | 95.45 | 97.4333 | 99.9 | 100.4 | 100.6
3.6 | 4.425 | 7.55 | 47.8 | 95.18 | 97.2667 | 99.9 | 100.4 | 100.6
3.7 | 4.3 | 7.2 | 37.5 | 94.7 | 97.075 | 99.9 | 100.4 | 100.675
3.8 | 4.125 | 7 | 28.7 | 94.26 | 96.8 | 99.8 | 100.4 | 100.7
3.9 | 4 | 6.8 | 22.1 | 93.8 | 96.5 | 99.8 | 100.4 | 100.7
4 | 3.9 | 6.5 | 18.3 | 93 | 96.2 | 99.7 | 100.4 | 100.7556
4.1 | 3.7 | 6.2667 | 15.75 | 91.95 | 96.05 | 99.7 | 100.4 | 100.8
4.2 | 3.6 | 6.1 | 14.3 | 90.4 | 95.8 | 99.7 | 100.4 | 100.8
4.3 | 3.5 | 5.9 | 13.35 | 87.8667 | 95.5333 | 99.6 | 100.4 | 100.8
4.4 | 3.4 | 5.7 | 12.6 | 83.55 | 95.175 | 99.5 | 100.4 | 100.8
4.5 | 3.3 | 5.575 | 12.1 | 76.7 | 94.88 | 99.5 | 100.4 | 100.8
4.6 | 3.2 | 5.4 | 11.6 | 68.9 | 94.475 | 99.5 | 100.4 | 100.8
4.7 | 3.1 | 5.2 | 11.2 | 58.6 | 94.08 | 99.4 | 100.4 | 100.8
4.8 | 3 | 5.1 | 10.8 | 49.0667 | 93.7 | 99.4 | 100.36 | 100.8
4.9 | 2.9 | 4.9 | 10.5 | 39.75 | 93.25 | 99.3143 | 100.3 | 100.8
5 | 2.8 | 4.7333 | 10.1 | 31.15 | 92.84 | 99.3 | 100.3 | 100.8
5.1 | 2.8 | 4.6 | 9.8 | 24.9 | 92.15 | 99.2 | 100.3 | 100.8
5.2 | 2.7 | 4.5 | 9.55 | 20.925 | 91.475 | 99.2 | 100.3 | 100.9
5.3 | 2.65 | 4.35 | 9.3 | 18.6167 | 90.5 | 99.1 | 100.3 | 100.9
5.4 | 2.6 | 4.2 | 9 | 17.175 | 89.2 | 99 | 100.375 | 100.9
5.5 | 2.5 | 4.1 | 8.725 | 16.08 | 86.975 | 98.9 | 100.4 | 100.9
5.6 | 2.5 | 4 | 8.475 | 15.4 | 83.4333 | 98.9 | 100.4 | 100.9
5.7 | 2.4 | 3.9 | 8.3 | 14.7667 | 77.4 | 98.8 | 100.3 | 100.9
5.8 | 2.3 | 3.8 | 7.975 | 14.3 | 69.5 | 98.725 | 100.3 | 100.9
5.9 | 2.275 | 3.7 | 7.75 | 13.75 | 60.35 | 98.5 | 100.25 | 100.9
6 | 2.2 | 3.6 | 7.6 | 13.375 | 50.925 | 98.4 | 100.2 | 100.9
6.1 | 2.2 | 3.6 | 7.4 | 12.9667 | 41.9 | 98.3 | 100.2 | 100.9
6.2 | 2.2 | 3.4 | 7.2 | 12.6333 | 33.6 | 98.3 | 100.2 | 100.9
6.3 | 2.1 | 3.3 | 7.0333 | 12.35 | 27.6 | 98.2 | 100.2 | 100.9
6.4 | 2.025 | 3.3 | 6.8 | 12 | 23.5 | 98.1 | 100.2 | 100.9
6.5 | 2 | 3.2 | 6.65 | 11.65 | 20.8 | 97.975 | 100.1 | 100.9
6.6 | 1.9333 | 3.1 | 6.5 | 11.3 | 19.25 | 97.825 | 100.1 | 100.9
6.7 | 1.9 | 3.1 | 6.3 | 11 | 18.2 | 97.7 | 100.05 | 100.8533
6.8 |  | 3 | 6.2 | 10.7667 | 17.4 | 97.6 | 100 | 100.8
6.9 |  | 2.9 | 6 | 10.5 | 16.8 | 97.5 | 99.9 | 100.8
7 |  | 2.8 | 5.9 | 10.2 | 16.28 | 97.375 | 99.9 | 100.85
7.1 |  | 2.8 | 5.7 | 9.875 | 15.84 | 97.225 | 99.85 | 100.9
7.2 |  | 2.8 | 5.6 | 9.7 | 15.4 | 97.1 | 99.8 | 100.9
7.3 |  | 2.7 | 5.4333 | 9.45 | 14.95 | 96.95 | 99.7333 | 100.9
7.4 |  |  | 5.3 | 9.2 | 14.625 | 96.8 | 99.7 | 100.8625
7.5 |  |  | 5.2 | 8.975 | 14.175 | 96.6 | 99.7 | 100.8
7.6 |  |  | 5.1 | 8.8 | 13.85 | 96.45 | 99.6429 | 100.8
7.7 |  |  | 5 | 8.575 | 13.475 | 96.2 | 99.6 | 100.8
7.8 |  |  | 4.85 | 8.325 | 13.1667 | 96.1 | 99.5167 | 100.8
7.9 |  |  | 4.75 | 8.1 | 12.9 | 96 | 99.45 | 100.7286
8 |  |  | 4.6 | 8 | 12.48 | 95.85 | 99.4 | 100.7
8.1 |  |  | 4.5 | 7.8 | 12.25 | 95.7 | 99.3 | 100.8
8.2 |  |  | 4.45 | 7.6 | 11.9 | 95.48 | 99.2667 | 100.8
8.3 |  |  | 4.4 | 7.4 | 11.65 | 95.2286 | 99.2 | 100.7
8.4 |  |  | 4.3 | 7.24 | 11.4 | 95.0714 | 99.2 | 100.7
8.5 |  |  | 4.2 | 7.1 | 11.075 | 94.7857 | 99.1 | 100.7
8.6 |  |  |  | 6.9 | 10.8 | 94.5875 | 99 | 100.6
8.7 |  |  |  | 6.8 | 10.5667 | 94.425 | 98.9 | 100.6
8.8 |  |  |  | 6.6 | 10.325 | 94.2 | 98.8667 | 100.6
8.9 |  |  |  | 6.5 | 10.175 | 93.925 | 98.8 | 100.5
9 |  |  |  | 6.325 | 9.925 | 93.6 | 98.75 | 100.5
9.1 |  |  |  | 6.2 | 9.7 | 93.44 | 98.7 | 100.5
9.2 |  |  |  | 6.1 | 9.525 | 93.175 | 98.6 | 100.5
9.3 |  |  |  | 5.975 | 9.325 | 92.85 | 98.6 | 100.5
9.4 |  |  |  | 5.825 | 9.075 | 92.5 | 98.5 | 100.4
9.5 |  |  |  | 5.7 | 8.85 | 92.325 | 98.4 | 100.4
9.6 |  |  |  | 5.625 | 8.7 | 92.075 | 98.3167 | 100.325
9.7 |  |  |  | 5.5 | 8.5143 | 91.825 | 98.2429 | 100.3
9.8 |  |  |  |  | 8.3 | 91.475 | 98.2 | 100.25
9.9 |  |  |  |  | 8.175 | 91.125 | 98.1 | 100.2
10 |  |  |  |  | 8.05 | 90.875 | 97.9833 | 100.1
10.1 |  |  |  |  | 7.85 | 90.55 | 97.8 | 100.1
10.2 |  |  |  |  | 7.7 | 90.1333 | 97.7 | 100
10.3 |  |  |  |  | 7.5875 | 89.875 | 97.6429 | 100
10.4 |  |  |  |  | 7.425 | 89.45 | 97.6 | 100
10.5 |  |  |  |  | 7.2375 | 89.075 | 97.5 | 100
10.6 |  |  |  |  | 7.1 | 88.725 | 97.5 | 100
10.7 |  |  |  |  | 6.9667 | 88.325 | 97.3333 | 99.9375
10.8 |  |  |  |  | 6.8 | 87.88 | 97.2 | 99.9
10.9 |  |  |  |  | 6.76 | 87.35 | 97.1 | 99.8625
11 | [5 empty] | 86.925 | 97.025 | 99.8
11.1 | [5 empty] | 86.35 | 96.9 | 99.8
11.2 | [5 empty] | 85.725 | 96.7833 | 99.7143
11.3 | [5 empty] | 84.8625 | 96.6167 | 99.7
11.4 | [5 empty] | 83.55 | 96.5 | 99.7
11.5 | [5 empty] | 81.7 | 96.3833 | 99.6125
11.6 | [5 empty] | 78.45 | 96.3 | 99.6
11.7 | [5 empty] | 73.5 | 96.2 | 99.5125
11.8 | [5 empty] | 67.3333 | 96.1 | 99.5
11.9 | [5 empty] | 60.225 | 95.9 | 99.4429
12 | [5 empty] | 52.4 | 95.8167 | 99.4
12.1 | [5 empty] | 44.6 | 95.65 | 99.3
12.2 | [5 empty] | 37.65 | 95.5833 | 99.3
12.3 | [5 empty] | 32.05 | 95.4286 | 99.3
12.4 | [5 empty] | 28.1 | 95.3 | 99.3
12.5 | [5 empty] | 25.875 | 95.1 | 99.2
12.6 | [5 empty] | 24.4 | 94.95 | 99.2
12.7 | [5 empty] | 23.35 | 94.8 | 99.1
12.8 | [5 empty] | 22.7 | 94.6167 | 99.02
12.9 | [5 empty] | 22 | 94.5 | 98.9333
13 | [5 empty] | 21.5 | 94.4 | 98.9
13.1 | [5 empty] | 20.9667 | 94.22 | 98.8571
13.2 | [5 empty] | 20.6 | 94.05 | 98.8
13.3 | [5 empty] | 20.1333 | 93.85 | 98.7
13.4 | [5 empty] | 19.7 | 93.6667 | 98.7
13.5 | [5 empty] | 19.4667 | 93.5 | 98.625
13.6 | [5 empty] | 19 | 93.2333 | 98.5875
13.7 | [5 empty] | 18.65 | 93.02 | 98.5
13.8 | [5 empty] | 18.3 | 92.875 | 98.3571
13.9 | [5 empty] | 17.95 | 92.68 | 98.3
14 | [5 empty] | 17.6 | 92.4714 | 98.2
14.1 | [5 empty] | 17.36 | 92.3 | 98.2
14.2 | [5 empty] | 16.95 | 92 | 98.05
14.3 | [5 empty] | 16.6 | 91.8429 | 98
14.4 | [5 empty] | 16.3 | 91.6 | 97.9
14.5 | [5 empty] | 16.05 | 91.4 | 97.9
14.6 | [5 empty] | 15.8 | 91.0667 | 97.8
14.7 | [5 empty] | 15.45 | 90.9333 | 97.7143
14.8 | [5 empty] | 15.15 | 90.7 | 97.6857
14.9 | [5 empty] | 14.9 | 90.4333 | 97.55
15 | [5 empty] | 14.6 | 90.22 | 97.4833
15.1 | [5 empty] | 14.4 | 89.94 | 97.4
15.2 | [5 empty] | 14.15 | 89.7 | 97.3
15.3 | [5 empty] | 13.925 | 89.45 | 97.25
15.4 | [5 empty] | 13.6667 | 89.1833 | 97.1143
15.5 | [5 empty] | 13.475 | 88.9333 | 97.0875
15.6 | [5 empty] | 13.1667 | 88.6333 | 96.9571
15.7 | [5 empty] | 13 | 88.3857 | 96.8143
15.8 | [5 empty] | 12.8 | 88.0667 | 96.7
15.9 | [5 empty] | 12.6 | 87.6857 | 96.6167
16 | [5 empty] | 12.4 | 87.4333 | 96.5
16.1 | [5 empty] | 12.2 | 87.15 | 96.4286
16.2 | [5 empty] | 12 | 86.8333 | 96.3833
16.3 | [5 empty] | 11.75 | 86.5 | 96.3
16.4 | [5 empty] | 11.6 | 86.1333 | 96.2
16.5 | [5 empty] | 11.3667 | 85.7429 | 96.0714
16.6 | [5 empty] | 11.2 | 85.38 | 95.9375
16.7 | [5 empty] | 11 | 85.02 | 95.8125
16.8 | [5 empty] | 10.85 | 84.62 | 95.7667
16.9 | [5 empty] | 10.68 | 84.1333 | 95.6143
17 | [6 empty] | 83.62 | 95.5
17.1 | [6 empty] | 82.9571 | 95.4375
17.2 | [6 empty] | 82.46 | 95.3125
17.3 | [6 empty] | 81.6 | 95.16
17.4 | [6 empty] | 80.6 | 95.0125
17.5 | [6 empty] | 78.7818 | 94.875
17.6 | [6 empty] | 75.9 | 94.7
17.7 | [6 empty] | 71.3111 | 94.5143
17.8 | [6 empty] | 65.6091 | 94.48
17.9 | [6 empty] | 58.8333 | 94.36
18 | [6 empty] | 51.78 | 94.1556
18.1 | [6 empty] | 44.325 | 93.925
18.2 | [6 empty] | 38.2143 | 93.7333
18.3 | [6 empty] | 33.1556 | 93.6
18.4 | [6 empty] | 29.9455 | 93.4667
18.5 | [6 empty] | 27.84 | 93.3286
18.6 | [6 empty] | 26.4 | 93.1818
18.7 | [6 empty] | 25.5111 | 93
18.8 | [6 empty] | 24.825 | 92.7333
18.9 | [6 empty] | 24.24 | 92.5333
19 | [6 empty] | 23.7429 | 92.4
19.1 | [6 empty] | 23.3429 | 92.1333
19.2 | [6 empty] | 22.8333 | 91.9333
19.3 | [6 empty] | 22.4143 | 91.8
19.4 | [6 empty] | 21.95 | 91.5778
19.5 | [6 empty] | 21.5875 | 91.4182
19.6 | [6 empty] | 21.2333 | 91.1556
19.7 | [6 empty] | 20.9143 | 90.925
19.8 | [6 empty] | 20.475 | 90.675
19.9 | [6 empty] | 20.2143 | 90.4571
20 | [6 empty] | 19.9 | 90.25
20.1 | [6 empty] | 19.55 | 89.8875
20.2 | [6 empty] | 19.24 | 89.7143
20.3 | [6 empty] | 18.94 | 89.45
20.4 | [6 empty] | 18.6 | 89.1875
20.5 | [6 empty] | 18.2667 | 89.04
20.6 | [6 empty] | 17.9333 | 88.8143
20.7 | [6 empty] | 17.725 | 88.475
20.8 | [6 empty] | 17.4333 | 88.2875
20.9 | [6 empty] | 17.16 | 88.0714
21 | [6 empty] | 16.9 | 87.675
21.1 | [6 empty] | 16.66 | 87.44
21.2 | [6 empty] | 16.4 | 87.14
21.3 | [6 empty] | 16.1667 | 86.94
21.4 | [6 empty] | 15.9 | 86.68
21.5 | [6 empty] | 15.65 | 86.32
21.6 | [6 empty] | 15.3833 | 85.9667
21.7 | [6 empty] | 15.2167 | 85.76
21.8 | [6 empty] | 14.92 | 85.3429
21.9 | [6 empty] | 14.75 | 84.94
22 | [6 empty] | 14.55 | 84.6714
22.1 | [6 empty] | 14.2667 | 84.2714
22.2 | [6 empty] | 14.1 | 84.0333
22.3 | [6 empty] | 13.8667 | 83.78
22.4 | [6 empty] | 13.6 | 83.25
22.5 | [6 empty] | 13.5 | 83
22.6 | [6 empty] | 13.3 | 82.65
22.7 | [6 empty] | 13.05 | 82.22
22.8 | [6 empty] | 12.8833 | 81.82
22.9 | [6 empty] | 12.7167 | 81.32
23 | [7 empty] | 80.86
23.1 | [7 empty] | 80.28
23.2 | [7 empty] | 79.64
23.3 | [7 empty] | 78.94
23.4 | [7 empty] | 77.86
23.5 | [7 empty] | 76.26
23.6 | [7 empty] | 73.66
23.7 | [7 empty] | 69.7
23.8 | [7 empty] | 64.24
23.9 | [7 empty] | 58.6
24 | [7 empty] | 52.02
24.1 | [7 empty] | 45.3857
24.2 | [7 empty] | 39.5
24.3 | [7 empty] | 34.5571
24.4 | [7 empty] | 31.175
24.5 | [7 empty] | 29
24.6 | [7 empty] | 27.6
24.7 | [7 empty] | 26.625
24.8 | [7 empty] | 25.8857
24.9 | [7 empty] | 25.3
25 | [7 empty] | 24.85
25.1 | [7 empty] | 24.35
25.2 | [7 empty] | 23.9571
25.3 | [7 empty] | 23.6
25.4 | [7 empty] | 23.0875
25.5 | [7 empty] | 22.6857
25.6 | [7 empty] | 22.4
25.7 | [7 empty] | 22.025
25.8 | [7 empty] | 21.6857
25.9 | [7 empty] | 21.3
26 | [7 empty] | 20.925
26.1 | [7 empty] | 20.6571
26.2 | [7 empty] | 20.4
26.3 | [7 empty] | 20.0875
26.4 | [7 empty] | 19.7125
26.5 | [7 empty] | 19.4
26.6 | [7 empty] | 19.16
26.7 | [7 empty] | 18.9
26.8 | [7 empty] | 18.6
26.9 | [7 empty] | 18.3
27 | [7 empty] | 18
27.1 | [7 empty] | 17.8167
27.2 | [7 empty] | 17.54
27.3 | [7 empty] | 17.2667
27.4 | [7 empty] | 17.0667
27.5 | [7 empty] | 16.8
27.6 | [7 empty] | 16.5
27.7 | [7 empty] | 16.32
27.8 | [7 empty] | 16.1
27.9 | [7 empty] | 15.88
28 | [7 empty] | 15.68
28.1 | [7 empty] | 15.48
28.2 | [7 empty] | 15.18
28.3 | [7 empty] | 15.0143
28.4 | [7 empty] | 14.7857
28.5 | [7 empty] | 14.65
28.6 | [7 empty] | 14.4286
28.7 | [7 empty] | 14.2714
28.8 | [7 empty] | 14.05
28.9 | [7 empty] | 13.85
29 | [7 empty] | 13.675
29.1 | [7 empty] | 13.4571
29.2 | [7 empty] | 13.325
29.3 | [7 empty] | 13.1875
29.4 | [7 empty] | 13.0143
29.5 | [7 empty] | 12.8143
29.6 | [7 empty] | 12.64
29.7 | [7 empty] | 12.4571

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 19.7231
-29.7 | [7 empty] | 19.9538
-29.6 | [7 empty] | 20.2
-29.5 | [7 empty] | 20.45
-29.4 | [7 empty] | 20.7
-29.3 | [7 empty] | 20.9143
-29.2 | [7 empty] | 21.16
-29.1 | [7 empty] | 21.42
-29 | [7 empty] | 21.62
-28.9 | [7 empty] | 21.88
-28.8 | [7 empty] | 22.16
-28.7 | [7 empty] | 22.43
-28.6 | [7 empty] | 22.7333
-28.5 | [7 empty] | 23.05
-28.4 | [7 empty] | 23.28
-28.3 | [7 empty] | 23.5714
-28.2 | [7 empty] | 23.875
-28.1 | [7 empty] | 24.175
-28 | [7 empty] | 24.5143
-27.9 | [7 empty] | 24.775
-27.8 | [7 empty] | 25.1286
-27.7 | [7 empty] | 25.4
-27.6 | [7 empty] | 25.725
-27.5 | [7 empty] | 26.0667
-27.4 | [6 empty] | 11.8667 | 26.4273
-27.3 | [6 empty] | 11.9778 | 26.8375
-27.2 | [6 empty] | 12.2 | 27.22
-27.1 | [6 empty] | 12.3143 | 27.6
-27 | [6 empty] | 12.45 | 28
-26.9 | [6 empty] | 12.5857 | 28.5
-26.8 | [6 empty] | 12.7 | 29
-26.7 | [6 empty] | 12.88 | 29.65
-26.6 | [6 empty] | 12.9667 | 30.4
-26.5 | [6 empty] | 13.1 | 31.7667
-26.4 | [6 empty] | 13.38 | 33.7
-26.3 | [6 empty] | 13.4375 | 36.5667
-26.2 | [6 empty] | 13.6 | 40.6231
-26.1 | [6 empty] | 13.7375 | 45.3923
-26 | [6 empty] | 13.8625 | 50.5778
-25.9 | [6 empty] | 14.1 | 55.4571
-25.8 | [6 empty] | 14.2286 | 60.4125
-25.7 | [6 empty] | 14.3625 | 64.8
-25.6 | [6 empty] | 14.6 | 67.8143
-25.5 | [6 empty] | 14.7 | 70.0375
-25.4 | [6 empty] | 14.9333 | 71.6
-25.3 | [6 empty] | 15.1714 | 72.6
-25.2 | [6 empty] | 15.325 | 73.2429
-25.1 | [6 empty] | 15.45 | 73.675
-25 | [6 empty] | 15.575 | 74.3
-24.9 | [6 empty] | 15.8 | 74.7714
-24.8 | [6 empty] | 15.975 | 75.25
-24.7 | [6 empty] | 16.1625 | 75.775
-24.6 | [6 empty] | 16.3 | 76.1143
-24.5 | [6 empty] | 16.525 | 76.3556
-24.4 | [6 empty] | 16.6571 | 76.8
-24.3 | [6 empty] | 16.875 | 77.1333
-24.2 | [6 empty] | 17.1714 | 77.5077
-24.1 | [6 empty] | 17.3875 | 77.8875
-24 | [6 empty] | 17.5286 | 78.175
-23.9 | [6 empty] | 17.7667 | 78.45
-23.8 | [6 empty] | 17.96 | 78.8667
-23.7 | [6 empty] | 18.16 | 79.15
-23.6 | [6 empty] | 18.3286 | 79.4286
-23.5 | [6 empty] | 18.6143 | 79.75
-23.4 | [6 empty] | 18.7875 | 80.0385
-23.3 | [6 empty] | 19.05 | 80.275
-23.2 | [6 empty] | 19.3 | 80.53
-23.1 | [6 empty] | 19.5714 | 80.8462
-23 | [6 empty] | 19.75 | 81.3077
-22.9 | [6 empty] | 20 | 81.5714
-22.8 | [6 empty] | 20.325 | 81.8
-22.7 | [6 empty] | 20.525 | 82.0714
-22.6 | [6 empty] | 20.7857 | 82.4
-22.5 | [6 empty] | 21.05 | 82.65
-22.4 | [6 empty] | 21.35 | 82.9
-22.3 | [6 empty] | 21.6 | 83.075
-22.2 | [6 empty] | 21.925 | 83.4125
-22.1 | [6 empty] | 22.125 | 83.575
-22 | [6 empty] | 22.4286 | 83.8818
-21.9 | [6 empty] | 22.7714 | 84.1727
-21.8 | [6 empty] | 23.0375 | 84.375
-21.7 | [6 empty] | 23.42 | 84.6429
-21.6 | [6 empty] | 23.7429 | 84.9
-21.5 | [6 empty] | 23.95 | 85.1375
-21.4 | [6 empty] | 24.325 | 85.3429
-21.3 | [6 empty] | 24.6286 | 85.56
-21.2 | [6 empty] | 25.025 | 85.7
-21.1 | [6 empty] | 25.3375 | 85.975
-21 | [6 empty] | 25.7364 | 86.2429
-20.9 | [6 empty] | 26.1444 | 86.4
-20.8 | [6 empty] | 26.4 | 86.6667
-20.7 | [6 empty] | 26.7714 | 86.8571
-20.6 | [6 empty] | 27.2 | 87.075
-20.5 | [6 empty] | 27.6375 | 87.325
-20.4 | [6 empty] | 28.1769 | 87.5143
-20.3 | [6 empty] | 28.8571 | 87.775
-20.2 | [6 empty] | 29.5 | 87.95
-20.1 | [6 empty] | 30.4857 | 88.2143
-20 | [6 empty] | 31.9286 | 88.475
-19.9 | [6 empty] | 34.4 | 88.6286
-19.8 | [6 empty] | 37.84 | 88.85
-19.7 | [6 empty] | 42.35 | 88.975
-19.6 | [6 empty] | 47.62 | 89.125
-19.5 | [6 empty] | 53.3909 | 89.3429
-19.4 | [6 empty] | 59.1571 | 89.55
-19.3 | [6 empty] | 64.425 | 89.7125
-19.2 | [6 empty] | 69.025 | 89.8857
-19.1 | [6 empty] | 72.1571 | 90.075
-19 | [6 empty] | 74.5857 | 90.2
-18.9 | [6 empty] | 75.9 | 90.48
-18.8 | [6 empty] | 76.8143 | 90.5286
-18.7 | [6 empty] | 77.475 | 90.6625
-18.6 | [6 empty] | 78.15 | 90.8
-18.5 | [6 empty] | 78.6714 | 91.0286
-18.4 | [6 empty] | 79.2 | 91.1714
-18.3 | [6 empty] | 79.7375 | 91.3286
-18.2 | [6 empty] | 80.075 | 91.4
-18.1 | [6 empty] | 80.4143 | 91.58
-18 | [6 empty] | 80.8 | 91.8286
-17.9 | [5 empty] | 14.1 | 81.1125 | 91.9833
-17.8 | [5 empty] | 14.36 | 81.5143 | 92.15
-17.7 | [5 empty] | 14.5 | 82 | 92.2875
-17.6 | [5 empty] | 14.7 | 82.2571 | 92.425
-17.5 | [5 empty] | 14.8667 | 82.6143 | 92.5
-17.4 | [5 empty] | 15.1 | 83 | 92.6222
-17.3 | [5 empty] | 15.3 | 83.2375 | 92.7
-17.2 | [5 empty] | 15.475 | 83.4429 | 92.7857
-17.1 | [5 empty] | 15.7 | 83.8 | 92.9125
-17 | [5 empty] | 15.9667 | 84.16 | 93.0429
-16.9 | [5 empty] | 16.15 | 84.4125 | 93.175
-16.8 | [5 empty] | 16.4 | 84.7429 | 93.31
-16.7 | [5 empty] | 16.7 | 85 | 93.425
-16.6 | [5 empty] | 17 | 85.2375 | 93.63
-16.5 | [5 empty] | 17.18 | 85.4429 | 93.7
-16.4 | [5 empty] | 17.4 | 85.8 | 93.775
-16.3 | [5 empty] | 17.6 | 85.975 | 93.88
-16.2 | [5 empty] | 17.9 | 86.3143 | 93.9667
-16.1 | [5 empty] | 18.1333 | 86.64 | 94.1
-16 | [5 empty] | 18.5 | 86.8778 | 94.2444
-15.9 | [5 empty] | 18.75 | 87.1 | 94.3625
-15.8 | [5 empty] | 19.05 | 87.3273 | 94.4
-15.7 | [5 empty] | 19.3 | 87.6 | 94.5
-15.6 | [5 empty] | 19.5714 | 87.7364 | 94.55
-15.5 | [5 empty] | 19.86 | 87.9333 | 94.6857
-15.4 | [5 empty] | 20.2 | 88.25 | 94.7818
-15.3 | [5 empty] | 20.5 | 88.5 | 94.9067
-15.2 | [5 empty] | 20.8 | 88.6857 | 95.04
-15.1 | [5 empty] | 21.1333 | 88.9286 | 95.1733
-15 | [5 empty] | 21.4 | 89.15 | 95.2615
-14.9 | [5 empty] | 21.78 | 89.3714 | 95.3625
-14.8 | [5 empty] | 22.1 | 89.54 | 95.4778
-14.7 | [5 empty] | 22.55 | 89.7667 | 95.6
-14.6 | [5 empty] | 22.9 | 90.05 | 95.7
-14.5 | [5 empty] | 23.175 | 90.3125 | 95.7625
-14.4 | [5 empty] | 23.6333 | 90.4429 | 95.9
-14.3 | [5 empty] | 24.025 | 90.65 | 96.025
-14.2 | [5 empty] | 24.45 | 90.825 | 96.1
-14.1 | [5 empty] | 24.8667 | 91.0375 | 96.1875
-14 | [5 empty] | 25.3333 | 91.1714 | 96.2286
-13.9 | [5 empty] | 25.85 | 91.3 | 96.3625
-13.8 | [5 empty] | 26.375 | 91.4375 | 96.4875
-13.7 | [5 empty] | 27.125 | 91.5833 | 96.6
-13.6 | [5 empty] | 28 | 91.875 | 96.6625
-13.5 | [5 empty] | 29.55 | 92.0714 | 96.8
-13.4 | [5 empty] | 31.6 | 92.2 | 96.8
-13.3 | [5 empty] | 35.15 | 92.3286 | 96.8667
-13.2 | [5 empty] | 39.9 | 92.525 | 96.925
-13.1 | [5 empty] | 45.6 | 92.7 | 97.05
-13 | [5 empty] | 52.5333 | 92.825 | 97.1857
-12.9 | [5 empty] | 59.2 | 93.0143 | 97.3
-12.8 | [5 empty] | 65.6 | 93.1 | 97.3
-12.7 | [5 empty] | 71.3 | 93.35 | 97.4
-12.6 | [5 empty] | 75.4 | 93.5571 | 97.4
-12.5 | [5 empty] | 78.55 | 93.6875 | 97.5
-12.4 | [5 empty] | 80.3 | 93.75 | 97.6286
-12.3 | [5 empty] | 81.3 | 93.9571 | 97.7
-12.2 | [5 empty] | 82.3 | 94.0875 | 97.8
-12.1 | [5 empty] | 82.9 | 94.125 | 97.8
-12 | [5 empty] | 83.45 | 94.2571 | 97.8
-11.9 | [5 empty] | 84.05 | 94.3875 | 97.8875
-11.8 | [5 empty] | 84.4667 | 94.525 | 97.925
-11.7 | [5 empty] | 85.0333 | 94.6889 | 98.0571
-11.6 | [5 empty] | 85.425 | 94.8 | 98.1
-11.5 | [5 empty] | 85.75 | 94.875 | 98.1571
-11.4 |  |  |  |  | 9.36 | 86.125 | 95.025 | 98.2
-11.3 |  |  |  |  | 9.575 | 86.475 | 95.26 | 98.2143
-11.2 |  |  |  |  | 9.7 | 86.925 | 95.4429 | 98.3
-11.1 |  |  |  |  | 9.88 | 87.175 | 95.5857 | 98.3857
-11 |  |  |  |  | 9.9667 | 87.525 | 95.725 | 98.5
-10.9 |  |  |  |  | 10.18 | 87.95 | 95.8571 | 98.5
-10.8 |  |  |  |  | 10.3 | 88.2 | 95.9875 | 98.5857
-10.7 |  |  |  |  | 10.475 | 88.5333 | 96.0286 | 98.6125
-10.6 |  |  |  |  | 10.625 | 88.85 | 96.1 | 98.7
-10.5 |  |  |  |  | 10.8 | 89.2667 | 96.2 | 98.7
-10.4 |  |  |  |  | 10.975 | 89.5 | 96.3 | 98.8
-10.3 |  |  |  |  | 11.2 | 89.725 | 96.3571 | 98.8429
-10.2 |  |  |  |  | 11.4 | 89.975 | 96.5 | 98.9
-10.1 |  |  |  | 7.8 | 11.625 | 90.3 | 96.6143 | 98.9
-10 |  |  |  | 7.9 | 11.8667 | 90.625 | 96.75 | 99
-9.9 |  |  |  | 8 | 12.025 | 90.9 | 96.875 | 99.0714
-9.8 |  |  |  | 8.15 | 12.2 | 91.25 | 96.9889 | 99.1
-9.7 |  |  |  | 8.3 | 12.4 | 91.5 | 97.0143 | 99.1375
-9.6 |  |  |  | 8.4 | 12.75 | 91.7 | 97.1 | 99.2
-9.5 |  |  |  | 8.6 | 12.9 | 91.94 | 97.15 | 99.3
-9.4 |  |  |  | 8.8 | 13.15 | 92.2 | 97.2 | 99.3
-9.3 |  |  |  | 8.9 | 13.4 | 92.4 | 97.3111 | 99.3
-9.2 |  |  |  | 9.1 | 13.7 | 92.8 | 97.4267 | 99.3
-9.1 |  |  |  | 9.25 | 13.975 | 93 | 97.56 | 99.4
-9 |  |  |  | 9.4 | 14.225 | 93.1 | 97.6467 | 99.4
-8.9 |  |  |  | 9.5 | 14.475 | 93.3 | 97.7182 | 99.5
-8.8 |  |  | 5.9 | 9.7333 | 14.825 | 93.45 | 97.8071 | 99.5
-8.7 |  |  | 6 | 9.9 | 15.075 | 93.775 | 97.8786 | 99.6
-8.6 |  |  | 6.1 | 10.15 | 15.35 | 93.9 | 97.9636 | 99.6
-8.5 |  |  | 6.3 | 10.3 | 15.66 | 94 | 98.06 | 99.7
-8.4 |  |  | 6.4 | 10.6 | 16 | 94.325 | 98.1 | 99.7
-8.3 |  |  | 6.5 | 10.7 | 16.3 | 94.46 | 98.25 | 99.76
-8.2 |  |  | 6.6 | 11 | 16.7 | 94.56 | 98.3467 | 99.8
-8.1 |  |  | 6.7 | 11.2 | 17 | 94.7667 | 98.4167 | 99.86
-8 |  |  | 6.8333 | 11.4 | 17.35 | 95 | 98.5 | 99.9125
-7.9 |  |  | 6.9 | 11.6333 | 17.6667 | 95.2 | 98.5222 | 100
-7.8 |  |  | 7.1 | 11.9 | 18.075 | 95.3 | 98.6 | 100
-7.7 |  |  | 7.2 | 12.15 | 18.5 | 95.425 | 98.675 | 100
-7.6 |  |  | 7.4 | 12.4 | 18.9 | 95.6 | 98.8 | 100
-7.5 |  | 3.7 | 7.6 | 12.7 | 19.375 | 95.825 | 98.8 | 100
-7.4 |  | 3.8 | 7.7 | 12.9 | 19.925 | 95.96 | 98.8 | 100.0625
-7.3 |  | 3.95 | 7.8 | 13.2333 | 20.625 | 96.15 | 98.9 | 100.1
-7.2 |  | 4 | 8.1 | 13.5 | 21.425 | 96.3 | 98.9714 | 100.1
-7.1 |  | 4.1 | 8.2 | 13.825 | 22.525 | 96.4 | 99 | 100.1625
-7 |  | 4.2 | 8.4 | 14.0667 | 24.175 | 96.5833 | 99.1 | 100.2
-6.9 |  | 4.3 | 8.5 | 14.4 | 26.68 | 96.8 | 99.1 | 100.2
-6.8 | 2.7 | 4.35 | 8.7 | 14.75 | 30.675 | 96.8625 | 99.2 | 100.2714
-6.7 | 2.8 | 4.4333 | 8.95 | 15.1 | 36.56 | 96.975 | 99.2 | 100.3
-6.6 | 2.9 | 4.5 | 9.15 | 15.45 | 43.64 | 97.1 | 99.2 | 100.3
-6.5 | 2.9 | 4.625 | 9.3 | 15.75 | 52.2 | 97.225 | 99.3 | 100.3
-6.4 | 3 | 4.7 | 9.5 | 16.175 | 60.95 | 97.3333 | 99.3 | 100.3
-6.3 | 3 | 4.8 | 9.75 | 16.6 | 69.175 | 97.5 | 99.36 | 100.3538
-6.2 | 3 | 4.9 | 9.9333 | 17 | 76.32 | 97.6 | 99.5 | 100.4
-6.1 | 3.1 | 5.1 | 10.1 | 17.5 | 81.4 | 97.66 | 99.52 | 100.4
-6 | 3.2 | 5.1 | 10.4 | 18.1333 | 85.3 | 97.8 | 99.6 | 100.4417
-5.9 | 3.2 | 5.2 | 10.8 | 19 | 87.8 | 97.9 | 99.6 | 100.5
-5.8 | 3.3 | 5.4 | 10.9 | 20 | 89.2 | 98 | 99.6 | 100.5
-5.7 | 3.4 | 5.5 | 11.2 | 21.575 | 90.2 | 98.1 | 99.65 | 100.5
-5.6 | 3.4 | 5.6667 | 11.4 | 24 | 90.8 | 98.1 | 99.7 | 100.5
-5.5 | 3.55 | 5.8 | 11.8 | 28.1 | 91.5 | 98.15 | 99.7 | 100.5
-5.4 | 3.6 | 5.9 | 12.1 | 34.2 | 91.9 | 98.3 | 99.7 | 100.5
-5.3 | 3.7 | 6 | 12.4 | 41.4 | 92.4 | 98.4333 | 99.7 | 100.4
-5.2 | 3.8 | 6.1667 | 12.7 | 50.1 | 92.7667 | 98.5 | 99.7143 | 100.4
-5.1 | 3.9 | 6.3 | 13 | 59.1 | 93.25 | 98.6 | 99.8 | 100.4
-5 | 4 | 6.5 | 13.45 | 67.9 | 93.625 | 98.7 | 99.8 | 100.5
-4.9 | 4.1 | 6.6 | 13.75 | 75.9 | 93.875 | 98.7333 | 99.9 | 100.5
-4.8 | 4.125 | 6.7333 | 14.25 | 82.25 | 94.2 | 98.8 | 99.9 | 100.5
-4.7 | 4.25 | 6.9 | 14.9 | 86.7 | 94.45 | 98.9 | 99.9 | 100.5
-4.6 | 4.3 | 7.2 | 15.6 | 89.1333 | 94.76 | 99 | 99.9429 | 100.5
-4.5 | 4.4667 | 7.4 | 16.7 | 91 | 95.2 | 99 | 100 | 100.4857
-4.4 | 4.6 | 7.55 | 18 | 92.1 | 95.4571 | 99.1 | 100 | 100.4
-4.3 | 4.7 | 7.8 | 20.5 | 92.9 | 95.7 | 99.1 | 100 | 100.4
-4.2 | 4.8 | 8.0333 | 24.7 | 93.4333 | 95.9 | 99.1 | 100 | 100.4
-4.1 | 5 | 8.3 | 30.75 | 93.9 | 96.1333 | 99.2 | 100 | 100.4
-4 | 5.15 | 8.5 | 39.2 | 94.4 | 96.45 | 99.25 | 100 | 100.4133
-3.9 | 5.3 | 8.75 | 47.9 | 94.8 | 96.66 | 99.3 | 100 | 100.48
-3.8 | 5.4 | 9.05 | 57.5 | 95.1 | 96.8 | 99.325 | 100 | 100.5
-3.7 | 5.5333 | 9.4 | 68 | 95.6333 | 96.95 | 99.475 | 100.1 | 100.5
-3.6 | 5.7 | 9.65 | 76.2 | 95.85 | 97.1 | 99.5 | 100.1 | 100.5
-3.5 | 5.85 | 10.1 | 83.8 | 96.1 | 97.3 | 99.575 | 100.1 | 100.5
-3.4 | 6.075 | 10.6 | 88.1 | 96.35 | 97.5 | 99.6 | 100.1 | 100.4571
-3.3 | 6.3 | 11.25 | 91.6 | 96.55 | 97.65 | 99.6 | 100.1 | 100.4
-3.2 | 6.5 | 12.3 | 93.2 | 96.8 | 97.9 | 99.6 | 100.125 | 100.3
-3.1 | 6.7 | 13.9 | 94.1 | 97.0333 | 98.1 | 99.8 | 100.2 | 100.3
-3 | 7 | 16.5 | 95.1 | 97.2 | 98.225 | 99.8 | 100.2 | 100.3556
-2.9 | 7.3 | 20.5 | 95.75 | 97.3333 | 98.35 | 99.8 | 100.2 | 100.4
-2.8 | 7.75 | 27.1 | 96.3 | 97.6 | 98.5 | 99.8 | 100.2 | 100.4
-2.7 | 8.3 | 36.15 | 96.6 | 97.8 | 98.5667 | 99.8 | 100.2 | 100.4
-2.6 | 9 | 45.8667 | 96.95 | 98.0667 | 98.74 | 99.85 | 100.2 | 100.4
-2.5 | 10.2 | 57.2 | 97.2 | 98.3 | 98.85 | 99.9 | 100.2 | 100.4
-2.4 | 12.2 | 67.8333 | 97.45 | 98.4667 | 98.9 | 99.925 | 100.2 | 100.4
-2.3 | 15.6 | 77.5 | 97.7 | 98.6 | 99 | 100 | 100.2 | 100.4
-2.2 | 20.8 | 85.1 | 97.9 | 98.75 | 99.1 | 100 | 100.14 | 100.4
-2.1 | 29.2 | 90.5 | 98.1333 | 98.9 | 99.2 | 99.95 | 100.1 | 100.3
-2 | 38.8 | 93.6 | 98.35 | 99 | 99.2 | 99.9 | 100.1 | 100.3
-1.9 | 50.5667 | 95.6 | 98.6 | 99.1 | 99.275 | 99.9 | 100.1 | 100.3
-1.8 | 61.9 | 96.8 | 98.8 | 99.2 | 99.3 | 99.9 | 100.2 | 100.2429
-1.7 | 72.3 | 97.4667 | 98.9 | 99.3 | 99.4 | 99.9 | 100.2 | 100.3
-1.6 | 81.8 | 97.9667 | 99 | 99.4 | 99.5 | 99.9 | 100.2 | 100.3
-1.5 | 89.1 | 98.3 | 99 | 99.5 | 99.6 | 99.9 | 100.2 | 100.2857
-1.4 | 93.3 | 98.72 | 99.1 | 99.5 | 99.7 | 99.9 | 100.2 | 100.2667
-1.3 | 95.9 | 98.9 | 99.2667 | 99.6 | 99.7 | 100 | 100.2 | 100.3
-1.2 | 97.1333 | 99.2 | 99.4 | 99.6 | 99.8 | 100 | 100.1 | 100.2
-1.1 | 98.1 | 99.3 | 99.5 | 99.6 | 99.8 | 100 | 100.1 | 100.2
-1 | 98.75 | 99.45 | 99.7 | 99.6333 | 99.8 | 100 | 100 | 100.2
-0.9 | 99.1 | 99.6 | 99.8 | 99.7 | 99.8 | 100.1 | 100 | 100.1857
-0.8 | 99.3 | 99.7 | 99.8 | 99.8 | 99.9 | 100.1 | 100 | 100.1
-0.7 | 99.4 | 99.8 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100.1
-0.6 | 99.55 | 99.8 | 99.9 | 100 | 99.88 | 100 | 100.1 | 100.1
-0.5 | 99.8 | 99.9 | 99.9 | 100 | 99.8 | 100 | 100.1 | 100.1
-0.4 | 99.9 | 100 | 99.9 | 100 | 99.8 | 100 | 100.1 | 100.1
-0.3 | 99.9 | 100 | 99.9 | 100 | 99.9 | 100 | 100.0733 | 100.1
-0.2 | 99.9 | 100 | 100 | 100 | 99.9 | 100 | 100.0067 | 100.1
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 99.9 | 100 | 100 | 100.1
0.3 | 99.9 | 100 | 99.9 | 100 | 99.9 | 100 | 100.0438 | 100.1
0.4 | 99.9 | 100 | 99.9 | 100 | 99.8833 | 100 | 100.1 | 100.1
0.5 | 99.8 | 99.9 | 99.9 | 100 | 99.8 | 100 | 100.1 | 100.1
0.6 | 99.5 | 99.8 | 99.9 | 100 | 99.8833 | 100 | 100.0222 | 100.1
0.7 | 99.4 | 99.8 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100.1
0.8 | 99.3 | 99.7 | 99.8 | 99.8 | 99.9 | 100.1 | 100 | 100.1
0.9 | 99.1 | 99.6 | 99.8 | 99.7 | 99.8 | 100.1 | 100 | 100.2
1 | 98.75 | 99.4333 | 99.7 | 99.6333 | 99.8 | 100.04 | 100 | 100.2
1.1 | 98.0333 | 99.3 | 99.5 | 99.6 | 99.8 | 100 | 100.1 | 100.2
1.2 | 97.15 | 99.2 | 99.4 | 99.6 | 99.8 | 100 | 100.1 | 100.2
1.3 | 95.9 | 98.9 | 99.2667 | 99.6 | 99.7 | 100 | 100.2 | 100.3
1.4 | 93.3 | 98.7 | 99.1 | 99.5 | 99.7 | 99.9 | 100.2 | 100.3
1.5 | 89.1 | 98.3 | 99 | 99.5 | 99.6 | 99.9 | 100.2 | 100.2143
1.6 | 81.8 | 97.9667 | 99 | 99.4 | 99.5 | 99.9 | 100.2 | 100.3
1.7 | 72.3 | 97.4667 | 98.9 | 99.3 | 99.4 | 99.9 | 100.2 | 100.3
1.8 | 61.9 | 96.8 | 98.8 | 99.2 | 99.3 | 99.9 | 100.18 | 100.2143
1.9 | 50.5667 | 95.6 | 98.6 | 99.1 | 99.25 | 99.9 | 100.1 | 100.25
2 | 38.8 | 93.6 | 98.35 | 99 | 99.2 | 99.9 | 100.1 | 100.3
2.1 | 29.2 | 90.5 | 98.125 | 98.9 | 99.2 | 99.95 | 100.1 | 100.3
2.2 | 20.8 | 85.1 | 97.9 | 98.75 | 99.1 | 100 | 100.2 | 100.4
2.3 | 15.6 | 77.375 | 97.7333 | 98.6 | 99 | 100 | 100.2 | 100.4
2.4 | 12.2 | 67.8 | 97.5 | 98.4667 | 98.9 | 99.925 | 100.2 | 100.4
2.5 | 10.2 | 57.2 | 97.2333 | 98.3 | 98.825 | 99.9 | 100.2 | 100.4
2.6 | 9 | 45.8667 | 96.95 | 98.0667 | 98.72 | 99.85 | 100.2 | 100.4
2.7 | 8.3 | 36.15 | 96.6 | 97.8 | 98.5667 | 99.8 | 100.2 | 100.4
2.8 | 7.75 | 27.1 | 96.3 | 97.5667 | 98.5 | 99.8 | 100.2 | 100.3286
2.9 | 7.3 | 20.5 | 95.75 | 97.3 | 98.35 | 99.8 | 100.2 | 100.4
3 | 7 | 16.6 | 95.1 | 97.2 | 98.22 | 99.8 | 100.2 | 100.33
3.1 | 6.7 | 13.9 | 94.1 | 97.0333 | 98.1 | 99.8 | 100.2 | 100.3
3.2 | 6.5333 | 12.3 | 93.2 | 96.8 | 97.875 | 99.6 | 100.1571 | 100.3
3.3 | 6.3 | 11.25 | 91.4 | 96.55 | 97.625 | 99.6 | 100.1 | 100.3625
3.4 | 6.075 | 10.6 | 88.1333 | 96.35 | 97.475 | 99.6 | 100.1 | 100.4
3.5 | 5.8 | 10.1 | 83.6 | 96.1333 | 97.3 | 99.55 | 100.1 | 100.5
3.6 | 5.7 | 9.65 | 76.2 | 95.85 | 97.18 | 99.5 | 100.1 | 100.5
3.7 | 5.5333 | 9.4 | 67.65 | 95.6333 | 96.95 | 99.475 | 100.0714 | 100.5
3.8 | 5.4 | 9.05 | 57.5 | 95.1 | 96.8 | 99.325 | 100 | 100.5
3.9 | 5.3 | 8.75 | 48.1 | 94.8 | 96.65 | 99.3 | 100 | 100.4467
4 | 5.15 | 8.45 | 39.2 | 94.4 | 96.45 | 99.3 | 100 | 100.4
4.1 | 5 | 8.3 | 30.75 | 93.8667 | 96.2 | 99.2 | 100 | 100.4
4.2 | 4.8 | 8.0333 | 24.7 | 93.4333 | 95.875 | 99.1 | 100 | 100.4
4.3 | 4.7 | 7.8 | 20.5 | 92.9 | 95.7 | 99.1 | 100 | 100.4
4.4 | 4.6 | 7.55 | 18 | 92.1 | 95.4286 | 99.1 | 100 | 100.4
4.5 | 4.4667 | 7.4 | 16.7 | 91 | 95.18 | 99 | 100 | 100.5
4.6 | 4.3 | 7.2 | 15.6 | 89.1333 | 94.84 | 99 | 99.975 | 100.5
4.7 | 4.25 | 6.9 | 14.9 | 86.65 | 94.45 | 98.9 | 99.9 | 100.5
4.8 | 4.1 | 6.7333 | 14.25 | 82.25 | 94.2 | 98.8 | 99.9 | 100.5
4.9 | 4.1 | 6.6 | 13.7 | 76.0667 | 93.85 | 98.7333 | 99.9 | 100.5
5 | 4 | 6.5 | 13.45 | 67.95 | 93.6 | 98.7 | 99.8 | 100.5
5.1 | 3.9 | 6.3 | 13 | 59.7 | 93.25 | 98.6 | 99.7857 | 100.4
5.2 | 3.8 | 6.2 | 12.7 | 50.1 | 92.85 | 98.5 | 99.7 | 100.4
5.3 | 3.7 | 6 | 12.4 | 41.4333 | 92.4 | 98.4333 | 99.7 | 100.4
5.4 | 3.6 | 5.9 | 12.1 | 34.1667 | 91.85 | 98.3 | 99.7 | 100.4286
5.5 | 3.55 | 5.8 | 11.8 | 28.1 | 91.5 | 98.125 | 99.7 | 100.5
5.6 | 3.4 | 5.65 | 11.4 | 24 | 90.8 | 98.1 | 99.7 | 100.5
5.7 | 3.4 | 5.5 | 11.2 | 21.6 | 90.2 | 98.1 | 99.6 | 100.5
5.8 | 3.3 | 5.4 | 10.9 | 20 | 89.2 | 98 | 99.6 | 100.5
5.9 | 3.2 | 5.2 | 10.8 | 19 | 87.85 | 97.95 | 99.6 | 100.4857
6 | 3.2 | 5.1 | 10.4 | 18.15 | 85.3 | 97.8 | 99.6 | 100.4143
6.1 | 3.1 | 5.1 | 10.1 | 17.6 | 81.5 | 97.65 | 99.56 | 100.4
6.2 | 3 | 4.95 | 9.9 | 17 | 76 | 97.6 | 99.4833 | 100.4
6.3 | 3 | 4.8 | 9.75 | 16.7 | 68.65 | 97.5 | 99.4 | 100.3231
6.4 | 3 | 4.7 | 9.5 | 16.1667 | 60.3 | 97.3 | 99.3 | 100.3
6.5 | 2.9 | 4.6 | 9.3 | 15.75 | 51.9333 | 97.2 | 99.3 | 100.3
6.6 | 2.9 | 4.5 | 9.1 | 15.4333 | 43.76 | 97.1 | 99.2286 | 100.3
6.7 | 2.8 | 4.4 | 8.9333 | 15.1 | 36.68 | 96.975 | 99.2 | 100.3
6.8 | 2.7 | 4.35 | 8.7 | 14.75 | 30.6667 | 96.8571 | 99.2 | 100.2429
6.9 |  | 4.3 | 8.5 | 14.4 | 26.7 | 96.8 | 99.1 | 100.2
7 |  | 4.2 | 8.4 | 14.0667 | 24.175 | 96.5833 | 99.1 | 100.2
7.1 |  | 4.1 | 8.2 | 13.8 | 22.525 | 96.4 | 99.0375 | 100.2
7.2 |  | 4 | 8.1 | 13.5 | 21.425 | 96.3 | 99 | 100.1
7.3 |  | 3.95 | 7.8 | 13.2 | 20.625 | 96.15 | 98.9 | 100.1
7.4 |  | 3.8 | 7.7 | 12.9 | 19.925 | 95.96 | 98.8 | 100.0286
7.5 |  | 3.7 | 7.6 | 12.7 | 19.375 | 95.825 | 98.8 | 100
7.6 |  |  | 7.4 | 12.4 | 18.9 | 95.6 | 98.8 | 100
7.7 |  |  | 7.3 | 12.1333 | 18.5 | 95.5 | 98.7 | 100
7.8 |  |  | 7.1 | 11.9 | 18.075 | 95.3 | 98.5714 | 100
7.9 |  |  | 6.9 | 11.6333 | 17.6667 | 95.2 | 98.5 | 100
8 |  |  | 6.8333 | 11.38 | 17.3333 | 95 | 98.4375 | 99.9
8.1 |  |  | 6.7 | 11.2 | 17 | 94.7667 | 98.4 | 99.9
8.2 |  |  | 6.6333 | 11 | 16.68 | 94.55 | 98.36 | 99.8
8.3 |  |  | 6.5 | 10.7 | 16.28 | 94.45 | 98.2267 | 99.7333
8.4 |  |  | 6.4 | 10.5667 | 16 | 94.3 | 98.1111 | 99.7
8.5 |  |  | 6.3 | 10.3 | 15.65 | 94 | 98 | 99.7
8.6 |  |  | 6.1 | 10.15 | 15.4 | 93.9 | 97.9923 | 99.6
8.7 |  |  | 6 | 9.9 | 15.05 | 93.76 | 97.9154 | 99.6
8.8 |  |  | 5.9 | 9.7 | 14.82 | 93.45 | 97.8333 | 99.5
8.9 |  |  |  | 9.5 | 14.55 | 93.275 | 97.7538 | 99.5
9 |  |  |  | 9.4 | 14.225 | 93.1 | 97.66 | 99.4375
9.1 |  |  |  | 9.2333 | 13.975 | 93 | 97.5267 | 99.4
9.2 |  |  |  | 9.1 | 13.7 | 92.7667 | 97.4429 | 99.3714
9.3 |  |  |  | 8.9 | 13.4 | 92.4 | 97.35 | 99.3
9.4 |  |  |  | 8.8 | 13.15 | 92.2 | 97.2143 | 99.3
9.5 |  |  |  | 8.6 | 12.9 | 91.95 | 97.1875 | 99.2714
9.6 |  |  |  | 8.4 | 12.75 | 91.7 | 97.1 | 99.2
9.7 |  |  |  | 8.3 | 12.4 | 91.5 | 97.05 | 99.1
9.8 |  |  |  | 8.15 | 12.2 | 91.25 | 96.925 | 99.1
9.9 |  |  |  | 8 | 12 | 90.9 | 96.8111 | 99.0375
10 |  |  |  | 7.9 | 11.85 | 90.6 | 96.7857 | 99
10.1 |  |  |  | 7.8 | 11.64 | 90.3 | 96.65 | 98.9714
10.2 |  |  |  |  | 11.4 | 89.975 | 96.525 | 98.9
10.3 |  |  |  |  | 11.225 | 89.7 | 96.3875 | 98.8125
10.4 |  |  |  |  | 10.975 | 89.5 | 96.3 | 98.8
10.5 |  |  |  |  | 10.8 | 89.2667 | 96.2286 | 98.7375
10.6 |  |  |  |  | 10.625 | 88.85 | 96.1 | 98.7
10.7 |  |  |  |  | 10.475 | 88.5333 | 96 | 98.6
10.8 |  |  |  |  | 10.3 | 88.2 | 95.925 | 98.6
10.9 |  |  |  |  | 10.175 | 87.9 | 95.7875 | 98.5143
11 |  |  |  |  | 10.0333 | 87.5 | 95.7 | 98.4857
11.1 |  |  |  |  | 9.88 | 87.1667 | 95.5286 | 98.4
11.2 |  |  |  |  | 9.7 | 86.925 | 95.375 | 98.3143
11.3 |  |  |  |  | 9.575 | 86.475 | 95.26 | 98.2
11.4 |  |  |  |  | 9.4 | 86.125 | 95.0857 | 98.2
11.5 | [5 empty] | 85.75 | 94.9125 | 98.2
11.6 | [5 empty] | 85.425 | 94.78 | 98.1
11.7 | [5 empty] | 85.05 | 94.6778 | 98.025
11.8 | [5 empty] | 84.4667 | 94.5571 | 97.9
11.9 | [5 empty] | 84 | 94.425 | 97.8571
12 | [5 empty] | 83.4 | 94.2875 | 97.8
12.1 | [5 empty] | 82.95 | 94.1571 | 97.8
12.2 | [5 empty] | 82.2333 | 94.1 | 97.7571
12.3 | [5 empty] | 81.325 | 93.975 | 97.7
12.4 | [5 empty] | 80.2 | 93.7571 | 97.6
12.5 | [5 empty] | 78.4 | 93.625 | 97.5
12.6 | [5 empty] | 75.34 | 93.4875 | 97.4375
12.7 | [5 empty] | 71.5 | 93.3143 | 97.38
12.8 | [5 empty] | 65.5333 | 93.125 | 97.3
12.9 | [5 empty] | 59 | 92.9875 | 97.2857
13 | [5 empty] | 52.4 | 92.8625 | 97.2
13.1 | [5 empty] | 45.6667 | 92.7286 | 97.125
13.2 | [5 empty] | 39.9 | 92.5 | 96.9
13.3 | [5 empty] | 35.15 | 92.3667 | 96.8333
13.4 | [5 empty] | 31.6667 | 92.2333 | 96.8
13.5 | [5 empty] | 29.55 | 92.1 | 96.7625
13.6 | [5 empty] | 28.0667 | 91.8429 | 96.7
13.7 | [5 empty] | 27.14 | 91.6 | 96.6
13.8 | [5 empty] | 26.375 | 91.4429 | 96.4625
13.9 | [5 empty] | 25.85 | 91.2375 | 96.3286
14 | [5 empty] | 25.3333 | 91.2 | 96.2
14.1 | [5 empty] | 24.8667 | 91.05 | 96.1625
14.2 | [5 empty] | 24.45 | 90.8429 | 96.1
14.3 | [5 empty] | 24.025 | 90.625 | 96
14.4 | [5 empty] | 23.6333 | 90.475 | 95.9
14.5 | [5 empty] | 23.175 | 90.2857 | 95.8286
14.6 | [5 empty] | 22.9 | 90.025 | 95.7
14.7 | [5 empty] | 22.55 | 89.8333 | 95.6625
14.8 | [5 empty] | 22.15 | 89.58 | 95.55
14.9 | [5 empty] | 21.78 | 89.3286 | 95.4286
15 | [5 empty] | 21.375 | 89.1714 | 95.2545
15.1 | [5 empty] | 21.125 | 88.9 | 95.1533
15.2 | [5 empty] | 20.8 | 88.725 | 95.0733
15.3 | [5 empty] | 20.475 | 88.4875 | 94.94
15.4 | [5 empty] | 20.2 | 88.2714 | 94.8462
15.5 | [5 empty] | 19.86 | 87.95 | 94.75
15.6 | [5 empty] | 19.5667 | 87.75 | 94.6143
15.7 | [5 empty] | 19.3 | 87.5667 | 94.5
15.8 | [5 empty] | 19.0333 | 87.3091 | 94.4
15.9 | [5 empty] | 18.7333 | 87.0667 | 94.3286
16 | [5 empty] | 18.5 | 86.86 | 94.2
16.1 | [5 empty] | 18.1333 | 86.625 | 94.0333
16.2 | [5 empty] | 17.9 | 86.3 | 93.9333
16.3 | [5 empty] | 17.6 | 85.9714 | 93.84
16.4 | [5 empty] | 17.4 | 85.775 | 93.7333
16.5 | [5 empty] | 17.18 | 85.4 | 93.7
16.6 | [5 empty] | 17 | 85.1714 | 93.5889
16.7 | [5 empty] | 16.6667 | 85.0375 | 93.48
16.8 | [5 empty] | 16.4 | 84.7 | 93.375
16.9 | [5 empty] | 16.15 | 84.4333 | 93.1857
17 | [5 empty] | 16 | 84.125 | 93.0125
17.1 | [5 empty] | 15.7 | 83.8125 | 92.8875
17.2 | [5 empty] | 15.4667 | 83.4 | 92.7571
17.3 | [5 empty] | 15.3 | 83.2429 | 92.7
17.4 | [5 empty] | 15.1 | 82.975 | 92.5875
17.5 | [5 empty] | 14.8667 | 82.6 | 92.5
17.6 | [5 empty] | 14.7 | 82.2625 | 92.3875
17.7 | [5 empty] | 14.5 | 82.0125 | 92.2571
17.8 | [5 empty] | 14.35 | 81.5 | 92.1143
17.9 | [5 empty] | 14.15 | 81.1143 | 91.9571
18 | [6 empty] | 80.775 | 91.7833
18.1 | [6 empty] | 80.4 | 91.6
18.2 | [6 empty] | 80.05 | 91.4571
18.3 | [6 empty] | 79.7286 | 91.3
18.4 | [6 empty] | 79.25 | 91.1333
18.5 | [6 empty] | 78.7 | 91
18.6 | [6 empty] | 78.1143 | 90.8625
18.7 | [6 empty] | 77.4625 | 90.6571
18.8 | [6 empty] | 76.85 | 90.5
18.9 | [6 empty] | 75.95 | 90.45
19 | [6 empty] | 74.4714 | 90.275
19.1 | [6 empty] | 72.125 | 90.1429
19.2 | [6 empty] | 68.9143 | 89.925
19.3 | [6 empty] | 64.475 | 89.75
19.4 | [6 empty] | 59.15 | 89.4857
19.5 | [6 empty] | 53.4667 | 89.3125
19.6 | [6 empty] | 47.8909 | 89.15
19.7 | [6 empty] | 42.4545 | 88.95
19.8 | [6 empty] | 38 | 88.8143
19.9 | [6 empty] | 34.25 | 88.6714
20 | [6 empty] | 31.9143 | 88.45
20.1 | [6 empty] | 30.4667 | 88.25
20.2 | [6 empty] | 29.5375 | 87.9875
20.3 | [6 empty] | 28.875 | 87.7714
20.4 | [6 empty] | 28.2 | 87.525
20.5 | [6 empty] | 27.6 | 87.375
20.6 | [6 empty] | 27.2 | 87.0714
20.7 | [6 empty] | 26.7917 | 86.825
20.8 | [6 empty] | 26.4 | 86.7
20.9 | [6 empty] | 26.08 | 86.4429
21 | [6 empty] | 25.74 | 86.2375
21.1 | [6 empty] | 25.3286 | 86
21.2 | [6 empty] | 25.025 | 85.6714
21.3 | [6 empty] | 24.625 | 85.5
21.4 | [6 empty] | 24.2857 | 85.3375
21.5 | [6 empty] | 23.9375 | 85.1
21.6 | [6 empty] | 23.675 | 84.8714
21.7 | [6 empty] | 23.44 | 84.675
21.8 | [6 empty] | 23.1 | 84.4
21.9 | [6 empty] | 22.7429 | 84.2
22 | [6 empty] | 22.4571 | 83.8667
22.1 | [6 empty] | 22.1429 | 83.6
22.2 | [6 empty] | 21.925 | 83.3667
22.3 | [6 empty] | 21.575 | 83.075
22.4 | [6 empty] | 21.3143 | 82.9125
22.5 | [6 empty] | 21.0286 | 82.5778
22.6 | [6 empty] | 20.8714 | 82.35
22.7 | [6 empty] | 20.6 | 82.1
22.8 | [6 empty] | 20.3286 | 81.8286
22.9 | [6 empty] | 19.9556 | 81.575
23 | [6 empty] | 19.7571 | 81.2667
23.1 | [6 empty] | 19.55 | 80.92
23.2 | [6 empty] | 19.3625 | 80.54
23.3 | [6 empty] | 19.0143 | 80.2467
23.4 | [6 empty] | 18.825 | 79.98
23.5 | [6 empty] | 18.5714 | 79.725
23.6 | [6 empty] | 18.35 | 79.475
23.7 | [6 empty] | 18.1 | 79.2091
23.8 | [6 empty] | 18 | 78.8429
23.9 | [6 empty] | 17.8 | 78.4714
24 | [6 empty] | 17.55 | 78.1667
24.1 | [6 empty] | 17.425 | 77.8571
24.2 | [6 empty] | 17.1667 | 77.5091
24.3 | [6 empty] | 16.9286 | 77.15
24.4 | [6 empty] | 16.6875 | 76.7875
24.5 | [6 empty] | 16.525 | 76.4
24.6 | [6 empty] | 16.3286 | 76.0571
24.7 | [6 empty] | 16.2 | 75.75
24.8 | [6 empty] | 15.9714 | 75.25
24.9 | [6 empty] | 15.8375 | 74.7857
25 | [6 empty] | 15.6 | 74.25
25.1 | [6 empty] | 15.475 | 73.7714
25.2 | [6 empty] | 15.2571 | 73.225
25.3 | [6 empty] | 15.125 | 72.5125
25.4 | [6 empty] | 14.8667 | 71.5625
25.5 | [6 empty] | 14.7143 | 70.0714
25.6 | [6 empty] | 14.55 | 67.9
25.7 | [6 empty] | 14.4 | 64.6125
25.8 | [6 empty] | 14.2571 | 60.4429
25.9 | [6 empty] | 14.0667 | 55.5
26 | [6 empty] | 13.9 | 50.2
26.1 | [6 empty] | 13.7714 | 45.2533
26.2 | [6 empty] | 13.6286 | 40.7867
26.3 | [6 empty] | 13.4714 | 36.8
26.4 | [6 empty] | 13.4 | 33.6
26.5 | [6 empty] | 13.175 | 31.77
26.6 | [6 empty] | 13 | 30.4333
26.7 | [6 empty] | 12.9 | 29.6143
26.8 | [6 empty] | 12.7429 | 28.98
26.9 | [6 empty] | 12.6125 | 28.45
27 | [6 empty] | 12.3857 | 27.95
27.1 | [6 empty] | 12.2429 | 27.54
27.2 | [6 empty] | 12.15 | 27.2
27.3 | [6 empty] | 12.0125 | 26.825
27.4 | [6 empty] | 11.9 | 26.4667
27.5 | [7 empty] | 26.0727
27.6 | [7 empty] | 25.7429
27.7 | [7 empty] | 25.4375
27.8 | [7 empty] | 25.1571
27.9 | [7 empty] | 24.8
28 | [7 empty] | 24.475
28.1 | [7 empty] | 24.1625
28.2 | [7 empty] | 23.8571
28.3 | [7 empty] | 23.575
28.4 | [7 empty] | 23.3222
28.5 | [7 empty] | 22.975
28.6 | [7 empty] | 22.71
28.7 | [7 empty] | 22.4357
28.8 | [7 empty] | 22.2214
28.9 | [7 empty] | 21.9
29 | [7 empty] | 21.6385
29.1 | [7 empty] | 21.3933
29.2 | [7 empty] | 21.1267
29.3 | [7 empty] | 20.92
29.4 | [7 empty] | 20.7111
29.5 | [7 empty] | 20.4636
29.6 | [7 empty] | 20.2111
29.7 | [7 empty] | 19.9636

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-42.1 |  |  | 2.6 | 4.9 | 8.5
-42 |  |  | 2.6 | 4.9 | 8.5358
-41.9 |  |  | 2.6 | 4.987 | 8.6645
-41.8 |  |  | 2.6 | 5.0899 | 8.7059
-41.7 |  |  | 2.6337 | 5.1 | 8.8
-41.6 |  |  | 2.7 | 5.1205 | 8.8859
-41.5 |  |  | 2.7 | 5.2 | 9
-41.4 |  |  | 2.7 | 5.2225 | 9.0894
-41.3 |  |  | 2.7235 | 5.3 | 9.1108
-41.2 |  |  | 2.8 | 5.3103 | 9.2287
-41.1 |  |  | 2.8 | 5.4 | 9.3
-41 |  |  | 2.8 | 5.4 | 9.3973
-40.9 |  | 1.9 | 2.8 | 5.4133 | 9.4483
-40.8 |  | 1.9429 | 2.8167 | 5.5126 | 9.5001
-40.7 |  | 2 | 2.9 | 5.6009 | 9.6519
-40.6 |  | 2 | 2.9 | 5.6794 | 9.8
-40.5 |  | 2 | 2.9 | 5.7 | 9.8
-40.4 |  | 2 | 2.9857 | 5.7058 | 9.9058
-40.3 |  | 2 | 3 | 5.8 | 10
-40.2 |  | 2 | 3 | 5.8608 | 10.1086
-40.1 |  | 2 | 3 | 5.9 | 10.2
-40 |  | 2.0822 | 3.051 | 5.9929 | 10.2915
-39.9 | 1.5129 | 2.1 | 3.1 | 6.0949 | 10.4
-39.8 | 1.6 | 2.1 | 3.1 | 6.1245 | 10.4943
-39.7 | 1.6 | 2.1 | 3.154 | 6.2 | 10.6
-39.6 | 1.6 | 2.1 | 3.2 | 6.2266 | 10.6972
-39.5 | 1.6 | 2.1357 | 3.2 | 6.3 | 10.8
-39.4 | 1.6 | 2.2 | 3.2 | 6.3286 | 10.8799
-39.3 | 1.6 | 2.2 | 3.2438 | 6.4259 | 11
-39.2 | 1.6 | 2.2 | 3.3 | 6.5 | 11.0829
-39.1 | 1.6202 | 2.2 | 3.3 | 6.5173 | 11.2369
-39 | 1.7 | 2.204 | 3.3469 | 6.6184 | 11.3656
-38.9 | 1.7 | 2.3 | 3.4 | 6.7194 | 11.5059
-38.8 | 1.7 | 2.3 | 3.4 | 6.8 | 11.6
-38.7 | 1.7 | 2.3 | 3.4187 | 6.8214 | 11.6875
-38.6 | 1.7 | 2.3 | 3.5 | 6.9224 | 11.8314
-38.5 | 1.7 | 2.3 | 3.5 | 7 | 11.9607
-38.4 | 1.7 | 2.3635 | 3.5959 | 7.0101 | 12.1141
-38.3 | 1.7797 | 2.4 | 3.6 | 7.1111 | 12.2397
-38.2 | 1.8 | 2.4 | 3.6 | 7.2122 | 12.357
-38.1 | 1.8 | 2.4 | 3.6846 | 7.3132 | 12.4703
-38 | 1.8 | 2.4571 | 3.7 | 7.4 | 12.6426
-37.9 | 1.8 | 2.5 | 3.7 | 7.4153 | 12.7
-37.8 | 1.8 | 2.5 | 3.7877 | 7.5163 | 12.7877
-37.7 | 1.8 | 2.5 | 3.8 | 7.6173 | 12.9774
-37.6 | 1.8 | 2.5 | 3.8325 | 7.7183 | 13.0897
-37.5 | 1.8907 | 2.567 | 3.9 | 7.8169 | 13.2814
-37.4 | 1.9 | 2.6 | 3.9302 | 7.906 | 13.3917
-37.3 | 1.9 | 2.6 | 4 | 8.0071 | 13.6498
-37.2 | 1.9 | 2.6 | 4 | 8.1081 | 13.7912
-37.1 | 1.9 | 2.6947 | 4.0233 | 8.2091 | 13.9272
-37 | 1.9362 | 2.7 | 4.1 | 8.3101 | 14.0387
-36.9 | 1.9906 | 2.7 | 4.1694 | 8.4111 | 14.1926
-36.8 | 2 | 2.7 | 4.2 | 8.5121 | 14.3549
-36.7 | 2 | 2.7865 | 4.273 | 8.6115 | 14.5567
-36.6 | 2 | 2.8 | 4.3198 | 8.6998 | 14.7711
-36.5 | 2 | 2.8 | 4.4 | 8.8009 | 14.8758
-36.4 | 2 | 2.9 | 4.4188 | 8.9019 | 15.0467
-36.3 | 2.0743 | 2.9 | 4.5 | 9.0041 | 15.2487
-36.2 | 2.1 | 2.9 | 4.5 | 9.1455 | 15.5039
-36.1 | 2.1 | 2.9168 | 4.6 | 9.3058 | 15.6148
-36 | 2.1 | 3 | 4.6236 | 9.4 | 15.906
-35.9 | 2.1 | 3 | 4.7 | 9.4697 | 16.0139
-35.8 | 2.1652 | 3.008 | 4.7593 | 9.6112 | 16.2186
-35.7 | 2.2 | 3.1 | 4.8 | 9.7657 | 16.4465
-35.6 | 2.2 | 3.1 | 4.8672 | 9.814 | 16.6485
-35.5 | 2.2 | 3.1128 | 4.9596 | 9.9693 | 16.8505
-35.4 | 2.2 | 3.2 | 5.0384 | 10.1168 | 17.0613
-35.3 | 2.2702 | 3.2 | 5.1 | 10.2728 | 17.2832
-35.2 | 2.3 | 3.2284 | 5.114 | 10.3998 | 17.4853
-35.1 | 2.3 | 3.3 | 5.2132 | 10.5016 | 17.7309
-35 | 2.3 | 3.3087 | 5.3016 | 10.7016 | 17.9893
-34.9 | 2.3314 | 3.3631 | 5.3899 | 10.7899 | 18.1913
-34.8 | 2.4 | 3.4162 | 5.4 | 11.0034 | 18.3933
-34.7 | 2.4 | 3.4667 | 5.4792 | 11.0918 | 18.643
-34.6 | 2.4 | 3.5 | 5.5773 | 11.2832 | 18.8974
-34.5 | 2.4124 | 3.5 | 5.6783 | 11.3926 | 19.116
-34.4 | 2.4713 | 3.5694 | 5.7 | 11.5873 | 19.4276
-34.3 | 2.5 | 3.6 | 5.7803 | 11.6946 | 19.7929
-34.2 | 2.5242 | 3.6522 | 5.8813 | 11.8913 | 20.1848
-34.1 | 2.6 | 3.7 | 5.9823 | 12.0933 | 20.5383
-34 | 2.6 | 3.7405 | 6.0834 | 12.2954 | 20.8536
-33.9 | 2.6273 | 3.8415 | 6.1844 | 12.3987 | 21.2453
-33.8 | 2.7 | 3.9 | 6.2195 | 12.5994 | 21.6979
-33.7 | 2.7 | 3.9435 | 6.361 | 12.7763 | 22.3612
-33.6 | 2.7303 | 4 | 6.403 | 12.9749 | 23.0122
-33.5 | 2.8 | 4.0456 | 6.5 | 13.177 | 23.7223
-33.4 | 2.8 | 4.1 | 6.6052 | 13.379 | 24.5446
-33.3 | 2.8333 | 4.1477 | 6.7467 | 13.581 | 25.3909
-33.2 | 2.9 | 4.2 | 6.8 | 13.783 | 26.4613
-33.1 | 2.9 | 4.2497 | 6.8925 | 13.9851 | 27.8378
-33 | 2.9 | 4.3507 | 6.9935 | 14.1871 | 29.4982
-32.9 | 2.9945 | 4.4362 | 7.1123 | 14.4483 | 31.3585
-32.8 | 3 | 4.5 | 7.2384 | 14.6626 | 33.5212
-32.7 | 3.0966 | 4.5 | 7.3394 | 14.8646 | 36.103
-32.6 | 3.1 | 4.6138 | 7.4809 | 15.15 | 38.9792
-32.5 | 3.1986 | 4.7173 | 7.6415 | 15.3687 | 41.917
-32.4 | 3.2 | 4.7816 | 7.7425 | 15.6561 | 44.8463
-32.3 | 3.3 | 4.836 | 7.8435 | 15.8727 | 47.7178
-32.2 | 3.3 | 4.8866 | 7.9891 | 16.1057 | 50.3413
-32.1 | 3.3577 | 4.9741 | 8.1456 | 16.4276 | 52.4506
-32 | 3.4 | 5.0752 | 8.2407 | 16.7157 | 54.1521
-31.9 | 3.4476 | 5.1762 | 8.3776 | 17.0582 | 55.3664
-31.8 | 3.5 | 5.2772 | 8.5567 | 17.3242 | 56.4401
-31.7 | 3.5496 | 5.3391 | 8.6745 | 17.7237 | 57.2167
-31.6 | 3.6 | 5.3896 | 8.8109 | 18.0327 | 57.8752
-31.5 | 3.6516 | 5.5021 | 8.9654 | 18.457 | 58.5614
-31.4 | 3.7526 | 5.6168 | 9.1137 | 18.9549 | 59.1873
-31.3 | 3.8 | 5.6876 | 9.2689 | 19.4655 | 59.8756
-31.2 | 3.8404 | 5.7729 | 9.4165 | 19.9864 | 60.4162
-31.1 | 3.9 | 5.87 | 9.558 | 20.69 | 61.0897
-31 | 3.9424 | 5.971 | 9.8194 | 21.3972 | 61.7224
-30.9 | 4.0434 | 6.072 | 10.0013 | 22.3574 | 62.4039
-30.8 | 4.1222 | 6.173 | 10.1685 | 23.4181 | 63.0812
-30.7 | 4.1727 | 6.274 | 10.3481 | 24.5934 | 63.6873
-30.6 | 4.2465 | 6.3875 | 10.4751 | 26.2611 | 64.3254
-30.5 | 4.3475 | 6.4703 | 10.7282 | 28.0515 | 65.0322
-30.4 | 4.4242 | 6.5942 | 10.9541 | 30.5505 | 65.7194
-30.3 | 4.4747 | 6.7458 | 11.1562 | 33.3277 | 66.3255
-30.2 | 4.5471 | 6.8907 | 11.3582 | 36.4973 | 66.9534
-30.1 | 4.6414 | 7.023 | 11.6403 | 39.9163 | 67.6284
-30 | 4.7191 | 7.124 | 11.8622 | 43.941 | 68.1942
-29.9 | 4.7696 | 7.2251 | 12.0642 | 47.764 | 68.7794
-29.8 | 4.8282 | 7.3405 | 12.3494 | 51.2266 | 69.3848
-29.7 | 4.8989 | 7.4976 | 12.5683 | 54.5093 | 69.9137
-29.6 | 4.9995 | 7.6741 | 12.8555 | 57.0033 | 70.3973
-29.5 | 5.1005 | 7.8291 | 13.1413 | 59.0382 | 70.9455
-29.4 | 5.2015 | 7.9602 | 13.4554 | 60.6216 | 71.4505
-29.3 | 5.3025 | 8.1311 | 13.747 | 61.797 | 71.8933
-29.2 | 5.4035 | 8.2749 | 14.0625 | 62.8071 | 72.2285
-29.1 | 5.5045 | 8.4947 | 14.3709 | 63.632 | 72.6657
-29 | 5.6056 | 8.6969 | 14.7928 | 64.4018 | 73.1366
-28.9 | 5.7066 | 8.8495 | 15.2642 | 65.2099 | 73.5055
-28.8 | 5.8076 | 9.0325 | 15.765 | 66.018 | 73.8446
-28.7 | 5.9086 | 9.2601 | 16.3602 | 66.7606 | 74.2487
-28.6 | 6.0096 | 9.4233 | 17.1055 | 67.5489 | 74.6003
-28.5 | 6.1083 | 9.6071 | 17.9527 | 68.4661 | 74.8748
-28.4 | 6.1869 | 9.8636 | 18.9422 | 69.293 | 75.1779
-28.3 | 6.3177 | 10.1111 | 20.177 | 70.1145 | 75.4809
-28.2 | 6.4538 | 10.3484 | 21.7178 | 70.996 | 75.7839
-28.1 | 6.5824 | 10.5866 | 23.6088 | 71.8545 | 76.058
-28 | 6.6939 | 10.8551 | 26.0416 | 72.6103 | 76.29
-27.9 | 6.8042 | 11.1392 | 29.0967 | 73.3175 | 76.6086
-27.8 | 6.9811 | 11.4622 | 32.7382 | 74.1245 | 76.9207
-27.7 | 7.0902 | 11.8158 | 36.8884 | 74.8317 | 77.0991
-27.6 | 7.2824 | 12.3039 | 41.7257 | 75.5388 | 77.3681
-27.5 | 7.3922 | 12.8564 | 47.2799 | 76.2459 | 77.5701
-27.4 | 7.5865 | 13.4866 | 52.8909 | 76.8836 | 77.8082
-27.3 | 7.78 | 14.3507 | 57.8859 | 77.4283 | 78.0973
-27.2 | 7.9594 | 15.2889 | 62.1667 | 78.0667 | 78.3477
-27.1 | 8.1212 | 16.4121 | 65.6657 | 78.7061 | 78.5745
-27 | 8.3232 | 17.8313 | 68.0828 | 79.233 | 78.8603
-26.9 | 8.5379 | 19.7658 | 69.9427 | 79.7226 | 79.096
-26.8 | 8.8273 | 22.241 | 71.4907 | 80.1402 | 79.3086
-26.7 | 9.0586 | 25.37 | 72.5436 | 80.5442 | 79.4894
-26.6 | 9.4783 | 29.286 | 73.6417 | 80.9483 | 79.6785
-26.5 | 10 | 34.1658 | 74.6634 | 81.3523 | 79.935
-26.4 | 10.6237 | 40.0658 | 75.6535 | 81.6771 | 80.1472
-26.3 | 11.4093 | 46.4549 | 76.8436 | 82.0461 | 80.3517
-26.2 | 12.5409 | 52.9965 | 77.8176 | 82.4504 | 80.5537
-26.1 | 13.9632 | 58.9391 | 78.8422 | 82.7909 | 80.7558
-26 | 15.4719 | 65.2234 | 79.9533 | 83.0939 | 81.0367
-25.9 | 17.7619 | 69.5936 | 81.0644 | 83.397 | 81.2598
-25.8 | 20.4726 | 73.1656 | 82.1756 | 83.7 | 81.4618
-25.7 | 24.4505 | 75.6427 | 83.2867 | 84.003 | 81.6638
-25.6 | 29.4636 | 77.4079 | 84.3978 | 84.2602 | 81.8659
-25.5 | 35.6935 | 78.7462 | 85.3982 | 84.398 | 82.096
-25.4 | 43.6881 | 79.9485 | 86.3073 | 84.5975 | 82.318
-25.3 | 51.1345 | 81.0889 | 87.2169 | 84.9008 | 82.5004
-25.2 | 58.8757 | 82.2276 | 88.0774 | 85.1419 | 82.6523
-25.1 | 65.8337 | 83.4192 | 88.8899 | 85.3664 | 82.8065
-25 | 71.6467 | 84.6467 | 89.6187 | 85.6352 | 83.0894
-24.9 | 76.283 | 86.0614 | 90.288 | 85.8372 | 83.3061
-24.8 | 79.5951 | 87.2359 | 90.8312 | 86.0458 | 83.4475
-24.7 | 81.7461 | 88.5111 | 91.4158 | 86.3048 | 83.589
-24.6 | 83.2783 | 89.7455 | 91.9639 | 86.536 | 83.7504
-24.5 | 84.508 | 90.9576 | 92.4273 | 86.7109 | 83.9836
-24.4 | 85.6322 | 92.0081 | 92.9195 | 86.938 | 84.2064
-24.3 | 86.9456 | 93.3426 | 93.2693 | 87.078 | 84.3946
-24.2 | 88.1949 | 94.3703 | 93.5901 | 87.32 | 84.7451
-24.1 | 89.4912 | 95.2847 | 93.9784 | 87.582 | 84.8813
-24 | 90.8893 | 96.2919 | 94.3824 | 87.784 | 85.0699
-23.9 | 92.3033 | 97.1016 | 94.7006 | 87.9861 | 85.186
-23.8 | 93.6429 | 97.9064 | 94.9695 | 88.1385 | 85.3739
-23.7 | 94.9311 | 98.4369 | 95.2353 | 88.2719 | 85.5539
-23.6 | 96.3455 | 99.1129 | 95.4923 | 88.4807 | 85.6871
-23.5 | 97.637 | 99.5255 | 95.7415 | 88.7433 | 85.8574
-23.4 | 98.8491 | 100.121 | 95.9964 | 88.9474 | 85.9767
-23.3 | 99.9978 | 100.5251 | 96.143 | 89.1088 | 86.1556
-23.2 | 101.0771 | 100.9291 | 96.2717 | 89.2419 | 86.417
-23.1 | 102.0182 | 101.3249 | 96.4737 | 89.3597 | 86.6317
-23 | 102.8465 | 101.6263 | 96.6379 | 89.4931 | 86.7731
-22.9 | 103.5624 | 101.8545 | 96.7778 | 89.6921 | 86.9243
-22.8 | 104.1699 | 102.0918 | 96.9559 | 89.8941 | 87.1513
-22.7 | 104.7883 | 102.3573 | 97.1623 | 89.9981 | 87.3267
-22.6 | 105.2337 | 102.4987 | 97.2989 | 90.1982 | 87.4277
-22.5 | 105.6578 | 102.6402 | 97.5001 | 90.4002 | 87.5574
-22.4 | 106.0464 | 102.8361 | 97.6297 | 90.6022 | 87.7297
-22.3 | 106.3329 | 102.9736 | 97.7614 | 90.8021 | 87.8269
-22.2 | 106.5349 | 103.0746 | 97.9317 | 90.9065 | 87.9349
-22.1 | 106.7518 | 103.2323 | 98.0327 | 91.1074 | 88.1185
-22 | 106.9962 | 103.3436 | 98.1338 | 91.2841 | 88.2585
-21.9 | 107.1491 | 103.4179 | 98.2406 | 91.4838 | 88.5205
-21.8 | 107.2501 | 103.5064 | 98.3501 | 91.5376 | 88.643
-21.7 | 107.3795 | 103.5947 | 98.4716 | 91.6915 | 88.8451
-21.6 | 107.5 | 103.6949 | 98.5807 | 91.8659 | 89.0412
-21.5 | 107.5 | 103.7 | 98.7634 | 92.062 | 89.2103
-21.4 | 107.5421 | 103.8132 | 98.8827 | 92.1685 | 89.3226
-21.3 | 107.6 | 103.9 | 98.9 | 92.339 | 89.5123
-21.2 | 107.6 | 103.9419 | 98.9989 | 92.541 | 89.6267
-21.1 | 107.6334 | 104 | 99.1201 | 92.6 | 89.8143
-21 | 107.7 | 104.0342 | 99.2439 | 92.7358 | 89.9154
-20.9 | 107.7 | 104.1095 | 99.3449 | 92.8059 | 90.0327
-20.8 | 107.7 | 104.1685 | 99.4 | 93.0555 | 90.2174
-20.7 | 107.7229 | 104.247 | 99.447 | 93.2 | 90.3322
-20.6 | 107.7936 | 104.3 | 99.596 | 93.2472 | 90.5052
-20.5 | 107.8 | 104.349 | 99.749 | 93.4086 | 90.6185
-20.4 | 107.8 | 104.4 | 99.8438 | 93.5834 | 90.9
-20.3 | 107.8 | 104.451 | 99.9368 | 93.751 | 90.9643
-20.2 | 107.8 | 104.5 | 100.0441 | 93.852 | 91.1129
-20.1 | 107.752 | 104.5 | 100.1743 | 93.953 | 91.2678
-20 | 107.7015 | 104.5473 | 100.2157 | 94.1082 | 91.4157
-19.9 | 107.7 | 104.6 | 100.3 | 94.2552 | 91.5571
-19.8 | 107.7 | 104.6418 | 100.3616 | 94.4123 | 91.7185
-19.7 | 107.7857 | 104.7 | 100.4 | 94.5572 | 91.86
-19.6 | 107.8 | 104.7012 | 100.4634 | 94.6509 | 92.0012
-19.5 | 107.8 | 104.8 | 100.6184 | 94.7448 | 92.119
-19.4 | 107.8 | 104.8 | 100.7602 | 94.8535 | 92.2553
-19.3 | 107.8 | 104.8184 | 100.8612 | 94.9714 | 92.4184
-19.2 | 107.7949 | 104.9 | 100.9 | 95.1529 | 92.5194
-19.1 | 107.7061 | 104.9 | 100.9428 | 95.2775 | 92.6204
-19 | 107.8 | 104.9464 | 101.0357 | 95.3 | 92.7333
-18.9 | 107.8 | 104.9969 | 101.1367 | 95.4592 | 92.8904
-18.8 | 107.8 | 105.0511 | 101.2377 | 95.6612 | 92.9948
-18.7 | 107.8 | 105.1 | 101.3387 | 95.7816 | 93.1678
-18.6 | 107.8 | 105.1 | 101.4397 | 95.8578 | 93.3119
-18.5 | 107.8 | 105.1761 | 101.5 | 95.957 | 93.359
-18.4 | 107.8 | 105.2 | 101.5275 | 96.0984 | 93.4154
-18.3 | 107.8 | 105.2 | 101.6333 | 96.2307 | 93.5333
-18.2 | 107.8 | 105.2 | 101.7 | 96.3395 | 93.6511
-18.1 | 107.8 | 105.2 | 101.8 | 96.4349 | 93.7517
-18 | 107.8 | 105.3 | 101.8 | 96.5173 | 93.8459
-17.9 | 107.8 | 105.3 | 101.8256 | 96.6 | 93.9
-17.8 | 107.8 | 105.327 | 101.967 | 96.6193 | 93.9958
-17.7 | 107.8 | 105.4 | 102.0678 | 96.7203 | 94.1489
-17.6 | 107.8 | 105.4 | 102.1 | 96.8213 | 94.2499
-17.5 | 107.8 | 105.4729 | 102.1509 | 96.9223 | 94.3509
-17.4 | 107.8 | 105.5 | 102.2 | 97.0467 | 94.4519
-17.3 | 107.8 | 105.5 | 102.2529 | 97.2 | 94.5529
-17.2 | 107.8 | 105.5539 | 102.3 | 97.2254 | 94.6539
-17.1 | 107.8 | 105.6 | 102.3549 | 97.3168 | 94.7549
-17 | 107.8 | 105.6 | 102.4 | 97.3811 | 94.856
-16.9 | 107.8141 | 105.6 | 102.457 | 97.4454 | 94.957
-16.8 | 107.9 | 105.6812 | 102.5 | 97.5076 | 95
-16.7 | 107.8838 | 105.7 | 102.559 | 97.5581 | 95.0413
-16.6 | 107.8 | 105.7 | 102.6 | 97.6172 | 95.11
-16.5 | 107.8 | 105.7 | 102.661 | 97.7182 | 95.1689
-16.4 | 107.8 | 105.7049 | 102.7621 | 97.8103 | 95.2669
-16.3 | 107.8 | 105.8 | 102.8 | 97.8647 | 95.406
-16.2 | 107.8498 | 105.8248 | 102.8 | 97.9311 | 95.5054
-16.1 | 107.9 | 105.9 | 102.8 | 98 | 95.584
-16 | 107.9 | 105.9 | 102.8325 | 98.0079 | 95.6563
-15.9 | 107.8059 | 105.9 | 102.9 | 98.0963 | 95.754
-15.8 | 107.9 | 105.9 | 102.9538 | 98.1968 | 95.8955
-15.7 | 107.9 | 105.9 | 103 | 98.2855 | 95.9
-15.6 | 107.9 | 105.9 | 103 | 98.3 | 95.9988
-15.5 | 107.9 | 105.9 | 103 | 98.3855 | 96
-15.4 | 107.9 | 105.9 | 103.0436 | 98.4 | 96.1008
-15.3 | 107.9 | 105.9 | 103.1 | 98.4875 | 96.2
-15.2 | 107.9 | 105.94 | 103.1 | 98.5 | 96.3099
-15.1 | 107.9 | 106 | 103.1 | 98.5253 | 96.4
-15 | 107.9 | 106 | 103.1 | 98.6 | 96.4468
-14.9 | 107.9 | 105.9432 | 103.1 | 98.6282 | 96.5082
-14.8 | 107.9 | 105.836 | 103.1 | 98.7 | 96.6496
-14.7 | 107.8921 | 105.8 | 103.1 | 98.7311 | 96.7651
-14.6 | 107.8 | 105.8 | 103.1 | 98.8725 | 96.8661
-14.5 | 107.8 | 105.7758 | 103.1 | 98.9 | 96.9587
-14.4 | 107.7646 | 105.7 | 103.1 | 98.9461 | 97.0538
-14.3 | 107.7 | 105.7 | 103.0595 | 99 | 97.1
-14.2 | 107.7 | 105.6419 | 103 | 99.0701 | 97.1559
-14.1 | 107.6003 | 105.6 | 103 | 99.1 | 97.2
-14 | 107.5824 | 105.5492 | 103 | 99.1 | 97.2579
-13.9 | 107.4646 | 105.5 | 103 | 99.1731 | 97.3589
-13.8 | 107.4 | 105.4476 | 103 | 99.2 | 97.4
-13.7 | 107.3123 | 105.4 | 103 | 99.2 | 97.4609
-13.6 | 107.3 | 105.4 | 103 | 99.2905 | 97.5542
-13.5 | 107.2085 | 105.3513 | 103 | 99.3915 | 97.6
-13.4 | 107.1 | 105.3 | 103 | 99.4079 | 97.6497
-13.3 | 107.0491 | 105.3 | 103 | 99.5 | 97.7507
-13.2 | 107 | 105.2483 | 103 | 99.5231 | 97.8603
-13.1 | 106.933 | 105.2 | 103 | 99.6 | 97.9
-13 | 106.8206 | 105.1463 | 103 | 99.6552 | 98
-12.9 | 106.7167 | 105.1 | 103 | 99.7 | 98
-12.8 | 106.7 | 105 | 103 | 99.7 | 98.0226
-12.7 | 106.6146 | 105 | 103 | 99.7795 | 98.1
-12.6 | 106.5992 | 105 | 103 | 99.8 | 98.1
-12.5 | 106.5 | 105 | 103 | 99.8874 | 98.1731
-12.4 | 106.5 | 104.9116 | 103 | 99.9 | 98.2741
-12.3 | 106.4124 | 104.9 | 103.0322 | 99.9 | 98.3752
-12.2 | 106.3953 | 104.9 | 103.1 | 99.9 | 98.4666
-12.1 | 106.3 | 104.9 | 103.1 | 99.9 | 98.4372
-12 | 106.3 | 104.9 | 103.0218 | 99.9809 | 98.4638
-11.9 | 106.2922 | 104.8491 | 103 | 100 | 98.5649
-11.8 | 106.2 | 104.8 | 103 | 100 | 98.6
-11.7 | 106.1719 | 104.8 | 103 | 100 | 98.667
-11.6 | 106.1 | 104.7874 | 103 | 100.0822 | 98.7
-11.5 | 106.1 | 104.7 | 103 | 100.1 | 98.7
-11.4 | 106.1 | 104.7 | 102.9158 | 100.1842 | 98.798
-11.3 | 106.0004 | 104.7 | 102.9 | 100.2 | 98.8
-11.2 | 106 | 104.683 | 102.9 | 100.2 | 98.9
-11.1 | 105.9984 | 104.6241 | 102.9 | 100.2873 | 98.9016
-11 | 105.9 | 104.6 | 102.9 | 100.3 | 99.0026
-10.9 | 105.8687 | 104.6 | 102.9 | 100.3 | 99.1
-10.8 | 105.8 | 104.5607 | 102.9 | 100.3761 | 99.1221
-10.7 | 105.8 | 104.4943 | 102.9 | 100.4 | 99.2
-10.6 | 105.7088 | 104.4 | 102.9 | 100.4 | 99.2209
-10.5 | 105.6066 | 104.4 | 102.9 | 100.4 | 99.3
-10.4 | 105.6 | 104.3913 | 102.9 | 100.4 | 99.3721
-10.3 | 105.5045 | 104.3 | 102.9 | 100.4 | 99.4
-10.2 | 105.5 | 104.3 | 102.9 | 100.4 | 99.4
-10.1 | 105.4242 | 104.2883 | 102.9 | 100.4137 | 99.4689
-10 | 105.3301 | 104.2 | 102.9 | 100.5 | 99.5699
-9.9 | 105.3 | 104.1862 | 102.8 | 100.5281 | 99.6
-9.8 | 105.2281 | 104.0852 | 102.7424 | 100.6 | 99.5281
-9.7 | 105.1271 | 103.9842 | 102.7 | 100.6388 | 99.5729
-9.6 | 105.0602 | 103.9 | 102.6164 | 100.7 | 99.6
-9.5 | 105.0058 | 103.9 | 102.6 | 100.7 | 99.675
-9.4 | 104.9425 | 103.8591 | 102.6 | 100.7451 | 99.7
-9.3 | 104.9 | 103.8047 | 102.6 | 100.7956 | 99.7627
-9.2 | 104.859 | 103.8 | 102.6 | 100.8 | 99.8446
-9.1 | 104.8 | 103.8 | 102.6 | 100.8067 | 99.9
-9 | 104.7485 | 103.756 | 102.6 | 100.8951 | 99.9
-8.9 | 104.7 | 103.7088 | 102.6 | 100.9 | 99.9096
-8.8 | 104.6771 | 103.7 | 102.6 | 100.8036 | 100
-8.7 | 104.6182 | 103.7 | 102.6 | 100.8974 | 100
-8.6 | 104.5302 | 103.7 | 102.6 | 100.9 | 100.0481
-8.5 | 104.4292 | 103.7 | 102.6 | 100.9 | 100.1
-8.4 | 104.4 | 103.6246 | 102.6 | 100.9 | 100.1
-8.3 | 104.4 | 103.6 | 102.6 | 100.9 | 100.1
-8.2 | 104.3353 | 103.6 | 102.5965 | 100.9022 | 100.1882
-8.1 | 104.3 | 103.5394 | 102.5 | 100.9905 | 100.2
-8 | 104.3 | 103.5 | 102.5 | 101 | 100.2591
-7.9 | 104.2373 | 103.4373 | 102.5 | 101 | 100.3
-7.8 | 104.2 | 103.4 | 102.4363 | 101 | 100.3
-7.7 | 104.1353 | 103.4 | 102.4 | 101 | 100.3
-7.6 | 104.1 | 103.32 | 102.4 | 101 | 100.3
-7.5 | 104.0889 | 103.3 | 102.4 | 101 | 100.3
-7.4 | 104.0417 | 103.3 | 102.4 | 101 | 100.3
-7.3 | 104 | 103.2835 | 102.4 | 101 | 100.3487
-7.2 | 104 | 103.2192 | 102.3765 | 101 | 100.3992
-7.1 | 103.9905 | 103.2 | 102.3 | 101 | 100.4497
-7 | 103.9197 | 103.2 | 102.3 | 101 | 100.5
-6.9 | 103.8151 | 103.1651 | 102.3 | 101 | 100.5
-6.8 | 103.7229 | 103.1 | 102.3 | 101.0452 | 100.5
-6.7 | 103.6345 | 103.0537 | 102.3 | 101.1 | 100.5
-6.6 | 103.6 | 103 | 102.3 | 101.1 | 100.5473
-6.5 | 103.5937 | 103 | 102.3 | 101.1 | 100.6
-6.4 | 103.5 | 102.9793 | 102.2935 | 101.1 | 100.6
-6.3 | 103.5 | 102.9 | 102.2 | 101.1 | 100.6
-6.2 | 103.4852 | 102.9 | 102.2 | 101.1 | 100.6
-6.1 | 103.4 | 102.8667 | 102.1905 | 101.1 | 100.6
-6 | 103.3181 | 102.8 | 102.0895 | 101.1 | 100.6
-5.9 | 103.3 | 102.8 | 102 | 101.1 | 100.6
-5.8 | 103.2446 | 102.7224 | 102 | 101.1 | 100.6
-5.7 | 103.2 | 102.7 | 102 | 101.1 | 100.6
-5.6 | 103.1553 | 102.6 | 101.9 | 101.1 | 100.6
-5.5 | 103.1 | 102.6 | 101.9 | 101.0416 | 100.5416
-5.4 | 103.0691 | 102.5459 | 101.9 | 101 | 100.5595
-5.3 | 102.9681 | 102.4953 | 101.9 | 101 | 100.6
-5.2 | 102.9 | 102.4 | 101.9 | 100.9539 | 100.6
-5.1 | 102.9 | 102.4 | 101.9 | 100.9 | 100.6
-5 | 102.8911 | 102.365 | 101.9 | 101 | 100.6492
-4.9 | 102.8 | 102.3 | 101.8497 | 101 | 100.7
-4.8 | 102.8 | 102.263 | 101.7641 | 101 | 100.7
-4.7 | 102.6763 | 102.1667 | 101.7 | 101 | 100.6816
-4.6 | 102.5753 | 102.1 | 101.7 | 100.9181 | 100.603
-4.5 | 102.4743 | 102.0743 | 101.685 | 100.9 | 100.6
-4.4 | 102.4 | 102 | 101.6261 | 100.9 | 100.6
-4.3 | 102.3865 | 101.9722 | 101.5606 | 100.9 | 100.6
-4.2 | 102.3 | 101.9 | 101.5 | 100.9 | 100.6
-4.1 | 102.2845 | 101.8652 | 101.4845 | 100.9 | 100.6
-4 | 102.1835 | 101.8 | 101.4 | 100.8974 | 100.6
-3.9 | 102.1 | 101.7295 | 101.4 | 100.8 | 100.6
-3.8 | 102.1 | 101.7 | 101.31 | 100.8 | 100.6
-3.7 | 102.0829 | 101.5954 | 101.3 | 100.8 | 100.6
-3.6 | 101.9926 | 101.507 | 101.3 | 100.8 | 100.6
-3.5 | 101.9 | 101.5 | 101.207 | 100.7498 | 100.6
-3.4 | 101.8631 | 101.4059 | 101.1059 | 100.7 | 100.6
-3.3 | 101.7621 | 101.4 | 101 | 100.7 | 100.6
-3.2 | 101.6611 | 101.3039 | 101 | 100.7 | 100.6
-3.1 | 101.6 | 101.3 | 101 | 100.7 | 100.6
-3 | 101.5522 | 101.2826 | 101 | 100.7 | 100.5766
-2.9 | 101.4437 | 101.2 | 100.9843 | 100.7 | 100.5
-2.8 | 101.4 | 101.1798 | 100.9 | 100.7 | 100.5
-2.7 | 101.3 | 101.1 | 100.9 | 100.6417 | 100.5
-2.6 | 101.3 | 100.9962 | 100.9 | 100.6 | 100.5
-2.5 | 101.2755 | 100.9 | 100.9 | 100.6 | 100.5
-2.4 | 101.2 | 100.9 | 100.8101 | 100.6 | 100.4838
-2.3 | 101.1519 | 100.9 | 100.8 | 100.5376 | 100.4052
-2.2 | 101.0509 | 100.9 | 100.7081 | 100.5 | 100.5
-2.1 | 100.9499 | 100.8356 | 100.7 | 100.5 | 100.5
-2 | 100.8489 | 100.8 | 100.6684 | 100.4428 | 100.5
-1.9 | 100.8 | 100.7419 | 100.6 | 100.4 | 100.4907
-1.8 | 100.8 | 100.7 | 100.6 | 100.4 | 100.4
-1.7 | 100.703 | 100.6368 | 100.5 | 100.4 | 100.4
-1.6 | 100.7 | 100.6 | 100.5 | 100.4 | 100.4
-1.5 | 100.6724 | 100.5152 | 100.4724 | 100.3438 | 100.3009
-1.4 | 100.5714 | 100.4799 | 100.3714 | 100.3 | 100.3
-1.3 | 100.5 | 100.3385 | 100.3 | 100.3 | 100.3
-1.2 | 100.4836 | 100.3 | 100.3 | 100.2408 | 100.3
-1.1 | 100.3826 | 100.2356 | 100.3 | 100.2 | 100.3
-1 | 100.3 | 100.2 | 100.2142 | 100.2 | 100.3
-0.9 | 100.2806 | 100.1607 | 100.1949 | 100.2 | 100.2949
-0.8 | 100.2 | 100.1 | 100.1 | 100.2 | 100.2
-0.7 | 100.2 | 100.1 | 100.1 | 100.1214 | 100.2
-0.6 | 100.1918 | 100.1 | 100.1 | 100.1115 | 100.2
-0.5 | 100.1 | 100.0736 | 100.1 | 100.1471 | 100.2
-0.4 | 100.1 | 100.0028 | 100.0714 | 100.1 | 100.1755
-0.3 | 100.1 | 100 | 100 | 100.1 | 100.1
-0.2 | 100 | 100 | 100 | 100.1 | 100.0877
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100.001
0.2 | 100 | 100 | 100 | 100.0628 | 100.1
0.3 | 100.0243 | 100 | 100 | 100.1 | 100.103
0.4 | 100.1 | 100 | 100.0457 | 100.1 | 100.2
0.5 | 100.1051 | 100.0484 | 100.0129 | 100.1089 | 100.2
0.6 | 100.2 | 100.1 | 100.1 | 100.1315 | 100.2
0.7 | 100.2 | 100.1 | 100.1 | 100.1 | 100.2
0.8 | 100.2081 | 100.1 | 100.1761 | 100.2 | 100.2095
0.9 | 100.3 | 100.1377 | 100.2 | 100.2 | 100.3
1 | 100.3102 | 100.2 | 100.2 | 100.2 | 100.3
1.1 | 100.4098 | 100.2156 | 100.3 | 100.213 | 100.3
1.2 | 100.4982 | 100.3 | 100.3 | 100.3 | 100.3
1.3 | 100.5 | 100.3 | 100.3418 | 100.3 | 100.3
1.4 | 100.5999 | 100.4599 | 100.4 | 100.3 | 100.3
1.5 | 100.7 | 100.5008 | 100.4438 | 100.3152 | 100.3258
1.6 | 100.7 | 100.5892 | 100.5 | 100.4 | 100.4
1.7 | 100.7201 | 100.6173 | 100.5 | 100.4 | 100.4
1.8 | 100.8 | 100.7 | 100.5856 | 100.4 | 100.4183
1.9 | 100.8 | 100.7193 | 100.6 | 100.4 | 100.5
2 | 100.8774 | 100.8 | 100.6484 | 100.4203 | 100.4797
2.1 | 100.9785 | 100.8213 | 100.7 | 100.5 | 100.4213
2.2 | 101.0795 | 100.9 | 100.7938 | 100.5195 | 100.4805
2.3 | 101.1704 | 100.9 | 100.8 | 100.6 | 100.4091
2.4 | 101.2 | 100.9 | 100.8 | 100.6 | 100.5
2.5 | 101.2963 | 100.9155 | 100.8825 | 100.6 | 100.5
2.6 | 101.3 | 101 | 100.9 | 100.6121 | 100.5
2.7 | 101.3 | 101.1 | 100.9 | 100.7 | 100.5
2.8 | 101.4 | 101.1598 | 100.9 | 100.7 | 100.5
2.9 | 101.4723 | 101.2 | 100.951 | 100.7 | 100.5009
3 | 101.5 | 101.2626 | 101 | 100.7 | 100.6
3.1 | 101.5886 | 101.3 | 101 | 100.7 | 100.6
3.2 | 101.6896 | 101.3 | 101 | 100.7 | 100.6
3.3 | 101.7906 | 101.3764 | 101.0669 | 100.7 | 100.6
3.4 | 101.8917 | 101.4 | 101.1 | 100.7 | 100.6
3.5 | 101.9 | 101.4784 | 101.1784 | 100.7297 | 100.6
3.6 | 102.0195 | 101.5794 | 101.2794 | 100.8 | 100.6
3.7 | 102.1 | 101.6804 | 101.3 | 100.8 | 100.6
3.8 | 102.1 | 101.7 | 101.395 | 100.8 | 100.6
3.9 | 102.111 | 101.7967 | 101.4 | 100.8539 | 100.6
4 | 102.2 | 101.8 | 101.4 | 100.9 | 100.6
4.1 | 102.2114 | 101.8416 | 101.4559 | 100.9 | 100.6
4.2 | 102.2998 | 101.9 | 101.5 | 100.9 | 100.6
4.3 | 102.4 | 101.9437 | 101.5338 | 100.9 | 100.6
4.4 | 102.4018 | 102 | 101.5927 | 100.9 | 100.6
4.5 | 102.5028 | 102.0457 | 101.6563 | 100.9 | 100.6
4.6 | 102.6038 | 102.1467 | 101.7 | 100.9038 | 100.6
4.7 | 102.7057 | 102.2 | 101.7 | 101 | 100.6038
4.8 | 102.8 | 102.2487 | 101.7441 | 101 | 100.6823
4.9 | 102.8 | 102.3 | 101.8212 | 101 | 100.7
5 | 102.8138 | 102.3444 | 101.9 | 101 | 100.635
5.1 | 102.9 | 102.4 | 101.9 | 100.9275 | 100.6
5.2 | 102.9 | 102.4 | 101.9 | 100.9339 | 100.6
5.3 | 102.9967 | 102.4753 | 101.9 | 101 | 100.6
5.4 | 103.0977 | 102.5168 | 101.9 | 101 | 100.5309
5.5 | 103.1 | 102.6 | 101.9 | 101.013 | 100.5701
5.6 | 103.1776 | 102.6 | 101.9 | 101.1 | 100.6
5.7 | 103.2 | 102.661 | 101.9722 | 101.1 | 100.6
5.8 | 103.2732 | 102.7854 | 102 | 101.1 | 100.6
5.9 | 103.3 | 102.8 | 102 | 101.1 | 100.6
6 | 103.3533 | 102.8 | 102.0752 | 101.1 | 100.6
6.1 | 103.4 | 102.8389 | 102.1667 | 101.1 | 100.6
6.2 | 103.5 | 102.9 | 102.2 | 101.1 | 100.6
6.3 | 103.5 | 102.9 | 102.2 | 101.1 | 100.6
6.4 | 103.5 | 102.9507 | 102.265 | 101.1 | 100.6
6.5 | 103.5324 | 103 | 102.3 | 101.1 | 100.6
6.6 | 103.6 | 103.0336 | 102.3 | 101.1 | 100.533
6.7 | 103.668 | 103.1 | 102.3 | 101.1 | 100.5
6.8 | 103.7537 | 103.1 | 102.3 | 101.0738 | 100.5
6.9 | 103.8484 | 103.1415 | 102.3 | 101 | 100.5
7 | 103.9 | 103.2 | 102.3 | 101 | 100.4788
7.1 | 103.9146 | 103.2 | 102.3 | 101 | 100.4282
7.2 | 103.9651 | 103.2 | 102.3569 | 101 | 100.4
7.3 | 104.0156 | 103.2563 | 102.4 | 101 | 100.4
7.4 | 104.0661 | 103.3 | 102.4 | 101 | 100.3767
7.5 | 104.1 | 103.3 | 102.4 | 101 | 100.3262
7.6 | 104.1 | 103.3057 | 102.4 | 101 | 100.3
7.7 | 104.1 | 103.4 | 102.4 | 101 | 100.3
7.8 | 104.1591 | 103.4 | 102.4078 | 101 | 100.3
7.9 | 104.2659 | 103.4088 | 102.5 | 101 | 100.3
8 | 104.3 | 103.5 | 102.5 | 101 | 100.2392
8.1 | 104.3 | 103.5108 | 102.5394 | 101 | 100.2
8.2 | 104.3603 | 103.6 | 102.6 | 100.9167 | 100.1596
8.3 | 104.4 | 103.6 | 102.6 | 100.9 | 100.1
8.4 | 104.4 | 103.6097 | 102.6 | 100.9 | 100.1
8.5 | 104.436 | 103.6804 | 102.6 | 100.9 | 100.1
8.6 | 104.4865 | 103.7 | 102.6 | 100.9 | 100.0178
8.7 | 104.5432 | 103.7 | 102.6 | 100.9 | 100
8.8 | 104.6028 | 103.7304 | 102.6 | 100.8249 | 99.9964
8.9 | 104.6814 | 103.7809 | 102.6 | 100.8761 | 99.9
9 | 104.7 | 103.8 | 102.6 | 100.9 | 99.9
9.1 | 104.7781 | 103.8 | 102.6 | 100.8317 | 99.8953
9.2 | 104.8791 | 103.8324 | 102.6 | 100.7676 | 99.8246
9.3 | 104.9 | 103.8829 | 102.6 | 100.7171 | 99.7424
9.4 | 104.971 | 103.9 | 102.6 | 100.7 | 99.7
9.5 | 105 | 103.9 | 102.6 | 100.7 | 99.6464
9.6 | 105 | 103.9 | 102.6975 | 100.7 | 99.6
9.7 | 105.0612 | 103.9556 | 102.7 | 100.6587 | 99.5444
9.8 | 105.1566 | 104.0661 | 102.7138 | 100.6 | 99.5566
9.9 | 105.2577 | 104.1719 | 102.8 | 100.5566 | 99.6
10 | 105.3587 | 104.2 | 102.8621 | 100.5 | 99.5487
10.1 | 105.4464 | 104.2597 | 102.9 | 100.4403 | 99.5
10.2 | 105.5 | 104.3 | 102.9 | 100.4 | 99.4536
10.3 | 105.5331 | 104.3 | 102.9 | 100.4 | 99.4
10.4 | 105.6 | 104.3627 | 102.9 | 100.4 | 99.3435
10.5 | 105.6352 | 104.4 | 102.9 | 100.4 | 99.3
10.6 | 105.7362 | 104.4647 | 102.9 | 100.4 | 99.2067
10.7 | 105.8 | 104.5 | 102.9 | 100.4 | 99.2
10.8 | 105.8 | 104.539 | 102.9 | 100.3046 | 99.1046
10.9 | 105.8949 | 104.5979 | 102.9 | 100.3 | 99.0751
11 | 105.926 | 104.6 | 102.9 | 100.3 | 98.974
11.1 | 106 | 104.6 | 102.9 | 100.3 | 98.9
11.2 | 106 | 104.6633 | 102.9 | 100.2006 | 98.8608
11.3 | 106 | 104.7 | 102.9 | 100.2 | 98.8
11.4 | 106.03 | 104.7 | 102.9 | 100.2 | 98.7724
11.5 | 106.1 | 104.7 | 103 | 100.1118 | 98.7
11.6 | 106.1 | 104.7649 | 103 | 100.1 | 98.7
11.7 | 106.1 | 104.8 | 103 | 100.0098 | 98.6384
11.8 | 106.2 | 104.8 | 103 | 100 | 98.6
11.9 | 106.2208 | 104.9 | 103 | 100 | 98.5364
12 | 106.3 | 104.9 | 103 | 100 | 98.4354
12.1 | 106.3228 | 104.9 | 103.0825 | 99.9058 | 98.4575
12.2 | 106.4 | 104.9 | 103.1 | 99.9 | 98.4476
12.3 | 106.4 | 104.9 | 103.0608 | 99.9 | 98.3466
12.4 | 106.4469 | 104.9 | 103 | 99.9 | 98.3
12.5 | 106.5 | 104.984 | 103 | 99.9 | 98.2445
12.6 | 106.6 | 105 | 103 | 99.8149 | 98.1435
12.7 | 106.6 | 105 | 103 | 99.8 | 98.1
12.8 | 106.63 | 105 | 103 | 99.7 | 98
12.9 | 106.731 | 105.1 | 103 | 99.7 | 98
13 | 106.8448 | 105.132 | 103 | 99.6752 | 98
13.1 | 106.9 | 105.2 | 103 | 99.6 | 97.9
13.2 | 106.9626 | 105.2197 | 103 | 99.5452 | 97.8324
13.3 | 107.0742 | 105.3 | 103 | 99.5 | 97.7318
13.4 | 107.1 | 105.3 | 103 | 99.4296 | 97.6354
13.5 | 107.1227 | 105.3227 | 103 | 99.3201 | 97.6
13.6 | 107.2238 | 105.4 | 103 | 99.2191 | 97.5333
13.7 | 107.3 | 105.4 | 103 | 99.2 | 97.4323
13.8 | 107.3634 | 105.4259 | 103 | 99.2 | 97.4
13.9 | 107.4976 | 105.5 | 103 | 99.2 | 97.3303
14 | 107.5158 | 105.5325 | 103 | 99.1007 | 97.2293
14.1 | 107.6289 | 105.6 | 103 | 99.0997 | 97.2
14.2 | 107.7 | 105.7 | 103 | 99 | 97.2
14.3 | 107.7 | 105.7 | 103.0452 | 98.9967 | 97.1355
14.4 | 107.7847 | 105.7 | 103.1 | 98.9 | 97.1
14.5 | 107.8 | 105.7472 | 103.1 | 98.9 | 96.977
14.6 | 107.8197 | 105.8 | 103.1 | 98.8925 | 96.8375
14.7 | 107.9 | 105.8 | 103.1 | 98.7511 | 96.7365
14.8 | 107.9 | 105.9 | 103.1 | 98.7 | 96.6247
14.9 | 107.9 | 105.9117 | 103.0618 | 98.6482 | 96.5
15 | 107.9 | 106 | 103.0028 | 98.6 | 96.4267
15.1 | 107.9 | 105.961 | 103.0842 | 98.5 | 96.4
15.2 | 107.9 | 105.9 | 103.1 | 98.5 | 96.3239
15.3 | 107.9 | 105.9 | 103.1 | 98.5 | 96.1854
15.4 | 107.9 | 105.9 | 103.0722 | 98.4151 | 96.0722
15.5 | 107.9 | 105.9 | 103 | 98.4 | 96
15.6 | 107.9 | 105.9 | 103 | 98.313 | 95.9702
15.7 | 107.9 | 105.9 | 103 | 98.3 | 95.9
15.8 | 107.8824 | 105.9 | 102.9721 | 98.211 | 95.8753
15.9 | 107.826 | 105.9 | 102.9 | 98.1212 | 95.7946
16 | 107.9 | 105.9 | 102.8605 | 98.0403 | 95.666
16.1 | 107.9 | 105.9 | 102.8 | 98 | 95.5617
16.2 | 107.9 | 105.8498 | 102.8 | 97.961 | 95.4784
16.3 | 107.8069 | 105.8 | 102.7802 | 97.8887 | 95.3824
16.4 | 107.8 | 105.7192 | 102.7 | 97.8382 | 95.3038
16.5 | 107.8 | 105.7 | 102.6896 | 97.7754 | 95.1879
16.6 | 107.8114 | 105.7 | 102.6 | 97.6743 | 95.0837
16.7 | 107.9 | 105.7 | 102.5876 | 97.5856 | 95.0194
16.8 | 107.8866 | 105.7 | 102.5 | 97.5312 | 95
16.9 | 107.8 | 105.6 | 102.4856 | 97.4726 | 94.9284
17 | 107.8 | 105.6 | 102.4 | 97.4084 | 94.8274
17.1 | 107.8 | 105.6 | 102.3835 | 97.3317 | 94.7264
17.2 | 107.8 | 105.5825 | 102.2825 | 97.2472 | 94.7
17.3 | 107.8 | 105.5 | 102.2 | 97.1529 | 94.6243
17.4 | 107.8 | 105.5 | 102.2 | 97.0519 | 94.5233
17.5 | 107.8 | 105.4945 | 102.1795 | 96.9509 | 94.4223
17.6 | 107.8 | 105.4062 | 102.1 | 96.8499 | 94.2426
17.7 | 107.8 | 105.4 | 102.0871 | 96.7489 | 94.1203
17.8 | 107.8 | 105.347 | 101.93 | 96.6479 | 94.0193
17.9 | 107.8 | 105.3 | 101.8364 | 96.6 | 93.9285
18 | 107.8 | 105.3 | 101.8 | 96.5459 | 93.8315
18.1 | 107.8 | 105.2027 | 101.8 | 96.4228 | 93.7305
18.2 | 107.8 | 105.2 | 101.7 | 96.2867 | 93.6295
18.3 | 107.8 | 105.2 | 101.6571 | 96.1857 | 93.5499
18.4 | 107.8 | 105.199 | 101.5561 | 96.0846 | 93.4703
18.5 | 107.8 | 105.1107 | 101.4551 | 95.9836 | 93.3693
18.6 | 107.8 | 105.1 | 101.354 | 95.873 | 93.2659
18.7 | 107.8 | 105.1 | 101.3 | 95.7157 | 93.1571
18.8 | 107.8 | 105.0688 | 101.2663 | 95.6091 | 93.0627
18.9 | 107.8 | 105.0183 | 101.1653 | 95.5081 | 92.9643
19 | 107.8 | 104.9678 | 101.0642 | 95.3141 | 92.7857
19.1 | 107.7775 | 104.9173 | 100.9714 | 95.2061 | 92.5918
19.2 | 107.7206 | 104.9 | 100.9 | 95.1051 | 92.492
19.3 | 107.8 | 104.8469 | 100.8898 | 95.004 | 92.4035
19.4 | 107.8 | 104.8 | 100.7776 | 94.8744 | 92.303
19.5 | 107.8 | 104.8 | 100.5878 | 94.7734 | 92.104
19.6 | 107.8 | 104.7296 | 100.4868 | 94.6724 | 91.9815
19.7 | 107.7714 | 104.7 | 100.3857 | 94.5428 | 91.84
19.8 | 107.7 | 104.6704 | 100.3 | 94.3741 | 91.6785
19.9 | 107.7 | 104.6 | 100.2771 | 94.2836 | 91.5372
20 | 107.7 | 104.5684 | 100.2 | 94.1653 | 91.4513
20.1 | 107.7 | 104.5 | 100.1 | 93.9816 | 91.2343
20.2 | 107.7311 | 104.4706 | 100.0327 | 93.8806 | 91.1912
20.3 | 107.7816 | 104.4 | 99.9 | 93.7592 | 91.0314
20.4 | 107.8 | 104.4 | 99.8643 | 93.57 | 90.8859
20.5 | 107.8 | 104.3776 | 99.7679 | 93.4286 | 90.6838
20.6 | 107.7643 | 104.3 | 99.6531 | 93.2841 | 90.4909
20.7 | 107.7099 | 104.2756 | 99.4756 | 93.1257 | 90.3796
20.8 | 107.7 | 104.1852 | 99.4 | 92.9843 | 90.1888
20.9 | 107.7 | 104.1262 | 99.3735 | 92.8287 | 90.0756
21 | 107.7 | 104.0509 | 99.2725 | 92.6868 | 89.8868
21.1 | 107.7 | 103.9572 | 99.1501 | 92.5858 | 89.7858
21.2 | 107.6133 | 103.9 | 99.0155 | 92.4695 | 89.6695
21.3 | 107.6 | 103.8552 | 98.8981 | 92.2837 | 89.4858
21.4 | 107.6 | 103.8 | 98.7971 | 92.1827 | 89.3941
21.5 | 107.5477 | 103.7245 | 98.6919 | 92.0634 | 89.1919
21.6 | 107.4865 | 103.7 | 98.4949 | 91.8663 | 88.9899
21.7 | 107.4158 | 103.6225 | 98.3929 | 91.7116 | 88.7939
21.8 | 107.243 | 103.5314 | 98.3 | 91.5702 | 88.6859
21.9 | 107.1205 | 103.443 | 98.2572 | 91.4287 | 88.4838
22 | 106.939 | 103.3636 | 98.1546 | 91.239 | 88.2841
22.1 | 106.8049 | 103.2823 | 98.0226 | 91.1185 | 88.1074
22.2 | 106.4805 | 103.1057 | 97.8603 | 90.9349 | 88.0032
22.3 | 106.3037 | 103.0021 | 97.7593 | 90.7329 | 87.8042
22.4 | 106.0475 | 102.8737 | 97.6027 | 90.626 | 87.7011
22.5 | 105.6505 | 102.6802 | 97.4335 | 90.3861 | 87.6001
22.6 | 105.2525 | 102.5235 | 97.3157 | 90.1554 | 87.3991
22.7 | 104.7399 | 102.3547 | 97.1248 | 90.0267 | 87.2962
22.8 | 104.1798 | 102.1359 | 96.98 | 89.9257 | 87.0941
22.9 | 103.4727 | 101.8491 | 96.8349 | 89.7493 | 86.9
23 | 102.7656 | 101.5662 | 96.6329 | 89.5219 | 86.8413
23.1 | 101.9609 | 101.2968 | 96.4309 | 89.3798 | 86.7097
23.2 | 101.0518 | 100.9434 | 96.2644 | 89.2703 | 86.3505
23.3 | 100.0068 | 100.5394 | 96.136 | 89.1289 | 86.1634
23.4 | 98.8061 | 100.0692 | 95.9535 | 88.9458 | 86.0624
23.5 | 97.5939 | 99.5641 | 95.727 | 88.6919 | 85.9228
23.6 | 96.3516 | 99.082 | 95.4493 | 88.5045 | 85.7208
23.7 | 95.0316 | 98.4994 | 95.2209 | 88.2693 | 85.5188
23.8 | 93.6177 | 97.8089 | 94.9233 | 88.1168 | 85.3168
23.9 | 92.2034 | 97.1016 | 94.6407 | 87.9574 | 85.1574
24 | 90.7891 | 96.3222 | 94.3968 | 87.7854 | 85.0564
24.1 | 89.4196 | 95.245 | 93.9927 | 87.5291 | 84.9479
24.2 | 88.1521 | 94.4644 | 93.5881 | 87.3201 | 84.732
24.3 | 86.8933 | 93.2822 | 93.2092 | 87.1352 | 84.4492
24.4 | 85.6756 | 92.1184 | 92.9263 | 86.9331 | 84.2165
24.5 | 84.5639 | 90.9922 | 92.4071 | 86.7311 | 83.9434
24.6 | 83.2713 | 89.7885 | 92.0207 | 86.5291 | 83.7304
24.7 | 81.7931 | 88.4827 | 91.4755 | 86.2906 | 83.6689
24.8 | 79.4918 | 87.2642 | 90.8372 | 85.995 | 83.5274
24.9 | 76.0842 | 86.0278 | 90.3307 | 85.7826 | 83.2827
25 | 71.3204 | 84.7744 | 89.6232 | 85.5744 | 83.0492
25.1 | 65.6839 | 83.5021 | 88.9163 | 85.3387 | 82.7832
25.2 | 58.6097 | 82.2167 | 88.1105 | 85.1636 | 82.6273
25.3 | 50.947 | 81.1033 | 87.1172 | 84.9255 | 82.4803
25.4 | 43.4615 | 79.9915 | 86.1657 | 84.6579 | 82.2473
25.5 | 36.54 | 78.776 | 85.3347 | 84.4219 | 82.055
25.6 | 30.1123 | 77.3616 | 84.4122 | 84.1936 | 81.8087
25.7 | 25.0225 | 75.6889 | 83.3192 | 83.9526 | 81.6067
25.8 | 20.9252 | 73.0329 | 82.2709 | 83.6858 | 81.4046
25.9 | 17.8842 | 69.2711 | 81.0789 | 83.3827 | 81.2513
26 | 15.7126 | 64.6035 | 79.9889 | 83.0794 | 81.0509
26.1 | 13.7934 | 59.1618 | 78.9567 | 82.7764 | 80.7986
26.2 | 12.5099 | 53.3164 | 77.8409 | 82.4644 | 80.5448
26.3 | 11.4208 | 46.6933 | 76.7832 | 82.0604 | 80.3538
26.4 | 10.6236 | 40.1232 | 75.736 | 81.6618 | 80.1964
26.5 | 9.9404 | 34.3382 | 74.603 | 81.3082 | 79.9267
26.6 | 9.4483 | 29.3693 | 73.6138 | 80.9196 | 79.6438
26.7 | 9.0582 | 25.4158 | 72.5543 | 80.5156 | 79.5253
26.8 | 8.7701 | 22.2823 | 71.4901 | 80.0755 | 79.3317
26.9 | 8.5521 | 19.7674 | 70.0127 | 79.6209 | 79.1064
27 | 8.2661 | 17.8618 | 68.0758 | 79.154 | 78.8928
27.1 | 8.0685 | 16.4552 | 65.3887 | 78.6813 | 78.6278
27.2 | 7.8917 | 15.252 | 61.8024 | 78.1668 | 78.3288
27.3 | 7.7149 | 14.2308 | 57.5422 | 77.4655 | 78.1172
27.4 | 7.5646 | 13.4528 | 53.1413 | 76.8594 | 77.8227
27.5 | 7.4273 | 12.8165 | 47.7556 | 76.2378 | 77.6131
27.6 | 7.2626 | 12.2658 | 42.1713 | 75.4473 | 77.4109
27.7 | 7.1232 | 11.8731 | 36.8492 | 74.8177 | 77.1133
27.8 | 6.9606 | 11.4711 | 32.4764 | 74.0793 | 76.9069
27.9 | 6.8529 | 11.2057 | 28.9036 | 73.417 | 76.7048
28 | 6.735 | 10.8457 | 26.1664 | 72.61 | 76.3057
28.1 | 6.5853 | 10.6155 | 23.6645 | 71.8034 | 76.1008
28.2 | 6.4222 | 10.3124 | 21.8037 | 70.9958 | 75.7984
28.3 | 6.3321 | 10.0396 | 20.2074 | 70.083 | 75.5334
28.4 | 6.2163 | 9.8454 | 18.9567 | 69.2077 | 75.1931
28.5 | 6.0821 | 9.6531 | 17.9422 | 68.3867 | 74.9279
28.6 | 5.9811 | 9.3919 | 17.1341 | 67.5768 | 74.6288
28.7 | 5.8801 | 9.2151 | 16.3945 | 66.7687 | 74.2343
28.8 | 5.779 | 9.0295 | 15.7885 | 66.0377 | 73.8303
28.9 | 5.678 | 8.8275 | 15.2309 | 65.1459 | 73.4263
29 | 5.577 | 8.6255 | 14.8196 | 64.4305 | 73.1167
29.1 | 5.476 | 8.4617 | 14.3881 | 63.6224 | 72.6227
29.2 | 5.3749 | 8.3214 | 14.1083 | 62.7929 | 72.2141
29.3 | 5.2739 | 8.1194 | 13.7307 | 61.7828 | 71.8587
29.4 | 5.1729 | 7.9587 | 13.4042 | 60.7091 | 71.463
29.5 | 5.0719 | 7.8259 | 13.1018 | 59.125 | 70.9029
29.6 | 5 | 7.6661 | 12.799 | 57.0191 | 70.3778
29.7 | 4.9766 | 7.5482 | 12.5955 | 54.3225 | 69.9535
29.8 | 4.8975 | 7.3806 | 12.2949 | 51.2704 | 69.4016
29.9 | 4.7965 | 7.2393 | 12.0938 | 47.664 | 68.8073
30 | 4.6955 | 7.1383 | 11.917 | 43.9345 | 68.2144
30.1 | 4.5944 | 6.9745 | 11.6261 | 40.0044 | 67.5667
30.2 | 4.4967 | 6.8256 | 11.4154 | 36.3793 | 66.9655
30.3 | 4.4462 | 6.7078 | 11.2133 | 33.1868 | 66.435
30.4 | 4.3919 | 6.5984 | 11.0113 | 30.5016 | 65.8339
30.5 | 4.2976 | 6.4915 | 10.7139 | 28.0307 | 65.1325
30.6 | 4.2034 | 6.4032 | 10.5073 | 26.2004 | 64.3291
30.7 | 4.1513 | 6.3026 | 10.4026 | 24.7155 | 63.7158
30.8 | 4.1008 | 6.2016 | 10.2032 | 23.3777 | 63.1097
30.9 | 4.0503 | 6.1006 | 9.9673 | 22.3846 | 62.4367
31 | 3.9997 | 5.9996 | 9.8393 | 21.4452 | 61.7967
31.1 | 3.9211 | 5.8986 | 9.5975 | 20.6897 | 61.0896
31.2 | 3.8261 | 5.7979 | 9.4365 | 20.046 | 60.4279
31.3 | 3.8 | 5.7095 | 9.2938 | 19.4253 | 59.8218
31.4 | 3.724 | 5.5852 | 9.1337 | 18.9347 | 59.1684
31.5 | 3.7 | 5.4783 | 8.9903 | 18.4168 | 58.5097
31.6 | 3.622 | 5.414 | 8.8308 | 18.0617 | 57.9036
31.7 | 3.521 | 5.321 | 8.6848 | 17.6735 | 57.2473
31.8 | 3.5 | 5.22 | 8.488 | 17.356 | 56.4634
31.9 | 3.419 | 5.1595 | 8.3618 | 17.0731 | 55.3925
32 | 3.4 | 5.109 | 8.2608 | 16.7629 | 54.2061
32.1 | 3.3354 | 5.017 | 8.1196 | 16.3837 | 52.4723
32.2 | 3.3 | 4.916 | 7.964 | 16.1846 | 50.319
32.3 | 3.2864 | 4.8149 | 7.8442 | 15.9349 | 47.8661
32.4 | 3.2 | 4.7139 | 7.671 | 15.6418 | 45.1332
32.5 | 3.1817 | 4.6446 | 7.57 | 15.4259 | 42.2039
32.6 | 3.1 | 4.5803 | 7.469 | 15.1358 | 38.9173
32.7 | 3.072 | 4.516 | 7.368 | 14.9218 | 36.0055
32.8 | 3 | 4.441 | 7.2615 | 14.6297 | 33.4475
32.9 | 3 | 4.4 | 7.1436 | 14.4176 | 31.2018
33 | 2.9693 | 4.3793 | 7.0258 | 14.2156 | 29.3901
33.1 | 2.9 | 4.2783 | 6.9211 | 14.0135 | 27.7028
33.2 | 2.9 | 4.2 | 6.8067 | 13.8352 | 26.4152
33.3 | 2.8047 | 4.1762 | 6.7 | 13.6382 | 25.4145
33.4 | 2.8 | 4.1 | 6.6316 | 13.3542 | 24.5019
33.5 | 2.8 | 4.0741 | 6.5 | 13.1341 | 23.6191
33.6 | 2.7017 | 3.9731 | 6.4424 | 12.9321 | 22.9697
33.7 | 2.7 | 3.9 | 6.3 | 12.7301 | 22.3878
33.8 | 2.6997 | 3.8711 | 6.2396 | 12.614 | 21.7837
33.9 | 2.6 | 3.8 | 6.0987 | 12.4259 | 21.2945
34 | 2.6 | 3.7691 | 6 | 12.2238 | 20.8702
34.1 | 2.5971 | 3.7 | 5.9971 | 12.0442 | 20.5306
34.2 | 2.5087 | 3.6744 | 5.9087 | 11.9242 | 20.1826
34.3 | 2.5 | 3.6 | 5.8089 | 11.7465 | 19.7775
34.4 | 2.5 | 3.5951 | 5.7079 | 11.6222 | 19.4184
34.5 | 2.4831 | 3.5165 | 5.6069 | 11.4424 | 19.1593
34.6 | 2.4046 | 3.5 | 5.5059 | 11.2402 | 18.9402
34.7 | 2.4 | 3.4953 | 5.5 | 11.1191 | 18.6573
34.8 | 2.4 | 3.4448 | 5.4038 | 10.9362 | 18.4362
34.9 | 2.3149 | 3.3927 | 5.3028 | 10.8171 | 18.1824
35 | 2.3 | 3.3284 | 5.2236 | 10.6321 | 17.9608
35.1 | 2.3 | 3.3 | 5.2 | 10.5151 | 17.7588
35.2 | 2.3 | 3.2426 | 5.1373 | 10.414 | 17.4852
35.3 | 2.2416 | 3.2 | 5.0416 | 10.2979 | 17.2304
35.4 | 2.2 | 3.2 | 5 | 10.1368 | 17.0933
35.5 | 2.2 | 3.1462 | 4.9824 | 9.9955 | 16.7968
35.6 | 2.2 | 3.1 | 4.8837 | 9.8425 | 16.6873
35.7 | 2.2 | 3.1 | 4.7938 | 9.7925 | 16.3938
35.8 | 2.1445 | 3.0366 | 4.7 | 9.6389 | 16.2588
35.9 | 2.1 | 3 | 4.6641 | 9.4897 | 16.0568
36 | 2.1 | 2.9345 | 4.6 | 9.3484 | 15.8547
36.1 | 2.1 | 2.9 | 4.6 | 9.3 | 15.6527
36.2 | 2.1 | 2.9 | 4.5 | 9.1655 | 15.4712
36.3 | 2.1 | 2.9 | 4.46 | 9.0367 | 15.32
36.4 | 2.0447 | 2.8447 | 4.4 | 8.9304 | 15.118
36.5 | 2 | 2.8 | 4.4 | 8.8294 | 14.916
36.6 | 2 | 2.8 | 4.3397 | 8.7284 | 14.7139
36.7 | 2 | 2.8 | 4.2979 | 8.6274 | 14.5486
36.8 | 2 | 2.7121 | 4.2 | 8.5356 | 14.3616
36.9 | 1.9707 | 2.7 | 4.1955 | 8.4397 | 14.2254
37 | 1.9064 | 2.7 | 4.1 | 8.3387 | 14.1118
37.1 | 1.9 | 2.6233 | 4.0519 | 8.2376 | 13.9182
37.2 | 1.9 | 2.6 | 4 | 8.1366 | 13.7713
37.3 | 1.9 | 2.6 | 4 | 8.0356 | 13.6124
37.4 | 1.9 | 2.6 | 3.9489 | 7.9345 | 13.3737
37.5 | 1.8621 | 2.5 | 3.9 | 7.8335 | 13.2621
37.6 | 1.8 | 2.5 | 3.8611 | 7.7325 | 13.1611
37.7 | 1.8 | 2.5 | 3.8 | 7.6401 | 13.0202
37.8 | 1.8 | 2.5 | 3.8 | 7.5448 | 12.8591
37.9 | 1.8 | 2.4758 | 3.7153 | 7.4438 | 12.7511
38 | 1.8 | 2.4 | 3.7 | 7.4 | 12.6427
38.1 | 1.8 | 2.4 | 3.7 | 7.2834 | 12.4834
38.2 | 1.8 | 2.4 | 3.6122 | 7.2 | 12.3309
38.3 | 1.754 | 2.4 | 3.6 | 7.1397 | 12.213
38.4 | 1.7 | 2.3857 | 3.5214 | 7.0387 | 12.0743
38.5 | 1.7 | 2.3071 | 3.5 | 7 | 11.9328
38.6 | 1.7 | 2.3 | 3.5 | 6.951 | 11.8094
38.7 | 1.7 | 2.3 | 3.4437 | 6.85 | 11.7
38.8 | 1.7 | 2.3 | 3.4 | 6.8 | 11.6085
38.9 | 1.7 | 2.3 | 3.4 | 6.7479 | 11.4871
39 | 1.7 | 2.2326 | 3.3755 | 6.6469 | 11.3457
39.1 | 1.6 | 2.2 | 3.3 | 6.5459 | 11.2042
39.2 | 1.6 | 2.2 | 3.3 | 6.5 | 11.1628
39.3 | 1.6 | 2.2 | 3.2724 | 6.4438 | 11.0214
39.4 | 1.6 | 2.2 | 3.2 | 6.3428 | 10.86
39.5 | 1.6 | 2.1577 | 3.2 | 6.2491 | 10.8
39.6 | 1.6 | 2.1 | 3.1836 | 6.2 | 10.7
39.7 | 1.6 | 2.1 | 3.1 | 6.154 | 10.6157
39.8 | 1.6 | 2.1 | 3.1 | 6.1 | 10.5
39.9 |  | 2.1 | 3.1 | 6.0274 | 10.4161
40 |  | 2.1 | 3.0796 | 6 | 10.2714
40.1 |  | 2.0072 | 3 | 5.9084 | 10.2
40.2 |  | 2 | 3 | 5.8887 | 10.1
40.3 |  | 2 | 3 | 5.8 | 10.0272
40.4 |  | 2 | 2.9072 | 5.7257 | 9.9
40.5 |  | 2 | 2.9 | 5.7 | 9.8243
40.6 |  | 2 | 2.9 | 5.7 | 9.7786
40.7 |  | 2 | 2.8679 | 5.6308 | 9.6215
40.8 |  | 1.9715 | 2.8 | 5.5429 | 9.5
40.9 |  |  | 2.8 | 5.4419 | 9.4233
41 |  |  | 2.8 | 5.4 | 9.3772
41.1 |  |  | 2.8 | 5.4 | 9.3
41.2 |  |  | 2.7531 | 5.3389 | 9.2
41.3 |  |  | 2.7 | 5.2379 | 9.1775
41.4 |  |  | 2.7 | 5.2 | 9.0516
41.5 |  |  | 2.7 | 5.2 | 9
41.6 |  |  | 2.7 | 5.1406 | 8.8609
41.7 |  |  | 2.6623 | 5.1 | 8.8
41.8 |  |  | 2.6 | 5.0185 | 8.7
41.9 |  |  | 2.6 | 5 | 8.6444
42 |  |  | 2.6 | 4.9 | 8.6
42.1 |  |  | 2.6 |  | 8.5013

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.866 | 0.879 | 0.887 | 0.898 | 0.906 | 0.912 | 0.915 | 0.919 | 0.92
| 4 | 0.879 | 0.895 | 0.905 | 0.919 | 0.929 | 0.937 | 0.942 | 0.947 | 0.948
| 5 | 0.889 | 0.907 | 0.92 | 0.935 | 0.948 | 0.958 | 0.963 | 0.969 | 0.97
| 7 | 0.903 | 0.924 | 0.94 | 0.959 | 0.975 | 0.989 | 0.996 | 1.004 | 1.006
| 10 | 0.914 | 0.938 | 0.956 | 0.98 | 1 | 1.018 | 1.029 | 1.039 | 1.042
| 15 | 0.922 | 0.948 | 0.969 | 0.996 | 1.022 | 1.045 | 1.059 | 1.073 | 1.078
| 20 | 0.926 | 0.954 | 0.976 | 1.005 | 1.033 | 1.06 | 1.077 | 1.095 | 1.101
| 30 | 0.931 | 0.96 | 0.983 | 1.015 | 1.047 | 1.078 | 1.099 | 1.121 | 1.129
| 40 | 0.933 | 0.963 | 0.986 | 1.02 | 1.053 | 1.086 | 1.109 | 1.132 | 1.141

## Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 55.3 | 59.9 | 67.6 | 73.3 | 75.5
0.1 | 59.625 | 64.1 | 71.3 | 77.06 | 78.96
0.2 | 66.7 | 70.48 | 77.3 | 82.3 | 83.8
0.3 | 76.85 | 79.38 | 85 | 88.6 | 89.5
0.4 | 84.64 | 86.675 | 90.38 | 92.96 | 93.82
0.5 | 90.5 | 91.925 | 94.64 | 96.1333 | 96.54
0.6 | 94.24 | 95.125 | 96.92 | 97.9333 | 98.2
0.7 | 96.9 | 97.4 | 98.32 | 99.04 | 99.2
0.8 | 98.6 | 98.8 | 99.36 | 99.68 | 99.85
0.9 | 99.4667 | 99.62 | 99.84 | 100 | 100
1 | 99.8 | 100 | 100 | 100 | 100
1.1 | 100 | 100 | 99.9833 | 100 | 100
1.2 | 100 | 100 | 99.9 | 99.9 | 99.875
1.3 | 99.925 | 99.8333 | 99.7 | 99.6667 | 99.6
1.4 | 99.58 | 99.6 | 99.5 | 99.3667 | 99.22
1.5 | 99.125 | 99.325 | 99.06 | 98.9667 | 98.94
1.6 | 98.5333 | 98.925 | 98.82 | 98.6 | 98.575
1.7 | 98.3 | 98.55 | 98.4 | 98.3 | 98.26
1.8 | 97.8 | 98.05 | 98 | 98 | 97.88
1.9 | 97.4 | 97.8 | 97.6 | 97.68 | 97.58
2 | 96.7 | 97.25 | 97.25 | 97.25 | 97.32
2.1 | 96.125 | 96.9667 | 96.9 | 96.96 | 96.88
2.2 | 95.64 | 96.2667 | 96.4 | 96.48 | 96.4714
2.3 | 95.25 | 95.9 | 95.96 | 96.18 | 96.14
2.4 | 94.575 | 95.5 | 95.6 | 95.76 | 95.7333
2.5 | 94.025 | 95.22 | 95.3143 | 95.37 | 95.46
2.6 | 93.52 | 94.68 | 94.9333 | 95.07 | 95.1286
2.7 | 93.04 | 94.32 | 94.6 | 94.75 | 94.74
2.8 | 92.45 | 93.85 | 94.1 | 94.28 | 94.1667
2.9 | 91.75 | 93.45 | 93.65 | 94 | 93.9333
3 | 91.4 | 92.9 | 93.4 | 93.6 | 93.58
3.1 | 90.78 | 92.36 | 92.9571 | 93.2 | 93.34
3.2 | 90.3 | 91.98 | 92.54 | 92.8 | 92.92
3.3 | 89.9333 | 91.5 | 92.14 | 92.5 | 92.58
3.4 | 89.28 | 91.05 | 91.7714 | 92.1 | 92.18
3.5 | 88.75 | 90.7 | 91.44 | 91.8 | 91.8667
3.6 | 88.325 | 90.3 | 90.95 | 91.3455 | 91.45
3.7 | 87.7 | 89.84 | 90.6 | 91.03 | 91.15
3.8 | 87.2 | 89.35 | 90.1 | 90.73 | 90.675
3.9 | 86.48 | 88.9 | 89.7 | 90.34 | 90.275
4 | 86.05 | 88.425 | 89.5 | 89.94 | 89.975
4.1 | 85.58 | 88 | 89.05 | 89.63 | 89.6
4.2 | 85.05 | 87.55 | 88.62 | 89.15 | 89.175
4.3 | 84.4 | 87.1 | 88.2333 | 88.74 | 88.9
4.4 | 83.8 | 86.7333 | 87.8 | 88.32 | 88.5
4.5 | 83.35 | 86.275 | 87.5 | 87.88 | 88.2
4.6 | 82.825 | 85.85 | 87.22 | 87.6 | 87.8
4.7 | 82.425 | 85.35 | 86.75 | 87.3 | 87.45
4.8 | 81.85 | 84.98 | 86.24 | 87 | 87.04
4.9 | 81.35 | 84.4 | 85.95 | 86.6 | 86.64
5 | 80.7 | 84 | 85.44 | 86.26 | 86.34
5.1 | 80.3 | 83.6 | 85.06 | 85.75 | 85.9333
5.2 | 79.8 | 83.18 | 84.68 | 85.52 | 85.58
5.3 | 79.25 | 82.8 | 84.36 | 85.14 | 85.28
5.4 | 78.8 | 82.35 | 83.96 | 84.825 | 84.95
5.5 | 78.225 | 81.9 | 83.575 | 84.3667 | 84.5
5.6 | 77.65 | 81.45 | 83.18 | 84.1 | 84.1
5.7 | 77.25 | 81.06 | 82.78 | 83.7 | 83.7
5.8 | 76.8 | 80.6333 | 82.45 | 83.3 | 83.4
5.9 | 76.2333 | 80.2333 | 82 | 82.875 | 83.1
6 | 75.9 | 79.6333 | 81.56 | 82.55 | 82.7
6.1 | 75.35 | 79.2 | 81.26 | 82.2 | 82.36
6.2 | 74.7 | 78.825 | 80.92 | 81.72 | 81.9333
6.3 | 74.3 | 78.35 | 80.52 | 81.38 | 81.6333
6.4 | 73.7 | 77.9 | 80.22 | 81.04 | 81.26
6.5 | 73.35 | 77.5 | 79.7333 | 80.74 | 80.86
6.6 | 72.925 | 77.2 | 79.36 | 80.34 | 80.46
6.7 | 72.425 | 76.7667 | 78.88 | 79.94 | 80.16
6.8 | 71.85 | 76.35 | 78.64 | 79.64 | 79.78
6.9 | 71.4 | 75.85 | 78.16 | 79.24 | 79.46
7 | 70.92 | 75.5 | 77.86 | 78.84 | 79.04
7.1 | 70.5 | 75 | 77.46 | 78.44 | 78.58
7.2 | 69.94 | 74.6 | 77.06 | 78.14 | 78.4
7.3 | 69.54 | 74.3 | 76.675 | 77.84 | 77.9857
7.4 | 68.98 | 73.7667 | 76.375 | 77.48 | 77.64
7.5 | 68.64 | 73.35 | 76 | 77.08 | 77.32
7.6 | 68.1333 | 72.925 | 75.6 | 76.68 | 76.92
7.7 | 67.65 | 72.6 | 75.2333 | 76.38 | 76.56
7.8 | 67.175 | 72.2 | 74.95 | 75.98 | 76.22
7.9 | 66.68 | 71.72 | 74.54 | 75.68 | 75.76
8 | 66.3 | 71.32 | 74.1 | 75.28 | 75.56
8.1 | 65.8 | 70.85 | 73.825 | 74.94 | 75.32
8.2 | 65.4 | 70.525 | 73.35 | 74.6 | 74.92
8.3 | 64.95 | 70.075 | 73.025 | 74.32 | 74.65
8.4 | 64.625 | 69.8 | 72.625 | 73.9 | 74.2333
8.5 | 64.05 | 69.35 | 72.225 | 73.66 | 73.7333
8.6 | 63.65 | 68.9333 | 71.925 | 73.22 | 73.54
8.7 | 63.3 | 68.475 | 71.525 | 72.92 | 73.16
8.8 | 62.8 | 68.15 | 71.2333 | 72.6333 | 72.88
8.9 | 62.4 | 67.76 | 70.8667 | 72.22 | 72.54
9 | 62.02 | 67.38 | 70.4667 | 71.8 | 72.16
9.1 | 61.65 | 67 | 70.15 | 71.45 | 71.86
9.2 | 61.225 | 66.55 | 69.76 | 71.22 | 71.46
9.3 | 60.75 | 66.2 | 69.4 | 70.94 | 71.08
9.4 | 60.38 | 65.8333 | 69 | 70.64 | 70.8667
9.5 | 60 | 65.45 | 68.7 | 70.24 | 70.4833
9.6 | 59.5 | 65.125 | 68.5 | 69.94 | 70.2
9.7 | 59.1 | 64.675 | 68.1 | 69.5667 | 69.8
9.8 | 58.7667 | 64.25 | 67.725 | 69.18 | 69.5
9.9 | 58.4 | 63.9333 | 67.4333 | 68.94 | 69.3
10 | 57.975 | 63.6 | 67.1 | 68.64 | 68.9
10.1 | 57.675 | 63.25 | 66.7667 | 68.28 | 68.6
10.2 | 57.25 | 62.8 | 66.4333 | 67.94 | 68.3
10.3 | 56.85 | 62.44 | 66.0333 | 67.52 | 67.8
10.4 | 56.5 | 62.125 | 65.6333 | 67.24 | 67.6
10.5 | 56.15 | 61.775 | 65.3667 | 67 | 67.225
10.6 | 55.7 | 61.35 | 65.05 | 66.68 | 66.95
10.7 | 55.2 | 60.975 | 64.8 | 66.3 | 66.65
10.8 | 54.925 | 60.625 | 64.35 | 66.06 | 66.35
10.9 | 54.5333 | 60.36 | 63.95 | 65.72 | 66.05
11 | 54.05 | 60 | 63.7 | 65.32 | 65.75
11.1 | 53.8333 | 59.5 | 63.3 | 65.06 | 65.4
11.2 | 53.375 | 59.15 | 63 | 64.72 | 65.0333
11.3 | 53.075 | 58.85 | 62.7 | 64.3333 | 64.7667
11.4 | 52.8 | 58.55 | 62.35 | 64.04 | 64.45
11.5 | 52.4 | 58.14 | 62.05 | 63.74 | 64.1636
11.6 | 52 | 57.85 | 61.72 | 63.4333 | 63.9
11.7 | 51.8 | 57.6 | 61.4 | 63.05 | 63.5429
11.8 | 51.4333 | 57.2333 | 61.125 | 62.8 | 63.14
11.9 | 51.05 | 56.8333 | 60.82 | 62.54 | 62.92
12 | 50.7667 | 56.5 | 60.525 | 62.2667 | 62.63
12.1 | 50.275 | 56.12 | 60.1333 | 61.9667 | 62.33
12.2 | 50.025 | 55.925 | 59.8667 | 61.7 | 62.03
12.3 | 49.58 | 55.45 | 59.5667 | 61.35 | 61.73
12.4 | 49.2 | 55.0667 | 59.175 | 60.9667 | 61.43
12.5 | 49 | 54.8 | 58.9 | 60.68 | 61.0727
12.6 | 48.6667 | 54.44 | 58.5333 | 60.46 | 60.92
12.7 | 48.2 | 54.2 | 58.275 | 60.175 | 60.6
12.8 | 47.9667 | 53.82 | 58 | 59.825 | 60.18
12.9 | 47.625 | 53.525 | 57.65 | 59.6 | 59.9
13 | 47.35 | 53.175 | 57.3 | 59.275 | 59.6
13.1 | 46.98 | 52.85 | 57.0667 | 58.9 | 59.3
13.2 | 46.7 | 52.575 | 56.68 | 58.6 | 59
13.3 | 46.35 | 52.22 | 56.4 | 58.275 | 58.7
13.4 | 46.1 | 51.9 | 56.08 | 58.04 | 58.5286
13.5 | 45.76 | 51.5 | 55.66 | 57.8667 | 58.22
13.6 | 45.46 | 51.2667 | 55.5 | 57.5333 | 57.92
13.7 | 45.2 | 50.95 | 55.3 | 57.1667 | 57.56
13.8 | 44.84 | 50.675 | 54.975 | 57 | 57.32
13.9 | 44.6 | 50.34 | 54.675 | 56.68 | 57.02
14 | 44.25 | 50.06 | 54.375 | 56.35 | 56.7667
14.1 | 43.9 | 49.725 | 54.075 | 56.16 | 56.475
14.2 | 43.65 | 49.475 | 53.75 | 55.8 | 56.24
14.3 | 43.36 | 49.2 | 53.52 | 55.5333 | 55.9333
14.4 | 43.075 | 49 | 53.32 | 55.3 | 55.68
14.5 | 42.825 | 48.65 | 52.9 | 55.0333 | 55.46
14.6 | 42.525 | 48.3 | 52.64 | 54.7333 | 55.18
14.7 | 42.275 | 47.96 | 52.375 | 54.36 | 54.88
14.8 | 41.9333 | 47.7 | 52.2 | 54.2333 | 54.66
14.9 | 41.66 | 47.4 | 51.8 | 53.8667 | 54.38
15 | 41.3 | 47.1 | 51.6 | 53.66 | 54.08
15.1 | 41.1 | 46.9 | 51.25 | 53.28 | 53.7833
15.2 | 40.8 | 46.525 | 51 | 53.18 | 53.6
15.3 | 40.5333 | 46.3 | 50.7 | 52.78 | 53.4
15.4 | 40.3 | 46.0667 | 50.5 | 52.58 | 53.1
15.5 | 40.025 | 45.7 | 50.2 | 52.38 | 52.8
15.6 | 39.675 | 45.45 | 50 | 52.18 | 52.5
15.7 | 39.44 | 45.2 | 49.6 | 51.86 | 52.24
15.8 | 39.2667 | 44.975 | 49.4 | 51.56 | 52
15.9 | 39 | 44.65 | 49.1 | 51.2667 | 51.8
16 | 38.7 | 44.375 | 48.9 | 50.9667 | 51.4333
16.1 | 38.4667 | 44.125 | 48.6 | 50.7 | 51.22
16.2 | 38.1667 | 43.78 | 48.4333 | 50.54 | 51.02
16.3 | 37.9 | 43.55 | 48.1333 | 50.14 | 50.72
16.4 | 37.675 | 43.3 | 47.7333 | 49.95 | 50.44
16.5 | 37.5 | 43.14 | 47.6 | 49.72 | 50.22
16.6 | 37.18 | 42.8 | 47.42 | 49.46 | 50
16.7 | 36.88 | 42.5 | 47.15 | 49.22 | 49.74
16.8 | 36.65 | 42.2 | 46.82 | 48.96 | 49.42
16.9 | 36.55 | 42.025 | 46.5 | 48.72 | 49.2333
17 | 36.3 | 41.775 | 46.325 | 48.46 | 48.94
17.1 | 35.96 | 41.525 | 46.04 | 48.22 | 48.74
17.2 | 35.775 | 41.3 | 45.74 | 47.96 | 48.44
17.3 | 35.525 | 40.9333 | 45.5333 | 47.68 | 48.18
17.4 | 35.275 | 40.64 | 45.38 | 47.48 | 47.9667
17.5 | 35.0333 | 40.5 | 45.16 | 47.1 | 47.74
17.6 | 34.8 | 40.2 | 44.86 | 46.98 | 47.48
17.7 | 34.55 | 39.8333 | 44.56 | 46.68 | 47.24
17.8 | 34.3 | 39.675 | 44.28 | 46.58 | 46.94
17.9 | 34.1 | 39.5 | 44.1 | 46.26 | 46.78
18 | 33.925 | 39.225 | 43.8 | 46.1 | 46.5
18.1 | 33.6 | 39.08 | 43.6 | 45.7667 | 46.3
18.2 | 33.38 | 38.85 | 43.4 | 45.6 | 46.1
18.3 | 33.2 | 38.6 | 43.1 | 45.4 | 45.9
18.4 | 33 | 38.35 | 42.9 | 45.2 | 45.6167
18.5 | 32.8333 | 38.12 | 42.7 | 44.8 | 45.42
18.6 | 32.6 | 37.925 | 42.46 | 44.7 | 45.22
18.7 | 32.3333 | 37.675 | 42.26 | 44.5 | 44.96
18.8 | 32.175 | 37.4 | 42.02 | 44.2 | 44.7667
18.9 | 31.9 | 37.225 | 41.76 | 44 | 44.5
19 | 31.7 | 36.88 | 41.56 | 43.8 | 44.3667
19.1 | 31.5 | 36.725 | 41.26 | 43.5 | 44.12
19.2 | 31.3 | 36.48 | 41.06 | 43.3167 | 43.8667
19.3 | 31.2 | 36.3 | 40.86 | 43.1167 | 43.6
19.4 | 30.95 | 36.0667 | 40.65 | 42.92 | 43.4
19.5 | 30.7 | 35.8667 | 40.4667 | 42.62 | 43.2
19.6 | 30.525 | 35.6 | 40.225 | 42.42 | 43
19.7 | 30.28 | 35.4 | 39.94 | 42.22 | 42.8
19.8 | 30 | 35.225 | 39.74 | 42.02 | 42.52
19.9 | 29.9 | 35.075 | 39.6 | 41.82 | 42.3
20 | 29.65 | 34.8 | 39.42 | 41.6333 | 42.12
20.1 | 29.5 | 34.58 | 39.12 | 41.34 | 41.92
20.2 | 29.375 | 34.4 | 38.925 | 41.15 | 41.72
20.3 | 29.125 | 34.2 | 38.68 | 40.94 | 41.52
20.4 | 28.9 | 33.9333 | 38.5 | 40.74 | 41.32
20.5 | 28.8 | 33.7 | 38.34 | 40.64 | 41.12
20.6 | 28.525 | 33.6 | 38.08 | 40.34 | 40.82
20.7 | 28.4 | 33.375 | 37.88 | 40.14 | 40.62
20.8 | 28.15 | 33.225 | 37.64 | 39.88 | 40.4
20.9 | 28 | 32.9 | 37.46 | 39.64 | 40.26
21 | 27.825 | 32.725 | 37.22 | 39.54 | 40
21.1 | 27.675 | 32.58 | 37.04 | 39.28 | 39.86
21.2 | 27.5 | 32.4 | 36.86 | 39.1 | 39.66
21.3 | 27.3 | 32.2 | 36.66 | 38.85 | 39.36
21.4 | 27.175 | 31.9667 | 36.46 | 38.74 | 39.1
21.5 | 26.925 | 31.8 | 36.26 | 38.425 | 39
21.6 | 26.7667 | 31.6 | 36 | 38.2333 | 38.9
21.7 | 26.6 | 31.4 | 35.86 | 38.05 | 38.6
21.8 | 26.5 | 31.18 | 35.66 | 37.85 | 38.4
21.9 | 26.25 | 31.04 | 35.46 | 37.75 | 38.1833
22 | 26.025 | 30.9 | 35.26 | 37.5667 | 38
22.1 | 25.9 | 30.6333 | 35 | 37.32 | 37.9
22.2 | 25.7 | 30.375 | 34.92 | 37.12 | 37.7
22.3 | 25.6 | 30.3 | 34.68 | 36.85 | 37.46
22.4 | 25.4 | 30.1 | 34.48 | 36.7333 | 37.26
22.5 | 25.25 | 29.825 | 34.2667 | 36.525 | 37.1
22.6 | 25.1 | 29.74 | 34.04 | 36.34 | 36.92
22.7 | 24.86 | 29.5 | 33.84 | 36.18 | 36.6667
22.8 | 24.725 | 29.4 | 33.7 | 35.94 | 36.475
22.9 | 24.575 | 29.15 | 33.5 | 35.74 | 36.28
23 | 24.4 | 29.02 | 33.28 | 35.64 | 36.18
23.1 | 24.2 | 28.8 | 33.1 | 35.44 | 35.98
23.2 | 24.1 | 28.625 | 32.9857 | 35.24 | 35.78
23.3 | 23.98 | 28.45 | 32.7667 | 35.04 | 35.58
23.4 | 23.8 | 28.3 | 32.54 | 34.84 | 35.4
23.5 | 23.6 | 28.14 | 32.44 | 34.66 | 35.2333
23.6 | 23.5 | 27.975 | 32.2571 | 34.5 | 35.1
23.7 | 23.4 | 27.82 | 32.1 | 34.26 | 34.94
23.8 | 23.175 | 27.6 | 31.9 | 34.16 | 34.74
23.9 | 23.1 | 27.4 | 31.7 | 33.96 | 34.5
24 | 22.9 | 27.28 | 31.5 | 33.76 | 34.4
24.1 | 22.8 | 27.14 | 31.4 | 33.6 | 34.25
24.2 | 22.65 | 26.98 | 31.3 | 33.46 | 34.05
24.3 | 22.5 | 26.8 | 31 | 33.26 | 33.8
24.4 | 22.3 | 26.64 | 30.9 | 33.1 | 33.6667
24.5 | 22.2 | 26.5 | 30.7 | 32.96 | 33.4667
24.6 | 22.06 | 26.34 | 30.5 | 32.76 | 33.3667
24.7 | 21.825 | 26.2 | 30.4 | 32.6 | 33.2
24.8 | 21.775 | 26.025 | 30.2 | 32.42 | 33
24.9 | 21.6 | 25.875 | 30 | 32.2667 | 32.8
25 | 21.5 | 25.725 | 29.9 | 32.08 | 32.6333
25.1 | 21.35 | 25.575 | 29.7167 | 31.9333 | 32.475
25.2 | 21.2 | 25.475 | 29.5333 | 31.78 | 32.3
25.3 | 21.1 | 25.3 | 29.4 | 31.59 | 32.2
25.4 | 20.9 | 25.075 | 29.2333 | 31.48 | 32
25.5 | 20.8 | 24.925 | 29.0333 | 31.28 | 31.775
25.6 | 20.6 | 24.8 | 28.9 | 31.09 | 31.62
25.7 | 20.6 | 24.65 | 28.76 | 30.98 | 31.52
25.8 | 20.425 | 24.55 | 28.6 | 30.78 | 31.32
25.9 | 20.28 | 24.4 | 28.4 | 30.64 | 31.2
26 | 20.1333 | 24.2 | 28.26 | 30.48 | 31.02
26.1 | 19.98 | 24.075 | 28.16 | 30.28 | 30.82
26.2 | 19.9 | 23.9 | 27.96 | 30.08 | 30.6667
26.3 | 19.8 | 23.8 | 27.8 | 29.94 | 30.5
26.4 | 19.6667 | 23.65 | 27.66 | 29.8 | 30.38
26.5 | 19.5 | 23.5 | 27.5 | 29.68 | 30.2333
26.6 | 19.3667 | 23.34 | 27.36 | 29.5 | 30.1
26.7 | 19.3 | 23.2 | 27.16 | 29.38 | 29.9333
26.8 | 19.14 | 23.1 | 27.06 | 29.2 | 29.8
26.9 | 19 | 23 | 26.88 | 29.08 | 29.6333
27 | 18.9 | 22.8 | 26.78 | 28.88 | 29.5
27.1 | 18.8 | 22.7 | 26.5333 | 28.78 | 29.3
27.2 | 18.7 | 22.575 | 26.42 | 28.6 | 29.2
27.3 | 18.6 | 22.425 | 26.3 | 28.5 | 29
27.4 | 18.475 | 22.3 | 26.15 | 28.3 | 28.9
27.5 | 18.3 | 22.16 | 26.04 | 28.2 | 28.7
27.6 | 18.2 | 22.02 | 25.84 | 28 | 28.6
27.7 | 18.1 | 21.9 | 25.7 | 27.9 | 28.5
27.8 | 18 | 21.7 | 25.6 | 27.7 | 28.26
27.9 | 17.9 | 21.6 | 25.5 | 27.62 | 28.16
28 | 17.8 | 21.5 | 25.32 | 27.4 | 28
28.1 | 17.7 | 21.4 | 25.2 | 27.3 | 27.86
28.2 | 17.52 | 21.2 | 25.02 | 27.12 | 27.7
28.3 | 17.4 | 21.1 | 24.82 | 27.02 | 27.5667
28.4 | 17.3 | 21 | 24.72 | 26.9 | 27.425
28.5 | 17.2 | 20.8667 | 24.62 | 26.7333 | 27.275
28.6 | 17.1 | 20.7667 | 24.5 | 26.6 | 27.1
28.7 | 17 | 20.6 | 24.4 | 26.44 | 26.98
28.8 | 16.9 | 20.475 | 24.22 | 26.3333 | 26.825
28.9 | 16.8 | 20.38 | 24.12 | 26.2 | 26.74
29 | 16.7 | 20.3 | 23.92 | 26.02 | 26.525
29.1 | 16.575 | 20.18 | 23.82 | 25.9 | 26.4
29.2 | 16.5 | 20.05 | 23.7333 | 25.74 | 26.34
29.3 | 16.3 | 19.9 | 23.6 | 25.64 | 26.2
29.4 | 16.2 | 19.75 | 23.44 | 25.54 | 26.04
29.5 | 16.15 | 19.6 | 23.3 | 25.4 | 25.9
29.6 | 16 | 19.5 | 23.2 | 25.24 | 25.8
29.7 | 15.96 | 19.425 | 23.04 | 25.15 | 25.7
29.8 | 15.8 | 19.275 | 22.925 | 25 | 25.5
29.9 | 15.7 | 19.1333 | 22.85 | 24.9 | 25.4
30 | 15.6 | 19.1 | 22.7 | 24.75 | 25.3

## SSD80 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 56.2 | 61.8 | 70.9 | 78.8 | 81.4
0.1 | 60.56 | 66.12 | 74.7333 | 82 | 84.4
0.2 | 67.9 | 71.9 | 80.48 | 86.54 | 88.3
0.3 | 77.9167 | 81.475 | 87.38 | 91.76 | 93.1
0.4 | 85.94 | 88.7 | 92.48 | 95.58 | 96.36
0.5 | 91.2429 | 93.08 | 95.74 | 97.84 | 98.24
0.6 | 94.7 | 96.05 | 97.92 | 99.06 | 99.3667
0.7 | 97.3143 | 98.025 | 99.3 | 99.65 | 99.92
0.8 | 98.725 | 99.25 | 99.8 | 100 | 100
0.9 | 99.6143 | 99.85 | 100 | 100 | 100
1 | 100 | 100 | 100 | 100 | 99.94
1.1 | 100 | 100 | 99.95 | 99.7 | 99.68
1.2 | 99.975 | 99.9 | 99.72 | 99.4 | 99.34
1.3 | 99.7444 | 99.62 | 99.2 | 99 | 98.92
1.4 | 99.35 | 99.3429 | 98.8 | 98.7 | 98.58
1.5 | 98.8857 | 98.9 | 98.3 | 98.2 | 98.25
1.6 | 98.3 | 98.4667 | 97.8 | 97.8 | 97.7
1.7 | 97.725 | 97.8714 | 97.38 | 97.3 | 97.28
1.8 | 97.24 | 97.4429 | 97 | 97 | 96.92
1.9 | 96.5857 | 97 | 96.54 | 96.52 | 96.42
2 | 96 | 96.4667 | 96.1 | 96.1333 | 96.02
2.1 | 95.42 | 96 | 95.7 | 95.65 | 95.68
2.2 | 94.8286 | 95.34 | 95.1 | 95.25 | 95.22
2.3 | 94.15 | 94.9 | 94.74 | 94.8 | 94.72
2.4 | 93.55 | 94.5 | 94.16 | 94.3 | 94.32
2.5 | 92.94 | 93.95 | 93.84 | 93.8 | 93.8
2.6 | 92.38 | 93.35 | 93.34 | 93.45 | 93.44
2.7 | 91.76 | 92.84 | 92.84 | 92.9667 | 92.9333
2.8 | 91.18 | 92.4 | 92.36 | 92.62 | 92.56
2.9 | 90.58 | 91.85 | 91.94 | 92.05 | 92.16
3 | 89.94 | 91.32 | 91.54 | 91.6667 | 91.76
3.1 | 89.36 | 90.85 | 90.96 | 91.2333 | 91.3333
3.2 | 88.7 | 90.36 | 90.55 | 90.8667 | 90.84
3.3 | 88.14 | 89.84 | 90.02 | 90.35 | 90.44
3.4 | 87.5333 | 89.38 | 89.6333 | 89.98 | 89.98
3.5 | 87 | 88.95 | 89.25 | 89.52 | 89.56
3.6 | 86.3 | 88.36 | 88.84 | 89.08 | 89.16
3.7 | 85.9 | 87.86 | 88.28 | 88.68 | 88.74
3.8 | 85.16 | 87.36 | 87.88 | 88.28 | 88.3
3.9 | 84.62 | 86.88 | 87.32 | 87.85 | 87.96
4 | 84.02 | 86.4 | 87 | 87.5 | 87.5
4.1 | 83.4 | 85.86 | 86.46 | 86.9667 | 87.1
4.2 | 82.75 | 85.45 | 86.1 | 86.5 | 86.6
4.3 | 82.2 | 84.9 | 85.625 | 86.1 | 86.3
4.4 | 81.8 | 84.4 | 85.15 | 85.6 | 85.8
4.5 | 81.14 | 83.9 | 84.625 | 85.3 | 85.4
4.6 | 80.7 | 83.34 | 84.26 | 84.8 | 85
4.7 | 79.98 | 82.95 | 83.8 | 84.5 | 84.5333
4.8 | 79.5 | 82.4125 | 83.38 | 83.95 | 84.1333
4.9 | 78.9286 | 81.9 | 82.9 | 83.54 | 83.675
5 | 78.38 | 81.45 | 82.5333 | 83.14 | 83.28
5.1 | 77.8 | 80.9571 | 82.05 | 82.7333 | 82.9
5.2 | 77.18 | 80.4 | 81.7 | 82.36 | 82.4667
5.3 | 76.68 | 80.0429 | 81.2333 | 82.0091 | 82
5.4 | 76.25 | 79.54 | 80.7 | 81.52 | 81.6333
5.5 | 75.675 | 79.1 | 80.2 | 81.12 | 81.2333
5.6 | 75.12 | 78.6429 | 79.85 | 80.6571 | 80.7667
5.7 | 74.65 | 78.1 | 79.5333 | 80.32 | 80.45
5.8 | 74.0333 | 77.65 | 79.1 | 79.8 | 79.9333
5.9 | 73.4556 | 77.2 | 78.6667 | 79.4 | 79.6
6 | 73.04 | 76.68 | 78.2667 | 79 | 79.2333
6.1 | 72.4286 | 76.25 | 77.72 | 78.62 | 78.7333
6.2 | 71.95 | 75.74 | 77.22 | 78.26 | 78.44
6.3 | 71.3 | 75.3333 | 76.86 | 77.86 | 78.04
6.4 | 70.74 | 74.82 | 76.42 | 77.4 | 77.62
6.5 | 70.275 | 74.24 | 76.02 | 76.96 | 77.24
6.6 | 69.85 | 73.85 | 75.5333 | 76.575 | 76.84
6.7 | 69.36 | 73.42 | 75.24 | 76.2 | 76.44
6.8 | 68.7 | 72.9 | 74.86 | 75.82 | 76.0667
6.9 | 68.36 | 72.45 | 74.36 | 75.3333 | 75.64
7 | 67.9 | 72.02 | 73.98 | 75.06 | 75.16
7.1 | 67.34 | 71.56 | 73.54 | 74.6 | 74.82
7.2 | 66.8 | 71.1625 | 73.0667 | 74.2333 | 74.48
7.3 | 66.34 | 70.72 | 72.6667 | 73.86 | 74.02
7.4 | 65.85 | 70.2286 | 72.3 | 73.46 | 73.725
7.5 | 65.35 | 69.8 | 71.9 | 73.06 | 73.3
7.6 | 65.02 | 69.45 | 71.5 | 72.58 | 72.9
7.7 | 64.38 | 68.98 | 71 | 72.24 | 72.5
7.8 | 63.94 | 68.4714 | 70.6 | 71.85 | 72.1
7.9 | 63.48 | 68.1286 | 70.3 | 71.4667 | 71.8
8 | 63.02 | 67.6556 | 70.02 | 71.1 | 71.42
8.1 | 62.4 | 67.22 | 69.54 | 70.7333 | 71
8.2 | 62.1 | 66.6714 | 69.12 | 70.34 | 70.6
8.3 | 61.8 | 66.3286 | 68.74 | 70.02 | 70.3
8.4 | 61.25 | 66 | 68.3333 | 69.64 | 69.9
8.5 | 60.76 | 65.48 | 67.975 | 69.28 | 69.54
8.6 | 60.5333 | 65.05 | 67.575 | 68.95 | 69.24
8.7 | 60 | 64.65 | 67.2 | 68.56 | 68.84
8.8 | 59.4667 | 64.24 | 66.85 | 68.2286 | 68.44
8.9 | 59.04 | 63.7857 | 66.5 | 67.7667 | 68.12
9 | 58.6 | 63.42 | 66.1 | 67.4143 | 67.74
9.1 | 58.25 | 63 | 65.7 | 67.0714 | 67.38
9.2 | 57.9 | 62.62 | 65.3 | 66.75 | 67.0429
9.3 | 57.48 | 62.2 | 64.9 | 66.36 | 66.64
9.4 | 57.05 | 61.8333 | 64.6 | 65.98 | 66.3545
9.5 | 56.58 | 61.4 | 64.2 | 65.6667 | 65.98
9.6 | 56.2 | 61.05 | 63.7667 | 65.3 | 65.6
9.7 | 55.78 | 60.62 | 63.42 | 65 | 65.3
9.8 | 55.4333 | 60.3 | 63.125 | 64.6 | 64.9
9.9 | 54.98 | 59.8667 | 62.75 | 64.2 | 64.5
10 | 54.6333 | 59.56 | 62.38 | 63.9 | 64.24
10.1 | 54.16 | 59.16 | 62.0667 | 63.5 | 63.94
10.2 | 53.7667 | 58.68 | 61.725 | 63.22 | 63.54
10.3 | 53.4 | 58.3667 | 61.46 | 62.9 | 63.22
10.4 | 52.96 | 57.88 | 61.0667 | 62.55 | 62.92
10.5 | 52.58 | 57.54 | 60.56 | 62.16 | 62.54
10.6 | 52.175 | 57.15 | 60.3 | 61.925 | 62.1667
10.7 | 51.9 | 56.8 | 59.8833 | 61.4 | 61.84
10.8 | 51.5 | 56.46 | 59.6 | 61.1333 | 61.54
10.9 | 51.1 | 56.0667 | 59.2667 | 60.8333 | 61.18
11 | 50.8 | 55.725 | 58.9 | 60.4667 | 60.88
11.1 | 50.3667 | 55.3667 | 58.5 | 60.14 | 60.56
11.2 | 49.9 | 55 | 58.2 | 59.84 | 60.22
11.3 | 49.6714 | 54.7 | 57.9 | 59.52 | 59.92
11.4 | 49.3286 | 54.32 | 57.5286 | 59.14 | 59.62
11.5 | 48.88 | 53.9333 | 57.1429 | 58.84 | 59.26
11.6 | 48.5 | 53.56 | 56.8636 | 58.56 | 59.02
11.7 | 48.2571 | 53.25 | 56.54 | 58.26 | 58.62
11.8 | 47.8857 | 52.92 | 56.14 | 57.95 | 58.22
11.9 | 47.5 | 52.45 | 55.78 | 57.6333 | 57.86
12 | 47.1714 | 52.22 | 55.5 | 57.3333 | 57.6667
12.1 | 46.8 | 51.8 | 55.2 | 56.9333 | 57.38
12.2 | 46.48 | 51.42 | 54.9 | 56.58 | 57.06
12.3 | 46.1 | 51.15 | 54.5 | 56.3 | 56.78
12.4 | 45.8 | 50.78 | 54.2 | 56.04 | 56.46
12.5 | 45.5333 | 50.475 | 53.9 | 55.68 | 56.12
12.6 | 45.1 | 50.24 | 53.6 | 55.38 | 55.8
12.7 | 44.86 | 49.84 | 53.3 | 55.08 | 55.4571
12.8 | 44.5 | 49.46 | 52.9 | 54.74 | 55.1625
12.9 | 44.12 | 49.15 | 52.6 | 54.52 | 54.84
13 | 43.82 | 48.75 | 52.3333 | 54.22 | 54.55
13.1 | 43.52 | 48.4 | 52.04 | 53.86 | 54.2833
13.2 | 43.225 | 48.2 | 51.72 | 53.56 | 54.1
13.3 | 42.94 | 47.9143 | 51.42 | 53.22 | 53.7
13.4 | 42.6 | 47.55 | 51.12 | 52.96 | 53.4
13.5 | 42.275 | 47.24 | 50.84 | 52.68 | 53.2
13.6 | 42 | 46.9 | 50.5 | 52.46 | 52.9
13.7 | 41.68 | 46.5333 | 50.26 | 52.08 | 52.5
13.8 | 41.4556 | 46.275 | 49.8667 | 51.78 | 52.2
13.9 | 41.15 | 46.0333 | 49.6 | 51.48 | 51.9
14 | 40.8667 | 45.74 | 49.4 | 51.34 | 51.7
14.1 | 40.6 | 45.45 | 49.04 | 50.98 | 51.3
14.2 | 40.275 | 45.125 | 48.75 | 50.74 | 51.12
14.3 | 40 | 44.8857 | 48.42 | 50.38 | 50.92
14.4 | 39.675 | 44.54 | 48.16 | 50.14 | 50.62
14.5 | 39.4 | 44.3333 | 47.92 | 49.94 | 50.32
14.6 | 39.175 | 44 | 47.6333 | 49.54 | 50.04
14.7 | 38.8 | 43.5875 | 47.35 | 49.28 | 49.74
14.8 | 38.58 | 43.38 | 47.02 | 48.98 | 49.52
14.9 | 38.35 | 43.0857 | 46.72 | 48.7 | 49.24
15 | 38.05 | 42.8 | 46.46 | 48.46 | 49.02
15.1 | 37.7714 | 42.5625 | 46.16 | 48.22 | 48.64
15.2 | 37.56 | 42.24 | 45.9667 | 47.96 | 48.42
15.3 | 37.25 | 41.9857 | 45.68 | 47.66 | 48.14
15.4 | 37 | 41.7 | 45.4333 | 47.36 | 47.82
15.5 | 36.72 | 41.4667 | 45.08 | 47.12 | 47.62
15.6 | 36.425 | 41.15 | 44.86 | 46.82 | 47.4
15.7 | 36.1714 | 40.9143 | 44.58 | 46.58 | 47.0667
15.8 | 35.9 | 40.6286 | 44.325 | 46.28 | 46.82
15.9 | 35.68 | 40.3 | 44.0667 | 46.08 | 46.64
16 | 35.4667 | 40.1 | 43.775 | 45.88 | 46.34
16.1 | 35.1 | 39.85 | 43.525 | 45.66 | 46.04
16.2 | 34.9667 | 39.55 | 43.3333 | 45.36 | 45.84
16.3 | 34.7 | 39.35 | 43.05 | 45.08 | 45.58
16.4 | 34.4667 | 39.0714 | 42.8333 | 44.86 | 45.3
16.5 | 34.3 | 38.82 | 42.5333 | 44.48 | 45.08
16.6 | 33.9667 | 38.5571 | 42.3 | 44.28 | 44.84
16.7 | 33.74 | 38.2667 | 42.0333 | 44.1 | 44.58
16.8 | 33.54 | 38.08 | 41.85 | 43.9 | 44.34
16.9 | 33.3 | 37.76 | 41.55 | 43.6 | 44.15
17 | 33.0286 | 37.55 | 41.3333 | 43.3 | 43.85
17.1 | 32.82 | 37.3 | 41 | 43.1 | 43.56
17.2 | 32.5889 | 37.12 | 40.8 | 42.8 | 43.3
17.3 | 32.36 | 36.7 | 40.5 | 42.5857 | 43.1
17.4 | 32.0714 | 36.52 | 40.2 | 42.32 | 42.8
17.5 | 31.82 | 36.3333 | 40.025 | 42.04 | 42.6
17.6 | 31.7 | 36.12 | 39.7833 | 41.92 | 42.4
17.7 | 31.425 | 35.8 | 39.525 | 41.62 | 42.1
17.8 | 31.2 | 35.5667 | 39.25 | 41.44 | 41.9
17.9 | 30.95 | 35.35 | 39.0667 | 41.1571 | 41.7
18 | 30.8 | 35.1 | 38.8 | 40.88 | 41.44
18.1 | 30.54 | 34.88 | 38.6 | 40.58 | 41.2
18.2 | 30.4 | 34.68 | 38.375 | 40.4 | 40.95
18.3 | 30.18 | 34.38 | 38.18 | 40.18 | 40.7
18.4 | 29.96 | 34.3 | 37.9333 | 39.98 | 40.5167
18.5 | 29.75 | 34 | 37.7 | 39.8667 | 40.32
18.6 | 29.525 | 33.825 | 37.52 | 39.6 | 40.1
18.7 | 29.3 | 33.6 | 37.2 | 39.3 | 39.9167
18.8 | 29.125 | 33.4333 | 37.125 | 39.1 | 39.72
18.9 | 28.9 | 33.1167 | 36.82 | 39 | 39.42
19 | 28.68 | 32.94 | 36.62 | 38.7 | 39.22
19.1 | 28.46 | 32.7 | 36.42 | 38.5 | 39.02
19.2 | 28.3 | 32.5 | 36.15 | 38.2 | 38.72
19.3 | 28.12 | 32.32 | 35.96 | 38 | 38.52
19.4 | 27.925 | 32.05 | 35.7 | 37.8 | 38.32
19.5 | 27.72 | 31.825 | 35.56 | 37.6 | 38.12
19.6 | 27.5 | 31.65 | 35.28 | 37.4 | 37.92
19.7 | 27.2857 | 31.4714 | 35.08 | 37.1333 | 37.72
19.8 | 27.15 | 31.2 | 34.88 | 37.02 | 37.5273
19.9 | 26.9625 | 31.0857 | 34.68 | 36.7 | 37.3143
20 | 26.8 | 30.9 | 34.48 | 36.6 | 37.1143
20.1 | 26.6143 | 30.6857 | 34.28 | 36.36 | 36.94
20.2 | 26.4 | 30.4857 | 34.08 | 36.16 | 36.68
20.3 | 26.2286 | 30.24 | 33.8667 | 35.925 | 36.44
20.4 | 26.06 | 30.0571 | 33.6 | 35.72 | 36.27
20.5 | 25.8625 | 29.85 | 33.5 | 35.55 | 36.1182
20.6 | 25.7 | 29.7 | 33.3 | 35.4 | 35.86
20.7 | 25.5167 | 29.52 | 33 | 35.22 | 35.66
20.8 | 25.3 | 29.2571 | 32.9 | 35.02 | 35.48
20.9 | 25.175 | 29.15 | 32.6 | 34.72 | 35.34
21 | 25.02 | 28.92 | 32.5 | 34.5333 | 35.06
21.1 | 24.8286 | 28.7 | 32.3 | 34.3 | 34.86
21.2 | 24.68 | 28.48 | 32.12 | 34.2 | 34.66
21.3 | 24.46 | 28.35 | 31.92 | 33.95 | 34.45
21.4 | 24.28 | 28.14 | 31.72 | 33.7667 | 34.3
21.5 | 24.1 | 27.9 | 31.52 | 33.52 | 34.1429
21.6 | 23.9 | 27.78 | 31.32 | 33.4 | 33.98
21.7 | 23.8 | 27.5667 | 31.14 | 33.24 | 33.78
21.8 | 23.7125 | 27.44 | 31.02 | 33.04 | 33.58
21.9 | 23.52 | 27.3 | 30.8 | 32.825 | 33.38
22 | 23.34 | 27.1 | 30.55 | 32.68 | 33.2
22.1 | 23.12 | 26.8 | 30.34 | 32.4833 | 33
22.2 | 23.02 | 26.6857 | 30.175 | 32.3 | 32.8
22.3 | 22.8 | 26.52 | 29.97 | 32.0833 | 32.6
22.4 | 22.66 | 26.35 | 29.84 | 31.9 | 32.5
22.5 | 22.48 | 26.1571 | 29.64 | 31.7333 | 32.3167
22.6 | 22.3143 | 25.9875 | 29.45 | 31.6 | 32.0714
22.7 | 22.2 | 25.84 | 29.3556 | 31.34 | 31.9
22.8 | 22.0875 | 25.72 | 29.15 | 31.2 | 31.72
22.9 | 21.925 | 25.6 | 28.96 | 31.1 | 31.575
23 | 21.775 | 25.34 | 28.76 | 30.875 | 31.325
23.1 | 21.6 | 25.14 | 28.625 | 30.65 | 31.2
23.2 | 21.5 | 25 | 28.5 | 30.56 | 30.9833
23.3 | 21.3667 | 24.9429 | 28.325 | 30.36 | 30.8
23.4 | 21.2 | 24.78 | 28.14 | 30.16 | 30.6
23.5 | 21 | 24.55 | 27.9667 | 30.02 | 30.5
23.6 | 20.88 | 24.425 | 27.82 | 29.86 | 30.3
23.7 | 20.7667 | 24.2667 | 27.6 | 29.66 | 30.2
23.8 | 20.64 | 24.1 | 27.475 | 29.46 | 30.0333
23.9 | 20.5 | 23.9429 | 27.32 | 29.26 | 29.9
24 | 20.3 | 23.75 | 27.1 | 29.16 | 29.7
24.1 | 20.2 | 23.65 | 26.98 | 28.96 | 29.5
24.2 | 20.1 | 23.5 | 26.78 | 28.78 | 29.4
24.3 | 19.94 | 23.2857 | 26.625 | 28.68 | 29.1333
24.4 | 19.8 | 23.2 | 26.48 | 28.48 | 29
24.5 | 19.6 | 23 | 26.36 | 28.38 | 28.82
24.6 | 19.5 | 22.9 | 26.1 | 28.2 | 28.78
24.7 | 19.5 | 22.76 | 26 | 28.08 | 28.58
24.8 | 19.225 | 22.55 | 25.88 | 27.88 | 28.38
24.9 | 19.1286 | 22.5 | 25.68 | 27.7333 | 28.28
25 | 19 | 22.3167 | 25.5 | 27.58 | 28.08
25.1 | 18.9 | 22.2 | 25.4 | 27.38 | 27.88
25.2 | 18.775 | 22 | 25.2 | 27.24 | 27.78
25.3 | 18.6 | 21.9 | 25.1 | 27.08 | 27.6
25.4 | 18.5 | 21.7333 | 25 | 26.88 | 27.44
25.5 | 18.4 | 21.65 | 24.8 | 26.74 | 27.3
25.6 | 18.2 | 21.45 | 24.7 | 26.64 | 27.14
25.7 | 18.12 | 21.32 | 24.5 | 26.48 | 26.94
25.8 | 18 | 21.2 | 24.4 | 26.34 | 26.8
25.9 | 17.8714 | 21.1 | 24.2 | 26.18 | 26.64
26 | 17.8 | 20.96 | 24.1 | 26.1 | 26.55
26.1 | 17.7 | 20.8 | 24 | 25.89 | 26.36
26.2 | 17.5286 | 20.7 | 23.8 | 25.7818 | 26.2
26.3 | 17.425 | 20.575 | 23.7 | 25.6 | 26.1
26.4 | 17.3 | 20.38 | 23.5 | 25.4 | 25.96
26.5 | 17.2 | 20.2667 | 23.42 | 25.3 | 25.86
26.6 | 17.05 | 20.1 | 23.22 | 25.2 | 25.7
26.7 | 16.975 | 20 | 23.12 | 25 | 25.56
26.8 | 16.8 | 19.9 | 23 | 24.9 | 25.36
26.9 | 16.74 | 19.74 | 22.82 | 24.8 | 25.2
27 | 16.6556 | 19.7 | 22.72 | 24.6 | 25.1167
27.1 | 16.48 | 19.5167 | 22.6 | 24.46 | 24.9
27.2 | 16.4 | 19.3714 | 22.42 | 24.3 | 24.82
27.3 | 16.3 | 19.2286 | 22.3 | 24.18 | 24.7429
27.4 | 16.2 | 19.1 | 22.2 | 24.08 | 24.55
27.5 | 16.1 | 19.0286 | 22.02 | 23.98 | 24.4167
27.6 | 15.9625 | 18.925 | 21.92 | 23.8 | 24.31
27.7 | 15.9 | 18.78 | 21.82 | 23.68 | 24.1
27.8 | 15.7833 | 18.6833 | 21.6333 | 23.48 | 24.0143
27.9 | 15.675 | 18.6 | 21.5 | 23.38 | 23.91
28 | 15.5857 | 18.475 | 21.34 | 23.28 | 23.81
28.1 | 15.42 | 18.3 | 21.24 | 23.1 | 23.6
28.2 | 15.34 | 18.1857 | 21.1 | 23 | 23.5143
28.3 | 15.2 | 18.1 | 21 | 22.8 | 23.32
28.4 | 15.15 | 17.94 | 20.88 | 22.7 | 23.21
28.5 | 15.1 | 17.8556 | 20.78 | 22.6167 | 23.11
28.6 | 14.9 | 17.7 | 20.68 | 22.5 | 22.92
28.7 | 14.8167 | 17.6143 | 20.58 | 22.3 | 22.81
28.8 | 14.7 | 17.5 | 20.4 | 22.2 | 22.7182
28.9 | 14.6167 | 17.425 | 20.3 | 22.1 | 22.62
29 | 14.56 | 17.3 | 20.2 | 22 | 22.44
29.1 | 14.475 | 17.1 | 20.1 | 21.9 | 22.32
29.2 | 14.3 | 17.0375 | 19.9 | 21.7 | 22.22
29.3 | 14.2167 | 16.94 | 19.8 | 21.6 | 22.12
29.4 | 14.2 | 16.9 | 19.7 | 21.5 | 22.02
29.5 | 14.04 | 16.7857 | 19.6 | 21.4 | 21.84
29.6 | 13.96 | 16.6429 | 19.4 | 21.2 | 21.72
29.7 | 13.9 | 16.54 | 19.3 | 21.1167 | 21.62
29.8 | 13.8 | 16.4 | 19.2375 | 21 | 21.44
29.9 | 13.7 | 16.4 | 19.16 | 20.9 | 21.3273
30 | 13.6 | 16.2 | 19 | 20.8 | 21.23

## SSD120 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 57.5 | 62 | 68.3 | 73.5 | 75.5
0.1 | 62.6 | 67.1 | 73 | 77.725 | 79.4
0.2 | 71.8 | 74.9333 | 80.1 | 83.5 | 85.1
0.3 | 80.7 | 83.125 | 86.3375 | 89.2 | 90.15
0.4 | 87.7 | 89.35 | 91.25 | 93.5 | 94.1
0.5 | 92.4 | 93.475 | 94.9375 | 96.25 | 96.65
0.6 | 95.4 | 96.1 | 97.2 | 98.2 | 98.2333
0.7 | 97.6 | 97.7 | 98.7 | 99.05 | 99.1
0.8 | 99 | 99 | 99.3875 | 99.8 | 99.7
0.9 | 99.65 | 99.4667 | 99.8 | 99.9333 | 100
1 | 100 | 99.9667 | 100 | 100 | 100
1.1 | 100 | 100 | 100 | 99.9 | 100
1.2 | 100 | 100 | 99.975 | 99.875 | 99.85
1.3 | 99.82 | 99.8 | 99.8 | 99.62 | 99.6
1.4 | 99.6 | 99.4 | 99.475 | 99.4 | 99.3667
1.5 | 99.175 | 99.125 | 99.25 | 99.15 | 99.075
1.6 | 98.675 | 98.95 | 98.975 | 98.875 | 98.825
1.7 | 98.25 | 98.575 | 98.55 | 98.625 | 98.475
1.8 | 97.85 | 98.1667 | 98.275 | 98.38 | 98.125
1.9 | 97.425 | 97.725 | 97.95 | 97.95 | 97.85
2 | 96.8 | 97.3667 | 97.6 | 97.58 | 97.65
2.1 | 96.525 | 96.925 | 97.225 | 97.325 | 97.3333
2.2 | 96.05 | 96.55 | 96.9 | 97.0333 | 97
2.3 | 95.58 | 96.3 | 96.55 | 96.7 | 96.6
2.4 | 95 | 95.8 | 96.1333 | 96.3 | 96.3
2.5 | 94.55 | 95.3 | 95.9 | 96 | 96.05
2.6 | 94.2 | 95.1 | 95.4857 | 95.85 | 95.8
2.7 | 93.5 | 94.5667 | 95.3 | 95.5 | 95.45
2.8 | 93.1 | 94.3 | 94.975 | 95.1 | 95.125
2.9 | 92.5 | 93.7667 | 94.6 | 94.8 | 94.8667
3 | 92.1 | 93.3333 | 94.15 | 94.45 | 94.44
3.1 | 91.6 | 93.1 | 93.825 | 94.1 | 94.1333
3.2 | 91 | 92.6667 | 93.58 | 93.775 | 93.85
3.3 | 90.6 | 92.05 | 93.0833 | 93.425 | 93.5
3.4 | 90.05 | 91.7 | 92.76 | 93.075 | 93.075
3.5 | 89.625 | 91.4 | 92.3429 | 92.725 | 92.825
3.6 | 89.125 | 91 | 92 | 92.45 | 92.45
3.7 | 88.65 | 90.55 | 91.65 | 92.125 | 92.175
3.8 | 88.125 | 90.1 | 91.4 | 91.75 | 91.8
3.9 | 87.55 | 89.75 | 91 | 91.35 | 91.5
4 | 87.12 | 89.34 | 90.7 | 90.98 | 91.1
4.1 | 86.4 | 89 | 90.175 | 90.6667 | 90.8
4.2 | 86 | 88.45 | 89.8 | 90.18 | 90.5
4.3 | 85.5 | 88.05 | 89.4 | 90 | 90.1
4.4 | 85 | 87.725 | 89.16 | 89.6667 | 89.7
4.5 | 84.5 | 87.275 | 88.7333 | 89.2 | 89.4
4.6 | 84.1 | 86.925 | 88.425 | 88.925 | 88.95
4.7 | 83.5 | 86.45 | 87.95 | 88.6 | 88.7
4.8 | 83.1 | 86.125 | 87.48 | 88.2 | 88.2667
4.9 | 82.6 | 85.7 | 87.175 | 88 | 87.94
5 | 82 | 85.18 | 86.8875 | 87.5 | 87.7
5.1 | 81.6333 | 84.8 | 86.5125 | 87.25 | 87.525
5.2 | 81.25 | 84.4 | 86.225 | 86.9 | 87.175
5.3 | 80.72 | 84 | 85.875 | 86.55 | 86.75
5.4 | 80.275 | 83.75 | 85.48 | 86.4 | 86.375
5.5 | 79.65 | 83.3 | 85.1667 | 85.9667 | 86.0667
5.6 | 79.15 | 82.9 | 84.8429 | 85.72 | 85.775
5.7 | 78.75 | 82.4 | 84.5 | 85.375 | 85.3333
5.8 | 78.225 | 82.05 | 84 | 84.95 | 84.9333
5.9 | 77.75 | 81.7 | 83.7 | 84.675 | 84.7
6 | 77.35 | 81.2 | 83.4333 | 84.275 | 84.4
6.1 | 76.875 | 80.725 | 83.15 | 83.9 | 84.1
6.2 | 76.45 | 80.475 | 82.8 | 83.55 | 83.75
6.3 | 76 | 80.15 | 82.3667 | 83.3 | 83.4667
6.4 | 75.6 | 79.65 | 81.9 | 82.825 | 83.2
6.5 | 75.1 | 79.25 | 81.62 | 82.5 | 82.7
6.6 | 74.5 | 78.7 | 81.2 | 82.24 | 82.45
6.7 | 74 | 78.3333 | 80.85 | 82 | 82.05
6.8 | 73.55 | 78.1 | 80.5 | 81.7 | 81.775
6.9 | 73.1 | 77.55 | 80.2667 | 81.2 | 81.35
7 | 72.82 | 77.2 | 79.9 | 80.9 | 80.975
7.1 | 72.35 | 76.8 | 79.55 | 80.4 | 80.725
7.2 | 71.85 | 76.4 | 79.125 | 80.16 | 80.475
7.3 | 71.4667 | 76.1 | 78.675 | 80 | 80.15
7.4 | 71.075 | 75.8 | 78.425 | 79.65 | 79.65
7.5 | 70.575 | 75.3 | 78.14 | 79.15 | 79.425
7.6 | 70.05 | 74.9 | 77.8 | 78.85 | 79.075
7.7 | 69.75 | 74.5 | 77.5 | 78.55 | 78.8
7.8 | 69.175 | 74.1333 | 77.05 | 78.22 | 78.5
7.9 | 68.875 | 73.9 | 76.75 | 77.875 | 78.15
8 | 68.375 | 73.54 | 76.4 | 77.625 | 77.9
8.1 | 67.9 | 73.025 | 75.9667 | 77.275 | 77.45
8.2 | 67.45 | 72.7 | 75.6 | 76.95 | 77.1
8.3 | 67.1 | 72.3 | 75.3625 | 76.675 | 76.9
8.4 | 66.6 | 71.875 | 74.9875 | 76.325 | 76.5
8.5 | 66.15 | 71.625 | 74.675 | 76.1 | 76.25
8.6 | 65.7667 | 71.225 | 74.425 | 75.7 | 75.9
8.7 | 65.5 | 70.825 | 74.075 | 75.5 | 75.5
8.8 | 65 | 70.45 | 73.7714 | 75.1 | 75.2
8.9 | 64.65 | 70.05 | 73.3 | 74.7 | 74.85
9 | 64.2 | 69.55 | 72.9 | 74.35 | 74.6667
9.1 | 63.8 | 69.24 | 72.55 | 74.1 | 74.25
9.2 | 63.45 | 68.88 | 72.2 | 73.75 | 74
9.3 | 62.85 | 68.6 | 71.875 | 73.4 | 73.725
9.4 | 62.575 | 68.1333 | 71.625 | 73.0667 | 73.25
9.5 | 62.225 | 67.7667 | 71.26 | 72.74 | 72.95
9.6 | 61.8 | 67.4 | 70.925 | 72.275 | 72.625
9.7 | 61.325 | 67.1 | 70.62 | 72.025 | 72.3
9.8 | 60.95 | 66.8 | 70.325 | 71.775 | 72.025
9.9 | 60.58 | 66.4 | 69.95 | 71.425 | 71.675
10 | 60.175 | 65.9 | 69.6333 | 71.175 | 71.475
10.1 | 59.8333 | 65.6 | 69.3667 | 70.8667 | 71.125
10.2 | 59.48 | 65.3 | 69.05 | 70.6 | 70.775
10.3 | 59 | 64.9 | 68.7 | 70.1667 | 70.5
10.4 | 58.76 | 64.55 | 68.45 | 69.925 | 70.175
10.5 | 58.275 | 64.2 | 68.1 | 69.65 | 69.925
10.6 | 58 | 63.95 | 67.85 | 69.3333 | 69.6
10.7 | 57.675 | 63.6 | 67.5 | 68.98 | 69.325
10.8 | 57.3 | 63.225 | 67.0667 | 68.725 | 68.975
10.9 | 56.875 | 62.85 | 66.8 | 68.4 | 68.65
11 | 56.45 | 62.55 | 66.45 | 68.05 | 68.35
11.1 | 56.075 | 62.125 | 66.15 | 67.8 | 68.04
11.2 | 55.725 | 61.775 | 65.85 | 67.5 | 67.8
11.3 | 55.375 | 61.45 | 65.6 | 67.25 | 67.45
11.4 | 55.025 | 61.1 | 65.1857 | 67 | 67.2
11.5 | 54.8 | 60.85 | 64.9222 | 66.6 | 66.85
11.6 | 54.4 | 60.55 | 64.6 | 66.3 | 66.7
11.7 | 54.05 | 60.05 | 64.2667 | 66.1 | 66.45
11.8 | 53.65 | 59.8 | 63.9 | 65.7 | 66.1
11.9 | 53.35 | 59.58 | 63.66 | 65.35 | 65.85
12 | 53.05 | 59.15 | 63.425 | 65.1 | 65.52
12.1 | 52.7 | 58.95 | 63.075 | 64.8 | 65.15
12.2 | 52.35 | 58.6 | 62.8333 | 64.5 | 64.8
12.3 | 52.1 | 58.24 | 62.58 | 64.26 | 64.56
12.4 | 51.7 | 57.95 | 62.3 | 63.9 | 64.325
12.5 | 51.3 | 57.575 | 61.84 | 63.65 | 64.05
12.6 | 51.0667 | 57.3 | 61.475 | 63.425 | 63.725
12.7 | 50.6 | 56.9333 | 61.225 | 63.15 | 63.5
12.8 | 50.475 | 56.7 | 61 | 62.825 | 63.25
12.9 | 50.125 | 56.4 | 60.68 | 62.5667 | 62.875
13 | 49.76 | 56.15 | 60.4 | 62.26 | 62.675
13.1 | 49.5 | 55.7 | 60.075 | 61.9286 | 62.4
13.2 | 49.15 | 55.4 | 59.825 | 61.6833 | 62.0333
13.3 | 48.85 | 55 | 59.55 | 61.4833 | 61.75
13.4 | 48.475 | 54.8 | 59.24 | 61.15 | 61.425
13.5 | 48.125 | 54.4 | 58.98 | 60.86 | 61.175
13.6 | 47.85 | 54.1333 | 58.7 | 60.5 | 60.94
13.7 | 47.525 | 53.9 | 58.3 | 60.28 | 60.8
13.8 | 47.25 | 53.5333 | 58.05 | 60.05 | 60.4
13.9 | 46.925 | 53.2 | 57.7667 | 59.6 | 60.1
14 | 46.66 | 52.875 | 57.4667 | 59.35 | 59.8
14.1 | 46.275 | 52.525 | 57.1667 | 59.15 | 59.5667
14.2 | 46.1 | 52.275 | 56.8333 | 58.9 | 59.3
14.3 | 45.7429 | 52.025 | 56.4875 | 58.55 | 59
14.4 | 45.4 | 51.8 | 56.325 | 58.3 | 58.85
14.5 | 45.15 | 51.55 | 56.075 | 58.05 | 58.54
14.6 | 44.9 | 51.175 | 55.825 | 57.8 | 58.2
14.7 | 44.5 | 50.925 | 55.5714 | 57.4667 | 57.975
14.8 | 44.3 | 50.575 | 55.3 | 57.2 | 57.6667
14.9 | 44 | 50.325 | 54.95 | 57.05 | 57.4
15 | 43.68 | 49.98 | 54.64 | 56.725 | 57.075
15.1 | 43.35 | 49.7333 | 54.4 | 56.525 | 56.9
15.2 | 43.12 | 49.5 | 54.2 | 56.275 | 56.725
15.3 | 42.8667 | 49.15 | 53.9 | 55.95 | 56.4
15.4 | 42.575 | 48.9 | 53.65 | 55.7 | 56.075
15.5 | 42.3 | 48.65 | 53.4 | 55.4333 | 55.825
15.6 | 42.15 | 48.3 | 53.1 | 55.1333 | 55.65
15.7 | 41.775 | 48.1 | 52.85 | 54.875 | 55.325
15.8 | 41.525 | 47.9 | 52.575 | 54.6 | 55.175
15.9 | 41.275 | 47.55 | 52.4 | 54.4 | 54.85
16 | 41.075 | 47.3 | 52.15 | 54.2 | 54.575
16.1 | 40.8667 | 46.95 | 51.825 | 53.95 | 54.4
16.2 | 40.5333 | 46.7 | 51.55 | 53.6 | 54.1
16.3 | 40.2857 | 46.45 | 51.325 | 53.5 | 53.8
16.4 | 40.0444 | 46.22 | 50.975 | 53.2 | 53.5667
16.5 | 39.775 | 45.975 | 50.8 | 52.95 | 53.4
16.6 | 39.525 | 45.725 | 50.475 | 52.6 | 53.05
16.7 | 39.275 | 45.5 | 50.2333 | 52.4 | 52.9
16.8 | 39.04 | 45.225 | 50 | 52.125 | 52.55
16.9 | 38.8 | 44.9667 | 49.75 | 51.925 | 52.3
17 | 38.55 | 44.7 | 49.5 | 51.7333 | 52.1
17.1 | 38.35 | 44.5 | 49.3 | 51.425 | 51.9
17.2 | 38.05 | 44.05 | 49.08 | 51.2 | 51.75
17.3 | 37.8 | 43.9 | 48.8 | 51 | 51.4
17.4 | 37.4 | 43.625 | 48.5 | 50.7 | 51.15
17.5 | 37.25 | 43.375 | 48.25 | 50.5 | 50.9
17.6 | 37.1333 | 43.2 | 48.075 | 50.2 | 50.76
17.7 | 36.925 | 43 | 47.85 | 50 | 50.5
17.8 | 36.675 | 42.725 | 47.6 | 49.75 | 50.1667
17.9 | 36.425 | 42.5 | 47.35 | 49.5 | 50
18 | 36.2167 | 42.25 | 47.1 | 49.3 | 49.86
18.1 | 35.9625 | 41.9 | 46.84 | 49 | 49.575
18.2 | 35.775 | 41.75 | 46.6625 | 48.8 | 49.325
18.3 | 35.4375 | 41.5 | 46.325 | 48.5333 | 49.075
18.4 | 35.2 | 41.3 | 46.0889 | 48.3667 | 48.825
18.5 | 35 | 41 | 45.9 | 48.0667 | 48.5667
18.6 | 34.8625 | 40.75 | 45.75 | 47.9 | 48.375
18.7 | 34.5667 | 40.5 | 45.45 | 47.65 | 48.1
18.8 | 34.4 | 40.25 | 45.325 | 47.475 | 47.9333
18.9 | 34.175 | 40 | 44.9 | 47.175 | 47.6
19 | 33.925 | 39.8 | 44.7778 | 47 | 47.4667
19.1 | 33.6571 | 39.7 | 44.55 | 46.8 | 47.3
19.2 | 33.5 | 39.4 | 44.32 | 46.575 | 47.0333
19.3 | 33.225 | 39.125 | 43.9857 | 46.3 | 46.8
19.4 | 33.1 | 38.875 | 43.8 | 46.1 | 46.5667
19.5 | 32.95 | 38.7 | 43.6 | 45.86 | 46.3667
19.6 | 32.6333 | 38.475 | 43.35 | 45.625 | 46.1
19.7 | 32.5 | 38.225 | 43.1 | 45.4429 | 45.95
19.8 | 32.2571 | 37.975 | 43 | 45.25 | 45.7
19.9 | 32.05 | 37.8 | 42.76 | 45.05 | 45.45
20 | 31.85 | 37.675 | 42.5 | 44.7143 | 45.2333
20.1 | 31.7 | 37.45 | 42.32 | 44.55 | 45.1
20.2 | 31.45 | 37.28 | 42.1 | 44.4 | 44.95
20.3 | 31.2 | 37.05 | 41.85 | 44.22 | 44.6
20.4 | 31.05 | 36.8 | 41.5833 | 43.925 | 44.4
20.5 | 30.9 | 36.65 | 41.4167 | 43.675 | 44.2
20.6 | 30.65 | 36.38 | 41.225 | 43.525 | 44
20.7 | 30.5 | 36.1 | 41.0375 | 43.26 | 43.825
20.8 | 30.3 | 35.9 | 40.825 | 43.1 | 43.6286
20.9 | 30.1 | 35.7333 | 40.55 | 42.88 | 43.4
21 | 29.96 | 35.6 | 40.375 | 42.7143 | 43.2167
21.1 | 29.7 | 35.35 | 40.18 | 42.5 | 43.02
21.2 | 29.575 | 35.1 | 40.0143 | 42.3 | 42.775
21.3 | 29.4 | 34.9 | 39.8 | 42.125 | 42.6125
21.4 | 29.175 | 34.7 | 39.55 | 41.875 | 42.4167
21.5 | 28.94 | 34.6 | 39.325 | 41.7 | 42.15
21.6 | 28.8 | 34.4 | 39.2 | 41.475 | 41.8833
21.7 | 28.65 | 34.15 | 38.9667 | 41.225 | 41.7167
21.8 | 28.5 | 34 | 38.75 | 41.0167 | 41.55
21.9 | 28.3 | 33.7 | 38.5 | 40.925 | 41.4125
22 | 28.1 | 33.55 | 38.35 | 40.675 | 41.18
22.1 | 27.9 | 33.4 | 38.1 | 40.425 | 41.0143
22.2 | 27.75 | 33.2 | 37.975 | 40.28 | 40.7857
22.3 | 27.5 | 32.925 | 37.8 | 40.1 | 40.62
22.4 | 27.4 | 32.675 | 37.625 | 39.9 | 40.38
22.5 | 27.2 | 32.575 | 37.4 | 39.65 | 40.1667
22.6 | 27.05 | 32.4333 | 37.275 | 39.6 | 40
22.7 | 26.9 | 32.28 | 37.1 | 39.35 | 39.86
22.8 | 26.7 | 32.1 | 36.85 | 39.15 | 39.55
22.9 | 26.6 | 31.9 | 36.6 | 38.95 | 39.5
23 | 26.45 | 31.65 | 36.45 | 38.7 | 39.225
23.1 | 26.2667 | 31.52 | 36.3 | 38.5667 | 39.075
23.2 | 26.0333 | 31.3 | 36.05 | 38.3 | 38.925
23.3 | 25.8667 | 31.15 | 35.8 | 38.15 | 38.775
23.4 | 25.8 | 30.9 | 35.7 | 37.9429 | 38.525
23.5 | 25.675 | 30.8 | 35.5 | 37.75 | 38.3
23.6 | 25.5 | 30.6333 | 35.275 | 37.58 | 38.125
23.7 | 25.325 | 30.4 | 35.1 | 37.4333 | 38
23.8 | 25.2 | 30.3 | 35 | 37.2 | 37.75
23.9 | 24.975 | 30.1 | 34.8167 | 37.05 | 37.5
24 | 24.8 | 29.9 | 34.6 | 36.9125 | 37.35
24.1 | 24.7 | 29.75 | 34.475 | 36.6875 | 37.3
24.2 | 24.475 | 29.625 | 34.3 | 36.5143 | 37.1
24.3 | 24.325 | 29.5 | 34.05 | 36.3 | 36.9
24.4 | 24.2 | 29.2667 | 33.8571 | 36.2 | 36.75
24.5 | 24.05 | 29.1 | 33.6857 | 36 | 36.62
24.6 | 23.9667 | 28.86 | 33.52 | 35.8 | 36.35
24.7 | 23.8 | 28.7 | 33.36 | 35.65 | 36.2
24.8 | 23.6 | 28.6 | 33.125 | 35.5 | 36.06
24.9 | 23.5 | 28.45 | 33 | 35.3 | 35.8
25 | 23.3 | 28.3 | 32.875 | 35.125 | 35.6
25.1 | 23.15 | 28.05 | 32.6875 | 34.9444 | 35.4
25.2 | 23 | 27.92 | 32.5 | 34.825 | 35.3
25.3 | 22.9 | 27.75 | 32.35 | 34.5889 | 35.15
25.4 | 22.82 | 27.6 | 32.125 | 34.45 | 35
25.5 | 22.65 | 27.4 | 32.0857 | 34.25 | 34.8
25.6 | 22.5 | 27.3 | 31.92 | 34.125 | 34.625
25.7 | 22.375 | 27.25 | 31.6875 | 34 | 34.5
25.8 | 22.2 | 27 | 31.525 | 33.775 | 34.325
25.9 | 22.075 | 26.9 | 31.4 | 33.65 | 34.175
26 | 21.925 | 26.75 | 31.175 | 33.5 | 33.925
26.1 | 21.775 | 26.5429 | 31 | 33.3 | 33.875
26.2 | 21.7 | 26.4125 | 30.8167 | 33.1 | 33.625
26.3 | 21.5667 | 26.2 | 30.7 | 32.9286 | 33.475
26.4 | 21.4 | 26.1 | 30.5429 | 32.8 | 33.325
26.5 | 21.3 | 25.9571 | 30.4 | 32.6333 | 33.175
26.6 | 21.15 | 25.8 | 30.275 | 32.5 | 33
26.7 | 21 | 25.7 | 30.0667 | 32.34 | 32.8
26.8 | 20.94 | 25.5 | 29.9 | 32.075 | 32.725
26.9 | 20.775 | 25.3333 | 29.725 | 31.925 | 32.4833
27 | 20.625 | 25.2 | 29.6 | 31.7875 | 32.3
27.1 | 20.5 | 25.05 | 29.4 | 31.625 | 32.24
27.2 | 20.4 | 24.9 | 29.1778 | 31.4875 | 32.1
27.3 | 20.3 | 24.8 | 29.1 | 31.3625 | 31.9
27.4 | 20.15 | 24.7 | 28.925 | 31.2286 | 31.8
27.5 | 20 | 24.5 | 28.88 | 31.1 | 31.6
27.6 | 19.9 | 24.4 | 28.7 | 30.9 | 31.4
27.7 | 19.8 | 24.24 | 28.5 | 30.8 | 31.25
27.8 | 19.66 | 24.1 | 28.4 | 30.6 | 31.2
27.9 | 19.5 | 23.88 | 28.3 | 30.5 | 31.05
28 | 19.4 | 23.8 | 28.15 | 30.375 | 30.9
28.1 | 19.3 | 23.7 | 27.9 | 30.2 | 30.6714
28.2 | 19.2 | 23.55 | 27.8 | 30 | 30.5375
28.3 | 19.075 | 23.4 | 27.62 | 29.9 | 30.4222
28.4 | 18.925 | 23.3 | 27.5 | 29.775 | 30.28
28.5 | 18.8 | 23.125 | 27.4 | 29.525 | 30.1143
28.6 | 18.725 | 22.9286 | 27.2 | 29.4 | 30
28.7 | 18.575 | 22.8 | 27.125 | 29.225 | 29.85
28.8 | 18.5 | 22.75 | 26.975 | 29.175 | 29.65
28.9 | 18.375 | 22.6 | 26.825 | 29 | 29.525
29 | 18.3 | 22.5 | 26.775 | 28.875 | 29.4
29.1 | 18.175 | 22.3333 | 26.525 | 28.725 | 29.3
29.2 | 18.0667 | 22.2 | 26.4 | 28.6 | 29.1
29.3 | 17.95 | 22.1 | 26.225 | 28.425 | 29
29.4 | 17.85 | 22 | 26.1 | 28.2 | 28.85
29.5 | 17.7 | 21.8 | 26 | 28.1 | 28.7
29.6 | 17.6 | 21.8 | 25.875 | 28 | 28.55
29.7 | 17.55 | 21.575 | 25.7 | 27.84 | 28.4
29.8 | 17.4 | 21.5 | 25.6 | 27.68 | 28.25
29.9 | 17.3 | 21.3667 | 25.5 | 27.5333 | 28.1
30 | 17.1833 | 21.2 | 25.3 | 27.425 | 27.95

## Wedge 15 Lateral Profiles
Lateral Profiles
Wedge 15
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.2 cm FieldSize:4x4 cm2 | Depth:1.2 cm FieldSize:10x10 cm2 | Depth:1.2 cm FieldSize:20x20 cm2 | Depth:1.2 cm FieldSize:30x30 cm2 | Depth:1.2 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-21.8 |  |  |  | 6.9 | 8.0571 |  |  |  | 8.5 | 9.7106 |  |  |  | 11.78 | 13.1 |  |  |  | 19.475 | 21.11 |  |  | 10.6 | 27.96 | 29.6259
-21.7 |  |  |  | 7 | 8.14 |  |  |  | 8.58 | 9.8314 |  |  |  | 12.0143 | 13.3 |  |  |  | 19.7333 | 21.4333 |  |  | 10.7 | 28.3143 | 30.026
-21.6 |  |  |  | 7 | 8.22 |  |  |  | 8.7 | 9.9647 |  |  |  | 12.18 | 13.4 |  |  |  | 20.0667 | 21.66 |  |  | 10.8 | 28.75 | 30.3661
-21.5 |  |  |  | 7.1 | 8.3 |  |  |  | 8.8 | 10.1147 |  |  |  | 12.28 | 13.64 |  |  |  | 20.26 | 21.92 |  |  | 10.9 | 28.9667 | 30.7262
-21.4 |  |  |  | 7.2 | 8.32 |  |  |  | 8.9 | 10.2123 |  |  |  | 12.48 | 13.84 |  |  |  | 20.66 | 22.22 |  |  | 11 | 29.3429 | 31.1562
-21.3 |  |  |  | 7.3 | 8.4 |  |  |  | 9.08 | 10.2979 |  |  |  | 12.7 | 13.94 |  |  |  | 20.96 | 22.46 |  |  | 11.14 | 29.7714 | 31.5263
-21.2 |  |  |  | 7.32 | 8.49 |  |  |  | 9.2 | 10.4579 |  |  |  | 12.84 | 14.24 |  |  |  | 21.22 | 22.78 |  |  | 11.3 | 30.11 | 31.7264
-21.1 |  |  |  | 7.42 | 8.5818 |  |  |  | 9.38 | 10.5979 |  |  |  | 13.04 | 14.34 |  |  |  | 21.52 | 23.12 |  |  | 11.4 | 30.48 | 32.1564
-21 |  |  |  | 7.5 | 8.68 |  |  |  | 9.48 | 10.7578 |  |  |  | 13.2 | 14.54 |  |  |  | 21.825 | 23.42 |  |  | 11.6 | 30.95 | 32.5665
-20.9 |  |  |  | 7.6 | 8.78 |  |  |  | 9.6 | 10.8578 |  |  |  | 13.38 | 14.7667 |  |  |  | 22.1 | 23.75 |  |  | 11.6667 | 31.3444 | 33.0066
-20.8 |  |  |  | 7.66 | 8.8889 |  |  |  | 9.7 | 10.9578 |  |  |  | 13.58 | 14.94 |  |  |  | 22.4333 | 24.02 |  |  | 11.7667 | 31.8 | 33.4567
-20.7 |  |  |  | 7.8 | 9 |  |  |  | 9.9 | 11.1578 |  |  |  | 13.78 | 15.1 |  |  |  | 22.75 | 24.36 |  |  | 11.8833 | 32.24 | 33.9068
-20.6 |  |  |  | 7.86 | 9.1 |  |  |  | 10.0333 | 11.2977 |  |  |  | 13.98 | 15.36 |  |  |  | 23.14 | 24.72 |  |  | 12.04 | 32.76 | 34.3068
-20.5 |  |  |  | 8 | 9.2 |  |  |  | 10.175 | 11.3977 |  |  |  | 14.18 | 15.56 |  |  |  | 23.45 | 25.02 |  |  | 12.2 | 33.24 | 34.8269
-20.4 |  |  |  | 8.16 | 9.3 |  |  |  | 10.3 | 11.5577 |  |  |  | 14.46 | 15.76 |  |  |  | 23.7667 | 25.36 |  |  | 12.38 | 33.82 | 35.4871
-20.3 |  |  |  | 8.26 | 9.36 |  |  |  | 10.45 | 11.6977 |  |  |  | 14.68 | 15.98 |  |  | 8.36 | 24.26 | 25.72 |  |  | 12.5 | 34.52 | 35.9272
-20.2 |  |  |  | 8.3 | 9.56 |  |  |  | 10.5 | 11.8976 |  |  |  | 14.78 | 16.18 |  |  | 8.54 | 24.56 | 26.075 |  |  | 12.6 | 35.28 | 36.7673
-20.1 |  |  |  | 8.4 | 9.7 |  |  |  | 10.7 | 11.9976 |  |  |  | 15.06 | 16.46 |  |  | 8.6 | 24.86 | 26.52 |  |  | 12.7 | 36.26 | 37.7475
-20 |  |  |  | 8.5 | 9.7167 |  |  |  | 10.86 | 12.0976 |  |  |  | 15.28 | 16.68 |  |  | 8.7 | 25.26 | 26.88 |  |  | 12.88 | 37.72 | 39.1278
-19.9 |  |  |  | 8.7 | 9.85 |  |  |  | 11 | 12.2975 |  |  |  | 15.5 | 16.98 |  |  | 8.8 | 25.66 | 27.2 |  |  | 13 | 40.02 | 41.1082
-19.8 |  |  |  | 8.8 | 9.96 |  |  |  | 11.14 | 12.4975 |  |  |  | 15.775 | 17.08 |  |  | 8.9 | 26.06 | 27.58 |  |  | 13.18 | 43.68 | 44.4689
-19.7 |  |  |  | 8.8667 | 10.06 |  |  |  | 11.34 | 12.5975 |  |  |  | 16.1 | 17.42 |  |  | 9.04 | 26.46 | 27.9667 |  |  | 13.3143 | 48.7 | 49.3698
-19.6 |  |  |  | 9 | 10.2 |  |  |  | 11.5 | 12.8141 |  |  |  | 16.3 | 17.72 |  |  | 9.14 | 26.9 | 28.4333 |  |  | 13.4778 | 54.98 | 55.5711
-19.5 |  |  |  | 9.1 | 10.36 |  |  |  | 11.65 | 12.9641 |  |  |  | 16.5 | 17.94 |  |  | 9.24 | 27.32 | 28.8 |  |  | 13.6571 | 62.38 | 62.3125
-19.4 |  |  |  | 9.2 | 10.46 |  |  |  | 11.84 | 13.1974 |  |  |  | 16.85 | 18.22 |  |  | 9.3 | 27.66 | 29.26 |  |  | 13.7143 | 69.4818 | 69.6939
-19.3 |  |  |  | 9.36 | 10.6 |  |  |  | 12 | 13.2973 |  |  |  | 17.16 | 18.44 |  |  | 9.44 | 28.24 | 29.7 |  |  | 13.8667 | 76.35 | 76.5353
-19.2 |  |  |  | 9.46 | 10.7 |  |  |  | 12.2 | 13.4973 |  |  |  | 17.46 | 18.85 |  |  | 9.54 | 28.66 | 30.1667 |  |  | 14.06 | 82.85 | 82.6499
-19.1 |  |  |  | 9.6 | 10.8 |  |  |  | 12.4 | 13.6973 |  |  |  | 17.7 | 19.08 |  |  | 9.7 | 29.15 | 30.64 |  |  | 14.26 | 88 | 87.8576
-19 |  |  |  | 9.6857 | 11 |  |  |  | 12.58 | 13.8972 |  |  |  | 18.04 | 19.38 |  |  | 9.8 | 29.74 | 31.1 |  |  | 14.36 | 91.55 | 91.0582
-18.9 |  |  |  | 9.83 | 11.1 |  |  |  | 12.78 | 14.0972 |  |  |  | 18.28 | 19.76 |  |  | 9.9 | 30.34 | 31.78 |  |  | 14.56 | 93.85 | 93.2187
-18.8 |  |  |  | 9.93 | 11.2 |  |  |  | 12.98 | 14.2971 |  |  |  | 18.66 | 20.04 |  |  | 10.02 | 30.94 | 32.6 |  |  | 14.76 | 95.25 | 94.5189
-18.7 |  |  |  | 10.06 | 11.3 |  |  |  | 13.18 | 14.4971 |  |  |  | 19.1 | 20.42 |  |  | 10.1 | 31.9 | 33.4 |  |  | 14.92 | 96.05 | 95.4591
-18.6 |  |  |  | 10.2333 | 11.5 |  |  |  | 13.38 | 14.6971 |  |  | 6.2333 | 19.38 | 20.74 |  |  | 10.2667 | 32.94 | 34.3 |  |  | 15.1167 | 96.7 | 95.9792
-18.5 |  |  |  | 10.3667 | 11.6 |  |  |  | 13.68 | 14.897 |  |  | 6.3 | 19.7667 | 21.12 |  |  | 10.3667 | 34.48 | 36.1 |  |  | 15.3 | 97.2 | 96.5593
-18.4 |  |  |  | 10.56 | 11.8 |  |  |  | 13.88 | 15.1684 |  |  | 6.4 | 20.15 | 21.44 |  |  | 10.5 | 37.2333 | 38.425 |  |  | 15.5 | 97.6 | 96.8394
-18.3 |  |  |  | 10.68 | 11.9 |  |  |  | 14.16 | 15.4369 |  |  | 6.44 | 20.4667 | 21.84 |  |  | 10.6667 | 41.16 | 42.375 |  |  | 15.6 | 97.9 | 97.0794
-18.2 |  |  |  | 10.8375 | 12.0833 |  |  |  | 14.38 | 15.7368 |  |  | 6.5 | 20.86 | 22.24 |  |  | 10.8 | 47.22 | 48.66 |  |  | 15.7 | 98.2 | 97.3795
-18.1 |  |  |  | 11 | 12.2 |  |  |  | 14.6 | 15.9568 |  |  | 6.6 | 21.26 | 22.64 |  |  | 10.9 | 55.3286 | 56.36 |  |  | 16 | 98.5 | 97.7396
-18 |  |  |  | 11.1667 | 12.4167 |  |  |  | 14.96 | 16.1568 |  |  | 6.6667 | 21.725 | 23.14 |  |  | 11.05 | 64.06 | 64.95 |  |  | 16.2 | 98.8 | 98.0796
-17.9 |  |  |  | 11.3 | 12.6 |  |  |  | 15.18 | 16.5167 |  |  | 6.7 | 22.15 | 23.46 |  |  | 11.2333 | 73.5571 | 74.26 |  |  | 16.4 | 99.1 | 98.2196
-17.8 |  |  |  | 11.4667 | 12.7 |  |  | 5.1 | 15.4714 | 16.7566 |  |  | 6.8 | 22.7 | 23.94 |  |  | 11.4 | 82.6 | 82.46 |  |  | 16.6 | 99.3 | 98.3797
-17.7 |  |  |  | 11.7 | 12.9 |  |  | 5.2 | 15.74 | 17.1166 |  |  | 6.9 | 23.08 | 24.36 |  |  | 11.5333 | 90.2 | 89.94 |  |  | 16.8 | 99.4833 | 98.6397
-17.6 |  |  |  | 11.8 | 13.1 |  |  | 5.2 | 16.04 | 17.3365 |  |  | 7 | 23.58 | 24.94 |  |  | 11.68 | 96.44 | 95.78 |  |  | 17 | 99.6444 | 98.8798
-17.5 |  |  |  | 12 | 13.3 |  |  | 5.3 | 16.42 | 17.7131 |  |  | 7.08 | 24.18 | 25.48 |  |  | 11.8333 | 100.0333 | 99.64 |  |  | 17.18 | 99.7625 | 99.0798
-17.4 |  |  |  | 12.2 | 13.5 |  |  | 5.3 | 16.74 | 18.0364 |  |  | 7.18 | 24.88 | 26.15 |  |  | 11.94 | 102.2 | 101.4 |  |  | 17.38 | 100 | 99.2799
-17.3 |  |  |  | 12.4 | 13.6 |  |  | 5.4 | 17.12 | 18.4363 |  |  | 7.3 | 25.64 | 26.9 |  |  | 12.1667 | 103.5 | 102.6 |  |  | 17.6333 | 100.18 | 99.3799
-17.2 |  |  |  | 12.6 | 13.9 |  |  | 5.46 | 17.52 | 18.7363 |  |  | 7.3833 | 26.5 | 27.82 |  |  | 12.3333 | 104.28 | 103.5 |  |  | 17.84 | 100.4 | 99.5199
-17.1 |  |  |  | 12.8 | 14.0333 |  |  | 5.55 | 17.84 | 19.1362 |  |  | 7.4444 | 27.82 | 28.9 |  |  | 12.52 | 104.8857 | 104 |  |  | 18.0667 | 100.54 | 99.6799
-17 |  |  | 4.6 | 13 | 14.3333 |  |  | 5.6 | 18.22 | 19.5361 |  |  | 7.5625 | 29.82 | 31.06 |  |  | 12.64 | 105.3 | 104.5 |  |  | 18.4 | 100.6 | 99.87
-16.9 |  |  | 4.7 | 13.2 | 14.5 |  |  | 5.6 | 18.65 | 19.956 |  |  | 7.7 | 33.42 | 34.18 |  |  | 12.9 | 105.7 | 104.8 |  |  | 18.5889 | 100.7 | 99.98
-16.8 |  |  | 4.7 | 13.5 | 14.7333 |  |  | 5.7 | 19.12 | 20.4559 |  |  | 7.7 | 38.95 | 39.3333 |  |  | 13.02 | 105.94 | 105 |  |  | 18.825 | 100.82 | 100.12
-16.7 |  |  | 4.74 | 13.7 | 14.88 |  |  | 5.76 | 19.6875 | 20.9758 |  |  | 7.88 | 47.8 | 48.18 |  |  | 13.22 | 106.1 | 105.2 |  |  | 19.1 | 100.9286 | 100.22
-16.6 |  |  | 4.8 | 13.9 | 15.18 |  |  | 5.8 | 20.2857 | 21.6557 |  |  | 7.9833 | 58.8 | 58.4667 |  |  | 13.4167 | 106.3333 | 105.4 |  |  | 19.24 | 101.1 | 100.36
-16.5 |  |  | 4.84 | 14.1 | 15.46 |  |  | 5.9 | 21.1 | 22.3755 |  |  | 8 | 70.16 | 70.4333 |  |  | 13.62 | 106.6 | 105.7 |  |  | 19.6 | 101.2 | 100.52
-16.4 |  |  | 4.9 | 14.4 | 15.68 |  |  | 5.9889 | 22.22 | 23.4153 |  |  | 8.2 | 82.34 | 82.8 |  |  | 13.8167 | 106.7 | 105.9 |  |  | 19.8889 | 101.28 | 100.56
-16.3 |  |  | 4.925 | 14.7 | 15.96 |  |  | 6 | 23.74 | 25.015 |  |  | 8.3 | 93.86 | 93.82 |  |  | 14.0333 | 106.8 | 106 |  |  | 20.1667 | 101.3 | 100.66
-16.2 |  |  | 5 | 15 | 16.24 |  |  | 6.1 | 26.5 | 27.7744 |  |  | 8.3833 | 103.6 | 103 |  |  | 14.3 | 106.9167 | 106.1 |  |  | 20.46 | 101.475 | 100.86
-16.1 |  |  | 5.04 | 15.38 | 16.64 |  |  | 6.2 | 31 | 32.7135 |  |  | 8.5 | 110.16 | 109.25 |  |  | 14.4 | 107.1 | 106.1 |  |  | 20.68 | 101.56 | 100.96
-16 |  |  | 5.1 | 15.94 | 17.16 |  |  | 6.26 | 39.9 | 41.1918 |  |  | 8.6 | 114.02 | 112.98 |  |  | 14.6 | 107.2 | 106.2 |  |  | 20.96 | 101.66 | 101.02
-15.9 |  |  | 5.1286 | 16.6333 | 17.78 |  |  | 6.3 | 51.9 | 52.7694 |  |  | 8.78 | 115.7625 | 115.08 |  |  | 14.9 | 107.3 | 106.3 |  |  | 21.38 | 101.8 | 101.12
-15.8 |  |  | 5.2 | 17.58 | 18.76 |  |  | 6.4 | 65.36 | 66.7266 |  |  | 8.9 | 116.9 | 116.06 |  |  | 15.05 | 107.4 | 106.4 |  |  | 21.66 | 101.9 | 101.12
-15.7 |  |  | 5.25 | 19.4167 | 20.54 |  |  | 6.5 | 79.9286 | 80.9038 |  |  | 9.0833 | 117.52 | 116.74 |  |  | 15.35 | 107.5 | 106.5 |  |  | 21.88 | 101.9 | 101.12
-15.6 |  |  | 5.3 | 22.9 | 24.1 |  |  | 6.56 | 93.82 | 94.2034 |  |  | 9.2 | 117.86 | 116.98 |  |  | 15.6 | 107.5 | 106.6 |  |  | 22.26 | 102 | 101.22
-15.5 |  |  | 5.35 | 30.38 | 30.8 |  |  | 6.66 | 106.5 | 106.1786 |  |  | 9.38 | 118.06 | 117.34 |  |  | 15.8 | 107.5 | 106.7 |  | 5.95 | 22.48 | 102.1 | 101.32
-15.4 |  |  | 5.45 | 41.82 | 42.2 |  |  | 6.76 | 115.7 | 115.4271 |  |  | 9.4 | 118.2 | 117.44 |  |  | 16.06 | 107.6 | 106.7 |  | 6 | 22.825 | 102.2 | 101.4367
-15.3 |  |  | 5.5 | 56.52 | 55.7 |  |  | 6.8 | 121.8 | 120.8189 |  |  | 9.6 | 118.2667 | 117.54 |  |  | 16.28 | 107.6 | 106.7 |  | 6.1 | 23.1625 | 102.2444 | 101.52
-15.2 |  |  | 5.5 | 72.76 | 71.6 |  |  | 6.9 | 124.7 | 123.4464 |  |  | 9.74 | 118.3 | 117.6 |  |  | 16.6 | 107.6 | 106.74 |  | 6.175 | 23.42 | 102.3 | 101.62
-15.1 |  |  | 5.59 | 89.5 | 88.1 |  |  | 7.0571 | 125.9 | 124.795 |  |  | 9.94 | 118.3857 | 117.6 |  |  | 16.8833 | 107.6 | 106.8 |  | 6.3 | 23.7571 | 102.4 | 101.72
-15 |  |  | 5.69 | 104.66 | 104 |  |  | 7.1875 | 126.6 | 125.415 |  |  | 10.1333 | 118.4 | 117.575 |  |  | 17.14 | 107.6 | 106.8 |  | 6.3 | 24.16 | 102.5667 | 101.86
-14.9 |  |  | 5.7143 | 118.02 | 117.12 |  |  | 7.2 | 126.8 | 125.675 |  |  | 10.3 | 118.4 | 117.5 |  |  | 17.28 | 107.6 | 106.8 |  | 6.4 | 24.46 | 102.7727 | 102.0867
-14.8 |  |  | 5.8 | 127.075 | 126.12 |  |  | 7.3429 | 126.8 | 125.775 |  |  | 10.38 | 118.3 | 117.5 |  |  | 17.64 | 107.6 | 106.8 |  | 6.5 | 24.86 | 102.96 | 102.26
-14.7 |  |  | 5.8857 | 132.16 | 130.76 |  |  | 7.46 | 126.8 | 125.775 |  |  | 10.58 | 118.2667 | 117.46 |  |  | 17.875 | 107.7 | 106.9 |  | 6.6 | 25.15 | 103.1 | 102.42
-14.6 |  |  | 5.94 | 133.9 | 132.5 |  |  | 7.52 | 126.8 | 125.7417 |  |  | 10.78 | 118.18 | 117.4 |  |  | 18.16 | 107.7 | 106.9 |  | 6.6 | 25.48 | 103.26 | 102.4453
-14.5 |  |  | 6 | 134.2667 | 133 |  |  | 7.675 | 126.8 | 125.555 |  |  | 10.98 | 117.975 | 117.3 |  |  | 18.48 | 107.7 | 107 |  | 6.7 | 26.02 | 103.36 | 102.5543
-14.4 |  |  | 6.1143 | 134.3 | 133.1 |  |  | 7.8 | 126.6 | 125.355 |  |  | 11.1333 | 117.88 | 117.2 |  |  | 18.82 | 107.7 | 107 |  | 6.8 | 26.36 | 103.4375 | 102.721
-14.3 |  |  | 6.24 | 134.2333 | 133.04 |  |  | 7.9 | 126.4 | 125.155 |  |  | 11.4 | 117.7 | 117.1 |  | 4.5 | 19.225 | 107.7 | 107 |  | 6.8 | 26.7714 | 103.6 | 102.841
-14.2 |  |  | 6.3 | 134.0333 | 132.8143 |  |  | 8 | 126.1 | 125.025 |  |  | 11.6 | 117.6 | 117 |  | 4.575 | 19.58 | 107.7 | 107 |  | 6.9 | 27.04 | 103.8 | 103.021
-14.1 |  |  | 6.32 | 133.74 | 132.4667 |  |  | 8.16 | 125.7 | 124.75 |  |  | 11.84 | 117.5 | 116.775 |  | 4.6 | 19.86 | 107.76 | 107 |  | 7 | 27.6 | 103.9 | 103.121
-14 |  |  | 6.425 | 133.28 | 132 |  |  | 8.36 | 125.5 | 124.455 |  |  | 12.04 | 117.375 | 116.6 |  | 4.7 | 20.16 | 107.8 | 107.1 |  | 7.1 | 28.1 | 104 | 103.221
-13.9 |  |  | 6.54 | 132.78 | 131.5 |  |  | 8.46 | 125.2 | 124.235 |  |  | 12.3 | 117.28 | 116.5 |  | 4.725 | 20.46 | 107.9 | 107.1 |  | 7.2 | 28.7 | 104 | 103.221
-13.8 |  |  | 6.6 | 132.32 | 131 |  |  | 8.6 | 124.8 | 123.935 |  |  | 12.5 | 117.06 | 116.3 |  | 4.8 | 20.98 | 108 | 107.2 |  | 7.3 | 29.3 | 104.08 | 103.321
-13.7 |  |  | 6.7 | 131.84 | 130.5 |  |  | 8.7 | 124.54 | 123.555 |  |  | 12.6833 | 116.86 | 116.2 |  | 4.9 | 21.36 | 108 | 107.2 |  | 7.4 | 30 | 104.2 | 103.441
-13.6 |  |  | 6.8 | 131.12 | 130 |  |  | 8.8667 | 124.24 | 123.255 |  |  | 12.98 | 116.58 | 116 |  | 4.95 | 21.74 | 108.04 | 107.4 |  | 7.44 | 31 | 104.3 | 103.641
-13.5 |  |  | 6.9 | 130.78 | 129.6 |  |  | 9.05 | 123.94 | 122.925 |  |  | 13.26 | 116.48 | 115.9 |  | 5 | 22.06 | 108.2 | 107.4 |  | 7.54 | 32.5 | 104.38 | 103.741
-13.4 |  |  | 7.04 | 130.22 | 128.9 |  |  | 9.28 | 123.62 | 122.695 |  |  | 13.58 | 116.3 | 115.64 |  | 5.05 | 22.56 | 108.3 | 107.48 |  | 7.64 | 34.7 | 104.4 | 103.841
-13.3 |  |  | 7.1 | 129.68 | 128.4 |  |  | 9.46 | 123.24 | 122.196 |  |  | 13.88 | 116.18 | 115.44 |  | 5.1 | 22.88 | 108.34 | 107.6 |  | 7.74 | 38.1 | 104.5 | 103.921
-13.2 |  |  | 7.2 | 129.08 | 128 |  |  | 9.66 | 122.92 | 121.9427 |  |  | 14.16 | 115.98 | 115.34 |  | 5.2 | 23.32 | 108.4 | 107.7 |  | 7.84 | 43 | 104.58 | 103.921
-13.1 |  |  | 7.4 | 128.62 | 127.5 |  |  | 9.82 | 122.54 | 121.7093 |  | 2.9992 | 14.38 | 115.8 | 115.2 |  | 5.275 | 23.82 | 108.44 | 107.7333 |  | 7.9 | 49.4 | 104.78 | 104.0377
-13 |  |  | 7.5 | 128.02 | 126.8 |  |  | 10.02 | 122.1667 | 121.2093 |  | 3.0992 | 14.76 | 115.7 | 115 |  | 5.325 | 24.32 | 108.6 | 107.8 |  | 8.04 | 56.9 | 104.8667 | 104.141
-12.9 |  |  | 7.6 | 127.48 | 126.3 |  |  | 10.2 | 121.8 | 120.851 |  | 3.0992 | 15.06 | 115.62 | 114.9 |  | 5.4 | 24.84 | 108.7 | 107.82 |  | 8.2 | 64.7 | 104.9 | 104.221
-12.8 |  |  | 7.72 | 126.96 | 125.875 |  |  | 10.36 | 121.4333 | 120.5093 |  | 3.1992 | 15.3833 | 115.52 | 114.8 |  | 5.5 | 25.54 | 108.7 | 107.92 |  | 8.3 | 71.9 | 104.9 | 104.321
-12.7 |  |  | 7.8857 | 126.25 | 125.325 |  |  | 10.64 | 121.04 | 120.201 |  | 3.1992 | 15.8 | 115.4 | 114.8 |  | 5.6 | 26.5 | 108.7 | 108 |  | 8.4 | 78.6 | 105 | 104.421
-12.6 |  |  | 8.02 | 125.75 | 124.7333 |  |  | 10.875 | 120.74 | 119.816 |  | 3.1992 | 16.1 | 115.3 | 114.7 |  | 5.7 | 27.5 | 108.8 | 108.02 |  | 8.5 | 83.6 | 105 | 104.521
-12.5 |  |  | 8.14 | 125.4333 | 124.44 |  | 2.3995 | 11.15 | 120.52 | 119.696 |  | 3.2992 | 16.38 | 115.3 | 114.7 |  | 5.8 | 28.94 | 108.8 | 108.2 |  | 8.54 | 87.5 | 105.16 | 104.621
-12.4 |  |  | 8.34 | 124.7667 | 123.94 |  | 2.3995 | 11.4143 | 120.24 | 119.376 |  | 3.2992 | 16.86 | 115.3 | 114.7 |  | 5.85 | 31.56 | 108.8 | 108.22 |  | 8.7333 | 89.8 | 105.26 | 104.681
-12.3 |  |  | 8.42 | 124.3667 | 123.28 |  | 2.3995 | 11.66 | 119.94 | 119.076 |  | 3.3992 | 17.26 | 115.28 | 114.7 |  | 5.9 | 35.54 | 108.9 | 108.3 |  | 8.84 | 91 | 105.3 | 104.781
-12.2 |  |  | 8.62 | 123.85 | 122.84 |  | 2.4195 | 11.96 | 119.72 | 118.976 |  | 3.3992 | 17.66 | 115.2 | 114.6 |  | 6 | 41.56 | 108.9 | 108.3 |  | 9.0333 | 91.8 | 105.4 | 104.821
-12.1 |  |  | 8.72 | 123.5 | 122.44 |  | 2.4995 | 12.24 | 119.52 | 118.856 |  | 3.4491 | 18.06 | 115.2 | 114.6 |  | 6.1 | 49.32 | 108.9 | 108.3 |  | 9.12 | 92.8 | 105.4 | 104.921
-12 |  | 2.0333 | 8.92 | 123.08 | 122.1 |  | 2.4995 | 12.54 | 119.32 | 118.676 |  | 3.4991 | 18.64 | 115.1 | 114.6 |  | 6.15 | 57.7 | 109 | 108.3 |  | 9.24 | 93.4 | 105.5 | 105.021
-11.9 |  | 2.1 | 9.12 | 122.75 | 121.72 |  | 2.4995 | 12.88 | 119.22 | 118.476 |  | 3.5491 | 19.22 | 115.1 | 114.6 |  | 6.3 | 66.58 | 109 | 108.3 |  | 9.34 | 93.7 | 105.5 | 105.021
-11.8 |  | 2.1 | 9.32 | 122.5 | 121.4333 |  | 2.4995 | 13.18 | 119.1333 | 118.376 |  | 3.5991 | 19.84 | 115.1 | 114.56 |  | 6.4 | 75.02 | 109 | 108.4 |  | 9.5333 | 94.06 | 105.5 | 105.021
-11.7 |  | 2.1 | 9.5333 | 122.2 | 121.1 |  | 2.5995 | 13.58 | 119 | 118.226 |  | 3.5991 | 20.72 | 115.1 | 114.5 |  | 6.5 | 82.32 | 109.1 | 108.4 |  | 9.64 | 94.5333 | 105.6 | 105.061
-11.6 |  | 2.175 | 9.62 | 121.96 | 120.95 |  | 2.5995 | 13.98 | 118.9 | 118.116 |  | 3.6991 | 21.68 | 115.02 | 114.46 |  | 6.575 | 88.26 | 109.1 | 108.4 | 2.4 | 9.72 | 95.1111 | 105.6 | 105.121
-11.5 |  | 2.2 | 9.925 | 121.72 | 120.75 |  | 2.5995 | 14.38 | 118.8 | 118.026 |  | 3.7741 | 23.4 | 115 | 114.36 |  | 6.6667 | 92.1 | 109.1 | 108.4 | 2.4 | 9.9 | 95.24 | 105.7 | 105.121
-11.4 |  | 2.2 | 10.14 | 121.46 | 120.66 |  | 2.6795 | 14.88 | 118.68 | 117.926 |  | 3.799 | 26.08 | 114.9143 | 114.3 |  | 6.725 | 94.1 | 109.1 | 108.4 | 2.5 | 10.08 | 95.5667 | 105.7 | 105.1496
-11.3 |  | 2.2 | 10.34 | 121.26 | 120.36 |  | 2.6995 | 15.45 | 118.5 | 117.736 |  | 3.874 | 31.2 | 114.8 | 114.2 |  | 6.9 | 95.5 | 109.1 | 108.5 | 2.5 | 10.2 | 95.94 | 105.7 | 105.221
-11.2 |  | 2.2 | 10.58 | 121.16 | 120.2 |  | 2.6995 | 16.3 | 118.4 | 117.616 |  | 3.899 | 38.66 | 114.71 | 114.16 |  | 7 | 96.5 | 109.1 | 108.5 | 2.5 | 10.3667 | 96.32 | 105.8 | 105.221
-11.1 |  | 2.225 | 10.88 | 120.96 | 120.1 |  | 2.7395 | 17.38 | 118.28 | 117.477 |  | 3.999 | 48.4 | 114.7 | 114.1 |  | 7.1 | 96.9 | 109.1 | 108.5 | 2.6 | 10.5 | 96.74 | 105.8 | 105.321
-11 |  | 2.3 | 11.26 | 120.86 | 119.8 |  | 2.7994 | 19.15 | 118.2 | 117.377 |  | 3.999 | 59.2 | 114.6 | 114.0333 |  | 7.2 | 97.7 | 109.1 | 108.5 | 2.5 | 10.66 | 97.04 | 105.88 | 105.321
-10.9 |  | 2.3 | 11.7 | 120.7 | 119.74 |  | 2.7994 | 22.3 | 118 | 117.277 |  | 4.099 | 70.175 | 114.5 | 114 |  | 7.32 | 98 | 109 | 108.5 | 2.6 | 10.86 | 97.42 | 105.9 | 105.3877
-10.8 |  | 2.3 | 12.38 | 120.475 | 119.54 |  | 2.8994 | 27.9 | 117.9 | 117.137 |  | 4.1989 | 80.9 | 114.42 | 113.9 |  | 7.4667 | 98.4 | 109 | 108.5 | 2.6 | 11.06 | 97.62 | 105.9 | 105.421
-10.7 |  | 2.3 | 13.3333 | 120.2333 | 119.44 |  | 2.8994 | 36.9 | 117.8 | 117.037 |  | 4.2989 | 89.0333 | 114.4 | 113.8 | 1.7 | 7.6 | 98.9 | 109 | 108.5 | 2.7 | 11.2 | 97.8 | 105.98 | 105.421
-10.6 |  | 2.4 | 15.38 | 120.1 | 119.3 |  | 2.9661 | 48 | 117.6 | 116.877 |  | 4.2989 | 95.88 | 114.4 | 113.7 | 1.7 | 7.7 | 99.3 | 109 | 108.5 | 2.7 | 11.3 | 97.92 | 106 | 105.421
-10.5 |  | 2.4 | 18.84 | 119.75 | 119 |  | 2.9994 | 60.5 | 117.4 | 116.677 |  | 4.3989 | 100.075 | 114.24 | 113.6 | 1.78 | 7.9 | 99.5833 | 109 | 108.4286 | 2.8 | 11.52 | 98.22 | 106 | 105.521
-10.4 |  | 2.4 | 26.24 | 119.6333 | 118.8 |  | 2.9994 | 73.2 | 117.3 | 116.477 |  | 4.4989 | 101.94 | 114.14 | 113.6 | 1.8 | 8.08 | 99.94 | 109 | 108.5 | 2.725 | 11.72 | 98.44 | 106 | 105.521
-10.3 |  | 2.5 | 37 | 119.5 | 118.7 |  | 3.0994 | 84.48 | 117.2 | 116.277 |  | 4.5789 | 103.3 | 114.04 | 113.5 | 1.8 | 8.2 | 100.18 | 109 | 108.5 | 2.8 | 11.96 | 98.7636 | 106.06 | 105.581
-10.2 |  | 2.5 | 50.16 | 119.28 | 118.5 |  | 3.0994 | 94.85 | 117.1 | 116.2437 |  | 4.6238 | 104.1 | 113.9 | 113.4 | 1.8 | 8.375 | 100.38 | 109 | 108.5 | 2.8 | 12.15 | 98.9833 | 106.1 | 105.581
-10.1 |  | 2.5 | 64.26 | 119.14 | 118.3 |  | 3.1994 | 101.9 | 117 | 116.1437 |  | 4.6988 | 104.7 | 113.8 | 113.3 | 1.8 | 8.425 | 100.76 | 108.9 | 108.5 | 2.8 | 12.3 | 99.3667 | 106.1 | 105.621
-10 |  | 2.5 | 78.6333 | 119 | 118.1 |  | 3.1994 | 106.2 | 116.75 | 116.017 |  | 4.8238 | 105.05 | 113.7 | 113.2 | 1.9 | 8.575 | 101.16 | 108.9 | 108.6 | 2.8 | 12.45 | 99.5 | 106.1 | 105.671
-9.9 |  | 2.575 | 91.42 | 118.78 | 118 |  | 3.2993 | 108.2 | 116.48 | 115.817 |  | 4.8988 | 105.375 | 113.65 | 113.1 | 1.9 | 8.9 | 101.38 | 108.9 | 108.5 | 2.9 | 12.65 | 99.62 | 106.125 | 105.721
-9.8 |  | 2.6 | 102.06 | 118.5 | 117.7 |  | 3.2993 | 109.3 | 116.28 | 115.717 | 1.1 | 4.9988 | 105.775 | 113.5 | 113.05 | 1.9 | 9 | 101.48 | 108.9 | 108.5 | 2.9 | 12.9 | 99.72 | 106.2 | 105.721
-9.7 |  | 2.6 | 109.14 | 118.34 | 117.6 |  | 3.3993 | 109.9 | 116.2 | 115.517 | 1.1 | 5.0987 | 106.075 | 113.4 | 113 | 1.9 | 9.1 | 101.68 | 108.9 | 108.4333 | 3 | 13.1 | 99.92 | 106.2 | 105.781
-9.6 |  | 2.7 | 112.525 | 118.14 | 117.3 |  | 3.3993 | 110.4 | 116.1 | 115.337 | 1.1 | 5.2987 | 106.36 | 113.3 | 112.92 | 2 | 9.3 | 101.96 | 109 | 108.4 | 3 | 13.38 | 100.1 | 106.2 | 105.721
-9.5 |  | 2.7 | 113.9 | 117.94 | 117.2167 |  | 3.4993 | 110.7 | 115.9 | 115.237 | 1.1 | 5.2987 | 106.66 | 113.16 | 112.82 | 2 | 9.5 | 102.08 | 109 | 108.38 | 3 | 13.64 | 100.22 | 106.24 | 105.721
-9.4 |  | 2.7 | 114.7 | 117.74 | 117 | 0.80016 | 3.5993 | 111 | 115.8 | 115.0326 | 1.1 | 5.4986 | 106.76 | 113.1 | 112.7 | 2 | 9.7 | 102.28 | 109 | 108.32 | 3.1 | 13.84 | 100.42 | 106.3 | 105.721
-9.3 |  | 2.8 | 114.95 | 117.6 | 116.875 | 0.80016 | 3.6743 | 111.2 | 115.6 | 114.847 | 1.1 | 5.5586 | 107.12 | 113 | 112.6 | 2 | 9.9 | 102.48 | 108.9 | 108.3 | 3.1 | 14.0667 | 100.62 | 106.3 | 105.721
-9.2 |  | 2.8 | 115.12 | 117.38 | 116.56 | 0.80016 | 3.6993 | 111.3 | 115.5 | 114.7395 | 1.1 | 5.6986 | 107.25 | 112.9 | 112.5 | 2 | 10.1 | 102.68 | 108.8 | 108.3 | 3.1 | 14.4 | 100.72 | 106.3 | 105.821
-9.1 |  | 2.9 | 115.2 | 117.24 | 116.4 | 0.80016 | 3.7992 | 111.5 | 115.4167 | 114.677 | 1.2 | 5.8319 | 107.4 | 112.8 | 112.4 | 2.1 | 10.3 | 102.8 | 108.78 | 108.3 | 3.2 | 14.7 | 100.8167 | 106.3 | 105.821
-9 |  | 2.9 | 115.2 | 117 | 116.3 | 0.80016 | 3.8992 | 111.64 | 115.1 | 114.5881 | 1.2 | 5.8985 | 107.5 | 112.7 | 112.32 | 2.1333 | 10.5 | 102.98 | 108.7 | 108.2 | 3.3 | 14.9 | 101 | 106.3 | 105.821
-8.9 | 0.7 | 2.94 | 115.2 | 116.88 | 116.1 | 0.80016 | 3.9492 | 111.7 | 114.9 | 114.4497 | 1.2 | 6.0985 | 107.65 | 112.6 | 112.2 | 2.2 | 10.825 | 103.08 | 108.7 | 108.14 | 3.3 | 15.2 | 101.1333 | 106.3 | 105.821
-8.8 | 0.7 | 3 | 115.2 | 116.68 | 115.98 | 0.80016 | 3.9992 | 111.8 | 114.8 | 114.3548 | 1.2 | 6.1985 | 107.84 | 112.5 | 112.1 | 2.2 | 10.975 | 103.2 | 108.6 | 108.1 | 3.3 | 15.5 | 101.4 | 106.3 | 105.801
-8.7 | 0.7 | 3.1 | 115.2 | 116.46 | 115.8 | 0.87518 | 4.0992 | 111.8444 | 114.7 | 114.187 | 1.2 | 6.3984 | 108.05 | 112.4 | 112 | 2.2 | 11.22 | 103.375 | 108.6 | 108.04 | 3.375 | 15.7667 | 101.5 | 106.3 | 105.721
-8.6 | 0.7 | 3.1 | 115.2 | 116.2 | 115.58 | 0.90018 | 4.1992 | 111.9 | 114.6 | 113.9225 | 1.3 | 6.4984 | 108.1 | 112.3 | 111.9 | 2.2 | 11.46 | 103.52 | 108.6 | 108 | 3.4 | 16.18 | 101.6 | 106.2 | 105.721
-8.5 | 0.7 | 3.2 | 115.1 | 116.0667 | 115.4 | 0.90018 | 4.2991 | 111.9 | 114.4286 | 113.7484 | 1.3 | 6.7233 | 108.2 | 112.2 | 111.78 | 2.3 | 11.82 | 103.6333 | 108.6 | 108 | 3.4 | 16.4 | 101.6833 | 106.2 | 105.721
-8.4 | 0.7 | 3.225 | 115 | 115.88 | 115.3 | 0.90018 | 4.3991 | 111.9 | 114.24 | 113.677 | 1.3 | 6.8983 | 108.24 | 112.12 | 111.7 | 2.3 | 12.06 | 103.8 | 108.5 | 107.96 | 3.5 | 16.68 | 101.8 | 106.2 | 105.721
-8.3 | 0.7 | 3.3 | 114.9 | 115.7143 | 115.04 | 0.90018 | 4.4991 | 111.99 | 114.1 | 113.477 | 1.3 | 7.0982 | 108.34 | 112.02 | 111.68 | 2.4 | 12.2 | 103.9 | 108.4 | 107.9 | 3.6 | 16.9333 | 101.9833 | 106.2 | 105.721
-8.2 | 0.7 | 3.4 | 114.8 | 115.5667 | 114.84 | 0.90018 | 4.5991 | 112 | 114 | 113.377 | 1.3 | 7.2982 | 108.4 | 111.92 | 111.5 | 2.4 | 12.5 | 104.025 | 108.4 | 107.9 | 3.6 | 17.2667 | 102.1 | 106.1 | 105.721
-8.1 | 0.7 | 3.4 | 114.7 | 115.38 | 114.74 | 0.90018 | 4.774 | 112 | 113.8 | 113.277 | 1.3 | 7.4581 | 108.4 | 111.88 | 111.4833 | 2.4 | 12.9 | 104.225 | 108.32 | 107.9 | 3.7 | 17.6667 | 102.28 | 106.1 | 105.621
-8 | 0.7 | 3.5 | 114.6 | 115.18 | 114.6 | 0.90018 | 4.874 | 111.92 | 113.7 | 113.177 | 1.4 | 7.6481 | 108.4 | 111.68 | 111.3 | 2.5 | 13.2 | 104.3 | 108.3 | 107.82 | 3.75 | 18.14 | 102.4 | 106.06 | 105.621
-7.9 | 0.7 | 3.55 | 114.5 | 115.04 | 114.5 | 0.90018 | 5.024 | 112 | 113.5 | 113.077 | 1.4 | 7.848 | 108.5 | 111.6 | 111.1833 | 2.5 | 13.5 | 104.3 | 108.2 | 107.8 | 3.8 | 18.4 | 102.5 | 106 | 105.621
-7.8 | 0.7 | 3.6 | 114.4 | 114.8667 | 114.36 | 0.90018 | 5.159 | 111.9 | 113.3 | 112.877 | 1.4 | 8.048 | 108.5 | 111.48 | 111.1 | 2.6 | 13.86 | 104.4 | 108.12 | 107.72 | 3.9 | 18.75 | 102.5 | 106 | 105.621
-7.7 | 0.7 | 3.7333 | 114.3 | 114.7125 | 114.1 | 1.0002 | 5.2989 | 111.8 | 113.2 | 112.777 | 1.45 | 8.2979 | 108.5 | 111.38 | 110.9 | 2.6 | 14.075 | 104.5 | 108.1 | 107.7 | 3.9 | 19.2667 | 102.5 | 106 | 105.521
-7.6 | 0.8 | 3.8 | 114 | 114.56 | 113.96 | 1.0002 | 5.4656 | 111.8 | 113.075 | 112.627 | 1.5 | 8.6312 | 108.5 | 111.275 | 110.8 | 2.7 | 14.475 | 104.6 | 108.02 | 107.6 | 4 | 19.68 | 102.6 | 106 | 105.621
-7.5 | 0.8 | 3.9 | 114 | 114.3667 | 113.76 | 1.0002 | 5.5989 | 111.7 | 112.9 | 112.403 | 1.5 | 8.8228 | 108.5 | 111.2 | 110.7 | 2.7 | 14.8333 | 104.6 | 107.92 | 107.6 | 4.1 | 20.225 | 102.6 | 106 | 105.621
-7.4 | 0.8 | 4 | 113.8 | 114.28 | 113.48 | 1.0002 | 5.7988 | 111.6 | 112.7 | 112.278 | 1.5 | 9.0727 | 108.5 | 111 | 110.6 | 2.775 | 15.125 | 104.66 | 107.82 | 107.5 | 4.1 | 20.75 | 102.7 | 106 | 105.621
-7.3 | 0.8 | 4.1 | 113.8 | 113.98 | 113.28 | 1.0002 | 5.9988 | 111.5 | 112.6 | 112.118 | 1.5333 | 9.3727 | 108.5 | 110.88 | 110.5 | 2.8 | 15.625 | 104.7 | 107.8 | 107.5 | 4.2 | 21.4 | 102.8 | 106 | 105.5377
-7.2 | 0.8 | 4.2 | 113.6 | 113.88 | 113.18 | 1.0002 | 6.1988 | 111.4 | 112.42 | 111.928 | 1.6 | 9.7226 | 108.6 | 110.78 | 110.5 | 2.9 | 16.025 | 104.7 | 107.7 | 107.4 | 4.3 | 22.18 | 102.96 | 105.9 | 105.521
-7.1 | 0.8 | 4.3 | 113.4 | 113.68 | 113.08 | 1.0002 | 6.3987 | 111.36 | 112.3 | 111.778 | 1.6 | 9.9725 | 108.52 | 110.68 | 110.4 | 2.9 | 16.3667 | 104.7 | 107.64 | 107.38 | 4.325 | 23 | 103 | 105.9 | 105.521
-7 | 0.8 | 4.46 | 113.2167 | 113.4833 | 112.9 | 1.1002 | 6.6387 | 111.2 | 112.075 | 111.678 | 1.7 | 10.3174 | 108.5 | 110.5667 | 110.1 | 3 | 17 | 104.8 | 107.54 | 107.3 | 4.4 | 24.95 | 103.1 | 105.8 | 105.441
-6.9 | 0.8 | 4.6 | 113.1 | 113.3 | 112.78 | 1.1002 | 6.7986 | 111.16 | 111.9 | 111.518 | 1.7 | 10.664 | 108.5 | 110.42 | 110 | 3.1 | 17.4 | 104.8 | 107.4571 | 107.2 | 4.5 | 27.2 | 103.1 | 105.7 | 105.421
-6.8 | 0.8 | 4.7 | 112.9 | 113.08 | 112.58 | 1.1002 | 7.0486 | 110.96 | 111.75 | 111.318 | 1.7 | 10.9973 | 108.4 | 110.3 | 109.9 | 3.1 | 18.05 | 104.8 | 107.4 | 107.1 | 4.52 | 31 | 103.1 | 105.7 | 105.421
-6.7 | 0.8 | 4.9 | 112.8 | 112.98 | 112.38 | 1.1002 | 7.3235 | 110.875 | 111.62 | 111.218 | 1.75 | 11.3472 | 108.38 | 110.18 | 109.82 | 3.2 | 18.9 | 104.8 | 107.34 | 107.1 | 4.6 | 36.9 | 103.1 | 105.6 | 105.341
-6.6 | 0.8 | 5 | 112.62 | 112.8 | 112.18 | 1.1002 | 7.6235 | 110.7667 | 111.42 | 110.978 | 1.8 | 11.7971 | 108.28 | 110 | 109.7 | 3.3 | 20 | 104.9 | 107.3 | 107 | 4.7 | 44.2 | 103.2 | 105.6 | 105.321
-6.5 | 0.8 | 5.15 | 112.52 | 112.6 | 111.98 | 1.2002 | 7.9234 | 110.675 | 111.3 | 110.878 | 1.8 | 12.247 | 108.2 | 109.84 | 109.62 | 3.3 | 21.425 | 104.8 | 107.3 | 106.9 | 4.85 | 52.9 | 103.1 | 105.6 | 105.221
-6.4 | 0.9 | 5.3 | 112.32 | 112.42 | 111.78 | 1.2002 | 8.1984 | 110.56 | 111.02 | 110.718 | 1.9 | 12.6468 | 108.1 | 109.64 | 109.42 | 3.4 | 24.12 | 104.8 | 107.3 | 106.8286 | 5 | 62.3 | 103.2 | 105.5 | 105.121
-6.3 | 0.9 | 5.5 | 112.12 | 112.12 | 111.58 | 1.2002 | 8.565 | 110.4 | 111 | 110.618 | 1.9 | 13.2967 | 108.1 | 109.54 | 109.3 | 3.5 | 28.075 | 104.8 | 107.22 | 106.7714 | 5 | 71.1 | 103.3 | 105.45 | 105.0543
-6.2 | 0.9 | 5.68 | 112.02 | 111.88 | 111.38 | 1.2002 | 8.9982 | 110.3 | 110.8333 | 110.418 | 2 | 14.1215 | 108.1 | 109.44 | 109.12 | 3.5 | 34.375 | 104.8833 | 107.1 | 106.7 | 5.1 | 79.2 | 103.3 | 105.36 | 105.021
-6.1 | 0.9 | 5.925 | 111.82 | 111.8 | 111.2143 | 1.2002 | 9.3981 | 110.26 | 110.7 | 110.318 | 2 | 15.1629 | 108 | 109.3 | 109 | 3.7 | 42.575 | 104.9 | 107 | 106.6167 | 5.1333 | 85.1 | 103.4 | 105.3 | 105.021
-6 | 0.9 | 6.15 | 111.62 | 111.68 | 111.16 | 1.3003 | 10.048 | 110.1 | 110.5333 | 110.118 | 2.1 | 16.7958 | 107.9 | 109.14 | 108.9 | 3.7 | 51.9 | 104.9 | 106.9 | 106.6 | 5.3 | 89.4 | 103.4 | 105.26 | 105.021
-5.9 | 0.9 | 6.475 | 111.52 | 111.475 | 110.9667 | 1.3003 | 10.8978 | 110.025 | 110.3 | 109.978 | 2.2 | 19.5618 | 107.9 | 109.04 | 108.86 | 3.82 | 62.275 | 104.9 | 106.8 | 106.42 | 5.5 | 92.0833 | 103.4 | 105.2 | 104.9781
-5.8 | 0.9 | 6.96 | 111.325 | 111.36 | 110.78 | 1.3253 | 12.0976 | 109.88 | 110.15 | 109.728 | 2.2 | 24.4689 | 107.8 | 108.84 | 108.625 | 3.9 | 71.5 | 104.9 | 106.7 | 106.4 | 5.58 | 93.8 | 103.4 | 105.1 | 104.921
-5.7 | 1 | 7.6 | 111.14 | 111.06 | 110.58 | 1.4003 | 14.0722 | 109.7833 | 110 | 109.6113 | 2.3 | 32.017 | 107.76 | 108.7 | 108.5 | 4 | 80.2 | 104.9 | 106.6 | 106.22 | 5.6667 | 95.06 | 103.4 | 105.1 | 104.921
-5.6 | 1 | 8.86 | 110.925 | 110.78 | 110.375 | 1.4003 | 17.6465 | 109.66 | 109.74 | 109.478 | 2.3 | 41.9096 | 107.66 | 108.64 | 108.38 | 4.1 | 86.8 | 104.8 | 106.56 | 106.2 | 5.8 | 95.84 | 103.4 | 105.08 | 104.8665
-5.5 | 1 | 11.05 | 110.6333 | 110.68 | 110.18 | 1.5003 | 24.4701 | 109.4 | 109.6 | 109.278 | 2.4 | 52.5269 | 107.5 | 108.44 | 108.2 | 4.2 | 91.4 | 104.8 | 106.46 | 106.1 | 5.92 | 96.18 | 103.4 | 104.9 | 104.771
-5.4 | 1 | 15.22 | 110.525 | 110.58 | 110.025 | 1.5003 | 33.8682 | 109.3 | 109.5 | 109.1447 | 2.5 | 64.264 | 107.4 | 108.3 | 108.14 | 4.2833 | 94.15 | 104.74 | 106.26 | 105.92 | 6.075 | 96.56 | 103.3167 | 104.88 | 104.671
-5.3 | 1 | 23.65 | 110.44 | 110.38 | 109.825 | 1.5003 | 45.3576 | 109.2 | 109.35 | 108.853 | 2.5 | 74.9812 | 107.26 | 108.2 | 107.94 | 4.4 | 95.6 | 104.7 | 106.16 | 105.9 | 6.2 | 97.1 | 103.3 | 104.78 | 104.571
-5.2 | 1.1 | 34 | 110.24 | 110.18 | 109.6333 | 1.6003 | 58.3634 | 108.9667 | 109.26 | 108.678 | 2.6 | 83.7791 | 107.16 | 108.04 | 107.84 | 4.54 | 96.5 | 104.7 | 106.06 | 105.9 | 6.3 | 97.3 | 103.3 | 104.68 | 104.471
-5.1 | 1.1 | 47 | 110.04 | 109.96 | 109.5 | 1.6003 | 70.5109 | 108.8429 | 109.06 | 108.578 | 2.7 | 91.0772 | 107.06 | 107.88 | 107.74 | 4.68 | 97.18 | 104.7 | 106 | 105.8 | 6.5 | 97.76 | 103.2 | 104.6 | 104.371
-5 | 1.1 | 60.9 | 109.84 | 109.78 | 109.325 | 1.7003 | 81.0838 | 108.6286 | 108.86 | 108.458 | 2.8333 | 95.2262 | 107 | 107.6833 | 107.625 | 4.84 | 97.8 | 104.62 | 105.86 | 105.8 | 6.6 | 97.98 | 103.2 | 104.6 | 104.321
-4.9 | 1.1 | 74.325 | 109.64 | 109.5833 | 109.1333 | 1.8004 | 90.5486 | 108.5273 | 108.66 | 108.158 | 2.9333 | 97.5756 | 106.9 | 107.5 | 107.5 | 5 | 98.2 | 104.6 | 105.8 | 105.62 | 6.75 | 98.18 | 103.2 | 104.6 | 104.321
-4.8 | 1.2 | 86.075 | 109.5 | 109.4 | 109 | 1.8004 | 96.5307 | 108.36 | 108.36 | 108.078 | 3 | 98.8753 | 106.775 | 107.3625 | 107.2 | 5.125 | 98.55 | 104.49 | 105.76 | 105.42 | 6.925 | 98.5 | 103.2 | 104.5 | 104.321
-4.7 | 1.2 | 95.925 | 109.24 | 109.24 | 108.8167 | 1.9004 | 99.98 | 108.16 | 108.16 | 107.998 | 3.1 | 99.7751 | 106.6667 | 107.2286 | 107 | 5.275 | 98.825 | 104.39 | 105.7 | 105.32 | 7.1 | 98.82 | 103.1 | 104.48 | 104.271
-4.6 | 1.2 | 101.525 | 109.04 | 109.0667 | 108.6 | 1.9254 | 101.58 | 108.03 | 107.9556 | 107.8113 | 3.2 | 100.45 | 106.48 | 107.1 | 106.8 | 5.4667 | 99.025 | 104.3 | 105.6 | 105.22 | 7.225 | 99.04 | 103.1 | 104.3 | 104.221
-4.5 | 1.3 | 104.225 | 108.9 | 108.8 | 108.42 | 2.0754 | 102.48 | 107.86 | 107.8 | 107.618 | 3.4 | 100.8 | 106.38 | 106.9 | 106.7 | 5.6 | 99.275 | 104.2 | 105.4667 | 105.02 | 7.55 | 99.3 | 103.2 | 104.2 | 104.171
-4.4 | 1.3 | 105.3 | 108.74 | 108.6 | 108.14 | 2.1254 | 102.979 | 107.66 | 107.7 | 107.3623 | 3.525 | 100.975 | 106.28 | 106.7375 | 106.6 | 5.775 | 99.475 | 104.125 | 105.24 | 105 | 7.7667 | 99.4 | 103.1 | 104.2 | 104.071
-4.3 | 1.4 | 105.8667 | 108.44 | 108.4 | 107.92 | 2.2004 | 103.429 | 107.53 | 107.44 | 107.159 | 3.7 | 101.275 | 106.1 | 106.6 | 106.4 | 6 | 99.675 | 104.04 | 105.1 | 104.92 | 7.9 | 99.5333 | 103.1 | 104.1 | 104.021
-4.2 | 1.4 | 106.1 | 108.24 | 108.2 | 107.8 | 2.3005 | 103.579 | 107.36 | 107.3 | 107.079 | 3.8 | 101.475 | 105.9857 | 106.46 | 106.4 | 6.16 | 99.975 | 104.0125 | 104.9667 | 104.9 | 8.1 | 99.78 | 103.125 | 104 | 103.981
-4.1 | 1.5 | 106.2 | 108.05 | 108 | 107.54 | 2.4205 | 103.804 | 107.16 | 107.1 | 106.919 | 3.9 | 101.675 | 105.9 | 106.36 | 106.2 | 6.35 | 100.1667 | 104.1 | 104.8 | 104.8 | 8.3333 | 99.9 | 103.2 | 104 | 103.921
-4 | 1.5 | 106.2 | 107.84 | 107.8167 | 107.42 | 2.5005 | 103.959 | 106.96 | 106.9 | 106.7699 | 4.16 | 101.9083 | 105.8 | 106.26 | 106.1 | 6.6 | 100.4 | 104 | 104.7 | 104.64 | 8.6 | 100.08 | 103.2 | 103.9 | 103.881
-3.9 | 1.6 | 106.2 | 107.6 | 107.62 | 107.32 | 2.6505 | 104.004 | 106.86 | 106.8 | 106.579 | 4.3 | 102.074 | 105.72 | 106 | 105.9 | 6.75 | 100.5 | 104 | 104.6667 | 104.6 | 8.85 | 100.18 | 103.1 | 103.9 | 103.741
-3.8 | 1.7 | 106.2 | 107.34 | 107.42 | 107.04 | 2.8006 | 104.079 | 106.66 | 106.6571 | 106.379 | 4.5 | 102.174 | 105.55 | 105.86 | 105.6 | 7 | 100.6 | 103.94 | 104.6 | 104.44 | 9.1 | 100.28 | 103 | 103.8 | 103.621
-3.7 | 1.7 | 106.15 | 107.14 | 107.22 | 106.82 | 2.9506 | 104.079 | 106.48 | 106.44 | 106.179 | 4.75 | 102.274 | 105.4143 | 105.6667 | 105.5 | 7.2333 | 100.75 | 103.9 | 104.5 | 104.3 | 9.34 | 100.375 | 102.9 | 103.7 | 103.581
-3.6 | 1.8 | 106 | 106.85 | 106.92 | 106.58 | 3.1006 | 104.179 | 106.3636 | 106.24 | 105.979 | 4.9 | 102.374 | 105.2364 | 105.48 | 105.4 | 7.5 | 100.9 | 103.8 | 104.3 | 104.2 | 9.725 | 100.5 | 102.8 | 103.6 | 103.441
-3.5 | 1.9 | 105.9 | 106.74 | 106.72 | 106.325 | 3.3007 | 104.179 | 106.1778 | 106.06 | 105.779 | 5.15 | 102.4073 | 105.12 | 105.38 | 105.22 | 7.7333 | 100.9 | 103.7 | 104.3 | 104.08 | 10.15 | 100.66 | 102.7 | 103.425 | 103.321
-3.4 | 2 | 105.825 | 106.56 | 106.5333 | 106.24 | 3.4007 | 104.279 | 105.9818 | 105.8833 | 105.579 | 5.38 | 102.499 | 105.02 | 105.2 | 105.12 | 8.175 | 101.0667 | 103.6 | 104.2 | 104 | 10.675 | 100.8 | 102.7 | 103.3 | 103.2543
-3.3 | 2.1 | 105.7 | 106.32 | 106.38 | 106.04 | 3.6007 | 104.279 | 105.88 | 105.7 | 105.429 | 5.68 | 102.574 | 104.84 | 104.98 | 105.02 | 8.55 | 101.2 | 103.5 | 104.1 | 103.94 | 11.225 | 100.9 | 102.6 | 103.2 | 103.171
-3.2 | 2.2 | 105.7 | 106.06 | 106.1857 | 105.84 | 3.8258 | 104.179 | 105.6857 | 105.5 | 105.279 | 5.98 | 102.574 | 104.72 | 104.88 | 104.9 | 8.975 | 101.225 | 103.44 | 104 | 103.9 | 12.4 | 100.9 | 102.6 | 103.175 | 103.021
-3.1 | 2.3 | 105.575 | 105.9 | 105.98 | 105.64 | 4.0758 | 104.179 | 105.45 | 105.4 | 105.079 | 6.325 | 102.574 | 104.6273 | 104.68 | 104.72 | 9.575 | 101.375 | 103.25 | 103.76 | 103.7 | 13.8 | 101 | 102.5 | 103.08 | 102.921
-3 | 2.5 | 105.45 | 105.76 | 105.7273 | 105.44 | 4.4259 | 104.179 | 105.225 | 105.2 | 104.8123 | 6.85 | 102.574 | 104.4444 | 104.48 | 104.52 | 10.575 | 101.4 | 103.16 | 103.7 | 103.5 | 16.325 | 101 | 102.5 | 102.94 | 102.921
-2.9 | 2.7 | 105.3667 | 105.4333 | 105.5333 | 105.24 | 4.801 | 104.079 | 104.99 | 105.0167 | 104.679 | 7.5 | 102.574 | 104.32 | 104.3 | 104.2667 | 11.975 | 101.5 | 103.06 | 103.5667 | 103.4 | 20.2 | 101.1 | 102.4 | 102.84 | 102.8543
-2.8 | 2.9 | 105.2 | 105.3 | 105.3333 | 105.04 | 5.4011 | 103.979 | 104.8875 | 104.82 | 104.579 | 8.3 | 102.574 | 104.14 | 104.3 | 104.1857 | 14.2667 | 101.6 | 102.9 | 103.48 | 103.2 | 25.9 | 101.1 | 102.32 | 102.8 | 102.821
-2.7 | 3.4 | 105.05 | 105.2 | 105.2571 | 104.8 | 6.2412 | 103.904 | 104.725 | 104.5714 | 104.479 | 9.75 | 102.674 | 104.0273 | 104.2 | 104.1 | 18.55 | 101.6 | 102.82 | 103.38 | 103.1 | 34.3 | 101.1 | 102.2 | 102.7 | 102.661
-2.6 | 4.3 | 104.82 | 105.02 | 105.06 | 104.7 | 7.5615 | 103.8457 | 104.6143 | 104.5 | 104.299 | 12.15 | 102.674 | 103.9222 | 103.98 | 103.9333 | 25.05 | 101.6 | 102.8 | 103.28 | 103.06 | 44.8 | 101.1 | 102.2 | 102.6 | 102.621
-2.5 | 5.8 | 104.675 | 104.82 | 104.78 | 104.5333 | 9.852 | 103.754 | 104.425 | 104.3 | 104.099 | 16.5 | 102.654 | 103.82 | 103.8 | 103.7571 | 33.65 | 101.6 | 102.8 | 103.1 | 102.9 | 55.9 | 101.1 | 102.1 | 102.6 | 102.521
-2.4 | 8.4333 | 104.5 | 104.62 | 104.5833 | 104.3333 | 14.2028 | 103.604 | 104.14 | 104.16 | 103.954 | 24.05 | 102.574 | 103.72 | 103.66 | 103.5875 | 44.4 | 101.6 | 102.675 | 102.98 | 102.8 | 67 | 101.2 | 102.1 | 102.5 | 102.4604
-2.3 | 13.8 | 104.3 | 104.4 | 104.4 | 104.05 | 22.1544 | 103.504 | 103.9833 | 104 | 103.754 | 34.2 | 102.474 | 103.5 | 103.56 | 103.4571 | 56.1 | 101.52 | 102.52 | 102.88 | 102.7 | 76.66 | 101.1 | 102.02 | 102.5 | 102.37
-2.2 | 23 | 104.175 | 104.2333 | 104.28 | 103.96 | 32.3065 | 103.379 | 103.8 | 103.74 | 103.559 | 46.2333 | 102.4407 | 103.2 | 103.26 | 103.28 | 67.55 | 101.5 | 102.45 | 102.68 | 102.6333 | 85.14 | 101.2 | 101.92 | 102.32 | 102.32
-2.1 | 36.1 | 104 | 104.1 | 103.98 | 103.8 | 45.009 | 103.219 | 103.6 | 103.4429 | 103.399 | 58.8 | 102.374 | 103.2 | 103.1 | 103.1455 | 77.575 | 101.5 | 102.3333 | 102.58 | 102.56 | 91.14 | 101.2 | 101.9 | 102.22 | 102.18
-2 | 49.7 | 103.775 | 103.8333 | 103.9 | 103.7 | 58.0116 | 103.079 | 103.4 | 103.4 | 103.099 | 70.8 | 102.374 | 103 | 102.96 | 102.9571 | 86.4667 | 101.4 | 102.26 | 102.5 | 102.5 | 94.875 | 101.1 | 101.9 | 102.1 | 102.08
-1.9 | 62.7 | 103.575 | 103.6 | 103.66 | 103.4 | 70.5141 | 102.979 | 103.3 | 103.22 | 102.979 | 80.625 | 102.274 | 103 | 102.86 | 102.8 | 92.1 | 101.4 | 102.16 | 102.5 | 102.3 | 97.2 | 101.1 | 101.8 | 101.92 | 101.88
-1.8 | 75.7 | 103.4 | 103.4 | 103.4 | 103.2 | 81.6163 | 102.879 | 103.2 | 103.1 | 102.879 | 89.475 | 102.174 | 102.8 | 102.72 | 102.8 | 95.7 | 101.4 | 102.06 | 102.4 | 102.2 | 98.2667 | 101.1 | 101.64 | 101.82 | 101.78
-1.7 | 87.3 | 103.1667 | 103.22 | 103.275 | 102.95 | 90.5514 | 102.729 | 102.9 | 102.82 | 102.7123 | 94.65 | 102.074 | 102.6 | 102.5 | 102.6 | 97.575 | 101.3 | 101.86 | 102.2 | 102.1 | 99.1333 | 101.1 | 101.5 | 101.8 | 101.72
-1.6 | 94.55 | 103 | 103.05 | 103.04 | 102.76 | 95.5691 | 102.579 | 102.7 | 102.62 | 102.4133 | 97.6 | 102.0493 | 102.5 | 102.36 | 102.35 | 98.6 | 101.3 | 101.8 | 102 | 102.06 | 99.65 | 101.1 | 101.4 | 101.8 | 101.62
-1.5 | 98.6 | 102.9 | 102.9 | 102.9 | 102.56 | 98.1696 | 102.38 | 102.52 | 102.42 | 102.28 | 98.9 | 101.85 | 102.4 | 102.26 | 102.2 | 99.275 | 101.2 | 101.7 | 101.9 | 101.925 | 99.8667 | 101.1 | 101.4 | 101.6 | 101.62
-1.4 | 100.3 | 102.7 | 102.6 | 102.725 | 102.4 | 99.4199 | 102.18 | 102.4 | 102.3 | 102.18 | 99.62 | 101.7 | 102.1 | 102.075 | 102.1 | 99.7 | 101.1 | 101.6 | 101.75 | 101.72 | 99.95 | 101.1 | 101.3 | 101.6 | 101.58
-1.3 | 101 | 102.4 | 102.4 | 102.4333 | 102.225 | 100.22 | 101.98 | 102.2 | 102.08 | 101.98 | 100 | 101.625 | 101.9 | 101.88 | 101.9 | 99.95 | 101.1 | 101.42 | 101.65 | 101.52 | 100.15 | 101.1 | 101.3 | 101.42 | 101.48
-1.2 | 101.25 | 102.1333 | 102.2 | 102.2333 | 102 | 100.4533 | 101.805 | 102.02 | 101.92 | 101.78 | 100.35 | 101.475 | 101.7 | 101.7667 | 101.74 | 100.1 | 101 | 101.32 | 101.5714 | 101.42 | 100.3 | 101.06 | 101.2 | 101.3 | 101.38
-1.1 | 101.3 | 102 | 102 | 102.075 | 101.725 | 100.64 | 101.7133 | 101.82 | 101.68 | 101.68 | 100.5 | 101.375 | 101.6 | 101.58 | 101.6 | 100.3 | 100.8667 | 101.14 | 101.4375 | 101.3 | 100.35 | 100.9 | 101.1 | 101.12 | 101.28
-1 | 101.3 | 101.75 | 101.9 | 101.8 | 101.6 | 100.72 | 101.6467 | 101.54 | 101.58 | 101.48 | 100.5 | 101.275 | 101.5 | 101.4143 | 101.5 | 100.35 | 100.8 | 101.02 | 101.26 | 101.2 | 100.4 | 100.8 | 100.9 | 101.02 | 101.12
-0.9 | 101.25 | 101.6333 | 101.58 | 101.58 | 101.4667 | 100.82 | 101.455 | 101.4 | 101.38 | 101.2967 | 100.6 | 101.115 | 101.25 | 101.2571 | 101.375 | 100.4 | 100.625 | 100.92 | 101.2 | 101.02 | 100.4 | 100.7 | 100.9 | 100.92 | 100.995
-0.8 | 101.1 | 101.45 | 101.4 | 101.38 | 101.2 | 100.82 | 101.205 | 101.3 | 101.2 | 101.18 | 100.6 | 101.0417 | 101.05 | 101.0857 | 101.26 | 100.3667 | 100.6 | 100.82 | 101.06 | 100.8333 | 100.4 | 100.7 | 100.8 | 100.9 | 100.8867
-0.7 | 101 | 101.225 | 101.3 | 101.18 | 101.1 | 100.82 | 101.08 | 101.24 | 101.1 | 100.88 | 100.6 | 100.775 | 100.9 | 100.9273 | 101.0333 | 100.4 | 100.6 | 100.72 | 100.9714 | 100.74 | 100.4 | 100.6 | 100.7 | 100.82 | 100.8
-0.6 | 100.95 | 100.975 | 101.2 | 101.08 | 101 | 100.72 | 100.98 | 101.04 | 100.9 | 100.7 | 100.6 | 100.675 | 100.8 | 100.77 | 100.92 | 100.3 | 100.5 | 100.64 | 100.8 | 100.54 | 100.4 | 100.52 | 100.62 | 100.72 | 100.72
-0.5 | 100.8 | 100.825 | 100.9 | 100.88 | 100.775 | 100.62 | 100.78 | 100.84 | 100.74 | 100.5 | 100.5 | 100.575 | 100.7 | 100.67 | 100.78 | 100.3 | 100.5 | 100.6 | 100.6 | 100.4571 | 100.4 | 100.42 | 100.5 | 100.6 | 100.6
-0.4 | 100.7 | 100.78 | 100.8 | 100.7 | 100.6 | 100.52 | 100.68 | 100.64 | 100.6 | 100.4 | 100.5 | 100.5417 | 100.5 | 100.57 | 100.675 | 100.2 | 100.34 | 100.54 | 100.5 | 100.3833 | 100.375 | 100.12 | 100.4 | 100.5 | 100.42
-0.3 | 100.5 | 100.525 | 100.5667 | 100.6 | 100.48 | 100.42 | 100.48 | 100.5 | 100.3 | 100.3 | 100.5 | 100.4 | 100.4 | 100.47 | 100.5 | 100.2 | 100.2 | 100.44 | 100.4 | 100.3 | 100.3 | 100.1 | 100.32 | 100.4 | 100.32
-0.2 | 100.3 | 100.3 | 100.3 | 100.3 | 100.36 | 100.32 | 100.18 | 100.34 | 100.26 | 100.2 | 100.225 | 100.275 | 100.3 | 100.34 | 100.4 | 100.1 | 100.1 | 100.24 | 100.2 | 100.3 | 100.2 | 100.1 | 100.22 | 100.225 | 100.22
-0.1 | 100.2 | 100.14 | 100.1 | 100.2 | 100.16 | 100.22 | 100.08 | 100.14 | 100.2 | 100.1 | 100.1 | 100.175 | 100.1333 | 100.14 | 100.175 | 100.1 | 100.1 | 100.14 | 100.1167 | 100.1 | 100.15 | 100.1 | 100.12 | 100.14 | 100.2
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9 | 99.8533 | 99.855 | 99.9 | 99.8167 | 99.8 | 99.9 | 99.7751 | 99.9 | 99.87 | 99.9 | 99.88 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9333 | 99.96 | 99.9 | 99.82
0.2 | 99.7 | 99.7 | 99.8 | 99.9 | 99.76 | 99.7199 | 99.7001 | 99.8 | 99.7 | 99.6801 | 99.8 | 99.6751 | 99.8 | 99.74 | 99.8 | 99.65 | 99.7 | 99.8 | 99.9 | 99.8 | 99.9 | 99.82 | 99.82 | 99.9 | 99.82
0.3 | 99.5333 | 99.55 | 99.7 | 99.8 | 99.6 | 99.6199 | 99.5551 | 99.6 | 99.52 | 99.5134 | 99.7 | 99.5751 | 99.7 | 99.6 | 99.7 | 99.5 | 99.6 | 99.7 | 99.75 | 99.7 | 99.8 | 99.725 | 99.7 | 99.8 | 99.82
0.4 | 99.3 | 99.4 | 99.46 | 99.7 | 99.5 | 99.5199 | 99.4051 | 99.5 | 99.5 | 99.4801 | 99.5 | 99.4085 | 99.6 | 99.5625 | 99.675 | 99.35 | 99.4333 | 99.6 | 99.6 | 99.66 | 99.7 | 99.6333 | 99.7 | 99.7 | 99.7199
0.5 | 99.2 | 99.3 | 99.3333 | 99.5 | 99.48 | 99.2198 | 99.3051 | 99.4 | 99.32 | 99.3051 | 99.2667 | 99.3752 | 99.4167 | 99.4 | 99.5 | 99.18 | 99.375 | 99.5 | 99.48 | 99.6 | 99.5667 | 99.44 | 99.58 | 99.7 | 99.7199
0.6 | 99.1 | 99.2 | 99.2333 | 99.4 | 99.38 | 98.9698 | 99.2551 | 99.3 | 99.22 | 99.1802 | 99.1 | 99.2752 | 99.3 | 99.28 | 99.4 | 99.0333 | 99.3 | 99.3444 | 99.4 | 99.5 | 99.4 | 99.3667 | 99.5 | 99.6 | 99.6199
0.7 | 98.95 | 99.1 | 99.2 | 99.3 | 99.26 | 98.8448 | 99.0802 | 99.1 | 99.1 | 99.0802 | 98.9 | 99.1752 | 99.3 | 99.16 | 99.4 | 98.9 | 99.225 | 99.2 | 99.3 | 99.36 | 99.175 | 99.3 | 99.38 | 99.5 | 99.5199
0.8 | 98.8 | 98.925 | 99.1 | 99.2167 | 99.16 | 98.6697 | 98.9135 | 99.1 | 99.06 | 99.0802 | 98.725 | 98.9253 | 99.12 | 99.1 | 99.3 | 98.7333 | 99.1 | 99.1 | 99.28 | 99.3 | 99 | 99.2 | 99.3 | 99.44 | 99.3199
0.9 | 98.6 | 98.8 | 98.96 | 99.1 | 99.0125 | 98.2696 | 98.7469 | 98.98 | 98.96 | 98.8802 | 98.5 | 98.7753 | 99 | 99 | 99.1 | 98.575 | 98.9333 | 99 | 99.2 | 99.18 | 98.775 | 99.025 | 99.2 | 99.4 | 99.2198
1 | 98.2667 | 98.675 | 98.86 | 99 | 98.9 | 98.0446 | 98.647 | 98.8 | 98.85 | 98.7802 | 98.1 | 98.7003 | 98.8 | 98.86 | 99 | 98.225 | 98.85 | 98.9 | 99.1 | 99.08 | 98.6 | 98.92 | 99.1 | 99.24 | 99.1198
1.1 | 97.9 | 98.6 | 98.8 | 98.9 | 98.8 | 97.6795 | 98.5603 | 98.8 | 98.775 | 98.6803 | 97.8333 | 98.5754 | 98.7 | 98.7 | 98.9 | 97.95 | 98.74 | 98.8429 | 99 | 99 | 98.475 | 98.7333 | 99.1 | 99.14 | 99.0198
1.2 | 97.6 | 98.5333 | 98.7 | 98.82 | 98.7 | 97.1861 | 98.4803 | 98.68 | 98.6 | 98.6803 | 97.3 | 98.4754 | 98.6 | 98.7 | 98.76 | 97.68 | 98.6 | 98.8 | 98.9 | 98.9 | 98.24 | 98.64 | 99 | 99.04 | 99.0198
1.3 | 97 | 98.5 | 98.6 | 98.8 | 98.6 | 96.6193 | 98.3303 | 98.5143 | 98.5333 | 98.5803 | 96.8 | 98.3504 | 98.52 | 98.6 | 98.7 | 97.28 | 98.4 | 98.72 | 98.8 | 98.88 | 98 | 98.54 | 98.9 | 99 | 99.0198
1.4 | 95.7 | 98.3333 | 98.5 | 98.7 | 98.6 | 95.8192 | 98.1804 | 98.41 | 98.5 | 98.4803 | 96.1 | 98.1755 | 98.42 | 98.5 | 98.6 | 96.88 | 98.25 | 98.5857 | 98.7 | 98.7 | 97.7 | 98.44 | 98.9 | 98.96 | 98.9198
1.5 | 93.7 | 98.2 | 98.4 | 98.6 | 98.4 | 94.4189 | 98.1304 | 98.31 | 98.375 | 98.3803 | 95.2 | 97.9755 | 98.38 | 98.4 | 98.5667 | 96.4 | 98.05 | 98.43 | 98.6 | 98.68 | 97.2333 | 98.3 | 98.8 | 98.86 | 98.9198
1.6 | 89.9 | 98.1 | 98.3 | 98.52 | 98.3 | 91.8683 | 98.0004 | 98.1667 | 98.2333 | 98.2803 | 93.6 | 97.8755 | 98.28 | 98.38 | 98.5 | 95.75 | 97.9 | 98.33 | 98.5 | 98.5 | 96.6 | 98.18 | 98.7 | 98.7 | 98.8198
1.7 | 82.2667 | 97.9 | 98.12 | 98.42 | 98.3 | 87.4175 | 97.8471 | 98.1 | 98.16 | 98.1804 | 90.95 | 97.7756 | 98.18 | 98.2 | 98.48 | 94.7 | 97.7 | 98.2364 | 98.4 | 98.4 | 96.02 | 98.1 | 98.6 | 98.6 | 98.7197
1.8 | 71.8667 | 97.8333 | 98 | 98.3 | 98.2 | 80.341 | 97.6805 | 98.1 | 98 | 98.0804 | 86.65 | 97.6256 | 98.08 | 98.08 | 98.28 | 93.25 | 97.525 | 98.1333 | 98.3 | 98.38 | 95.5 | 97.9 | 98.5 | 98.55 | 98.5197
1.9 | 59.4667 | 97.66 | 97.92 | 98.2 | 98.1 | 69.764 | 97.5055 | 98 | 97.9333 | 97.9804 | 78.82 | 97.3757 | 97.9 | 97.98 | 98.1 | 90.72 | 97.375 | 98.1 | 98.12 | 98.28 | 94.4333 | 97.8 | 98.4 | 98.5 | 98.4364
2 | 46.15 | 97.6 | 97.9 | 98.14 | 97.9833 | 57.8616 | 97.3555 | 97.925 | 97.9 | 97.8804 | 69.2 | 97.1757 | 97.8 | 97.9 | 98.08 | 86.2 | 97.225 | 97.9818 | 98.1 | 98.1 | 92.8 | 97.7 | 98.3 | 98.4 | 98.4197
2.1 | 33.85 | 97.46 | 97.8 | 98 | 97.9 | 45.6092 | 97.2805 | 97.8 | 97.8 | 97.7604 | 57.7 | 97.0757 | 97.7 | 97.8 | 97.98 | 79.9 | 97.1 | 97.89 | 98.02 | 98 | 90.125 | 97.5 | 98.1 | 98.22 | 98.3197
2.2 | 21.9667 | 97.3 | 97.66 | 97.88 | 97.74 | 33.7568 | 97.0556 | 97.6286 | 97.76 | 97.6205 | 46.15 | 96.9425 | 97.5833 | 97.74 | 97.88 | 70.76 | 96.88 | 97.7833 | 97.92 | 97.9 | 85.3 | 97.3 | 98 | 98.2 | 98.2197
2.3 | 13.7 | 97.1 | 97.56 | 97.74 | 97.64 | 23.0546 | 96.9056 | 97.52 | 97.66 | 97.5205 | 35 | 96.6425 | 97.4 | 97.7 | 97.8 | 60.2 | 96.75 | 97.5889 | 97.82 | 97.8 | 79.3667 | 97.1 | 97.9 | 98.1 | 98.0446
2.4 | 7.8 | 97 | 97.42 | 97.6571 | 97.56 | 15.5031 | 96.6807 | 97.34 | 97.46 | 97.4805 | 24.8 | 96.3959 | 97.3 | 97.525 | 97.74 | 48.875 | 96.6 | 97.44 | 97.72 | 97.7 | 70.6667 | 96.9 | 97.8 | 98.1 | 97.8767
2.5 | 5.2333 | 96.9 | 97.22 | 97.54 | 97.45 | 10.3021 | 96.5057 | 97.22 | 97.38 | 97.3805 | 17.3 | 96.076 | 97.2 | 97.4 | 97.6 | 38.3 | 96.35 | 97.325 | 97.7 | 97.7 | 59.95 | 96.8 | 97.7 | 97.975 | 97.8196
2.6 | 3.95 | 96.7 | 97.12 | 97.44 | 97.4 | 7.6615 | 96.3807 | 97.12 | 97.275 | 97.2805 | 12.45 | 95.901 | 97.1 | 97.3 | 97.5 | 28.325 | 96 | 97.1571 | 97.52 | 97.6 | 50.16 | 96.6 | 97.6 | 97.8333 | 97.7445
2.7 | 3.2 | 96.44 | 97.08 | 97.4 | 97.3 | 6.2012 | 96.2208 | 97.02 | 97.18 | 97.1806 | 9.7 | 95.7761 | 96.975 | 97.24 | 97.4 | 21.02 | 95.85 | 97.04 | 97.42 | 97.4 | 39.6333 | 96.4 | 97.5 | 97.7 | 97.6104
2.8 | 2.8 | 96.3 | 96.875 | 97.3 | 97.2 | 5.3011 | 95.8808 | 96.84 | 97.08 | 97.0306 | 8.2 | 95.5561 | 96.775 | 97.15 | 97.27 | 15.7 | 95.625 | 96.9 | 97.34 | 97.3 | 30.6 | 96.2 | 97.3 | 97.5333 | 97.5195
2.9 | 2.5 | 96.15 | 96.7667 | 97.1 | 97 | 4.7009 | 95.6809 | 96.72 | 96.98 | 96.9206 | 7.3 | 95.3762 | 96.675 | 97.1 | 97.14 | 12.7 | 95.3667 | 96.8714 | 97.2 | 97.2 | 23.2 | 95.9 | 97.22 | 97.44 | 97.5195
3 | 2.3 | 96 | 96.62 | 97 | 96.9 | 4.3009 | 95.5309 | 96.6273 | 96.88 | 96.8406 | 6.6 | 95.0163 | 96.575 | 96.96 | 97.02 | 11 | 95.15 | 96.7375 | 97.1 | 97.02 | 17.8667 | 95.6333 | 97.12 | 97.4 | 97.4195
3.1 | 2.2 | 95.75 | 96.6 | 96.95 | 96.8 | 4.0008 | 95.2809 | 96.53 | 96.7833 | 96.6807 | 6.25 | 94.7763 | 96.4 | 96.76 | 96.9273 | 9.9 | 94.85 | 96.6 | 97.025 | 96.92 | 14.7 | 95.42 | 97.02 | 97.34 | 97.3195
3.2 | 2.1 | 95.6 | 96.5 | 96.85 | 96.7 | 3.7007 | 95.031 | 96.43 | 96.6 | 96.5807 | 5.8 | 94.5264 | 96.3 | 96.66 | 96.83 | 9.1 | 94.625 | 96.5 | 96.9 | 96.9 | 12.85 | 95.1 | 96.9 | 97.3 | 97.2194
3.3 | 2 | 95.2667 | 96.425 | 96.75 | 96.6 | 3.5007 | 94.781 | 96.33 | 96.5 | 96.4807 | 5.56 | 94.2364 | 96.12 | 96.56 | 96.66 | 8.5667 | 94.4 | 96.4 | 96.8 | 96.8 | 11.6 | 94.9 | 96.72 | 97.1625 | 97.1194
3.4 | 1.9 | 95.1 | 96.28 | 96.7 | 96.6 | 3.3257 | 94.5144 | 96.16 | 96.4 | 96.3807 | 5.2 | 93.8766 | 96.02 | 96.5 | 96.53 | 8.175 | 94.14 | 96.2 | 96.7 | 96.62 | 10.7 | 94.9 | 96.52 | 97 | 97.0194
3.5 | 1.8 | 95 | 96.1 | 96.575 | 96.5 | 3.1756 | 94.2561 | 96.0364 | 96.3 | 96.2807 | 5 | 93.4766 | 95.92 | 96.36 | 96.43 | 7.85 | 93.7 | 96.12 | 96.6 | 96.52 | 10.26 | 94.5 | 96.4 | 96.8875 | 96.9194
3.6 | 1.7 | 94.7667 | 96 | 96.4 | 96.4 | 3.0006 | 94.0062 | 95.94 | 96.2 | 96.1808 | 4.825 | 93.1767 | 95.8 | 96.26 | 96.33 | 7.55 | 93.38 | 95.92 | 96.5 | 96.42 | 9.75 | 94.16 | 96.32 | 96.8 | 96.8194
3.7 | 1.6 | 94.48 | 95.9 | 96.3 | 96.2667 | 2.8756 | 93.6563 | 95.8333 | 96.1 | 96.0808 | 4.5667 | 92.9435 | 95.7 | 96.06 | 96.23 | 7.15 | 93.1 | 95.8 | 96.36 | 96.3 | 9.3 | 93.9 | 96.2333 | 96.7286 | 96.7193
3.8 | 1.6 | 94.24 | 95.8 | 96.2091 | 96.16 | 2.7005 | 93.3213 | 95.7 | 95.98 | 95.9808 | 4.375 | 92.6768 | 95.58 | 95.925 | 96.13 | 6.95 | 92.9 | 95.7 | 96.2 | 96.22 | 9.1333 | 93.7 | 96.14 | 96.6 | 96.6193
3.9 | 1.5 | 93.96 | 95.74 | 96.1 | 96.06 | 2.5755 | 92.9564 | 95.64 | 95.9 | 95.8808 | 4.225 | 92.2436 | 95.44 | 95.82 | 96.03 | 6.725 | 92.6 | 95.52 | 96.1 | 96.125 | 8.8333 | 93.32 | 95.925 | 96.5 | 96.5193
4 | 1.45 | 93.75 | 95.64 | 96 | 95.9 | 2.4255 | 92.7065 | 95.5 | 95.8 | 95.7808 | 4 | 91.927 | 95.3 | 95.72 | 95.9364 | 6.5 | 92.22 | 95.4 | 95.925 | 96.04 | 8.575 | 93.1 | 95.74 | 96.48 | 96.4993
4.1 | 1.4 | 93.5 | 95.5 | 95.9 | 95.8 | 2.3338 | 92.3149 | 95.4 | 95.74 | 95.6409 | 3.8333 | 91.5271 | 95.14 | 95.6 | 95.8143 | 6.3333 | 92 | 95.22 | 95.8 | 95.84 | 8.375 | 92.84 | 95.6 | 96.36 | 96.2993
4.2 | 1.3667 | 93 | 95.4 | 95.8 | 95.78 | 2.2404 | 91.9149 | 95.2 | 95.64 | 95.4809 | 3.7 | 91.0273 | 95 | 95.6 | 95.69 | 6.1 | 91.58 | 95.1 | 95.7 | 95.74 | 8.075 | 92.52 | 95.44 | 96.18 | 96.1992
4.3 | 1.3 | 92.475 | 95.3 | 95.7 | 95.6 | 2.2004 | 91.3484 | 95.1 | 95.5 | 95.4809 | 3.55 | 90.6773 | 95 | 95.5 | 95.575 | 5.95 | 91.175 | 94.9333 | 95.6 | 95.6 | 7.85 | 92.22 | 95.3 | 96.08 | 96.0992
4.4 | 1.2 | 91.7 | 95.14 | 95.6 | 95.5 | 2.0404 | 90.8319 | 95 | 95.4 | 95.3409 | 3.4 | 90.0975 | 94.8 | 95.4875 | 95.4 | 5.7 | 90.825 | 94.8 | 95.5 | 95.54 | 7.7 | 91.86 | 95.22 | 95.98 | 95.9192
4.5 | 1.2 | 90.2 | 94.94 | 95.5 | 95.4 | 2.0004 | 89.982 | 94.9 | 95.2 | 95.2809 | 3.3 | 89.6776 | 94.6 | 95.3571 | 95.28 | 5.5 | 90.4333 | 94.68 | 95.42 | 95.4 | 7.45 | 91.34 | 95.04 | 95.8818 | 95.8192
4.6 | 1.2 | 87.6 | 94.8 | 95.4 | 95.3 | 1.9004 | 89.1822 | 94.76 | 95.0833 | 95.161 | 3.2 | 89.1277 | 94.47 | 95.18 | 95.2 | 5.4 | 90.125 | 94.48 | 95.34 | 95.34 | 7.3 | 91.12 | 94.92 | 95.78 | 95.7191
4.7 | 1.1 | 82.325 | 94.7 | 95.3143 | 95.18 | 1.8004 | 87.4825 | 94.6 | 95 | 95.061 | 3 | 88.3779 | 94.34 | 95.1 | 95.14 | 5.25 | 89.6 | 94.3 | 95.22 | 95.2 | 7.05 | 90.82 | 94.66 | 95.59 | 95.6024
4.8 | 1.1 | 74.025 | 94.6 | 95.21 | 95.0833 | 1.8004 | 84.6831 | 94.46 | 94.9 | 94.861 | 3 | 87.678 | 94.17 | 94.98 | 95.04 | 5.1 | 89.1 | 94.18 | 95.14 | 95.1 | 6.9 | 90.5 | 94.52 | 95.49 | 95.4941
4.9 | 1.1 | 63.375 | 94.44 | 95.0833 | 95 | 1.7003 | 79.6341 | 94.36 | 94.8 | 94.756 | 2.8333 | 86.1785 | 94.04 | 94.78 | 94.94 | 4.975 | 88.6 | 93.98 | 95 | 94.96 | 6.7333 | 90.08 | 94.4273 | 95.39 | 95.4191
5 | 1.1 | 52.6 | 94.3 | 94.9286 | 94.8 | 1.6253 | 72.4855 | 94.16 | 94.7 | 94.6811 | 2.775 | 83.9791 | 93.87 | 94.6 | 94.74 | 4.82 | 88.0667 | 93.725 | 94.9 | 94.8714 | 6.6 | 89.68 | 94.33 | 95.28 | 95.3191
5.1 | 1 | 40.7 | 94.2 | 94.9 | 94.7 | 1.6003 | 63.1207 | 94 | 94.66 | 94.5811 | 2.6333 | 80.2799 | 93.77 | 94.5833 | 94.6 | 4.675 | 87.5333 | 93.6 | 94.84 | 94.8 | 6.42 | 89.22 | 94.16 | 95.075 | 95.1523
5.2 | 1 | 29.7 | 94.025 | 94.74 | 94.6 | 1.5253 | 52.2646 | 93.86 | 94.56 | 94.4811 | 2.58 | 74.1815 | 93.65 | 94.5 | 94.5 | 4.525 | 86.9 | 93.4667 | 94.7 | 94.66 | 6.225 | 88.78 | 94 | 94.9 | 95.019
5.3 | 1 | 20.05 | 93.8333 | 94.6 | 94.5 | 1.5003 | 41.2417 | 93.66 | 94.36 | 94.3811 | 2.5 | 66.9083 | 93.46 | 94.4 | 94.45 | 4.4 | 85.8 | 93.3 | 94.6 | 94.42 | 6.1 | 88.42 | 93.76 | 94.8 | 94.8523
5.4 | 1 | 13.4 | 93.75 | 94.44 | 94.3 | 1.4336 | 31.4687 | 93.56 | 94.2 | 94.2811 | 2.4 | 57.5256 | 93.2429 | 94.3 | 94.3 | 4.3 | 84.6333 | 93.225 | 94.45 | 94.23 | 5.9 | 87.92 | 93.56 | 94.7 | 94.7189
5.5 | 0.9 | 9.6 | 93.66 | 94.24 | 94.2 | 1.4003 | 22.7954 | 93.4 | 94.1 | 94.1612 | 2.3 | 48.138 | 93.1 | 94.1 | 94.2 | 4.2 | 82.5 | 93 | 94.3 | 94.13 | 5.9 | 87.28 | 93.36 | 94.6833 | 94.6189
5.6 | 0.9 | 7.8 | 93.5 | 94.2 | 94.1167 | 1.4003 | 16.5967 | 93.3 | 94 | 93.9612 | 2.3 | 37.9705 | 92.9375 | 94 | 94.1 | 4.1 | 79.45 | 92.8 | 94.15 | 94.1 | 5.725 | 86.72 | 93.175 | 94.6 | 94.4689
5.7 | 0.9 | 6.75 | 93.26 | 94.2 | 94.0167 | 1.3003 | 13.0974 | 93.18 | 93.92 | 93.7812 | 2.2 | 29.3927 | 92.76 | 93.8 | 94.06 | 4 | 74.35 | 92.65 | 93.96 | 94 | 5.6 | 86 | 93.05 | 94.5 | 94.3189
5.8 | 0.9 | 6.2 | 93.16 | 93.9 | 93.92 | 1.3003 | 11.0378 | 92.9 | 93.82 | 93.6813 | 2.1 | 22.4694 | 92.575 | 93.7 | 93.86 | 3.9 | 67.74 | 92.425 | 93.86 | 93.96 | 5.5 | 85.2 | 92.8667 | 94.4 | 94.2188
5.9 | 0.9 | 5.7333 | 93.05 | 93.8 | 93.7 | 1.3003 | 9.948 | 92.78 | 93.62 | 93.5813 | 2.1 | 18.2754 | 92.42 | 93.6 | 93.66 | 3.8 | 59.6333 | 92.3 | 93.76 | 93.7889 | 5.4333 | 83.72 | 92.73 | 94.24 | 94.1188
6 | 0.9 | 5.5 | 92.9 | 93.66 | 93.5 | 1.2002 | 9.1982 | 92.68 | 93.5 | 93.4646 | 2 | 15.4361 | 92.24 | 93.4182 | 93.5 | 3.7 | 50.9 | 92.16 | 93.7 | 93.6333 | 5.3 | 81.65 | 92.5727 | 94.15 | 94.1188
6.1 | 0.8 | 5.3 | 92.6167 | 93.5 | 93.4 | 1.2002 | 8.5983 | 92.48 | 93.32 | 93.3613 | 1.975 | 13.9965 | 91.9857 | 93.31 | 93.3667 | 3.6 | 42.175 | 91.9 | 93.56 | 93.4 | 5.2 | 78.4 | 92.38 | 93.94 | 93.9855
6.2 | 0.8 | 5.1 | 92.52 | 93.3667 | 93.3 | 1.2002 | 8.1984 | 92.28 | 93.22 | 93.1814 | 1.9 | 12.9968 | 91.8 | 93.12 | 93.2 | 3.525 | 34.84 | 91.7 | 93.36 | 93.3 | 5.1 | 73.32 | 92.18 | 93.8 | 93.8855
6.3 | 0.8 | 4.94 | 92.4 | 93.225 | 93.2 | 1.2002 | 7.7485 | 92.1 | 93.1 | 93.0814 | 1.875 | 12.247 | 91.68 | 93.01 | 93.075 | 3.425 | 28.1667 | 91.5 | 93.2 | 93.2 | 5 | 67.65 | 91.98 | 93.76 | 93.7187
6.4 | 0.8 | 4.8 | 92.3 | 93.1 | 93.08 | 1.1002 | 7.4985 | 91.94 | 92.92 | 92.9814 | 1.8 | 11.6971 | 91.48 | 92.82 | 92.9333 | 3.3 | 23.6 | 91.36 | 93.1 | 93.1 | 4.9 | 60.15 | 91.75 | 93.7 | 93.6187
6.5 | 0.8 | 4.7 | 92.12 | 93.02 | 92.9143 | 1.1002 | 7.1652 | 91.825 | 92.84 | 92.8814 | 1.8 | 11.2472 | 91.28 | 92.71 | 92.84 | 3.3 | 20.6333 | 91.12 | 92.96 | 92.98 | 4.75 | 52.76 | 91.58 | 93.56 | 93.5187
6.6 | 0.8 | 4.5 | 91.82 | 92.925 | 92.7667 | 1.1002 | 6.8736 | 91.54 | 92.74 | 92.6815 | 1.7333 | 10.8306 | 91.08 | 92.6 | 92.8 | 3.2 | 19.1 | 90.8333 | 92.6714 | 92.88 | 4.7 | 44.84 | 91.39 | 93.3 | 93.3937
6.7 | 0.8 | 4.4 | 91.72 | 92.725 | 92.6333 | 1.1002 | 6.7237 | 91.4 | 92.54 | 92.4815 | 1.7 | 10.3974 | 90.78 | 92.4857 | 92.6333 | 3.125 | 17.9 | 90.46 | 92.5375 | 92.68 | 4.6 | 37.62 | 91.27 | 93.175 | 93.1936
6.8 | 0.8 | 4.2667 | 91.62 | 92.6 | 92.54 | 1.0002 | 6.4237 | 91.3 | 92.34 | 92.3815 | 1.7 | 10.1225 | 90.58 | 92.38 | 92.54 | 3.1 | 17 | 90.3 | 92.5 | 92.48 | 4.6 | 31.5 | 90.975 | 93.06 | 93.0186
6.9 | 0.8 | 4.12 | 91.42 | 92.425 | 92.4111 | 1.0002 | 6.2238 | 91.0667 | 92.3 | 92.3815 | 1.6 | 9.7642 | 90.4 | 92.2667 | 92.34 | 3 | 16.5 | 90.18 | 92.4 | 92.38 | 4.5 | 27.18 | 90.7143 | 92.96 | 92.8386
7 | 0.8 | 4 | 91.3 | 92.28 | 92.2364 | 1.0002 | 6.0488 | 90.9 | 92.14 | 92.1816 | 1.6 | 9.4726 | 90.3 | 92.0667 | 92.2 | 3 | 15.875 | 89.98 | 92.28 | 92.18 | 4.4 | 24.2 | 90.6 | 92.8 | 92.6385
7.1 | 0.7 | 3.925 | 91.0333 | 92.1 | 92.1111 | 1.0002 | 5.8988 | 90.74 | 91.94 | 92.0149 | 1.6 | 9.1977 | 89.9833 | 91.9 | 92.04 | 2.9 | 15.55 | 89.78 | 92.08 | 92 | 4.3 | 22.42 | 90.3833 | 92.6 | 92.5385
7.2 | 0.7 | 3.8 | 90.925 | 92 | 91.92 | 1.0002 | 5.6489 | 90.54 | 91.84 | 91.8816 | 1.5 | 8.8978 | 89.8 | 91.8 | 91.94 | 2.9 | 15.15 | 89.56 | 91.9143 | 91.9 | 4.28 | 21.25 | 90.1889 | 92.5 | 92.4185
7.3 | 0.7 | 3.7 | 90.74 | 91.9 | 91.81 | 0.96019 | 5.3989 | 90.34 | 91.75 | 91.7816 | 1.5 | 8.6478 | 89.5 | 91.6167 | 91.74 | 2.8 | 14.625 | 89.26 | 91.72 | 91.8 | 4.2 | 20.3 | 89.95 | 92.3333 | 92.2518
7.4 | 0.7 | 3.6 | 90.52 | 91.7 | 91.7182 | 0.90018 | 5.2989 | 90.14 | 91.66 | 91.6017 | 1.5 | 8.3979 | 89.26 | 91.4 | 91.6 | 2.7 | 14.275 | 88.98 | 91.62 | 91.64 | 4.1 | 19.78 | 89.75 | 92.2 | 92.1384
7.5 | 0.7 | 3.6 | 90.34 | 91.6 | 91.54 | 0.90018 | 5.099 | 89.94 | 91.5 | 91.4817 | 1.4 | 8.098 | 89.15 | 91.3333 | 91.5 | 2.7 | 14.025 | 88.78 | 91.5 | 91.46 | 4.1 | 19.275 | 89.54 | 91.9 | 92.0184
7.6 | 0.7 | 3.5 | 90.1 | 91.5 | 91.34 | 0.90018 | 5.0657 | 89.7 | 91.3 | 91.3017 | 1.4 | 7.898 | 88.82 | 91.2 | 91.32 | 2.625 | 13.65 | 88.56 | 91.32 | 91.36 | 3.94 | 18.8333 | 89.3 | 91.8 | 91.8584
7.7 | 0.7 | 3.4 | 89.98 | 91.3 | 91.22 | 0.90018 | 4.799 | 89.44 | 91.2 | 91.1018 | 1.4 | 7.6581 | 88.55 | 91.06 | 91.1714 | 2.6 | 13.38 | 88.28 | 91.225 | 91.16 | 3.9 | 18.4333 | 89.1 | 91.7 | 91.7183
7.8 | 0.7 | 3.3 | 89.78 | 91.2 | 91.04 | 0.90018 | 4.6991 | 89.14 | 91 | 90.9818 | 1.4 | 7.3982 | 88.3 | 90.9 | 91 | 2.6 | 13 | 88.0667 | 91.04 | 91 | 3.8 | 17.9 | 88.9 | 91.5667 | 91.6583
7.9 | 0.7 | 3.24 | 89.58 | 91.1 | 91 | 0.90018 | 4.5991 | 88.96 | 90.86 | 90.8818 | 1.3 | 7.2232 | 88.1 | 90.75 | 90.8 | 2.5 | 12.75 | 87.84 | 90.84 | 90.875 | 3.8 | 17.5667 | 88.7 | 91.4 | 91.3583
8 | 0.7 | 3.2 | 89.46 | 90.9 | 90.8 | 0.90018 | 4.4991 | 88.7 | 90.7 | 90.7019 | 1.3 | 7.0316 | 87.9 | 90.54 | 90.65 | 2.5 | 12.4 | 87.42 | 90.7 | 90.76 | 3.7 | 17.2667 | 88.5 | 91.3 | 91.1582
8.1 | 0.7 | 3.1 | 89.1833 | 90.7 | 90.6 | 0.80016 | 4.3991 | 88.4 | 90.56 | 90.5019 | 1.3 | 6.8583 | 87.62 | 90.34 | 90.6 | 2.4 | 12.15 | 87.14 | 90.54 | 90.62 | 3.7 | 16.9333 | 88.16 | 91.225 | 91.0582
8.2 | 0.7 | 3 | 88.9667 | 90.5556 | 90.5 | 0.80016 | 4.2991 | 88.1 | 90.4 | 90.3019 | 1.3 | 6.6983 | 87.25 | 90.25 | 90.5091 | 2.4 | 11.86 | 86.94 | 90.4 | 90.4 | 3.6 | 16.575 | 87.8167 | 91.1 | 90.8582
8.3 | 0.6 | 3 | 88.78 | 90.4 | 90.36 | 0.80016 | 4.1992 | 88 | 90.2714 | 90.202 | 1.3 | 6.4984 | 86.94 | 90.16 | 90.3111 | 2.3 | 11.6 | 86.76 | 90.2667 | 90.2 | 3.6 | 16.3 | 87.52 | 90.98 | 90.7782
8.4 | 0.6 | 3 | 88.6 | 90.3167 | 90.16 | 0.80016 | 4.0992 | 87.72 | 90.1167 | 90.082 | 1.2 | 6.3984 | 86.6667 | 89.86 | 90.1545 | 2.3 | 11.32 | 86.56 | 90.1 | 90.1143 | 3.5 | 15.875 | 87.32 | 90.8 | 90.6781
8.5 | 0.6 | 2.9 | 88.3 | 90.1429 | 90 | 0.80016 | 3.9992 | 87.4 | 90.02 | 89.9153 | 1.2 | 6.1985 | 86.5 | 89.76 | 90.02 | 2.2 | 11.1 | 86.22 | 89.975 | 89.925 | 3.4 | 15.6 | 87.12 | 90.7 | 90.4781
8.6 | 0.6 | 2.875 | 88 | 89.9 | 89.8 | 0.80016 | 3.8992 | 87.1 | 89.82 | 89.782 | 1.2 | 6.0985 | 86.1667 | 89.66 | 89.82 | 2.2 | 10.8 | 85.95 | 89.8 | 89.74 | 3.4 | 15.4571 | 86.92 | 90.4 | 90.3848
8.7 | 0.6 | 2.8 | 87.8 | 89.8 | 89.775 | 0.80016 | 3.7992 | 86.9 | 89.62 | 89.6821 | 1.2 | 5.9319 | 85.85 | 89.5 | 89.62 | 2.2 | 10.6 | 85.66 | 89.7 | 89.6 | 3.4 | 15.1 | 86.64 | 90.3 | 90.218
8.8 | 0.6 | 2.8 | 87.6 | 89.73 | 89.7 | 0.80016 | 3.7659 | 86.6 | 89.51 | 89.4621 | 1.2 | 5.7985 | 85.525 | 89.36 | 89.46 | 2.1 | 10.4333 | 85.32 | 89.5 | 89.44 | 3.3 | 14.8286 | 86.42 | 90 | 90.118
8.9 | 0.6 | 2.7 | 87.4 | 89.63 | 89.5 | 0.80016 | 3.6743 | 86.18 | 89.32 | 89.2222 | 1.1 | 5.5986 | 85.2333 | 89.16 | 89.32 | 2.1 | 10.2 | 85.0429 | 89.5 | 89.24 | 3.25 | 14.56 | 86.12 | 89.9 | 89.943
9 | 0.6 | 2.7 | 87.2 | 89.53 | 89.4 | 0.80016 | 3.5993 | 85.85 | 89.21 | 89.0822 | 1.1 | 5.4986 | 84.8 | 88.9667 | 89.12 | 2.1 | 9.9333 | 84.8167 | 89.25 | 89 | 3.2 | 14.2818 | 85.88 | 89.8 | 89.7379
9.1 |  | 2.65 | 86.9 | 89.43 | 89.3 | 0.80016 | 3.4993 | 85.54 | 89.11 | 89.0222 | 1.1 | 5.4236 | 84.4 | 88.8667 | 88.96 | 2.1 | 9.7 | 84.425 | 89 | 88.9 | 3.2 | 14 | 85.6667 | 89.6 | 89.5579
9.2 |  | 2.6 | 86.6 | 89.4 | 89.3 | 0.70014 | 3.4243 | 85.22 | 88.9364 | 88.8822 | 1.1 | 5.2987 | 84.1 | 88.68 | 88.86 | 2 | 9.6 | 84.04 | 88.75 | 88.74 | 3.1 | 13.76 | 85.2667 | 89.4 | 89.4179
9.3 |  | 2.55 | 86.4 | 89.32 | 89.2 | 0.70014 | 3.3743 | 85.02 | 88.82 | 88.7822 | 1.1 | 5.0987 | 83.9 | 88.5 | 88.76 | 2 | 9.45 | 83.6455 | 88.5 | 88.55 | 3.1 | 13.56 | 84.9 | 89.3 | 89.2578
9.4 |  | 2.5 | 86 | 89.3 | 89.2 | 0.70014 | 3.2993 | 84.64 | 88.72 | 88.7822 | 1.1 | 5.0237 | 83.4333 | 88.4 | 88.6 | 2 | 9.2 | 83.3778 | 88.4 | 88.38 | 3.1 | 13.28 | 84.55 | 89.1 | 89.0578
9.5 |  | 2.5 | 85.0833 | 89.3 | 89.2 |  | 3.2743 | 84.32 | 88.7 | 88.6823 | 1.05 | 4.9788 | 83.1 | 88.3 | 88.3857 | 1.9 | 9.05 | 83.16 | 88.175 | 88.225 | 3.1 | 12.96 | 84.2667 | 88.9 | 88.8578
9.6 |  | 2.42 | 83.9 | 89.24 | 89.1 |  | 3.1994 | 83.8 | 88.62 | 88.5823 | 1 | 4.8238 | 82.65 | 88.12 | 88.26 | 1.9 | 8.8 | 82.82 | 88.06 | 88 | 3 | 12.78 | 84.1 | 88.7167 | 88.6577
9.7 |  | 2.4 | 81.2 | 89.2 | 89.1 |  | 3.1994 | 83.3 | 88.52 | 88.4823 | 1 | 4.6988 | 82.3 | 87.92 | 88.12 | 1.9 | 8.65 | 82.475 | 87.9 | 87.9 | 3 | 12.58 | 83.84 | 88.52 | 88.6177
9.8 |  | 2.4 | 76.1 | 89.2 | 89.1 |  | 3.0994 | 82.6 | 88.5 | 88.4823 | 1 | 4.5988 | 81.9 | 87.82 | 87.9636 | 1.9 | 8.5 | 82.12 | 87.76 | 87.7 | 2.9 | 12.38 | 83.625 | 88.32 | 88.3677
9.9 |  | 2.375 | 68.3 | 89.2 | 89.1 |  | 2.9994 | 81.6 | 88.5 | 88.3823 |  | 4.5988 | 81.5 | 87.7333 | 87.87 | 1.85 | 8.35 | 81.7333 | 87.52 | 87.5 | 2.9 | 12.18 | 83.18 | 88.22 | 88.1676
10 |  | 2.3 | 59.5333 | 89.2 | 89.1 |  | 2.9994 | 80.05 | 88.42 | 88.3823 |  | 4.4989 | 81.1 | 87.7 | 87.77 | 1.8 | 8.2 | 81.425 | 87.3 | 87.3 | 2.8 | 11.94 | 82.84 | 88.02 | 87.9776
10.1 |  | 2.3 | 49.1 | 89.2 | 89 |  | 2.9994 | 77.2 | 88.4 | 88.2823 |  | 4.3989 | 80.6 | 87.54 | 87.67 | 1.8 | 8.0333 | 80.9667 | 87.175 | 87.2167 | 2.8 | 11.8 | 82.58 | 87.74 | 87.7776
10.2 |  | 2.3 | 38.48 | 89.2 | 89 |  | 2.8994 | 72.14 | 88.32 | 88.2823 |  | 4.2989 | 80.1 | 87.4 | 87.57 | 1.7 | 7.9 | 80.5667 | 86.98 | 87.1 | 2.8 | 11.65 | 82.2 | 87.62 | 87.6175
10.3 |  | 2.3 | 29.25 | 89.2 | 89 |  | 2.8994 | 64.8333 | 88.3 | 88.1824 |  | 4.2489 | 79.5 | 87.4 | 87.47 | 1.7 | 7.74 | 80.25 | 86.8 | 86.9 | 2.8 | 11.4 | 81.9 | 87.52 | 87.4775
10.4 |  | 2.2 | 20.98 | 89.1 | 89 |  | 2.7994 | 56.2667 | 88.24 | 88.1824 |  | 4.1989 | 78.4 | 87.3 | 87.4 | 1.7 | 7.68 | 79.7 | 86.68 | 86.6333 | 2.7 | 11.25 | 81.45 | 87.34 | 87.2775
10.5 |  | 2.2 | 15.64 | 89.1 | 89 |  | 2.7994 | 46.8667 | 88.1714 | 88.0824 |  | 4.099 | 76.85 | 87.24 | 87.3 | 1.7 | 7.5333 | 79.4 | 86.5 | 86.5 | 2.7 | 10.98 | 81.18 | 87.12 | 86.9774
10.6 |  | 2.2 | 12.8 | 89.1 | 88.9 |  | 2.7994 | 37.8667 | 88.1 | 87.9824 |  | 4.099 | 74.1 | 87.14 | 87.16 | 1.7 | 7.4 | 79 | 86.38 | 86.4 | 2.6 | 10.875 | 80.88 | 87 | 86.7774
10.7 |  | 2.2 | 11.2 | 88.96 | 88.9 |  | 2.6995 | 29.45 | 88 | 87.9824 |  | 3.999 | 69.64 | 86.96 | 87.1 | 1.6333 | 7.28 | 78.5 | 86.18 | 86.2667 | 2.6 | 10.7 | 80.6 | 86.74 | 86.684
10.8 |  | 2.125 | 10.38 | 88.86 | 88.8 |  | 2.6995 | 22.675 | 87.95 | 87.8824 |  | 3.899 | 63.28 | 86.86 | 87.1 |  | 7.1 | 78.2 | 86 | 86.0667 | 2.5 | 10.5 | 80.12 | 86.525 | 86.4773
10.9 |  | 2.1 | 9.85 | 88.8 | 88.8 |  | 2.6245 | 18.125 | 87.88 | 87.8824 |  | 3.824 | 55.9667 | 86.8 | 87 |  | 7 | 77.7 | 85.9 | 85.8667 | 2.5 | 10.4 | 79.72 | 86.3 | 86.2923
11 |  | 2.1 | 9.5 | 88.8 | 88.8 |  | 2.5995 | 15.66 | 87.8 | 87.8424 |  | 3.799 | 47.5 | 86.7 | 86.88 |  | 6.9 | 77.2 | 85.8 | 85.7667 | 2.5 | 10.3 | 79.36 | 86.2 | 86.0972
11.1 |  | 2.1 | 9.2333 | 88.8 | 88.7 |  | 2.5795 | 14.3 | 87.78 | 87.7824 |  | 3.6991 | 39.6 | 86.66 | 86.78 |  | 6.75 | 76.775 | 85.7 | 85.6 | 2.5 | 10.1 | 79.02 | 85.9 | 85.9172
11.2 |  | 2.0333 | 9.02 | 88.7 | 88.66 |  | 2.4995 | 13.4 | 87.7 | 87.6825 |  | 3.5991 | 32.44 | 86.6 | 86.7 |  | 6.625 | 76.2 | 85.6 | 85.58 | 2.4 | 9.9 | 78.58 | 85.8 | 85.8172
11.3 |  | 2 | 8.8667 | 88.7 | 88.58 |  | 2.4995 | 12.78 | 87.68 | 87.5825 |  | 3.5991 | 26.32 | 86.5 | 86.56 |  | 6.5667 | 75.6 | 85.5 | 85.4 | 2.4 | 9.7 | 78.15 | 85.62 | 85.6546
11.4 |  | 2 | 8.64 | 88.6 | 88.5 |  | 2.4995 | 12.32 | 87.58 | 87.5425 |  | 3.4991 | 22.08 | 86.4 | 86.46 |  | 6.4 | 74.5 | 85.3 | 85.2667 | 2.4 | 9.5667 | 77.7333 | 85.4 | 85.4171
11.5 |  | 2 | 8.5 | 88.6 | 88.48 |  | 2.4495 | 11.94 | 87.5 | 87.4425 |  | 3.4991 | 19.54 | 86.3 | 86.38 |  | 6.3 | 72.9 | 85.1 | 85.06 | 2.4 | 9.5 | 77.36 | 85.28 | 85.217
11.6 |  | 2 | 8.3333 | 88.5 | 88.38 |  | 2.3995 | 11.58 | 87.5 | 87.3825 |  | 3.3992 | 18.02 | 86.1857 | 86.3 |  | 6.225 | 70.6667 | 85.1 | 84.9 | 2.3333 | 9.3 | 76.9667 | 85.08 | 85.077
11.7 |  | 1.9667 | 8.2 | 88.5 | 88.3 |  | 2.3995 | 11.28 | 87.4 | 87.2825 |  | 3.3992 | 17.2 | 86.1 | 86.18 |  | 6.1 | 66.82 | 84.9 | 84.82 |  | 9.14 | 76.76 | 84.925 | 84.877
11.8 |  | 1.9 | 8.04 | 88.48 | 88.2 |  | 2.3995 | 11.04 | 87.3125 | 87.2625 |  | 3.2992 | 16.425 | 85.9429 | 86 |  | 6.04 | 61.775 | 84.8 | 84.72 |  | 9.1 | 76.34 | 84.74 | 84.6769
11.9 |  | 1.9 | 7.9 | 88.34 | 88.2 |  | 2.2995 | 10.74 | 87.2333 | 87.1626 |  | 3.2992 | 15.95 | 85.89 | 85.9 |  | 6 | 55.3 | 84.68 | 84.52 |  | 8.9 | 75.88 | 84.54 | 84.6169
12 |  | 1.9 | 7.8 | 88.3 | 88.1 |  | 2.2995 | 10.48 | 87.14 | 86.9826 |  | 3.2992 | 15.48 | 85.7889 | 85.8 |  | 5.9 | 48.725 | 84.4 | 84.425 |  | 8.8 | 75.46 | 84.44 | 84.4769
12.1 |  |  | 7.6 | 88.2 | 88.1 |  | 2.2995 | 10.24 | 87.04 | 86.8826 |  | 3.1992 | 15.05 | 85.6818 | 85.7 |  | 5.8 | 42.1 | 84.225 | 84.3 |  | 8.7 | 74.84 | 84.34 | 84.2769
12.2 |  |  | 7.5 | 88.2 | 88.1 |  | 2.2745 | 10.04 | 87 | 86.8826 |  | 3.1992 | 14.7667 | 85.58 | 85.6 |  | 5.65 | 36.06 | 84.1667 | 84.14 |  | 8.6 | 74.4286 | 84.24 | 84.1835
12.3 |  |  | 7.4 | 88.1 | 88 |  | 2.1996 | 9.85 | 86.9 | 86.7826 |  | 3.0992 | 14.44 | 85.39 | 85.475 |  | 5.6 | 30.96 | 84.1 | 84 |  | 8.42 | 73.5 | 84.04 | 84.0768
12.4 |  |  | 7.3 | 88.1 | 87.9 |  | 2.1996 | 9.65 | 86.8 | 86.6827 |  | 3.0992 | 14.1 | 85.28 | 85.3 |  | 5.52 | 27.18 | 83.925 | 83.84 |  | 8.32 | 72.5 | 83.84 | 83.8968
12.5 |  |  | 7.1 | 88 | 87.8 |  | 2.1996 | 9.425 | 86.7 | 86.4827 |  | 2.9992 | 13.7667 | 85.09 | 85.14 |  | 5.4667 | 24.9 | 83.7 | 83.7 |  | 8.2 | 71.18 | 83.66 | 83.6967
12.6 |  |  | 6.9833 | 88 | 87.8 |  |  | 9.2667 | 86.625 | 86.4827 |  | 2.9992 | 13.48 | 84.9909 | 85 |  | 5.4 | 23.5 | 83.54 | 83.5 |  | 8.12 | 69.34 | 83.46 | 83.5967
12.7 |  |  | 6.9 | 87.9 | 87.7 |  |  | 9.08 | 86.49 | 86.3827 |  | 2.9992 | 13.25 | 84.9 | 84.9 |  | 5.3 | 22.3333 | 83.34 | 83.3 |  | 8 | 65.7571 | 83.36 | 83.4967
12.8 |  |  | 6.7889 | 87.85 | 87.65 |  |  | 8.88 | 86.375 | 86.2827 |  | 2.8993 | 13 | 84.7 | 84.7 |  | 5.2 | 21.8 | 83.16 | 83.1 |  | 7.9 | 61.7 | 83.2 | 83.2967
12.9 |  |  | 6.6818 | 87.8 | 87.6 |  |  | 8.7 | 86.2 | 86.1661 |  | 2.8993 | 12.675 | 84.6 | 84.6 |  | 5.1 | 21.2 | 83.06 | 82.9 |  | 7.8 | 56.14 | 83.0818 | 83.0966
13 |  |  | 6.59 | 87.8 | 87.6 |  |  | 8.5667 | 86.14 | 86.0828 |  | 2.866 | 12.42 | 84.4 | 84.4 |  | 5 | 20.6667 | 82.86 | 82.8 |  | 7.7 | 50.58 | 82.9889 | 82.7665
13.1 |  |  | 6.49 | 87.7714 | 87.5 |  |  | 8.35 | 86.04 | 85.9828 |  | 2.7993 | 12.1 | 84.25 | 84.3 |  | 5 | 20.3 | 82.76 | 82.6 |  | 7.6 | 45.2 | 82.86 | 82.6165
13.2 |  |  | 6.4 | 87.65 | 87.4286 |  |  | 8.26 | 85.94 | 85.8828 |  |  | 11.975 | 84.1375 | 84.125 |  | 4.9 | 19.9 | 82.56 | 82.5 |  | 7.5 | 39.5286 | 82.7125 | 82.4965
13.3 |  |  | 6.38 | 87.6 | 87.4 |  |  | 8.1 | 85.78 | 85.7204 |  |  | 11.75 | 84.0714 | 83.96 |  | 4.84 | 19.55 | 82.46 | 82.4 |  | 7.4 | 34.82 | 82.48 | 82.2564
13.4 |  |  | 6.2 | 87.6 | 87.3 |  |  | 8 | 85.7 | 85.5829 |  |  | 11.4 | 83.875 | 83.8429 |  | 4.8 | 19.18 | 82.26 | 82.2 |  | 7.24 | 31.34 | 82.28 | 82.1564
13.5 |  |  | 6.19 | 87.5091 | 87.25 |  |  | 7.8333 | 85.58 | 85.4829 |  |  | 11.2 | 83.66 | 83.6727 |  | 4.75 | 18.86 | 81.975 | 82 |  | 7.2 | 28.89 | 82.09 | 82.0564
13.6 |  |  | 6.09 | 87.41 | 87.2 |  |  | 7.7 | 85.5 | 85.2829 |  |  | 11 | 83.5 | 83.48 |  | 4.7 | 18.52 | 81.8667 | 81.8 |  | 7.1 | 27.35 | 81.98 | 81.8664
13.7 |  |  | 5.99 | 87.31 | 87.1 |  |  | 7.6 | 85.3909 | 85.183 |  |  | 10.8167 | 83.28 | 83.34 |  | 4.6 | 18.2 | 81.6667 | 81.62 |  | 7.1 | 26.28 | 81.7833 | 81.7163
13.8 |  |  | 5.9 | 87.21 | 87.05 |  |  | 7.46 | 85.3 | 84.983 |  |  | 10.7 | 83.1 | 83.12 |  | 4.5 | 17.86 | 81.48 | 81.42 |  | 7 | 25.54 | 81.5333 | 81.6163
13.9 |  |  | 5.89 | 87.11 | 86.875 |  |  | 7.3875 | 85.1 | 84.9205 |  |  | 10.4 | 82.98 | 82.94 |  | 4.5 | 17.5 | 81.28 | 81.22 |  | 6.9 | 24.98 | 81.325 | 81.3763
14 |  |  | 5.7909 | 87.01 | 86.72 |  |  | 7.24 | 84.96 | 84.743 |  |  | 10.3 | 82.78 | 82.78 |  | 4.4 | 17.3 | 81 | 80.92 |  | 6.8 | 24.5333 | 81.14 | 81.2162
14.1 |  |  | 5.7 | 86.91 | 86.56 |  |  | 7.15 | 84.8 | 84.5831 |  |  | 10.12 | 82.58 | 82.58 |  | 4.4 | 17.025 | 80.88 | 80.8167 |  | 6.7 | 24.15 | 80.9667 | 81.0495
14.2 |  |  | 5.6 | 86.63 | 86.38 |  |  | 7 | 84.7 | 84.3831 |  |  | 9.9667 | 82.38 | 82.32 |  | 4.3 | 16.68 | 80.675 | 80.7 |  | 6.6 | 23.78 | 80.7 | 80.9495
14.3 |  |  | 5.6 | 86.42 | 86.08 |  |  | 6.97 | 84.4 | 84.2831 |  |  | 9.7833 | 82.18 | 82.14 |  | 4.3 | 16.4 | 80.425 | 80.36 |  | 6.5667 | 23.3333 | 80.5 | 80.6494
14.4 |  |  | 5.5 | 86.1286 | 85.78 |  |  | 6.8625 | 84.14 | 84.0232 |  |  | 9.6 | 82 | 81.98 |  |  | 16.2 | 80.3 | 80.16 |  | 6.5 | 22.96 | 80.4 | 80.3661
14.5 |  |  | 5.4 | 85.42 | 85.3 |  |  | 6.8 | 83.92 | 83.8032 |  |  | 9.5 | 81.86 | 81.78 |  |  | 15.9 | 80.1 | 79.9667 |  | 6.4 | 22.58 | 80.2167 | 80.1493
14.6 |  |  | 5.4 | 84.48 | 84.4 |  |  | 6.6 | 83.64 | 83.5133 |  |  | 9.3 | 81.58 | 81.52 |  |  | 15.7 | 79.82 | 79.7667 |  | 6.3 | 22.24 | 80.075 | 79.966
14.7 |  |  | 5.3 | 82.66 | 82.77 |  |  | 6.5167 | 83.34 | 83.2134 |  |  | 9.15 | 81.28 | 81.2909 |  |  | 15.4 | 79.5667 | 79.52 |  | 6.2 | 21.94 | 79.875 | 79.6659
14.8 |  |  | 5.3 | 79.1 | 79.63 |  |  | 6.4167 | 83.02 | 82.8234 |  |  | 9.0167 | 81.08 | 81.1 |  |  | 15.2 | 79.32 | 79.3143 |  | 6.2 | 21.7 | 79.7 | 79.3826
14.9 |  |  | 5.2 | 73.14 | 73.99 |  |  | 6.4 | 82.56 | 82.4235 |  |  | 8.9 | 80.88 | 80.8 |  |  | 15 | 79.08 | 79.1 |  | 6.1 | 21.4 | 79.5 | 79.2825
15 |  |  | 5.2 | 64.98 | 65.93 |  |  | 6.3 | 82.06 | 81.9336 |  |  | 8.7167 | 80.6667 | 80.4889 |  |  | 14.7167 | 78.9667 | 78.86 |  | 6.05 | 21.1 | 79.2 | 79.0825
15.1 |  |  | 5.1 | 55.6545 | 56.73 |  |  | 6.2125 | 81.3 | 81.1638 |  |  | 8.62 | 80.25 | 80.2364 |  |  | 14.5 | 78.74 | 78.59 |  | 6 | 20.8 | 79.025 | 78.8358
15.2 |  |  | 5.0714 | 46.25 | 46.89 |  |  | 6.12 | 80.34 | 80.094 |  |  | 8.5 | 80.0333 | 79.95 |  |  | 14.24 | 78.54 | 78.48 |  | 5.86 | 20.5 | 78.75 | 78.6907
15.3 |  |  | 5 | 36.4 | 37.88 |  |  | 6.1 | 78.48 | 78.1026 |  |  | 8.35 | 79.5667 | 79.6 |  |  | 14.02 | 78.24 | 78.27 |  | 5.8 | 20.2333 | 78.4667 | 78.4157
15.4 |  |  | 5 | 27.94 | 29.26 |  |  | 6 | 74.76 | 74.6851 |  |  | 8.2 | 79.2333 | 79.25 |  |  | 13.82 | 78.04 | 77.97 |  | 5.76 | 19.92 | 78.28 | 78.2907
15.5 |  |  | 4.9 | 21.3 | 22.6333 |  |  | 5.94 | 69.04 | 69.1719 |  |  | 8.1143 | 78.9 | 78.9 |  |  | 13.62 | 77.8 | 77.675 |  | 5.7 | 19.6 | 78.1 | 78.0356
15.6 |  |  | 4.8545 | 17.0556 | 18.42 |  |  | 5.84 | 61.7 | 61.8076 |  |  | 8.0333 | 78.46 | 78.45 |  |  | 13.42 | 77.4 | 77.44 |  |  | 19.42 | 77.96 | 77.8556
15.7 |  |  | 4.8 | 14.86 | 16.2 |  |  | 5.8 | 53.5 | 53.3893 |  |  | 7.9 | 78.02 | 77.9 |  |  | 13.28 | 77.3 | 77.17 |  |  | 19.1667 | 77.6833 | 77.5555
15.8 |  |  | 4.8 | 13.78 | 15.06 |  |  | 5.7 | 44.75 | 44.971 |  |  | 7.8 | 77.3 | 77.2 |  |  | 13.08 | 77.1 | 76.8818 |  |  | 18.84 | 77.5 | 77.3405
15.9 |  |  | 4.75 | 13.12 | 14.48 |  |  | 5.64 | 36.24 | 36.9326 |  |  | 7.64 | 76.2833 | 76.2545 |  |  | 12.78 | 76.8 | 76.7 |  |  | 18.64 | 77.3 | 77.0954
16 |  |  | 4.7 | 12.68 | 14 |  |  | 5.6 | 29 | 29.6741 |  |  | 7.55 | 74.78 | 74.88 |  |  | 12.6833 | 76.5 | 76.4 |  |  | 18.34 | 77 | 76.8554
16.1 |  |  | 4.7 | 12.36 | 13.66 |  |  | 5.5 | 23.28 | 24.2552 |  |  | 7.44 | 72.4 | 72.55 |  |  | 12.45 | 76.3 | 75.96 |  |  | 18.1 | 76.7 | 76.5778
16.2 |  |  | 4.6 | 12.14 | 13.46 |  |  | 5.5 | 19.8 | 20.8558 |  |  | 7.36 | 68.1 | 68.5571 |  |  | 12.3714 | 75.9857 | 75.86 |  |  | 17.9 | 76.5 | 76.3724
16.3 |  |  | 4.6 | 11.9 | 13.22 |  |  | 5.4 | 18 | 19.0562 |  |  | 7.25 | 62.0333 | 63.06 |  |  | 12.15 | 75.5714 | 75.52 |  |  | 17.72 | 76.275 | 76.1952
16.4 |  |  | 4.5857 | 11.74 | 13.02 |  |  | 5.34 | 16.7333 | 17.9364 |  |  | 7.16 | 55.4 | 56.3333 |  |  | 11.95 | 75.3333 | 75.2429 |  |  | 17.46 | 76.0429 | 75.9652
16.5 |  |  | 4.5 | 11.5625 | 12.92 |  |  | 5.3 | 15.9714 | 17.2965 |  |  | 7.06 | 47.7 | 48.7 |  |  | 11.8167 | 74.9 | 74.95 |  |  | 17.12 | 75.775 | 75.7551
16.6 |  |  | 4.5 | 11.4286 | 12.7 |  |  | 5.25 | 15.54 | 16.73 |  |  | 6.96 | 40.1 | 41.82 |  |  | 11.62 | 74.6222 | 74.55 |  |  | 16.96 | 75.62 | 75.5551
16.7 |  |  | 4.44 | 11.2545 | 12.5667 |  |  | 5.2 | 15.1571 | 16.3567 |  |  | 6.9 | 33.6 | 35.06 |  |  | 11.5 | 74.2 | 74.1 |  |  | 16.8 | 75.3 | 75.2751
16.8 |  |  | 4.4 | 11.13 | 12.4 |  |  | 5.16 | 14.7545 | 15.9768 |  |  | 6.8 | 28 | 29.66 |  |  | 11.42 | 73.9091 | 73.8 |  |  | 16.6 | 75.04 | 75.055
16.9 |  |  | 4.34 | 11.0222 | 12.28 |  |  | 5.1 | 14.45 | 15.7568 |  |  | 6.7 | 24.2 | 25.74 |  |  | 11.22 | 73.6 | 73.5333 |  |  | 16.4333 | 74.82 | 74.755
17 |  |  | 4.3 | 10.9 | 12.18 |  |  | 5.1 | 14.16 | 15.4569 |  |  | 6.7 | 21.9 | 23.46 |  |  | 11.02 | 73.1 | 73.02 |  |  | 16.175 | 74.44 | 74.4749
17.1 |  |  | 4.3 | 10.7 | 12.08 |  |  | 5 | 13.86 | 15.177 |  |  | 6.6 | 20.6 | 22 |  |  | 10.86 | 72.56 | 72.5667 |  |  | 15.98 | 74.22 | 74.2749
17.2 |  |  |  | 10.6 | 11.98 |  |  | 5 | 13.625 | 14.877 |  |  | 6.5 | 19.78 | 21.16 |  |  | 10.8 | 71.96 | 72.14 |  |  | 15.78 | 74.12 | 74.0748
17.3 |  |  |  | 10.5 | 11.8667 |  |  | 4.9 | 13.42 | 14.6771 |  |  | 6.4 | 19.24 | 20.52 |  |  | 10.67 | 71.45 | 71.32 |  |  | 15.58 | 73.6667 | 73.5897
17.4 |  |  |  | 10.48 | 11.74 |  |  | 4.9 | 13.2143 | 14.4771 |  |  | 6.38 | 18.65 | 20.05 |  |  | 10.54 | 70.58 | 70.5 |  |  | 15.38 | 73.46 | 73.2747
17.5 |  |  |  | 10.3 | 11.6 |  |  | 4.8667 | 12.95 | 14.2805 |  |  | 6.3 | 18.2 | 19.58 |  |  | 10.35 | 69.36 | 69.225 |  |  | 15.2 | 73.0667 | 73.0003
17.6 |  |  |  | 10.1833 | 11.5 |  |  | 4.8 | 12.82 | 14.0772 |  |  | 6.3 | 17.8 | 19.22 |  |  | 10.2 | 67.24 | 67.22 |  |  | 15.025 | 72.8 | 72.7545
17.7 |  |  |  | 10.1 | 11.4 |  |  | 4.8 | 12.64 | 13.8805 |  |  | 6.1833 | 17.4333 | 18.88 |  |  | 10.05 | 64.14 | 64.16 |  |  | 14.8286 | 72.4667 | 72.3945
17.8 |  |  |  | 10 | 11.3222 |  |  | 4.7 | 12.44 | 13.7972 |  |  | 6.1 | 17.1 | 18.62 |  |  | 9.96 | 59.46 | 59.86 |  |  | 14.64 | 72.2 | 72.1544
17.9 |  |  |  | 9.9 | 11.1714 |  |  |  | 12.24 | 13.5973 |  |  | 6.1 | 16.92 | 18.2667 |  |  | 9.86 | 54.38 | 54.58 |  |  | 14.45 | 71.9 | 71.8544
18 |  |  |  | 9.8 | 11.1 |  |  |  | 12.04 | 13.3973 |  |  | 6 | 16.54 | 17.975 |  |  | 9.76 | 48.38 | 48.72 |  |  | 14.36 | 71.38 | 71.4943
18.1 |  |  |  | 9.7167 | 11 |  |  |  | 11.92 | 13.1974 |  |  | 5.9 | 16.32 | 17.66 |  |  | 9.66 | 42.36 | 43.06 |  |  | 14.16 | 71.1 | 71.1742
18.2 |  |  |  | 9.6 | 10.9 |  |  |  | 11.74 | 12.9974 |  |  | 5.8 | 16.0667 | 17.46 |  |  | 9.46 | 37.1 | 38.1 |  |  | 13.975 | 70.8 | 70.8742
18.3 |  |  |  | 9.54 | 10.8 |  |  |  | 11.62 | 12.8974 |  |  | 5.8 | 15.775 | 17.16 |  |  | 9.375 | 32.4 | 33.52 |  |  | 13.8429 | 70.52 | 70.5941
18.4 |  |  |  | 9.4 | 10.7 |  |  |  | 11.48 | 12.735 |  |  | 5.7 | 15.5333 | 16.86 |  |  | 9.26 | 29 | 30.22 |  |  | 13.66 | 70.12 | 70.134
18.5 |  |  |  | 9.3 | 10.64 |  |  |  | 11.3 | 12.6575 |  |  | 5.65 | 15.26 | 16.6667 |  |  | 9.18 | 26.75 | 28.08 |  |  | 13.53 | 69.68 | 69.6339
18.6 |  |  |  | 9.3 | 10.54 |  |  |  | 11.2 | 12.4975 |  |  | 5.6 | 15.08 | 16.475 |  |  | 9.08 | 25.46 | 26.8 |  |  | 13.36 | 69.2333 | 69.2739
18.7 |  |  |  | 9.2 | 10.4333 |  |  |  | 11 | 12.3575 |  |  |  | 14.88 | 16.2333 |  |  | 8.98 | 24.5333 | 25.94 |  |  | 13.18 | 68.725 | 68.8538
18.8 |  |  |  | 9.1 | 10.34 |  |  |  | 10.9 | 12.1576 |  |  |  | 14.66 | 15.98 |  |  | 8.8 | 23.8 | 25.34 |  |  | 13 | 67.96 | 68.0936
18.9 |  |  |  | 9 | 10.24 |  |  |  | 10.8 | 12.0576 |  |  |  | 14.38 | 15.78 |  |  | 8.78 | 23.375 | 24.8 |  |  | 12.79 | 67.125 | 67.3135
19 |  |  |  | 8.9 | 10.14 |  |  |  | 10.64 | 11.9576 |  |  |  | 14.18 | 15.58 |  |  | 8.7 | 22.96 | 24.48 |  |  | 12.7 | 65.78 | 66.0332
19.1 |  |  |  | 8.8 | 10.1 |  |  |  | 10.54 | 11.7976 |  |  |  | 13.98 | 15.38 |  |  | 8.5 | 22.55 | 23.96 |  |  | 12.5833 | 63.9 | 64.1529
19.2 |  |  |  | 8.7 | 9.94 |  |  |  | 10.44 | 11.6977 |  |  |  | 13.7833 | 15.1875 |  |  | 8.4 | 22.175 | 23.66 |  |  | 12.5 | 61.18 | 61.4873
19.3 |  |  |  | 8.6 | 9.84 |  |  |  | 10.34 | 11.6644 |  |  |  | 13.58 | 15.04 |  |  | 8.3 | 21.82 | 23.2714 |  |  | 12.3 | 57.5 | 57.8916
19.4 |  |  |  | 8.5 | 9.8 |  |  |  | 10.2 | 11.4577 |  |  |  | 13.4 | 14.84 |  |  | 8.3 | 21.5333 | 22.95 |  |  | 12.1167 | 52.5667 | 53.3821
19.5 |  |  |  | 8.5 | 9.7 |  |  |  | 10 | 11.2977 |  |  |  | 13.2 | 14.64 |  |  | 8.2 | 21.14 | 22.7571 |  |  | 12 | 47.66 | 48.9098
19.6 |  |  |  | 8.4 | 9.6 |  |  |  | 10 | 11.1978 |  |  |  | 13.1 | 14.4 |  |  | 8.1 | 20.9333 | 22.35 |  |  | 11.9 | 42.84 | 43.9888
19.7 |  |  |  | 8.3 | 9.5 |  |  |  | 9.9 | 11.0978 |  |  |  | 12.9 | 14.27 |  |  | 8 | 20.62 | 22.05 |  |  | 11.8 | 38.6 | 39.808
19.8 |  |  |  | 8.2 | 9.4 |  |  |  | 9.7286 | 11.0178 |  |  |  | 12.8 | 14.1182 |  |  | 7.9 | 20.28 | 21.7333 |  |  | 11.7 | 34.66 | 36.1672
19.9 |  |  |  | 8.07 | 9.3 |  |  |  | 9.65 | 10.8978 |  |  |  | 12.6 | 13.875 |  |  | 7.8333 | 20.1 | 21.5364 |  |  | 11.6 | 31.68 | 33.3867
20 |  |  |  | 8 | 9.26 |  |  |  | 9.5222 | 10.7978 |  |  |  | 12.4429 | 13.8 |  |  | 7.74 | 19.78 | 21.26 |  |  | 11.475 | 29.8857 | 31.5863
20.1 |  |  |  | 7.93 | 9.16 |  |  |  | 9.3778 | 10.6179 |  |  |  | 12.28 | 13.625 |  |  | 7.64 | 19.5 | 21 |  |  | 11.34 | 28.6 | 30.086
20.2 |  |  |  | 7.8222 | 9.05 |  |  |  | 9.2727 | 10.5312 |  |  |  | 12.1 | 13.52 |  |  | 7.6 | 19.2667 | 20.64 |  |  | 11.2 | 27.6889 | 29.3859
20.3 |  |  |  | 7.72 | 8.96 |  |  |  | 9.2 | 10.4312 |  |  |  | 11.9 | 13.32 |  |  | 7.5 | 19 | 20.425 |  |  | 11 | 27.15 | 28.7858
20.4 |  |  |  | 7.7 | 8.86 |  |  |  | 9.1 | 10.3379 |  |  |  | 11.8 | 13.1667 |  |  |  | 18.65 | 20.2 |  |  | 10.9 | 26.6889 | 28.1857
20.5 |  |  |  | 7.6 | 8.76 |  |  |  | 9 | 10.2313 |  |  |  | 11.7 | 13 |  |  |  | 18.46 | 19.9 |  |  | 10.8 | 26.2444 | 27.7722
20.6 |  |  |  | 7.46 | 8.66 |  |  |  | 8.9 | 10.1313 |  |  |  | 11.5 | 12.9 |  |  |  | 18.22 | 19.66 |  |  | 10.725 | 25.84 | 27.4655
20.7 |  |  |  | 7.4 | 8.6 |  |  |  | 8.8333 | 10.048 |  |  |  | 11.4 | 12.74 |  |  |  | 17.96 | 19.4 |  |  | 10.68 | 25.4857 | 27.0854
20.8 |  |  |  | 7.33 | 8.5 |  |  |  | 8.75 | 9.928 |  |  |  | 11.3 | 12.6 |  |  |  | 17.8 | 19.3 |  |  | 10.575 | 25.225 | 26.8054
20.9 |  |  |  | 7.275 | 8.4 |  |  |  | 8.6 | 9.798 |  |  |  | 11.1 | 12.44 |  |  |  | 17.52 | 19.02 |  |  | 10.5 | 24.84 | 26.4053
21 |  |  |  | 7.2 | 8.38 |  |  |  | 8.5 | 9.7617 |  |  |  | 11 | 12.34 |  |  |  | 17.2857 | 18.82 |  |  | 10.34 | 24.58 | 26.2052
21.1 |  |  |  | 7.1 | 8.28 |  |  |  | 8.4 | 9.5881 |  |  |  | 10.9 | 12.225 |  |  |  | 17.075 | 18.54 |  |  | 10.24 | 24.18 | 25.8052
21.2 |  |  |  | 7.05 | 8.2 |  |  |  | 8.3 | 9.4881 |  |  |  | 10.7 | 12.1 |  |  |  | 16.84 | 18.32 |  |  | 10.1 | 23.94 | 25.6051
21.3 |  |  |  | 6.9875 | 8.1 |  |  |  | 8.2 | 9.3881 |  |  |  | 10.6 | 11.8833 |  |  |  | 16.7 | 18.12 |  |  | 10 | 23.6333 | 25.205
21.4 |  |  |  | 6.9 | 8.1 |  |  |  | 8.1 | 9.2781 |  |  |  | 10.5 | 11.8 |  |  |  | 16.52 | 17.9333 |  |  | 9.9 | 23.38 | 25.105
21.5 |  |  |  | 6.8 | 8 |  |  |  | 8 | 9.1982 |  |  |  | 10.36 | 11.7 |  |  |  | 16.2667 | 17.8 |  |  | 9.8 | 23.04 | 24.7413
21.6 |  |  |  | 6.78 | 7.94 |  |  |  | 7.9167 | 9.1382 |  |  |  | 10.16 | 11.5167 |  |  |  | 16.1 | 17.5833 |  |  | 9.7 | 22.85 | 24.4049
21.7 |  |  |  | 6.7 | 7.9 |  |  |  | 7.8 | 9.0382 |  |  |  | 10.1 | 11.4 |  |  |  | 15.9 | 17.3 |  |  | 9.7 | 22.56 | 24.2249
21.8 | [22 empty] | 9.6

## Wedge 15 Longitudinal Profiles
Longitudinal Profiles
Wedge 15
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 13.5714
-34.9 |  |  |  |  | 13.7
-34.8 |  |  |  |  | 13.8375
-34.7 |  |  |  |  | 14
-34.6 |  |  |  |  | 14
-34.5 |  |  |  |  | 14.0714
-34.4 |  |  |  |  | 14.16
-34.3 |  |  |  |  | 14.36
-34.2 |  |  |  |  | 14.5
-34.1 |  |  |  |  | 14.5571
-34 |  |  |  |  | 14.6167
-33.9 |  |  |  |  | 14.7833
-33.8 |  |  |  |  | 14.85
-33.7 |  |  |  |  | 15.0571
-33.6 |  |  |  |  | 15.2
-33.5 |  |  |  |  | 15.2571
-33.4 |  |  |  |  | 15.3875
-33.3 |  |  |  |  | 15.5143
-33.2 |  |  |  |  | 15.65
-33.1 |  |  |  |  | 15.7857
-33 |  |  |  |  | 15.9143
-32.9 |  |  |  |  | 16.05
-32.8 |  |  |  |  | 16.1857
-32.7 |  |  |  |  | 16.3125
-32.6 |  |  |  |  | 16.4375
-32.5 |  |  |  |  | 16.5714
-32.4 |  |  |  |  | 16.7
-32.3 |  |  |  | 11.525 | 16.8375
-32.2 |  |  |  | 11.6 | 17.0429
-32.1 |  |  |  | 11.7 | 17.2
-32 |  |  |  | 11.8 | 17.3
-31.9 |  |  |  | 11.8875 | 17.3714
-31.8 |  |  |  | 12.08 | 17.6
-31.7 |  |  |  | 12.1143 | 17.8375
-31.6 |  |  |  | 12.2 | 17.9714
-31.5 |  |  |  | 12.2857 | 18.1
-31.4 |  |  |  | 12.4143 | 18.275
-31.3 |  |  |  | 12.6 | 18.4625
-31.2 |  |  |  | 12.7 | 18.6
-31.1 |  |  |  | 12.8 | 18.8
-31 |  |  |  | 12.9 | 18.8714
-30.9 |  |  |  | 13 | 19.1
-30.8 |  |  |  | 13.18 | 19.2286
-30.7 |  |  |  | 13.3 | 19.425
-30.6 |  |  |  | 13.4286 | 19.5875
-30.5 |  |  |  | 13.5625 | 19.825
-30.4 |  |  |  | 13.6875 | 20.0143
-30.3 |  |  |  | 13.8 | 20.1875
-30.2 |  |  |  | 13.94 | 20.425
-30.1 |  |  |  | 14.0875 | 20.6
-30 |  |  |  | 14.225 | 20.7857
-29.9 |  |  |  | 14.3 | 21.0143
-29.8 |  |  |  | 14.3875 | 21.2
-29.7 |  |  |  | 14.6143 | 21.3857
-29.6 |  |  | 9.5 | 14.75 | 21.6125
-29.5 |  |  | 9.5 | 14.8857 | 21.7857
-29.4 |  |  | 9.6091 | 15.0125 | 22.05
-29.3 |  |  | 9.7 | 15.1857 | 22.325
-29.2 |  |  | 9.7875 | 15.375 | 22.5429
-29.1 |  |  | 9.9 | 15.5143 | 22.75
-29 |  |  | 9.95 | 15.6571 | 22.925
-28.9 |  |  | 10 | 15.9 | 23.175
-28.8 |  |  | 10.1143 | 16.05 | 23.4429
-28.7 |  |  | 10.2667 | 16.1857 | 23.8
-28.6 |  |  | 10.3 | 16.4 | 23.975
-28.5 |  |  | 10.4 | 16.6 | 24.1714
-28.4 |  |  | 10.5 | 16.7 | 24.5
-28.3 |  | 8.4 | 10.56 | 16.8857 | 24.7571
-28.2 |  | 8.45 | 10.64 | 17.1125 | 25.025
-28.1 |  | 8.5 | 10.74 | 17.2857 | 25.4
-28 |  | 8.6 | 10.9 | 17.475 | 25.625
-27.9 |  | 8.6375 | 10.94 | 17.625 | 25.8143
-27.8 |  | 8.7 | 11.04 | 17.875 | 26.1625
-27.7 |  | 8.8 | 11.2 | 18.0714 | 26.45
-27.6 |  | 8.8375 | 11.34 | 18.325 | 26.7143
-27.5 |  | 8.9 | 11.4 | 18.5375 | 27.0625
-27.4 |  | 9 | 11.54 | 18.7429 | 27.475
-27.3 |  | 9.0286 | 11.7 | 18.9 | 27.8143
-27.2 | 7.8375 | 9.1625 | 11.8 | 19.1571 | 28.075
-27.1 | 7.9 | 9.2 | 11.94 | 19.425 | 28.475
-27 | 7.9231 | 9.2286 | 12.04 | 19.7 | 28.8714
-26.9 | 8 | 9.3625 | 12.24 | 19.8571 | 29.35
-26.8 | 8.025 | 9.4 | 12.34 | 20.1875 | 29.9111
-26.7 | 8.1 | 9.525 | 12.4333 | 20.4 | 30.5727
-26.6 | 8.1 | 9.6 | 12.62 | 20.6571 | 31.5
-26.5 | 8.2 | 9.6875 | 12.72 | 20.925 | 32.6714
-26.4 | 8.22 | 9.7286 | 12.92 | 21.175 | 34.4455
-26.3 | 8.3 | 9.8625 | 13.02 | 21.4286 | 36.5667
-26.2 | 8.3857 | 9.9 | 13.2 | 21.7 | 39.15
-26.1 | 8.4 | 10.025 | 13.3286 | 21.9714 | 42.66
-26 | 8.5 | 10.1 | 13.525 | 22.3286 | 46.5333
-25.9 | 8.54 | 10.1875 | 13.7 | 22.6 | 50.875
-25.8 | 8.6 | 10.3143 | 13.8286 | 22.8714 | 55.1333
-25.7 | 8.68 | 10.45 | 14.025 | 23.2286 | 59.425
-25.6 | 8.7 | 10.5857 | 14.1875 | 23.5 | 63.0692
-25.5 | 8.78 | 10.6143 | 14.3286 | 23.8571 | 66.1125
-25.4 | 8.88 | 10.75 | 14.55 | 24.2286 | 68.7375
-25.3 | 8.9 | 10.8857 | 14.7714 | 24.5778 | 70.38
-25.2 | 9 | 11.0143 | 14.9143 | 24.9429 | 71.7533
-25.1 | 9.0857 | 11.15 | 15.1 | 25.4 | 72.8867
-25 | 9.1667 | 11.2857 | 15.2857 | 25.7933 | 73.54
-24.9 | 9.2 | 11.4125 | 15.6154 | 26.26 | 74.0867
-24.8 | 9.3143 | 11.5 | 15.7692 | 26.8667 | 74.5533
-24.7 | 9.4 | 11.575 | 15.9778 | 27.6286 | 75.05
-24.6 | 9.5 | 11.7125 | 16.2143 | 28.56 | 75.5273
-24.5 | 9.56 | 11.8429 | 16.4 | 30.16 | 75.9091
-24.4 | 9.6429 | 12.05 | 16.6714 | 32.32 | 76.2385
-24.3 | 9.775 | 12.2 | 16.9222 | 34.9667 | 76.7
-24.2 | 9.8167 | 12.3429 | 17.1667 | 38.6333 | 77.1286
-24.1 | 9.9714 | 12.475 | 17.5 | 43.02 | 77.4
-24 | 10 | 12.625 | 17.74 | 47.9267 | 77.6714
-23.9 | 10.1375 | 12.8375 | 17.96 | 53.1933 | 77.9429
-23.8 | 10.2714 | 12.9714 | 18.26 | 58.42 | 78.4
-23.7 | 10.4 | 13.2 | 18.56 | 63.5 | 78.7714
-23.6 | 10.4375 | 13.3375 | 18.8667 | 67.5 | 79.0143
-23.5 | 10.5625 | 13.5429 | 19.2125 | 71.1667 | 79.25
-23.4 | 10.7 | 13.7 | 19.5429 | 73.9182 | 79.55
-23.3 | 10.7286 | 13.9375 | 19.9 | 75.8667 | 79.825
-23.2 | 10.8625 | 14.0714 | 20.3125 | 77.4 | 80.0857
-23.1 | 11 | 14.3 | 20.6875 | 78.4 | 80.425
-23 | 11.1286 | 14.575 | 21.1 | 79.0667 | 80.8
-22.9 | 11.2625 | 14.7833 | 21.6143 | 79.75 | 81.0286
-22.8 | 11.4 | 15.0333 | 22.275 | 80.2286 | 81.225
-22.7 | 11.525 | 15.3 | 23.2 | 80.625 | 81.5
-22.6 | 11.6571 | 15.55 | 24.325 | 81.0692 | 81.7571
-22.5 | 11.7875 | 15.8143 | 26.3 | 81.4111 | 82.0875
-22.4 | 11.925 | 16.1625 | 28.75 | 81.74 | 82.4
-22.3 | 12.1143 | 16.35 | 32.4143 | 82.2 | 82.68
-22.2 | 12.3143 | 16.6714 | 37.025 | 82.5 | 82.9467
-22.1 | 12.45 | 17.15 | 42.775 | 82.7571 | 83.16
-22 | 12.5857 | 17.475 | 49.3429 | 83.0875 | 83.375
-21.9 | 12.725 | 17.9667 | 56.475 | 83.5 | 83.625
-21.8 | 12.9143 | 18.58 | 63.625 | 83.7571 | 83.8867
-21.7 | 13.0875 | 19.4 | 70.2286 | 84.025 | 84.14
-21.6 | 13.325 | 20.6857 | 75.875 | 84.3 | 84.34
-21.5 | 13.5143 | 22.6875 | 80.15 | 84.6571 | 84.54
-21.4 | 13.6875 | 25.6667 | 83 | 84.925 | 84.7867
-21.3 | 13.975 | 29.9143 | 84.9286 | 85.2625 | 85.04
-21.2 | 14.3143 | 35.3286 | 86.2143 | 85.65 | 85.25
-21.1 | 14.6625 | 42.45 | 87.15 | 85.9143 | 85.5
-21 | 15.25 | 50.5143 | 87.8286 | 86.175 | 85.7375
-20.9 | 16.1571 | 58.725 | 88.25 | 86.425 | 85.95
-20.8 | 17.55 | 67.0286 | 88.7125 | 86.6714 | 86.1714
-20.7 | 19.75 | 75.3 | 89.1143 | 86.975 | 86.3
-20.6 | 23.6571 | 81.65 | 89.5 | 87.25 | 86.5286
-20.5 | 29.125 | 86.2 | 89.975 | 87.45 | 86.6833
-20.4 | 36.3429 | 89.4167 | 90.2429 | 87.7571 | 87.05
-20.3 | 45.7143 | 91.38 | 90.6 | 87.925 | 87.2571
-20.2 | 54.5667 | 92.56 | 90.975 | 88.1857 | 87.3875
-20.1 | 64.65 | 93.1667 | 91.225 | 88.3857 | 87.44
-20 | 73.8385 | 93.875 | 91.5 | 88.525 | 87.7625
-19.9 | 82.3 | 94.25 | 91.65 | 88.7857 | 88.05
-19.8 | 88.6333 | 94.7 | 91.9143 | 89.125 | 88.25
-19.7 | 92.6692 | 95 | 92.2625 | 89.3125 | 88.3857
-19.6 | 95.4 | 95.2667 | 92.525 | 89.4857 | 88.5286
-19.5 | 96.8 | 95.525 | 92.6571 | 89.675 | 88.8
-19.4 | 97.64 | 95.7857 | 92.7875 | 89.84 | 88.975
-19.3 | 98.2 | 96.05 | 93.1 | 90 | 89.2143
-19.2 | 98.425 | 96.125 | 93.3727 | 90.25 | 89.35
-19.1 | 98.6429 | 96.375 | 93.5929 | 90.4333 | 89.55
-19 | 98.775 | 96.6429 | 93.8111 | 90.725 | 89.725
-18.9 | 98.925 | 96.8 | 94.05 | 90.9 | 89.9
-18.8 | 99.1429 | 96.975 | 94.3111 | 91.0571 | 89.9714
-18.7 | 99.275 | 97.1714 | 94.4444 | 91.325 | 90.2
-18.6 | 99.4125 | 97.3 | 94.6692 | 91.5556 | 90.3375
-18.5 | 99.5333 | 97.4667 | 94.9 | 91.7 | 90.4714
-18.4 | 99.6 | 97.7 | 94.9833 | 91.83 | 90.6
-18.3 | 99.75 | 97.8429 | 95.1846 | 92.1143 | 90.76
-18.2 | 99.8 | 97.9 | 95.3625 | 92.25 | 91
-18.1 | 99.9286 | 98.1375 | 95.4875 | 92.4714 | 91.1375
-18 | 100.0625 | 98.2625 | 95.68 | 92.5818 | 91.2
-17.9 | 100.1 | 98.5 | 95.8727 | 92.7333 | 91.3
-17.8 | 100.2 | 98.6286 | 96.0429 | 92.9385 | 91.5286
-17.7 | 100.2625 | 98.7625 | 96.1857 | 93.1187 | 91.6625
-17.6 | 100.3875 | 98.9 | 96.3286 | 93.1813 | 91.7875
-17.5 | 100.5 | 99 | 96.4625 | 93.34 | 91.9
-17.4 | 100.5571 | 99.0625 | 96.675 | 93.5267 | 92.05
-17.3 | 100.6778 | 99.1 | 96.825 | 93.66 | 92.2
-17.2 | 100.7615 | 99.3286 | 96.9571 | 93.84 | 92.3
-17.1 | 100.8455 | 99.4625 | 97.0875 | 94.025 | 92.4
-17 | 100.9 | 99.52 | 97.2167 | 94.15 | 92.5333
-16.9 | 100.9286 | 99.6 | 97.3833 | 94.28 | 92.7
-16.8 | 101.0625 | 99.7 | 97.5143 | 94.4077 | 92.82
-16.7 | 101.1538 | 99.7625 | 97.65 | 94.4846 | 92.9
-16.6 | 101.2286 | 99.9 | 97.775 | 94.5889 | 93.0167
-16.5 | 101.3 | 100.0286 | 97.9 | 94.69 | 93.1714
-16.4 | 101.3909 | 100.1625 | 97.9429 | 94.79 | 93.3
-16.3 | 101.475 | 100.2 | 98.075 | 94.925 | 93.4
-16.2 | 101.5 | 100.3 | 98.3 | 95.1143 | 93.4625
-16.1 | 101.6 | 100.3571 | 98.3429 | 95.2 | 93.5875
-16 | 101.65 | 100.4875 | 98.4 | 95.3 | 93.7
-15.9 | 101.7125 | 100.6 | 98.5 | 95.35 | 93.7571
-15.8 | 101.85 | 100.65 | 98.6286 | 95.4857 | 93.8875
-15.7 | 101.9143 | 100.7 | 98.7625 | 95.6 | 94.025
-15.6 | 102 | 100.7889 | 98.8 | 95.7 | 94.1571
-15.5 | 102.0667 | 100.9143 | 98.9 | 95.7 | 94.2875
-15.4 | 102.2 | 101 | 98.9 | 95.7333 | 94.4125
-15.3 | 102.2 | 101.0857 | 99 | 95.9125 | 94.5429
-15.2 | 102.2714 | 101.1 | 99.1 | 96.0429 | 94.675
-15.1 | 102.3 | 101.1429 | 99.1 | 96.1857 | 94.8125
-15 | 102.3286 | 101.2 | 99.2 | 96.2 | 94.925
-14.9 | 102.4 | 101.2 | 99.2625 | 96.3 | 95.0111
-14.8 | 102.4 | 101.2375 | 99.3 | 96.4 | 95.1154
-14.7 | 102.4 | 101.3 | 99.4 | 96.5 | 95.1923
-14.6 | 102.4625 | 101.3 | 99.4 | 96.5625 | 95.32
-14.5 | 102.5 | 101.3375 | 99.4857 | 96.7 | 95.4267
-14.4 | 102.6 | 101.4 | 99.5 | 96.725 | 95.4933
-14.3 | 102.6 | 101.4 | 99.5429 | 96.8571 | 95.5562
-14.2 | 102.7 | 101.4286 | 99.6 | 96.9 | 95.64
-14.1 | 102.6714 | 101.5 | 99.7 | 97.025 | 95.7733
-14 | 102.6 | 101.5 | 99.7 | 97.1571 | 95.8
-13.9 | 102.6 | 101.5 | 99.7 | 97.2 | 95.9
-13.8 | 102.6 | 101.5 | 99.8 | 97.2143 | 95.9429
-13.7 | 102.6 | 101.5 | 99.8 | 97.3 | 96.075
-13.6 | 102.6 | 101.525 | 99.8714 | 97.4 | 96.2125
-13.5 | 102.6 | 101.6 | 99.9 | 97.4571 | 96.3375
-13.4 | 102.6 | 101.6 | 100 | 97.5778 | 96.4714
-13.3 | 102.6 | 101.6 | 100.0714 | 97.7125 | 96.6
-13.2 | 102.6 | 101.6 | 100.1 | 97.8 | 96.6375
-13.1 | 102.6 | 101.6 | 100.1 | 97.8308 | 96.7625
-13 | 102.6 | 101.6 | 100.1625 | 97.91 | 96.8
-12.9 | 102.5625 | 101.6 | 100.2 | 98.0083 | 96.9
-12.8 | 102.5 | 101.66 | 100.225 | 98.0917 | 96.9
-12.7 | 102.5 | 101.7 | 100.3 | 98.1 | 97
-12.6 | 102.5 | 101.7 | 100.3 | 98.1583 | 97
-12.5 | 102.6 | 101.7 | 100.3 | 98.2625 | 97.025
-12.4 | 102.6 | 101.7 | 100.35 | 98.37 | 97.1
-12.3 | 102.6 | 101.7 | 100.4 | 98.4 | 97.2
-12.2 | 102.6 | 101.7 | 100.5 | 98.5 | 97.3
-12.1 | 102.6 | 101.8 | 100.5 | 98.5625 | 97.3
-12 | 102.6 | 101.8 | 100.5 | 98.6 | 97.3
-11.9 | 102.6875 | 101.84 | 100.6 | 98.6286 | 97.4
-11.8 | 102.6857 | 101.9 | 100.6 | 98.7556 | 97.48
-11.7 | 102.6 | 101.9 | 100.675 | 98.8 | 97.5667
-11.6 | 102.6857 | 101.9 | 100.7 | 98.9 | 97.7
-11.5 | 102.6 | 102 | 100.7429 | 98.9 | 97.8
-11.4 | 102.6 | 102 | 100.8 | 98.9857 | 97.9
-11.3 | 102.6 | 102 | 100.8 | 99 | 97.9
-11.2 | 102.6 | 102 | 100.8 | 99.05 | 97.96
-11.1 | 102.69 | 102 | 100.8 | 99.1 | 98.1
-11 | 102.7 | 102 | 100.8 | 99.1 | 98.2
-10.9 | 102.7 | 102 | 100.8 | 99.1 | 98.2
-10.8 | 102.7 | 102 | 100.8714 | 99.175 | 98.2875
-10.7 | 102.7 | 102 | 100.9 | 99.1125 | 98.4
-10.6 | 102.7 | 102 | 101 | 99.2 | 98.4
-10.5 | 102.7 | 102 | 101 | 99.275 | 98.4
-10.4 | 102.7 | 102 | 101 | 99.3 | 98.5
-10.3 | 102.7 | 102 | 101 | 99.3375 | 98.55
-10.2 | 102.7 | 102 | 101 | 99.4 | 98.6
-10.1 | 102.7 | 102 | 101 | 99.5 | 98.6143
-10 | 102.6571 | 102 | 100.9571 | 99.5 | 98.7
-9.9 | 102.6 | 101.975 | 100.9 | 99.5 | 98.7
-9.8 | 102.6 | 101.9 | 100.9 | 99.5375 | 98.8
-9.7 | 102.6 | 101.9875 | 100.9667 | 99.6625 | 98.8
-9.6 | 102.6 | 102 | 101 | 99.7 | 98.9
-9.5 | 102.6 | 102 | 101 | 99.8 | 98.9429
-9.4 | 102.5 | 102 | 101.0429 | 99.8 | 99
-9.3 | 102.5 | 102 | 101.1 | 99.8 | 99.1
-9.2 | 102.5 | 102 | 101.1 | 99.8 | 99.1
-9.1 | 102.5 | 102 | 101.1 | 99.8625 | 99.2
-9 | 102.5 | 102 | 101.2 | 99.9875 | 99.2
-8.9 | 102.5 | 102 | 101.2 | 100.1 | 99.24
-8.8 | 102.5 | 102 | 101.2 | 100.1 | 99.3
-8.7 | 102.5 | 102 | 101.2 | 100.1857 | 99.3
-8.6 | 102.5 | 102 | 101.2 | 100.2 | 99.4375
-8.5 | 102.5 | 102 | 101.2 | 100.2 | 99.5
-8.4 | 102.4 | 102.0125 | 101.2 | 100.2 | 99.6
-8.3 | 102.5 | 102.1 | 101.2 | 100.2533 | 99.6
-8.2 | 102.4 | 102.1 | 101.2 | 100.3 | 99.6
-8.1 | 102.4 | 102.1 | 101.2 | 100.3 | 99.6
-8 | 102.4 | 102.1 | 101.2 | 100.3 | 99.6286
-7.9 | 102.4 | 102.1 | 101.2875 | 100.3 | 99.7
-7.8 | 102.4 | 102.1 | 101.3 | 100.3 | 99.7
-7.7 | 102.4 | 102.1 | 101.3 | 100.4 | 99.7
-7.6 | 102.4 | 102.1 | 101.3 | 100.37 | 99.7
-7.5 | 102.4286 | 102.1 | 101.3 | 100.3 | 99.7875
-7.4 | 102.4 | 102.1 | 101.3 | 100.3625 | 99.825
-7.3 | 102.4 | 102.1 | 101.3 | 100.4 | 99.9
-7.2 | 102.4 | 102.1 | 101.3143 | 100.4 | 99.9
-7.1 | 102.4 | 102.1 | 101.35 | 100.4 | 100
-7 | 102.4 | 102 | 101.3 | 100.4875 | 100
-6.9 | 102.4 | 102 | 101.3 | 100.5 | 100
-6.8 | 102.3857 | 102 | 101.3 | 100.5 | 100.0875
-6.7 | 102.3 | 101.9 | 101.3 | 100.5 | 100.0429
-6.6 | 102.3 | 101.9857 | 101.3 | 100.6 | 100.175
-6.5 | 102.3 | 101.95 | 101.3 | 100.6 | 100.2
-6.4 | 102.3 | 101.9857 | 101.3 | 100.6 | 100.2429
-6.3 | 102.3 | 101.9 | 101.4 | 100.7 | 100.3
-6.2 | 102.3 | 101.9 | 101.4 | 100.7 | 100.3
-6.1 | 102.3 | 101.9 | 101.4 | 100.625 | 100.3
-6 | 102.2 | 101.9 | 101.3 | 100.6 | 100.3
-5.9 | 102.2 | 101.9 | 101.3 | 100.7 | 100.3
-5.8 | 102.2 | 101.9 | 101.3 | 100.7 | 100.3
-5.7 | 102.2 | 101.9 | 101.3 | 100.7 | 100.3
-5.6 | 102.1 | 101.8375 | 101.3 | 100.7 | 100.3
-5.5 | 102.1 | 101.8 | 101.2375 | 100.7 | 100.3
-5.4 | 102.1 | 101.8 | 101.3 | 100.7 | 100.3
-5.3 | 102 | 101.8 | 101.3 | 100.7 | 100.3
-5.2 | 102 | 101.8 | 101.3 | 100.7 | 100.3
-5.1 | 102 | 101.8 | 101.3 | 100.7143 | 100.3
-5 | 102 | 101.7 | 101.3 | 100.8 | 100.3
-4.9 | 102 | 101.7 | 101.2375 | 100.8 | 100.3
-4.8 | 101.9 | 101.6467 | 101.2 | 100.8 | 100.3
-4.7 | 101.9 | 101.6 | 101.2 | 100.8 | 100.3125
-4.6 | 101.8 | 101.6 | 101.2 | 100.8 | 100.4
-4.5 | 101.8 | 101.6 | 101.2 | 100.7125 | 100.4
-4.4 | 101.8 | 101.5867 | 101.1 | 100.7571 | 100.4
-4.3 | 101.75 | 101.52 | 101.1 | 100.7 | 100.4
-4.2 | 101.7 | 101.4533 | 101.1 | 100.7 | 100.325
-4.1 | 101.7 | 101.4 | 101.0857 | 100.7 | 100.3
-4 | 101.6857 | 101.4 | 101 | 100.7 | 100.3
-3.9 | 101.6 | 101.35 | 101 | 100.6 | 100.3
-3.8 | 101.5 | 101.3 | 101 | 100.6833 | 100.3
-3.7 | 101.4 | 101.2833 | 101 | 100.6 | 100.3
-3.6 | 101.4 | 101.2 | 101 | 100.6 | 100.3
-3.5 | 101.3 | 101.2 | 101 | 100.6 | 100.3
-3.4 | 101.3 | 101.2 | 101 | 100.6 | 100.3
-3.3 | 101.3 | 101.175 | 101 | 100.6 | 100.3
-3.2 | 101.2 | 101.1 | 100.9625 | 100.6 | 100.3
-3.1 | 101.2 | 101.1 | 100.9 | 100.58 | 100.3
-3 | 101.2 | 101.1 | 100.9 | 100.5 | 100.3
-2.9 | 101.1 | 101.0429 | 100.8167 | 100.5 | 100.3
-2.8 | 101.0625 | 101 | 100.7875 | 100.48 | 100.3
-2.7 | 101 | 101 | 100.7 | 100.4 | 100.3
-2.6 | 100.9 | 100.9429 | 100.6583 | 100.4 | 100.3375
-2.5 | 100.9 | 100.9 | 100.6 | 100.4 | 100.4
-2.4 | 100.8286 | 100.8 | 100.625 | 100.32 | 100.4
-2.3 | 100.8 | 100.8 | 100.6933 | 100.325 | 100.4
-2.2 | 100.8 | 100.7125 | 100.6267 | 100.4 | 100.3
-2.1 | 100.7375 | 100.7 | 100.5571 | 100.4 | 100.3
-2 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
-1.9 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
-1.8 | 100.6429 | 100.6 | 100.5 | 100.3 | 100.2571
-1.7 | 100.6 | 100.6 | 100.4615 | 100.3 | 100.3
-1.6 | 100.5625 | 100.6 | 100.4 | 100.3 | 100.26
-1.5 | 100.5 | 100.5 | 100.3571 | 100.3 | 100.2
-1.4 | 100.5 | 100.5 | 100.3 | 100.2857 | 100.1
-1.3 | 100.4 | 100.5 | 100.3 | 100.2 | 100.1
-1.2 | 100.4 | 100.425 | 100.3 | 100.2 | 100.1
-1.1 | 100.3125 | 100.3875 | 100.3 | 100.2 | 100.1
-1 | 100.275 | 100.3 | 100.2 | 100.2 | 100.1
-0.9 | 100.2 | 100.2286 | 100.2 | 100.2 | 100.1
-0.8 | 100.2 | 100.2 | 100.16 | 100.2 | 100.1
-0.7 | 100.1333 | 100.1714 | 100.1 | 100.125 | 100
-0.6 | 100.1 | 100.1 | 100.1 | 100.0875 | 100
-0.5 | 100.0583 | 100.1 | 100.025 | 100 | 100
-0.4 | 100 | 100.1 | 100 | 100 | 100
-0.3 | 100 | 100.1 | 100 | 100 | 100
-0.2 | 100 | 100.1 | 100 | 100 | 100
-0.1 | 100 | 100.075 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100.075 | 100 | 100 | 100
0.2 | 100 | 100.1 | 100 | 100 | 100
0.3 | 100 | 100.1 | 100 | 100 | 100
0.4 | 100 | 100.1 | 100 | 100 | 100
0.5 | 100.0286 | 100.1 | 100.0833 | 100 | 100
0.6 | 100.1 | 100.1 | 100.1 | 100.0571 | 100
0.7 | 100.1 | 100.1714 | 100.1 | 100.1 | 100
0.8 | 100.1909 | 100.2 | 100.1333 | 100.2 | 100.1
0.9 | 100.2125 | 100.2286 | 100.2 | 100.2 | 100.1
1 | 100.3 | 100.3 | 100.2 | 100.2 | 100.1
1.1 | 100.3 | 100.3875 | 100.2833 | 100.2 | 100.1
1.2 | 100.4 | 100.425 | 100.3 | 100.2 | 100.1
1.3 | 100.4 | 100.5 | 100.3 | 100.2 | 100.1
1.4 | 100.4714 | 100.5 | 100.3 | 100.25 | 100.1
1.5 | 100.5 | 100.5 | 100.4 | 100.3 | 100.2
1.6 | 100.6 | 100.6 | 100.4 | 100.3 | 100.26
1.7 | 100.6 | 100.6 | 100.4333 | 100.35 | 100.3
1.8 | 100.7 | 100.6 | 100.5 | 100.325 | 100.2571
1.9 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
2 | 100.7 | 100.7 | 100.5455 | 100.3 | 100.3
2.1 | 100.7 | 100.7 | 100.6 | 100.375 | 100.3
2.2 | 100.8 | 100.7125 | 100.6 | 100.4 | 100.3
2.3 | 100.8 | 100.8 | 100.66 | 100.3625 | 100.4
2.4 | 100.9 | 100.8 | 100.66 | 100.3833 | 100.4
2.5 | 100.9 | 100.9 | 100.6364 | 100.4 | 100.4
2.6 | 100.9625 | 100.9429 | 100.7 | 100.4 | 100.3375
2.7 | 101 | 101 | 100.7 | 100.4 | 100.3
2.8 | 101.0286 | 101 | 100.7556 | 100.5 | 100.3
2.9 | 101.1 | 101.0429 | 100.8 | 100.5 | 100.3
3 | 101.1833 | 101.1 | 100.9 | 100.5 | 100.3
3.1 | 101.2 | 101.1 | 100.9 | 100.54 | 100.3
3.2 | 101.26 | 101.1 | 101 | 100.6 | 100.3
3.3 | 101.3 | 101.175 | 101 | 100.6 | 100.3
3.4 | 101.3 | 101.2 | 101 | 100.6 | 100.3
3.5 | 101.3 | 101.2 | 101 | 100.6 | 100.3
3.6 | 101.4 | 101.2 | 101 | 100.6 | 100.3
3.7 | 101.4 | 101.2833 | 101 | 100.6 | 100.3
3.8 | 101.48 | 101.3 | 101 | 100.65 | 100.3
3.9 | 101.58 | 101.35 | 101 | 100.6 | 100.3
4 | 101.6571 | 101.4 | 101.0143 | 100.6857 | 100.3
4.1 | 101.7 | 101.4 | 101.1 | 100.7 | 100.3
4.2 | 101.7 | 101.4533 | 101.1 | 100.7 | 100.325
4.3 | 101.8 | 101.52 | 101.1 | 100.7 | 100.4
4.4 | 101.8 | 101.5867 | 101.1 | 100.8 | 100.4
4.5 | 101.8 | 101.6 | 101.1857 | 100.7429 | 100.4
4.6 | 101.8 | 101.6 | 101.2 | 100.7875 | 100.4
4.7 | 101.8833 | 101.6 | 101.2 | 100.8 | 100.3125
4.8 | 101.9 | 101.6467 | 101.2 | 100.8 | 100.3
4.9 | 101.96 | 101.7 | 101.2125 | 100.8 | 100.3
5 | 102 | 101.7 | 101.3 | 100.8 | 100.3
5.1 | 102 | 101.8 | 101.3 | 100.7571 | 100.3
5.2 | 102 | 101.8 | 101.3 | 100.7 | 100.3
5.3 | 102.0375 | 101.8 | 101.3 | 100.7 | 100.3
5.4 | 102.1 | 101.8 | 101.2286 | 100.7 | 100.3
5.5 | 102.1 | 101.8 | 101.3 | 100.7 | 100.3
5.6 | 102.1286 | 101.8375 | 101.3 | 100.7 | 100.3
5.7 | 102.2 | 101.9 | 101.3 | 100.7 | 100.3
5.8 | 102.2 | 101.9 | 101.3 | 100.7 | 100.3
5.9 | 102.2 | 101.9 | 101.3 | 100.6714 | 100.3
6 | 102.2625 | 101.9 | 101.3 | 100.6 | 100.3
6.1 | 102.3 | 101.9 | 101.4 | 100.7 | 100.3
6.2 | 102.3 | 101.9 | 101.4 | 100.7 | 100.3
6.3 | 102.3 | 101.9 | 101.4 | 100.6375 | 100.3
6.4 | 102.3 | 101.9857 | 101.3125 | 100.6 | 100.2429
6.5 | 102.3 | 101.95 | 101.3 | 100.6 | 100.2
6.6 | 102.3 | 101.9857 | 101.3 | 100.6 | 100.175
6.7 | 102.3 | 101.9 | 101.3 | 100.5143 | 100.0429
6.8 | 102.35 | 102 | 101.3 | 100.5 | 100.0875
6.9 | 102.4 | 102 | 101.3 | 100.5 | 100
7 | 102.4 | 102 | 101.3 | 100.425 | 100
7.1 | 102.4 | 102.1 | 101.3857 | 100.4 | 100
7.2 | 102.4 | 102.1 | 101.3 | 100.4 | 99.9
7.3 | 102.4 | 102.1 | 101.3 | 100.4 | 99.9
7.4 | 102.4 | 102.1 | 101.3143 | 100.4 | 99.825
7.5 | 102.4 | 102.1 | 101.35 | 100.33 | 99.7875
7.6 | 102.4 | 102.1 | 101.3 | 100.4 | 99.7
7.7 | 102.4 | 102.1 | 101.3 | 100.3375 | 99.7
7.8 | 102.4 | 102.1 | 101.3 | 100.3 | 99.7
7.9 | 102.4 | 102.1 | 101.3 | 100.3 | 99.7
8 | 102.4 | 102.1 | 101.2 | 100.3 | 99.6286
8.1 | 102.4 | 102.1 | 101.2 | 100.3 | 99.6
8.2 | 102.4 | 102.1 | 101.2 | 100.2533 | 99.6
8.3 | 102.48 | 102.1 | 101.2 | 100.2 | 99.6
8.4 | 102.4 | 102.0125 | 101.2 | 100.2 | 99.6
8.5 | 102.5 | 102 | 101.2 | 100.2 | 99.5
8.6 | 102.5 | 102 | 101.2 | 100.2 | 99.4375
8.7 | 102.5 | 102 | 101.2625 | 100.2 | 99.3
8.8 | 102.5 | 102 | 101.2 | 100.1 | 99.3
8.9 | 102.5 | 102 | 101.2 | 100.0571 | 99.24
9 | 102.5 | 102 | 101.2 | 100 | 99.2
9.1 | 102.5 | 102 | 101.1167 | 99.8875 | 99.2
9.2 | 102.5 | 102 | 101.1 | 99.8 | 99.1
9.3 | 102.5 | 102 | 101.1 | 99.8 | 99.1
9.4 | 102.5 | 102 | 101.075 | 99.8 | 99
9.5 | 102.5714 | 102 | 101 | 99.7625 | 98.9429
9.6 | 102.6 | 102 | 101 | 99.7 | 98.9
9.7 | 102.6 | 101.9875 | 100.9 | 99.6 | 98.8
9.8 | 102.6 | 101.9 | 100.9 | 99.5625 | 98.8
9.9 | 102.6 | 101.975 | 100.9 | 99.5 | 98.7
10 | 102.625 | 102 | 100.9125 | 99.5 | 98.7
10.1 | 102.7 | 102 | 101 | 99.5 | 98.6143
10.2 | 102.7 | 102 | 101 | 99.4 | 98.6
10.3 | 102.7 | 102 | 101 | 99.3714 | 98.55
10.4 | 102.7 | 102 | 101 | 99.3 | 98.5
10.5 | 102.7 | 102 | 101 | 99.3 | 98.4
10.6 | 102.7 | 102 | 100.9429 | 99.175 | 98.4
10.7 | 102.7 | 102 | 100.9 | 99.1 | 98.4
10.8 | 102.7 | 102 | 100.9 | 99.1875 | 98.2875
10.9 | 102.7 | 102 | 100.8 | 99.1 | 98.2
11 | 102.7 | 102 | 100.8 | 99.1 | 98.2
11.1 | 102.62 | 102 | 100.8 | 99.1 | 98.1
11.2 | 102.6 | 102 | 100.8 | 99 | 97.96
11.3 | 102.6 | 102 | 100.8 | 99 | 97.9
11.4 | 102.6 | 102 | 100.8 | 99 | 97.9
11.5 | 102.6 | 102 | 100.775 | 98.9 | 97.8
11.6 | 102.6 | 101.9 | 100.7 | 98.9 | 97.7
11.7 | 102.6143 | 101.9 | 100.7 | 98.8143 | 97.5667
11.8 | 102.7 | 101.9 | 100.6 | 98.7875 | 97.48
11.9 | 102.7 | 101.84 | 100.6 | 98.6571 | 97.4
12 | 102.6 | 101.8 | 100.5125 | 98.6 | 97.3
12.1 | 102.6 | 101.8 | 100.5 | 98.6 | 97.3
12.2 | 102.6 | 101.7 | 100.5 | 98.5 | 97.3
12.3 | 102.6 | 101.7 | 100.4125 | 98.4286 | 97.2
12.4 | 102.6 | 101.7 | 100.3857 | 98.3 | 97.1
12.5 | 102.5375 | 101.7 | 100.3 | 98.2 | 97.025
12.6 | 102.5 | 101.7 | 100.3 | 98.1917 | 97
12.7 | 102.5 | 101.7 | 100.2875 | 98.1083 | 97
12.8 | 102.5 | 101.66 | 100.2 | 98.0308 | 96.9
12.9 | 102.5286 | 101.6 | 100.2 | 98 | 96.9
13 | 102.6 | 101.6 | 100.2 | 97.9455 | 96.8
13.1 | 102.6 | 101.6 | 100.1 | 97.8615 | 96.7625
13.2 | 102.6 | 101.6 | 100.1 | 97.775 | 96.6375
13.3 | 102.6 | 101.6 | 100 | 97.7 | 96.6
13.4 | 102.6 | 101.6 | 100 | 97.6125 | 96.4714
13.5 | 102.6 | 101.6 | 99.9375 | 97.4857 | 96.3375
13.6 | 102.56 | 101.525 | 99.9 | 97.4 | 96.2125
13.7 | 102.6 | 101.5 | 99.8 | 97.3 | 96.075
13.8 | 102.6 | 101.5 | 99.7375 | 97.2571 | 95.9429
13.9 | 102.6 | 101.5 | 99.7 | 97.2 | 95.9
14 | 102.6 | 101.5 | 99.7 | 97.1 | 95.8
14.1 | 102.6375 | 101.5 | 99.7 | 97.0571 | 95.7733
14.2 | 102.6286 | 101.4286 | 99.6125 | 96.925 | 95.64
14.3 | 102.6 | 101.4 | 99.5 | 96.8 | 95.5562
14.4 | 102.6 | 101.4 | 99.5 | 96.7625 | 95.4933
14.5 | 102.5286 | 101.3375 | 99.5 | 96.6286 | 95.4267
14.6 | 102.5 | 101.3 | 99.4 | 96.6 | 95.32
14.7 | 102.4 | 101.3 | 99.4 | 96.44 | 95.1923
14.8 | 102.4 | 101.2375 | 99.3143 | 96.4 | 95.1154
14.9 | 102.4 | 101.2 | 99.3 | 96.3 | 95.0111
15 | 102.3625 | 101.2 | 99.1625 | 96.2 | 94.925
15.1 | 102.3 | 101.1429 | 99.1 | 96.1125 | 94.8125
15.2 | 102.3 | 101.1 | 99.1 | 96.05 | 94.675
15.3 | 102.2 | 101.0857 | 99 | 95.8429 | 94.5429
15.4 | 102.1375 | 101 | 98.9 | 95.7667 | 94.4125
15.5 | 102.1 | 100.9143 | 98.9 | 95.7 | 94.2875
15.6 | 102 | 100.7889 | 98.8286 | 95.7 | 94.1571
15.7 | 101.9429 | 100.7 | 98.7 | 95.5571 | 94.025
15.8 | 101.8714 | 100.65 | 98.6625 | 95.4143 | 93.8875
15.9 | 101.7 | 100.6 | 98.5375 | 95.3857 | 93.7571
16 | 101.6857 | 100.4875 | 98.4 | 95.3 | 93.7
16.1 | 101.55 | 100.3571 | 98.375 | 95.225 | 93.5875
16.2 | 101.5 | 100.3 | 98.2429 | 95.0875 | 93.4625
16.3 | 101.4 | 100.2 | 98.1125 | 94.9571 | 93.4
16.4 | 101.3231 | 100.1625 | 97.975 | 94.75 | 93.3
16.5 | 101.3 | 100.0286 | 97.8429 | 94.7 | 93.1714
16.6 | 101.22 | 99.9 | 97.8 | 94.6222 | 93.0167
16.7 | 101.0909 | 99.7625 | 97.675 | 94.5111 | 92.9
16.8 | 101 | 99.7 | 97.55 | 94.38 | 92.82
16.9 | 100.9625 | 99.6 | 97.4143 | 94.25 | 92.7
17 | 100.8643 | 99.52 | 97.25 | 94.125 | 92.5333
17.1 | 100.8 | 99.4625 | 97.0286 | 93.9933 | 92.4
17.2 | 100.7923 | 99.3286 | 96.8875 | 93.84 | 92.3
17.3 | 100.7154 | 99.1 | 96.8 | 93.64 | 92.2
17.4 | 100.6 | 99.0625 | 96.65 | 93.4933 | 92.05
17.5 | 100.5 | 99 | 96.49 | 93.36 | 91.9
17.6 | 100.425 | 98.9 | 96.38 | 93.2267 | 91.7875
17.7 | 100.3 | 98.7625 | 96.18 | 93.1 | 91.6625
17.8 | 100.2 | 98.6286 | 96 | 92.9455 | 91.5286
17.9 | 100.1286 | 98.5 | 95.8385 | 92.6929 | 91.3
18 | 100 | 98.2625 | 95.6867 | 92.6214 | 91.2
18.1 | 99.9625 | 98.1375 | 95.5533 | 92.425 | 91.1375
18.2 | 99.8286 | 97.9 | 95.3875 | 92.2857 | 91
18.3 | 99.7111 | 97.8429 | 95.2 | 92.1111 | 90.76
18.4 | 99.58 | 97.7 | 95.0182 | 91.8889 | 90.6
18.5 | 99.5 | 97.4667 | 94.8714 | 91.725 | 90.4714
18.6 | 99.4429 | 97.3 | 94.6636 | 91.5222 | 90.3375
18.7 | 99.3125 | 97.1714 | 94.4556 | 91.3 | 90.2
18.8 | 99.15 | 96.975 | 94.275 | 91.0625 | 89.9714
18.9 | 98.9429 | 96.8 | 94.025 | 90.9286 | 89.9
19 | 98.8125 | 96.6429 | 93.7818 | 90.7222 | 89.725
19.1 | 98.575 | 96.375 | 93.6 | 90.4667 | 89.55
19.2 | 98.3857 | 96.125 | 93.35 | 90.2667 | 89.35
19.3 | 98.16 | 96.05 | 93.1 | 90.0375 | 89.2143
19.4 | 97.7111 | 95.7857 | 92.775 | 89.86 | 88.975
19.5 | 96.8889 | 95.525 | 92.6875 | 89.625 | 88.8
19.6 | 95.4167 | 95.2667 | 92.5571 | 89.475 | 88.5286
19.7 | 93 | 95 | 92.275 | 89.3429 | 88.3857
19.8 | 88.9167 | 94.7 | 91.875 | 89.125 | 88.25
19.9 | 82.1308 | 94.25 | 91.6625 | 88.825 | 88.05
20 | 73.9 | 93.875 | 91.475 | 88.55 | 87.7625
20.1 | 64.5 | 93.1667 | 91.2 | 88.4143 | 87.44
20.2 | 54.4125 | 92.56 | 90.9429 | 88.1714 | 87.3875
20.3 | 45.7 | 91.38 | 90.6125 | 87.95 | 87.2571
20.4 | 36.9571 | 89.4167 | 90.3 | 87.7286 | 87.05
20.5 | 29.3286 | 86.2 | 89.9857 | 87.4857 | 86.6833
20.6 | 23.9286 | 81.65 | 89.5125 | 87.25 | 86.5286
20.7 | 19.95 | 75.3 | 89.1375 | 86.9286 | 86.3
20.8 | 17.6 | 67.0286 | 88.725 | 86.675 | 86.1714
20.9 | 16.1 | 58.725 | 88.3 | 86.4143 | 85.95
21 | 15.2286 | 50.5143 | 87.7571 | 86.15 | 85.7375
21.1 | 14.6571 | 42.45 | 87.1286 | 85.9 | 85.5
21.2 | 14.275 | 35.3286 | 86.25 | 85.6875 | 85.25
21.3 | 14.0143 | 29.9143 | 84.9429 | 85.2857 | 85.04
21.4 | 13.75 | 25.6667 | 83.125 | 84.9 | 84.7867
21.5 | 13.4875 | 22.6875 | 80.0857 | 84.625 | 84.54
21.6 | 13.3143 | 20.6857 | 75.9 | 84.3571 | 84.34
21.7 | 13.125 | 19.4 | 70.175 | 84.1 | 84.14
21.8 | 12.8875 | 18.58 | 63.6571 | 83.725 | 83.8867
21.9 | 12.7571 | 17.9667 | 56.525 | 83.4571 | 83.625
22 | 12.6143 | 17.475 | 49.25 | 83.1 | 83.375
22.1 | 12.4714 | 17.15 | 42.7714 | 82.7875 | 83.16
22.2 | 12.25 | 16.6714 | 37.125 | 82.475 | 82.9467
22.3 | 12.0875 | 16.35 | 32.325 | 82.225 | 82.68
22.4 | 11.9571 | 16.1625 | 28.875 | 81.7667 | 82.4
22.5 | 11.825 | 15.8143 | 26.2143 | 81.4333 | 82.0875
22.6 | 11.6875 | 15.55 | 24.4 | 81.1 | 81.7571
22.7 | 11.5571 | 15.3 | 23.1333 | 80.6455 | 81.5
22.8 | 11.425 | 15.0333 | 22.3 | 80.2 | 81.225
22.9 | 11.3 | 14.7833 | 21.6857 | 79.6714 | 81.0286
23 | 11.1625 | 14.575 | 21.15 | 79.0625 | 80.8
23.1 | 11.0286 | 14.3 | 20.7 | 78.325 | 80.425
23.2 | 10.9 | 14.0714 | 20.3143 | 77.2833 | 80.0857
23.3 | 10.7625 | 13.9375 | 19.9125 | 75.8667 | 79.825
23.4 | 10.6375 | 13.7 | 19.5 | 74.2 | 79.55
23.5 | 10.6 | 13.5429 | 19.2143 | 71.2385 | 79.25
23.6 | 10.4625 | 13.3375 | 18.85 | 67.9308 | 79.0143
23.7 | 10.4 | 13.2 | 18.55 | 63.4 | 78.7714
23.8 | 10.3 | 12.9714 | 18.25 | 58.4933 | 78.4
23.9 | 10.1714 | 12.8375 | 17.95 | 53.2267 | 77.9429
24 | 10.0375 | 12.625 | 17.7 | 47.96 | 77.6714
24.1 | 10 | 12.475 | 17.4889 | 43.0133 | 77.4
24.2 | 9.8429 | 12.3429 | 17.2 | 38.5467 | 77.1286
24.3 | 9.8 | 12.2 | 16.8857 | 35.16 | 76.7
24.4 | 9.6714 | 12.05 | 16.7125 | 32.32 | 76.2385
24.5 | 9.6 | 11.8429 | 16.3857 | 30.12 | 75.9091
24.6 | 9.5 | 11.7125 | 16.2 | 28.7 | 75.5273
24.7 | 9.3857 | 11.575 | 16.0143 | 27.5667 | 75.05
24.8 | 9.3 | 11.5 | 15.775 | 26.7857 | 74.5533
24.9 | 9.2 | 11.4125 | 15.5727 | 26.2143 | 74.0867
25 | 9.2 | 11.2857 | 15.3286 | 25.7308 | 73.54
25.1 | 9.1 | 11.15 | 15.0857 | 25.375 | 72.8867
25.2 | 9 | 11.0143 | 14.95 | 25.0375 | 71.7533
25.3 | 8.9 | 10.8857 | 14.7286 | 24.525 | 70.38
25.4 | 8.8 | 10.75 | 14.5714 | 24.2 | 68.7375
25.5 | 8.8 | 10.6143 | 14.35 | 23.9286 | 66.1125
25.6 | 8.7 | 10.5857 | 14.2143 | 23.4857 | 63.0692
25.7 | 8.7 | 10.45 | 14 | 23.25 | 59.425
25.8 | 8.6 | 10.3143 | 13.825 | 22.8429 | 55.1333
25.9 | 8.5 | 10.1875 | 13.6286 | 22.5714 | 50.875
26 | 8.5 | 10.1 | 13.5 | 22.3 | 46.5333
26.1 | 8.4 | 10.025 | 13.3625 | 22.0286 | 42.66
26.2 | 8.4 | 9.9 | 13.2375 | 21.7571 | 39.15
26.3 | 8.3 | 9.8625 | 13.06 | 21.4 | 36.5667
26.4 | 8.26 | 9.7286 | 12.96 | 21.225 | 34.4455
26.5 | 8.2 | 9.6875 | 12.76 | 20.9 | 32.6714
26.6 | 8.1143 | 9.6 | 12.6 | 20.6625 | 31.5
26.7 | 8.1 | 9.525 | 12.46 | 20.4571 | 30.5727
26.8 | 8.0556 | 9.4 | 12.36 | 20.2 | 29.9111
26.9 | 8 | 9.3625 | 12.16 | 19.8625 | 29.35
27 | 7.9533 | 9.2286 | 12.06 | 19.6571 | 28.8714
27.1 | 7.9 | 9.2 | 11.9 | 19.4 | 28.475
27.2 |  | 9.1625 | 11.8 | 19.125 | 28.075
27.3 |  | 9.0286 | 11.68 | 18.875 | 27.8143
27.4 |  | 9 | 11.58 | 18.7 | 27.475
27.5 |  | 8.9 | 11.4 | 18.5429 | 27.0625
27.6 |  | 8.8375 | 11.3 | 18.275 | 26.7143
27.7 |  | 8.8 | 11.2 | 18.025 | 26.45
27.8 |  | 8.7 | 11.08 | 17.875 | 26.1625
27.9 |  | 8.6375 | 10.98 | 17.6857 | 25.8143
28 |  | 8.6 | 10.88 | 17.5125 | 25.625
28.1 |  | 8.5 | 10.78 | 17.275 | 25.4
28.2 |  | 8.45 | 10.68 | 17.1 | 25.025
28.3 |  | 8.4 | 10.58 | 16.9143 | 24.7571
28.4 |  |  | 10.48 | 16.6857 | 24.5
28.5 |  |  | 10.4 | 16.55 | 24.1714
28.6 |  |  | 10.3 | 16.4333 | 23.975
28.7 |  |  | 10.2 | 16.225 | 23.8
28.8 |  |  | 10.15 | 16.0714 | 23.4429
28.9 |  |  | 10.0143 | 15.8429 | 23.175
29 |  |  | 9.9857 | 15.6857 | 22.925
29.1 |  |  | 9.9 | 15.5429 | 22.75
29.2 |  |  | 9.825 | 15.325 | 22.5429
29.3 |  |  | 9.7 | 15.175 | 22.325
29.4 |  |  | 9.59 | 15.0429 | 22.05
29.5 |  |  | 9.5 | 14.9125 | 21.7857
29.6 |  |  | 9.45 | 14.7857 | 21.6125
29.7 |  |  |  | 14.6 | 21.3857
29.8 |  |  |  | 14.4143 | 21.2
29.9 |  |  |  | 14.2875 | 21.0143
30 |  |  |  | 14.2 | 20.7857
30.1 |  |  |  | 14.05 | 20.6
30.2 |  |  |  | 13.9 | 20.425
30.3 |  |  |  | 13.7667 | 20.1875
30.4 |  |  |  | 13.625 | 20.0143
30.5 |  |  |  | 13.5 | 19.825
30.6 |  |  |  | 13.3625 | 19.5875
30.7 |  |  |  | 13.3 | 19.425
30.8 |  |  |  | 13.12 | 19.2286
30.9 |  |  |  | 13.02 | 19.1
31 |  |  |  | 12.92 | 18.8714
31.1 |  |  |  | 12.8 | 18.8
31.2 |  |  |  | 12.7125 | 18.6
31.3 |  |  |  | 12.5857 | 18.4625
31.4 |  |  |  | 12.4 | 18.275
31.5 |  |  |  | 12.3 | 18.1
31.6 |  |  |  | 12.2 | 17.9714
31.7 |  |  |  | 12.15 | 17.8375
31.8 |  |  |  | 12.0333 | 17.6
31.9 |  |  |  | 11.9125 | 17.3714
32 |  |  |  | 11.8 | 17.3
32.1 |  |  |  | 11.7 | 17.2
32.2 |  |  |  | 11.6167 | 17.0429
32.3 |  |  |  | 11.5571 | 16.8375
32.4 |  |  |  |  | 16.7
32.5 |  |  |  |  | 16.5714
32.6 |  |  |  |  | 16.4375
32.7 |  |  |  |  | 16.3125
32.8 |  |  |  |  | 16.1857
32.9 |  |  |  |  | 16.05
33 |  |  |  |  | 15.9143
33.1 |  |  |  |  | 15.7857
33.2 |  |  |  |  | 15.65
33.3 |  |  |  |  | 15.5143
33.4 |  |  |  |  | 15.3875
33.5 |  |  |  |  | 15.2571
33.6 |  |  |  |  | 15.2
33.7 |  |  |  |  | 15.0571
33.8 |  |  |  |  | 14.85
33.9 |  |  |  |  | 14.7833
34 |  |  |  |  | 14.6167
34.1 |  |  |  |  | 14.5571
34.2 |  |  |  |  | 14.5
34.3 |  |  |  |  | 14.36
34.4 |  |  |  |  | 14.16
34.5 |  |  |  |  | 14.0714
34.6 |  |  |  |  | 14
34.7 |  |  |  |  | 14
34.8 |  |  |  |  | 13.8375
34.9 |  |  |  |  | 13.7
35 |  |  |  |  | 13.5714

## Wedge 15 Output factors
Output factors
Wedge 15 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.866 | 0.878 | 0.887 | 0.897 | 0.905 | 0.913 | 0.917 | 0.923
| 4 | 0.878 | 0.894 | 0.904 | 0.918 | 0.928 | 0.938 | 0.943 | 0.951
| 5 | 0.887 | 0.905 | 0.917 | 0.934 | 0.946 | 0.958 | 0.965 | 0.974
| 7 | 0.901 | 0.922 | 0.938 | 0.959 | 0.975 | 0.99 | 1.001 | 1.013
| 10 | 0.912 | 0.937 | 0.955 | 0.98 | 1 | 1.022 | 1.036 | 1.054
| 15 | 0.923 | 0.949 | 0.97 | 0.998 | 1.024 | 1.053 | 1.074 | 1.098
| 20 | 0.928 | 0.955 | 0.978 | 1.01 | 1.04 | 1.074 | 1.1 | 1.13
| 30 | 0.934 | 0.964 | 0.988 | 1.024 | 1.061 | 1.102 | 1.134 | 1.174
| 40 | 0.937 | 0.97 | 0.995 | 1.032 | 1.072 | 1.118 | 1.154 | 1.2

## Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 53.4 | 58.4 | 66.1 | 72.6 | 75.2
0.1 | 57.8 | 62.8 | 70.08 | 76.2 | 78.74
0.2 | 66.0667 | 69.94 | 75.94 | 81.78 | 83.4
0.3 | 76.38 | 79.42 | 84.04 | 88.1 | 89.18
0.4 | 84 | 86.72 | 90 | 92.76 | 93.56
0.5 | 90.0444 | 91.4 | 94.08 | 95.88 | 96.7
0.6 | 93.9 | 95.05 | 96.6 | 97.92 | 98.22
0.7 | 96.55 | 97.2 | 98.36 | 99.1 | 99.34
0.8 | 98.5 | 98.6 | 99.44 | 99.8 | 99.86
0.9 | 99.375 | 99.44 | 99.9 | 100 | 100
1 | 99.8625 | 99.9 | 100 | 100 | 100
1.1 | 100 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 99.9 | 99.9 | 99.9
1.3 | 99.82 | 99.8 | 99.8 | 99.6333 | 99.66
1.4 | 99.56 | 99.65 | 99.6 | 99.34 | 99.36
1.5 | 99.36 | 99.35 | 99.2 | 99.0667 | 98.92
1.6 | 98.8667 | 98.9 | 98.8 | 98.56 | 98.62
1.7 | 98.36 | 98.5 | 98.45 | 98.22 | 98.26
1.8 | 97.78 | 98.2 | 98.125 | 97.88 | 98.02
1.9 | 97.3667 | 97.74 | 97.58 | 97.54 | 97.6333
2 | 96.8667 | 97.3 | 97.275 | 97.2667 | 97.2
2.1 | 96.1667 | 96.84 | 96.95 | 96.88 | 96.8
2.2 | 95.95 | 96.5 | 96.72 | 96.6 | 96.5
2.3 | 95.05 | 95.96 | 96.15 | 96.16 | 96.2
2.4 | 94.62 | 95.34 | 95.82 | 95.78 | 95.8
2.5 | 94.06 | 95.1 | 95.34 | 95.38 | 95.42
2.6 | 93.56 | 94.64 | 94.9 | 95 | 95.06
2.7 | 93.22 | 94.3 | 94.64 | 94.8 | 94.64
2.8 | 92.54 | 93.7333 | 94.32 | 94.4 | 94.24
2.9 | 92.04 | 93.3333 | 93.94 | 94.1 | 93.9
3 | 91.52 | 92.9333 | 93.4667 | 93.6 | 93.6
3.1 | 90.76 | 92.5333 | 93.1 | 93.2 | 93.1
3.2 | 90.26 | 92 | 92.7 | 92.9 | 92.825
3.3 | 89.8667 | 91.52 | 92.4 | 92.4 | 92.4333
3.4 | 89.4333 | 91 | 91.8 | 91.9 | 92.2333
3.5 | 88.8 | 90.6667 | 91.6 | 91.7167 | 91.7
3.6 | 88.2333 | 90.22 | 91.075 | 91.3 | 91.46
3.7 | 87.75 | 89.8 | 90.64 | 91.02 | 91.05
3.8 | 87.3333 | 89.34 | 90.34 | 90.64 | 90.725
3.9 | 86.775 | 88.84 | 89.94 | 90.24 | 90.3
4 | 86.3 | 88.54 | 89.55 | 89.82 | 89.775
4.1 | 85.8 | 88.24 | 89.2 | 89.52 | 89.44
4.2 | 85.2 | 87.66 | 88.8 | 89.14 | 89.15
4.3 | 84.6 | 87.0667 | 88.3667 | 88.7 | 88.82
4.4 | 84.3 | 86.725 | 87.9667 | 88.28 | 88.42
4.5 | 83.6 | 86.38 | 87.6667 | 87.94 | 87.98
4.6 | 83.2 | 85.92 | 87.225 | 87.64 | 87.68
4.7 | 82.8 | 85.48 | 86.9 | 87.22 | 87.38
4.8 | 82.075 | 85.14 | 86.6 | 86.78 | 86.86
4.9 | 81.55 | 84.6 | 86 | 86.48 | 86.46
5 | 81.18 | 84.2 | 85.7 | 86.1 | 86.2
5.1 | 80.74 | 83.8 | 85.2 | 85.78 | 85.8
5.2 | 80.14 | 83.2 | 84.9 | 85.35 | 85.62
5.3 | 79.66 | 82.9 | 84.3667 | 85.02 | 85.14
5.4 | 79.05 | 82.5333 | 84.075 | 84.58 | 84.82
5.5 | 78.6 | 82.1167 | 83.74 | 84.22 | 84.4333
5.6 | 78.1 | 81.65 | 83.38 | 83.86 | 84.08
5.7 | 77.6 | 81.12 | 82.92 | 83.5667 | 83.6667
5.8 | 77.1 | 80.82 | 82.62 | 83.2333 | 83.38
5.9 | 76.7 | 80.28 | 82.18 | 82.95 | 83.08
6 | 76.3 | 79.8 | 81.78 | 82.52 | 82.58
6.1 | 75.6 | 79.38 | 81.3667 | 82 | 82.2
6.2 | 75.325 | 79.06 | 81 | 81.76 | 82
6.3 | 74.72 | 78.74 | 80.6 | 81.28 | 81.6
6.4 | 74.28 | 78.26 | 80.4 | 80.98 | 81.2
6.5 | 73.82 | 77.76 | 79.9 | 80.66 | 80.8667
6.6 | 73.22 | 77.46 | 79.6 | 80.18 | 80.4667
6.7 | 72.86 | 77 | 79.2 | 79.86 | 80.1333
6.8 | 72.3 | 76.6 | 78.8 | 79.6 | 79.75
6.9 | 71.94 | 76.2 | 78.4333 | 79.2 | 79.38
7 | 71.4 | 75.7 | 78.04 | 78.9 | 78.98
7.1 | 70.98 | 75.3 | 77.72 | 78.5 | 78.62
7.2 | 70.58 | 74.9 | 77.16 | 78.2 | 78.24
7.3 | 70.08 | 74.45 | 76.94 | 77.8 | 78
7.4 | 69.58 | 74.025 | 76.54 | 77.5 | 77.62
7.5 | 69.1 | 73.76 | 76.18 | 77.1 | 77.32
7.6 | 68.6 | 73.26 | 75.84 | 76.8 | 76.98
7.7 | 68.2 | 72.88 | 75.3667 | 76.35 | 76.46
7.8 | 67.7 | 72.46 | 75.04 | 75.92 | 76.12
7.9 | 67.3 | 72.04 | 74.64 | 75.56 | 75.76
8 | 66.8 | 71.58 | 74.24 | 75.3714 | 75.56
8.1 | 66.34 | 71.26 | 73.88 | 74.96 | 75.175
8.2 | 65.96 | 70.9 | 73.54 | 74.62 | 74.86
8.3 | 65.62 | 70.3 | 73.22 | 74.28 | 74.46
8.4 | 65.32 | 70 | 72.9 | 73.94 | 74.08
8.5 | 64.78 | 69.7 | 72.52 | 73.58 | 73.78
8.6 | 64.26 | 69.3 | 72.22 | 73.28 | 73.48
8.7 | 63.8333 | 68.9 | 71.82 | 72.88 | 73.06
8.8 | 63.38 | 68.44 | 71.46 | 72.48 | 72.7667
8.9 | 62.9714 | 68.12 | 71.18 | 72.18 | 72.4571
9 | 62.6 | 67.74 | 70.675 | 71.92 | 72.24
9.1 | 62.3 | 67.34 | 70.42 | 71.62 | 71.94
9.2 | 61.9 | 67.02 | 70.12 | 71.32 | 71.48
9.3 | 61.4 | 66.64 | 69.66 | 70.82 | 71.08
9.4 | 61.02 | 66.24 | 69.28 | 70.56 | 70.86
9.5 | 60.62 | 65.85 | 68.98 | 70.26 | 70.52
9.6 | 60.22 | 65.5333 | 68.58 | 69.88 | 70.12
9.7 | 59.86 | 65.075 | 68.36 | 69.58 | 69.92
9.8 | 59.48 | 64.74 | 68.06 | 69.28 | 69.56
9.9 | 58.98 | 64.48 | 67.65 | 68.96 | 69.32
10 | 58.58 | 64.02 | 67.3 | 68.7 | 68.8714
10.1 | 58.1333 | 63.58 | 66.9 | 68.28 | 68.5222
10.2 | 57.875 | 63.38 | 66.5 | 68.1 | 68.2167
10.3 | 57.52 | 62.88 | 66.2 | 67.7 | 67.925
10.4 | 57.08 | 62.525 | 66 | 67.4 | 67.63
10.5 | 56.66 | 62.2 | 65.7 | 67.1 | 67.35
10.6 | 56.28 | 61.8 | 65.3 | 66.8 | 67.0571
10.7 | 55.94 | 61.5 | 64.9 | 66.4 | 66.6667
10.8 | 55.48 | 61.2 | 64.7 | 66.1 | 66.38
10.9 | 55.24 | 60.8 | 64.3 | 65.7 | 65.98
11 | 54.98 | 60.4167 | 64 | 65.4 | 65.65
11.1 | 54.55 | 60.05 | 63.74 | 65.1167 | 65.3556
11.2 | 54.16 | 59.76 | 63.3333 | 64.9167 | 65.1
11.3 | 53.8 | 59.34 | 63.1 | 64.44 | 64.85
11.4 | 53.5 | 59.02 | 62.75 | 64.22 | 64.5667
11.5 | 53.1 | 58.64 | 62.4333 | 63.92 | 64.26
11.6 | 52.7 | 58.34 | 61.96 | 63.64 | 63.86
11.7 | 52.4 | 58.04 | 61.7333 | 63.3 | 63.56
11.8 | 52.1 | 57.82 | 61.5 | 62.86 | 63.36
11.9 | 51.6857 | 57.35 | 61.25 | 62.7 | 63.06
12 | 51.45 | 56.9667 | 60.8333 | 62.44 | 62.76
12.1 | 51.12 | 56.62 | 60.4667 | 62.12 | 62.36
12.2 | 50.7 | 56.34 | 60.22 | 61.82 | 62.08
12.3 | 50.3125 | 55.98 | 59.95 | 61.46 | 61.86
12.4 | 50.06 | 55.64 | 59.6 | 61.14 | 61.48
12.5 | 49.79 | 55.32 | 59.2667 | 60.9333 | 61.18
12.6 | 49.42 | 55.04 | 58.96 | 60.58 | 60.96
12.7 | 49.07 | 54.78 | 58.66 | 60.22 | 60.5833
12.8 | 48.76 | 54.44 | 58.2 | 59.88 | 60.4
12.9 | 48.38 | 54.15 | 58.08 | 59.64 | 60
13 | 48.16 | 53.72 | 57.7667 | 59.4 | 59.8
13.1 | 47.8 | 53.36 | 57.4 | 59.075 | 59.4
13.2 | 47.47 | 53.1 | 57.2 | 58.8333 | 59.2
13.3 | 47.16 | 52.9 | 56.9 | 58.52 | 58.8
13.4 | 46.7833 | 52.5 | 56.5 | 58.3 | 58.5167
13.5 | 46.55 | 52.24 | 56.3 | 57.96 | 58.4
13.6 | 46.25 | 51.9 | 55.9 | 57.7667 | 58
13.7 | 45.95 | 51.62 | 55.6 | 57.4 | 57.8
13.8 | 45.6 | 51.32 | 55.36 | 57.125 | 57.5
13.9 | 45.2 | 50.86 | 55.12 | 56.925 | 57.2
14 | 45 | 50.7 | 54.76 | 56.56 | 56.9
14.1 | 44.7 | 50.44 | 54.52 | 56.32 | 56.7
14.2 | 44.4 | 50.04 | 54.32 | 56.06 | 56.46
14.3 | 44.1222 | 49.74 | 53.92 | 55.86 | 56.16
14.4 | 43.8 | 49.55 | 53.66 | 55.52 | 55.82
14.5 | 43.62 | 49.22 | 53.38 | 55.3 | 55.46
14.6 | 43.2667 | 48.94 | 53.18 | 54.875 | 55.3
14.7 | 42.9333 | 48.6667 | 52.86 | 54.65 | 55.1333
14.8 | 42.64 | 48.35 | 52.56 | 54.3 | 54.78
14.9 | 42.52 | 47.9667 | 52.28 | 54 | 54.58
15 | 42.2 | 47.74 | 52.06 | 53.8 | 54.26
15.1 | 41.84 | 47.44 | 51.78 | 53.6167 | 54.08
15.2 | 41.62 | 47.18 | 51.48 | 53.34 | 53.76
15.3 | 41.42 | 46.78 | 51.3 | 53.04 | 53.5
15.4 | 41.14 | 46.6 | 51 | 52.82 | 53.2
15.5 | 40.84 | 46.38 | 50.7 | 52.62 | 53
15.6 | 40.44 | 46.05 | 50.3571 | 52.32 | 52.7
15.7 | 40.22 | 45.84 | 50.15 | 52.02 | 52.5
15.8 | 39.94 | 45.5667 | 49.84 | 51.74 | 52.2
15.9 | 39.82 | 45.32 | 49.6 | 51.52 | 51.8833
16 | 39.54 | 45.05 | 49.3143 | 51.24 | 51.7
16.1 | 39.22 | 44.7 | 49 | 51.04 | 51.4
16.2 | 39 | 44.5 | 48.8 | 50.84 | 51.2
16.3 | 38.74 | 44.275 | 48.6 | 50.48 | 50.9
16.4 | 38.4667 | 43.96 | 48.4 | 50.2714 | 50.8
16.5 | 38.28 | 43.7 | 48.15 | 49.9667 | 50.5
16.6 | 37.98 | 43.4667 | 47.8125 | 49.74 | 50.2
16.7 | 37.76 | 43.26 | 47.56 | 49.54 | 50
16.8 | 37.46 | 42.9 | 47.3 | 49.34 | 49.7
16.9 | 37.2 | 42.68 | 47.02 | 48.98 | 49.5
17 | 36.98 | 42.48 | 46.8 | 48.725 | 49.2
17.1 | 36.78 | 42.28 | 46.66 | 48.5286 | 49
17.2 | 36.58 | 41.96 | 46.3 | 48.26 | 48.7
17.3 | 36.28 | 41.5857 | 46.02 | 48.04 | 48.4333
17.4 | 35.98 | 41.4333 | 45.82 | 47.76 | 48.2
17.5 | 35.78 | 41.18 | 45.5333 | 47.56 | 48.02
17.6 | 35.475 | 40.94 | 45.34 | 47.29 | 47.74
17.7 | 35.3 | 40.7 | 45.08 | 47.06 | 47.52
17.8 | 35.04 | 40.44 | 44.8333 | 46.86 | 47.22
17.9 | 34.84 | 40.18 | 44.54 | 46.6091 | 47.02
18 | 34.64 | 39.9 | 44.34 | 46.3 | 46.82
18.1 | 34.4 | 39.7 | 44.1 | 46.0727 | 46.5333
18.2 | 34.22 | 39.46 | 43.8667 | 45.82 | 46.44
18.3 | 33.96 | 39.26 | 43.7 | 45.64 | 46.14
18.4 | 33.76 | 39.06 | 43.4 | 45.48 | 45.9
18.5 | 33.56 | 38.8333 | 43.2 | 45.25 | 45.5833
18.6 | 33.32 | 38.55 | 43 | 45.0143 | 45.5
18.7 | 33.1 | 38.36 | 42.7 | 44.78 | 45.3
18.8 | 32.9 | 38.08 | 42.5 | 44.58 | 45.1
18.9 | 32.72 | 37.88 | 42.2 | 44.38 | 44.8
19 | 32.5429 | 37.68 | 42.02 | 44.1909 | 44.6
19.1 | 32.2333 | 37.48 | 41.7 | 43.95 | 44.4
19.2 | 32.02 | 37.1857 | 41.6 | 43.7 | 44.1
19.3 | 31.8667 | 36.94 | 41.4667 | 43.5 | 43.8667
19.4 | 31.5333 | 36.8 | 41.14 | 43.25 | 43.7
19.5 | 31.4 | 36.62 | 40.92 | 43 | 43.5
19.6 | 31.22 | 36.34 | 40.72 | 42.8 | 43.275
19.7 | 31.02 | 36.14 | 40.6 | 42.6 | 43.075
19.8 | 30.88 | 35.88 | 40.375 | 42.4 | 42.775
19.9 | 30.6286 | 35.6667 | 40.16 | 42.15 | 42.6
20 | 30.42 | 35.4286 | 39.86 | 41.95 | 42.42
20.1 | 30.2714 | 35.24 | 39.6 | 41.8091 | 42.22
20.2 | 30.06 | 35.05 | 39.44 | 41.5667 | 42.02
20.3 | 29.96 | 34.86 | 39.1 | 41.35 | 41.78
20.4 | 29.7909 | 34.66 | 39 | 41.15 | 41.58
20.5 | 29.54 | 34.42 | 38.8667 | 40.9167 | 41.4
20.6 | 29.34 | 34.2 | 38.6 | 40.7 | 41.18
20.7 | 29.14 | 34.06 | 38.38 | 40.5 | 40.98
20.8 | 28.9625 | 33.86 | 38.2 | 40.3182 | 40.78
20.9 | 28.8 | 33.7 | 37.9833 | 40.03 | 40.5833
21 | 28.6 | 33.5 | 37.8 | 39.82 | 40.4
21.1 | 28.4 | 33.2 | 37.54 | 39.71 | 40.2
21.2 | 28.2286 | 33.1 | 37.34 | 39.52 | 39.9
21.3 | 28 | 32.82 | 37.14 | 39.2857 | 39.7667
21.4 | 27.9 | 32.6 | 36.98 | 39.12 | 39.6
21.5 | 27.64 | 32.4667 | 36.74 | 38.92 | 39.4
21.6 | 27.46 | 32.2545 | 36.56 | 38.72 | 39.2
21.7 | 27.34 | 32.12 | 36.4 | 38.52 | 38.9
21.8 | 27.1 | 31.9111 | 36.2 | 38.32 | 38.8
21.9 | 26.96 | 31.6636 | 36.02 | 38.08 | 38.6
22 | 26.78 | 31.46 | 35.76 | 37.8636 | 38.42
22.1 | 26.6429 | 31.32 | 35.56 | 37.7625 | 38.16
22.2 | 26.5 | 31.16 | 35.4 | 37.6286 | 38.06
22.3 | 26.38 | 30.9857 | 35.24 | 37.28 | 37.82
22.4 | 26.1778 | 30.725 | 34.92 | 37.12 | 37.6
22.5 | 25.9556 | 30.575 | 34.82 | 36.9 | 37.5
22.6 | 25.7571 | 30.4125 | 34.62 | 36.8 | 37.2
22.7 | 25.58 | 30.2364 | 34.4667 | 36.6 | 37
22.8 | 25.44 | 29.94 | 34.24 | 36.42 | 36.86
22.9 | 25.325 | 29.77 | 34.04 | 36.2 | 36.66
23 | 25.15 | 29.67 | 33.84 | 36 | 36.4857
23.1 | 24.95 | 29.54 | 33.7 | 35.88 | 36.3333
23.2 | 24.89 | 29.35 | 33.5 | 35.68 | 36.14
23.3 | 24.7727 | 29.16 | 33.4 | 35.48 | 35.9
23.4 | 24.5 | 28.98 | 33.2 | 35.3 | 35.8
23.5 | 24.3818 | 28.78 | 33 | 35.1 | 35.65
23.6 | 24.2 | 28.6778 | 32.8 | 34.96 | 35.46
23.7 | 24.1 | 28.48 | 32.7 | 34.78 | 35.35
23.8 | 23.9 | 28.36 | 32.55 | 34.6 | 35.15
23.9 | 23.7909 | 28.16 | 32.3 | 34.3833 | 34.96
24 | 23.7 | 27.96 | 32.16 | 34.28 | 34.76
24.1 | 23.5 | 27.78 | 31.92 | 34.08 | 34.6
24.2 | 23.3 | 27.66 | 31.82 | 33.925 | 34.46
24.3 | 23.2 | 27.4667 | 31.62 | 33.74 | 34.225
24.4 | 23.1 | 27.3333 | 31.42 | 33.54 | 34.12
24.5 | 22.9 | 27.2125 | 31.22 | 33.34 | 33.92
24.6 | 22.7 | 27.02 | 31.1 | 33.1875 | 33.72
24.7 | 22.6 | 26.82 | 30.92 | 33.1 | 33.5
24.8 | 22.4 | 26.62 | 30.8 | 32.84 | 33.4
24.9 | 22.2167 | 26.52 | 30.62 | 32.74 | 33.22
25 | 22.1 | 26.4 | 30.42 | 32.54 | 33
25.1 | 22 | 26.225 | 30.2 | 32.3571 | 32.825
25.2 | 21.9 | 26.1 | 30.1 | 32.2 | 32.64
25.3 | 21.72 | 25.9333 | 30 | 32 | 32.5143
25.4 | 21.62 | 25.74 | 29.84 | 31.9 | 32.38
25.5 | 21.42 | 25.6 | 29.64 | 31.7 | 32.1875
25.6 | 21.3 | 25.5 | 29.54 | 31.6 | 32.04
25.7 | 21.12 | 25.34 | 29.34 | 31.4 | 31.84
25.8 | 21.02 | 25.14 | 29.14 | 31.3 | 31.74
25.9 | 20.92 | 25 | 29 | 31 | 31.54
26 | 20.82 | 24.9 | 28.84 | 30.925 | 31.4
26.1 | 20.7 | 24.74 | 28.7 | 30.8 | 31.3
26.2 | 20.52 | 24.54 | 28.54 | 30.62 | 31.15
26.3 | 20.42 | 24.34 | 28.34 | 30.42 | 30.95
26.4 | 20.32 | 24.26 | 28.24 | 30.3 | 30.75
26.5 | 20.2 | 24.15 | 28.04 | 30.2 | 30.6
26.6 | 20.02 | 23.94 | 27.9 | 30.04 | 30.46
26.7 | 19.9 | 23.84 | 27.7 | 29.84 | 30.3
26.8 | 19.8 | 23.74 | 27.6333 | 29.64 | 30.2
26.9 | 19.65 | 23.64 | 27.5 | 29.54 | 30
27 | 19.5 | 23.44 | 27.3333 | 29.44 | 29.9
27.1 | 19.4 | 23.3 | 27.2 | 29.24 | 29.7
27.2 | 19.3 | 23.16 | 27.06 | 29.1 | 29.62
27.3 | 19.2 | 23.06 | 26.9 | 29 | 29.5
27.4 | 19.06 | 22.96 | 26.725 | 28.78 | 29.32
27.5 | 18.96 | 22.76 | 26.6 | 28.7 | 29.12
27.6 | 18.8 | 22.56 | 26.5 | 28.54 | 29.02
27.7 | 18.66 | 22.44 | 26.32 | 28.35 | 28.92
27.8 | 18.5889 | 22.36 | 26.2 | 28.2 | 28.68
27.9 | 18.46 | 22.26 | 26.0333 | 28.06 | 28.5
28 | 18.36 | 22.1 | 25.8667 | 28 | 28.4
28.1 | 18.2667 | 22 | 25.7 | 27.82 | 28.3
28.2 | 18.0667 | 21.86 | 25.6 | 27.6 | 28.1
28.3 | 17.96 | 21.76 | 25.475 | 27.5 | 28
28.4 | 17.9 | 21.5667 | 25.3 | 27.36 | 27.8
28.5 | 17.8 | 21.48 | 25.2 | 27.2 | 27.7
28.6 | 17.7 | 21.3 | 25.1 | 27.06 | 27.6
28.7 | 17.6 | 21.18 | 24.9 | 26.9667 | 27.4667
28.8 | 17.5 | 21.1 | 24.8 | 26.8 | 27.36
28.9 | 17.4 | 20.98 | 24.66 | 26.6667 | 27.16
29 | 17.2333 | 20.88 | 24.55 | 26.58 | 27.06
29.1 | 17.14 | 20.78 | 24.44 | 26.38 | 26.96
29.2 | 17.025 | 20.58 | 24.3 | 26.2 | 26.76
29.3 | 16.9 | 20.48 | 24.16 | 26.1 | 26.56
29.4 | 16.8 | 20.3 | 24 | 26 | 26.46
29.5 | 16.7 | 20.2 | 23.86 | 25.78 | 26.36
29.6 | 16.6 | 20.1 | 23.76 | 25.68 | 26.2
29.7 | 16.5 | 19.98 | 23.5667 | 25.58 | 26.0333
29.8 | 16.4 | 19.8333 | 23.48 | 25.5 | 25.92
29.9 | 16.3 | 19.7364 | 23.38 | 25.3 | 25.8
30 | 16.2 | 19.64 | 23.28 | 25.2 | 25.7

## SSD80 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 53.7 | 60.1 | 70.5 | 79.1 | 81.9
0.1 | 57.725 | 64.3 | 74.125 | 82.36 | 84.8
0.2 | 65.6 | 71.1 | 79.4333 | 86.6 | 88.85
0.3 | 76.3 | 80.5 | 86.62 | 91.7 | 93.4
0.4 | 84.1333 | 87.6 | 92.2 | 95.5 | 96.4857
0.5 | 89.82 | 92.1333 | 95.6333 | 97.74 | 98.5
0.6 | 94.04 | 95.5333 | 97.72 | 99.16 | 99.4571
0.7 | 96.86 | 97.72 | 99.025 | 99.88 | 99.94
0.8 | 98.54 | 99.08 | 99.84 | 100 | 100
0.9 | 99.62 | 99.84 | 100 | 100 | 100
1 | 99.98 | 100 | 100 | 99.94 | 99.86
1.1 | 100 | 100 | 100 | 99.75 | 99.66
1.2 | 100 | 99.92 | 99.7 | 99.44 | 99.32
1.3 | 99.84 | 99.72 | 99.36 | 98.98 | 98.85
1.4 | 99.4667 | 99.44 | 98.98 | 98.58 | 98.49
1.5 | 99.18 | 99.04 | 98.375 | 98.22 | 98.0818
1.6 | 98.52 | 98.46 | 98.1333 | 97.72 | 97.66
1.7 | 97.92 | 98.02 | 97.58 | 97.32 | 97.2
1.8 | 97.5 | 97.56 | 97.2 | 96.82 | 96.7
1.9 | 96.8 | 97.14 | 96.7 | 96.38 | 96.26
2 | 96.3 | 96.64 | 96.3 | 95.92 | 95.86
2.1 | 95.6 | 96.0667 | 95.7 | 95.42 | 95.4
2.2 | 95.1 | 95.54 | 95.3 | 94.98 | 94.96
2.3 | 94.45 | 95.05 | 94.8 | 94.52 | 94.5
2.4 | 93.9 | 94.46 | 94.4 | 94.1 | 94.06
2.5 | 93.24 | 94.0667 | 93.8333 | 93.64 | 93.66
2.6 | 92.58 | 93.6 | 93.4286 | 93.225 | 93.2333
2.7 | 92.08 | 93 | 92.82 | 92.7333 | 92.78
2.8 | 91.3667 | 92.4 | 92.35 | 92.4 | 92.22
2.9 | 90.95 | 92 | 91.9444 | 91.9 | 91.88
3 | 90.24 | 91.4 | 91.425 | 91.3 | 91.32
3.1 | 89.6 | 90.9 | 91.04 | 90.875 | 90.92
3.2 | 89.1 | 90.6 | 90.5714 | 90.4667 | 90.46
3.3 | 88.55 | 89.9 | 90.15 | 90.0333 | 90.04
3.4 | 87.7667 | 89.5 | 89.7 | 89.68 | 89.58
3.5 | 87.34 | 89.02 | 89.175 | 89.18 | 89.26
3.6 | 86.775 | 88.4667 | 88.72 | 88.7333 | 88.74
3.7 | 86.1667 | 88.1 | 88.4 | 88.26 | 88.26
3.8 | 85.56 | 87.58 | 87.8 | 87.88 | 87.86
3.9 | 84.9667 | 87.08 | 87.3 | 87.45 | 87.44
4 | 84.52 | 86.52 | 86.9 | 87 | 86.96
4.1 | 84.02 | 86.06 | 86.4333 | 86.6 | 86.6
4.2 | 83.36 | 85.45 | 86.04 | 86.2 | 86.1
4.3 | 82.7 | 85.04 | 85.5667 | 85.7 | 85.7
4.4 | 82.22 | 84.56 | 84.96 | 85.2 | 85.3
4.5 | 81.6 | 84.04 | 84.5667 | 84.82 | 84.9125
4.6 | 81.1 | 83.56 | 84.14 | 84.35 | 84.4
4.7 | 80.4 | 83.1333 | 83.725 | 83.86 | 83.95
4.8 | 79.9 | 82.68 | 83.3 | 83.58 | 83.6
4.9 | 79.4 | 82.28 | 82.9 | 83.18 | 83.2
5 | 78.9333 | 81.62 | 82.35 | 82.7667 | 82.75
5.1 | 78.3429 | 81.1 | 81.88 | 82.3 | 82.34
5.2 | 77.82 | 80.72 | 81.6 | 81.94 | 81.9
5.3 | 77.25 | 80.2 | 81.05 | 81.35 | 81.5364
5.4 | 76.68 | 79.75 | 80.7 | 81 | 81.05
5.5 | 76.22 | 79.35 | 80.3 | 80.62 | 80.73
5.6 | 75.68 | 78.7667 | 79.9 | 80.12 | 80.2909
5.7 | 75.1 | 78.32 | 79.4 | 79.82 | 79.8
5.8 | 74.55 | 77.78 | 78.9 | 79.42 | 79.3444
5.9 | 74.08 | 77.25 | 78.5 | 79.06 | 78.98
6 | 73.64 | 76.9 | 78.1 | 78.66 | 78.58
6.1 | 73.04 | 76.5333 | 77.6 | 78.16 | 78.26
6.2 | 72.56 | 76.04 | 77.1 | 77.66 | 77.78
6.3 | 72.12 | 75.45 | 76.74 | 77.28 | 77.36
6.4 | 71.56 | 75 | 76.44 | 76.86 | 77.02
6.5 | 71.04 | 74.7 | 75.84 | 76.44 | 76.6
6.6 | 70.54 | 74.1 | 75.46 | 75.98 | 76.18
6.7 | 69.9667 | 73.7 | 75.04 | 75.66 | 75.86
6.8 | 69.56 | 73.2 | 74.64 | 75.26 | 75.44
6.9 | 68.95 | 72.8 | 74.1 | 74.78 | 74.96
7 | 68.42 | 72.25 | 73.86 | 74.38 | 74.66
7.1 | 67.98 | 71.82 | 73.54 | 74.0667 | 74.2667
7.2 | 67.6 | 71.36 | 72.98 | 73.7 | 73.8
7.3 | 67.05 | 70.94 | 72.52 | 73.26 | 73.4
7.4 | 66.5333 | 70.54 | 72.1667 | 72.9 | 73.0667
7.5 | 66.0333 | 70.04 | 71.6667 | 72.4 | 72.72
7.6 | 65.625 | 69.6667 | 71.25 | 72 | 72.3333
7.7 | 65.05 | 69.18 | 70.8667 | 71.7 | 71.92
7.8 | 64.625 | 68.78 | 70.5333 | 71.3 | 71.54
7.9 | 64.15 | 68.3 | 70.1333 | 70.9 | 71.06
8 | 63.8333 | 67.82 | 69.7 | 70.5 | 70.78
8.1 | 63.3333 | 67.48 | 69.3333 | 70.24 | 70.46
8.2 | 62.8667 | 66.98 | 68.9667 | 69.84 | 70.0667
8.3 | 62.4333 | 66.62 | 68.5 | 69.4333 | 69.66
8.4 | 62.0333 | 66.28 | 68.2 | 69.14 | 69.35
8.5 | 61.5667 | 65.78 | 67.875 | 68.74 | 69
8.6 | 61.1 | 65.38 | 67.52 | 68.34 | 68.6
8.7 | 60.6 | 65.02 | 67.12 | 68.02 | 68.2
8.8 | 60.25 | 64.52 | 66.7 | 67.54 | 67.8
8.9 | 59.7667 | 64.2 | 66.4 | 67.28 | 67.5
9 | 59.32 | 63.82 | 66.02 | 66.98 | 67.1
9.1 | 58.9 | 63.32 | 65.5333 | 66.4667 | 66.7
9.2 | 58.5667 | 62.92 | 65.2 | 66.14 | 66.5
9.3 | 58.225 | 62.6 | 64.82 | 65.88 | 66
9.4 | 57.6667 | 62.14 | 64.4455 | 65.48 | 65.7167
9.5 | 57.34 | 61.74 | 64.06 | 65.08 | 65.34
9.6 | 56.925 | 61.42 | 63.66 | 64.74 | 64.94
9.7 | 56.48 | 61 | 63.32 | 64.4 | 64.64
9.8 | 56.125 | 60.6667 | 62.96 | 64.02 | 64.24
9.9 | 55.65 | 60.3333 | 62.62 | 63.68 | 63.94
10 | 55.2 | 59.8333 | 62.32 | 63.35 | 63.54
10.1 | 54.88 | 59.44 | 61.9625 | 62.98 | 63.22
10.2 | 54.58 | 59.04 | 61.58 | 62.72 | 62.8667
10.3 | 54.24 | 58.64 | 61.15 | 62.32 | 62.62
10.4 | 53.82 | 58.32 | 60.85 | 62.02 | 62.225
10.5 | 53.35 | 58.02 | 60.55 | 61.72 | 61.92
10.6 | 53 | 57.64 | 60.1 | 61.22 | 61.48
10.7 | 52.64 | 57.34 | 59.8 | 60.96 | 61.24
10.8 | 52.28 | 56.92 | 59.4 | 60.62 | 60.88
10.9 | 51.92 | 56.45 | 59.1333 | 60.36 | 60.64
11 | 51.5 | 56.1667 | 58.8 | 59.92 | 60.28
11.1 | 51.16 | 55.74 | 58.44 | 59.66 | 59.94
11.2 | 50.82 | 55.38 | 58.04 | 59.3333 | 59.58
11.3 | 50.36 | 55.04 | 57.74 | 58.975 | 59.3
11.4 | 50.02 | 54.74 | 57.4 | 58.7333 | 59.05
11.5 | 49.72 | 54.38 | 57.1167 | 58.36 | 58.7
11.6 | 49.42 | 53.94 | 56.74 | 58.1333 | 58.325
11.7 | 49.02 | 53.74 | 56.4667 | 57.76 | 58.08
11.8 | 48.68 | 53.38 | 56.125 | 57.48 | 57.675
11.9 | 48.36 | 52.9143 | 55.82 | 57.16 | 57.375
12 | 48.02 | 52.6 | 55.54 | 56.78 | 57.075
12.1 | 47.62 | 52.3143 | 55.18 | 56.46 | 56.7
12.2 | 47.26 | 51.95 | 54.78 | 56.2 | 56.45
12.3 | 46.96 | 51.6 | 54.44 | 55.78 | 56.1
12.4 | 46.58 | 51.2 | 54.14 | 55.5 | 55.85
12.5 | 46.36 | 50.9111 | 53.86 | 55.18 | 55.55
12.6 | 45.96 | 50.6125 | 53.52 | 54.9 | 55.2
12.7 | 45.58 | 50.4 | 53.26 | 54.6667 | 54.925
12.8 | 45.28 | 49.9667 | 52.96 | 54.32 | 54.6
12.9 | 44.98 | 49.6 | 52.62 | 53.96 | 54.3
13 | 44.7333 | 49.2636 | 52.32 | 53.66 | 54
13.1 | 44.325 | 49.02 | 51.96 | 53.35 | 53.76
13.2 | 44.0333 | 48.74 | 51.68 | 53.12 | 53.46
13.3 | 43.7333 | 48.43 | 51.36 | 52.82 | 53.1333
13.4 | 43.5 | 48.04 | 51.075 | 52.52 | 52.875
13.5 | 43.3 | 47.7667 | 50.825 | 52.22 | 52.56
13.6 | 42.9 | 47.4636 | 50.5545 | 51.94 | 52.2714
13.7 | 42.6 | 47.14 | 50.18 | 51.68 | 52
13.8 | 42.3 | 46.78 | 49.94 | 51.38 | 51.76
13.9 | 41.9667 | 46.48 | 49.58 | 51.14 | 51.38
14 | 41.6667 | 46.18 | 49.34 | 50.78 | 51.15
14.1 | 41.4 | 45.9 | 49.1 | 50.54 | 50.86
14.2 | 41.0667 | 45.7 | 48.81 | 50.24 | 50.56
14.3 | 40.8 | 45.3 | 48.51 | 49.98 | 50.31
14.4 | 40.4667 | 45.1 | 48.21 | 49.74 | 50.0625
14.5 | 40.1667 | 44.7 | 47.925 | 49.44 | 49.7667
14.6 | 39.9667 | 44.5 | 47.7286 | 49.14 | 49.5
14.7 | 39.68 | 44.2 | 47.4 | 48.84 | 49.2
14.8 | 39.425 | 43.9 | 47.075 | 48.54 | 49
14.9 | 39.14 | 43.7 | 46.82 | 48.35 | 48.6571
15 | 38.84 | 43.4 | 46.56 | 48.12 | 48.4429
15.1 | 38.6667 | 43.0857 | 46.22 | 47.76 | 48.1
15.2 | 38.3 | 42.7818 | 45.9667 | 47.56 | 47.8875
15.3 | 38.1 | 42.5444 | 45.72 | 47.26 | 47.6375
15.4 | 37.8 | 42.26 | 45.42 | 47.02 | 47.39
15.5 | 37.5 | 41.96 | 45.15 | 46.76 | 47.16
15.6 | 37.26 | 41.7333 | 44.92 | 46.5333 | 46.89
15.7 | 37.0333 | 41.475 | 44.64 | 46.2 | 46.66
15.8 | 36.8 | 41.16 | 44.34 | 46.0333 | 46.36
15.9 | 36.5 | 40.8714 | 44.12 | 45.68 | 46.1333
16 | 36.3 | 40.62 | 43.84 | 45.38 | 45.88
16.1 | 35.94 | 40.32 | 43.64 | 45.18 | 45.6286
16.2 | 35.72 | 40.02 | 43.35 | 44.96 | 45.34
16.3 | 35.5 | 39.825 | 43.04 | 44.68 | 45.12
16.4 | 35.275 | 39.58 | 42.86 | 44.48 | 44.84
16.5 | 35.06 | 39.325 | 42.58 | 44.18 | 44.64
16.6 | 34.78 | 39.0833 | 42.36 | 43.98 | 44.44
16.7 | 34.56 | 38.825 | 42.08 | 43.78 | 44.18
16.8 | 34.275 | 38.6 | 41.8 | 43.4833 | 43.94
16.9 | 34.1 | 38.35 | 41.58 | 43.3 | 43.64
17 | 33.84 | 38.05 | 41.38 | 42.98 | 43.44
17.1 | 33.5727 | 37.76 | 41.1 | 42.8 | 43.18
17.2 | 33.3 | 37.55 | 40.9 | 42.5 | 42.94
17.3 | 32.9833 | 37.34 | 40.6 | 42.3167 | 42.7
17.4 | 32.9 | 37.08 | 40.3 | 42.04 | 42.45
17.5 | 32.7 | 36.84 | 40.1 | 41.8 | 42.24
17.6 | 32.5 | 36.64 | 39.9 | 41.54 | 41.96
17.7 | 32.2 | 36.34 | 39.6 | 41.32 | 41.65
17.8 | 32 | 36.1 | 39.4 | 41.02 | 41.46
17.9 | 31.7 | 35.9 | 39.1 | 40.85 | 41.26
18 | 31.5 | 35.6667 | 38.9 | 40.66 | 41.05
18.1 | 31.3 | 35.4 | 38.6 | 40.38 | 40.8
18.2 | 31.1182 | 35.2 | 38.4 | 40.18 | 40.575
18.3 | 30.8857 | 35 | 38.2 | 39.88 | 40.36
18.4 | 30.7 | 34.8 | 38.05 | 39.7 | 40.16
18.5 | 30.48 | 34.5 | 37.8 | 39.56 | 39.96
18.6 | 30.28 | 34.3 | 37.55 | 39.28 | 39.775
18.7 | 30.08 | 34.1 | 37.35 | 39.08 | 39.46
18.8 | 29.88 | 33.9 | 37.1 | 38.88 | 39.275
18.9 | 29.58 | 33.675 | 36.9167 | 38.7 | 39.075
19 | 29.44 | 33.46 | 36.775 | 38.4 | 38.875
19.1 | 29.24 | 33.22 | 36.525 | 38.2 | 38.62
19.2 | 29.0667 | 33.06 | 36.225 | 37.9833 | 38.4667
19.3 | 28.9 | 32.82 | 36.0667 | 37.8 | 38.1667
19.4 | 28.7 | 32.56 | 35.86 | 37.6 | 37.98
19.5 | 28.4 | 32.46 | 35.5667 | 37.375 | 37.8667
19.6 | 28.2 | 32.16 | 35.4667 | 37.15 | 37.58
19.7 | 28.1 | 31.96 | 35.1667 | 36.88 | 37.38
19.8 | 27.9 | 31.76 | 35 | 36.74 | 37.1667
19.9 | 27.7286 | 31.56 | 34.775 | 36.54 | 37
20 | 27.5714 | 31.3667 | 34.7 | 36.35 | 36.7
20.1 | 27.4 | 31.2 | 34.45 | 36.15 | 36.5833
20.2 | 27.08 | 30.95 | 34.15 | 35.96 | 36.3
20.3 | 26.95 | 30.72 | 33.94 | 35.75 | 36.0833
20.4 | 26.8 | 30.62 | 33.8 | 35.46 | 36
20.5 | 26.62 | 30.32 | 33.5833 | 35.36 | 35.8
20.6 | 26.46 | 30.22 | 33.4 | 35.16 | 35.6667
20.7 | 26.26 | 30.0333 | 33.2 | 34.96 | 35.3
20.8 | 26.1 | 29.84 | 33.04 | 34.76 | 35.175
20.9 | 25.96 | 29.64 | 32.8 | 34.6 | 34.975
21 | 25.76 | 29.44 | 32.6 | 34.42 | 34.8
21.1 | 25.56 | 29.24 | 32.4 | 34.16 | 34.65
21.2 | 25.32 | 29.04 | 32.16 | 33.96 | 34.44
21.3 | 25.16 | 28.84 | 32 | 33.76 | 34.25
21.4 | 24.96 | 28.725 | 31.825 | 33.56 | 33.96
21.5 | 24.86 | 28.5 | 31.6333 | 33.46 | 33.85
21.6 | 24.7 | 28.3 | 31.5 | 33.18 | 33.66
21.7 | 24.5667 | 28.2 | 31.3 | 33.08 | 33.46
21.8 | 24.3 | 28 | 31.12 | 32.78 | 33.26
21.9 | 24.18 | 27.8 | 30.92 | 32.68 | 33.1
22 | 24.08 | 27.6 | 30.72 | 32.3833 | 32.9
22.1 | 23.88 | 27.4 | 30.58 | 32.2333 | 32.72
22.2 | 23.68 | 27.25 | 30.3667 | 32.1 | 32.56
22.3 | 23.58 | 27.05 | 30.2 | 32 | 32.32
22.4 | 23.38 | 26.96 | 30 | 31.74 | 32.1
22.5 | 23.1667 | 26.66 | 29.88 | 31.5 | 32
22.6 | 23 | 26.56 | 29.68 | 31.4 | 31.78
22.7 | 22.9 | 26.42 | 29.48 | 31.2 | 31.68
22.8 | 22.7833 | 26.16 | 29.3 | 31 | 31.48
22.9 | 22.6 | 26.06 | 29.1333 | 30.8667 | 31.28
23 | 22.4 | 25.8667 | 28.9333 | 30.6667 | 31.1
23.1 | 22.3 | 25.7 | 28.7333 | 30.5 | 30.98
23.2 | 22.14 | 25.6 | 28.6 | 30.36 | 30.78
23.3 | 22.04 | 25.4 | 28.48 | 30.2 | 30.6
23.4 | 21.84 | 25.225 | 28.3 | 30 | 30.48
23.5 | 21.7 | 25.12 | 28.1 | 29.86 | 30.2833
23.6 | 21.575 | 24.9333 | 28 | 29.66 | 30.1
23.7 | 21.44 | 24.8 | 27.8 | 29.5 | 29.9833
23.8 | 21.3 | 24.68 | 27.7 | 29.42 | 29.8
23.9 | 21.2 | 24.5 | 27.5 | 29.2 | 29.66
24 | 21 | 24.3333 | 27.3 | 29 | 29.46
24.1 | 20.9 | 24.175 | 27.2 | 28.88 | 29.3
24.2 | 20.7 | 24 | 27 | 28.775 | 29.15
24.3 | 20.6 | 23.8 | 26.8 | 28.56 | 28.9667
24.4 | 20.4333 | 23.7 | 26.7 | 28.38 | 28.825
24.5 | 20.32 | 23.575 | 26.5 | 28.28 | 28.68
24.6 | 20.2 | 23.4 | 26.4 | 28.08 | 28.4667
24.7 | 20.06 | 23.2 | 26.2 | 27.98 | 28.38
24.8 | 19.9 | 23.1 | 26 | 27.7 | 28.18
24.9 | 19.8 | 23 | 25.92 | 27.58 | 28.08
25 | 19.66 | 22.8 | 25.7167 | 27.46 | 27.88
25.1 | 19.5 | 22.7 | 25.6 | 27.28 | 27.7
25.2 | 19.4 | 22.5 | 25.42 | 27.16 | 27.58
25.3 | 19.26 | 22.4 | 25.3 | 27 | 27.38
25.4 | 19.1 | 22.3 | 25.12 | 26.8 | 27.28
25.5 | 19 | 22.1 | 25 | 26.7 | 27.1
25.6 | 18.875 | 22 | 24.82 | 26.5 | 27
25.7 | 18.8 | 21.9 | 24.72 | 26.4 | 26.79
25.8 | 18.62 | 21.72 | 24.59 | 26.26 | 26.69
25.9 | 18.6 | 21.575 | 24.4714 | 26.1 | 26.5714
26 | 18.42 | 21.4 | 24.3 | 25.925 | 26.4
26.1 | 18.22 | 21.28 | 24.15 | 25.82 | 26.3
26.2 | 18.14 | 21.1667 | 24.04 | 25.62 | 26.1
26.3 | 18.04 | 21 | 23.88 | 25.5 | 25.9571
26.4 | 17.94 | 20.88 | 23.7 | 25.4 | 25.78
26.5 | 17.8 | 20.78 | 23.58 | 25.22 | 25.64
26.6 | 17.7 | 20.6 | 23.4 | 25.1 | 25.54
26.7 | 17.5 | 20.5 | 23.28 | 25 | 25.3909
26.8 | 17.5 | 20.38 | 23.18 | 24.82 | 25.2
26.9 | 17.3 | 20.28 | 23.0833 | 24.62 | 25.0444
27 | 17.2167 | 20.09 | 22.8571 | 24.5333 | 24.9455
27.1 | 17.1 | 19.9833 | 22.78 | 24.4 | 24.8
27.2 | 17 | 19.9 | 22.68 | 24.24 | 24.7
27.3 | 16.9 | 19.8 | 22.48 | 24.14 | 24.56
27.4 | 16.8 | 19.66 | 22.38 | 24 | 24.4
27.5 | 16.68 | 19.5 | 22.2 | 23.84 | 24.3
27.6 | 16.5 | 19.36 | 22.1 | 23.74 | 24.12
27.7 | 16.48 | 19.26 | 21.98 | 23.6 | 24
27.8 | 16.38 | 19.1667 | 21.88 | 23.475 | 23.9
27.9 | 16.28 | 19 | 21.7 | 23.36 | 23.8
28 | 16.18 | 18.875 | 21.58 | 23.22 | 23.6333
28.1 | 16.08 | 18.8 | 21.48 | 23.06 | 23.5
28.2 | 15.9143 | 18.66 | 21.3 | 22.95 | 23.34
28.3 | 15.81 | 18.5 | 21.2 | 22.8 | 23.24
28.4 | 15.74 | 18.4 | 21 | 22.6833 | 23.14
28.5 | 15.64 | 18.3 | 21 | 22.58 | 23
28.6 | 15.5 | 18.2 | 20.8 | 22.4143 | 22.9
28.7 | 15.4 | 18.08 | 20.7 | 22.3 | 22.74
28.8 | 15.3 | 17.98 | 20.6 | 22.17 | 22.6
28.9 | 15.2 | 17.88 | 20.5 | 22.07 | 22.44
29 | 15.1 | 17.7833 | 20.4 | 21.97 | 22.34
29.1 | 15.05 | 17.7 | 20.3 | 21.84 | 22.2571
29.2 | 14.96 | 17.5 | 20.14 | 21.7 | 22.1143
29.3 | 14.8 | 17.4 | 20 | 21.62 | 22.01
29.4 | 14.7 | 17.3 | 19.9 | 21.52 | 21.84
29.5 | 14.6 | 17.2 | 19.74 | 21.34 | 21.7
29.6 | 14.54 | 17.1 | 19.7 | 21.2 | 21.6
29.7 | 14.37 | 17 | 19.5 | 21.12 | 21.5
29.8 | 14.3 | 16.9 | 19.4 | 21 | 21.4
29.9 | 14.25 | 16.8 | 19.3 | 20.9 | 21.3
30 | 14.2 | 16.7 | 19.2 | 20.8 | 21.2

## SSD120 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 55.2 | 60 | 66.8 | 72.6 | 74.9
0.1 | 60.6 | 64.9 | 71.3 | 76.975 | 78.82
0.2 | 69.58 | 73.4 | 78.475 | 82.8 | 84.25
0.3 | 79.55 | 82.0333 | 85.825 | 88.7 | 90.1
0.4 | 87 | 88.4 | 91.075 | 93.1 | 93.95
0.5 | 91.325 | 92.9 | 94.575 | 96.2 | 96.9
0.6 | 94.725 | 95.6 | 96.9 | 97.9 | 98.55
0.7 | 97.125 | 97.5 | 98.24 | 99.2 | 99.4
0.8 | 98.7 | 98.8 | 99.22 | 99.7 | 99.74
0.9 | 99.525 | 99.6 | 99.8 | 99.9 | 100
1 | 99.95 | 99.9 | 100 | 100 | 100
1.1 | 100 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 99.95 | 100 | 99.775
1.3 | 99.775 | 99.76 | 99.8 | 99.85 | 99.625
1.4 | 99.6 | 99.525 | 99.45 | 99.625 | 99.4
1.5 | 99.25 | 99.375 | 99.16 | 99.075 | 99.225
1.6 | 98.825 | 98.95 | 99 | 98.9 | 98.975
1.7 | 98.46 | 98.45 | 98.65 | 98.575 | 98.65
1.8 | 98 | 98.125 | 98.225 | 98.425 | 98.375
1.9 | 97.5 | 97.85 | 97.975 | 98.025 | 98.025
2 | 97.15 | 97.5 | 97.75 | 97.6 | 97.775
2.1 | 96.8 | 97.1 | 97.3286 | 97.38 | 97.38
2.2 | 96.1 | 96.65 | 96.9 | 97.1 | 97.08
2.3 | 95.6 | 96.4 | 96.6 | 96.6 | 96.8
2.4 | 95.2 | 95.95 | 96.2 | 96.45 | 96.3333
2.5 | 94.7 | 95.5 | 95.95 | 96.2 | 96.15
2.6 | 94.3333 | 95.175 | 95.6 | 95.85 | 95.8
2.7 | 93.675 | 94.8 | 95.15 | 95.6 | 95.4
2.8 | 93.35 | 94.475 | 94.9 | 95.2 | 95.16
2.9 | 92.75 | 94.05 | 94.5333 | 94.7 | 94.8
3 | 92.325 | 93.65 | 94.1 | 94.45 | 94.475
3.1 | 91.95 | 93.225 | 93.9 | 94.2 | 94.14
3.2 | 91.375 | 92.775 | 93.4 | 93.86 | 93.85
3.3 | 90.85 | 92.425 | 93.1 | 93.525 | 93.4714
3.4 | 90.45 | 92.075 | 92.7333 | 93.2 | 93.22
3.5 | 89.975 | 91.575 | 92.5 | 92.825 | 92.7667
3.6 | 89.6 | 91.1 | 92.1 | 92.55 | 92.475
3.7 | 89.1 | 90.9333 | 91.8333 | 92.15 | 92.15
3.8 | 88.4 | 90.4 | 91.425 | 91.825 | 91.875
3.9 | 87.9 | 90.05 | 91.0667 | 91.425 | 91.6
4 | 87.5 | 89.48 | 90.675 | 91.2 | 91.1667
4.1 | 87.1 | 89.2 | 90.3 | 90.85 | 90.7857
4.2 | 86.55 | 88.75 | 90 | 90.425 | 90.4667
4.3 | 85.94 | 88.3 | 89.6 | 90.2 | 90.1
4.4 | 85.42 | 88.05 | 89.1 | 89.8 | 89.78
4.5 | 84.96 | 87.5 | 88.75 | 89.5 | 89.5
4.6 | 84.55 | 87.1 | 88.6 | 89.2 | 89.2333
4.7 | 84.15 | 86.7 | 88.15 | 88.8 | 88.8667
4.8 | 83.45 | 86.3 | 87.825 | 88.3667 | 88.55
4.9 | 83.05 | 85.85 | 87.45 | 88.05 | 88.1333
5 | 82.65 | 85.575 | 87.125 | 87.7 | 87.775
5.1 | 82.225 | 85.25 | 86.775 | 87.35 | 87.5
5.2 | 81.675 | 84.7333 | 86.525 | 87.12 | 87.15
5.3 | 81.32 | 84.3833 | 86.175 | 86.85 | 86.92
5.4 | 80.8333 | 84.05 | 85.75 | 86.525 | 86.5857
5.5 | 80.4 | 83.65 | 85.375 | 86.05 | 86.125
5.6 | 79.9 | 83.2625 | 85.04 | 85.825 | 85.8333
5.7 | 79.25 | 82.85 | 84.76 | 85.5667 | 85.475
5.8 | 78.9 | 82.475 | 84.35 | 85.15 | 85.25
5.9 | 78.5333 | 82.0111 | 83.9 | 84.7333 | 84.85
6 | 77.9667 | 81.6 | 83.62 | 84.35 | 84.475
6.1 | 77.5333 | 81.2 | 83.15 | 84.025 | 84.2
6.2 | 77.1 | 80.8 | 82.82 | 83.675 | 83.9
6.3 | 76.64 | 80.525 | 82.525 | 83.4 | 83.575
6.4 | 76.225 | 80.15 | 82.1667 | 83 | 83.25
6.5 | 75.65 | 79.85 | 81.775 | 82.7 | 82.875
6.6 | 75.25 | 79.4 | 81.425 | 82.4 | 82.55
6.7 | 74.75 | 78.9 | 81.18 | 82.1 | 82.175
6.8 | 74.2667 | 78.4667 | 80.9 | 81.85 | 81.75
6.9 | 73.975 | 78.0857 | 80.5333 | 81.54 | 81.475
7 | 73.45 | 77.8 | 80.05 | 81.1333 | 81.225
7.1 | 72.925 | 77.3625 | 79.8333 | 80.85 | 80.9
7.2 | 72.5 | 76.95 | 79.425 | 80.45 | 80.55
7.3 | 72.1 | 76.5125 | 78.975 | 80.175 | 80.3
7.4 | 71.7 | 76.05 | 78.65 | 79.825 | 79.95
7.5 | 71.2 | 75.6667 | 78.375 | 79.475 | 79.65
7.6 | 70.9 | 75.4167 | 78.025 | 79.05 | 79.3222
7.7 | 70.3 | 75.1 | 77.6 | 78.775 | 78.85
7.8 | 69.8667 | 74.6333 | 77.35 | 78.525 | 78.6167
7.9 | 69.45 | 74.3 | 77.1 | 78.15 | 78.3625
8 | 69.1 | 74 | 76.8 | 77.8 | 78.025
8.1 | 68.72 | 73.52 | 76.4 | 77.6 | 77.7125
8.2 | 68.2 | 73.025 | 76 | 77.2 | 77.425
8.3 | 67.925 | 72.76 | 75.64 | 76.88 | 77.1
8.4 | 67.5 | 72.525 | 75.375 | 76.45 | 76.775
8.5 | 67.02 | 72.15 | 74.95 | 76.2 | 76.4
8.6 | 66.7 | 71.75 | 74.675 | 75.9 | 76.1667
8.7 | 66.25 | 71.4 | 74.45 | 75.6 | 75.825
8.8 | 65.85 | 70.9 | 74 | 75.25 | 75.475
8.9 | 65.4333 | 70.6 | 73.725 | 75 | 75.225
9 | 64.96 | 70.2 | 73.425 | 74.8 | 74.825
9.1 | 64.6333 | 69.82 | 73.025 | 74.4 | 74.5
9.2 | 64.15 | 69.65 | 72.75 | 74.06 | 74.25
9.3 | 63.725 | 69.15 | 72.325 | 73.75 | 73.8
9.4 | 63.35 | 68.75 | 71.975 | 73.2667 | 73.4667
9.5 | 62.8 | 68.5 | 71.625 | 73 | 73.2
9.6 | 62.525 | 68.15 | 71.375 | 72.825 | 72.9571
9.7 | 62.175 | 67.9125 | 71.025 | 72.575 | 72.6111
9.8 | 61.7 | 67.4125 | 70.76 | 72.125 | 72.375
9.9 | 61.44 | 67.0375 | 70.2667 | 71.875 | 72.0375
10 | 61.1 | 66.66 | 69.975 | 71.525 | 71.775
10.1 | 60.75 | 66.32 | 69.74 | 71.25 | 71.4875
10.2 | 60.3 | 65.95 | 69.5 | 70.925 | 71.2
10.3 | 59.9333 | 65.64 | 69.15 | 70.65 | 70.8333
10.4 | 59.4 | 65.28 | 68.9 | 70.325 | 70.575
10.5 | 59.2 | 64.95 | 68.6 | 69.98 | 70.25
10.6 | 58.7 | 64.6 | 68.2 | 69.7 | 69.9375
10.7 | 58.3 | 64.25 | 67.9 | 69.4 | 69.7
10.8 | 58 | 64.1 | 67.6 | 69.2 | 69.3
10.9 | 57.65 | 63.5 | 67.4 | 68.8 | 69.1
11 | 57.3 | 63.2 | 67 | 68.6 | 68.6
11.1 | 57.05 | 62.92 | 66.6 | 68.2 | 68.45
11.2 | 56.7333 | 62.5 | 66.3 | 67.85 | 68.2
11.3 | 56.2 | 62.12 | 66.05 | 67.6 | 67.875
11.4 | 55.7833 | 61.975 | 65.725 | 67.35 | 67.6
11.5 | 55.5 | 61.52 | 65.45 | 67 | 67.225
11.6 | 55.275 | 61.1 | 65.125 | 66.85 | 66.9
11.7 | 54.95 | 60.85 | 64.875 | 66.4 | 66.725
11.8 | 54.6 | 60.575 | 64.6 | 66.1 | 66.4
11.9 | 54.15 | 60.275 | 64.175 | 65.9 | 66.1
12 | 53.875 | 60.025 | 63.925 | 65.7 | 65.8
12.1 | 53.65 | 59.66 | 63.575 | 65.35 | 65.55
12.2 | 53.275 | 59.3667 | 63.325 | 65.06 | 65.25
12.3 | 52.86 | 58.98 | 63.06 | 64.725 | 64.9667
12.4 | 52.6333 | 58.65 | 62.68 | 64.45 | 64.6
12.5 | 52.3 | 58.4 | 62.4 | 64.2333 | 64.3667
12.6 | 52.04 | 58.1 | 62.05 | 63.96 | 64.1
12.7 | 51.7 | 57.6 | 61.7 | 63.6333 | 63.84
12.8 | 51.35 | 57.35 | 61.45 | 63.3 | 63.6222
12.9 | 51 | 57 | 61.2 | 63 | 63.3667
13 | 50.65 | 56.75 | 60.95 | 62.7667 | 63.075
13.1 | 50.3 | 56.5 | 60.7 | 62.5 | 62.75
13.2 | 49.95 | 56.15 | 60.35 | 62.2 | 62.3833
13.3 | 49.7 | 55.9 | 60.1 | 62 | 62.2167
13.4 | 49.45 | 55.6 | 59.85 | 61.66 | 61.875
13.5 | 49.2 | 55.2 | 59.5 | 61.425 | 61.575
13.6 | 48.8 | 54.96 | 59.25 | 60.98 | 61.3333
13.7 | 48.35 | 54.6333 | 58.9 | 60.825 | 61.15
13.8 | 48.1667 | 54.38 | 58.55 | 60.475 | 60.875
13.9 | 47.9 | 54.05 | 58.3 | 60.25 | 60.6
14 | 47.6667 | 53.7 | 58 | 59.9 | 60.325
14.1 | 47.275 | 53.45 | 57.72 | 59.725 | 59.9857
14.2 | 46.925 | 53.1 | 57.45 | 59.4 | 59.66
14.3 | 46.675 | 52.925 | 57.125 | 59.15 | 59.5143
14.4 | 46.35 | 52.575 | 56.875 | 58.8 | 59.25
14.5 | 46.15 | 52.25 | 56.625 | 58.65 | 58.975
14.6 | 45.75 | 51.875 | 56.375 | 58.4 | 58.6714
14.7 | 45.475 | 51.625 | 56.125 | 58.05 | 58.4
14.8 | 45.225 | 51.275 | 55.875 | 57.8 | 58.125
14.9 | 44.875 | 51.14 | 55.525 | 57.5 | 57.875
15 | 44.725 | 50.8667 | 55.375 | 57.2 | 57.65
15.1 | 44.425 | 50.5333 | 55.1333 | 57.05 | 57.3
15.2 | 44.1 | 50.3 | 54.825 | 56.7 | 57.075
15.3 | 43.85 | 50.0667 | 54.5 | 56.55 | 56.825
15.4 | 43.7 | 49.7 | 54.175 | 56.2 | 56.575
15.5 | 43.2667 | 49.5 | 54 | 55.96 | 56.4125
15.6 | 43 | 49.2 | 53.7333 | 55.7 | 56.175
15.7 | 42.8 | 48.9 | 53.5 | 55.46 | 55.925
15.8 | 42.4667 | 48.6 | 53.2 | 55.3 | 55.6125
15.9 | 42.25 | 48.4 | 52.95 | 54.95 | 55.2375
16 | 41.9 | 48.15 | 52.7333 | 54.725 | 54.9875
16.1 | 41.7 | 47.9 | 52.45 | 54.475 | 54.825
16.2 | 41.5 | 47.55 | 52.075 | 54.225 | 54.575
16.3 | 41.2 | 47.4 | 51.9 | 54.0667 | 54.325
16.4 | 40.975 | 47.1 | 51.7 | 53.675 | 54.2
16.5 | 40.7 | 46.8 | 51.4 | 53.425 | 53.9
16.6 | 40.475 | 46.4667 | 51.15 | 53.25 | 53.55
16.7 | 40.225 | 46.3 | 51 | 53.025 | 53.35
16.8 | 39.975 | 45.975 | 50.75 | 52.675 | 53.1
16.9 | 39.725 | 45.825 | 50.3667 | 52.425 | 52.95
17 | 39.475 | 45.575 | 50.14 | 52.2 | 52.7
17.1 | 39.225 | 45.3 | 49.8 | 52.025 | 52.45
17.2 | 38.975 | 45.075 | 49.7 | 51.78 | 52.2
17.3 | 38.725 | 44.825 | 49.475 | 51.55 | 51.95
17.4 | 38.475 | 44.475 | 49.15 | 51.4 | 51.7
17.5 | 38.225 | 44.225 | 48.875 | 51.05 | 51.45
17.6 | 37.9 | 43.975 | 48.725 | 50.8 | 51.3
17.7 | 37.8 | 43.74 | 48.575 | 50.55 | 51.05
17.8 | 37.6 | 43.5 | 48.325 | 50.35 | 50.8
17.9 | 37.35 | 43.25 | 48.075 | 50 | 50.5
18 | 37 | 43 | 47.825 | 49.85 | 50.4
18.1 | 36.8 | 42.85 | 47.575 | 49.6 | 50.05
18.2 | 36.64 | 42.6 | 47.325 | 49.45 | 49.8
18.3 | 36.5 | 42.3 | 47.1 | 49.2 | 49.65
18.4 | 36.1667 | 42.1 | 46.875 | 49.06 | 49.3
18.5 | 36 | 41.95 | 46.5429 | 48.74 | 49.15
18.6 | 35.8 | 41.6 | 46.3 | 48.475 | 48.9
18.7 | 35.5 | 41.35 | 46.05 | 48.3 | 48.775
18.8 | 35.3 | 41.1 | 45.85 | 48.1 | 48.525
18.9 | 35 | 40.9 | 45.7 | 47.825 | 48.25
19 | 34.8 | 40.6 | 45.45 | 47.525 | 48.025
19.1 | 34.65 | 40.5 | 45.2 | 47.375 | 47.775
19.2 | 34.3 | 40.25 | 45.1 | 47.1 | 47.625
19.3 | 34.1 | 40 | 44.9 | 46.9 | 47.375
19.4 | 34 | 39.8 | 44.6 | 46.825 | 47.125
19.5 | 33.775 | 39.625 | 44.3 | 46.575 | 46.9
19.6 | 33.6 | 39.35 | 44.1 | 46.325 | 46.68
19.7 | 33.35 | 39.125 | 43.95 | 46.075 | 46.5143
19.8 | 33.175 | 38.875 | 43.7 | 45.925 | 46.3
19.9 | 32.94 | 38.725 | 43.46 | 45.675 | 46.125
20 | 32.75 | 38.4333 | 43.275 | 45.425 | 45.9
20.1 | 32.5 | 38.275 | 43.125 | 45.18 | 45.65
20.2 | 32.325 | 38.2 | 42.74 | 45.025 | 45.5333
20.3 | 32.1167 | 37.9 | 42.56 | 44.8 | 45.2
20.4 | 31.9857 | 37.65 | 42.375 | 44.6 | 45.0375
20.5 | 31.7875 | 37.4 | 42.225 | 44.4 | 44.9125
20.6 | 31.625 | 37.225 | 42 | 44.15 | 44.6625
20.7 | 31.375 | 36.9 | 41.8 | 44 | 44.325
20.8 | 31.125 | 36.8 | 41.5 | 43.85 | 44.1375
20.9 | 31 | 36.58 | 41.3333 | 43.6 | 44.0125
21 | 30.825 | 36.325 | 41.15 | 43.45 | 43.875
21.1 | 30.6 | 36.1 | 40.85 | 43.15 | 43.625
21.2 | 30.4 | 36.04 | 40.6333 | 43 | 43.4375
21.3 | 30.25 | 35.7 | 40.4 | 42.7 | 43.225
21.4 | 30.025 | 35.55 | 40.3 | 42.6 | 43
21.5 | 29.875 | 35.3 | 40.1 | 42.35 | 42.8286
21.6 | 29.7 | 35.15 | 39.9 | 42.1 | 42.575
21.7 | 29.5667 | 35 | 39.7 | 41.9 | 42.3625
21.8 | 29.375 | 34.8 | 39.45 | 41.7 | 42.2
21.9 | 29.2 | 34.6 | 39.25 | 41.525 | 41.9667
22 | 28.98 | 34.35 | 39.1 | 41.3 | 41.8
22.1 | 28.8 | 34.2 | 38.9 | 41.125 | 41.6
22.2 | 28.6 | 33.95 | 38.7 | 41 | 41.35
22.3 | 28.45 | 33.8 | 38.5 | 40.825 | 41.225
22.4 | 28.2 | 33.6 | 38.3 | 40.575 | 41
22.5 | 28.1 | 33.5 | 38.15 | 40.425 | 40.85
22.6 | 27.9 | 33.25 | 37.9 | 40.275 | 40.65
22.7 | 27.7333 | 33 | 37.76 | 40.025 | 40.45
22.8 | 27.5667 | 32.875 | 37.525 | 39.775 | 40.2
22.9 | 27.4 | 32.7 | 37.375 | 39.525 | 40.075
23 | 27.24 | 32.5 | 37.2 | 39.4 | 39.9
23.1 | 27.1 | 32.3 | 36.975 | 39.25 | 39.725
23.2 | 26.925 | 32.1 | 36.8 | 39 | 39.6
23.3 | 26.8 | 32 | 36.675 | 38.9 | 39.35
23.4 | 26.6 | 31.8 | 36.425 | 38.775 | 39.1
23.5 | 26.42 | 31.625 | 36.275 | 38.525 | 38.925
23.6 | 26.275 | 31.475 | 36.025 | 38.4 | 38.75
23.7 | 26.1 | 31.3 | 35.875 | 38.15 | 38.5333
23.8 | 25.975 | 31.1 | 35.7 | 37.9 | 38.425
23.9 | 25.825 | 30.875 | 35.575 | 37.8 | 38.2556
24 | 25.675 | 30.725 | 35.4 | 37.6 | 38.0571
24.1 | 25.425 | 30.575 | 35.2 | 37.45 | 37.9
24.2 | 25.3 | 30.425 | 34.95 | 37.2 | 37.725
24.3 | 25.175 | 30.225 | 34.9 | 37.05 | 37.5875
24.4 | 25.05 | 30.1 | 34.6333 | 36.9 | 37.425
24.5 | 24.9 | 29.9667 | 34.5 | 36.8 | 37.275
24.6 | 24.75 | 29.8 | 34.35 | 36.55 | 37
24.7 | 24.5 | 29.55 | 34.1 | 36.4 | 36.9
24.8 | 24.4 | 29.4 | 33.95 | 36.16 | 36.7
24.9 | 24.3 | 29.3 | 33.8 | 36.025 | 36.6
25 | 24.15 | 29.1 | 33.55 | 35.9 | 36.375
25.1 | 24 | 28.95 | 33.4 | 35.8 | 36.22
25.2 | 23.9 | 28.7 | 33.2667 | 35.5 | 36.0429
25.3 | 23.75 | 28.65 | 33 | 35.4 | 35.8
25.4 | 23.6 | 28.42 | 32.9 | 35.2 | 35.65
25.5 | 23.4 | 28.2667 | 32.7667 | 35.025 | 35.5
25.6 | 23.3 | 28.1 | 32.6 | 34.875 | 35.25
25.7 | 23.1 | 27.975 | 32.4667 | 34.7 | 35.05
25.8 | 23 | 27.825 | 32.3 | 34.4667 | 34.925
25.9 | 22.8 | 27.675 | 32.125 | 34.375 | 34.75
26 | 22.75 | 27.5 | 31.9 | 34.2 | 34.65
26.1 | 22.6 | 27.2667 | 31.725 | 34 | 34.4
26.2 | 22.5 | 27.1 | 31.5 | 33.825 | 34.275
26.3 | 22.325 | 26.975 | 31.4 | 33.68 | 34.1111
26.4 | 22.175 | 26.825 | 31.275 | 33.55 | 33.9286
26.5 | 22.025 | 26.7 | 31.1 | 33.4 | 33.8111
26.6 | 21.9 | 26.6 | 30.975 | 33.24 | 33.6857
26.7 | 21.8 | 26.5 | 30.74 | 33.08 | 33.5111
26.8 | 21.575 | 26.3667 | 30.6 | 32.95 | 33.3375
26.9 | 21.5 | 26.2 | 30.5 | 32.8 | 33.2125
27 | 21.4 | 26 | 30.3 | 32.6 | 33.0875
27.1 | 21.225 | 25.8333 | 30.15 | 32.4 | 32.9625
27.2 | 21.175 | 25.65 | 30 | 32.25 | 32.775
27.3 | 21.04 | 25.5 | 29.85 | 32.1 | 32.6125
27.4 | 20.9 | 25.35 | 29.7 | 31.95 | 32.4875
27.5 | 20.8 | 25.25 | 29.6 | 31.8 | 32.325
27.6 | 20.7 | 25.1 | 29.5 | 31.6667 | 32.1
27.7 | 20.55 | 25 | 29.35 | 31.52 | 32.0143
27.8 | 20.4 | 24.9 | 29.2 | 31.375 | 31.78
27.9 | 20.25 | 24.7 | 29 | 31.2 | 31.6143
28 | 20.1 | 24.6 | 28.8 | 31 | 31.5
28.1 | 20 | 24.4667 | 28.7 | 30.85 | 31.3
28.2 | 19.9 | 24.3 | 28.5 | 30.675 | 31.2
28.3 | 19.8 | 24.1667 | 28.4 | 30.525 | 31.1
28.4 | 19.7 | 24 | 28.3 | 30.4 | 30.9
28.5 | 19.6 | 23.9 | 28.1 | 30.3 | 30.725
28.6 | 19.5 | 23.725 | 28 | 30.2 | 30.6
28.7 | 19.3 | 23.675 | 27.8 | 30 | 30.4
28.8 | 19.225 | 23.5 | 27.7 | 29.9 | 30.3
28.9 | 19.1 | 23.375 | 27.575 | 29.775 | 30.2
29 | 18.925 | 23.3 | 27.425 | 29.6 | 30
29.1 | 18.875 | 23.1 | 27.275 | 29.5 | 29.875
29.2 | 18.7 | 23 | 27.125 | 29.2667 | 29.725
29.3 | 18.625 | 22.875 | 27.075 | 29.1 | 29.675
29.4 | 18.5 | 22.725 | 26.84 | 29.05 | 29.5
29.5 | 18.425 | 22.6 | 26.7 | 28.875 | 29.3667
29.6 | 18.375 | 22.5 | 26.7 | 28.8 | 29.175
29.7 | 18.2 | 22.4 | 26.45 | 28.7 | 29.0333
29.8 | 18.075 | 22.2 | 26.3 | 28.5 | 28.9
29.9 | 18 | 22.1 | 26.1 | 28.3333 | 28.8
30 | 17.8667 | 21.9375 | 26 | 28.175 | 28.66

## Wedge 30 Lateral Profiles
Lateral Profiles
Wedge 30
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.2 cm FieldSize:4x4 cm2 | Depth:1.2 cm FieldSize:10x10 cm2 | Depth:1.2 cm FieldSize:20x20 cm2 | Depth:1.2 cm FieldSize:30x30 cm2 | Depth:1.2 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-21.7 | [18 empty] | 24.6216 | 26.4 |  |  |  | 34.2 | 36
-21.6 | [18 empty] | 24.905 | 26.7 |  |  |  | 34.76 | 36.5333
-21.5 | [18 empty] | 25.2551 | 27.1 |  |  |  | 35.1 | 37
-21.4 | [13 empty] | 16.0571 | 17.8429 |  |  |  | 25.5451 | 27.3429 |  |  |  | 35.5 | 37.34
-21.3 | [13 empty] | 16.3333 | 18.04 |  |  |  | 25.9852 | 27.66 |  |  |  | 35.92 | 37.76
-21.2 | [13 empty] | 16.575 | 18.28 |  |  |  | 26.2652 | 28.1333 |  |  |  | 36.4333 | 38.26
-21.1 | [13 empty] | 16.74 | 18.5 |  |  |  | 26.6253 | 28.46 |  |  |  | 36.92 | 38.75
-21 | [13 empty] | 16.975 | 18.72 |  |  |  | 26.961 | 28.8333 |  |  |  | 37.4 | 39.1
-20.9 | [13 empty] | 17.2333 | 18.92 |  |  |  | 27.4855 | 29.22 |  |  |  | 37.9 | 39.8
-20.8 | [13 empty] | 17.46 | 19.2333 |  |  |  | 27.8856 | 29.58 |  |  |  | 38.3 | 40.2
-20.7 | [8 empty] | 13.18 | 14.775 |  |  |  | 17.7 | 19.425 |  |  |  | 28.2056 | 30.02 |  |  |  | 39 | 40.8
-20.6 | [8 empty] | 13.29 | 14.96 |  |  |  | 18.02 | 19.6667 |  |  |  | 28.6724 | 30.38 |  |  |  | 39.6 | 41.4
-20.5 | [8 empty] | 13.5 | 15.15 |  |  |  | 18.25 | 19.88 |  |  |  | 29.0458 | 30.92 |  |  |  | 40.1 | 41.7
-20.4 | [8 empty] | 13.6 | 15.3 |  |  |  | 18.6 | 20.26 |  |  |  | 29.5459 | 31.22 |  |  |  | 40.8 | 42.6
-20.3 | [8 empty] | 13.8167 | 15.5 |  |  |  | 18.8 | 20.42 |  |  |  | 29.926 | 31.68 |  |  |  | 41.7 | 43.3
-20.2 | [8 empty] | 14 | 15.7 |  |  |  | 19.05 | 20.74 |  |  |  | 30.256 | 32.06 |  |  |  | 42.6 | 44.1
-20.1 |  |  |  | 11.48 | 13.1033 |  |  |  | 14.2 | 15.9 |  |  |  | 19.3 | 21 |  |  |  | 30.7861 | 32.48 |  |  |  | 43.9 | 45.5
-20 |  |  |  | 11.64 | 13.3033 |  |  |  | 14.4 | 16.1 |  |  |  | 19.64 | 21.28 |  |  |  | 31.0862 | 32.92 |  |  |  | 45.6 | 47.1
-19.9 |  |  |  | 11.84 | 13.4034 |  |  |  | 14.6 | 16.3 |  |  |  | 19.9667 | 21.62 |  |  |  | 31.7863 | 33.52 |  |  |  | 48.5 | 50
-19.8 |  |  |  | 11.94 | 13.5534 |  |  |  | 14.8 | 16.4 |  |  |  | 20.22 | 21.94 |  |  |  | 32.1864 | 33.96 |  |  |  | 52.9 | 54
-19.7 |  |  |  | 12.06 | 13.7034 |  |  |  | 15 | 16.6 |  |  |  | 20.54 | 22.24 |  |  |  | 32.6265 | 34.36 |  |  |  | 58.9 | 60
-19.6 |  |  |  | 12.2 | 13.8035 |  |  |  | 15.2 | 16.9 |  |  |  | 20.94 | 22.58 |  |  |  | 33.2066 | 34.98 |  |  |  | 66.8 | 67.5
-19.5 |  |  |  | 12.3273 | 13.9635 |  |  |  | 15.5 | 17.1 |  |  |  | 21.22 | 22.88 |  |  |  | 33.6667 | 35.48 |  |  |  | 75.8333 | 76.1
-19.4 |  |  |  | 12.52 | 14.1636 |  |  |  | 15.6 | 17.3 |  |  |  | 21.54 | 23.24 |  |  |  | 34.2469 | 35.98 |  |  |  | 85.8 | 85
-19.3 |  |  |  | 12.66 | 14.3036 |  |  |  | 15.8 | 17.54 |  |  |  | 21.8714 | 23.52 |  |  |  | 34.767 | 36.58 |  |  |  | 94.3333 | 94
-19.2 |  |  |  | 12.82 | 14.4636 |  |  |  | 16.1 | 17.74 |  |  |  | 22.275 | 23.9 |  |  |  | 35.3471 | 37 |  |  |  | 102.5 | 101
-19.1 |  |  |  | 12.975 | 14.6037 |  |  |  | 16.4 | 17.94 |  |  |  | 22.66 | 24.3 |  |  |  | 36.0472 | 37.64 |  |  |  | 108.58 | 107.6
-19 |  |  |  | 13.11 | 14.8037 |  |  |  | 16.6 | 18.24 |  |  |  | 23.08 | 24.7333 |  |  |  | 36.7474 | 38.34 |  |  |  | 113.2333 | 111.7
-18.9 |  |  |  | 13.22 | 14.9037 |  |  |  | 16.8 | 18.44 |  |  |  | 23.46 | 25.1667 |  |  |  | 37.3875 | 39.04 |  |  |  | 115.8 | 113.9875
-18.8 |  |  |  | 13.42 | 15.1038 |  |  |  | 17.0667 | 18.7 |  |  |  | 23.86 | 25.5 |  |  |  | 38.1362 | 39.86 |  |  |  | 117.58 | 115.8
-18.7 |  |  |  | 13.62 | 15.3038 |  |  |  | 17.375 | 19.02 |  |  |  | 24.3 | 25.9667 |  |  |  | 39.2579 | 40.9 |  |  |  | 118.5 | 116.8571
-18.6 |  |  |  | 13.82 | 15.5039 |  |  |  | 17.66 | 19.3 |  |  |  | 24.75 | 26.35 |  |  |  | 40.4881 | 42.14 |  |  |  | 119.08 | 117.74
-18.5 |  |  |  | 14.01 | 15.7039 |  |  |  | 17.92 | 19.6 |  |  |  | 25.16 | 26.8 |  |  |  | 42.8371 | 44.32 |  |  |  | 119.6 | 118.1857
-18.4 |  |  |  | 14.1333 | 15.804 |  |  |  | 18.26 | 19.84 |  |  |  | 25.64 | 27.2 |  |  |  | 46.1492 | 47.38 |  |  |  | 120.3 | 118.4
-18.3 |  |  |  | 14.4 | 16.0183 |  |  |  | 18.52 | 20.1667 |  |  |  | 26.06 | 27.7 |  |  |  | 51.4603 | 52.72 |  |  |  | 120.6833 | 118.9
-18.2 |  |  |  | 14.6 | 16.2441 |  |  |  | 18.84 | 20.44 |  |  |  | 26.48 | 28.26 |  |  |  | 59.3918 | 60.52 |  |  |  | 120.94 | 119.26
-18.1 |  |  |  | 14.7 | 16.4441 |  |  |  | 19.22 | 20.8333 |  |  |  | 27.1 | 28.7333 |  |  |  | 69.4839 | 70 |  |  |  | 121.3 | 119.5286
-18 |  |  |  | 14.975 | 16.6375 |  |  |  | 19.5 | 21.14 |  |  |  | 27.54 | 29.2 |  |  |  | 80.9562 | 80.52 |  |  |  | 121.48 | 119.7429
-17.9 |  |  |  | 15.16 | 16.8375 |  |  |  | 19.82 | 21.48 |  |  |  | 28.1 | 29.7 |  |  |  | 92.7386 | 92.16 |  |  | 18.74 | 121.72 | 119.96
-17.8 |  |  |  | 15.38 | 17.0443 |  |  |  | 20.225 | 21.86 |  |  |  | 28.7 | 30.3333 |  |  |  | 103.9108 | 102.9 |  |  | 19 | 121.86 | 120.24
-17.7 |  |  |  | 15.58 | 17.2443 |  |  |  | 20.625 | 22.2 |  |  |  | 29.25 | 30.95 |  |  |  | 113.3928 | 112.24 |  |  | 19.3 | 122.12 | 120.4455
-17.6 |  |  |  | 15.86 | 17.4844 |  |  |  | 21.025 | 22.66 |  |  |  | 30.02 | 31.4 |  |  |  | 120.4243 | 119.24 |  |  | 19.56 | 122.4 | 120.5778
-17.5 |  |  |  | 16.08 | 17.7544 |  |  |  | 21.44 | 23.06 |  |  |  | 30.6 | 32.2 |  |  |  | 125.4254 | 123.9 |  |  | 19.78 | 122.56 | 120.7455
-17.4 |  |  |  | 16.28 | 17.9902 |  |  |  | 21.8333 | 23.44 |  |  |  | 31.5714 | 32.9 |  |  |  | 128.1442 | 126.5 |  |  | 20.08 | 122.86 | 120.94
-17.3 |  |  |  | 16.58 | 18.2546 |  |  |  | 22.25 | 23.86 |  |  |  | 32.45 | 33.8 |  |  |  | 129.526 | 127.86 |  |  | 20.18 | 123.06 | 121.28
-17.2 |  |  |  | 16.78 | 18.5046 |  |  |  | 22.725 | 24.36 |  |  |  | 33.6 | 35 |  |  |  | 130.4593 | 128.74 |  |  | 20.48 | 123.2 | 121.47
-17.1 |  |  |  | 17.06 | 18.738 |  |  |  | 23.24 | 24.8 |  |  |  | 35.3 | 36.5 |  |  |  | 131.126 | 129.38 |  |  | 20.78 | 123.36 | 121.5
-17 |  |  |  | 17.38 | 19.0048 |  |  |  | 23.66 | 25.34 |  |  |  | 37.7375 | 38.9 |  |  |  | 131.626 | 129.68 |  |  | 21 | 123.5 | 121.57
-16.9 |  |  |  | 17.58 | 19.3248 |  |  |  | 24.2857 | 25.85 |  |  |  | 41.8 | 42.8 |  |  | 15.1 | 131.946 | 130.14 |  |  | 21.2714 | 123.5 | 121.64
-16.8 |  |  |  | 17.88 | 19.5449 |  |  |  | 24.86 | 26.44 |  |  |  | 48.92 | 49.7 |  |  | 15.3333 | 132.3163 | 130.4333 |  |  | 21.48 | 123.6 | 121.7875
-16.7 |  |  |  | 18.2 | 19.925 |  |  |  | 25.5375 | 27.13 |  |  |  | 59.8 | 60.8 |  |  | 15.6 | 132.587 | 130.62 |  |  | 21.7 | 123.7 | 121.84
-16.6 |  |  |  | 18.46 | 20.145 |  |  |  | 26.32 | 27.92 |  |  |  | 74.4 | 74.7 |  |  | 15.84 | 132.787 | 130.825 |  |  | 22.1 | 123.8 | 122
-16.5 |  |  |  | 18.8 | 20.4451 |  |  |  | 27.28 | 28.9 |  |  |  | 89.9 | 88.7 |  |  | 16 | 132.987 | 131.0333 |  |  | 22.4 | 123.925 | 122.02
-16.4 |  |  |  | 19.12 | 20.7452 |  |  |  | 28.64 | 30.17 |  |  |  | 104.5 | 103.6 |  |  | 16.3 | 133.202 | 131.28 |  |  | 22.6 | 124 | 122.12
-16.3 |  |  |  | 19.42 | 21.1453 |  |  |  | 30.66 | 32.28 |  |  |  | 118.8 | 116.6909 |  |  | 16.5 | 133.3984 | 131.42 |  |  | 23 | 124.1 | 122.2
-16.2 |  |  |  | 19.82 | 21.5454 |  |  |  | 34 | 35.74 |  |  |  | 131.1 | 129.5 |  |  | 16.7 | 133.427 | 131.525 |  |  | 23.2667 | 124.2 | 122.22
-16.1 |  |  |  | 20.32 | 22.0655 |  |  |  | 40.6 | 41.77 |  |  |  | 139.9 | 137.8286 |  |  | 16.9 | 133.507 | 131.6333 |  |  | 23.5667 | 124.2 | 122.3
-16 |  |  |  | 20.98 | 22.6657 |  |  |  | 50.9 | 52.2667 |  |  |  | 144.8 | 142.6 |  |  | 17.12 | 133.607 | 131.7 |  |  | 23.88 | 124.1 | 122.34
-15.9 |  |  |  | 21.8 | 23.5059 |  |  |  | 66.275 | 66.9667 |  |  | 10.9022 | 147.5 | 145.175 |  |  | 17.4 | 133.627 | 131.725 |  |  | 24.16 | 124.1 | 122.5182
-15.8 |  |  |  | 23.14 | 24.7662 |  |  |  | 84.9 | 84.74 |  |  | 11.0736 | 148.7 | 146.42 |  |  | 17.5833 | 133.627 | 131.8 |  |  | 24.46 | 124.2 | 122.6
-15.7 |  |  |  | 25.56 | 27.0268 |  |  |  | 103.4 | 102.2 |  |  | 11.2623 | 149.4 | 147.16 |  |  | 17.96 | 133.627 | 131.9 |  |  | 24.78 | 124.3 | 122.6
-15.6 |  |  |  | 29.425 | 31.4745 |  |  |  | 121.5 | 119.2 |  |  | 11.4623 | 150 | 147.64 |  |  | 18.1667 | 133.702 | 131.9 |  |  | 25.1 | 124.3 | 122.6
-15.5 |  |  |  | 38.4 | 40.07 |  |  |  | 136.9667 | 135.36 |  |  | 11.6023 | 150.1 | 147.84 |  |  | 18.44 | 133.727 | 131.92 |  |  | 25.54 | 124.3 | 122.64
-15.4 |  |  |  | 52.3 | 54.5936 |  |  | 8.8714 | 149.1333 | 146.82 |  |  | 11.7273 | 150.3 | 148 |  |  | 18.75 | 133.667 | 131.98 |  |  | 25.96 | 124.3 | 122.6
-15.3 |  |  |  | 70.3 | 72.4181 |  |  | 8.95 | 156.4 | 153.5333 |  |  | 11.9224 | 150.4 | 148.14 |  |  | 19 | 133.627 | 131.9 |  |  | 26.2833 | 124.3 | 122.6
-15.2 |  |  |  | 91.7 | 92.7832 |  |  | 9.06 | 159.6 | 156.92 |  |  | 12.0224 | 150.4 | 148.2 |  |  | 19.32 | 133.627 | 131.8 |  |  | 26.6 | 124.38 | 122.6
-15.1 |  |  | 7.7013 | 111.8 | 112.5284 |  |  | 9.2 | 160.9667 | 158.5 |  |  | 12.2224 | 150.4 | 148.2 |  |  | 19.62 | 133.567 | 131.78 |  |  | 26.8 | 124.4667 | 122.7
-15 |  |  | 7.8013 | 132.44 | 131.9827 |  |  | 9.2833 | 161.6 | 159.25 |  |  | 12.4025 | 150.3 | 148.2 |  |  | 19.92 | 133.527 | 131.7 |  |  | 27.4 | 124.5 | 122.7
-14.9 |  |  | 7.8613 | 149.62 | 148.4172 |  |  | 9.4 | 162.12 | 159.56 |  |  | 12.6225 | 150.2 | 148.1 |  |  | 20.22 | 133.427 | 131.7 |  |  | 27.7 | 124.54 | 122.8
-14.8 |  |  | 8.0013 | 162.18 | 160.3401 |  |  | 9.5 | 162.2 | 159.6 |  |  | 12.8426 | 150.1 | 147.96 |  |  | 20.475 | 133.427 | 131.7 |  |  | 28.15 | 124.5 | 122.825
-14.7 |  |  | 8.0613 | 168.82 | 165.9414 |  |  | 9.7 | 162.2 | 159.6 |  |  | 13.0026 | 150 | 147.8 |  |  | 20.8 | 133.327 | 131.7 |  |  | 28.48 | 124.34 | 122.9
-14.6 |  |  | 8.1814 | 171.24 | 168.402 |  |  | 9.8 | 162.1 | 159.6 |  |  | 13.2226 | 149.8 | 147.6 |  |  | 21.1333 | 133.327 | 131.5 |  |  | 28.86 | 124.24 | 122.8
-14.5 |  |  | 8.2014 | 171.8 | 168.982 |  |  | 10 | 161.9 | 159.5 |  |  | 13.5227 | 149.6 | 147.5 |  |  | 21.5333 | 133.307 | 131.5 |  |  | 29.26 | 124.1 | 122.8
-14.4 |  |  | 8.3614 | 171.8 | 169.042 |  |  | 10.1 | 161.66 | 159.22 |  |  | 13.7227 | 149.4 | 147.3 |  |  | 21.84 | 133.227 | 131.4 |  |  | 29.68 | 124.1 | 122.7
-14.3 |  |  | 8.4014 | 171.72 | 168.9563 |  |  | 10.2333 | 161.4 | 158.86 |  |  | 13.9361 | 149.2 | 147.16 |  |  | 22.2333 | 133.227 | 131.3 |  |  | 30.26 | 124.1 | 122.7
-14.2 |  |  | 8.5014 | 171.42 | 168.762 |  |  | 10.4 | 161.1 | 158.35 |  |  | 14.1528 | 149 | 146.96 |  |  | 22.64 | 133.107 | 131.2 |  |  | 30.74 | 124.1 | 122.6
-14.1 |  |  | 8.6014 | 170.82 | 168.382 |  |  | 10.5 | 160.5 | 158.02 |  |  | 14.3629 | 148.8 | 146.82 |  |  | 22.9 | 133.027 | 131.2 |  |  | 31.24 | 124 | 122.5
-14 |  |  | 8.7014 | 170.22 | 167.802 |  |  | 10.64 | 160.0667 | 157.7667 |  |  | 14.6629 | 148.5 | 146.46 |  |  | 23.3 | 132.827 | 131.2 |  |  | 31.82 | 124 | 122.4
-13.9 |  |  | 8.8215 | 169.82 | 167.402 |  |  | 10.82 | 159.7667 | 157.32 |  |  | 14.863 | 148.2 | 146.1667 |  |  | 23.6 | 132.707 | 131.1 |  |  | 32.44 | 123.94 | 122.3
-13.8 |  |  | 8.9215 | 169.3 | 166.542 |  |  | 11.02 | 159.25 | 156.94 |  |  | 15.153 | 147.9556 | 145.98 |  |  | 24.2 | 132.607 | 131 |  |  | 33.0667 | 123.85 | 122.3
-13.7 |  |  | 9.1015 | 168.175 | 165.9012 |  |  | 11.2 | 158.9 | 156.46 |  |  | 15.4531 | 147.7 | 145.76 |  |  | 24.6 | 132.426 | 130.8 |  |  | 34.08 | 123.7 | 122.2
-13.6 |  |  | 9.2215 | 167.7333 | 165.301 |  |  | 11.4 | 158.3 | 156.04 |  |  | 15.7431 | 147.3667 | 145.38 |  |  | 25 | 132.326 | 130.7 |  |  | 35.26 | 123.6 | 122.1
-13.5 |  |  | 9.3216 | 166.9333 | 164.501 |  |  | 11.575 | 157.9 | 155.54 |  |  | 16.0632 | 147.12 | 145.28 |  |  | 25.4 | 132.176 | 130.6 |  |  | 37.075 | 123.6 | 122.1
-13.4 |  |  | 9.4016 | 166.2 | 163.901 |  |  | 11.7 | 157.5 | 155.22 |  |  | 16.4033 | 146.7667 | 144.96 |  |  | 25.8 | 132.026 | 130.6 |  |  | 39.9 | 123.46 | 122
-13.3 |  |  | 9.5616 | 165.46 | 163.221 |  |  | 12 | 156.9 | 154.74 |  |  | 16.7033 | 146.5429 | 144.58 |  |  | 26.3 | 131.8927 | 130.5 |  |  | 43.78 | 123.36 | 121.9
-13.2 |  |  | 9.6616 | 164.94 | 162.481 |  |  | 12.175 | 156.6 | 154.22 |  |  | 17.0034 | 146.3091 | 144.28 |  |  | 26.8 | 131.826 | 130.3 |  |  | 49.62 | 123.3 | 121.9
-13.1 |  |  | 9.8391 | 164.3 | 161.7202 |  |  | 12.375 | 156 | 153.84 |  |  | 17.3635 | 145.975 | 144.08 |  |  | 27.4 | 131.726 | 130.0429 |  |  | 57.72 | 123.2 | 121.76
-13 |  |  | 10.0017 | 163.5 | 161.1 |  |  | 12.675 | 155.38 | 153.2571 |  |  | 17.6702 | 145.62 | 143.74 |  |  | 27.9 | 131.501 | 129.8167 |  |  | 66.6 | 123.1 | 121.7
-12.9 |  |  | 10.1017 | 162.64 | 160.52 |  |  | 12.8 | 154.92 | 152.6857 |  |  | 18.1036 | 145.12 | 143.2667 |  |  | 28.7 | 131.301 | 129.6429 |  |  | 75.76 | 123.1 | 121.65
-12.8 |  |  | 10.2017 | 162.0333 | 159.96 |  |  | 13.1667 | 154.4 | 152.32 |  |  | 18.4637 | 144.76 | 143.025 |  |  | 29.4 | 131.206 | 129.5125 |  |  | 84.34 | 123.08 | 121.55
-12.7 |  |  | 10.4518 | 161.12 | 159.16 |  |  | 13.35 | 153.9 | 151.7 |  |  | 18.8638 | 144.52 | 142.78 |  |  | 30.34 | 131.0593 | 129.375 |  |  | 91.92 | 123 | 121.3
-12.6 |  |  | 10.6018 | 160.54 | 158.44 |  |  | 13.65 | 153.46 | 151.26 |  |  | 19.1838 | 144.2 | 142.36 |  |  | 31.58 | 130.806 | 129.22 |  |  | 98.04 | 122.7375 | 121.26
-12.5 |  |  | 10.8018 | 159.74 | 157.8392 |  |  | 13.94 | 152.96 | 150.86 |  |  | 19.6639 | 143.925 | 142.06 |  |  | 33.48 | 130.606 | 128.92 |  |  | 101.96 | 122.5 | 121.12
-12.4 |  |  | 10.9685 | 159.14 | 156.989 |  |  | 14.24 | 152.42 | 150.22 |  |  | 20.064 | 143.55 | 141.68 |  |  | 36.48 | 130.406 | 128.7 |  |  | 104.3 | 122.36 | 121
-12.3 |  |  | 11.1619 | 158.56 | 156.519 |  |  | 14.54 | 151.85 | 149.725 |  |  | 20.6441 | 143.2 | 141.3 |  |  | 41.5 | 130.306 | 128.5667 |  |  | 105.86 | 122.3 | 120.9
-12.2 |  |  | 11.3019 | 158.02 | 155.839 |  |  | 14.84 | 151.36 | 149.35 |  |  | 21.0642 | 142.9333 | 140.9545 |  |  | 48.5 | 130.106 | 128.4 |  |  | 106.74 | 122.2143 | 120.8
-12.1 |  |  | 11.5019 | 157.3111 | 155.019 |  |  | 15.18 | 150.84 | 148.92 |  |  | 21.5643 | 142.5 | 140.61 |  |  | 58 | 129.926 | 128.3 |  |  | 107.38 | 122.2 | 120.66
-12 |  |  | 11.7306 | 156.4667 | 154.3723 |  |  | 15.5 | 150.24 | 148.22 |  |  | 22.1444 | 142.1 | 140.2875 |  |  | 68.42 | 129.626 | 128.125 |  |  | 107.92 | 122.11 | 120.5
-11.9 |  |  | 11.902 | 156 | 153.858 |  |  | 15.88 | 149.66 | 147.78 |  |  | 22.8246 | 141.72 | 139.9571 |  |  | 79.08 | 129.326 | 127.9 |  |  | 108.35 | 121.92 | 120.3857
-11.8 |  |  | 12.1163 | 155.1 | 153.088 |  |  | 16.25 | 149.25 | 147.2 |  |  | 23.6047 | 141.36 | 139.61 |  |  | 89.6 | 129.126 | 127.8 |  |  | 108.72 | 121.7364 | 120.22
-11.7 |  |  | 12.3592 | 154.6 | 152.738 |  |  | 16.7 | 148.64 | 146.62 |  |  | 24.5049 | 141.0111 | 139.31 |  |  | 97.5 | 128.826 | 127.6333 |  |  | 108.94 | 121.4333 | 120.2
-11.6 |  |  | 12.5854 | 153.9 | 151.938 |  |  | 17.2 | 148.15 | 146.3 |  |  | 25.8452 | 140.5875 | 139.01 |  |  | 104.35 | 128.526 | 127.3 |  |  | 109.24 | 121.31 | 120.2
-11.5 |  |  | 12.7821 | 153.25 | 151.398 |  |  | 17.64 | 147.6333 | 145.8333 |  |  | 27.8656 | 140.1857 | 138.6364 |  |  | 108.74 | 128.401 | 127.16 |  | 10.5 | 109.34 | 121.2 | 120.06
-11.4 |  |  | 13.0622 | 152.65 | 150.788 |  |  | 18.22 | 147.05 | 145.18 |  |  | 31.3863 | 139.78 | 138.1333 |  |  | 111.14 | 128.186 | 126.96 |  | 10.75 | 109.55 | 121.17 | 120
-11.3 |  |  | 13.3622 | 152 | 150.138 |  |  | 18.92 | 146.55 | 144.52 |  |  | 37.4075 | 139.32 | 137.75 |  |  | 112.6 | 127.966 | 126.66 |  | 11 | 109.7 | 121.07 | 119.7429
-11.2 |  |  | 13.6623 | 151.3 | 149.437 |  |  | 19.8 | 146.05 | 144.16 |  |  | 46.5893 | 139.02 | 137.3 |  |  | 113.34 | 127.826 | 126.46 |  | 11.1 | 109.86 | 120.97 | 119.61
-11.1 |  |  | 13.9623 | 150.8 | 148.837 |  |  | 21.08 | 145.475 | 143.74 |  |  | 58.2117 | 138.58 | 136.92 |  |  | 113.84 | 127.5456 | 126.22 |  | 11.2 | 110.2 | 120.84 | 119.51
-11 |  |  | 14.4167 | 150.1 | 148.237 |  |  | 23.2333 | 145.05 | 143.14 |  |  | 71.0714 | 138.08 | 136.56 |  | 7.9 | 114.24 | 127.225 | 125.975 |  | 11.3667 | 110.5 | 120.64 | 119.3364
-10.9 |  |  | 14.9525 | 149.3667 | 147.637 |  |  | 27.0333 | 144.54 | 142.56 |  |  | 84.8169 | 137.76 | 136.22 |  | 8.05 | 114.64 | 126.965 | 125.76 |  | 11.6 | 110.5 | 120.44 | 119.22
-10.8 |  |  | 15.7026 | 148.9 | 147.137 |  |  | 33.8333 | 143.88 | 142.16 |  |  | 97.3394 | 137.38 | 135.86 |  | 8.14 | 115.02 | 126.745 | 125.48 |  | 11.75 | 110.7 | 120.24 | 119.2
-10.7 |  |  | 17.1029 | 148.225 | 146.4037 |  |  | 44.7333 | 143.18 | 141.66 |  |  | 107.1214 | 136.95 | 135.62 |  | 8.24 | 115.2 | 126.545 | 125.28 |  | 11.9 | 110.7 | 120.0727 | 119.12
-10.6 |  |  | 19.2032 | 147.475 | 145.936 |  |  | 58.0333 | 142.74 | 141.24 |  |  | 115.1733 | 136.52 | 135.1 |  | 8.46 | 115.32 | 126.365 | 124.98 |  | 12 | 110.8 | 119.9778 | 119.02
-10.5 |  |  | 23.704 | 146.72 | 145.2693 |  |  | 73.2667 | 142.34 | 140.74 |  | 5.1 | 119.5811 | 136.2 | 134.46 |  | 8.6 | 115.42 | 126.045 | 124.78 |  | 12.2 | 110.9 | 119.84 | 118.8556
-10.4 |  |  | 32.2054 | 146.24 | 144.616 |  |  | 89.4 | 141.66 | 139.92 |  | 5.2 | 121.784 | 135.45 | 134.06 |  | 8.8 | 115.52 | 125.845 | 124.6 |  | 12.4 | 111.025 | 119.6727 | 118.64
-10.3 |  |  | 45.1075 | 145.6 | 143.956 |  |  | 103.3 | 141.16 | 139.38 |  | 5.2 | 123.0249 | 135.22 | 133.68 |  | 8.94 | 115.7111 | 125.765 | 124.48 |  | 12.6 | 111.1 | 119.56 | 118.41
-10.2 |  |  | 61.2102 | 144.88 | 143.236 |  | 3.9 | 115.1 | 140.74 | 139.08 |  | 5.4 | 123.845 | 134.96 | 133.36 |  | 9.1 | 115.82 | 125.405 | 124.16 |  | 12.8333 | 111.1 | 119.36 | 118.15
-10.1 |  |  | 78.113 | 144.2 | 142.636 |  | 3.9 | 123.2 | 140.3 | 138.4333 |  | 5.4 | 124.35 | 134.58 | 133.075 |  | 9.2 | 115.92 | 125.065 | 123.8667 |  | 13.1 | 111.2 | 119.16 | 118.0714
-10 |  | 3.35 | 95.5159 | 143.72 | 142.136 |  | 4 | 127.92 | 139.52 | 137.85 |  | 5.5 | 124.505 | 134.08 | 132.55 |  | 9.4 | 116 | 124.725 | 123.7 |  | 13.2333 | 111.14 | 118.98 | 117.875
-9.9 |  | 3.4 | 111.219 | 142.95 | 141.435 |  | 4.1 | 130.06 | 139.02 | 137.4333 |  | 5.6 | 124.665 | 133.48 | 132.175 |  | 9.54 | 116 | 124.465 | 123.4 |  | 13.4 | 111.1 | 118.86 | 117.73
-9.8 |  | 3.4 | 124.221 | 142.45 | 140.7714 |  | 4.1 | 131.04 | 138.42 | 136.95 |  | 5.7 | 124.825 | 133.1 | 131.98 |  | 9.725 | 116.02 | 124.265 | 123.0667 |  | 13.7 | 111.15 | 118.68 | 117.6125
-9.7 |  | 3.4 | 131.9931 | 141.74 | 140.235 |  | 4.16 | 131.56 | 137.8125 | 136.54 |  | 5.8 | 124.825 | 132.7 | 131.46 |  | 9.94 | 116.12 | 124.065 | 122.875 |  | 14 | 111.15 | 118.56 | 117.3857
-9.6 |  | 3.5 | 136.023 | 141.04 | 139.6572 |  | 4.2 | 131.76 | 137.4 | 135.8 |  | 5.9 | 124.895 | 132.4 | 131.06 |  | 10.1 | 116.2 | 123.805 | 122.64 |  | 14.1 | 111.2 | 118.3636 | 117.23
-9.5 |  | 3.5 | 137.223 | 140.42 | 138.9259 |  | 4.3 | 131.8 | 136.9 | 135.4 |  | 6 | 124.925 | 131.9 | 130.64 |  | 10.225 | 116.2 | 123.465 | 122.4 |  | 14.4 | 111.2 | 118.18 | 117.1125
-9.4 |  | 3.5 | 137.623 | 139.78 | 138.4948 |  | 4.3333 | 131.84 | 136.3 | 134.86 |  | 6.1 | 124.855 | 131.46 | 130.0571 |  | 10.425 | 116.2 | 123.1917 | 122.3 |  | 14.6 | 111.3 | 117.98 | 116.9667
-9.3 |  | 3.6 | 137.623 | 139.225 | 137.814 |  | 4.4 | 131.74 | 135.7429 | 134.42 |  | 6.3 | 124.755 | 131.18 | 129.64 |  | 10.6 | 116.2 | 123.025 | 122.1 |  | 14.9 | 111.3 | 117.78 | 116.775
-9.2 |  | 3.6 | 137.523 | 138.48 | 137.214 |  | 4.5 | 131.54 | 135.4333 | 133.86 |  | 6.4 | 124.655 | 130.6429 | 129.51 |  | 10.925 | 116.2 | 122.8917 | 121.7 |  | 15.2 | 111.3 | 117.5875 | 116.6
-9.1 |  | 3.7 | 137.323 | 138.04 | 136.614 |  | 4.6 | 131.34 | 134.9 | 133.24 |  | 6.5 | 124.555 | 130.3 | 129.32 |  | 11.025 | 116.1 | 122.545 | 121.4 |  | 15.5 | 111.3 | 117.4571 | 116.5
-9 |  | 3.8 | 136.933 | 137.46 | 136.054 |  | 4.6 | 131.1 | 134.44 | 132.74 |  | 6.675 | 124.455 | 130.0167 | 128.76 |  | 11.225 | 116.1 | 122.074 | 121.2 |  | 15.7 | 111.3 | 117.34 | 116.4
-8.9 |  | 3.8 | 136.663 | 137.02 | 135.514 |  | 4.7 | 130.92 | 133.9 | 132.26 |  | 6.8 | 124.355 | 129.3 | 128.07 |  | 11.55 | 116.1 | 121.849 | 120.925 |  | 15.9 | 111.3 | 117.18 | 116.26
-8.8 |  | 3.8667 | 136.223 | 136.16 | 134.894 |  | 4.8 | 130.65 | 133.15 | 131.78 |  | 6.975 | 124.185 | 128.9 | 127.82 |  | 11.8 | 116 | 121.564 | 120.7 |  | 16.3 | 111.3 | 116.98 | 115.98
-8.7 |  | 3.9 | 135.883 | 135.62 | 134.354 |  | 4.9 | 130.45 | 132.62 | 131.38 |  | 7.1 | 123.925 | 128.8167 | 127.62 |  | 11.9 | 115.9 | 121.324 | 120.55 |  | 16.575 | 111.3 | 116.72 | 115.8
-8.6 |  | 3.9 | 135.3894 | 134.9333 | 133.8336 |  | 5 | 130.2 | 132.2 | 130.92 |  | 7.3 | 123.825 | 128.36 | 127.1091 |  | 12.3 | 115.9 | 121.124 | 120.18 |  | 16.85 | 111.16 | 116.525 | 115.68
-8.5 |  | 4 | 134.772 | 134.44 | 133.193 |  | 5.1 | 129.9 | 131.6 | 130.5 |  | 7.4 | 123.625 | 127.84 | 126.76 |  | 12.5 | 115.72 | 120.924 | 119.98 |  | 17.1 | 111 | 116.34 | 115.6
-8.4 |  | 4.1 | 134.4506 | 133.88 | 132.773 |  | 5.2 | 129.5 | 131.25 | 129.88 |  | 7.6 | 123.525 | 127.44 | 126.3 |  | 12.7 | 115.62 | 120.624 | 119.7667 |  | 17.45 | 111 | 116.08 | 115.48
-8.3 |  | 4.1 | 134.102 | 133.375 | 132.173 |  | 5.3 | 129.1667 | 130.65 | 129.48 |  | 7.775 | 123.325 | 127.04 | 125.96 |  | 13 | 115.42 | 120.324 | 119.46 |  | 17.8 | 110.9 | 115.8875 | 115.18
-8.2 |  | 4.2 | 133.4553 | 132.9333 | 131.653 |  | 5.4 | 128.88 | 130.05 | 128.98 |  | 7.975 | 123.1417 | 126.54 | 125.66 |  | 13.3 | 115.3 | 120.124 | 119.16 |  | 18.15 | 110.9 | 115.8 | 114.88
-8.1 |  | 4.3 | 132.682 | 132.3 | 131.1997 |  | 5.55 | 128.42 | 129.65 | 128.55 |  | 8.2 | 122.845 | 126.04 | 125.25 |  | 13.7 | 115.22 | 119.824 | 118.88 |  | 18.55 | 110.88 | 115.54 | 114.76
-8 |  | 4.3 | 132.402 | 131.825 | 130.613 |  | 5.7 | 128.0333 | 129.3333 | 128.12 |  | 8.3667 | 122.4408 | 125.72 | 124.75 |  | 14 | 115.1 | 119.424 | 118.8 |  | 18.9 | 110.8 | 115.28 | 114.4833
-7.9 |  | 4.4 | 131.942 | 131.2 | 130.2184 |  | 5.8 | 127.7333 | 128.8333 | 127.62 |  | 8.65 | 122.244 | 125.34 | 124.31 |  | 14.3 | 115 | 119.124 | 118.38 |  | 19.2 | 110.7 | 115.04 | 114.4
-7.8 |  | 4.5 | 131.382 | 130.6 | 129.682 |  | 6 | 127.25 | 128.2 | 127.28 |  | 8.875 | 122.044 | 124.9667 | 124.01 |  | 14.6 | 114.82 | 118.924 | 118.1 |  | 19.7 | 110.6 | 114.84 | 114.1
-7.7 |  | 4.6 | 130.982 | 130.1 | 129.132 |  | 6.1 | 126.9333 | 127.8 | 126.74 |  | 9.0667 | 121.874 | 124.56 | 123.68 |  | 15 | 114.7 | 118.724 | 117.9 |  | 20.1 | 110.525 | 114.6 | 113.9
-7.6 |  | 4.7 | 130.372 | 129.6 | 128.582 |  | 6.3 | 126.52 | 127.4 | 126.12 |  | 9.4 | 121.5526 | 124.18 | 123.2636 |  | 15.25 | 114.6 | 118.324 | 117.7 | 3.9 | 20.6 | 110.5 | 114.46 | 113.8
-7.5 |  | 4.8 | 129.9553 | 128.9 | 127.932 |  | 6.4 | 126.0333 | 126.88 | 125.58 |  | 9.575 | 121.214 | 123.68 | 122.8 |  | 15.55 | 114.44 | 117.984 | 117.4 | 4.3 | 21 | 110.4 | 114.36 | 113.5
-7.4 |  | 4.9 | 129.422 | 128.5 | 127.332 |  | 6.7 | 125.58 | 126.48 | 125.26 |  | 9.875 | 120.984 | 123.32 | 122.3333 | 2.9007 | 16 | 114.24 | 117.784 | 117.1 | 4.4 | 21.5 | 110.4 | 114.06 | 113.4
-7.3 |  | 5 | 128.821 | 127.98 | 126.992 |  | 6.8 | 125.28 | 125.74 | 124.82 |  | 10.0667 | 120.749 | 122.88 | 122.05 | 3.0008 | 16.45 | 114.04 | 117.584 | 116.7 | 4.5 | 22.3 | 110.275 | 113.8667 | 113.3
-7.2 |  | 5.1 | 128.321 | 127.36 | 126.472 |  | 7 | 124.94 | 125.36 | 124.38 | 1.7 | 10.5667 | 120.3513 | 122.52 | 121.72 | 3.0008 | 16.85 | 113.9 | 117.383 | 116.3 | 4.5 | 22.9 | 110.2 | 113.5667 | 113
-7.1 |  | 5.2 | 127.821 | 126.94 | 126.0457 | 1.3 | 7.3 | 124.32 | 124.92 | 123.96 | 1.8 | 10.8 | 119.994 | 121.98 | 121.28 | 3.1008 | 17.3 | 113.8 | 117.043 | 116.1 | 4.5 | 24.2 | 110.2 | 113.36 | 112.8
-7 | 1 | 5.2667 | 127.241 | 126.44 | 125.531 | 1.2 | 7.5 | 123.78 | 124.55 | 123.4333 | 1.8 | 11.1667 | 119.704 | 121.56 | 120.92 | 3.2008 | 17.7 | 113.66 | 116.783 | 115.9167 | 4.5 | 26.18 | 110.1 | 113.2 | 112.7
-6.9 | 1 | 5.5 | 126.681 | 125.84 | 124.911 | 1.3 | 7.7 | 123.48 | 124.1667 | 122.9333 | 1.8 | 11.5 | 119.484 | 121.18 | 120.58 | 3.3008 | 18.36 | 113.5 | 116.5897 | 115.5 | 4.6 | 28.7 | 110 | 113.16 | 112.5167
-6.8 | 1 | 5.6 | 126.301 | 125.28 | 124.391 | 1.3 | 7.9 | 123.36 | 123.46 | 122.4333 | 1.9 | 11.8 | 119.094 | 120.88 | 120.08 | 3.3008 | 19 | 113.25 | 116.343 | 115.2667 | 4.7 | 32.7 | 109.9 | 113 | 112.26
-6.7 | 1 | 5.8 | 125.671 | 124.825 | 123.911 | 1.3 | 8.2 | 122.82 | 122.78 | 121.95 | 1.9 | 12.2 | 118.784 | 120.48 | 119.52 | 3.4009 | 19.9 | 113.1 | 115.973 | 115.1 | 4.8 | 38.9 | 109.8 | 112.8333 | 112.16
-6.6 | 1 | 5.94 | 125.221 | 124.18 | 123.431 | 1.3 | 8.5 | 122.2667 | 122.38 | 121.5333 | 1.925 | 12.6 | 118.394 | 120 | 119.26 | 3.5009 | 21 | 113 | 115.673 | 115 | 4.9 | 47.4 | 109.7 | 112.56 | 111.96
-6.5 | 1 | 6.025 | 124.681 | 123.68 | 122.781 | 1.3 | 8.775 | 121.9 | 122.06 | 120.9667 | 2 | 13.04 | 118.084 | 119.58 | 118.96 | 3.5009 | 22.8 | 112.8167 | 115.403 | 114.76 | 5 | 55.9 | 109.6 | 112.48 | 111.76
-6.4 | 1.1 | 6.24 | 124.071 | 123.24 | 122.331 | 1.4 | 9.0667 | 121.55 | 121.56 | 120.5 | 2 | 13.625 | 117.7058 | 119.2 | 118.62 | 3.6009 | 25.8 | 112.7 | 115.223 | 114.46 | 5 | 66.1 | 109.5 | 112.38 | 111.6
-6.3 | 1.1 | 6.425 | 123.661 | 122.9 | 121.87 | 1.4 | 9.4333 | 121.05 | 120.98 | 120.2 | 2.1 | 14.24 | 117.524 | 118.94 | 118.2 | 3.6509 | 30.2 | 112.7 | 115.003 | 114.2 | 5.1 | 74.85 | 109.5 | 112.18 | 111.56
-6.2 | 1.1 | 6.6 | 123.0313 | 122.2714 | 121.43 | 1.4 | 9.8 | 120.58 | 120.64 | 119.85 | 2.14 | 15.025 | 117.223 | 118.65 | 117.75 | 3.7009 | 37.4 | 112.52 | 114.7897 | 113.98 | 5.3 | 83.5 | 109.4 | 112.06 | 111.3667
-6.1 | 1.1 | 6.8 | 122.44 | 121.7 | 121.0633 | 1.5 | 10.35 | 120.02 | 120.16 | 119.3333 | 2.2 | 16.26 | 116.823 | 118.3 | 117.46 | 3.8009 | 45.9 | 112.4 | 114.5897 | 113.8 | 5.4 | 89.6 | 109.3 | 111.9 | 111.2
-6 | 1.1 | 7.1 | 121.92 | 121.1857 | 120.53 | 1.5 | 11.1 | 119.76 | 119.86 | 118.8667 | 2.225 | 17.84 | 116.573 | 117.9 | 117.16 | 3.901 | 55.9 | 112.22 | 114.2897 | 113.7 | 5.5 | 94.8 | 109.2 | 111.8 | 111.1
-5.9 | 1.1 | 7.45 | 121.62 | 120.6111 | 119.99 | 1.5 | 11.9 | 119.42 | 119.5 | 118.5 | 2.3 | 20.82 | 116.323 | 117.5 | 116.88 | 4.001 | 66.6 | 112.1 | 114.083 | 113.4833 | 5.65 | 97.5 | 109.1 | 111.6 | 110.98
-5.8 | 1.1 | 7.9 | 121.12 | 120.2 | 119.43 | 1.5 | 13.35 | 119.02 | 119 | 118.1667 | 2.3333 | 26.06 | 116.073 | 117.36 | 116.56 | 4.101 | 77.1 | 111.92 | 113.823 | 113.2 | 5.7 | 99 | 108.9 | 111.2 | 110.88
-5.7 | 1.1333 | 8.6 | 120.62 | 119.82 | 119.005 | 1.5 | 15.8 | 118.66 | 118.5667 | 117.7333 | 2.4 | 33.84 | 115.823 | 117 | 116.16 | 4.2011 | 85.9 | 111.8 | 113.623 | 113 | 5.7 | 100.15 | 108.8 | 110.9 | 110.6714
-5.6 | 1.2 | 9.8 | 120.24 | 119.34 | 118.65 | 1.6 | 19.85 | 118.28 | 118.2333 | 117.42 | 2.5 | 44.4 | 115.623 | 116.6222 | 115.86 | 4.3011 | 93 | 111.7 | 113.323 | 112.7 | 5.9 | 101 | 108.8 | 110.825 | 110.42
-5.5 | 1.2 | 12 | 119.8343 | 118.98 | 118.2586 | 1.6 | 27.6 | 117.82 | 117.925 | 117.1333 | 2.525 | 56.2 | 115.3866 | 116.25 | 115.58 | 4.4011 | 97.6 | 111.5333 | 113.0563 | 112.5 | 6.1 | 101.4 | 108.8 | 110.64 | 110.2333
-5.4 | 1.2 | 16.3 | 119.27 | 118.42 | 117.8957 | 1.7 | 37.3 | 117.42 | 117.4 | 116.76 | 2.6 | 68.28 | 115.1659 | 115.95 | 115.2667 | 4.5011 | 100 | 111.24 | 112.923 | 112.2333 | 6.2 | 101.7 | 108.65 | 110.6 | 109.9875
-5.3 | 1.2 | 24.3 | 118.8867 | 118.02 | 117.5957 | 1.7 | 49.4 | 117.16 | 117.1 | 116.36 | 2.7 | 80.02 | 114.943 | 115.6182 | 114.98 | 4.6011 | 101.3 | 111.14 | 112.643 | 112 | 6.38 | 101.9 | 108.5 | 110.425 | 109.825
-5.2 | 1.2 | 35.5 | 118.54 | 117.82 | 117.3004 | 1.7 | 63 | 116.86 | 116.9 | 115.98 | 2.7 | 89.3 | 114.598 | 115.28 | 114.8 | 4.7345 | 102.6 | 110.94 | 112.342 | 111.8 | 6.475 | 102.2667 | 108.425 | 110.2333 | 109.7333
-5.1 | 1.3 | 49.9 | 118.145 | 117.45 | 116.969 | 1.8 | 75.8 | 116.56 | 116.4333 | 115.66 | 2.9 | 97.38 | 114.298 | 114.98 | 114.44 | 4.9012 | 103.1 | 110.8 | 112.022 | 111.625 | 6.575 | 102.5 | 108.28 | 109.9333 | 109.54
-5 | 1.3 | 64.5 | 117.58 | 117.12 | 116.489 | 1.8 | 88 | 116.28 | 115.95 | 115.38 | 3 | 101.58 | 114.143 | 114.68 | 114.2 | 5.0012 | 103.5 | 110.64 | 111.922 | 111.425 | 6.78 | 102.7 | 108.1 | 109.7333 | 109.3833
-4.9 | 1.3 | 78.9 | 117.38 | 116.62 | 116.149 | 1.9 | 97.6 | 115.95 | 115.64 | 115.06 | 3 | 104.02 | 113.863 | 114.38 | 113.86 | 5.1013 | 103.9333 | 110.54 | 111.742 | 111.1 | 7 | 102.9 | 108.04 | 109.6 | 109.2444
-4.8 | 1.4 | 91.8 | 117.0396 | 116.22 | 115.689 | 2 | 103.6 | 115.58 | 115.3 | 114.58 | 3.1 | 105.38 | 113.543 | 114.08 | 113.6 | 5.2263 | 104.3 | 110.34 | 111.362 | 110.98 | 7.1 | 103 | 108 | 109.5 | 109.1
-4.7 | 1.4 | 102 | 116.769 | 115.85 | 115.349 | 2.1 | 106.7 | 115.26 | 115.02 | 114.3833 | 3.3 | 106.18 | 113.273 | 113.7909 | 113.3 | 5.4013 | 104.4 | 110.1 | 111.142 | 110.78 | 7.275 | 103 | 107.94 | 109.4 | 108.8
-4.6 | 1.4 | 108.2 | 116.394 | 115.5 | 115.009 | 2.1 | 108.4 | 114.96 | 114.72 | 114 | 3.4 | 106.74 | 112.9897 | 113.4875 | 112.9 | 5.5014 | 104.5667 | 109.86 | 110.842 | 110.44 | 7.45 | 103.1 | 107.74 | 109.2 | 108.5
-4.5 | 1.45 | 111.3 | 116.039 | 115.12 | 114.679 | 2.2 | 109.3 | 114.56 | 114.32 | 113.7 | 3.5 | 107.1 | 112.863 | 113.1333 | 112.6571 | 5.7014 | 104.7 | 109.56 | 110.642 | 110.3 | 7.675 | 103.2 | 107.65 | 108.8 | 108.4
-4.4 | 1.5 | 112.5 | 115.539 | 114.72 | 114.279 | 2.3 | 109.8 | 114.2 | 113.96 | 113.44 | 3.6 | 107.3 | 112.648 | 112.88 | 112.34 | 5.9015 | 104.8 | 109.36 | 110.442 | 110.1 | 7.875 | 103.3 | 107.44 | 108.6 | 108.2
-4.3 | 1.5667 | 112.85 | 115.239 | 114.32 | 113.9783 | 2.4 | 109.9 | 113.95 | 113.54 | 113.12 | 3.7 | 107.4 | 112.262 | 112.64 | 112.04 | 6.1015 | 104.9 | 109.16 | 110.142 | 109.85 | 8.075 | 103.6 | 107.25 | 108.5 | 108.1
-4.2 | 1.6 | 113 | 114.869 | 114.06 | 113.678 | 2.5 | 110 | 113.5667 | 113.25 | 112.76 | 4 | 107.4 | 111.962 | 112.28 | 111.75 | 6.2016 | 104.9333 | 109 | 109.962 | 109.525 | 8.375 | 103.7 | 107.06 | 108.325 | 107.9
-4.1 | 1.65 | 113 | 114.459 | 113.72 | 113.278 | 2.6 | 110 | 113.2667 | 112.92 | 112.52 | 4.1 | 107.4 | 111.722 | 111.98 | 111.42 | 6.5016 | 105 | 108.86 | 109.762 | 109.4 | 8.5 | 103.7 | 106.96 | 108.2333 | 107.7
-4 | 1.7 | 112.9 | 114.094 | 113.4 | 112.8947 | 2.7 | 110 | 112.875 | 112.44 | 112.22 | 4.25 | 107.4 | 111.522 | 111.62 | 111.16 | 6.7017 | 105.1 | 108.7 | 109.502 | 109.15 | 8.65 | 103.7 | 106.76 | 108.0333 | 107.5
-3.9 | 1.8 | 112.8 | 113.599 | 113.12 | 112.503 | 2.8 | 109.9 | 112.525 | 112.24 | 111.82 | 4.45 | 107.4 | 111.222 | 111.35 | 110.96 | 6.9017 | 105.1 | 108.56 | 109.162 | 108.8667 | 8.9667 | 103.7 | 106.56 | 107.8333 | 107.4
-3.8 | 1.8333 | 112.6 | 113.139 | 112.74 | 112.303 | 2.9 | 109.8 | 112.1333 | 111.92 | 111.52 | 4.6 | 107.4 | 110.822 | 111.02 | 110.66 | 7.2018 | 105.2 | 108.28 | 109.062 | 108.6 | 9.2 | 103.7 | 106.46 | 107.5 | 107.3
-3.7 | 1.9 | 112.3667 | 112.839 | 112.22 | 112.103 | 3.1 | 109.8 | 111.9 | 111.54 | 111.22 | 4.8 | 107.4 | 110.622 | 110.7 | 110.32 | 7.5519 | 105.2 | 108.08 | 108.862 | 108.425 | 9.5 | 103.8 | 106.3 | 107.4 | 107.1182
-3.6 | 2 | 112.1 | 112.519 | 111.92 | 111.568 | 3.2 | 109.7 | 111.6 | 111.14 | 110.82 | 4.95 | 107.4 | 110.122 | 110.56 | 110.02 | 7.7019 | 105.2 | 107.8 | 108.562 | 108.2 | 9.7667 | 103.9 | 106.2 | 107.15 | 106.92
-3.5 | 2.1 | 111.9 | 112.119 | 111.66 | 111.1613 | 3.4 | 109.6 | 111.1333 | 110.84 | 110.46 | 5.2333 | 107.4 | 109.972 | 110.1667 | 109.8 | 7.902 | 105.3 | 107.7 | 108.372 | 107.9667 | 10.25 | 103.9 | 106.16 | 106.92 | 106.63
-3.4 | 2.2 | 111.7 | 111.759 | 111.28 | 110.808 | 3.6 | 109.5 | 110.8333 | 110.58 | 110.02 | 5.5 | 107.4 | 109.6887 | 109.92 | 109.58 | 8.3021 | 105.2 | 107.58 | 108.082 | 107.72 | 10.82 | 104 | 105.96 | 106.7 | 106.42
-3.3 | 2.225 | 111.3 | 111.449 | 110.8333 | 110.578 | 3.8 | 109.3333 | 110.675 | 110.3 | 109.76 | 5.8333 | 107.3 | 109.397 | 109.42 | 109.08 | 8.6021 | 105.2 | 107.35 | 107.842 | 107.4667 | 11.8 | 104 | 105.7667 | 106.6 | 106.22
-3.2 | 2.325 | 110.9 | 111.119 | 110.48 | 110.268 | 4 | 109.1 | 110.3375 | 109.9333 | 109.56 | 6.1333 | 107.2 | 109.072 | 109.12 | 108.88 | 9.2023 | 105.2 | 107.1143 | 107.4816 | 107.22 | 12.7 | 104 | 105.6 | 106.4667 | 106.0333
-3.1 | 2.5 | 110.7 | 110.718 | 110.28 | 110.028 | 4.2 | 108.95 | 109.9 | 109.6 | 109.18 | 6.4333 | 107.1 | 108.847 | 108.96 | 108.66 | 10.1025 | 105.1 | 106.9 | 107.281 | 106.9333 | 13.8667 | 104 | 105.48 | 106.24 | 105.94
-3 | 2.625 | 110.4 | 110.318 | 110.06 | 109.547 | 4.5 | 108.7333 | 109.66 | 109.4167 | 108.88 | 6.9 | 107 | 108.5887 | 108.7 | 108.36 | 10.8277 | 105 | 106.8 | 107.081 | 106.8 | 16.7667 | 103.9 | 105.38 | 106.14 | 105.74
-2.9 | 2.8 | 110 | 109.958 | 109.54 | 109.167 | 5 | 108.35 | 109.44 | 109.06 | 108.58 | 7.5333 | 106.8 | 108.242 | 108.4 | 107.88 | 12.5031 | 105 | 106.6 | 106.921 | 106.6 | 20.8 | 103.8 | 105.2 | 105.9667 | 105.6
-2.8 | 3.1 | 109.8 | 109.518 | 109.0333 | 108.867 | 5.6 | 108.2 | 108.96 | 108.64 | 108.36 | 8.5333 | 106.75 | 107.882 | 108 | 107.7 | 14.9787 | 104.9 | 106.2 | 106.781 | 106.2 | 26.9667 | 103.8 | 105.06 | 105.7833 | 105.38
-2.7 | 3.6 | 109.45 | 109.218 | 108.68 | 108.507 | 6.4 | 108 | 108.62 | 108.24 | 107.925 | 10.1 | 106.55 | 107.622 | 107.8167 | 107.48 | 19.6299 | 104.8 | 106 | 106.401 | 106.1 | 35.5667 | 103.7 | 104.7714 | 105.6 | 105.14
-2.6 | 4.4 | 109.05 | 109.018 | 108.36 | 108.207 | 7.8 | 107.8667 | 108.4333 | 108.02 | 107.725 | 12.6 | 106.4 | 107.542 | 107.45 | 107.18 | 26.4066 | 104.7 | 105.8 | 106.081 | 105.7 | 45.8 | 103.7 | 104.5333 | 105.4 | 104.94
-2.5 | 5.9 | 108.7 | 108.718 | 108.22 | 107.887 | 10.5 | 107.4667 | 107.92 | 107.64 | 107.425 | 17.2 | 106.08 | 107.221 | 107.06 | 106.8667 | 35.6089 | 104.6 | 105.6 | 105.801 | 105.5 | 57.2333 | 103.6 | 104.4571 | 105.1 | 104.74
-2.4 | 8.9 | 108.4 | 108.318 | 107.86 | 107.527 | 15.2 | 107.1667 | 107.66 | 107.34 | 107.125 | 24.3 | 105.825 | 106.781 | 106.82 | 106.7 | 47.1118 | 104.5 | 105.5 | 105.601 | 105.325 | 68.0333 | 103.5333 | 104.3 | 104.8 | 104.6
-2.3 | 14.5 | 108.1 | 107.8513 | 107.5 | 107.177 | 23.1 | 106.9667 | 107.36 | 107 | 106.74 | 34.7 | 105.7 | 106.441 | 106.54 | 106.5 | 58.3813 | 104.4 | 105.1 | 105.421 | 105.14 | 78.2 | 103.4 | 104.1333 | 104.6 | 104.3333
-2.2 | 24 | 107.7 | 107.518 | 107.2 | 106.927 | 34.45 | 106.5667 | 107.02 | 106.66 | 106.45 | 46.6 | 105.575 | 106.201 | 106.375 | 106.1 | 70.5176 | 104.4 | 104.9429 | 105.281 | 104.98 | 86.2 | 103.3 | 103.94 | 104.4 | 104.2
-2.1 | 36.8 | 107.4 | 107.218 | 106.7 | 106.627 | 47.15 | 106.4 | 106.72 | 106.3 | 105.95 | 59.2 | 105.4 | 105.961 | 105.95 | 105.8 | 81.1203 | 104.2 | 104.66 | 105.001 | 104.725 | 92.8667 | 103.1 | 103.74 | 104.3 | 104.1
-2 | 50.2 | 107.1 | 106.938 | 106.4 | 106.2413 | 60.1 | 106.1667 | 106.32 | 106.08 | 105.58 | 71.7 | 105.2 | 105.701 | 105.65 | 105.4 | 89.3223 | 104.2 | 104.46 | 104.701 | 104.5 | 96.8667 | 102.9 | 103.6 | 104.2 | 103.75
-1.9 | 64.1 | 106.8 | 106.478 | 106.2 | 105.8862 | 72.3 | 105.7667 | 106.0333 | 105.7667 | 105.375 | 82.5 | 104.975 | 105.301 | 105.3 | 105.15 | 94.4036 | 103.95 | 104.16 | 104.481 | 104.3 | 98.6 | 102.8 | 103.44 | 104 | 103.56
-1.8 | 77.15 | 106.52 | 106.118 | 105.8333 | 105.6078 | 83.65 | 105.575 | 105.6667 | 105.4333 | 105.1333 | 91.3 | 104.8 | 105.001 | 105.1 | 104.92 | 98.1245 | 103.7 | 103.9667 | 104.1043 | 103.9 | 99.7 | 102.75 | 103.2333 | 103.82 | 103.46
-1.7 | 87.8 | 106.1 | 105.678 | 105.4 | 105.316 | 92.75 | 105.3 | 105.275 | 105.14 | 104.85 | 96.55 | 104.55 | 104.661 | 104.86 | 104.6 | 99.925 | 103.5 | 103.8333 | 103.896 | 103.7 | 100.5 | 102.6 | 102.9455 | 103.44 | 103.2667
-1.6 | 96.1 | 105.65 | 105.3513 | 105.0667 | 105.016 | 97.8 | 105.075 | 105.0333 | 104.8 | 104.58 | 99.3 | 104.275 | 104.461 | 104.5 | 104.36 | 100.625 | 103.3 | 103.62 | 103.721 | 103.5 | 100.9 | 102.5 | 102.78 | 103.12 | 103.18
-1.5 | 100.5 | 105.26 | 104.9978 | 104.7667 | 104.786 | 100.08 | 104.7 | 104.76 | 104.54 | 104.36 | 100.82 | 103.975 | 104.101 | 104.22 | 104.12 | 101.1583 | 103.1 | 103.325 | 103.521 | 103.35 | 101.1 | 102.4 | 102.68 | 102.92 | 102.98
-1.4 | 102.1 | 105.025 | 104.797 | 104.4667 | 104.446 | 101.42 | 104.4 | 104.28 | 104.0667 | 103.88 | 101.325 | 103.775 | 103.801 | 103.92 | 103.72 | 101.4917 | 102.9 | 103.2 | 103.221 | 103.02 | 101.4 | 102.3 | 102.54 | 102.8 | 102.76
-1.3 | 103 | 104.55 | 104.377 | 104.1667 | 103.976 | 101.78 | 104 | 103.98 | 103.8 | 103.58 | 101.7333 | 103.46 | 103.561 | 103.6667 | 103.44 | 101.6917 | 102.75 | 103 | 102.941 | 102.82 | 101.4 | 102.2 | 102.26 | 102.7 | 102.48
-1.2 | 103.1 | 104.1667 | 103.977 | 103.8333 | 103.616 | 101.94 | 103.5 | 103.7 | 103.6 | 103.38 | 101.9 | 103.275 | 103.261 | 103.38 | 103.22 | 101.725 | 102.5 | 102.725 | 102.641 | 102.68 | 101.5 | 102 | 102.06 | 102.46 | 102.26
-1.1 | 103.1 | 103.825 | 103.597 | 103.55 | 103.2831 | 102 | 103.3 | 103.3 | 103.3 | 103.08 | 102 | 102.975 | 103.041 | 102.94 | 102.94 | 101.725 | 102.4 | 102.5333 | 102.32 | 102.36 | 101.5 | 101.9 | 101.88 | 102.18 | 102.025
-1 | 102.9 | 103.64 | 103.277 | 103.1 | 102.946 | 102 | 103 | 103.1 | 102.96 | 102.74 | 102 | 102.68 | 102.681 | 102.74 | 102.62 | 101.725 | 102.2 | 102.34 | 102.2367 | 101.98 | 101.4 | 101.8 | 101.7667 | 102.0667 | 101.94
-0.9 | 102.7 | 103.06 | 102.9313 | 102.9 | 102.746 | 101.86 | 102.7 | 103 | 102.6 | 102.36 | 101.9 | 102.48 | 102.34 | 102.44 | 102.34 | 101.625 | 102.1 | 102.08 | 102.02 | 101.7 | 101.2 | 101.6 | 101.58 | 101.82 | 101.84
-0.8 | 102.4 | 102.74 | 102.597 | 102.55 | 102.366 | 101.66 | 102.4 | 102.4 | 102.3 | 102 | 101.76 | 102.18 | 102.18 | 102.3 | 102.04 | 101.425 | 101.9 | 101.74 | 101.87 | 101.58 | 101.1 | 101.4 | 101.28 | 101.78 | 101.58
-0.7 | 102 | 102.3667 | 102.397 | 102.3 | 102.066 | 101.6 | 102.05 | 102.1 | 102 | 101.8667 | 101.4 | 102 | 101.98 | 101.98 | 101.84 | 101.2917 | 101.6 | 101.5 | 101.62 | 101.3833 | 101 | 101.1 | 101.14 | 101.58 | 101.3
-0.6 | 101.7 | 102.025 | 101.957 | 101.9 | 101.765 | 101.48 | 101.85 | 101.8167 | 101.5 | 101.5571 | 101.2 | 101.7 | 101.64 | 101.64 | 101.58 | 101.0917 | 101.3667 | 101.3 | 101.2867 | 101.2 | 100.7 | 101 | 100.98 | 101.38 | 101.06
-0.5 | 101.4 | 101.65 | 101.477 | 101.6 | 101.545 | 101.28 | 101.5 | 101.5 | 101.42 | 101.4 | 101.1 | 101.5 | 101.395 | 101.28 | 101.24 | 100.925 | 101.1667 | 101.1 | 101.02 | 101.0833 | 100.5 | 100.7 | 100.78 | 101.15 | 100.86
-0.4 | 101 | 101.25 | 101.097 | 101.2 | 101.2583 | 101.08 | 101.15 | 101.1 | 101.15 | 101 | 100.9 | 101.26 | 101.17 | 101.04 | 100.98 | 100.7917 | 100.8667 | 100.9 | 100.84 | 100.9 | 100.3 | 100.6 | 100.58 | 100.92 | 100.76
-0.3 | 100.8 | 100.95 | 100.897 | 100.94 | 100.95 | 100.86 | 100.9 | 100.8 | 100.9 | 100.6286 | 100.8 | 100.95 | 100.7533 | 100.74 | 100.6 | 100.625 | 100.7 | 100.56 | 100.58 | 100.5 | 100.2 | 100.425 | 100.38 | 100.7 | 100.56
-0.2 | 100.5333 | 100.65 | 100.697 | 100.64 | 100.685 | 100.46 | 100.4 | 100.5 | 100.5667 | 100.44 | 100.6 | 100.5 | 100.5 | 100.52 | 100.45 | 100.5 | 100.5 | 100.4 | 100.3867 | 100.3 | 100.1 | 100.3 | 100.24 | 100.3 | 100.38
-0.1 | 100.3 | 100.35 | 100.357 | 100.32 | 100.305 | 100.18 | 100.2 | 100.3 | 100.2 | 100.24 | 100.3 | 100.3 | 100.28 | 100.26 | 100.32 | 100.2 | 100.2 | 100.16 | 100.2 | 100.1 | 100.1 | 100.025 | 100.2 | 100.2 | 100.18
0 | 100 | 100 | 100.0003 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.6333 | 99.75 | 99.7166 | 99.76 | 99.7249 | 99.7 | 99.6 | 99.7 | 99.7667 | 99.7 | 99.82 | 99.8 | 99.78 | 99.75 | 99.86 | 99.8 | 99.9 | 99.82 | 99.7533 | 99.7 | 99.9 | 99.9 | 99.7364 | 99.875 | 99.78
0.2 | 99.24 | 99.5333 | 99.3966 | 99.56 | 99.4999 | 99.4 | 99.4 | 99.52 | 99.5667 | 99.44 | 99.62 | 99.5 | 99.4999 | 99.6333 | 99.5667 | 99.5249 | 99.7 | 99.66 | 99.4199 | 99.4 | 99.7 | 99.7333 | 99.5 | 99.66 | 99.58
0.3 | 98.94 | 99.3 | 99.2165 | 99.26 | 99.3248 | 99.06 | 99.1 | 99.24 | 99.2667 | 99.15 | 99.26 | 99.35 | 99.3199 | 99.5 | 99.28 | 99.2248 | 99.4 | 99.36 | 99.3199 | 99.2167 | 99.3333 | 99.4333 | 99.3 | 99.425 | 99.36
0.4 | 98.7 | 99.0667 | 99.0165 | 99.02 | 99.0648 | 98.7 | 98.8 | 98.92 | 99.05 | 98.95 | 99.02 | 99.15 | 99.0198 | 99.1 | 98.9714 | 98.9247 | 99.2 | 99.1667 | 99.1698 | 99.02 | 98.98 | 99.2333 | 99 | 99.26 | 99.2
0.5 | 98.34 | 98.75 | 98.8165 | 98.72 | 98.8247 | 98.4 | 98.7 | 98.72 | 98.7 | 98.68 | 98.7 | 98.9 | 98.7197 | 98.8333 | 98.8167 | 98.7247 | 99.1 | 98.98 | 98.9598 | 98.84 | 98.74 | 98.9333 | 98.9 | 99.1 | 99
0.6 | 98.08 | 98.475 | 98.3164 | 98.52 | 98.6247 | 98.1 | 98.3333 | 98.52 | 98.5333 | 98.44 | 98.3 | 98.5 | 98.5197 | 98.6333 | 98.7 | 98.4246 | 98.8 | 98.86 | 98.6864 | 98.62 | 98.54 | 98.7 | 98.7 | 99 | 98.9
0.7 | 97.75 | 98.1 | 98.1164 | 98.42 | 98.4446 | 97.6 | 98.1 | 98.24 | 98.2333 | 98.3 | 98 | 98.3 | 98.4197 | 98.4333 | 98.36 | 98.0245 | 98.6 | 98.38 | 98.4597 | 98.42 | 98.34 | 98.6 | 98.6 | 98.9 | 98.7833
0.8 | 97.36 | 97.8 | 97.9163 | 98.14 | 98.1645 | 97.4 | 97.825 | 98.02 | 98.1 | 98.1 | 97.4 | 98.1 | 98.1196 | 98.1667 | 98.1667 | 97.8245 | 98.4 | 98.18 | 98.3197 | 98.22 | 98.05 | 98.4 | 98.4 | 98.675 | 98.6
0.9 | 96.94 | 97.5 | 97.7363 | 97.82 | 98.0245 | 96.9333 | 97.52 | 97.72 | 98.02 | 97.85 | 97.2 | 97.7 | 97.9196 | 98 | 98.05 | 97.5744 | 98.15 | 98.08 | 98.1196 | 98.02 | 97.75 | 98.25 | 98.3091 | 98.48 | 98.4
1 | 96.58 | 97.2333 | 97.5363 | 97.64 | 97.7444 | 96.6 | 97.25 | 97.5333 | 97.675 | 97.66 | 96.75 | 97.4 | 97.5195 | 97.82 | 97.76 | 97.3243 | 97.9667 | 97.88 | 97.8796 | 97.825 | 97.3 | 98.1667 | 98.12 | 98.3 | 98.3
1.1 | 96.12 | 96.875 | 97.3362 | 97.34 | 97.4744 | 96 | 97.1 | 97.4 | 97.4333 | 97.425 | 96.3 | 97.2 | 97.3195 | 97.5 | 97.4667 | 97.1243 | 97.7 | 97.68 | 97.6795 | 97.5667 | 97.05 | 97.85 | 97.92 | 98.16 | 98.0333
1.2 | 95.68 | 96.575 | 97.0562 | 97.1 | 97.1643 | 95.4 | 96.72 | 97.14 | 97.2 | 97.14 | 95.7 | 96.9 | 97.1194 | 97.325 | 97.3 | 96.5241 | 97.4 | 97.5 | 97.4445 | 97.44 | 96.72 | 97.75 | 97.82 | 97.9 | 97.8167
1.3 | 95 | 96.4 | 96.6361 | 97 | 97.0243 | 94.6 | 96.3333 | 96.92 | 96.85 | 97 | 95.1 | 96.7 | 96.8194 | 97.14 | 97.15 | 96.2241 | 97.1 | 97.3 | 97.2528 | 97.25 | 96.36 | 97.3667 | 97.6 | 97.8 | 97.62
1.4 | 93.74 | 96.3 | 96.4361 | 96.8 | 96.8442 | 93.65 | 96.1333 | 96.62 | 96.74 | 96.75 | 94.4 | 96.45 | 96.6943 | 96.94 | 96.95 | 95.7239 | 96.8 | 97.0667 | 97.1194 | 97.1 | 96.05 | 97.175 | 97.4571 | 97.6 | 97.4
1.5 | 91.66 | 95.8667 | 96.156 | 96.5 | 96.6041 | 91.95 | 95.88 | 96.25 | 96.625 | 96.52 | 93.3 | 96.2 | 96.5393 | 96.725 | 96.6667 | 94.9237 | 96.7 | 96.9 | 96.9527 | 96.9 | 95.45 | 96.95 | 97.1333 | 97.5 | 97.3
1.6 | 87.6 | 95.5667 | 95.936 | 96.2 | 96.3641 | 89.25 | 95.64 | 96.14 | 96.26 | 96.32 | 91.7 | 96.1 | 96.2192 | 96.4 | 96.4667 | 94.1235 | 96.4 | 96.6 | 96.7193 | 96.725 | 94.975 | 96.575 | 96.94 | 97.15 | 97.18
1.7 | 80.5 | 95.4 | 95.7159 | 96 | 96.064 | 84.35 | 95.45 | 95.94 | 95.98 | 96.075 | 89.1 | 95.7 | 96.0525 | 96.15 | 96.2667 | 92.6232 | 96.2 | 96.425 | 96.6193 | 96.525 | 94.45 | 96.3 | 96.74 | 96.925 | 96.98
1.8 | 70.4 | 95.1 | 95.4826 | 95.7667 | 95.864 | 76.7 | 95.2 | 95.58 | 95.78 | 95.78 | 84.5 | 95.4 | 95.8525 | 96 | 96.02 | 91.1828 | 95.9 | 96.175 | 96.4193 | 96.325 | 93.55 | 96.075 | 96.6 | 96.8 | 96.68
1.9 | 58.4 | 94.9 | 95.2826 | 95.6 | 95.6639 | 67.5 | 94.75 | 95.44 | 95.58 | 95.5 | 77.3 | 95.1 | 95.4991 | 95.875 | 95.825 | 88.3554 | 95.6 | 95.96 | 96.1192 | 96.1333 | 92.425 | 95.875 | 96.44 | 96.675 | 96.5
2 | 45.8 | 94.7 | 94.9158 | 95.225 | 95.5239 | 55.225 | 94.45 | 95.14 | 95.4 | 95.375 | 67.6 | 94.8 | 95.2791 | 95.6 | 95.625 | 83.971 | 95.35 | 95.775 | 96.0192 | 95.9333 | 90.4 | 95.6 | 96.24 | 96.5 | 96.38
2.1 | 33.6 | 94.3 | 94.6558 | 95.1 | 95.2738 | 44.1 | 94.2 | 94.84 | 95.14 | 95.18 | 56.6 | 94.5 | 94.999 | 95.3 | 95.5 | 77.386 | 95.05 | 95.575 | 95.7191 | 95.65 | 88 | 95.4 | 96.04 | 96.32 | 96.2
2.2 | 22 | 94 | 94.4557 | 94.9667 | 95.0838 | 32.8 | 93.95 | 94.54 | 94.825 | 95.06 | 45.5 | 94.2 | 94.799 | 95.1 | 95.3333 | 68.5171 | 94.8 | 95.3 | 95.4441 | 95.4 | 83.3 | 95.1 | 95.85 | 96.2 | 96.08
2.3 | 13.8333 | 93.8 | 94.1957 | 94.7333 | 94.7837 | 21.9 | 93.6667 | 94.36 | 94.6 | 94.66 | 34.4 | 93.9 | 94.5189 | 95 | 95.1333 | 57.6144 | 94.3 | 95.1333 | 95.119 | 95.3 | 76.8 | 94.8 | 95.55 | 95.94 | 95.88
2.4 | 8.1333 | 93.4 | 93.9657 | 94.42 | 94.5236 | 14.5 | 93.25 | 94.15 | 94.4333 | 94.4 | 24.5 | 93.5 | 94.3189 | 94.8 | 94.7333 | 46.9784 | 94 | 94.9 | 95.044 | 95.1 | 68.1 | 94.66 | 95.36 | 95.875 | 95.68
2.5 | 5.4 | 93.1 | 93.7756 | 94.14 | 94.2836 | 9.8 | 92.75 | 93.92 | 94.2 | 94.28 | 17 | 93.3 | 94.2188 | 94.58 | 94.45 | 36.5091 | 93.9 | 94.75 | 94.794 | 94.825 | 58.1 | 94.4 | 95.2 | 95.6 | 95.55
2.6 | 4.05 | 92.8 | 93.4656 | 93.8833 | 94.0835 | 7.45 | 92.6 | 93.66 | 94 | 94.06 | 12.2 | 92.7667 | 93.8188 | 94.28 | 94.3 | 27.1401 | 93.5 | 94.5 | 94.4589 | 94.6667 | 48 | 94.1 | 95.06 | 95.34 | 95.4
2.7 | 3.28 | 92.5333 | 93.1755 | 93.6 | 93.8835 | 6.1 | 92.4333 | 93.32 | 93.8 | 93.7286 | 9.6 | 92.5667 | 93.6187 | 94.0667 | 94.1 | 19.905 | 93.3 | 94.2667 | 94.2388 | 94.52 | 37.9 | 93.9 | 94.86 | 95.1 | 95.24
2.8 | 2.86 | 92.2333 | 93.0355 | 93.5 | 93.5834 | 5.2333 | 91.875 | 93.06 | 93.5286 | 93.6 | 8.1 | 92.0667 | 93.3187 | 93.84 | 93.9 | 15.3038 | 92.9 | 93.9667 | 94.0388 | 94.375 | 30.025 | 93.7 | 94.56 | 95.1 | 95
2.9 | 2.6 | 91.8 | 92.6354 | 93.2667 | 93.3233 | 4.6667 | 91.6333 | 92.7 | 93.3 | 93.34 | 7.2 | 91.7667 | 93.0186 | 93.625 | 93.75 | 12.4031 | 92.65 | 93.7667 | 93.9021 | 94.125 | 22.7333 | 93.4 | 94.3667 | 94.825 | 94.8
3 | 2.4 | 91.5333 | 92.3754 | 92.9 | 93.0633 | 4.2333 | 91.275 | 92.5 | 93 | 93.2 | 6.65 | 91.3667 | 92.9186 | 93.3333 | 93.5667 | 10.8027 | 92.35 | 93.4667 | 93.7187 | 93.875 | 17.7 | 93 | 94.18 | 94.6 | 94.6667
3.1 | 2.28 | 91.2333 | 92.1754 | 92.775 | 92.9232 | 3.9667 | 91.025 | 92.2 | 92.76 | 92.88 | 6.0667 | 91 | 92.6185 | 93.0333 | 93.32 | 9.8025 | 92.0667 | 93.2 | 93.4987 | 93.7 | 14.55 | 92.7 | 93.98 | 94.4 | 94.4
3.2 | 2.2 | 90.85 | 91.9753 | 92.5 | 92.6899 | 3.7667 | 90.475 | 92.02 | 92.56 | 92.64 | 5.8 | 90.8 | 92.3185 | 92.85 | 93.04 | 9.0023 | 91.55 | 93.025 | 93.3187 | 93.52 | 12.6667 | 92.3 | 93.88 | 94.2 | 94.225
3.3 | 2.1 | 90.5333 | 91.6753 | 92.3 | 92.4898 | 3.5667 | 90.2 | 91.92 | 92.32 | 92.34 | 5.4667 | 90.2 | 91.9184 | 92.6 | 92.82 | 8.4021 | 91.1667 | 92.825 | 93.0986 | 93.2 | 11.55 | 92 | 93.58 | 94 | 94.06
3.4 | 1.98 | 90.35 | 91.3752 | 92.0333 | 92.3231 | 3.375 | 89.8 | 91.72 | 92.06 | 92.1 | 5.25 | 89.9 | 91.8184 | 92.4 | 92.62 | 8.102 | 90.9 | 92.7 | 92.9586 | 93.075 | 10.6 | 91.64 | 93.38 | 93.8 | 93.9
3.5 | 1.9 | 89.95 | 91.1352 | 91.82 | 92.063 | 3.175 | 89.4 | 91.44 | 91.8667 | 91.96 | 4.95 | 89.4 | 91.6383 | 92.3 | 92.42 | 7.8019 | 90.5 | 92.425 | 92.7185 | 92.9 | 10.1667 | 91.34 | 93.1667 | 93.6 | 93.7
3.6 | 1.8 | 89.55 | 90.8902 | 91.52 | 91.803 | 3 | 89 | 91.04 | 91.6667 | 91.76 | 4.6667 | 89.0333 | 91.3383 | 92.1 | 92.08 | 7.4018 | 90.2 | 92.225 | 92.5185 | 92.8 | 9.8333 | 90.94 | 93 | 93.4 | 93.56
3.7 | 1.78 | 89.1 | 90.4751 | 91.275 | 91.6029 | 2.875 | 88.6 | 90.75 | 91.3667 | 91.5667 | 4.5667 | 88.6333 | 91.1382 | 91.8 | 91.875 | 7.2018 | 89.5 | 91.925 | 92.3185 | 92.4667 | 9.5 | 90.72 | 92.8 | 93.22 | 93.2333
3.8 | 1.7 | 88.8333 | 90.3951 | 91 | 91.3895 | 2.7 | 88.3 | 90.55 | 91.175 | 91.4 | 4.375 | 88.1667 | 90.8932 | 91.6 | 91.58 | 7.0018 | 89.2 | 91.8 | 92.0184 | 92.14 | 9.1 | 90.26 | 92.6 | 92.94 | 92.98
3.9 | 1.6 | 88.38 | 90.095 | 90.88 | 91.0978 | 2.6 | 87.8 | 90.35 | 90.9143 | 91.16 | 4.2 | 87.6 | 90.7381 | 91.42 | 91.4 | 6.7017 | 88.7 | 91.4 | 91.8784 | 92 | 8.8 | 89.74 | 92.4 | 92.72 | 92.9
4 | 1.6 | 88.1 | 89.895 | 90.6 | 90.9227 | 2.5 | 87.2 | 90.05 | 90.725 | 90.7667 | 4.1 | 87.15 | 90.3581 | 91.22 | 91.2 | 6.4016 | 88.5 | 91.2 | 91.6783 | 91.8 | 8.6 | 89.52 | 92 | 92.52 | 92.7
4.1 | 1.5 | 87.6 | 89.6949 | 90.38 | 90.7227 | 2.4 | 86.8 | 89.7667 | 90.57 | 90.58 | 3.8 | 86.8333 | 90.138 | 90.84 | 91.0833 | 6.3016 | 88 | 90.9 | 91.4783 | 91.64 | 8.3 | 89.125 | 91.8 | 92.32 | 92.6
4.2 | 1.5 | 87 | 89.4749 | 90.28 | 90.4976 | 2.3 | 86.4 | 89.6 | 90.2714 | 90.38 | 3.7 | 86.4667 | 89.9513 | 90.52 | 90.85 | 6.1015 | 87.45 | 90.7 | 91.2783 | 91.475 | 8.1 | 88.78 | 91.5 | 92.2 | 92.4
4.3 | 1.4 | 86.5 | 89.1749 | 89.98 | 90.2426 | 2.2 | 85.7333 | 89.38 | 90.02 | 90.28 | 3.5 | 85.7 | 89.6429 | 90.3333 | 90.54 | 5.9015 | 86.9333 | 90.45 | 91.0782 | 91.16 | 7.9 | 88.4 | 91.4167 | 92.02 | 92.3
4.4 | 1.4 | 85.4 | 88.8948 | 89.78 | 90.0625 | 2.1 | 85.05 | 89.0667 | 89.84 | 89.98 | 3.4 | 85.3167 | 89.3512 | 90.14 | 90.3 | 5.6514 | 86.5 | 90.15 | 90.8982 | 90.9 | 7.7 | 87.95 | 91.14 | 91.82 | 92
4.5 | 1.3 | 84.15 | 88.6148 | 89.38 | 89.6224 | 2.1 | 84.26 | 88.74 | 89.58 | 89.8 | 3.35 | 84.6 | 89.1578 | 89.88 | 90.1 | 5.5014 | 86.2 | 89.975 | 90.6981 | 90.76 | 7.5 | 87.58 | 91.02 | 91.54 | 91.7
4.6 | 1.3 | 81.6 | 88.3947 | 89.28 | 89.4891 | 2 | 83.25 | 88.575 | 89.25 | 89.7 | 3.2 | 84.1 | 88.9578 | 89.64 | 89.775 | 5.3613 | 85.6 | 89.76 | 90.3981 | 90.46 | 7.3 | 87.25 | 90.72 | 91.4333 | 91.4
4.7 | 1.3 | 77.05 | 88.0814 | 89.05 | 89.3223 | 1.9 | 81.475 | 88.375 | 89.04 | 89.3 | 3.1 | 83.1 | 88.6577 | 89.44 | 89.55 | 5.2013 | 85.125 | 89.55 | 90.098 | 90.3 | 7.2 | 86.88 | 90.44 | 91.24 | 91.25
4.8 | 1.22 | 70 | 87.7646 | 88.7 | 89.0023 | 1.8 | 78.82 | 87.96 | 88.84 | 89.05 | 3 | 82.2 | 88.3177 | 89.24 | 89.32 | 5.0763 | 84.675 | 89.24 | 89.898 | 90.02 | 7 | 86.45 | 90.14 | 91.04 | 91.1
4.9 | 1.2 | 60.1 | 87.5146 | 88.4667 | 88.8022 | 1.8 | 73.84 | 87.82 | 88.58 | 88.8 | 2.9 | 80.85 | 88.1576 | 89.14 | 89.12 | 5.0012 | 84.1 | 89 | 89.6979 | 89.82 | 6.8 | 85.9 | 89.92 | 90.9 | 90.96
5 | 1.2 | 49.8 | 87.3145 | 88.15 | 88.5822 | 1.7 | 66.36 | 87.56 | 88.4 | 88.7 | 2.8 | 78.65 | 87.8976 | 88.9 | 89.02 | 4.7762 | 83.4 | 88.7333 | 89.4179 | 89.62 | 6.6 | 85.25 | 89.72 | 90.8 | 90.8
5.1 | 1.1 | 39.1 | 87.0645 | 88.02 | 88.3221 | 1.6 | 57.86 | 87.32 | 88.1 | 88.52 | 2.7 | 75.3 | 87.6175 | 88.65 | 88.82 | 4.7012 | 83 | 88.5333 | 89.2078 | 89.44 | 6.5 | 84.62 | 89.32 | 90.4 | 90.45
5.2 | 1.1 | 28.8 | 86.7544 | 87.75 | 88.1053 | 1.6 | 48.1 | 87 | 87.94 | 88.25 | 2.6 | 69.95 | 87.4775 | 88.4 | 88.5333 | 4.6011 | 82.1 | 88.32 | 89.0978 | 89.24 | 6.4 | 84.4 | 89.22 | 90.2 | 90.2667
5.3 | 1.1 | 20.3 | 86.5144 | 87.6 | 87.822 | 1.6 | 38.26 | 86.8 | 87.75 | 87.76 | 2.5 | 62.35 | 87.1374 | 88.22 | 88.2333 | 4.5011 | 81.1 | 87.975 | 88.8905 | 89.1 | 6.2 | 84 | 89.0333 | 90 | 89.85
5.4 | 1.1 | 13.7 | 86.3144 | 87.36 | 87.7219 | 1.5 | 29.06 | 86.55 | 87.6 | 87.56 | 2.5 | 53.9 | 86.7173 | 87.85 | 88.15 | 4.4011 | 79.8 | 87.8 | 88.6177 | 88.94 | 6.1 | 83.46 | 88.72 | 89.8 | 89.6667
5.5 | 1.1 | 10.05 | 85.9476 | 87.12 | 87.4219 | 1.5 | 21.26 | 86.2 | 87.42 | 87.3 | 2.4 | 45 | 86.4923 | 87.6 | 87.925 | 4.2011 | 77.6 | 87.46 | 88.5177 | 88.78 | 6 | 82.85 | 88.44 | 89.56 | 89.5
5.6 | 1.05 | 8.3 | 85.6643 | 86.72 | 87.2218 | 1.5 | 16.1 | 85.94 | 87.06 | 87.16 | 2.3 | 36.2 | 86.2373 | 87.5 | 87.6 | 4.101 | 74.3 | 87.1333 | 88.2605 | 88.48 | 5.85 | 82.3 | 88.24 | 89.46 | 89.4667
5.7 | 1 | 7.2 | 85.4342 | 86.375 | 86.8217 | 1.4 | 12.5 | 85.6 | 86.8 | 86.9667 | 2.2333 | 28.3 | 85.8772 | 87.32 | 87.4 | 4.001 | 69.6 | 86.86 | 88.0343 | 88.24 | 5.6 | 81.4 | 88.04 | 89.25 | 89.2667
5.8 | 1 | 6.6 | 85.0742 | 86.28 | 86.5216 | 1.4 | 10.85 | 85.42 | 86.5429 | 86.7 | 2.2 | 22.05 | 85.6171 | 87.08 | 87.1833 | 3.901 | 63 | 86.675 | 87.8287 | 88.04 | 5.5 | 80.5 | 87.84 | 89.02 | 89.1
5.9 | 1 | 6.2 | 84.8942 | 86.08 | 86.4216 | 1.4 | 9.7333 | 85.1 | 86.36 | 86.56 | 2.1333 | 17.6 | 85.4671 | 86.76 | 86.9833 | 3.901 | 55.6 | 86.38 | 87.6175 | 87.8667 | 5.5 | 79.1 | 87.54 | 88.82 | 89
6 | 1 | 5.95 | 84.5941 | 85.825 | 86.1415 | 1.4 | 9 | 84.86 | 86.16 | 86.28 | 2.1 | 15.1667 | 85.2371 | 86.46 | 86.7 | 3.8009 | 47.3 | 86.125 | 87.3811 | 87.6 | 5.3 | 77.3 | 87.24 | 88.6333 | 88.65
6.1 | 1 | 5.75 | 84.2941 | 85.6333 | 85.8415 | 1.3 | 8.7 | 84.6 | 85.82 | 85.98 | 2.1 | 13.7 | 84.967 | 86.18 | 86.4 | 3.7009 | 39.3 | 85.84 | 87.1174 | 87.4 | 5.3 | 74 | 87.06 | 88.42 | 88.4
6.2 | 1 | 5.5333 | 83.974 | 85.4 | 85.6381 | 1.3 | 8.2 | 84.375 | 85.46 | 85.66 | 2 | 12.64 | 84.4969 | 85.98 | 86.2 | 3.6009 | 32.2 | 85.58 | 86.9317 | 87.2 | 5.2 | 69.6 | 86.72 | 88.26 | 88.3
6.3 | 1 | 5.4 | 83.794 | 85.08 | 85.4414 | 1.3 | 7.9 | 84 | 85.3 | 85.38 | 1.9 | 12.04 | 84.2454 | 85.8 | 86 | 3.5509 | 26.95 | 85.2 | 86.6473 | 86.9 | 5.05 | 63.9167 | 86.46 | 87.975 | 88.1
6.4 | 0.9 | 5.25 | 83.4739 | 84.8 | 85.1613 | 1.3 | 7.6 | 83.7 | 85.02 | 85.18 | 1.9 | 11.6 | 83.9468 | 85.5833 | 85.9 | 3.4509 | 22.5 | 84.94 | 86.4373 | 86.62 | 5 | 56.3 | 86.22 | 87.8 | 87.9833
6.5 | 0.9 | 5.1 | 83.1939 | 84.48 | 84.8612 | 1.2 | 7.3 | 83.4 | 84.8 | 85.08 | 1.8 | 11.0833 | 83.65 | 85.4 | 85.5 | 3.3008 | 19.95 | 84.7667 | 86.2773 | 86.42 | 4.8667 | 49.72 | 85.975 | 87.6 | 87.9
6.6 | 0.9 | 5.075 | 82.9138 | 84.24 | 84.5611 | 1.2 | 7.1 | 83.1 | 84.6 | 84.88 | 1.8 | 10.7 | 83.3166 | 85.2 | 85.3167 | 3.3008 | 18.375 | 84.38 | 86.0672 | 86.22 | 4.8 | 42 | 85.76 | 87.2667 | 87.7
6.7 | 0.9 | 4.9 | 82.6138 | 84.15 | 84.2711 | 1.2 | 6.9 | 82.8 | 84.38 | 84.6833 | 1.8 | 10.4 | 82.9966 | 84.9 | 85.0333 | 3.2008 | 17.4667 | 84.2 | 85.8315 | 85.9333 | 4.7 | 35.6 | 85.575 | 87.1 | 87.4
6.8 | 0.9 | 4.78 | 82.3137 | 83.85 | 84.041 | 1.1 | 6.6 | 82.6 | 84.16 | 84.2833 | 1.7 | 10.1 | 82.7565 | 84.6 | 84.82 | 3.2008 | 16.675 | 83.875 | 85.5771 | 85.72 | 4.6 | 30.2 | 85.3333 | 86.9 | 87.2
6.9 | 0.9 | 4.7 | 82.1137 | 83.62 | 83.8543 | 1.1 | 6.5 | 82.3 | 83.88 | 84 | 1.7 | 9.8 | 82.4915 | 84.3 | 84.54 | 3.1008 | 16.025 | 83.55 | 85.2456 | 85.5333 | 4.5 | 26.1667 | 85.02 | 86.7 | 86.9333
7 | 0.8 | 4.58 | 81.7136 | 83.36 | 83.5609 | 1.1 | 6.3 | 81.925 | 83.6286 | 83.8 | 1.7 | 9.5 | 82.2564 | 84.0333 | 84.32 | 3.0008 | 15.55 | 83.24 | 85.127 | 85.28 | 4.5 | 23.5667 | 84.48 | 86.4 | 86.72
7.1 |  | 4.5 | 81.3136 | 82.92 | 83.3608 |  | 6.1 | 81.625 | 83.26 | 83.6 | 1.6 | 9.2 | 81.7963 | 83.8167 | 84.12 | 3.0008 | 15.225 | 82.84 | 84.797 | 85.025 | 4.4 | 21.8 | 84.28 | 86.2 | 86.52
7.2 |  | 4.4 | 81.1135 | 82.65 | 83.1008 |  | 5.9 | 81.325 | 82.925 | 83.3 | 1.6 | 8.9 | 81.5563 | 83.62 | 83.92 | 2.9007 | 14.7667 | 82.6 | 84.5455 | 84.84 | 4.3 | 20.7333 | 83.96 | 86.1 | 86.3
7.3 |  | 4.3 | 80.8135 | 82.4 | 82.8607 |  | 5.75 | 81.1 | 82.74 | 83 |  | 8.7 | 81.2162 | 83.32 | 83.54 | 2.8007 | 14.4333 | 82.42 | 84.2714 | 84.64 | 4.3 | 19.9667 | 83.68 | 85.9 | 86.1
7.4 |  | 4.1833 | 80.6134 | 82.25 | 82.5039 |  | 5.54 | 80.7667 | 82.56 | 82.7 |  | 8.45 | 80.9162 | 83.02 | 83.3333 |  | 14.1333 | 81.94 | 84.1168 | 84.34 | 4.2 | 19.4667 | 83.5 | 85.6333 | 85.88
7.5 |  | 4.1 | 80.2134 | 81.9667 | 82.3206 |  | 5.4 | 80.4333 | 82.3455 | 82.5 |  | 8.3 | 80.4161 | 82.82 | 83.14 |  | 13.8 | 81.64 | 84.0168 | 84.14 | 4 | 18.8333 | 83.34 | 85.4333 | 85.66
7.6 |  | 4.1 | 79.9333 | 81.7667 | 82.0205 |  | 5.35 | 80.04 | 82.0625 | 82.4 |  | 8 | 80.116 | 82.54 | 82.82 |  | 13.5 | 81.32 | 83.8168 | 83.94 |  | 18.4333 | 83.04 | 85.3167 | 85.3875
7.7 |  | 4 | 79.6133 | 81.5333 | 81.8205 |  | 5.2 | 79.7333 | 81.74 | 82.1167 |  | 7.8 | 79.8731 | 82.22 | 82.7 |  | 13.15 | 81.0333 | 83.5167 | 83.74 |  | 17.95 | 82.725 | 85.1 | 85.3
7.8 |  | 3.9 | 79.3299 | 81.18 | 81.5204 |  | 5.1 | 79.45 | 81.4875 | 81.875 |  | 7.6 | 79.5759 | 82.08 | 82.54 |  | 12.8 | 80.6714 | 83.1166 | 83.54 |  | 17.6 | 82.38 | 84.7 | 85.14
7.9 |  | 3.8 | 79.1332 | 80.88 | 81.3203 |  | 5 | 79.12 | 81.34 | 81.575 |  | 7.4 | 79.3759 | 81.8667 | 82.08 |  | 12.65 | 80.32 | 82.9166 | 83.24 |  | 17.3333 | 82.025 | 84.5 | 84.78
8 |  | 3.8 | 78.7131 | 80.56 | 81.1203 |  | 4.85 | 78.78 | 80.975 | 81.42 |  | 7.3 | 79.0358 | 81.58 | 81.84 |  | 12.3 | 80.08 | 82.7165 | 82.95 |  | 16.9 | 81.7 | 84.3 | 84.54
8.1 |  | 3.7 | 78.3131 | 80.3667 | 80.9203 |  | 4.7 | 78.4667 | 80.7571 | 81.04 |  | 7.1 | 78.549 | 81.4 | 81.68 |  | 12 | 79.6667 | 82.5165 | 82.75 |  | 16.7 | 81.325 | 84.08 | 84.35
8.2 |  | 3.7 | 78.113 | 80.1 | 80.4201 |  | 4.7 | 78.2 | 80.52 | 80.775 |  | 6.9 | 78.2357 | 81.1 | 81.44 |  | 11.8 | 79.2857 | 82.3365 | 82.56 |  | 16.3 | 81.1 | 83.88 | 84.15
8.3 |  | 3.6 | 77.7129 | 79.78 | 80.145 |  | 4.6 | 77.8667 | 80.22 | 80.5857 |  | 6.8 | 77.9156 | 80.85 | 81.24 |  | 11.5 | 79.0143 | 82.1564 | 82.25 |  | 16.1 | 80.86 | 83.74 | 83.9
8.4 |  | 3.5 | 77.4529 | 79.56 | 80.02 |  | 4.5 | 77.5333 | 80.0333 | 80.3333 |  | 6.6 | 77.4726 | 80.58 | 80.88 |  | 11.3 | 78.8 | 81.9764 | 82.1 |  | 15.8 | 80.65 | 83.54 | 83.56
8.5 |  | 3.5 | 77.1129 | 79.28 | 79.6999 |  | 4.4 | 77.2167 | 79.7 | 79.94 |  | 6.5 | 77.1011 | 80.36 | 80.58 |  | 11.1 | 78.5 | 81.5496 | 81.82 |  | 15.4 | 80.34 | 83.225 | 83.36
8.6 |  | 3.4 | 76.8461 | 78.9833 | 79.4999 |  | 4.3 | 76.65 | 79.4 | 79.64 |  | 6.3 | 76.7608 | 80.2 | 80.35 |  | 10.9 | 78.1 | 81.1995 | 81.6333 |  | 15.15 | 79.94 | 82.98 | 83.1
8.7 |  | 3.4 | 76.3794 | 78.8 | 79.1798 |  | 4.2 | 76.3 | 79.2 | 79.42 |  | 6.2 | 76.4153 | 79.9 | 80.12 |  | 10.6 | 77.9 | 81.0662 | 81.275 |  | 14.9167 | 79.72 | 82.8 | 83.06
8.8 |  | 3.4 | 75.9794 | 78.5667 | 78.8864 |  | 4.2 | 76 | 79 | 79.22 |  | 6 | 76.1152 | 79.65 | 79.86 |  | 10.4 | 77.52 | 80.8162 | 81.0571 |  | 14.64 | 79.3 | 82.58 | 82.8
8.9 |  | 3.3 | 75.6876 | 78.4 | 78.7197 |  | 4.1 | 75.6333 | 78.7 | 79.02 |  | 5.9 | 75.5751 | 79.2667 | 79.62 |  | 10.2 | 77.2 | 80.6161 | 80.82 |  | 14.4 | 78.8667 | 82.4 | 82.6
9 |  | 3.3 | 75.3726 | 78 | 78.5196 |  | 4 | 75.3 | 78.5 | 78.8 |  | 5.825 | 75.115 | 79 | 79.36 |  | 10 | 76.825 | 80.3661 | 80.58 |  | 14.1 | 78.58 | 82.14 | 82.38
9.1 |  | 3.2 | 74.9725 | 77.8 | 78.3196 |  | 4 | 74.85 | 78.2714 | 78.5667 |  | 5.725 | 74.935 | 78.925 | 79.16 |  | 9.8 | 76.36 | 80.2661 | 80.375 |  | 13.9 | 78.22 | 81.95 | 82.1286
9.2 |  | 3.2 | 74.5791 | 77.6 | 78.0695 |  | 3.9 | 74.42 | 77.9857 | 78.38 |  | 5.6 | 74.3749 | 78.825 | 78.96 |  | 9.6 | 75.88 | 80.056 | 80.1143 |  | 13.62 | 77.925 | 81.66 | 81.86
9.3 |  | 3.2 | 74.2124 | 77.42 | 77.9195 |  | 3.9 | 74.0333 | 77.74 | 78.16 |  | 5.5 | 74.0548 | 78.4333 | 78.62 |  | 9.425 | 75.64 | 79.7959 | 79.9 |  | 13.42 | 77.74 | 81.46 | 81.58
9.4 |  | 3.1 | 73.7123 | 77.3 | 77.7194 |  | 3.8 | 73.6333 | 77.58 | 77.88 |  | 5.4 | 73.6547 | 78.12 | 78.4667 |  | 9.325 | 75.28 | 79.3358 | 79.6 |  | 13.22 | 77.42 | 81.06 | 81.48
9.5 |  | 3.1 | 72.9121 | 77.1167 | 77.5194 |  | 3.7 | 73.16 | 77.3333 | 77.65 |  | 5.3 | 73.3547 | 77.82 | 78.22 |  | 9.2 | 74.94 | 79.1158 | 79.4 |  | 13.02 | 77.02 | 80.78 | 81.26
9.6 |  | 3 | 71.662 | 76.9 | 77.4194 |  | 3.7 | 72.74 | 77.14 | 77.4333 |  | 5.1333 | 72.7545 | 77.64 | 77.98 |  | 8.925 | 74.5 | 78.9158 | 79.1714 |  | 12.9 | 76.6 | 80.6 | 80.96
9.7 |  | 3 | 69.5116 | 76.82 | 77.2193 |  | 3.6 | 72.06 | 76.9333 | 77.28 |  | 5.1 | 72.3345 | 77.44 | 77.7667 |  | 8.725 | 74.04 | 78.6657 | 79 |  | 12.64 | 76.3 | 80.58 | 80.7
9.8 |  | 3 | 65.2109 | 76.8 | 77.1193 |  | 3.5 | 71.38 | 76.7636 | 77.1 |  | 5 | 71.8644 | 77.14 | 77.58 |  | 8.6 | 73.725 | 78.4824 | 78.8 |  | 12.42 | 76 | 80.36 | 80.5
9.9 |  | 3 | 58.8098 | 76.62 | 77.0193 |  | 3.5 | 70.5 | 76.67 | 77 |  | 4.9 | 71.3743 | 76.98 | 77.36 |  | 8.4333 | 73.3 | 78.2757 | 78.5 |  | 12.125 | 75.6 | 80.18 | 80.2
10 |  | 2.9 | 51.2085 | 76.42 | 76.9192 |  | 3.5 | 69.06 | 76.54 | 76.8286 |  | 4.8 | 70.8742 | 76.74 | 76.9667 |  | 8.3 | 73.025 | 77.9906 | 78.3 |  | 12 | 75.2 | 79.96 | 80
10.1 |  |  | 42.3071 | 76.2875 | 76.7192 |  | 3.4 | 66.52 | 76.34 | 76.7 |  | 4.75 | 70.3541 | 76.6 | 76.7333 |  | 8.1333 | 72.575 | 77.7155 | 78.0333 |  | 11.84 | 74.92 | 79.7 | 79.7
10.2 |  |  | 34.0556 | 76.1571 | 76.5391 |  | 3.4 | 62.1 | 76.2 | 76.48 |  | 4.7 | 69.754 | 76.34 | 76.6 |  | 8 | 72.22 | 77.3155 | 77.74 |  | 11.6 | 74.52 | 79.4 | 79.5
10.3 |  |  | 26.5044 | 76.025 | 76.4991 |  |  | 55.96 | 76.04 | 76.4 |  | 4.6 | 69.0338 | 76.14 | 76.44 |  | 7.9 | 71.78 | 77.1487 | 77.45 |  | 11.45 | 74.2 | 79.1 | 79.3
10.4 |  |  | 20.0433 | 75.94 | 76.3991 |  |  | 48.78 | 75.8778 | 76.2 |  | 4.5667 | 68.1136 | 76.05 | 76.3 |  | 7.8 | 71.25 | 76.9154 | 77.22 |  | 11.2667 | 73.76 | 78.7 | 79.1
10.5 |  |  | 15.7776 | 75.85 | 76.1991 |  |  | 41.1 | 75.7667 | 76.04 |  | 4.5 | 66.5933 | 75.8 | 76.15 |  | 7.6667 | 70.82 | 76.7653 | 77.02 |  | 11.2 | 73.46 | 78.5167 | 78.8333
10.6 |  |  | 13.2772 | 75.7 | 76.099 |  |  | 33.98 | 75.7 | 75.9 |  |  | 64.2128 | 75.66 | 75.96 |  | 7.5333 | 70.2833 | 76.4153 | 76.78 |  | 10.9667 | 73.1 | 78.3 | 78.62
10.7 |  |  | 12.002 | 75.5 | 75.9023 |  |  | 27.08 | 75.5333 | 75.8 |  |  | 60.4321 | 75.56 | 75.82 |  | 7.4 | 70 | 76.2819 | 76.55 |  | 10.775 | 72.68 | 78.02 | 78.34
10.8 |  |  | 11.3519 | 75.4 | 75.7189 |  |  | 21.56 | 75.32 | 75.6286 |  |  | 55.011 | 75.36 | 75.6 |  | 7.3 | 69.54 | 76.0152 | 76.3286 |  | 10.6 | 72.4333 | 77.875 | 78.12
10.9 |  |  | 10.9018 | 75.3 | 75.5189 |  |  | 17.74 | 75.12 | 75.5 |  |  | 48.6098 | 75.2 | 75.46 |  | 7.2 | 69.2 | 75.8152 | 76.2182 |  | 10.5 | 71.92 | 77.5375 | 78.02
11 |  |  | 10.6268 | 75.1 | 75.4189 |  |  | 15.56 | 74.92 | 75.35 |  |  | 41.7484 | 75.075 | 75.3 |  | 7 | 68.5 | 75.6151 | 75.96 |  | 10.3 | 71.52 | 77.32 | 77.9
11.1 |  |  | 10.3517 | 74.9 | 75.2855 |  |  | 14.32 | 74.72 | 75.1 |  |  | 35.1671 | 74.86 | 75.16 |  |  | 67.9 | 75.4551 | 75.66 |  | 10.2 | 71.2 | 77.2182 | 77.62
11.2 |  |  | 10.1017 | 74.8 | 75.0188 |  |  | 13.54 | 74.58 | 75.02 |  |  | 28.9658 | 74.7 | 74.96 |  |  | 67.4 | 75.255 | 75.52 |  | 10.1 | 70.8667 | 76.96 | 77.22
11.3 |  |  | 9.935 | 74.7091 | 75.0188 |  |  | 13.1 | 74.44 | 74.82 |  |  | 24.0248 | 74.4857 | 74.76 |  |  | 66.6 | 75.04 | 75.34 |  | 9.9333 | 70.26 | 76.66 | 76.8667
11.4 |  |  | 9.8216 | 74.6 | 74.8854 |  |  | 12.68 | 74.34 | 74.72 |  |  | 20.6208 | 74.32 | 74.6667 |  |  | 65.7 | 74.84 | 75.08 |  | 9.7 | 69.84 | 76.44 | 76.68
11.5 |  |  | 9.6016 | 74.4182 | 74.6587 |  |  | 12.34 | 74.24 | 74.52 |  |  | 18.7037 | 74.12 | 74.4667 |  |  | 64.3 | 74.6149 | 74.8222 |  | 9.8 | 69.54 | 76.24 | 76.5333
11.6 |  |  | 9.5216 | 74.22 | 74.4586 |  |  | 12.04 | 74.04 | 74.32 |  |  | 17.5035 | 73.98 | 74.28 |  |  | 62.2 | 74.4149 | 74.64 |  |  | 68.94 | 76.04 | 76.3333
11.7 |  |  | 9.4016 | 74.1 | 74.3786 |  |  | 11.78 | 73.94 | 74.2333 |  |  | 16.7033 | 73.8143 | 74.08 |  |  | 58.9 | 74.1148 | 74.44 |  |  | 68.5667 | 75.84 | 76.08
11.8 |  |  | 9.3016 | 73.92 | 74.1786 |  |  | 11.5 | 73.74 | 74.04 |  |  | 16.1032 | 73.7333 | 73.875 |  |  | 54.5 | 74.0148 | 74.24 |  |  | 68.175 | 75.6 | 75.84
11.9 |  |  | 9.1215 | 73.7636 | 74.0785 |  |  | 11.3 | 73.54 | 73.78 |  |  | 15.6031 | 73.5273 | 73.7857 |  |  | 49.2 | 73.8148 | 74.04 |  |  | 67.7667 | 75.32 | 75.64
12 |  |  | 9.0215 | 73.65 | 73.9185 |  |  | 11.08 | 73.4 | 73.7 |  |  | 15.203 | 73.31 | 73.6 |  |  | 43.7 | 73.6147 | 73.94 |  |  | 67.34 | 75.06 | 75.54
12.1 |  |  | 8.9015 | 73.42 | 73.7785 |  |  | 11 | 73.2 | 73.54 |  |  | 14.903 | 73.04 | 73.4 |  |  | 37.9 | 73.3147 | 73.8 |  |  | 66.95 | 74.86 | 75.34
12.2 |  |  | 8.7615 | 73.2 | 73.6184 |  |  | 10.8 | 73.1 | 73.34 |  |  | 14.6029 | 72.84 | 73.3 |  |  | 32.8 | 73.2146 | 73.66 |  |  | 66.46 | 74.66 | 75.1
12.3 |  |  | 8.6614 | 73.1 | 73.4898 |  |  | 10.6 | 73 | 73.14 |  |  | 14.4229 | 72.6625 | 73.08 |  |  | 28.7 | 73.0146 | 73.46 |  |  | 65.72 | 74.58 | 74.86
12.4 |  |  | 8.5614 | 73.025 | 73.2516 |  |  | 10.4 | 72.7667 | 73.04 |  |  | 14.1028 | 72.4667 | 72.9 |  |  | 25.5 | 72.8396 | 73.15 |  |  | 64.6 | 74.38 | 74.66
12.5 |  |  | 8.4614 | 72.8 | 73.1183 |  |  | 10.25 | 72.575 | 72.86 |  |  | 13.8361 | 72.24 | 72.6833 |  |  | 23.65 | 72.6745 | 72.96 |  |  | 63.24 | 74.0833 | 74.5333
12.6 |  |  | 8.3814 | 72.775 | 72.9808 |  |  | 10.06 | 72.42 | 72.66 |  |  | 13.5227 | 72.04 | 72.4667 |  |  | 22.3 | 72.5145 | 72.7333 |  |  | 61.175 | 73.9 | 74.3
12.7 |  |  | 8.2764 | 72.5 | 72.8982 |  |  | 9.85 | 72.2333 | 72.46 |  |  | 13.2426 | 71.94 | 72.2286 |  |  | 21.4 | 72.2644 | 72.44 |  |  | 58.6 | 73.7833 | 74.1
12.8 |  |  | 8.1814 | 72.4 | 72.7432 |  |  | 9.74 | 72.08 | 72.26 |  |  | 13.1226 | 71.74 | 72.04 |  |  | 20.8 | 72.1144 | 72.22 |  |  | 54.7 | 73.6 | 73.92
12.9 |  |  | 8.0813 | 72.3 | 72.5581 |  |  | 9.6 | 71.9 | 72.06 |  |  | 12.8426 | 71.56 | 71.9 |  |  | 20.4167 | 71.8544 | 72.12 |  |  | 50.025 | 73.3 | 73.72
13 |  |  | 7.9813 | 72.075 | 72.3752 |  |  | 9.44 | 71.74 | 71.9667 |  |  | 12.6358 | 71.36 | 71.7429 |  |  | 20.04 | 71.7343 | 72 |  |  | 45.02 | 73.1 | 73.52
13.1 |  |  | 7.8813 | 71.92 | 72.2181 |  |  | 9.3 | 71.54 | 71.675 |  |  | 12.3692 | 71.1 | 71.5167 |  |  | 19.62 | 71.3943 | 71.84 |  |  | 40.28 | 72.9167 | 73.22
13.2 |  |  | 7.7846 | 71.8 | 72.118 |  |  | 9.2 | 71.34 | 71.5667 |  |  | 12.2024 | 70.9167 | 71.3 |  |  | 19.3167 | 71.1142 | 71.48 |  |  | 35.72 | 72.8 | 73.12
13.3 |  |  | 7.7013 | 71.6333 | 71.9805 |  |  | 9.05 | 71.24 | 71.4 |  |  | 12.0224 | 70.7 | 71.1 |  |  | 18.94 | 70.9975 | 71.3 |  |  | 32 | 72.6 | 72.9333
13.4 |  |  | 7.6013 | 71.5 | 71.918 |  |  | 8.96 | 70.94 | 71.22 |  |  | 11.8224 | 70.5333 | 70.8 |  |  | 18.6 | 70.8975 | 71.14 |  |  | 29.14 | 72.375 | 72.7333
13.5 |  |  | 7.5012 | 71.3333 | 71.6179 |  |  | 8.8 | 70.84 | 71.0333 |  |  | 11.6223 | 70.3714 | 70.6 |  |  | 18.2333 | 70.6141 | 70.88 |  |  | 27.3 | 72.16 | 72.4667
13.6 |  |  | 7.4612 | 71.15 | 71.4179 |  |  | 8.7 | 70.625 | 70.8 |  |  | 11.5023 | 70.2 | 70.4 |  |  | 17.94 | 70.4141 | 70.58 |  |  | 25.9 | 71.96 | 72.3
13.7 |  |  | 7.4012 | 70.95 | 71.3178 |  |  | 8.66 | 70.5 | 70.68 |  |  | 11.2422 | 69.9571 | 70.2 |  |  | 17.66 | 70.214 | 70.44 |  |  | 25 | 71.75 | 72.14
13.8 |  |  | 7.3012 | 70.75 | 71.1178 |  |  | 8.46 | 70.35 | 70.4 |  |  | 11.1355 | 69.75 | 70 |  |  | 17.45 | 70.0473 | 70.24 |  |  | 24.4 | 71.4 | 71.94
13.9 |  |  | 7.3012 | 70.55 | 70.8177 |  |  | 8.36 | 70.1 | 70.2 |  |  | 10.9422 | 69.46 | 69.775 |  |  | 17.14 | 69.8473 | 70 |  |  | 23.8 | 71.275 | 71.54
14 |  |  | 7.2012 | 70.4 | 70.6177 |  |  | 8.26 | 69.76 | 70 |  |  | 10.7421 | 69.22 | 69.52 |  |  | 16.9667 | 69.6339 | 69.75 |  |  | 23.5333 | 71.1 | 71.3
14.1 |  |  | 7.1012 | 70.2 | 70.4176 |  |  | 8.1667 | 69.6 | 69.8 |  |  | 10.6271 | 69.06 | 69.26 |  |  | 16.7667 | 69.3139 | 69.55 |  |  | 23.1667 | 70.8667 | 71.16
14.2 |  |  | 7.0012 | 69.85 | 70.2176 |  |  | 8.1 | 69.4 | 69.5833 |  |  | 10.5021 | 68.82 | 69.1333 |  |  | 16.4667 | 69.0538 | 69.2 |  |  | 22.7333 | 70.66 | 71.06
14.3 |  |  | 6.9262 | 69.55 | 69.8175 |  |  | 8 | 69.1 | 69.3 |  |  | 10.242 | 68.5667 | 68.84 |  |  | 16.22 | 68.9138 | 69 |  |  | 22.5 | 70.4 | 70.82
14.4 |  |  | 6.9012 | 69.22 | 69.6508 |  |  | 7.9 | 68.8857 | 69.1 |  |  | 10.162 | 68.36 | 68.62 |  |  | 15.9667 | 68.6137 | 68.7333 |  |  | 22.2 | 70.2 | 70.62
14.5 |  |  | 6.8011 | 68.65 | 69.0773 |  |  | 7.8 | 68.72 | 68.8 |  |  | 10.042 | 68.16 | 68.42 |  |  | 15.78 | 68.3137 | 68.6 |  |  | 21.8333 | 69.975 | 70.4
14.6 |  |  | 6.7411 | 67.96 | 68.3971 |  |  | 7.7 | 68.34 | 68.5667 |  |  | 9.902 | 67.78 | 68.04 |  |  | 15.6 | 68.1736 | 68.28 |  |  | 21.46 | 69.7667 | 70.2
14.7 |  |  | 6.7011 | 66.625 | 66.9167 |  |  | 7.6 | 68.02 | 68.2 |  |  | 9.762 | 67.56 | 67.82 |  |  | 15.3 | 67.8803 | 68.1 |  |  | 21.3 | 69.58 | 69.84
14.8 |  |  | 6.6011 | 63.9 | 64.5161 |  |  | 7.5333 | 67.62 | 67.9 |  |  | 9.6619 | 67.38 | 67.5667 |  |  | 15 | 67.6024 | 67.88 |  |  | 21 | 69.34 | 69.56
14.9 |  |  | 6.6011 | 59.7 | 60.2351 |  |  | 7.5 | 67.2667 | 67.4 |  |  | 9.5419 | 67.14 | 67.325 |  |  | 14.8 | 67.4735 | 67.58 |  |  | 20.65 | 69.14 | 69.4
15 |  |  | 6.5011 | 53.6 | 54.3469 |  |  | 7.4 | 66.76 | 66.9 |  |  | 9.4019 | 66.68 | 67.04 |  |  | 14.6 | 67.2535 | 67.46 |  |  | 20.4 | 68.9 | 69.3
15.1 |  |  |  | 46.5333 | 47.6119 |  |  | 7.3 | 65.95 | 66.2 |  |  | 9.3185 | 66.46 | 66.7 |  |  | 14.4 | 66.9801 | 67.14 |  |  | 20.16 | 68.7 | 69.12
15.2 |  |  |  | 39.36 | 40.2434 |  |  | 7.3 | 65 | 65.3 |  |  | 9.2018 | 66.16 | 66.3 |  |  | 14.3375 | 66.7801 | 66.7833 |  |  | 19.92 | 68.48 | 68.74
15.3 |  |  |  | 32.12 | 33.3417 |  |  | 7.2 | 63.48 | 63.66 |  |  | 9.1018 | 65.8 | 66.1 |  |  | 14.1 | 66.4466 | 66.7 |  |  | 19.66 | 68.1 | 68.44
15.4 |  |  |  | 25.7 | 26.9067 |  |  | 7.1 | 60.3 | 60.84 |  |  | 9.0018 | 65.5 | 65.6 |  |  | 13.9333 | 66.1418 | 66.4 |  |  | 19.46 | 67.8333 | 68.2
15.5 |  |  |  | 20.4667 | 22.0055 |  |  |  | 56.02 | 56.6 |  |  | 8.9018 | 65.1 | 65.34 |  |  | 13.74 | 65.9532 | 66.1 |  |  | 19.1333 | 67.625 | 68
15.6 |  |  |  | 17.42 | 19.1381 |  |  |  | 50.4667 | 50.72 |  |  | 8.7818 | 64.6 | 64.96 |  |  | 13.48 | 65.6331 | 65.9167 |  |  | 18.92 | 67.45 | 67.88
15.7 |  |  |  | 15.74 | 17.5644 |  |  |  | 43.9 | 44.82 |  |  | 8.6517 | 64.2 | 64.44 |  |  | 13.3 | 65.3531 | 65.6 |  |  | 18.7 | 67.3125 | 67.58
15.8 |  |  |  | 14.8 | 16.5842 |  |  |  | 37.3333 | 38.4167 |  |  | 8.5017 | 63.6 | 63.86 |  |  | 13.2 | 65.093 | 65.4 |  |  | 18.4667 | 67.12 | 67.28
15.9 |  |  |  | 14.2667 | 16.104 |  |  |  | 30.9667 | 32.4286 |  |  | 8.4017 | 62.7 | 63.08 |  |  | 13.1 | 64.833 | 65.1429 |  |  | 18.2 | 66.8333 | 67.08
16 |  |  |  | 14 | 15.7439 |  |  |  | 25.58 | 27 |  |  |  | 61.5 | 61.76 |  |  | 12.85 | 64.5329 | 64.82 |  |  | 18 | 66.54 | 66.94
16.1 |  |  |  | 13.7 | 15.5239 |  |  |  | 21.5667 | 23.025 |  |  |  | 59.4 | 59.94 |  |  | 12.7 | 64.2729 | 64.46 |  |  | 17.775 | 66.325 | 66.6429
16.2 |  |  |  | 13.5 | 15.3238 |  |  |  | 19 | 20.5857 |  |  |  | 56.1 | 56.68 |  |  | 12.6 | 64.0328 | 64.2455 |  |  | 17.55 | 66.0333 | 66.375
16.3 |  |  |  | 13.3 | 15.0788 |  |  |  | 17.54 | 19.24 |  |  |  | 51.7 | 52.22 |  |  | 12.425 | 63.6794 | 63.97 |  |  | 17.375 | 65.78 | 66.09
16.4 |  |  |  | 13.2 | 15.0038 |  |  |  | 16.72 | 18.38 |  |  |  | 46.5833 | 47.08 |  |  | 12.2333 | 63.3727 | 63.67 |  |  | 17.22 | 65.58 | 65.79
16.5 |  |  |  | 13.1 | 14.8037 |  |  |  | 16.12 | 17.875 |  |  |  | 40.6857 | 41.64 |  |  | 12.1 | 63.0681 | 63.36 |  |  | 17.02 | 65.3 | 65.56
16.6 |  |  |  | 12.9 | 14.7037 |  |  |  | 15.75 | 17.5 |  |  |  | 34.95 | 36.16 |  |  | 11.9 | 62.658 | 63.02 |  |  | 16.8333 | 65.18 | 65.28
16.7 |  |  |  | 12.8 | 14.6037 |  |  |  | 15.46 | 17.2 |  |  |  | 29.8 | 31.125 |  |  | 11.8 | 62.3825 | 62.72 |  |  | 16.7 | 64.95 | 65.15
16.8 |  |  |  | 12.7 | 14.5036 |  |  |  | 15.16 | 16.8 |  |  |  | 25.8182 | 27.2 |  |  | 11.7 | 62.0624 | 62.36 |  |  | 16.4333 | 64.5 | 64.92
16.9 |  |  |  | 12.6 | 14.3036 |  |  |  | 14.96 | 16.6 |  |  |  | 22.84 | 24.25 |  |  | 11.6 | 61.6123 | 62.02 |  |  | 16.25 | 64.3 | 64.54
17 |  |  |  | 12.5 | 14.2036 |  |  |  | 14.76 | 16.5 |  |  |  | 20.9 | 22.5 |  |  |  | 61.3123 | 61.7 |  |  | 16.06 | 63.98 | 64.32
17.1 |  |  |  | 12.35 | 14.0835 |  |  |  | 14.56 | 16.2 |  |  |  | 19.76 | 21.45 |  |  |  | 60.9122 | 61.2 |  |  | 15.8667 | 63.6833 | 63.96
17.2 |  |  |  | 12.3 | 14.0035 |  |  |  | 14.36 | 16 |  |  |  | 19.04 | 20.675 |  |  |  | 60.3621 | 60.75 |  |  | 15.68 | 63.5 | 63.72
17.3 |  |  |  | 12.16 | 13.9035 |  |  |  | 14.16 | 15.85 |  |  |  | 18.5 | 20.1857 |  |  |  | 59.7231 | 60.1 |  |  | 15.46 | 63.3 | 63.44
17.4 |  |  |  | 12.06 | 13.8035 |  |  |  | 13.96 | 15.7 |  |  |  | 18.12 | 19.8375 |  |  |  | 58.8718 | 59.4167 |  |  | 15.38 | 63 | 63.2
17.5 |  |  |  | 12 | 13.7034 |  |  |  | 13.86 | 15.5 |  |  |  | 17.82 | 19.44 |  |  |  | 57.8116 | 58.2222 |  |  | 15.18 | 62.7 | 63.025
17.6 |  |  |  | 11.9 | 13.6034 |  |  |  | 13.66 | 15.3 |  |  |  | 17.475 | 19.16 |  |  |  | 56.0737 | 56.4667 |  |  | 15.0143 | 62.3667 | 62.6667
17.7 |  |  |  | 11.86 | 13.5034 |  |  |  | 13.46 | 15.2091 |  |  |  | 17.24 | 18.94 |  |  |  | 53.4907 | 53.8167 |  |  | 14.9 | 61.88 | 62.4333
17.8 |  |  |  | 11.7 | 13.4034 |  |  |  | 13.3 | 15.02 |  |  |  | 16.95 | 18.6455 |  |  |  | 49.9099 | 50.4 |  |  | 14.72 | 61.6 | 62.1333
17.9 |  |  |  | 11.66 | 13.3433 |  |  |  | 13.25 | 14.91 |  |  |  | 16.72 | 18.38 |  |  |  | 45.6491 | 46.3 |  |  | 14.4667 | 61.4 | 61.72
18 |  |  |  | 11.6 | 13.3033 |  |  |  | 13.0667 | 14.72 |  |  |  | 16.46 | 18.1857 |  |  |  | 41.0282 | 41.8 |  |  |  | 61.1 | 61.4333
18.1 |  |  |  | 11.5 | 13.2033 |  |  |  | 12.9 | 14.61 |  |  |  | 16.3 | 18 |  |  |  | 36.4873 | 37.4 |  |  |  | 60.8 | 61.15
18.2 |  |  |  | 11.4222 | 13.1033 |  |  |  | 12.85 | 14.4857 |  |  |  | 16.06 | 17.78 |  |  |  | 32.2064 | 33.6 |  |  |  | 60.5 | 60.8
18.3 |  |  |  | 11.4 | 13.1033 |  |  |  | 12.7 | 14.36 |  |  |  | 15.86 | 17.58 |  |  |  | 28.8258 | 30.2 |  |  |  | 60.1 | 60.4
18.4 |  |  |  | 11.3 | 13.0033 |  |  |  | 12.6 | 14.26 |  |  |  | 15.66 | 17.38 |  |  |  | 26.1652 | 27.64 |  |  |  | 59.7 | 60.1667
18.5 |  |  |  | 11.2 | 12.9032 |  |  |  | 12.46 | 14.16 |  |  |  | 15.46 | 17.18 |  |  |  | 24.5449 | 26.12 |  |  |  | 59.45 | 59.94
18.6 |  |  |  | 11.2 | 12.8032 |  |  |  | 12.38 | 14.02 |  |  |  | 15.26 | 16.98 |  |  |  | 23.4047 | 25.08 |  |  |  | 59 | 59.4
18.7 |  |  |  | 11.1 | 12.8032 |  |  |  | 12.28 | 13.9 |  |  |  | 15.1 | 16.78 |  |  |  | 22.7795 | 24.46 |  |  |  | 58.5333 | 58.8667
18.8 |  |  |  | 11.1 | 12.7032 |  |  |  | 12.18 | 13.8333 |  |  |  | 14.95 | 16.58 |  |  |  | 22.3244 | 24.04 |  |  |  | 57.8 | 58.44
18.9 |  |  |  | 11 | 12.7032 |  |  |  | 12.08 | 13.7 |  |  |  | 14.68 | 16.44 |  |  |  | 21.9711 | 23.62 |  |  |  | 57 | 57.7
19 |  |  |  | 10.9 | 12.6032 |  |  |  | 11.9833 | 13.6 |  |  |  | 14.5667 | 16.22 |  |  |  | 21.5443 | 23.24 |  |  |  | 55.96 | 56.58
19.1 |  |  |  | 10.88 | 12.5031 |  |  |  | 11.9 | 13.54 |  |  |  | 14.4 | 16.0333 |  |  |  | 21.1442 | 22.94 |  |  |  | 54.45 | 55.12
19.2 |  |  |  | 10.8 | 12.4031 |  |  |  | 11.8 | 13.4 |  |  |  | 14.28 | 15.9333 |  |  |  | 20.9442 | 22.62 |  |  |  | 52.08 | 52.92
19.3 |  |  |  | 10.7 | 12.3031 |  |  |  | 11.7 | 13.3 |  |  |  | 14.08 | 15.8 |  |  |  | 20.6441 | 22.38 |  |  |  | 49 | 50.02
19.4 |  |  |  | 10.6 | 12.2698 |  |  |  | 11.6 | 13.2 |  |  |  | 13.98 | 15.6 |  |  |  | 20.4441 | 22.04 |  |  |  | 45.24 | 46.5
19.5 |  |  |  | 10.5 | 12.1697 |  |  |  | 11.55 | 13.1 |  |  |  | 13.88 | 15.5 |  |  |  | 20.084 | 21.85 |  |  |  | 41.75 | 42.92
19.6 |  |  |  | 10.48 | 12.103 |  |  |  | 11.5 | 13.1 |  |  |  | 13.6833 | 15.3375 |  |  |  | 19.844 | 21.6 |  |  |  | 37.94 | 39.12
19.7 |  |  |  | 10.4 | 12.003 |  |  |  | 11.4 | 13 |  |  |  | 13.5 | 15.2 |  |  |  | 19.6439 | 21.3429 |  |  |  | 34.1 | 35.74
19.8 |  |  |  | 10.3 | 11.903 |  |  |  | 11.275 | 12.9 |  |  |  | 13.4 | 15.1 |  |  |  | 19.3639 | 21.075 |  |  |  | 31.14 | 32.92
19.9 |  |  |  | 10.28 | 11.803 |  |  |  | 11.2 | 12.8 |  |  |  | 13.3 | 15 |  |  |  | 19.2038 | 20.9 |  |  |  | 29 | 30.66
20 |  |  |  | 10.18 | 11.6529 |  |  |  | 11.1 | 12.7 |  |  |  | 13.14 | 14.8 |  |  |  | 19.0038 | 20.72 |  |  |  | 27.5 | 29.24
20.1 |  |  |  | 10.08 | 11.6029 |  |  |  | 11 | 12.7 |  |  |  | 13.04 | 14.7 |  |  |  | 18.7323 | 20.52 |  |  |  | 26.55 | 28.36
20.2 | [8 empty] | 10.9333 | 12.5 |  |  |  | 12.9 | 14.6 |  |  |  | 18.4894 | 20.28 |  |  |  | 25.76 | 27.7
20.3 | [8 empty] | 10.8 | 12.42 |  |  |  | 12.84 | 14.4 |  |  |  | 18.3237 | 20.05 |  |  |  | 25.375 | 27.1
20.4 | [8 empty] | 10.78 | 12.3 |  |  |  | 12.64 | 14.3 |  |  |  | 18.1236 | 19.8286 |  |  |  | 25 | 26.7
20.5 | [8 empty] | 10.7 | 12.28 |  |  |  | 12.6 | 14.2167 |  |  |  | 17.9236 | 19.71 |  |  |  | 24.7 | 26.4
20.6 | [8 empty] | 10.6 | 12.1 |  |  |  | 12.5 | 14.1 |  |  |  | 17.7035 | 19.43 |  |  |  | 24.25 | 26.2
20.7 | [8 empty] | 10.5 | 12 |  |  |  | 12.3 | 14 |  |  |  | 17.4892 | 19.22 |  |  |  | 23.98 | 25.8
20.8 | [13 empty] | 12.2 | 13.9 |  |  |  | 17.3146 | 19.04 |  |  |  | 23.8 | 25.5
20.9 | [13 empty] | 12.175 | 13.8 |  |  |  | 17.1034 | 18.9111 |  |  |  | 23.38 | 25.3
21 | [13 empty] | 12.06 | 13.68 |  |  |  | 16.8867 | 18.72 |  |  |  | 23.0667 | 25
21.1 | [13 empty] | 11.9 | 13.58 |  |  |  | 16.7033 | 18.44 |  |  |  | 22.8 | 24.8429
21.2 | [13 empty] | 11.8 | 13.5 |  |  |  | 16.5033 | 18.34 |  |  |  | 22.7 | 24.52
21.3 | [13 empty] | 11.8 | 13.38 |  |  |  | 16.4033 | 18.04 |  |  |  | 22.4667 | 24.26
21.4 | [13 empty] | 11.6833 | 13.3 |  |  |  | 16.2199 | 17.94 |  |  |  | 22.3 | 24.06
21.5 | [18 empty] | 16.1232 | 17.84 |  |  |  | 22.0667 | 23.92
21.6 | [18 empty] | 15.9232 | 17.66 |  |  |  | 21.84 | 23.6667
21.7 | [18 empty] | 15.8032 | 17.45 |  |  |  | 21.64 | 23.42

## Wedge 30 Longitudinal Profiles
Longitudinal Profiles
Wedge 30
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-34.8 |  |  |  |  | 15.2222
-34.7 |  |  |  |  | 15.3375
-34.6 |  |  |  |  | 15.4556
-34.5 |  |  |  |  | 15.5
-34.4 |  |  |  |  | 15.6143
-34.3 |  |  |  |  | 15.75
-34.2 |  |  |  |  | 15.8857
-34.1 |  |  |  |  | 15.925
-34 |  |  |  |  | 16.0444
-33.9 |  |  |  |  | 16.1714
-33.8 |  |  |  |  | 16.3
-33.7 |  |  |  |  | 16.4
-33.6 |  |  |  |  | 16.55
-33.5 |  |  |  |  | 16.6
-33.4 |  |  |  |  | 16.74
-33.3 |  |  |  |  | 16.9
-33.2 |  |  |  |  | 17
-33.1 |  |  |  |  | 17.1833
-33 |  |  |  |  | 17.25
-32.9 |  |  |  |  | 17.4571
-32.8 |  |  |  |  | 17.5875
-32.7 |  |  |  |  | 17.725
-32.6 |  |  |  |  | 17.85
-32.5 |  |  |  |  | 17.9
-32.4 |  |  |  |  | 18.0143
-32.3 |  |  |  |  | 18.15
-32.2 |  |  |  |  | 18.35
-32.1 |  |  |  | 13.2 | 18.525
-32 |  |  |  | 13.36 | 18.7
-31.9 |  |  |  | 13.5 | 18.7857
-31.8 |  |  |  | 13.6 | 18.98
-31.7 |  |  |  | 13.7 | 19.1429
-31.6 |  |  |  | 13.7 | 19.3
-31.5 |  |  |  | 13.8286 | 19.4667
-31.4 |  |  |  | 13.9625 | 19.5667
-31.3 |  |  |  | 14.0875 | 19.7
-31.2 |  |  |  | 14.2 | 19.9
-31.1 |  |  |  | 14.3 | 20.1
-31 |  |  |  | 14.3833 | 20.2714
-30.9 |  |  |  | 14.5143 | 20.4143
-30.8 |  |  |  | 14.6571 | 20.55
-30.7 |  |  |  | 14.7 | 20.6857
-30.6 |  |  |  | 14.8143 | 20.9125
-30.5 |  |  |  | 14.95 | 21.0429
-30.4 |  |  |  | 15.075 | 21.175
-30.3 |  |  |  | 15.3 | 21.425
-30.2 |  |  |  | 15.3429 | 21.6375
-30.1 |  |  |  | 15.55 | 21.8429
-30 |  |  |  | 15.6125 | 22.0125
-29.9 |  |  |  | 15.7375 | 22.1375
-29.8 |  |  |  | 15.8714 | 22.3429
-29.7 |  |  |  | 16 | 22.5
-29.6 |  |  |  | 16.175 | 22.775
-29.5 |  |  |  | 16.3714 | 23.0429
-29.4 |  |  | 11.4 | 16.5 | 23.2
-29.3 |  |  | 11.525 | 16.6286 | 23.475
-29.2 |  |  | 11.6 | 16.7625 | 23.6714
-29.1 |  |  | 11.675 | 16.9 | 23.9
-29 |  |  | 11.8 | 17.075 | 24.05
-28.9 |  |  | 11.8444 | 17.2556 | 24.3
-28.8 |  |  | 11.9455 | 17.4 | 24.5429
-28.7 |  |  | 12.0444 | 17.5889 | 24.7167
-28.6 |  |  | 12.1625 | 17.7769 | 24.9
-28.5 |  |  | 12.2583 | 17.9267 | 25.3
-28.4 |  |  | 12.3 | 18.06 | 25.4
-28.3 |  |  | 12.4 | 18.2167 | 25.6
-28.2 |  |  | 12.5 | 18.3714 | 25.9
-28.1 |  | 10.54 | 12.5714 | 18.6 | 26.2429
-28 |  | 10.6286 | 12.7 | 18.7375 | 26.5
-27.9 |  | 10.7 | 12.8 | 18.8714 | 26.7571
-27.8 |  | 10.78 | 12.85 | 19.1182 | 27.025
-27.7 |  | 10.8286 | 12.9714 | 19.3556 | 27.3
-27.6 |  | 10.9 | 13.1 | 19.5 | 27.6571
-27.5 |  | 10.9875 | 13.23 | 19.7308 | 27.9875
-27.4 |  | 11.0286 | 13.3375 | 19.94 | 28.275
-27.3 |  | 11.1 | 13.4625 | 20.1571 | 28.55
-27.2 |  | 11.2 | 13.5875 | 20.4429 | 28.8714
-27.1 |  | 11.2222 | 13.725 | 20.5889 | 29.2625
-27 | 10.2125 | 11.3 | 13.8 | 20.8154 | 29.65
-26.9 | 10.3 | 11.375 | 13.8875 | 20.9692 | 29.9714
-26.8 | 10.3 | 11.5 | 14.025 | 21.2667 | 30.3625
-26.7 | 10.4 | 11.5429 | 14.1571 | 21.45 | 30.9714
-26.6 | 10.4375 | 11.6 | 14.2875 | 21.7143 | 31.6857
-26.5 | 10.5 | 11.7125 | 14.425 | 21.975 | 32.4286
-26.4 | 10.6 | 11.8429 | 14.5571 | 22.25 | 33.8
-26.3 | 10.6 | 11.9 | 14.6875 | 22.4571 | 35.54
-26.2 | 10.6133 | 11.9125 | 14.85 | 22.7625 | 37.68
-26.1 | 10.68 | 12.0429 | 15.0571 | 23.05 | 40.75
-26 | 10.7467 | 12.175 | 15.1875 | 23.3143 | 44.05
-25.9 | 10.8133 | 12.2 | 15.25 | 23.575 | 48.1429
-25.8 | 10.88 | 12.3286 | 15.4571 | 23.825 | 52.5286
-25.7 | 10.9467 | 12.4714 | 15.5875 | 24.0857 | 56.75
-25.6 | 11.0133 | 12.5167 | 15.7286 | 24.425 | 60.8857
-25.5 | 11.08 | 12.6714 | 15.95 | 24.7375 | 64.275
-25.4 | 11.1437 | 12.7 | 16.0857 | 25.1286 | 67.0143
-25.3 | 11.2077 | 12.8286 | 16.2125 | 25.375 | 69
-25.2 | 11.2846 | 13 | 16.3857 | 25.7375 | 70.6
-25.1 | 11.4 | 13.1 | 16.575 | 26.1286 | 71.6625
-25 | 11.5 | 13.2 | 16.8125 | 26.45 | 72.4571
-24.9 | 11.5333 | 13.3 | 16.9857 | 26.85 | 73
-24.8 | 11.6 | 13.4667 | 17.175 | 27.35 | 73.575
-24.7 | 11.6875 | 13.56 | 17.4125 | 27.9571 | 73.9857
-24.6 | 11.725 | 13.66 | 17.5857 | 28.7 | 74.4
-24.5 | 11.8571 | 13.76 | 17.775 | 29.775 | 74.775
-24.4 | 11.9875 | 13.84 | 18.1 | 31.2571 | 75.1143
-24.3 | 12.025 | 13.94 | 18.3375 | 33.2 | 75.5
-24.2 | 12.1364 | 14.1333 | 18.5429 | 36.0429 | 75.8125
-24.1 | 12.2333 | 14.275 | 18.7 | 39.7125 | 76.2143
-24 | 12.3308 | 14.4 | 19.0286 | 44.3 | 76.6
-23.9 | 12.4067 | 14.5429 | 19.225 | 49.2429 | 76.9375
-23.8 | 12.4733 | 14.7 | 19.5 | 54.6625 | 77.125
-23.7 | 12.55 | 14.9375 | 19.7571 | 60.2 | 77.5
-23.6 | 12.65 | 15.0714 | 20.025 | 64.625 | 77.7857
-23.5 | 12.78 | 15.2 | 20.4 | 68.7 | 78.1875
-23.4 | 12.9067 | 15.375 | 20.6571 | 71.7375 | 78.5
-23.3 | 12.9733 | 15.5714 | 20.9875 | 74.05 | 78.65
-23.2 | 13.0429 | 15.8 | 21.4 | 75.5 | 78.9714
-23.1 | 13.1364 | 15.9375 | 21.7571 | 76.5 | 79.3625
-23 | 13.3067 | 16.1429 | 22.0875 | 77.35 | 79.65
-22.9 | 13.3733 | 16.375 | 22.5 | 77.8286 | 79.9143
-22.8 | 13.4667 | 16.6429 | 22.9429 | 78.375 | 80.175
-22.7 | 13.6273 | 16.8 | 23.675 | 78.9417 | 80.45
-22.6 | 13.775 | 17.0571 | 24.5 | 79.3583 | 80.7143
-22.5 | 13.9 | 17.2714 | 25.7143 | 79.64 | 81.0625
-22.4 | 14.0125 | 17.55 | 27.5625 | 79.96 | 81.275
-22.3 | 14.1857 | 17.8143 | 30.4 | 80.36 | 81.5571
-22.2 | 14.375 | 18.1625 | 34.125 | 80.7 | 81.775
-22.1 | 14.5125 | 18.45 | 38.7 | 81.0154 | 81.95
-22 | 14.6429 | 18.7714 | 44.5714 | 81.3867 | 82.15
-21.9 | 14.85 | 19.1625 | 51.0857 | 81.8533 | 82.3714
-21.8 | 14.98 | 19.6 | 58 | 82.1813 | 82.6375
-21.7 | 15.0889 | 20.1857 | 65 | 82.48 | 82.9857
-21.6 | 15.3455 | 21.0125 | 71.1714 | 82.7467 | 83.175
-21.5 | 15.56 | 22.225 | 75.7 | 83.0667 | 83.425
-21.4 | 15.7 | 24.2857 | 79.5429 | 83.38 | 83.6429
-21.3 | 15.9 | 27.2 | 81.7714 | 83.6467 | 83.85
-21.2 | 16.2077 | 31.8286 | 83.25 | 83.8778 | 84.1375
-21.1 | 16.4 | 37.4375 | 84.2143 | 84.1143 | 84.5
-21 | 16.8143 | 44.0714 | 84.8 | 84.35 | 84.6143
-20.9 | 17.4125 | 52.25 | 85.3875 | 84.6714 | 84.7
-20.8 | 18.3 | 60.225 | 86 | 84.9286 | 84.86
-20.7 | 19.8571 | 68.06 | 86.3667 | 85.2 | 85.1286
-20.6 | 22.4375 | 75.2625 | 86.68 | 85.3857 | 85.35
-20.5 | 26.3 | 80.6571 | 87.06 | 85.6429 | 85.6571
-20.4 | 31.875 | 84.55 | 87.46 | 85.95 | 85.925
-20.3 | 39.0714 | 86.95 | 87.78 | 86.1714 | 86.1857
-20.2 | 47.375 | 88.4 | 88.06 | 86.4286 | 86.375
-20.1 | 56.2714 | 89.4 | 88.38 | 86.7 | 86.5125
-20 | 65.7 | 89.9625 | 88.68 | 86.8857 | 86.675
-19.9 | 74.7 | 90.5286 | 88.96 | 87.1 | 86.8714
-19.8 | 82.2 | 90.85 | 89.26 | 87.4 | 87.1
-19.7 | 87.1 | 91.14 | 89.48 | 87.65 | 87.275
-19.6 | 90.4143 | 91.6 | 89.76 | 87.8 | 87.4625
-19.5 | 92.1 | 91.875 | 90.08 | 88.06 | 87.6
-19.4 | 93.0143 | 92.0778 | 90.28 | 88.2571 | 87.7571
-19.3 | 93.6 | 92.3077 | 90.475 | 88.4833 | 88.0875
-19.2 | 94.025 | 92.52 | 90.675 | 88.625 | 88.4
-19.1 | 94.2857 | 92.75 | 90.9333 | 88.8 | 88.425
-19 | 94.475 | 92.98 | 91.16 | 89.0714 | 88.5571
-18.9 | 94.7 | 93.2 | 91.36 | 89.2143 | 88.6875
-18.8 | 94.8375 | 93.4286 | 91.56 | 89.4 | 88.8143
-18.7 | 94.9714 | 93.625 | 91.76 | 89.6714 | 89.05
-18.6 | 95.2 | 93.8 | 92.02 | 89.7286 | 89.35
-18.5 | 95.3429 | 93.95 | 92.2 | 90 | 89.5286
-18.4 | 95.475 | 94.1889 | 92.46 | 90.2714 | 89.75
-18.3 | 95.7 | 94.3625 | 92.6571 | 90.4286 | 89.8
-18.2 | 95.8 | 94.4 | 92.8 | 90.6571 | 89.975
-18.1 | 95.8714 | 94.5571 | 92.9571 | 90.8125 | 90.1714
-18 | 96 | 94.825 | 93.0875 | 90.975 | 90.4
-17.9 | 96.1286 | 94.9875 | 93.25 | 91.1714 | 90.5
-17.8 | 96.2625 | 95.125 | 93.5143 | 91.3 | 90.5625
-17.7 | 96.4 | 95.3 | 93.6875 | 91.4375 | 90.6875
-17.6 | 96.5286 | 95.4857 | 93.8143 | 91.7143 | 90.825
-17.5 | 96.6 | 95.6 | 93.9571 | 91.9 | 90.9571
-17.4 | 96.7 | 95.7 | 94.175 | 92.0375 | 91.1
-17.3 | 96.8 | 95.8857 | 94.3 | 92.1714 | 91.2833
-17.2 | 96.8 | 96.0286 | 94.5429 | 92.2727 | 91.325
-17.1 | 96.9 | 96.25 | 94.62 | 92.4556 | 91.4571
-17 | 97.0571 | 96.3 | 94.8 | 92.7 | 91.54
-16.9 | 97.2625 | 96.4143 | 94.9 | 92.8111 | 91.65
-16.8 | 97.3 | 96.55 | 95.0571 | 92.9182 | 91.88
-16.7 | 97.425 | 96.6857 | 95.2625 | 93.0167 | 92.05
-16.6 | 97.55 | 96.8125 | 95.3 | 93.1833 | 92.1
-16.5 | 97.6 | 96.9 | 95.5286 | 93.3385 | 92.16
-16.4 | 97.725 | 96.9857 | 95.6 | 93.4 | 92.325
-16.3 | 97.8571 | 97.0125 | 95.6875 | 93.4636 | 92.45
-16.2 | 98 | 97.1375 | 95.925 | 93.58 | 92.5857
-16.1 | 98.025 | 97.2714 | 96.0571 | 93.7133 | 92.7143
-16 | 98.1 | 97.4 | 96.1 | 93.8467 | 92.85
-15.9 | 98.125 | 97.6 | 96.2143 | 93.98 | 92.9857
-15.8 | 98.2667 | 97.6714 | 96.35 | 94.1125 | 93.1125
-15.7 | 98.4 | 97.8 | 96.4857 | 94.2375 | 93.2429
-15.6 | 98.5 | 97.8375 | 96.5143 | 94.3667 | 93.375
-15.5 | 98.6 | 97.9 | 96.6571 | 94.5 | 93.5125
-15.4 | 98.7 | 98 | 96.7 | 94.5833 | 93.6429
-15.3 | 98.8 | 98.1 | 96.8286 | 94.61 | 93.7
-15.2 | 98.9 | 98.1625 | 97 | 94.7125 | 93.8
-15.1 | 98.98 | 98.3 | 97 | 94.8333 | 93.9
-15 | 99.1 | 98.4 | 97.125 | 94.9286 | 93.9714
-14.9 | 99.18 | 98.4625 | 97.2571 | 95 | 94.1
-14.8 | 99.2 | 98.6 | 97.3 | 95.0714 | 94.2375
-14.7 | 99.3 | 98.6571 | 97.425 | 95.2091 | 94.3
-14.6 | 99.36 | 98.7 | 97.5 | 95.4 | 94.4
-14.5 | 99.44 | 98.8 | 97.5 | 95.5625 | 94.5375
-14.4 | 99.5 | 98.8 | 97.6125 | 95.67 | 94.6
-14.3 | 99.6 | 98.9 | 97.7 | 95.7 | 94.7
-14.2 | 99.64 | 98.9 | 97.7 | 95.7 | 94.9286
-14.1 | 99.7 | 98.9 | 97.7667 | 95.7625 | 95
-14 | 99.74 | 99 | 97.875 | 95.9 | 95.0875
-13.9 | 99.8 | 99 | 97.9 | 96.0286 | 95.125
-13.8 | 99.8 | 99.0857 | 97.925 | 96.1625 | 95.2
-13.7 | 99.8 | 99.1 | 98 | 96.2 | 95.2
-13.6 | 99.8 | 99.1 | 98.0625 | 96.225 | 95.2833
-13.5 | 99.8 | 99.1 | 98.1 | 96.4 | 95.5167
-13.4 | 99.8 | 99.2 | 98.1286 | 96.5 | 95.7
-13.3 | 99.8375 | 99.3 | 98.2 | 96.6143 | 95.8
-13.2 | 99.9 | 99.3625 | 98.3 | 96.75 | 95.9
-13.1 | 99.9 | 99.4 | 98.3286 | 96.8 | 95.9
-13 | 100 | 99.4 | 98.4 | 96.8125 | 96
-12.9 | 100 | 99.5 | 98.5 | 96.9 | 96.1
-12.8 | 100 | 99.5 | 98.6 | 96.975 | 96.1833
-12.7 | 100.0875 | 99.54 | 98.6571 | 97.1125 | 96.2
-12.6 | 100.1 | 99.6067 | 98.7 | 97.2 | 96.3286
-12.5 | 100.1 | 99.6733 | 98.8143 | 97.2714 | 96.4
-12.4 | 100.2 | 99.7 | 98.9 | 97.4 | 96.5
-12.3 | 100.2 | 99.7067 | 98.9 | 97.4375 | 96.6
-12.2 | 100.3 | 99.7733 | 98.9125 | 97.5 | 96.6571
-12.1 | 100.3 | 99.84 | 99 | 97.5 | 96.7
-12 | 100.3 | 99.9 | 99.0857 | 97.6 | 96.7571
-11.9 | 100.4 | 99.9 | 99.1143 | 97.6625 | 96.8
-11.8 | 100.4 | 99.9333 | 99.2 | 97.72 | 96.8125
-11.7 | 100.4 | 100 | 99.275 | 97.9 | 96.9429
-11.6 | 100.4125 | 100 | 99.3 | 97.9286 | 97.15
-11.5 | 100.5 | 100.0333 | 99.3 | 98 | 97.2
-11.4 | 100.5 | 100.1 | 99.375 | 98.0875 | 97.3
-11.3 | 100.6 | 100.1667 | 99.4 | 98.1571 | 97.3556
-11.2 | 100.6 | 100.2 | 99.4 | 98.2 | 97.4
-11.1 | 100.6 | 100.2 | 99.4833 | 98.2 | 97.4533
-11 | 100.6125 | 100.2667 | 99.6 | 98.2857 | 97.5867
-10.9 | 100.7 | 100.3 | 99.7 | 98.3 | 97.66
-10.8 | 100.7 | 100.3 | 99.7 | 98.3 | 97.7
-10.7 | 100.8 | 100.3714 | 99.7 | 98.3714 | 97.7
-10.6 | 100.8 | 100.4 | 99.7714 | 98.4 | 97.7533
-10.5 | 100.875 | 100.4667 | 99.8 | 98.5 | 97.82
-10.4 | 100.9 | 100.5 | 99.8 | 98.5375 | 97.8867
-10.3 | 100.9 | 100.6 | 99.8 | 98.6 | 97.9533
-10.2 | 100.9 | 100.6 | 99.8 | 98.7 | 98.0214
-10.1 | 100.9 | 100.6 | 99.8 | 98.7 | 98.0929
-10 | 101 | 100.6 | 99.9 | 98.7571 | 98.1818
-9.9 | 101 | 100.6 | 100 | 98.8 | 98.2667
-9.8 | 101 | 100.6 | 100 | 98.825 | 98.3
-9.7 | 101 | 100.6 | 100 | 98.9571 | 98.3636
-9.6 | 101 | 100.6 | 100 | 99 | 98.475
-9.5 | 101 | 100.6 | 100 | 99.1 | 98.6125
-9.4 | 101 | 100.6625 | 100.1 | 99.1 | 98.7
-9.3 | 101.1 | 100.7 | 100.1 | 99.1875 | 98.7
-9.2 | 101.02 | 100.7 | 100.1 | 99.2 | 98.7
-9.1 | 101.0143 | 100.7 | 100.1875 | 99.2 | 98.7375
-9 | 101.1 | 100.7 | 100.3 | 99.2875 | 98.8
-8.9 | 101.1 | 100.7 | 100.3 | 99.3 | 98.9
-8.8 | 101.0125 | 100.8 | 100.3875 | 99.35 | 98.9
-8.7 | 101.1 | 100.8 | 100.4 | 99.475 | 99
-8.6 | 101.1 | 100.8 | 100.4 | 99.5 | 99.1
-8.5 | 101.1 | 100.8 | 100.4875 | 99.55 | 99.1429
-8.4 | 101.2 | 100.825 | 100.5 | 99.6 | 99.2
-8.3 | 101.2 | 100.9 | 100.5 | 99.7 | 99.2
-8.2 | 101.2 | 100.9 | 100.5 | 99.7 | 99.2375
-8.1 | 101.2286 | 101 | 100.6 | 99.7 | 99.3714
-8 | 101.2375 | 101 | 100.6 | 99.7 | 99.4
-7.9 | 101.2 | 101 | 100.6 | 99.7 | 99.4
-7.8 | 101.3 | 101 | 100.6 | 99.775 | 99.5
-7.7 | 101.3 | 101.06 | 100.6 | 99.9 | 99.5
-7.6 | 101.3 | 101.1 | 100.6 | 99.9 | 99.5571
-7.5 | 101.3857 | 101.1 | 100.6 | 100 | 99.6
-7.4 | 101.4111 | 101.1 | 100.7 | 100.0286 | 99.64
-7.3 | 101.5 | 101.1 | 100.7 | 100.1 | 99.7
-7.2 | 101.5 | 101.2 | 100.7714 | 100.1 | 99.7
-7.1 | 101.5 | 101.2 | 100.8 | 100.0286 | 99.725
-7 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
-6.9 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7143
-6.8 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
-6.7 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
-6.6 | 101.4364 | 101.2 | 100.925 | 100.2 | 99.8
-6.5 | 101.4429 | 101.2 | 101 | 100.2 | 99.825
-6.4 | 101.5 | 101.2 | 101 | 100.2 | 99.9
-6.3 | 101.5 | 101.2 | 101 | 100.3 | 99.9
-6.2 | 101.5 | 101.2 | 101 | 100.3 | 99.9
-6.1 | 101.5 | 101.2 | 101 | 100.3 | 99.9
-6 | 101.5 | 101.3 | 100.9 | 100.325 | 99.8875
-5.9 | 101.5 | 101.3 | 100.95 | 100.4 | 99.85
-5.8 | 101.5 | 101.3 | 101 | 100.4 | 99.9
-5.7 | 101.5 | 101.3 | 100.9875 | 100.4 | 99.9
-5.6 | 101.5 | 101.2625 | 100.9375 | 100.4 | 99.9
-5.5 | 101.5 | 101.2 | 101 | 100.475 | 99.9125
-5.4 | 101.5 | 101.2 | 101 | 100.4 | 100
-5.3 | 101.4 | 101.2 | 101 | 100.4429 | 100.075
-5.2 | 101.4 | 101.2 | 101.0714 | 100.5 | 100.1
-5.1 | 101.4 | 101.2 | 101.1 | 100.5 | 100.1429
-5 | 101.4 | 101.2 | 101 | 100.5 | 100.2
-4.9 | 101.4 | 101.2333 | 101 | 100.4667 | 100.2
-4.8 | 101.4 | 101.2143 | 101 | 100.4 | 100.2
-4.7 | 101.4 | 101.2 | 101 | 100.5 | 100.2
-4.6 | 101.3 | 101.2 | 101 | 100.4667 | 100.1
-4.5 | 101.3 | 101.2 | 101 | 100.4 | 100.1
-4.4 | 101.3 | 101.2 | 101 | 100.4 | 100.1857
-4.3 | 101.3 | 101.1 | 101 | 100.4 | 100.1083
-4.2 | 101.2875 | 101.1 | 101 | 100.4 | 100.1917
-4.1 | 101.2 | 101.1 | 100.9 | 100.4 | 100.2143
-4 | 101.2 | 101.1 | 100.9 | 100.4 | 100.3
-3.9 | 101.2 | 101.1 | 100.8667 | 100.4 | 100.375
-3.8 | 101.1571 | 101.0167 | 100.8 | 100.4 | 100.4
-3.7 | 101.1 | 101 | 100.8 | 100.4 | 100.3625
-3.6 | 101.1875 | 101 | 100.8 | 100.4 | 100.3
-3.5 | 101.1 | 101 | 100.8 | 100.4 | 100.3
-3.4 | 101.1 | 100.9 | 100.7778 | 100.4 | 100.4
-3.3 | 101 | 100.9 | 100.7 | 100.4 | 100.4
-3.2 | 101 | 100.9 | 100.7 | 100.4625 | 100.4
-3.1 | 100.925 | 100.9 | 100.7 | 100.5 | 100.4
-3 | 100.9 | 100.85 | 100.7 | 100.5 | 100.4
-2.9 | 100.9 | 100.8 | 100.7 | 100.4714 | 100.3
-2.8 | 100.8286 | 100.8 | 100.7 | 100.4 | 100.3
-2.7 | 100.8 | 100.7625 | 100.6571 | 100.4 | 100.2375
-2.6 | 100.8 | 100.7 | 100.6 | 100.39 | 100.2
-2.5 | 100.7286 | 100.7 | 100.5875 | 100.3 | 100.26
-2.4 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
-2.3 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
-2.2 | 100.7 | 100.6 | 100.5 | 100.3533 | 100.4
-2.1 | 100.6 | 100.6 | 100.5 | 100.4 | 100.4
-2 | 100.6 | 100.6 | 100.5 | 100.3867 | 100.4
-1.9 | 100.6 | 100.6 | 100.5 | 100.32 | 100.4
-1.8 | 100.6 | 100.6 | 100.5 | 100.2533 | 100.4
-1.7 | 100.6 | 100.6 | 100.5 | 100.2 | 100.4
-1.6 | 100.5 | 100.5 | 100.5 | 100.2 | 100.32
-1.5 | 100.4714 | 100.5 | 100.4714 | 100.2 | 100.3143
-1.4 | 100.4 | 100.4 | 100.4 | 100.2 | 100.35
-1.3 | 100.3125 | 100.4 | 100.375 | 100.2 | 100.4
-1.2 | 100.3 | 100.375 | 100.3 | 100.2 | 100.3
-1.1 | 100.3 | 100.3 | 100.3 | 100.2 | 100.3
-1 | 100.2125 | 100.3 | 100.2429 | 100.1167 | 100.2
-0.9 | 100.2 | 100.275 | 100.2 | 100.1167 | 100.2
-0.8 | 100.2 | 100.2 | 100.2 | 100.1 | 100.2
-0.7 | 100.2 | 100.1125 | 100.1 | 100.1 | 100.18
-0.6 | 100.2 | 100.1 | 100.1 | 100.1 | 100.1
-0.5 | 100.2 | 100.0429 | 100.1 | 100.1 | 100
-0.4 | 100.1 | 100 | 100.1 | 100 | 100
-0.3 | 100.1 | 100 | 100.1 | 100 | 100
-0.2 | 100.0875 | 100 | 100.1 | 100 | 100
-0.1 | 100 | 100 | 100.0571 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100.0167 | 100 | 100
0.2 | 100.0875 | 100 | 100.1 | 100 | 100
0.3 | 100.1 | 100 | 100.1 | 100 | 100
0.4 | 100.1 | 100 | 100.1 | 100 | 100
0.5 | 100.2 | 100.075 | 100.1 | 100.1 | 100
0.6 | 100.2 | 100.1 | 100.1 | 100.1 | 100.1
0.7 | 100.2 | 100.1 | 100.1667 | 100.1 | 100.14
0.8 | 100.2 | 100.175 | 100.2 | 100.1 | 100.2
0.9 | 100.2 | 100.3 | 100.2 | 100.1167 | 100.2
1 | 100.2125 | 100.3 | 100.2143 | 100.1167 | 100.2
1.1 | 100.3 | 100.3 | 100.3 | 100.2 | 100.3
1.2 | 100.3 | 100.4 | 100.3 | 100.2 | 100.3
1.3 | 100.3125 | 100.4 | 100.3429 | 100.2 | 100.3875
1.4 | 100.4 | 100.4 | 100.4 | 100.2 | 100.3
1.5 | 100.4714 | 100.5 | 100.4375 | 100.2 | 100.3
1.6 | 100.5 | 100.5 | 100.5 | 100.2 | 100.35
1.7 | 100.6 | 100.6 | 100.5 | 100.2 | 100.4
1.8 | 100.6 | 100.6 | 100.5 | 100.2533 | 100.4
1.9 | 100.6 | 100.6 | 100.5 | 100.32 | 100.4
2 | 100.6 | 100.6 | 100.5 | 100.3867 | 100.4
2.1 | 100.6 | 100.6 | 100.5 | 100.4 | 100.4
2.2 | 100.7 | 100.6167 | 100.5 | 100.3533 | 100.4
2.3 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
2.4 | 100.7 | 100.7 | 100.5 | 100.3 | 100.3
2.5 | 100.7286 | 100.7 | 100.5571 | 100.3 | 100.3
2.6 | 100.8 | 100.7 | 100.6 | 100.39 | 100.2
2.7 | 100.8 | 100.7875 | 100.625 | 100.4 | 100.2
2.8 | 100.8286 | 100.8 | 100.7 | 100.4 | 100.3
2.9 | 100.9 | 100.8 | 100.7 | 100.4714 | 100.3
3 | 100.9 | 100.8 | 100.7 | 100.5 | 100.4
3.1 | 100.925 | 100.9 | 100.7 | 100.5 | 100.4
3.2 | 101 | 100.9 | 100.7 | 100.4625 | 100.4
3.3 | 101 | 100.9 | 100.7 | 100.4 | 100.4
3.4 | 101.1 | 100.9 | 100.7429 | 100.4 | 100.3714
3.5 | 101.1 | 100.95 | 100.8 | 100.4 | 100.3
3.6 | 101.1875 | 101 | 100.8 | 100.4 | 100.3
3.7 | 101.1 | 101 | 100.8 | 100.4 | 100.4
3.8 | 101.1571 | 101.05 | 100.8667 | 100.4 | 100.4
3.9 | 101.2 | 101.1 | 100.9 | 100.4 | 100.3125
4 | 101.2 | 101.1 | 100.9 | 100.4 | 100.2857
4.1 | 101.2 | 101.1 | 100.9 | 100.4 | 100.2
4.2 | 101.2875 | 101.1 | 100.975 | 100.4 | 100.2
4.3 | 101.3 | 101.1 | 101 | 100.4 | 100.1364
4.4 | 101.3 | 101.2 | 101 | 100.4 | 100.1857
4.5 | 101.3 | 101.2 | 101 | 100.4 | 100.1
4.6 | 101.3 | 101.2 | 101 | 100.4667 | 100.1429
4.7 | 101.4 | 101.2 | 101 | 100.5 | 100.2
4.8 | 101.4 | 101.2429 | 101 | 100.4 | 100.2
4.9 | 101.4 | 101.2875 | 101 | 100.4667 | 100.2
5 | 101.4 | 101.2 | 101.0286 | 100.5 | 100.2
5.1 | 101.4 | 101.2 | 101.1 | 100.5 | 100.1
5.2 | 101.4 | 101.2 | 101.1 | 100.5 | 100.1
5.3 | 101.4 | 101.2 | 101 | 100.4429 | 100.1
5.4 | 101.5 | 101.2 | 101 | 100.4 | 100
5.5 | 101.5 | 101.2 | 101 | 100.475 | 99.9429
5.6 | 101.5 | 101.2 | 100.975 | 100.4 | 99.9
5.7 | 101.5 | 101.3 | 100.9571 | 100.4 | 99.9
5.8 | 101.5 | 101.3 | 101 | 100.4 | 99.9
5.9 | 101.5 | 101.3 | 100.975 | 100.4 | 99.9
6 | 101.5 | 101.3 | 100.9 | 100.325 | 99.9
6.1 | 101.5 | 101.2429 | 100.9857 | 100.3 | 99.9
6.2 | 101.5 | 101.2 | 101 | 100.3 | 99.9
6.3 | 101.5 | 101.2 | 101 | 100.3 | 99.9
6.4 | 101.5 | 101.2 | 101 | 100.2 | 99.9
6.5 | 101.4429 | 101.2 | 100.9875 | 100.2 | 99.8571
6.6 | 101.4364 | 101.2 | 100.9 | 100.2 | 99.8
6.7 | 101.5 | 101.2 | 100.8333 | 100.1 | 99.8
6.8 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
6.9 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7
7 | 101.5 | 101.2 | 100.8 | 100.1 | 99.8
7.1 | 101.5 | 101.2 | 100.8 | 100.0286 | 99.7571
7.2 | 101.5 | 101.2 | 100.8 | 100.1 | 99.7
7.3 | 101.5 | 101.1 | 100.7 | 100.1 | 99.7
7.4 | 101.4111 | 101.1 | 100.7 | 100.0286 | 99.6
7.5 | 101.3857 | 101.1 | 100.6 | 100 | 99.6
7.6 | 101.3 | 101.1 | 100.6 | 99.9 | 99.5875
7.7 | 101.3 | 101.1 | 100.6 | 99.9 | 99.5
7.8 | 101.3 | 101 | 100.6 | 99.775 | 99.5
7.9 | 101.2 | 101 | 100.6 | 99.7 | 99.4
8 | 101.2375 | 101 | 100.6 | 99.7 | 99.4
8.1 | 101.2286 | 101 | 100.6 | 99.7 | 99.3
8.2 | 101.2 | 100.94 | 100.5143 | 99.7 | 99.2714
8.3 | 101.2 | 100.9 | 100.5 | 99.7 | 99.2
8.4 | 101.2 | 100.8875 | 100.5 | 99.6 | 99.2
8.5 | 101.1 | 100.8 | 100.5 | 99.55 | 99.1
8.6 | 101.1 | 100.8 | 100.4 | 99.5 | 99.1
8.7 | 101.1 | 100.8 | 100.4 | 99.475 | 99
8.8 | 101.0125 | 100.8 | 100.325 | 99.35 | 98.92
8.9 | 101.1 | 100.7 | 100.3 | 99.3 | 98.82
9 | 101.1 | 100.7 | 100.3 | 99.2875 | 98.8
9.1 | 101.0143 | 100.7 | 100.225 | 99.2 | 98.7
9.2 | 101.02 | 100.7 | 100.1 | 99.2 | 98.7
9.3 | 101.1 | 100.7 | 100.1 | 99.1875 | 98.7
9.4 | 101 | 100.7 | 100.0333 | 99.1 | 98.675
9.5 | 101 | 100.6 | 100 | 99.1 | 98.5429
9.6 | 101 | 100.6 | 100 | 99 | 98.5
9.7 | 101 | 100.6 | 100 | 98.9571 | 98.3917
9.8 | 101 | 100.6 | 100 | 98.825 | 98.3083
9.9 | 101 | 100.6 | 99.9286 | 98.8 | 98.2923
10 | 101 | 100.6 | 99.9375 | 98.7571 | 98.2154
10.1 | 100.9 | 100.6 | 99.8 | 98.7 | 98.1273
10.2 | 100.9 | 100.6 | 99.8 | 98.7 | 98.0067
10.3 | 100.9 | 100.6 | 99.8 | 98.6 | 97.8867
10.4 | 100.9 | 100.5 | 99.8 | 98.5375 | 97.82
10.5 | 100.875 | 100.5 | 99.8 | 98.5 | 97.8
10.6 | 100.8 | 100.4273 | 99.8 | 98.4 | 97.7867
10.7 | 100.8 | 100.3417 | 99.7 | 98.3714 | 97.72
10.8 | 100.7 | 100.3 | 99.7 | 98.3 | 97.7
10.9 | 100.7 | 100.3 | 99.625 | 98.3 | 97.6867
11 | 100.6125 | 100.2333 | 99.55 | 98.2857 | 97.5533
11.1 | 100.6 | 100.2 | 99.4125 | 98.2 | 97.46
11.2 | 100.6 | 100.2 | 99.4 | 98.2 | 97.39
11.3 | 100.6 | 100.1333 | 99.4 | 98.1571 | 97.3
11.4 | 100.5 | 100.0667 | 99.4 | 98.0875 | 97.3
11.5 | 100.5 | 100 | 99.3 | 98 | 97.2375
11.6 | 100.4125 | 100 | 99.3 | 97.9286 | 97.1125
11.7 | 100.4 | 100 | 99.2125 | 97.9 | 96.95
11.8 | 100.4 | 100 | 99.2 | 97.72 | 96.8
11.9 | 100.4 | 99.9357 | 99.15 | 97.6625 | 96.8
12 | 100.3 | 99.8688 | 99.1 | 97.6 | 96.7875
12.1 | 100.3 | 99.8062 | 99 | 97.5 | 96.7
12.2 | 100.3 | 99.8 | 98.95 | 97.5 | 96.6
12.3 | 100.2 | 99.7733 | 98.9 | 97.4375 | 96.6
12.4 | 100.2 | 99.7067 | 98.9 | 97.4 | 96.4571
12.5 | 100.1 | 99.64 | 98.85 | 97.2714 | 96.4
12.6 | 100.1 | 99.575 | 98.725 | 97.2 | 96.3714
12.7 | 100.0875 | 99.5125 | 98.6 | 97.1125 | 96.2167
12.8 | 100 | 99.5 | 98.5556 | 96.975 | 96.2
12.9 | 100 | 99.5 | 98.5 | 96.9 | 96.1
13 | 100 | 99.4625 | 98.4 | 96.8125 | 96
13.1 | 99.9 | 99.4 | 98.3625 | 96.8 | 95.9
13.2 | 99.9 | 99.33 | 98.2286 | 96.75 | 95.9
13.3 | 99.8375 | 99.3 | 98.2 | 96.6143 | 95.8
13.4 | 99.8 | 99.2 | 98.1 | 96.5 | 95.64
13.5 | 99.8 | 99.1444 | 98.1 | 96.4 | 95.4571
13.6 | 99.8 | 99.1 | 98 | 96.225 | 95.325
13.7 | 99.8 | 99.1 | 98 | 96.2 | 95.2
13.8 | 99.8 | 99.05 | 97.9714 | 96.1625 | 95.2
13.9 | 99.8 | 99 | 97.9 | 96.0286 | 95.1571
14 | 99.74 | 99 | 97.9 | 95.9 | 95.025
14.1 | 99.7 | 98.95 | 97.8 | 95.7625 | 95
14.2 | 99.64 | 98.9 | 97.7091 | 95.7 | 94.8625
14.3 | 99.6 | 98.8857 | 97.675 | 95.7 | 94.7375
14.4 | 99.5 | 98.8 | 97.6 | 95.67 | 94.6
14.5 | 99.44 | 98.8 | 97.5125 | 95.5625 | 94.5
14.6 | 99.36 | 98.7 | 97.4875 | 95.4 | 94.4375
14.7 | 99.3 | 98.6286 | 97.4 | 95.2091 | 94.3
14.8 | 99.2 | 98.6 | 97.325 | 95.0714 | 94.2714
14.9 | 99.18 | 98.4286 | 97.2 | 95 | 94.075
15 | 99.1 | 98.3875 | 97.1571 | 94.9286 | 94
15.1 | 98.98 | 98.2625 | 97.0143 | 94.8333 | 93.9
15.2 | 98.9 | 98.2 | 96.9857 | 94.7125 | 93.8375
15.3 | 98.8 | 98.1 | 96.85 | 94.61 | 93.7
15.4 | 98.7 | 98.0625 | 96.7143 | 94.5833 | 93.575
15.5 | 98.6 | 97.9375 | 96.6875 | 94.5 | 93.5
15.6 | 98.5 | 97.8 | 96.5571 | 94.3667 | 93.4125
15.7 | 98.4 | 97.775 | 96.4143 | 94.2375 | 93.275
15.8 | 98.2667 | 97.7 | 96.3857 | 94.1125 | 93.1
15.9 | 98.125 | 97.6 | 96.25 | 93.98 | 92.9143
16 | 98.1 | 97.4429 | 96.1143 | 93.8467 | 92.8857
16.1 | 98.025 | 97.2375 | 96 | 93.7133 | 92.75
16.2 | 98 | 97.2 | 95.9143 | 93.58 | 92.625
16.3 | 97.8571 | 97.075 | 95.725 | 93.4636 | 92.4875
16.4 | 97.725 | 96.95 | 95.6 | 93.4 | 92.3571
16.5 | 97.6 | 96.9 | 95.4571 | 93.3385 | 92.2
16.6 | 97.55 | 96.7857 | 95.325 | 93.1833 | 92.1
16.7 | 97.425 | 96.65 | 95.2 | 93.0167 | 92.0833
16.8 | 97.3 | 96.5143 | 95.0625 | 92.9182 | 91.9
16.9 | 97.2625 | 96.4857 | 94.9286 | 92.8111 | 91.7143
17 | 97.0571 | 96.35 | 94.74 | 92.7 | 91.58
17.1 | 96.9 | 96.2143 | 94.6 | 92.4556 | 91.4875
17.2 | 96.8 | 96.0714 | 94.475 | 92.2727 | 91.3571
17.3 | 96.8 | 95.85 | 94.32 | 92.1714 | 91.225
17.4 | 96.7 | 95.7143 | 94.1333 | 92.0375 | 91.04
17.5 | 96.6 | 95.5857 | 93.9857 | 91.9 | 90.9875
17.6 | 96.5286 | 95.45 | 93.85 | 91.7143 | 90.8143
17.7 | 96.4 | 95.3143 | 93.725 | 91.4375 | 90.625
17.8 | 96.2625 | 95.1 | 93.475 | 91.3 | 90.6
17.9 | 96.1286 | 94.95 | 93.2571 | 91.1714 | 90.5
18 | 96 | 94.825 | 93.125 | 90.975 | 90.375
18.1 | 95.8714 | 94.6 | 92.9857 | 90.8125 | 90.2
18.2 | 95.8 | 94.4625 | 92.82 | 90.6571 | 89.9714
18.3 | 95.7 | 94.3286 | 92.6571 | 90.4286 | 89.8286
18.4 | 95.475 | 94.2 | 92.38 | 90.2714 | 89.7
18.5 | 95.3429 | 94 | 92.18 | 90 | 89.6
18.6 | 95.2 | 93.7625 | 91.9833 | 89.7286 | 89.3286
18.7 | 94.9714 | 93.6286 | 91.8 | 89.6714 | 89.0714
18.8 | 94.8375 | 93.5 | 91.6 | 89.4 | 88.85
18.9 | 94.7 | 93.1833 | 91.4 | 89.2143 | 88.725
19 | 94.475 | 93.0167 | 91.2 | 89.0714 | 88.5875
19.1 | 94.2857 | 92.6727 | 90.9 | 88.8 | 88.4625
19.2 | 94.025 | 92.52 | 90.7 | 88.625 | 88.3286
19.3 | 93.6 | 92.3 | 90.5 | 88.4833 | 88.1
19.4 | 93.0143 | 92.0909 | 90.3 | 88.2571 | 87.7625
19.5 | 92.1 | 91.9375 | 90.1 | 88.06 | 87.5571
19.6 | 90.4143 | 91.5182 | 89.7 | 87.8 | 87.5
19.7 | 87.1 | 91.12 | 89.5 | 87.65 | 87.225
19.8 | 82.2 | 90.8667 | 89.2 | 87.4 | 87.0375
19.9 | 74.7 | 90.5375 | 89.0167 | 87.1 | 86.9
20 | 65.7 | 89.925 | 88.64 | 86.8857 | 86.675
20.1 | 56.2714 | 89.3667 | 88.42 | 86.7 | 86.5429
20.2 | 47.375 | 88.44 | 88.04 | 86.4286 | 86.4125
20.3 | 39.0714 | 87 | 87.72 | 86.1714 | 86.15
20.4 | 31.875 | 84.6 | 87.52 | 85.95 | 85.9
20.5 | 26.3 | 80.725 | 87.04 | 85.6429 | 85.6286
20.6 | 22.4375 | 75.175 | 86.74 | 85.3857 | 85.3714
20.7 | 19.8571 | 68.2222 | 86.34 | 85.2 | 85.1
20.8 | 18.3 | 60.1875 | 85.94 | 84.9286 | 84.9
20.9 | 17.4125 | 51.9125 | 85.375 | 84.6714 | 84.7
21 | 16.8143 | 44.6625 | 84.8286 | 84.35 | 84.6429
21.1 | 16.4 | 37.5714 | 84.175 | 84.1143 | 84.4571
21.2 | 16.2077 | 31.775 | 83.2714 | 83.8778 | 84.0857
21.3 | 15.9 | 27.42 | 81.8 | 83.6467 | 83.825
21.4 | 15.7 | 24.25 | 79.5571 | 83.38 | 83.65
21.5 | 15.56 | 22.175 | 75.7571 | 83.0667 | 83.3857
21.6 | 15.3455 | 21 | 71 | 82.7467 | 83.125
21.7 | 15.0889 | 20.225 | 64.9571 | 82.48 | 82.95
21.8 | 14.98 | 19.5625 | 58 | 82.1813 | 82.65
21.9 | 14.85 | 19.1714 | 51.1 | 81.8533 | 82.3286
22 | 14.6429 | 18.775 | 44.4571 | 81.3867 | 82.0875
22.1 | 14.5125 | 18.475 | 38.7375 | 81.0154 | 81.9571
22.2 | 14.375 | 18.1714 | 34.225 | 80.7 | 81.75
22.3 | 14.1857 | 17.85 | 30.3857 | 80.36 | 81.575
22.4 | 14.0125 | 17.6 | 27.5 | 79.96 | 81.3143
22.5 | 13.9 | 17.2667 | 25.6875 | 79.64 | 81.05
22.6 | 13.775 | 17.1 | 24.5 | 79.3583 | 80.675
22.7 | 13.6273 | 16.825 | 23.7 | 78.9417 | 80.4571
22.8 | 13.4667 | 16.6375 | 23.0125 | 78.375 | 80.175
22.9 | 13.3733 | 16.4 | 22.4857 | 77.8286 | 79.875
23 | 13.3067 | 16.1429 | 22.1 | 77.35 | 79.625
23.1 | 13.1364 | 16 | 21.725 | 76.5 | 79.3571
23.2 | 13.0429 | 15.7714 | 21.3857 | 75.5 | 78.975
23.3 | 12.9733 | 15.575 | 21 | 74.05 | 78.725
23.4 | 12.9067 | 15.4 | 20.725 | 71.7375 | 78.4571
23.5 | 12.78 | 15.1714 | 20.3857 | 68.7 | 78.2
23.6 | 12.65 | 15.0375 | 20 | 64.625 | 77.7875
23.7 | 12.55 | 14.9 | 19.825 | 60.2 | 77.4571
23.8 | 12.4733 | 14.7 | 19.5571 | 54.6625 | 77.2
23.9 | 12.4067 | 14.6 | 19.3 | 49.2429 | 76.925
24 | 12.3308 | 14.4714 | 19.025 | 44.3 | 76.6125
24.1 | 12.2333 | 14.275 | 18.7571 | 39.7125 | 76.2
24.2 | 12.1364 | 14.1 | 18.5 | 36.0429 | 75.8143
24.3 | 12.025 | 14 | 18.3429 | 33.2 | 75.475
24.4 | 11.9875 | 13.82 | 18.075 | 31.2571 | 75.1
24.5 | 11.8571 | 13.8 | 17.825 | 29.775 | 74.7429
24.6 | 11.725 | 13.62 | 17.575 | 28.7 | 74.4125
24.7 | 11.6875 | 13.52 | 17.4429 | 27.9571 | 74
24.8 | 11.6 | 13.4333 | 17.225 | 27.35 | 73.5857
24.9 | 11.5333 | 13.3 | 17.05 | 26.85 | 73.05
25 | 11.5 | 13.2 | 16.7857 | 26.45 | 72.4
25.1 | 11.4 | 13.06 | 16.6125 | 26.1286 | 71.7
25.2 | 11.2846 | 12.96 | 16.45 | 25.7375 | 70.575
25.3 | 11.2077 | 12.9 | 16.25 | 25.375 | 69.0625
25.4 | 11.1437 | 12.7625 | 16.1143 | 25.1286 | 67
25.5 | 11.08 | 12.6375 | 15.9714 | 24.7375 | 64.3
25.6 | 11.0133 | 12.5833 | 15.75 | 24.425 | 60.7857
25.7 | 10.9467 | 12.4375 | 15.6143 | 24.0857 | 56.8429
25.8 | 10.88 | 12.3 | 15.3857 | 23.825 | 52.5
25.9 | 10.8133 | 12.275 | 15.25 | 23.575 | 47.9857
26 | 10.7467 | 12.1429 | 15.2 | 23.3143 | 44.1143
26.1 | 10.68 | 12.0125 | 14.9875 | 23.05 | 40.6
26.2 | 10.6133 | 11.975 | 14.8571 | 22.7625 | 37.8
26.3 | 10.6 | 11.9 | 14.725 | 22.4571 | 35.6
26.4 | 10.6 | 11.8125 | 14.5875 | 22.25 | 33.8
26.5 | 10.5 | 11.7 | 14.4571 | 21.975 | 32.4429
26.6 | 10.4375 | 11.6429 | 14.325 | 21.7143 | 31.7143
26.7 | 10.4 | 11.5125 | 14.1875 | 21.45 | 30.95
26.8 | 10.3 | 11.475 | 14.0571 | 21.2667 | 30.4
26.9 | 10.3 | 11.3429 | 13.925 | 20.9692 | 29.9625
27 | 10.2125 | 11.3 | 13.7875 | 20.8154 | 29.6143
27.1 |  | 11.2889 | 13.6571 | 20.5889 | 29.275
27.2 |  | 11.1667 | 13.6 | 20.4429 | 28.8625
27.3 |  | 11.1 | 13.3889 | 20.1571 | 28.5143
27.4 |  | 11 | 13.275 | 19.94 | 28.325
27.5 |  | 10.9571 | 13.2 | 19.7308 | 28
27.6 |  | 10.9 | 13.1375 | 19.5 | 27.6875
27.7 |  | 10.8 | 13 | 19.3556 | 27.2857
27.8 |  | 10.75 | 12.88 | 19.1182 | 27.1
27.9 |  | 10.7 | 12.7714 | 18.8714 | 26.725
28 |  | 10.6 | 12.7 | 18.7375 | 26.475
28.1 |  | 10.5 | 12.6 | 18.6 | 26.2
28.2 |  |  | 12.5 | 18.3714 | 25.94
28.3 |  |  | 12.4375 | 18.2167 | 25.64
28.4 |  |  | 12.3 | 18.06 | 25.44
28.5 |  |  | 12.1909 | 17.9267 | 25.24
28.6 |  |  | 12.1 | 17.7769 | 24.98
28.7 |  |  | 12.0778 | 17.5889 | 24.74
28.8 |  |  | 11.9455 | 17.4 | 24.5714
28.9 |  |  | 11.8 | 17.2556 | 24.375
29 |  |  | 11.75 | 17.075 | 24.0667
29.1 |  |  | 11.7 | 16.9 | 23.8375
29.2 |  |  | 11.6 | 16.7625 | 23.7
29.3 |  |  | 11.5556 | 16.6286 | 23.5429
29.4 |  |  | 11.4375 | 16.5 | 23.275
29.5 |  |  |  | 16.3714 | 23.1
29.6 |  |  |  | 16.175 | 22.7429
29.7 |  |  |  | 16 | 22.5375
29.8 |  |  |  | 15.8714 | 22.3
29.9 |  |  |  | 15.7375 | 22.175
30 |  |  |  | 15.6125 | 22.0429
30.1 |  |  |  | 15.55 | 21.8
30.2 |  |  |  | 15.3429 | 21.65
30.3 |  |  |  | 15.3 | 21.3857
30.4 |  |  |  | 15.075 | 21.2125
30.5 |  |  |  | 14.95 | 21
30.6 |  |  |  | 14.8143 | 20.95
30.7 |  |  |  | 14.7 | 20.7286
30.8 |  |  |  | 14.6571 | 20.475
30.9 |  |  |  | 14.5143 | 20.35
31 |  |  |  | 14.3833 | 20.2143
31.1 |  |  |  | 14.3 | 20.06
31.2 |  |  |  | 14.2 | 19.8833
31.3 |  |  |  | 14.0875 | 19.7
31.4 |  |  |  | 13.9625 | 19.6
31.5 |  |  |  | 13.8286 | 19.5
31.6 |  |  |  | 13.7 | 19.3333
31.7 |  |  |  | 13.7 | 19.1429
31.8 |  |  |  | 13.6 | 18.9167
31.9 |  |  |  | 13.5 | 18.8
32 |  |  |  | 13.36 | 18.675
32.1 |  |  |  | 13.2 | 18.5429
32.2 |  |  |  |  | 18.325
32.3 |  |  |  |  | 18.175
32.4 |  |  |  |  | 18.05
32.5 |  |  |  |  | 17.9143
32.6 |  |  |  |  | 17.8
32.7 |  |  |  |  | 17.75
32.8 |  |  |  |  | 17.625
32.9 |  |  |  |  | 17.3875
33 |  |  |  |  | 17.2571
33.1 |  |  |  |  | 17.125
33.2 |  |  |  |  | 17.04
33.3 |  |  |  |  | 16.94
33.4 |  |  |  |  | 16.7
33.5 |  |  |  |  | 16.6
33.6 |  |  |  |  | 16.5
33.7 |  |  |  |  | 16.38
33.8 |  |  |  |  | 16.3
33.9 |  |  |  |  | 16.2
34 |  |  |  |  | 16.075
34.1 |  |  |  |  | 15.95
34.2 |  |  |  |  | 15.9
34.3 |  |  |  |  | 15.7857
34.4 |  |  |  |  | 15.65
34.5 |  |  |  |  | 15.5143
34.6 |  |  |  |  | 15.4875
34.7 |  |  |  |  | 15.34
34.8 |  |  |  |  | 15.1571

## Wedge 30 Output factors
Output factors
Wedge 30 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.862 | 0.875 | 0.883 | 0.894 | 0.903 | 0.911 | 0.916 | 0.924
| 4 | 0.874 | 0.889 | 0.9 | 0.914 | 0.925 | 0.936 | 0.944 | 0.953
| 5 | 0.883 | 0.9 | 0.914 | 0.929 | 0.943 | 0.957 | 0.966 | 0.978
| 7 | 0.896 | 0.918 | 0.933 | 0.954 | 0.972 | 0.99 | 1.003 | 1.019
| 10 | 0.908 | 0.932 | 0.951 | 0.976 | 1 | 1.024 | 1.042 | 1.065
| 15 | 0.918 | 0.945 | 0.967 | 0.998 | 1.027 | 1.061 | 1.086 | 1.119
| 20 | 0.924 | 0.953 | 0.977 | 1.011 | 1.046 | 1.087 | 1.118 | 1.159
| 30 | 0.933 | 0.965 | 0.991 | 1.03 | 1.072 | 1.123 | 1.163 | 1.218
| 40 | 0.937 | 0.97 | 0.998 | 1.04 | 1.087 | 1.146 | 1.193 | 1.254

## Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 49 | 54.5 | 61.8 | 66.7
0.1 | 53.55 | 59.22 | 66.35 | 70.76
0.2 | 60.825 | 66.65 | 72.7 | 76.3
0.3 | 71.9 | 77.6 | 81.1857 | 84.2
0.4 | 81.325 | 84.9 | 87.9571 | 89.95
0.5 | 87.68 | 90.84 | 92.6286 | 94.06
0.6 | 92.6 | 94.6 | 95.7 | 96.68
0.7 | 95.9 | 96.9 | 97.5857 | 98.2
0.8 | 97.82 | 98.4833 | 98.7 | 99.1
0.9 | 99 | 99.3 | 99.6 | 99.78
1 | 99.7 | 99.76 | 99.925 | 100
1.1 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 100 | 99.9667
1.3 | 99.9 | 99.95 | 99.84 | 99.74
1.4 | 99.7 | 99.72 | 99.48 | 99.5
1.5 | 99.475 | 99.4 | 99.22 | 99.18
1.6 | 99.05 | 99.15 | 98.8571 | 98.825
1.7 | 98.5333 | 98.74 | 98.54 | 98.425
1.8 | 97.9667 | 98.4 | 98.2 | 98.2
1.9 | 97.45 | 97.86 | 97.7333 | 97.7667
2 | 97.04 | 97.6 | 97.4 | 97.3333
2.1 | 96.5 | 97.0429 | 97.06 | 97
2.2 | 96 | 96.62 | 96.65 | 96.7
2.3 | 95.52 | 96.0833 | 96.34 | 96.2667
2.4 | 95.1 | 95.8 | 95.8714 | 95.8
2.5 | 94.6 | 95.3429 | 95.575 | 95.56
2.6 | 93.9333 | 94.9 | 95.2 | 95.08
2.7 | 93.55 | 94.5286 | 94.68 | 94.84
2.8 | 93.04 | 94.05 | 94.3143 | 94.5333
2.9 | 92.4333 | 93.575 | 93.9 | 93.9667
3 | 91.94 | 93.18 | 93.6857 | 93.6
3.1 | 91.44 | 92.75 | 93.1857 | 93.3
3.2 | 90.9 | 92.28 | 92.75 | 92.9
3.3 | 90.44 | 91.8444 | 92.3667 | 92.52
3.4 | 90.04 | 91.5571 | 92.05 | 92.175
3.5 | 89.5 | 91.06 | 91.65 | 91.82
3.6 | 88.92 | 90.6714 | 91.32 | 91.52
3.7 | 88.525 | 90.26 | 91 | 91.14
3.8 | 87.88 | 89.8 | 90.5429 | 90.74
3.9 | 87.42 | 89.4143 | 90.0667 | 90.32
4 | 86.8 | 88.98 | 89.78 | 89.9333
4.1 | 86.22 | 88.45 | 89.3 | 89.56
4.2 | 85.9 | 88.125 | 89.08 | 89.2667
4.3 | 85.5 | 87.72 | 88.75 | 88.9
4.4 | 84.8 | 87.2333 | 88.2875 | 88.5
4.5 | 84.4 | 86.75 | 87.825 | 88.2
4.6 | 83.85 | 86.425 | 87.4 | 87.78
4.7 | 83.46 | 85.98 | 87.14 | 87.38
4.8 | 83.125 | 85.65 | 86.84 | 87
4.9 | 82.58 | 85.3 | 86.3857 | 86.6
5 | 81.875 | 84.7 | 86 | 86.32
5.1 | 81.3857 | 84.35 | 85.65 | 86.025
5.2 | 81 | 83.98 | 85.34 | 85.6
5.3 | 80.45 | 83.5 | 84.95 | 85.3571
5.4 | 79.98 | 83.02 | 84.56 | 84.88
5.5 | 79.4286 | 82.64 | 84.2 | 84.48
5.6 | 79.125 | 82.3333 | 83.8 | 84.12
5.7 | 78.45 | 81.875 | 83.32 | 83.74
5.8 | 78.14 | 81.42 | 82.88 | 83.42
5.9 | 77.65 | 81.06 | 82.6 | 83.02
6 | 77.1 | 80.62 | 82.24 | 82.725
6.1 | 76.68 | 80.16 | 81.72 | 82.34
6.2 | 76 | 79.55 | 81.5 | 82.06
6.3 | 75.74 | 79.36 | 81.24 | 81.6429
6.4 | 75.275 | 78.975 | 80.8 | 81.3
6.5 | 74.8 | 78.68 | 80.4429 | 80.9857
6.6 | 74.4 | 78.24 | 80.1333 | 80.44
6.7 | 73.8 | 77.7667 | 79.76 | 80.2
6.8 | 73.45 | 77.475 | 79.4667 | 79.76
6.9 | 72.7 | 77.1 | 79.12 | 79.45
7 | 72.3 | 76.6 | 78.8 | 79.15
7.1 | 71.92 | 76.08 | 78.25 | 78.78
7.2 | 71.52 | 75.72 | 77.8429 | 78.38
7.3 | 71.1 | 75.34 | 77.5571 | 78.0429
7.4 | 70.46 | 74.84 | 77.04 | 77.8286
7.5 | 70.22 | 74.52 | 76.7 | 77.35
7.6 | 69.88 | 74.12 | 76.5143 | 76.9
7.7 | 69.3 | 73.7 | 76.1571 | 76.625
7.8 | 68.84 | 73.38 | 75.6875 | 76.3
7.9 | 68.3 | 72.95 | 75.4 | 75.8857
8 | 67.85 | 72.54 | 74.9667 | 75.64
8.1 | 67.5 | 72.1 | 74.6125 | 75.2714
8.2 | 67.12 | 71.78 | 74.34 | 74.86
8.3 | 66.65 | 71.3667 | 73.94 | 74.5
8.4 | 66.3 | 71.05 | 73.575 | 74.22
8.5 | 65.9 | 70.75 | 73.25 | 73.85
8.6 | 65.4 | 70.28 | 72.9 | 73.575
8.7 | 64.95 | 69.9 | 72.6333 | 73.3
8.8 | 64.6 | 69.42 | 72.375 | 72.925
8.9 | 64.32 | 69 | 72.025 | 72.6
9 | 63.8286 | 68.7 | 71.64 | 72.24
9.1 | 63.375 | 68.38 | 71.2 | 72.02
9.2 | 63.05 | 67.94 | 70.92 | 71.48
9.3 | 62.75 | 67.55 | 70.5143 | 71.2
9.4 | 62.2143 | 67.3167 | 70.175 | 70.88
9.5 | 61.75 | 66.975 | 69.8429 | 70.55
9.6 | 61.3667 | 66.62 | 69.58 | 70.225
9.7 | 60.95 | 66.175 | 69.2143 | 69.85
9.8 | 60.5 | 65.78 | 68.8667 | 69.6
9.9 | 60.2 | 65.34 | 68.6375 | 69.3
10 | 59.8143 | 65.02 | 68.3333 | 69
10.1 | 59.4556 | 64.66 | 67.9714 | 68.65
10.2 | 59.06 | 64.4 | 67.7 | 68.2
10.3 | 58.725 | 64.06 | 67.14 | 68.0143
10.4 | 58.32 | 63.62 | 66.9 | 67.8
10.5 | 57.98 | 63.28 | 66.5714 | 67.4286
10.6 | 57.4714 | 62.9 | 66.3 | 67
10.7 | 57.2 | 62.58 | 65.875 | 66.7714
10.8 | 56.875 | 62.3 | 65.6 | 66.4
10.9 | 56.4714 | 61.86 | 65.34 | 66.08
11 | 56.0429 | 61.4667 | 65.05 | 65.8333
11.1 | 55.8 | 61.2 | 64.66 | 65.54
11.2 | 55.35 | 60.8667 | 64.3667 | 65.2
11.3 | 55.06 | 60.6 | 64.1 | 64.94
11.4 | 54.66 | 60.3 | 63.8667 | 64.55
11.5 | 54.275 | 60 | 63.5667 | 64.24
11.6 | 54.025 | 59.5 | 63.2 | 63.925
11.7 | 53.7 | 59.1 | 62.8 | 63.65
11.8 | 53.325 | 58.8 | 62.45 | 63.45
11.9 | 52.9667 | 58.52 | 62.15 | 63.12
12 | 52.725 | 58.26 | 61.76 | 62.8
12.1 | 52.3 | 57.84 | 61.55 | 62.375
12.2 | 52.06 | 57.4 | 61.275 | 62.1
12.3 | 51.6 | 57.12 | 60.96 | 61.84
12.4 | 51.26 | 56.84 | 60.725 | 61.55
12.5 | 50.9 | 56.48 | 60.4857 | 61.25
12.6 | 50.56 | 56.2667 | 60.1 | 60.9333
12.7 | 50.2 | 55.925 | 59.725 | 60.64
12.8 | 49.8714 | 55.64 | 59.35 | 60.4333
12.9 | 49.5857 | 55.225 | 59.1429 | 60.175
13 | 49.35 | 54.92 | 58.85 | 59.8
13.1 | 49.05 | 54.64 | 58.54 | 59.52
13.2 | 48.6 | 54.28 | 58.3 | 59.2667
13.3 | 48.4714 | 54.0333 | 57.9143 | 59
13.4 | 48.1 | 53.725 | 57.7 | 58.58
13.5 | 47.9 | 53.54 | 57.4571 | 58.3333
13.6 | 47.5 | 53.15 | 57.16 | 57.98
13.7 | 47.16 | 52.8667 | 56.9143 | 57.775
13.8 | 46.85 | 52.56 | 56.62 | 57.525
13.9 | 46.575 | 52.3 | 56.2667 | 57.3
14 | 46.34 | 51.975 | 56.05 | 57.04
14.1 | 46 | 51.74 | 55.7429 | 56.725
14.2 | 45.675 | 51.38 | 55.375 | 56.5
14.3 | 45.4 | 50.9 | 55.1333 | 56.2
14.4 | 45.075 | 50.7 | 54.88 | 55.8
14.5 | 44.9 | 50.5 | 54.6333 | 55.6
14.6 | 44.58 | 50.18 | 54.34 | 55.4
14.7 | 44.35 | 49.775 | 53.9833 | 55.0667
14.8 | 44.025 | 49.45 | 53.76 | 54.74
14.9 | 43.75 | 49.26 | 53.5 | 54.6
15 | 43.425 | 49 | 53.2 | 54.3
15.1 | 43.16 | 48.725 | 52.95 | 54
15.2 | 42.84 | 48.4333 | 52.725 | 53.6857
15.3 | 42.5429 | 48.16 | 52.4714 | 53.52
15.4 | 42.3 | 47.92 | 52.3 | 53.2571
15.5 | 42.1 | 47.675 | 51.975 | 53
15.6 | 41.7333 | 47.42 | 51.64 | 52.6714
15.7 | 41.5 | 47.1333 | 51.3667 | 52.42
15.8 | 41.2286 | 46.74 | 51.15 | 52.2
15.9 | 41 | 46.48 | 50.85 | 51.96
16 | 40.7167 | 46.18 | 50.6 | 51.6
16.1 | 40.4 | 45.9333 | 50.3333 | 51.3571
16.2 | 40.26 | 45.8333 | 50.1 | 51.22
16.3 | 39.9 | 45.5 | 49.7571 | 50.9571
16.4 | 39.68 | 45.18 | 49.55 | 50.72
16.5 | 39.4714 | 44.9 | 49.28 | 50.48
16.6 | 39.2571 | 44.68 | 49.02 | 50.22
16.7 | 38.94 | 44.44 | 48.78 | 49.96
16.8 | 38.78 | 44.28 | 48.55 | 49.6714
16.9 | 38.55 | 43.9 | 48.35 | 49.4
17 | 38.18 | 43.7 | 48.125 | 49.1714
17.1 | 37.9667 | 43.44 | 47.9 | 48.9429
17.2 | 37.7 | 43.2 | 47.5333 | 48.64
17.3 | 37.4143 | 42.875 | 47.3 | 48.46
17.4 | 37.26 | 42.66 | 47 | 48.24
17.5 | 36.94 | 42.3333 | 46.8 | 47.9333
17.6 | 36.8 | 42.12 | 46.55 | 47.75
17.7 | 36.5 | 41.8 | 46.32 | 47.4333
17.8 | 36.2286 | 41.56 | 46.05 | 47.2
17.9 | 36.1 | 41.36 | 45.82 | 46.95
18 | 35.8286 | 41.15 | 45.6 | 46.76
18.1 | 35.6 | 40.95 | 45.46 | 46.525
18.2 | 35.3857 | 40.72 | 45.25 | 46.4
18.3 | 35.2 | 40.48 | 44.925 | 46.1333
18.4 | 34.95 | 40.2667 | 44.72 | 45.92
18.5 | 34.675 | 40.14 | 44.5 | 45.6
18.6 | 34.52 | 39.68 | 44.28 | 45.5
18.7 | 34.3286 | 39.5333 | 44.05 | 45.2571
18.8 | 34.15 | 39.325 | 43.82 | 44.98
18.9 | 33.8333 | 39.14 | 43.62 | 44.75
19 | 33.525 | 38.82 | 43.38 | 44.525
19.1 | 33.4333 | 38.6833 | 43.2 | 44.3333
19.2 | 33.18 | 38.45 | 42.9 | 44.05
19.3 | 32.9667 | 38.24 | 42.675 | 43.8429
19.4 | 32.775 | 38.1 | 42.44 | 43.6
19.5 | 32.6 | 37.8 | 42.2 | 43.45
19.6 | 32.375 | 37.6 | 42 | 43.225
19.7 | 32.14 | 37.425 | 41.8 | 42.975
19.8 | 31.9429 | 37.2 | 41.5167 | 42.8
19.9 | 31.8 | 36.9 | 41.36 | 42.5833
20 | 31.4889 | 36.7 | 41.2 | 42.4
20.1 | 31.35 | 36.55 | 41 | 42.1714
20.2 | 31.15 | 36.3 | 40.8 | 42
20.3 | 31 | 35.95 | 40.54 | 41.6857
20.4 | 30.8778 | 35.8 | 40.22 | 41.52
20.5 | 30.66 | 35.6714 | 40.075 | 41.3
20.6 | 30.4571 | 35.45 | 39.9 | 41.06
20.7 | 30.3143 | 35.22 | 39.64 | 40.94
20.8 | 30.1 | 34.98 | 39.5 | 40.7
20.9 | 29.82 | 34.8 | 39.3667 | 40.56
21 | 29.675 | 34.68 | 39.125 | 40.3
21.1 | 29.45 | 34.4 | 38.9333 | 40.18
21.2 | 29.2857 | 34.1857 | 38.74 | 39.925
21.3 | 29.12 | 34 | 38.54 | 39.68
21.4 | 28.925 | 33.88 | 38.2333 | 39.46
21.5 | 28.7 | 33.65 | 38.08 | 39.3
21.6 | 28.5 | 33.45 | 37.8 | 39.2
21.7 | 28.3714 | 33.25 | 37.7 | 38.9
21.8 | 28.175 | 32.9571 | 37.4667 | 38.7714
21.9 | 28 | 32.775 | 37.24 | 38.58
22 | 27.8 | 32.6571 | 37.1 | 38.3
22.1 | 27.7 | 32.475 | 36.88 | 38.18
22.2 | 27.54 | 32.3 | 36.64 | 37.9667
22.3 | 27.34 | 32.15 | 36.5 | 37.7286
22.4 | 27.15 | 31.8667 | 36.4 | 37.6
22.5 | 27.04 | 31.675 | 36.1 | 37.3857
22.6 | 26.8556 | 31.5 | 35.9 | 37.1
22.7 | 26.7 | 31.3571 | 35.8 | 36.9333
22.8 | 26.4143 | 31.125 | 35.6 | 36.725
22.9 | 26.2857 | 31 | 35.4 | 36.6
23 | 26.125 | 30.875 | 35.3 | 36.38
23.1 | 26 | 30.625 | 34.96 | 36.26
23.2 | 25.8167 | 30.5333 | 34.76 | 36.04
23.3 | 25.7 | 30.325 | 34.56 | 35.88
23.4 | 25.52 | 30.0571 | 34.4 | 35.7
23.5 | 25.4 | 29.9667 | 34.3 | 35.52
23.6 | 25.22 | 29.7857 | 34.04 | 35.3286
23.7 | 25.0167 | 29.54 | 33.8667 | 35.24
23.8 | 24.9 | 29.35 | 33.7 | 34.9667
23.9 | 24.7286 | 29.2625 | 33.575 | 34.775
24 | 24.625 | 29.1 | 33.38 | 34.64
24.1 | 24.4333 | 28.9143 | 33.18 | 34.48
24.2 | 24.225 | 28.7 | 33 | 34.3
24.3 | 24.1286 | 28.6167 | 32.84 | 34.1
24.4 | 23.925 | 28.4 | 32.72 | 34
24.5 | 23.8 | 28.3286 | 32.5667 | 33.8
24.6 | 23.7 | 28.075 | 32.4 | 33.7
24.7 | 23.5667 | 27.875 | 32.2 | 33.475
24.8 | 23.4 | 27.72 | 32.05 | 33.26
24.9 | 23.2778 | 27.6 | 31.8 | 33.1
25 | 23.2 | 27.42 | 31.7 | 32.9
25.1 | 22.9857 | 27.3 | 31.5 | 32.7
25.2 | 22.8 | 27.12 | 31.4 | 32.6
25.3 | 22.6857 | 26.98 | 31.1 | 32.4
25.4 | 22.6 | 26.8 | 31 | 32.32
25.5 | 22.375 | 26.675 | 30.775 | 32.2
25.6 | 22.25 | 26.425 | 30.66 | 31.92
25.7 | 22.1857 | 26.325 | 30.5 | 31.7
25.8 | 22.02 | 26.2 | 30.32 | 31.58
25.9 | 21.8167 | 26.1 | 30.2 | 31.45
26 | 21.7 | 25.85 | 30.02 | 31.25
26.1 | 21.5444 | 25.74 | 29.88 | 31.15
26.2 | 21.4 | 25.6667 | 29.7 | 30.98
26.3 | 21.28 | 25.44 | 29.625 | 30.86
26.4 | 21.2 | 25.3 | 29.42 | 30.7
26.5 | 21.1286 | 25.2 | 29.2 | 30.575
26.6 | 20.925 | 25.05 | 29.075 | 30.34
26.7 | 20.78 | 24.85 | 28.925 | 30.16
26.8 | 20.6833 | 24.725 | 28.8 | 30.04
26.9 | 20.5286 | 24.66 | 28.7 | 29.9
27 | 20.425 | 24.4333 | 28.5 | 29.8
27.1 | 20.3111 | 24.36 | 28.35 | 29.625
27.2 | 20.1857 | 24.15 | 28.2 | 29.4
27.3 | 20.0429 | 24.06 | 28.0833 | 29.2571
27.4 | 19.94 | 23.925 | 27.85 | 29.125
27.5 | 19.82 | 23.8 | 27.7333 | 29
27.6 | 19.7 | 23.6 | 27.62 | 28.86
27.7 | 19.54 | 23.4667 | 27.5 | 28.75
27.8 | 19.4111 | 23.3 | 27.32 | 28.4833
27.9 | 19.3 | 23.175 | 27.2167 | 28.4
28 | 19.1833 | 23 | 27 | 28.2857
28.1 | 19.1 | 22.9 | 26.86 | 28.2
28.2 | 18.9429 | 22.74 | 26.725 | 27.925
28.3 | 18.86 | 22.68 | 26.62 | 27.8
28.4 | 18.7667 | 22.5 | 26.4 | 27.7
28.5 | 18.6 | 22.4 | 26.32 | 27.6
28.6 | 18.5 | 22.3 | 26.18 | 27.45
28.7 | 18.4 | 22.2 | 26.075 | 27.24
28.8 | 18.325 | 22.1 | 25.94 | 27.1
28.9 | 18.2333 | 21.9 | 25.725 | 26.94
29 | 18.1 | 21.8 | 25.6 | 26.9
29.1 | 17.9714 | 21.65 | 25.5 | 26.74
29.2 | 17.8 | 21.54 | 25.3 | 26.575
29.3 | 17.7125 | 21.5 | 25.2 | 26.4
29.4 | 17.6 | 21.3 | 25.1 | 26.275
29.5 | 17.5 | 21.2 | 24.95 | 26.14
29.6 | 17.4333 | 21.1 | 24.8 | 26.075
29.7 | 17.3 | 20.9 | 24.7 | 25.9333
29.8 | 17.1714 | 20.8 | 24.5 | 25.8
29.9 | 17.1 | 20.7 | 24.4 | 25.6667
30 | 17 | 20.5 | 24.3 | 25.56

## SSD80 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 49.8 | 55.7 | 65.3 | 71.4
0.1 | 54.4 | 59.8 | 69.35 | 74.9
0.2 | 62.25 | 67.1333 | 75.2333 | 79.9
0.3 | 73.65 | 77.4 | 83.475 | 86.95
0.4 | 82.875 | 85.3 | 89.4286 | 92
0.5 | 89.35 | 90.95 | 93.98 | 95.5
0.6 | 93.4 | 94.9 | 96.6 | 97.8
0.7 | 96.475 | 97.275 | 98.4333 | 99.05
0.8 | 98.2 | 98.7 | 99.375 | 99.86
0.9 | 99.3 | 99.58 | 99.96 | 100
1 | 99.9 | 99.9 | 100 | 100
1.1 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 99.78 | 99.825
1.3 | 99.85 | 99.9 | 99.64 | 99.55
1.4 | 99.5 | 99.5667 | 99.3 | 99.125
1.5 | 99.1 | 99.1333 | 98.85 | 98.675
1.6 | 98.5 | 98.88 | 98.4 | 98.275
1.7 | 98.12 | 98.2333 | 98.02 | 97.8667
1.8 | 97.5 | 97.8 | 97.65 | 97.55
1.9 | 96.95 | 97.22 | 97.15 | 96.98
2 | 96.5 | 96.825 | 96.65 | 96.65
2.1 | 95.925 | 96.3 | 96.25 | 96.1
2.2 | 95.35 | 95.825 | 95.72 | 95.7
2.3 | 94.65 | 95.35 | 95.3 | 95.16
2.4 | 94.05 | 94.85 | 94.8667 | 94.7
2.5 | 93.6333 | 94.45 | 94.4 | 94.3333
2.6 | 93 | 93.75 | 93.9 | 93.9
2.7 | 92.3333 | 93.25 | 93.4 | 93.4
2.8 | 91.8333 | 93 | 92.94 | 93.1
2.9 | 91.26 | 92.4375 | 92.55 | 92.52
3 | 90.7 | 91.8667 | 92.075 | 92.15
3.1 | 90.1 | 91.4 | 91.56 | 91.675
3.2 | 89.55 | 90.9 | 91.15 | 91.25
3.3 | 89 | 90.4667 | 90.7 | 90.675
3.4 | 88.3 | 89.975 | 90.4 | 90.38
3.5 | 87.74 | 89.5 | 89.9 | 89.95
3.6 | 87.35 | 89.025 | 89.3667 | 89.55
3.7 | 86.65 | 88.625 | 89 | 89.1
3.8 | 86.125 | 88.05 | 88.52 | 88.7
3.9 | 85.45 | 87.55 | 88.1 | 88.3
4 | 84.925 | 87.1 | 87.65 | 87.8
4.1 | 84.425 | 86.6667 | 87.225 | 87.4
4.2 | 83.9 | 86.1 | 86.775 | 86.95
4.3 | 83.25 | 85.7333 | 86.32 | 86.6
4.4 | 82.86 | 85.1 | 85.95 | 86
4.5 | 82.4 | 84.8 | 85.54 | 85.6
4.6 | 81.7 | 84.35 | 85.1 | 85.3
4.7 | 81.15 | 83.75 | 84.55 | 84.76
4.8 | 80.6 | 83.2667 | 84.2 | 84.35
4.9 | 80.1 | 82.8 | 83.75 | 83.975
5 | 79.6333 | 82.35 | 83.3 | 83.55
5.1 | 78.9667 | 81.9 | 82.8 | 83.2333
5.2 | 78.4125 | 81.4 | 82.4 | 82.68
5.3 | 77.95 | 80.9667 | 82.1 | 82.4
5.4 | 77.38 | 80.475 | 81.72 | 81.85
5.5 | 76.84 | 79.975 | 81.2 | 81.525
5.6 | 76.4 | 79.5 | 80.95 | 81.1
5.7 | 75.9 | 79.05 | 80.4 | 80.7
5.8 | 75.3 | 78.7 | 79.95 | 80.3375
5.9 | 74.8 | 78.2 | 79.5 | 80.0375
6 | 74.44 | 77.8 | 79.0667 | 79.45
6.1 | 73.925 | 77.25 | 78.65 | 79.08
6.2 | 73.35 | 76.8 | 78.3 | 78.7333
6.3 | 72.85 | 76.42 | 78 | 78.3286
6.4 | 72.375 | 75.825 | 77.5 | 77.9375
6.5 | 71.75 | 75.375 | 77 | 77.5
6.6 | 71.35 | 75.05 | 76.6 | 77.1
6.7 | 70.94 | 74.65 | 76.425 | 76.65
6.8 | 70.45 | 74.05 | 75.85 | 76.32
6.9 | 69.9 | 73.675 | 75.475 | 75.9
7 | 69.35 | 73.25 | 75.025 | 75.525
7.1 | 68.9 | 72.9 | 74.64 | 75.1
7.2 | 68.5 | 72.45 | 74.15 | 74.75
7.3 | 68 | 71.9 | 73.8667 | 74.325
7.4 | 67.4 | 71.5 | 73.4 | 73.95
7.5 | 67 | 71.1 | 73.15 | 73.625
7.6 | 66.55 | 70.6 | 72.65 | 73.25
7.7 | 66.15 | 70.14 | 72.2125 | 72.74
7.8 | 65.55 | 69.775 | 71.8375 | 72.4
7.9 | 65.15 | 69.425 | 71.4625 | 72.1
8 | 64.75 | 68.95 | 71.05 | 71.6
8.1 | 64.225 | 68.45 | 70.6333 | 71.2333
8.2 | 63.75 | 68.2 | 70.275 | 70.9667
8.3 | 63.35 | 67.65 | 69.8667 | 70.5
8.4 | 62.86 | 67.325 | 69.5667 | 70.1
8.5 | 62.55 | 66.9333 | 69.2 | 69.75
8.6 | 62.1 | 66.5333 | 68.875 | 69.525
8.7 | 61.7 | 66.05 | 68.525 | 69.06
8.8 | 61.4 | 65.6667 | 68.12 | 68.725
8.9 | 60.75 | 65.3 | 67.7375 | 68.45
9 | 60.44 | 64.8 | 67.375 | 68.05
9.1 | 60.12 | 64.4 | 66.975 | 67.675
9.2 | 59.65 | 64.1 | 66.55 | 67.35
9.3 | 59.2286 | 63.7 | 66.15 | 66.875
9.4 | 58.85 | 63.32 | 65.8333 | 66.6
9.5 | 58.3875 | 62.95 | 65.56 | 66.25
9.6 | 57.9429 | 62.5 | 65.15 | 65.9
9.7 | 57.4667 | 62.1667 | 64.9 | 65.55
9.8 | 57.075 | 61.7 | 64.5 | 65.2
9.9 | 56.7875 | 61.3 | 64.1 | 64.95
10 | 56.4571 | 60.8667 | 63.8 | 64.5
10.1 | 56 | 60.7 | 63.3333 | 64.15
10.2 | 55.65 | 60.25 | 63 | 63.8
10.3 | 55.275 | 59.9 | 62.75 | 63.5667
10.4 | 54.9 | 59.55 | 62.375 | 63.1
10.5 | 54.5 | 59.2 | 61.9 | 62.8
10.6 | 54.05 | 58.8333 | 61.675 | 62.45
10.7 | 53.675 | 58.375 | 61.325 | 62.15
10.8 | 53.3 | 58.025 | 60.9 | 61.8
10.9 | 52.925 | 57.7 | 60.65 | 61.4
11 | 52.55 | 57.34 | 60.35 | 61.15
11.1 | 52.25 | 56.9 | 60.0333 | 60.8667
11.2 | 52 | 56.55 | 59.6333 | 60.525
11.3 | 51.5571 | 56.25 | 59.3 | 60.2
11.4 | 51.15 | 55.875 | 59.0333 | 59.875
11.5 | 50.775 | 55.5 | 58.725 | 59.5333
11.6 | 50.45 | 55.25 | 58.4667 | 59.25
11.7 | 50.1571 | 54.95 | 57.95 | 58.9
11.8 | 49.8111 | 54.5429 | 57.6 | 58.6333
11.9 | 49.5125 | 54.3 | 57.5333 | 58.2714
12 | 49.1375 | 53.9 | 57.05 | 57.9286
12.1 | 48.7333 | 53.4 | 56.75 | 57.625
12.2 | 48.4 | 53.1 | 56.4 | 57.3
12.3 | 48 | 52.8 | 56.05 | 57
12.4 | 47.64 | 52.45 | 55.8 | 56.65
12.5 | 47.4 | 52.2 | 55.45 | 56.4
12.6 | 47.05 | 51.75 | 55.15 | 55.975
12.7 | 46.8 | 51.575 | 54.8 | 55.725
12.8 | 46.45 | 51.1 | 54.5 | 55.5
12.9 | 46.1 | 50.9 | 54.2333 | 55.175
13 | 45.8 | 50.5 | 53.9 | 54.925
13.1 | 45.5 | 50.225 | 53.575 | 54.575
13.2 | 45.2 | 49.9 | 53.25 | 54.25
13.3 | 44.825 | 49.6667 | 53 | 53.9
13.4 | 44.55 | 49.3 | 52.825 | 53.725
13.5 | 44.2 | 49.05 | 52.375 | 53.45
13.6 | 43.975 | 48.6 | 52.125 | 53.05
13.7 | 43.7 | 48.4 | 51.8 | 52.875
13.8 | 43.45 | 48.175 | 51.64 | 52.64
13.9 | 43 | 47.825 | 51.2 | 52.2286
14 | 42.725 | 47.475 | 50.95 | 52
14.1 | 42.55 | 47.125 | 50.6 | 51.675
14.2 | 42.2667 | 46.85 | 50.3714 | 51.425
14.3 | 41.9 | 46.55 | 50.1333 | 51.075
14.4 | 41.6 | 46.4 | 49.85 | 50.84
14.5 | 41.375 | 46 | 49.6 | 50.58
14.6 | 41.15 | 45.7 | 49.275 | 50.325
14.7 | 40.9 | 45.5 | 49 | 49.975
14.8 | 40.55 | 45.2 | 48.75 | 49.7714
14.9 | 40.2 | 44.85 | 48.4667 | 49.48
15 | 39.85 | 44.6 | 48.175 | 49.25
15.1 | 39.725 | 44.35 | 47.925 | 48.9
15.2 | 39.35 | 44 | 47.5857 | 48.7
15.3 | 39.125 | 43.75 | 47.4 | 48.4
15.4 | 38.875 | 43.525 | 47.1 | 48.15
15.5 | 38.625 | 43.275 | 46.9 | 47.8
15.6 | 38.375 | 42.95 | 46.55 | 47.65
15.7 | 38.125 | 42.6667 | 46.2 | 47.4
15.8 | 37.875 | 42.3667 | 46.05 | 47.15
15.9 | 37.625 | 42.1 | 45.8 | 46.92
16 | 37.375 | 41.8667 | 45.575 | 46.675
16.1 | 37.125 | 41.5 | 45.2667 | 46.4
16.2 | 36.8857 | 41.3 | 44.9667 | 46.1
16.3 | 36.6333 | 41.1 | 44.675 | 45.86
16.4 | 36.3333 | 40.8667 | 44.5 | 45.575
16.5 | 36.1143 | 40.575 | 44.25 | 45.275
16.6 | 35.9 | 40.325 | 44.0333 | 45.025
16.7 | 35.65 | 40.1 | 43.8 | 44.875
16.8 | 35.4 | 39.8429 | 43.45 | 44.625
16.9 | 35.15 | 39.6286 | 43.225 | 44.3667
17 | 34.9 | 39.3667 | 42.975 | 44.05
17.1 | 34.7333 | 39.1667 | 42.825 | 43.9
17.2 | 34.45 | 38.85 | 42.5 | 43.7
17.3 | 34.25 | 38.625 | 42.325 | 43.3333
17.4 | 33.95 | 38.275 | 42.075 | 43.1
17.5 | 33.65 | 38.025 | 41.8 | 42.88
17.6 | 33.5 | 37.825 | 41.475 | 42.6
17.7 | 33.1875 | 37.5 | 41.325 | 42.35
17.8 | 33.0143 | 37.4 | 41.075 | 42.2
17.9 | 32.8 | 37.1333 | 40.85 | 41.95
18 | 32.5375 | 36.8857 | 40.5 | 41.72
18.1 | 32.275 | 36.7 | 40.3 | 41.45
18.2 | 32.125 | 36.46 | 40.18 | 41.25
18.3 | 31.9375 | 36.2 | 39.88 | 41.05
18.4 | 31.7833 | 35.96 | 39.7 | 40.875
18.5 | 31.4857 | 35.7444 | 39.45 | 40.625
18.6 | 31.3143 | 35.6 | 39.3 | 40.375
18.7 | 31.1 | 35.35 | 39.05 | 40.125
18.8 | 30.85 | 35.1 | 38.9 | 39.975
18.9 | 30.7 | 34.85 | 38.6 | 39.725
19 | 30.55 | 34.7 | 38.4 | 39.55
19.1 | 30.3 | 34.46 | 38.16 | 39.225
19.2 | 30.05 | 34.2 | 37.9 | 39.0667
19.3 | 29.9 | 34.05 | 37.7 | 38.875
19.4 | 29.65 | 33.8 | 37.525 | 38.625
19.5 | 29.5 | 33.5667 | 37.275 | 38.35
19.6 | 29.375 | 33.325 | 37.025 | 38.2
19.7 | 29.12 | 33.175 | 36.875 | 38
19.8 | 28.8667 | 32.9625 | 36.625 | 37.775
19.9 | 28.72 | 32.775 | 36.4 | 37.58
20 | 28.475 | 32.6125 | 36.225 | 37.325
20.1 | 28.325 | 32.375 | 36 | 37.15
20.2 | 28.175 | 32.1625 | 35.9 | 36.9
20.3 | 27.925 | 31.9889 | 35.6667 | 36.7667
20.4 | 27.775 | 31.75 | 35.4 | 36.6
20.5 | 27.525 | 31.5 | 35.2 | 36.4
20.6 | 27.375 | 31.375 | 35 | 36.125
20.7 | 27.225 | 31.15 | 34.7667 | 35.975
20.8 | 27.1 | 30.925 | 34.6 | 35.85
20.9 | 26.8 | 30.8 | 34.4 | 35.6
21 | 26.68 | 30.65 | 34.3 | 35.4
21.1 | 26.5 | 30.4 | 34.1 | 35.2
21.2 | 26.3 | 30.225 | 33.8 | 35
21.3 | 26.15 | 30 | 33.6 | 34.7667
21.4 | 25.9 | 29.8556 | 33.46 | 34.62
21.5 | 25.85 | 29.6571 | 33.2571 | 34.475
21.6 | 25.6 | 29.4667 | 33.075 | 34.225
21.7 | 25.45 | 29.225 | 32.9 | 34.075
21.8 | 25.3 | 29.075 | 32.66 | 33.825
21.9 | 25.1 | 28.8625 | 32.425 | 33.6
22 | 24.9333 | 28.675 | 32.275 | 33.5
22.1 | 24.76 | 28.5125 | 32.075 | 33.275
22.2 | 24.625 | 28.35 | 31.85 | 33.025
22.3 | 24.475 | 28.2 | 31.7125 | 32.9
22.4 | 24.3 | 28 | 31.5 | 32.8
22.5 | 24.16 | 27.825 | 31.4125 | 32.6
22.6 | 24.025 | 27.675 | 31.2 | 32.35
22.7 | 23.8 | 27.425 | 30.975 | 32.2
22.8 | 23.7 | 27.3 | 30.85 | 32
22.9 | 23.4667 | 27.125 | 30.65 | 31.9
23 | 23.3429 | 27 | 30.5 | 31.65
23.1 | 23.2 | 26.775 | 30.35 | 31.5
23.2 | 23.075 | 26.625 | 30.15 | 31.3
23.3 | 22.9 | 26.425 | 30.025 | 31.15
23.4 | 22.75 | 26.275 | 29.8 | 31
23.5 | 22.625 | 26.15 | 29.575 | 30.8
23.6 | 22.4 | 26 | 29.4 | 30.6667
23.7 | 22.275 | 25.8 | 29.2875 | 30.45
23.8 | 22.2 | 25.625 | 29.1625 | 30.325
23.9 | 22 | 25.5 | 28.9571 | 30.1
24 | 21.85 | 25.4125 | 28.8 | 29.925
24.1 | 21.7 | 25.2 | 28.5667 | 29.8
24.2 | 21.6 | 25.0333 | 28.5 | 29.625
24.3 | 21.4 | 24.9 | 28.3 | 29.475
24.4 | 21.3 | 24.75 | 28.125 | 29.3
24.5 | 21.15 | 24.58 | 27.98 | 29.1
24.6 | 21 | 24.4 | 27.85 | 29
24.7 | 20.9 | 24.3 | 27.7 | 28.8
24.8 | 20.72 | 24.15 | 27.5143 | 28.65
24.9 | 20.5667 | 24 | 27.375 | 28.5
25 | 20.5 | 23.8 | 27.1667 | 28.35
25.1 | 20.35 | 23.7 | 27 | 28.2
25.2 | 20.2 | 23.52 | 26.9 | 28
25.3 | 20.1 | 23.4 | 26.7 | 27.9
25.4 | 20 | 23.22 | 26.6 | 27.76
25.5 | 19.825 | 23.06 | 26.475 | 27.525
25.6 | 19.675 | 23 | 26.3 | 27.4
25.7 | 19.525 | 22.8 | 26.1 | 27.325
25.8 | 19.4 | 22.7 | 26.025 | 27.075
25.9 | 19.3625 | 22.575 | 25.775 | 26.9875
26 | 19.1875 | 22.4 | 25.7 | 26.8625
26.1 | 19.0625 | 22.2 | 25.575 | 26.675
26.2 | 18.8857 | 22.1 | 25.34 | 26.5125
26.3 | 18.8 | 21.975 | 25.3 | 26.3875
26.4 | 18.7 | 21.825 | 25.15 | 26.3
26.5 | 18.6 | 21.7 | 25 | 26.0875
26.6 | 18.46 | 21.6 | 24.8 | 25.9625
26.7 | 18.3 | 21.4 | 24.6 | 25.8375
26.8 | 18.2 | 21.35 | 24.5 | 25.6444
26.9 | 18.1 | 21.2 | 24.4 | 25.5
27 | 18 | 21.05 | 24.3 | 25.3714
27.1 | 17.9 | 20.9 | 24.1 | 25.2
27.2 | 17.75 | 20.8 | 24 | 25.025
27.3 | 17.6 | 20.6 | 23.78 | 25
27.4 | 17.5 | 20.55 | 23.7 | 24.8
27.5 | 17.4 | 20.4 | 23.525 | 24.7
27.6 | 17.3 | 20.3 | 23.3875 | 24.6
27.7 | 17.2 | 20.2 | 23.3 | 24.4143
27.8 | 17.1 | 20 | 23.1875 | 24.3
27.9 | 17 | 19.9 | 23.025 | 24.2
28 | 16.9 | 19.8 | 22.95 | 24
28.1 | 16.775 | 19.625 | 22.8125 | 23.9
28.2 | 16.625 | 19.5875 | 22.625 | 23.7667
28.3 | 16.5 | 19.4625 | 22.5 | 23.6
28.4 | 16.4 | 19.3375 | 22.4 | 23.4857
28.5 | 16.275 | 19.2125 | 22.275 | 23.32
28.6 | 16.2 | 19.0889 | 22.125 | 23.1875
28.7 | 16 | 18.975 | 21.975 | 23.05
28.8 | 15.95 | 18.85 | 21.9 | 22.925
28.9 | 15.9 | 18.725 | 21.8 | 22.875
29 | 15.75 | 18.6 | 21.65 | 22.725
29.1 | 15.7 | 18.55 | 21.5 | 22.6
29.2 | 15.6 | 18.4 | 21.4 | 22.525
29.3 | 15.5 | 18.325 | 21.3 | 22.3
29.4 | 15.35 | 18.2 | 21.15 | 22.2
29.5 | 15.3 | 18.075 | 21 | 22.1
29.6 | 15.2 | 17.9429 | 20.9 | 22
29.7 | 15.125 | 17.825 | 20.7 | 21.9
29.8 | 15 | 17.7 | 20.65 | 21.72
29.9 | 14.9 | 17.6 | 20.5 | 21.6
30 | 14.8 | 17.5125 | 20.45 | 21.5

## SSD120 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 51.4 | 56.4 | 63.3 | 67.2
0.1 | 57.05 | 61.9 | 68.1 | 71.6
0.2 | 66.9 | 71.3 | 75.9 | 79
0.3 | 77.15 | 79.94 | 83.9 | 86.1
0.4 | 85.4 | 87.3 | 89.475 | 91.3
0.5 | 90.8 | 91.65 | 93.625 | 94.56
0.6 | 94.3 | 95.25 | 96.16 | 96.875
0.7 | 96.9 | 97.325 | 98.125 | 98.45
0.8 | 98.38 | 98.9 | 98.9 | 99.4333
0.9 | 99.15 | 99.5286 | 99.5 | 99.7667
1 | 99.85 | 99.9 | 99.9 | 99.925
1.1 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 100 | 100
1.3 | 99.9 | 99.875 | 99.85 | 99.76
1.4 | 99.65 | 99.625 | 99.7 | 99.525
1.5 | 99.3 | 99.48 | 99.325 | 99.3
1.6 | 98.85 | 99.15 | 99.0111 | 98.9286
1.7 | 98.45 | 98.7 | 98.7125 | 98.7
1.8 | 98.2 | 98.5 | 98.325 | 98.2667
1.9 | 97.85 | 98.1 | 98.08 | 98.05
2 | 97.3 | 97.85 | 97.9 | 97.825
2.1 | 96.8 | 97.4 | 97.4 | 97.5333
2.2 | 96.3 | 96.95 | 97.15 | 97.125
2.3 | 95.8 | 96.5 | 96.8 | 96.7667
2.4 | 95.4 | 96.075 | 96.5 | 96.425
2.5 | 95.05 | 95.85 | 96.2 | 96.2
2.6 | 94.5 | 95.475 | 95.75 | 95.95
2.7 | 94.05 | 95.15 | 95.425 | 95.475
2.8 | 93.65 | 94.65 | 95.175 | 95.18
2.9 | 93.16 | 94.15 | 94.675 | 94.8
3 | 92.825 | 93.86 | 94.375 | 94.55
3.1 | 92.275 | 93.425 | 94.125 | 94.2
3.2 | 91.775 | 93.1 | 93.7 | 93.8
3.3 | 91.3333 | 92.85 | 93.2667 | 93.4
3.4 | 90.8 | 92.2 | 92.975 | 93.22
3.5 | 90.4333 | 92.02 | 92.7 | 92.8
3.6 | 89.85 | 91.6 | 92.3 | 92.52
3.7 | 89.5 | 91.3571 | 91.9 | 92.075
3.8 | 89.0333 | 91.05 | 91.575 | 91.75
3.9 | 88.5333 | 90.475 | 91.3875 | 91.35
4 | 87.95 | 90.1125 | 90.95 | 91.1667
4.1 | 87.45 | 89.6125 | 90.5375 | 90.9
4.2 | 86.875 | 89.2375 | 90.2375 | 90.45
4.3 | 86.58 | 88.7625 | 89.725 | 90
4.4 | 86.1 | 88.4 | 89.475 | 89.76
4.5 | 85.7 | 88.1 | 89.1375 | 89.4286
4.6 | 85.1 | 87.7 | 88.775 | 88.9833
4.7 | 84.7 | 87.2 | 88.45 | 88.7333
4.8 | 84.2 | 86.8 | 88.0333 | 88.35
4.9 | 83.7333 | 86.4 | 87.75 | 87.9833
5 | 83.175 | 86.225 | 87.5 | 87.65
5.1 | 82.98 | 85.825 | 87.125 | 87.35
5.2 | 82.42 | 85.25 | 86.85 | 87.0875
5.3 | 81.95 | 84.9 | 86.425 | 86.7333
5.4 | 81.5 | 84.65 | 86.1 | 86.45
5.5 | 81.1 | 84.2 | 85.775 | 86.0667
5.6 | 80.7 | 83.7 | 85.4 | 85.6
5.7 | 80.125 | 83.4 | 85.0625 | 85.34
5.8 | 79.65 | 83.1 | 84.725 | 84.95
5.9 | 79.2 | 82.7 | 84.4125 | 84.725
6 | 78.75 | 82.175 | 84.0375 | 84.475
6.1 | 78.3 | 81.85 | 83.7625 | 84.1286
6.2 | 77.85 | 81.475 | 83.3667 | 83.675
6.3 | 77.5 | 81.225 | 82.95 | 83.35
6.4 | 77 | 81 | 82.65 | 83.0167
6.5 | 76.55 | 80.5 | 82.3 | 82.775
6.6 | 76.1 | 80.15 | 81.975 | 82.45
6.7 | 75.6 | 79.7429 | 81.7 | 82.175
6.8 | 75.2333 | 79.2 | 81.35 | 81.7
6.9 | 74.84 | 78.8714 | 80.95 | 81.4
7 | 74.3 | 78.6 | 80.65 | 81.075
7.1 | 73.9333 | 78.3 | 80.3 | 80.7
7.2 | 73.575 | 77.84 | 80.0143 | 80.55
7.3 | 73.1 | 77.45 | 79.7 | 80.2
7.4 | 72.5667 | 77 | 79.425 | 79.85
7.5 | 72.2333 | 76.575 | 79.04 | 79.45
7.6 | 71.775 | 76.2 | 78.7 | 79.15
7.7 | 71.45 | 76.0571 | 78.45 | 78.7
7.8 | 71.06 | 75.64 | 77.88 | 78.6
7.9 | 70.55 | 75.15 | 77.64 | 78.1667
8 | 70 | 74.825 | 77.4 | 77.85
8.1 | 69.7 | 74.5125 | 77.15 | 77.65
8.2 | 69.325 | 74.1714 | 76.8 | 77.4
8.3 | 68.75 | 73.65 | 76.45 | 77
8.4 | 68.4333 | 73.3429 | 76.1 | 76.6667
8.5 | 67.96 | 73.075 | 75.9 | 76.36
8.6 | 67.65 | 72.6111 | 75.4333 | 76.04
8.7 | 67.2 | 72.2 | 75.0667 | 75.6667
8.8 | 66.925 | 71.925 | 74.775 | 75.4
8.9 | 66.35 | 71.55 | 74.425 | 75.15
9 | 66.125 | 71.2286 | 74.15 | 74.66
9.1 | 65.575 | 70.74 | 73.8 | 74.35
9.2 | 65.2 | 70.475 | 73.5 | 74.06
9.3 | 64.8 | 70.05 | 73.1 | 73.825
9.4 | 64.3667 | 69.65 | 72.85 | 73.525
9.5 | 64.1 | 69.35 | 72.5 | 73.15
9.6 | 63.75 | 69.0667 | 72.16 | 72.9333
9.7 | 63.3 | 68.8 | 71.85 | 72.65
9.8 | 63 | 68.45 | 71.4667 | 72.15
9.9 | 62.55 | 68.0375 | 71.1 | 71.875
10 | 62.1 | 67.6667 | 70.8 | 71.4
10.1 | 61.7 | 67.3333 | 70.475 | 71.225
10.2 | 61.4 | 67.0714 | 70.325 | 70.975
10.3 | 61.1 | 66.74 | 70.0167 | 70.8
10.4 | 60.75 | 66.3 | 69.65 | 70.5
10.5 | 60.4 | 66.025 | 69.44 | 70.2
10.6 | 59.95 | 65.6857 | 69.075 | 69.88
10.7 | 59.6 | 65.225 | 68.825 | 69.25
10.8 | 59.25 | 65 | 68.55 | 69.025
10.9 | 58.8 | 64.7 | 68.1625 | 68.8
11 | 58.5667 | 64.4 | 67.75 | 68.64
11.1 | 58.14 | 64 | 67.46 | 68.325
11.2 | 57.7667 | 63.5667 | 67.1286 | 68.1
11.3 | 57.44 | 63.3 | 66.8714 | 67.8
11.4 | 57.075 | 63 | 66.54 | 67.45
11.5 | 56.925 | 62.65 | 66.2857 | 67.15
11.6 | 56.475 | 62.35 | 66.1 | 66.9
11.7 | 56.075 | 62.025 | 65.75 | 66.72
11.8 | 55.875 | 61.575 | 65.4 | 66.3
11.9 | 55.525 | 61.325 | 65.05 | 65.96
12 | 55.175 | 61.15 | 64.825 | 65.82
12.1 | 54.85 | 60.825 | 64.675 | 65.375
12.2 | 54.55 | 60.4 | 64.4 | 65.2
12.3 | 54.15 | 60.15 | 63.9667 | 64.95
12.4 | 53.775 | 59.8 | 63.725 | 64.625
12.5 | 53.425 | 59.5 | 63.45 | 64.3
12.6 | 53.1667 | 59.2 | 63.175 | 63.95
12.7 | 52.875 | 58.9 | 62.8 | 63.6714
12.8 | 52.5 | 58.55 | 62.5 | 63.35
12.9 | 52.2 | 58.275 | 62.3 | 63.0667
13 | 51.9 | 57.85 | 62.05 | 62.8
13.1 | 51.6 | 57.56 | 61.74 | 62.6556
13.2 | 51.35 | 57.4 | 61.5 | 62.425
13.3 | 51 | 57.0833 | 61.1 | 62.175
13.4 | 50.75 | 56.85 | 60.9 | 61.8
13.5 | 50.4 | 56.5 | 60.6 | 61.425
13.6 | 50.1 | 56.1375 | 60.3 | 61.2
13.7 | 49.7 | 55.8778 | 60 | 61
13.8 | 49.36 | 55.6429 | 59.7 | 60.7
13.9 | 49.125 | 55.3333 | 59.4 | 60.35
14 | 48.8 | 55 | 59.2 | 60.1
14.1 | 48.4667 | 54.625 | 58.86 | 59.85
14.2 | 48.175 | 54.375 | 58.54 | 59.6
14.3 | 47.825 | 54.1 | 58.2667 | 59.3
14.4 | 47.6 | 53.8 | 58.1 | 59
14.5 | 47.25 | 53.48 | 57.8 | 58.75
14.6 | 47 | 53.2 | 57.475 | 58.5
14.7 | 46.65 | 52.9 | 57.225 | 58.2
14.8 | 46.525 | 52.6 | 56.975 | 57.925
14.9 | 46.275 | 52.3667 | 56.725 | 57.775
15 | 46 | 52.1333 | 56.5 | 57.525
15.1 | 45.6333 | 51.8 | 56.225 | 57.275
15.2 | 45.35 | 51.56 | 55.975 | 56.95
15.3 | 44.98 | 51.3 | 55.725 | 56.675
15.4 | 44.85 | 50.9857 | 55.45 | 56.425
15.5 | 44.5 | 50.6333 | 55.1333 | 56.15
15.6 | 44.25 | 50.4625 | 54.9 | 55.925
15.7 | 44.1 | 50.125 | 54.8 | 55.8
15.8 | 43.8 | 49.875 | 54.46 | 55.45
15.9 | 43.5 | 49.7125 | 54.15 | 55.25
16 | 43.26 | 49.375 | 53.88 | 54.95
16.1 | 43 | 49.125 | 53.6 | 54.6
16.2 | 42.75 | 48.8125 | 53.4 | 54.35
16.3 | 42.425 | 48.6125 | 53.15 | 54.175
16.4 | 42.175 | 48.4 | 52.9 | 53.95
16.5 | 41.9 | 48.0875 | 52.55 | 53.7
16.6 | 41.675 | 47.7889 | 52.4 | 53.45
16.7 | 41.425 | 47.5286 | 52.2 | 53.3
16.8 | 41.175 | 47.25 | 51.9 | 52.9833
16.9 | 40.825 | 47.025 | 51.65 | 52.7833
17 | 40.675 | 46.8 | 51.4 | 52.5
17.1 | 40.425 | 46.55 | 51.15 | 52.225
17.2 | 40.175 | 46.3 | 50.925 | 52
17.3 | 39.925 | 46.05 | 50.675 | 51.85
17.4 | 39.675 | 45.7 | 50.42 | 51.64
17.5 | 39.425 | 45.55 | 50.1 | 51.4
17.6 | 39.28 | 45.3 | 49.975 | 50.9857
17.7 | 38.925 | 45.05 | 49.725 | 50.7667
17.8 | 38.775 | 44.9 | 49.45 | 50.55
17.9 | 38.55 | 44.5714 | 49.2 | 50.3333
18 | 38.4 | 44.3 | 48.975 | 50.2
18.1 | 38.05 | 44.1333 | 48.7333 | 49.95
18.2 | 37.8 | 43.8375 | 48.48 | 49.7
18.3 | 37.65 | 43.575 | 48.3333 | 49.4
18.4 | 37.4 | 43.3 | 48.075 | 49.25
18.5 | 37.15 | 43.075 | 47.8 | 49
18.6 | 36.9 | 42.85 | 47.55 | 48.65
18.7 | 36.75 | 42.55 | 47.4 | 48.5
18.8 | 36.55 | 42.4 | 47.1 | 48.25
18.9 | 36.3 | 42.2 | 46.9 | 48.1
19 | 36 | 42.025 | 46.7167 | 47.9
19.1 | 35.8 | 41.675 | 46.55 | 47.6667
19.2 | 35.56 | 41.5 | 46.3 | 47.4
19.3 | 35.425 | 41.275 | 46.075 | 47.225
19.4 | 35.075 | 41.0667 | 45.8 | 47.05
19.5 | 34.925 | 40.8667 | 45.575 | 46.775
19.6 | 34.675 | 40.575 | 45.32 | 46.54
19.7 | 34.525 | 40.4 | 45.2 | 46.3429
19.8 | 34.375 | 40.2 | 44.9667 | 46.125
19.9 | 34.125 | 40 | 44.75 | 46
20 | 34 | 39.7333 | 44.4 | 45.7143
20.1 | 33.775 | 39.45 | 44.25 | 45.4333
20.2 | 33.5 | 39.3 | 44.0667 | 45.3
20.3 | 33.275 | 39.15 | 43.9 | 45.0333
20.4 | 33.1 | 38.9 | 43.7 | 44.7857
20.5 | 32.9 | 38.7 | 43.44 | 44.6333
20.6 | 32.675 | 38.46 | 43.25 | 44.375
20.7 | 32.475 | 38.225 | 43.025 | 44.2
20.8 | 32.3 | 37.975 | 42.8 | 43.9667
20.9 | 32.1 | 37.8 | 42.55 | 43.8
21 | 31.85 | 37.6167 | 42.375 | 43.65
21.1 | 31.7 | 37.4 | 42.225 | 43.5
21.2 | 31.55 | 37.125 | 42.05 | 43.25
21.3 | 31.4 | 36.95 | 41.725 | 42.9
21.4 | 31.2 | 36.8 | 41.6 | 42.7571
21.5 | 31 | 36.6 | 41.3667 | 42.6
21.6 | 30.8 | 36.425 | 41.2 | 42.45
21.7 | 30.6 | 36.275 | 41 | 42.2
21.8 | 30.45 | 36.025 | 40.75 | 41.9571
21.9 | 30.3 | 35.775 | 40.6 | 41.7625
22 | 30.0667 | 35.6 | 40.4 | 41.6167
22.1 | 29.9 | 35.4 | 40.2 | 41.3625
22.2 | 29.775 | 35.225 | 40 | 41.175
22.3 | 29.52 | 35.1 | 39.8 | 40.9625
22.4 | 29.375 | 34.875 | 39.4714 | 40.8167
22.5 | 29.225 | 34.7 | 39.3 | 40.6571
22.6 | 28.975 | 34.475 | 39.25 | 40.46
22.7 | 28.825 | 34.2 | 38.9667 | 40.3125
22.8 | 28.6 | 34.1 | 38.8 | 39.9875
22.9 | 28.5 | 33.875 | 38.6 | 39.825
23 | 28.375 | 33.775 | 38.5 | 39.6375
23.1 | 28.125 | 33.45 | 38.225 | 39.425
23.2 | 28 | 33.3 | 38 | 39.3
23.3 | 27.825 | 33.125 | 37.9 | 39.18
23.4 | 27.675 | 32.95 | 37.7167 | 39
23.5 | 27.5 | 32.8 | 37.375 | 38.75
23.6 | 27.325 | 32.65 | 37.2 | 38.55
23.7 | 27.18 | 32.4 | 37.175 | 38.3
23.8 | 26.95 | 32.25 | 36.95 | 38.175
23.9 | 26.8 | 32.1 | 36.725 | 38.025
24 | 26.65 | 31.95 | 36.6 | 37.8
24.1 | 26.5 | 31.7 | 36.45 | 37.55
24.2 | 26.4 | 31.6 | 36.3 | 37.4
24.3 | 26.2 | 31.5 | 36.1 | 37.2
24.4 | 26.1 | 31.25 | 35.875 | 37.025
24.5 | 25.9 | 31 | 35.7 | 36.86
24.6 | 25.8 | 30.85 | 35.5667 | 36.725
24.7 | 25.6 | 30.7 | 35.3 | 36.6
24.8 | 25.4 | 30.55 | 35.2 | 36.425
24.9 | 25.3 | 30.3 | 35 | 36.2286
25 | 25.2 | 30.2 | 34.92 | 36.025
25.1 | 25 | 30.025 | 34.775 | 35.875
25.2 | 24.85 | 29.775 | 34.5 | 35.7
25.3 | 24.72 | 29.7 | 34.3667 | 35.475
25.4 | 24.575 | 29.575 | 34.16 | 35.3667
25.5 | 24.4 | 29.4 | 34.025 | 35.2
25.6 | 24.3 | 29.2 | 33.875 | 35.1
25.7 | 24.1 | 29.1 | 33.625 | 34.8333
25.8 | 23.925 | 28.96 | 33.475 | 34.6333
25.9 | 23.8 | 28.725 | 33.325 | 34.475
26 | 23.6 | 28.575 | 33.2 | 34.4
26.1 | 23.5 | 28.425 | 33.075 | 34.2
26.2 | 23.4 | 28.225 | 32.825 | 34.025
26.3 | 23.275 | 28.075 | 32.675 | 33.88
26.4 | 23.1 | 27.925 | 32.5 | 33.7
26.5 | 23 | 27.8 | 32.275 | 33.6
26.6 | 22.9 | 27.625 | 32.125 | 33.45
26.7 | 22.775 | 27.475 | 31.975 | 33.28
26.8 | 22.64 | 27.4 | 31.825 | 33.05
26.9 | 22.4 | 27.275 | 31.675 | 32.9
27 | 22.3 | 27.1 | 31.5 | 32.8
27.1 | 22.2 | 26.875 | 31.4 | 32.6
27.2 | 22.1 | 26.8 | 31.24 | 32.4333
27.3 | 21.9 | 26.675 | 31.08 | 32.35
27.4 | 21.8 | 26.425 | 30.95 | 32.2
27.5 | 21.68 | 26.3 | 30.74 | 32
27.6 | 21.5 | 26.2 | 30.6 | 31.8
27.7 | 21.45 | 26 | 30.45 | 31.7
27.8 | 21.3 | 25.8667 | 30.3 | 31.525
27.9 | 21.15 | 25.7 | 30.2 | 31.42
28 | 21 | 25.6 | 30.1 | 31.225
28.1 | 20.95 | 25.5 | 29.85 | 31.1
28.2 | 20.825 | 25.35 | 29.7 | 31
28.3 | 20.6667 | 25.3 | 29.5667 | 30.875
28.4 | 20.52 | 25 | 29.4 | 30.625
28.5 | 20.4 | 24.9 | 29.2 | 30.5
28.6 | 20.325 | 24.8 | 29.1 | 30.4
28.7 | 20.2 | 24.65 | 28.975 | 30.275
28.8 | 20.125 | 24.525 | 28.825 | 30.1
28.9 | 19.975 | 24.36 | 28.7 | 29.975
29 | 19.825 | 24.225 | 28.525 | 29.8
29.1 | 19.775 | 24.1 | 28.375 | 29.675
29.2 | 19.625 | 24.025 | 28.225 | 29.5
29.3 | 19.5 | 23.875 | 28.1333 | 29.3
29.4 | 19.34 | 23.725 | 27.9 | 29.2
29.5 | 19.275 | 23.6 | 27.8 | 29.08
29.6 | 19.2 | 23.5 | 27.7 | 28.95
29.7 | 19.0333 | 23.375 | 27.5 | 28.8
29.8 | 18.9 | 23.225 | 27.4 | 28.7
29.9 | 18.8 | 23.1 | 27.275 | 28.5
30 | 18.7 | 22.94 | 27.1 | 28.35

## Wedge 45 Lateral Profiles
Lateral Profiles
Wedge 45
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.2 cm FieldSize:4x4 cm2 | Depth:1.2 cm FieldSize:10x10 cm2 | Depth:1.2 cm FieldSize:20x20 cm2 | Depth:1.2 cm FieldSize:40x20 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:40x20 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:40x20 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:40x20 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:40x20 cm2
-17.9 | [18 empty] | 19.1 | 23.3889
-17.8 | [18 empty] | 19.3 | 23.65
-17.7 | [18 empty] | 19.475 | 23.92
-17.6 | [18 empty] | 19.78 | 24.2
-17.5 | [18 empty] | 20.05 | 24.5286
-17.4 | [18 empty] | 20.28 | 24.78
-17.3 | [18 empty] | 20.5 | 25.1
-17.2 | [18 empty] | 20.775 | 25.275
-17.1 | [18 empty] | 21.05 | 25.7
-17 | [18 empty] | 21.275 | 25.9857
-16.9 | [14 empty] | 15.0833 | 19.0629 |  |  | 21.6 | 26.25
-16.8 | [14 empty] | 15.425 | 19.2711 |  |  | 21.92 | 26.4778
-16.7 | [14 empty] | 15.66 | 19.4961 |  |  | 22.2571 | 26.86
-16.6 | [14 empty] | 15.8 | 19.796 |  |  | 22.56 | 27.1857
-16.5 | [14 empty] | 16.08 | 20.0293 |  |  | 22.9143 | 27.4667
-16.4 | [14 empty] | 16.35 | 20.2559 |  |  | 23.26 | 27.78
-16.3 | [14 empty] | 16.64 | 20.5559 |  |  | 23.4571 | 28.02
-16.2 | [14 empty] | 16.9 | 20.8758 |  |  | 23.88 | 28.35
-16.1 | [14 empty] | 17.08 | 21.0958 |  |  | 24.18 | 28.775
-16 | [14 empty] | 17.35 | 21.2957 |  |  | 24.44 | 29.05
-15.9 | [10 empty] | 10.6143 | 13.875 |  |  | 17.66 | 21.4707 |  |  | 24.7714 | 29.36
-15.8 | [10 empty] | 10.7 | 14.0333 |  |  | 18 | 21.8813 |  |  | 25.2 | 29.75
-15.7 | [10 empty] | 10.85 | 14.2333 |  |  | 18.22 | 22.1956 |  |  | 25.52 | 30.15
-15.6 | [10 empty] | 11 | 14.375 |  |  | 18.4667 | 22.4355 |  |  | 25.875 | 30.62
-15.5 | [10 empty] | 11.2857 | 14.6 |  |  | 18.82 | 22.7454 |  |  | 26.3 | 30.9
-15.4 | [6 empty] | 8.2833 | 11.3355 |  |  | 11.44 | 14.7286 |  |  | 19.15 | 23.0354 |  |  | 26.7333 | 31.36
-15.3 | [6 empty] | 8.4 | 11.3954 |  |  | 11.675 | 14.9571 |  |  | 19.4167 | 23.3753 |  |  | 27.1 | 31.6333
-15.2 | [6 empty] | 8.5 | 11.5954 |  |  | 11.8 | 15.16 |  |  | 19.675 | 23.6353 |  |  | 27.55 | 32.0143
-15.1 |  |  | 7.101 |  |  |  | 8.7 | 11.7096 |  |  | 12.0833 | 15.4 |  |  | 20.075 | 23.9523 |  |  | 28 | 32.4125
-15 |  |  | 7.101 | 10.0667 |  |  | 8.8 | 11.8952 |  |  | 12.3 | 15.5833 |  |  | 20.2 | 24.2702 |  |  | 28.1667 | 32.8
-14.9 |  |  | 7.201 | 10.2 |  |  | 8.9 | 12.0452 |  |  | 12.425 | 15.74 |  |  | 20.62 | 24.7575 |  |  | 28.9 | 33.2
-14.8 |  |  | 7.301 | 10.3 |  |  | 9.1 | 12.1452 |  |  | 12.625 | 15.9667 |  |  | 20.94 | 24.975 |  |  | 29.225 | 33.7
-14.7 |  |  | 7.4011 | 10.4 |  |  | 9.18 | 12.355 |  |  | 12.78 | 16.2143 |  |  | 21.32 | 25.3949 |  |  | 29.68 | 34.16
-14.6 |  |  | 7.5011 | 10.5 |  |  | 9.3 | 12.495 |  |  | 13.0857 | 16.55 |  |  | 21.75 | 25.7449 |  |  | 30.05 | 34.6
-14.5 |  |  | 7.6411 | 10.6 |  |  | 9.5 | 12.595 |  |  | 13.25 | 16.725 |  |  | 22.06 | 26.0948 |  |  | 30.62 | 35.05
-14.4 |  |  | 7.7011 | 10.7 |  |  | 9.65 | 12.7349 |  |  | 13.62 | 17 |  |  | 22.5 | 26.5322 |  |  | 31.0667 | 35.6
-14.3 |  |  | 7.8011 | 10.85 |  |  | 9.78 | 12.8449 |  |  | 13.8667 | 17.25 |  |  | 23.0125 | 26.9446 |  |  | 31.72 | 36.1167
-14.2 |  |  | 7.9411 | 10.94 |  |  | 9.92 | 13.1348 |  |  | 14.1 | 17.48 |  |  | 23.3429 | 27.2946 |  |  | 32.2 | 36.4
-14.1 |  |  | 8.0511 | 11.1 |  |  | 10.12 | 13.3447 |  |  | 14.3667 | 17.7 |  |  | 23.7125 | 27.7144 |  |  | 32.62 | 37.1167
-14 |  |  | 8.1762 | 11.2 |  |  | 10.22 | 13.4446 |  |  | 14.7 | 18.1 |  |  | 24.1 | 28.1744 |  |  | 33.35 | 37.78
-13.9 |  |  | 8.3012 | 11.34 |  |  | 10.5 | 13.6945 |  |  | 14.95 | 18.3 |  |  | 24.6857 | 28.6543 |  |  | 34.02 | 38.425
-13.8 |  |  | 8.4012 | 11.4667 |  |  | 10.7 | 13.8145 |  |  | 15.2 | 18.58 |  |  | 25.125 | 29.0942 |  |  | 34.9 | 39.2
-13.7 |  |  | 8.5212 | 11.6 |  |  | 10.85 | 14.023 |  |  | 15.45 | 18.94 |  |  | 25.4667 | 29.2941 |  |  | 35.725 | 40.1833
-13.6 |  |  | 8.6012 | 11.74 |  |  | 11.075 | 14.2276 |  |  | 15.88 | 19.24 |  |  | 26.06 | 30.044 |  |  | 37.15 | 41.42
-13.5 |  |  | 8.7512 | 11.9 |  |  | 11.25 | 14.4942 |  |  | 16.2 | 19.6 |  |  | 26.58 | 30.5439 |  |  | 39.1 | 43.36
-13.4 |  |  | 8.8813 | 12.08 |  |  | 11.475 | 14.6275 |  |  | 16.5 | 19.85 |  |  | 27.08 | 31.0938 |  |  | 42 | 46.025
-13.3 |  |  | 9.0413 | 12.1333 |  |  | 11.725 | 14.894 |  |  | 16.825 | 20.2 |  |  | 27.65 | 31.5937 |  |  | 46.8 | 50.74
-13.2 |  |  | 9.2013 | 12.3333 |  |  | 11.94 | 15.134 |  |  | 17.25 | 20.725 |  |  | 28.125 | 32.0686 |  |  | 54 | 57.1
-13.1 |  |  | 9.3347 | 12.475 |  |  | 12.075 | 15.4688 |  |  | 17.66 | 21 |  |  | 28.78 | 32.6135 |  |  | 62.9 | 65.1
-13 |  |  | 9.4847 | 12.625 |  |  | 12.38 | 15.6338 |  |  | 18.02 | 21.45 |  |  | 29.44 | 33.2933 |  |  | 73.02 | 75.2857
-12.9 |  |  | 9.6014 | 12.8333 |  |  | 12.6333 | 15.9436 |  |  | 18.4333 | 21.85 |  |  | 30.2 | 33.9332 |  |  | 84.125 | 84.75
-12.8 |  |  | 9.7414 | 12.925 |  |  | 12.88 | 16.0736 |  |  | 18.82 | 22.35 |  |  | 31.0833 | 34.7681 |  |  | 94.68 | 94.4
-12.7 |  |  | 9.9214 | 13.1556 |  |  | 13.225 | 16.4684 |  |  | 19.26 | 22.8 |  |  | 32.05 | 35.7729 |  |  | 103.8 | 102.55
-12.6 |  |  | 10.1014 | 13.3571 |  |  | 13.4667 | 16.7333 |  |  | 19.74 | 23.15 |  |  | 33.74 | 37.1925 |  |  | 111.15 | 109.35
-12.5 |  |  | 10.3015 | 13.5 |  |  | 13.78 | 17.0932 |  |  | 20.3 | 23.65 |  |  | 35.775 | 39.2122 |  |  | 116.4 | 113.84
-12.4 |  |  | 10.4765 | 13.6571 |  |  | 14.1 | 17.3731 |  |  | 20.66 | 24.14 |  |  | 39.6333 | 42.5415 |  |  | 119.56 | 116.6375
-12.3 |  |  | 10.6015 | 13.92 |  |  | 14.4333 | 17.6929 |  |  | 21.2 | 24.7 |  |  | 45.06 | 47.7904 |  |  | 121.18 | 118.1
-12.2 |  |  | 10.9016 | 14.125 |  |  | 14.82 | 18.0728 |  |  | 21.72 | 25.2 |  |  | 53.5 | 56.3287 |  |  | 121.9 | 119.3143
-12.1 |  |  | 11.0516 | 14.275 |  |  | 15.2 | 18.4127 |  |  | 22.4333 | 25.8143 |  |  | 65 | 66.3068 |  |  | 122.6 | 119.88
-12 |  |  | 11.3016 | 14.5 |  |  | 15.5333 | 18.8725 |  |  | 22.975 | 26.4 |  |  | 77.3714 | 78.0343 |  |  | 122.95 | 120.1556
-11.9 |  |  | 11.4266 | 14.72 |  |  | 16 | 19.1673 |  |  | 23.7667 | 26.9833 |  |  | 90.7 | 89.782 |  |  | 123.32 | 120.48
-11.8 |  |  | 11.7217 | 15.025 |  |  | 16.48 | 19.6921 |  |  | 24.6 | 27.9 |  |  | 102.425 | 100.48 |  |  | 123.5 | 120.8857
-11.7 |  |  | 11.9588 | 15.275 |  |  | 16.86 | 20.112 |  |  | 25.76 | 29.0333 |  |  | 113 | 111.0777 |  |  | 123.8 | 121.025
-11.6 |  |  | 12.1767 | 15.5 |  |  | 17.3714 | 20.6118 |  |  | 27.225 | 30.5 |  |  | 121.2 | 118.6763 |  |  | 124.05 | 121.1778
-11.5 |  |  | 12.4589 | 15.775 |  |  | 18.025 | 21.1249 |  |  | 29.7 | 32.8571 |  |  | 126.525 | 123.215 |  |  | 124.26 | 121.38
-11.4 |  |  | 12.7351 | 16.0333 |  |  | 18.62 | 21.8112 |  |  | 33.8 | 36.925 |  |  | 129.2333 | 126.3417 |  | 10.7 | 124.36 | 121.5286
-11.3 |  |  | 13.0519 | 16.36 |  |  | 19.475 | 22.641 |  |  | 41 | 43.8429 |  |  | 130.8 | 127.5077 |  | 10.7 | 124.5 | 121.6
-11.2 |  |  | 13.4019 | 16.72 |  |  | 20.6 | 23.6705 |  |  | 51.6 | 54.625 |  |  | 131.48 | 128.474 |  | 10.8 | 124.5 | 121.675
-11.1 |  |  | 13.702 | 17.06 |  |  | 22.18 | 25.11 |  |  | 65.9 | 67.8 |  |  | 132.02 | 129.094 |  | 10.9222 | 124.5 | 121.7
-11 |  |  | 14.1853 | 17.5 |  |  | 24.7 | 27.529 |  |  | 81 | 82.3 |  | 7.5 | 132.24 | 129.2073 |  | 11.1 | 124.5 | 121.7667
-10.9 |  |  | 14.8021 | 18.1 |  |  | 29.22 | 31.9872 |  |  | 97.26 | 98.4 |  | 7.7 | 132.4 | 129.374 |  | 11.2429 | 124.6 | 121.8
-10.8 |  |  | 15.8023 | 18.9 |  |  | 37.7143 | 39.4442 |  |  | 113.8 | 111.7 |  | 7.825 | 132.5286 | 129.499 |  | 11.4333 | 124.6 | 121.8
-10.7 |  |  | 17.2168 | 20.5 |  |  | 50.5 | 52.3291 |  |  | 125.88 | 124.5143 |  | 8 | 132.7 | 129.674 |  | 11.7 | 124.64 | 121.875
-10.6 |  |  | 19.7029 | 23.28 |  |  | 68.28 | 68.5126 |  |  | 136.1 | 133.4 |  | 8.125 | 132.7286 | 129.749 |  | 11.875 | 124.7 | 121.9
-10.5 |  |  | 25.2608 | 28.3 |  |  | 86 | 85.026 |  | 4.7981 | 141.5 | 138.3556 |  | 8.2667 | 132.8 | 129.774 |  | 12.0111 | 124.625 | 121.9
-10.4 |  |  | 35.4051 | 38.3 |  |  | 105.9 | 103.0584 |  | 4.8731 | 144.1 | 141 |  | 8.4 | 132.7778 | 129.834 |  | 12.225 | 124.5 | 121.9571
-10.3 |  |  | 50.8448 | 53.1 |  |  | 122 | 119.2522 |  | 4.9647 | 145.14 | 142.02 |  | 8.525 | 132.7 | 129.774 |  | 12.4 | 124.5 | 122
-10.2 |  |  | 70.26 | 72.2 |  | 3.5009 | 137.1 | 134.2864 |  | 4.998 | 146.05 | 142.5 |  | 8.7 | 132.6429 | 129.674 |  | 12.625 | 124.5 | 121.9
-10.1 |  |  | 90.4415 | 91.9 |  | 3.6009 | 146.75 | 143.7426 |  | 5.098 | 146.2 | 142.8571 |  | 8.95 | 132.5 | 129.674 |  | 12.8 | 124.4 | 121.9
-10 |  | 3 | 112.816 | 111.5 |  | 3.7009 | 152.375 | 149.08 |  | 5.1729 | 146.2 | 142.9 |  | 9.1 | 132.5 | 129.674 |  | 13.075 | 124.35 | 121.85
-9.9 |  | 3.1 | 131.9187 | 130.3 |  | 3.7009 | 154.8 | 151.2798 |  | 5.2979 | 146.2 | 142.8111 |  | 9.2333 | 132.4 | 129.574 |  | 13.3 | 124.3 | 121.72
-9.8 |  | 3.1 | 147.021 | 144.78 |  | 3.8009 | 155.5667 | 152.079 |  | 5.3978 | 146.1 | 142.675 |  | 9.4 | 132.2833 | 129.534 |  | 13.55 | 124.2 | 121.6
-9.7 |  | 3.1 | 157.0028 | 154.3333 |  | 3.8009 | 155.7 | 152.399 |  | 5.4978 | 145.88 | 142.6 |  | 9.5 | 132.08 | 129.454 |  | 13.7 | 124.1 | 121.5667
-9.6 |  | 3.175 | 161.773 | 158.8 |  | 3.901 | 155.7 | 152.439 |  | 5.5978 | 145.62 | 142.45 |  | 9.8 | 131.9 | 129.224 |  | 13.925 | 124 | 121.4
-9.5 |  | 3.2 | 163.2087 | 160 |  | 3.901 | 155.5 | 152.319 |  | 5.6977 | 145.42 | 142.275 |  | 10 | 131.8 | 129.074 |  | 14.14 | 123.88 | 121.3
-9.4 |  | 3.2 | 163.323 | 160 |  | 4.001 | 155.2 | 151.999 |  | 5.7977 | 145.12 | 142.06 |  | 10.2 | 131.76 | 128.974 |  | 14.42 | 123.75 | 121.28
-9.3 |  | 3.3 | 163.2516 | 159.9167 |  | 4.101 | 154.7333 | 151.799 |  | 5.8976 | 144.78 | 141.7 |  | 10.3333 | 131.5 | 128.8311 |  | 14.7571 | 123.5333 | 121.1667
-9.2 |  | 3.3 | 162.823 | 159.6 |  | 4.1344 | 154.4 | 151.2893 |  | 6.0976 | 144.5333 | 141.4 |  | 10.52 | 131.28 | 128.634 |  | 15.05 | 123.4 | 121.0333
-9.1 |  | 3.35 | 162.363 | 158.825 |  | 4.2211 | 153.95 | 150.715 |  | 6.1975 | 144.175 | 141.0857 |  | 10.8 | 131.08 | 128.424 |  | 15.2 | 123.38 | 120.9
-9 |  | 3.4 | 161.523 | 158.16 |  | 4.3011 | 153.4 | 150.22 |  | 6.2975 | 143.8375 | 140.75 |  | 11 | 130.96 | 128.174 |  | 15.4 | 123.3 | 120.8
-8.9 |  | 3.5 | 160.463 | 157.54 |  | 4.4011 | 152.55 | 149.64 |  | 6.4974 | 143.35 | 140.5333 |  | 11.2 | 130.8 | 128.0311 |  | 15.78 | 123.04 | 120.7
-8.8 |  | 3.5 | 159.748 | 156.5667 |  | 4.5011 | 151.9143 | 149.22 |  | 6.6473 | 143.02 | 140.28 |  | 11.5 | 130.45 | 127.934 |  | 16.12 | 122.8 | 120.7
-8.7 |  | 3.6 | 159.023 | 155.8667 |  | 4.6011 | 151.375 | 148.541 |  | 6.7973 | 142.52 | 139.8571 |  | 11.8 | 130.24 | 127.7883 |  | 16.3714 | 122.575 | 120.6
-8.6 |  | 3.6 | 158.023 | 155.1 |  | 4.6011 | 150.82 | 148.041 |  | 6.8972 | 142.175 | 139.28 |  | 12 | 129.975 | 127.5344 |  | 16.84 | 122.46 | 120.46
-8.5 |  | 3.7 | 157.142 | 154.125 |  | 4.7012 | 150.25 | 147.461 |  | 7.0972 | 141.92 | 138.95 |  | 12.2 | 129.7 | 127.335 |  | 17.1 | 122.4 | 120.26
-8.4 |  | 3.7 | 156.322 | 153.3 |  | 4.9012 | 149.6 | 146.9267 |  | 7.2638 | 141.24 | 138.5 |  | 12.5 | 129.25 | 127.015 |  | 17.5 | 122.1 | 120.1
-8.3 |  | 3.8 | 155.422 | 152.6 |  | 5.0012 | 148.85 | 146.341 |  | 7.4637 | 140.84 | 138.1 |  | 12.8 | 128.95 | 126.675 |  | 17.76 | 122.1 | 120
-8.2 |  | 3.9 | 154.647 | 151.7667 |  | 5.1013 | 148.12 | 145.6753 |  | 7.6469 | 140.34 | 137.6333 |  | 13.1 | 128.62 | 126.4464 |  | 18.14 | 121.76 | 119.8
-8.1 |  | 4 | 153.772 | 150.8 |  | 5.2013 | 147.64 | 144.902 |  | 7.7969 | 139.8 | 137.18 |  | 13.475 | 128.3 | 126.2321 |  | 18.54 | 121.525 | 119.5667
-8 |  | 4.05 | 152.897 | 150.28 |  | 5.3013 | 147 | 144.242 |  | 8.0301 | 139.16 | 136.6667 |  | 13.8 | 127.8714 | 125.815 |  | 18.94 | 121.32 | 119.25
-7.9 |  | 4.1 | 151.972 | 149.3 |  | 5.5014 | 146.2 | 143.7427 |  | 8.2967 | 138.6667 | 136.2 |  | 14.15 | 127.7 | 125.475 |  | 19.28 | 121.2 | 119.12
-7.8 |  | 4.2 | 151.102 | 148.2 |  | 5.6014 | 145.4857 | 143.143 |  | 8.5366 | 138.3 | 135.72 |  | 14.5 | 127.375 | 125.2583 |  | 19.7667 | 120.92 | 118.94
-7.7 |  | 4.3 | 150.1547 | 147.5714 |  | 5.7014 | 144.94 | 142.303 |  | 8.7632 | 137.8 | 135.25 |  | 14.775 | 127.2333 | 125.05 |  | 20.08 | 120.7 | 118.725
-7.6 |  | 4.4 | 149.241 | 146.7 |  | 5.9015 | 144.0429 | 141.823 |  | 9.0964 | 137.2 | 134.84 |  | 15.2 | 126.8714 | 124.755 | 3.7987 | 20.6 | 120.46 | 118.525
-7.5 |  | 4.5 | 148.421 | 145.94 |  | 6.1682 | 143.56 | 141.044 |  | 9.3962 | 136.86 | 134.4 |  | 15.4667 | 126.6 | 124.4972 | 4.0986 | 21.075 | 120.0833 | 118.3833
-7.4 |  | 4.6 | 147.501 | 145.1 |  | 6.3016 | 142.725 | 140.5107 |  | 9.5962 | 136.26 | 133.82 | 3 | 16 | 126.38 | 124.275 | 4.1986 | 21.65 | 119.9 | 118.275
-7.3 |  | 4.7 | 146.7543 | 144.14 |  | 6.4683 | 141.8 | 139.844 |  | 9.896 | 135.62 | 133.3 | 2.9 | 16.25 | 126 | 123.9893 | 4.2986 | 22.4 | 119.64 | 118.04
-7.2 |  | 4.825 | 145.971 | 143.5286 |  | 6.7017 | 141.1714 | 139.119 | 1.7 | 10.096 | 135.0667 | 132.5 | 3 | 16.75 | 125.7 | 123.735 | 4.3985 | 23.3 | 119.44 | 117.74
-7.1 |  | 4.9 | 144.9654 | 142.86 | 1.1996 | 6.8617 | 140.52 | 138.145 | 1.7 | 10.4958 | 134.52 | 132.1333 | 3 | 17.2333 | 125.3333 | 123.3464 | 4.3985 | 24.34 | 119.2 | 117.56
-7 | 0.96699 | 5 | 144.021 | 141.7 | 1.1996 | 7.1351 | 139.7 | 137.695 | 1.7 | 10.9289 | 134.05 | 131.825 | 3.1 | 17.7 | 125 | 123.025 | 4.4985 | 26.1875 | 118.9 | 117.3
-6.9 | 1.0003 | 5.2 | 143.3493 | 140.95 | 1.1996 | 7.3618 | 139.02 | 137.045 | 1.8 | 11.1955 | 133.62 | 131.4 | 3.1 | 18.2667 | 124.6143 | 122.775 | 4.4985 | 28.8 | 118.56 | 117.16
-6.8 | 1.0003 | 5.4 | 142.445 | 140.3714 | 1.1996 | 7.6352 | 138.2 | 136.4117 | 1.8 | 11.5954 | 132.9667 | 130.55 | 3.2 | 19.075 | 124.2 | 122.4755 | 4.6984 | 33.4286 | 118.25 | 116.95
-6.7 | 1.0003 | 5.5 | 141.82 | 139.625 | 1.2329 | 7.962 | 137.6 | 135.746 | 1.8 | 11.8952 | 132.36 | 130.14 | 3.3 | 19.7 | 123.92 | 122.101 | 4.7984 | 39.8 | 118 | 116.775
-6.6 | 1.0003 | 5.7 | 140.62 | 138.6 | 1.2996 | 8.202 | 137 | 134.986 | 1.9 | 12.495 | 131.94 | 129.6 | 3.35 | 20.975 | 123.6 | 121.7093 | 4.7984 | 48.55 | 117.8 | 116.55
-6.5 | 1.0003 | 5.9 | 139.82 | 137.8333 | 1.2996 | 8.5521 | 136 | 134.346 | 1.9 | 12.8449 | 131.35 | 128.9 | 3.4 | 22.9 | 123.15 | 121.326 | 4.8734 | 58.35 | 117.42 | 116.35
-6.4 | 1.0003 | 6.0667 | 139.08 | 137 | 1.2996 | 8.8689 | 135.3714 | 133.587 | 1.9 | 13.428 | 130.78 | 128.5 | 3.5 | 26 | 122.6 | 120.9093 | 4.9983 | 68.44 | 117.22 | 115.96
-6.3 | 1.0003 | 6.2 | 138.46 | 136.14 | 1.2996 | 9.2356 | 134.7714 | 132.972 | 2 | 14.0944 | 130.14 | 127.94 | 3.5667 | 30.8 | 122.08 | 120.651 | 5.0583 | 78.6 | 116.94 | 115.72
-6.2 | 1.1004 | 6.4 | 137.5057 | 135.6 | 1.2996 | 9.6024 | 134 | 132.397 | 2.05 | 14.8941 | 129.46 | 127.5 | 3.6 | 38 | 121.65 | 120.276 | 5.2316 | 87.9 | 116.8 | 115.52
-6.1 | 1.1004 | 6.6 | 136.5334 | 134.8 | 1.3995 | 10.2026 | 133.42 | 131.947 | 2.1 | 16.2435 | 129.15 | 127.05 | 3.7 | 48 | 121.28 | 120.136 | 5.2982 | 95.0333 | 116.525 | 115.38
-6 | 1.1004 | 6.8667 | 135.719 | 134 | 1.3995 | 10.8027 | 132.7429 | 131.0623 | 2.1 | 17.9428 | 128.375 | 126.5 | 3.8 | 59.25 | 121.0667 | 119.756 | 5.3982 | 100.4 | 116.4 | 115.1667
-5.9 | 1.1004 | 7.2667 | 134.879 | 133.26 | 1.3995 | 11.7029 | 132.16 | 130.228 | 2.2 | 21.2515 | 127.94 | 126.22 | 3.9 | 70.75 | 120.82 | 119.536 | 5.4582 | 103.54 | 116.08 | 114.88
-5.8 | 1.1004 | 7.8 | 134.2333 | 132.5 | 1.4995 | 13.3033 | 131.5143 | 129.768 | 2.2 | 26.6394 | 127.48 | 125.575 | 4 | 81.675 | 120.44 | 119.126 | 5.5981 | 105.1333 | 115.74 | 114.68
-5.7 | 1.1004 | 8.5 | 133.5904 | 131.84 | 1.4995 | 15.7039 | 130.86 | 128.8487 | 2.3 | 35.2859 | 126.8 | 125.02 | 4 | 91.45 | 120.08 | 118.716 | 5.7581 | 106.42 | 115.55 | 114.44
-5.6 | 1.1004 | 9.8 | 132.819 | 131.16 | 1.4995 | 20.2051 | 129.94 | 128.199 | 2.3 | 46.9312 | 126.3 | 124.5667 | 4.1 | 98.4 | 119.7 | 118.276 | 5.898 | 106.98 | 115.34 | 114.14
-5.5 | 1.1004 | 12.3 | 131.919 | 130.5 | 1.4995 | 27.907 | 129.425 | 127.749 | 2.4 | 59.3096 | 125.65 | 123.9667 | 4.3 | 103.5 | 119.24 | 118.056 | 5.998 | 107.28 | 115.2 | 114
-5.4 | 1.2004 | 17.2 | 131.2761 | 129.7 | 1.5995 | 39.0098 | 128.8 | 126.989 | 2.5 | 72.3711 | 125.08 | 123.54 | 4.4 | 106.75 | 118.975 | 117.826 | 6.098 | 107.64 | 114.9 | 113.9
-5.3 | 1.2004 | 26 | 130.319 | 128.975 | 1.5995 | 52.8132 | 127.8286 | 126.349 | 2.6 | 85.0327 | 124.64 | 122.85 | 4.5 | 108.1 | 118.64 | 117.457 | 6.2979 | 107.825 | 114.6 | 113.68
-5.2 | 1.2004 | 38.9 | 129.619 | 128.1 | 1.6994 | 66.8167 | 127.06 | 125.85 | 2.65 | 96.0116 | 123.72 | 122.44 | 4.6 | 109 | 118.3 | 116.877 | 6.3979 | 107.9667 | 114.36 | 113.45
-5.1 | 1.2004 | 53.3 | 128.938 | 127.325 | 1.6994 | 81.3203 | 126.6 | 125.25 | 2.7 | 103.992 | 123.425 | 122.02 | 4.7333 | 109.7 | 118 | 116.657 | 6.4978 | 108 | 114.06 | 113.1
-5 | 1.3004 | 69.7 | 127.818 | 126.65 | 1.7994 | 94.3236 | 126 | 124.47 | 2.8 | 108.7065 | 122.925 | 121.52 | 4.9 | 109.8 | 117.45 | 116.4437 | 6.5978 | 108.02 | 113.72 | 112.7
-4.9 | 1.3004 | 85.2 | 127.298 | 126 | 1.8494 | 104.226 | 125.3 | 123.95 | 2.9 | 111.655 | 122.52 | 121.05 | 5 | 110 | 117.12 | 116.077 | 6.7977 | 108.12 | 113.5 | 112.6
-4.8 | 1.3004 | 99.2 | 126.668 | 125.4 | 1.8994 | 111.128 | 124.65 | 123.1843 | 3 | 112.755 | 122 | 120.66 | 5.1 | 110.2 | 116.8 | 115.702 | 6.8977 | 108.2 | 113.36 | 112.5
-4.7 | 1.3504 | 110.3 | 125.958 | 124.68 | 1.9993 | 114.829 | 124.1 | 122.691 | 3.1 | 113.455 | 121.5 | 120.1667 | 5.3 | 110.2 | 116.6 | 115.477 | 7.131 | 108.3 | 113.05 | 112.3
-4.6 | 1.4005 | 116.6 | 125.318 | 124.1 | 2.0993 | 116.3623 | 123.4571 | 122.251 | 3.2 | 113.755 | 120.9333 | 119.4 | 5.5 | 110.3 | 116.12 | 115.237 | 7.2976 | 108.3 | 112.66 | 111.98
-4.5 | 1.4005 | 119.7 | 124.758 | 123.38 | 2.0993 | 117.029 | 122.7 | 121.701 | 3.3 | 113.854 | 120.36 | 119.0333 | 5.6667 | 110.26 | 115.8 | 114.837 | 7.6308 | 108.3 | 112.46 | 111.55
-4.4 | 1.5005 | 120.7 | 123.918 | 122.6 | 2.1993 | 117.229 | 122.3286 | 121.1659 | 3.45 | 113.954 | 119.96 | 118.56 | 5.8 | 110.2 | 115.4 | 114.427 | 7.8374 | 108.3 | 112.18 | 111.275
-4.3 | 1.5005 | 120.8 | 123.318 | 122.02 | 2.2992 | 117.329 | 121.5667 | 120.552 | 3.625 | 113.854 | 119.55 | 118.225 | 6 | 110.2 | 115.1 | 114.1103 | 7.9973 | 108.3 | 112 | 111.1
-4.2 | 1.6005 | 120.8 | 122.317 | 121.45 | 2.3992 | 117.329 | 121 | 120.092 | 3.7667 | 113.755 | 118.8 | 117.78 | 6.2 | 110.1 | 114.825 | 113.752 | 8.0973 | 108.1667 | 111.84 | 111
-4.1 | 1.6005 | 120.55 | 121.817 | 120.94 | 2.4492 | 117.179 | 120.4667 | 119.532 | 3.94 | 113.655 | 118.32 | 117 | 6.4333 | 110 | 114.5 | 113.597 | 8.2972 | 108.02 | 111.38 | 110.82
-4 | 1.7006 | 120.25 | 121.477 | 120.3 | 2.5741 | 117.004 | 120.0333 | 118.952 | 4.1 | 113.555 | 117.78 | 116.6 | 6.6 | 109.9 | 114.02 | 113.227 | 8.5971 | 107.9 | 111.1333 | 110.3333
-3.9 | 1.7006 | 119.95 | 120.817 | 119.66 | 2.7491 | 116.829 | 119.2714 | 118.473 | 4.225 | 113.38 | 117.55 | 116.18 | 6.8333 | 109.7 | 113.6 | 112.797 | 8.897 | 107.7667 | 110.78 | 110
-3.8 | 1.8006 | 119.4 | 120.077 | 119.1 | 2.899 | 116.529 | 118.6 | 117.873 | 4.5 | 113.155 | 116.9 | 115.7 | 7.075 | 109.5 | 113.275 | 112.378 | 9.147 | 107.7 | 110.34 | 109.9
-3.7 | 1.9006 | 118.8 | 119.317 | 118.55 | 2.979 | 116.029 | 118.2 | 117.213 | 4.7 | 112.98 | 116.3 | 115.08 | 7.25 | 109.4 | 112.82 | 112.138 | 9.4219 | 107.6 | 110.2 | 109.7
-3.6 | 2.0007 | 118.25 | 118.817 | 118.0667 | 3.1989 | 115.729 | 117.55 | 116.7959 | 4.8 | 112.7217 | 115.88 | 114.75 | 7.5667 | 109.3 | 112.3667 | 111.828 | 9.8301 | 107.5 | 110.06 | 109.5
-3.5 | 2.0007 | 117.8667 | 118.357 | 117.5 | 3.2989 | 115.429 | 117.12 | 116.254 | 5.2 | 112.48 | 115.3 | 114.4 | 7.86 | 109.3 | 112.12 | 111.598 | 10.2216 | 107.5 | 109.8 | 109.15
-3.4 | 2.2007 | 117.5667 | 117.6456 | 116.8 | 3.4738 | 115.079 | 116.46 | 115.654 | 5.3 | 112.1217 | 114.9 | 114.05 | 8.2 | 109.1 | 111.8 | 111.278 | 10.7631 | 107.4 | 109.52 | 108.8
-3.3 | 2.2007 | 117.0667 | 116.957 | 116.1333 | 3.6988 | 114.629 | 115.95 | 115.204 | 5.6 | 111.855 | 114.4 | 113.4714 | 8.64 | 108.9 | 111.475 | 110.838 | 11.2963 | 107.3 | 109.15 | 108.5333
-3.2 | 2.3008 | 116.34 | 116.4545 | 115.5333 | 3.7987 | 114.229 | 115.35 | 114.694 | 5.9333 | 111.615 | 113.95 | 113 | 9.0667 | 108.8 | 111.1 | 110.528 | 12.3292 | 107.1 | 108.82 | 108.3167
-3.1 | 2.4008 | 115.8 | 115.917 | 114.8333 | 4.1653 | 113.928 | 114.7143 | 114.134 | 6.25 | 111.355 | 113.4 | 112.5571 | 9.8 | 108.7 | 110.66 | 110.228 | 14.1203 | 106.9 | 108.4 | 108.175
-3 | 2.6009 | 115.35 | 115.236 | 114.4 | 4.4652 | 113.328 | 114.075 | 113.595 | 6.7333 | 110.976 | 113.05 | 112 | 10.9333 | 108.5 | 110.34 | 109.8447 | 16.6278 | 106.76 | 108.2 | 107.96
-2.9 | 2.8009 | 114.8333 | 114.616 | 113.86 | 4.8984 | 112.928 | 113.64 | 113.015 | 7.4 | 110.5893 | 112.5 | 111.5429 | 12.625 | 108.3 | 109.98 | 109.458 | 20.7931 | 106.65 | 107.9 | 107.725
-2.8 | 3.051 | 114.35 | 113.916 | 113.35 | 5.5481 | 112.528 | 113.3 | 112.555 | 8.3 | 110.306 | 111.92 | 111.2 | 15.325 | 107.8 | 109.6 | 109.078 | 27.3242 | 106.4 | 107.7667 | 107.42
-2.7 | 3.5345 | 113.775 | 113.356 | 112.85 | 6.5645 | 112.028 | 112.6714 | 112.1121 | 9.8 | 110.056 | 111.66 | 110.86 | 19.7 | 107.6333 | 109.2 | 108.878 | 35.8881 | 106.3 | 107.35 | 106.96
-2.6 | 4.4015 | 113.3667 | 112.836 | 112.28 | 7.9973 | 111.6613 | 112.2 | 111.615 | 12.4 | 109.706 | 111.2667 | 110.3571 | 26.55 | 107.4 | 108.8333 | 108.528 | 45.6848 | 106.2 | 106.9333 | 106.66
-2.5 | 5.852 | 112.46 | 112.2827 | 111.82 | 10.7631 | 111.328 | 111.75 | 111.136 | 16.9 | 109.256 | 110.6 | 109.94 | 35.8 | 107.3 | 108.46 | 107.998 | 56.4812 | 105.9 | 106.65 | 106.36
-2.4 | 8.7529 | 112.1 | 111.666 | 111.5 | 15.9947 | 110.8947 | 111.2 | 110.781 | 24 | 109.006 | 110.12 | 109.4143 | 47.1 | 107.2 | 108.05 | 107.638 | 68.2272 | 105.56 | 106.42 | 106.075
-2.3 | 14.6549 | 111.7 | 111.266 | 110.66 | 24.2919 | 110.278 | 110.54 | 110.0131 | 35 | 108.657 | 109.74 | 108.95 | 60 | 106.9 | 107.7167 | 107.419 | 78.2739 | 105.36 | 106.14 | 105.9
-2.2 | 23.958 | 110.9 | 110.6327 | 110.26 | 35.7881 | 109.9022 | 110.02 | 109.556 | 47.2333 | 108.257 | 109.36 | 108.6667 | 71.2 | 106.7 | 107.3333 | 107.179 | 87.4708 | 105.3 | 105.88 | 105.68
-2.1 | 36.162 | 110.4 | 110.016 | 109.75 | 47.7174 | 109.527 | 109.525 | 109.116 | 59.8 | 108.007 | 108.88 | 108.425 | 81.75 | 106.25 | 107.02 | 106.779 | 93.7687 | 105.12 | 105.7 | 105.5667
-2 | 50.6669 | 109.9 | 109.6589 | 109.16 | 61.4462 | 108.8937 | 109.02 | 108.7565 | 71.5 | 107.707 | 108.45 | 107.7 | 90.6 | 105.9 | 106.75 | 106.479 | 97.7674 | 105 | 105.42 | 105.34
-1.9 | 64.5882 | 109.5 | 109.116 | 108.7 | 74.6351 | 108.427 | 108.48 | 108.057 | 83.5 | 107.357 | 108.0667 | 107.2857 | 96.8 | 105.7333 | 106.42 | 106.079 | 100.067 | 104.7333 | 105.125 | 105.2
-1.8 | 78.3261 | 108.95 | 108.382 | 108.3 | 85.5715 | 107.9937 | 108.1 | 107.707 | 91.9 | 106.882 | 107.55 | 106.9 | 99.8333 | 105.5333 | 106.1 | 105.829 | 101.666 | 104.55 | 104.9 | 104.8
-1.7 | 89.3631 | 108.275 | 107.835 | 107.46 | 94.6084 | 107.6937 | 107.6 | 107.157 | 97.5 | 106.437 | 107.08 | 106.675 | 101.35 | 105.3 | 105.54 | 105.379 | 102.166 | 104.16 | 104.58 | 104.4667
-1.6 | 97.5325 | 107.8 | 107.435 | 107.1667 | 99.6667 | 107.227 | 107.1 | 106.817 | 100.7 | 106.158 | 106.68 | 106.2 | 102.375 | 104.8333 | 105.2 | 104.999 | 102.366 | 103.92 | 104.35 | 104.2
-1.5 | 102.334 | 107.25 | 106.995 | 106.8 | 102.166 | 106.8937 | 106.6 | 106.3822 | 101.9 | 105.958 | 106.28 | 105.92 | 102.8 | 104.5 | 104.9 | 104.779 | 102.566 | 103.72 | 104.04 | 103.94
-1.4 | 104.1017 | 106.8 | 106.495 | 106.25 | 103.366 | 106.247 | 106.18 | 105.7247 | 102.5 | 105.458 | 105.775 | 105.4 | 102.9 | 104.2667 | 104.65 | 104.559 | 102.766 | 103.58 | 103.86 | 103.5
-1.3 | 104.635 | 106.4 | 106.065 | 105.8 | 103.665 | 105.926 | 105.7 | 105.3866 | 102.6 | 105.058 | 105.38 | 105.0571 | 102.9 | 104.1 | 104.3571 | 103.979 | 102.766 | 103.4 | 103.45 | 103.2
-1.2 | 104.635 | 105.7 | 105.415 | 105.3 | 103.765 | 105.451 | 105.16 | 104.9247 | 102.7 | 104.558 | 105 | 104.6 | 102.9 | 103.7333 | 104 | 103.819 | 102.666 | 103.18 | 103.2 | 103.06
-1.1 | 104.635 | 105.4 | 105.015 | 104.8 | 103.665 | 104.776 | 104.725 | 104.5437 | 102.7 | 104.158 | 104.6 | 104.3 | 102.7667 | 103.4 | 103.64 | 103.419 | 102.366 | 102.84 | 102.8 | 102.84
-1 | 104.4017 | 104.9 | 104.415 | 104.5 | 103.565 | 104.326 | 104.3 | 104.183 | 102.6 | 103.759 | 104.3 | 103.82 | 102.58 | 103.2 | 103.3 | 103.029 | 102.191 | 102.5 | 102.62 | 102.5857
-0.9 | 103.885 | 104.3 | 104.165 | 103.9 | 103.066 | 103.976 | 103.7333 | 103.6757 | 102.5 | 103.359 | 103.76 | 103.4571 | 102.4667 | 102.85 | 103.1 | 102.829 | 101.966 | 102.2667 | 102.32 | 102.25
-0.8 | 103.5013 | 103.8 | 103.465 | 103.5 | 102.866 | 103.426 | 103.35 | 103.0368 | 102.2 | 102.859 | 103.15 | 103 | 102.275 | 102.5667 | 102.62 | 102.3467 | 101.866 | 102 | 102.06 | 101.9857
-0.7 | 103.159 | 103.3 | 102.9983 | 103 | 102.666 | 103.006 | 102.94 | 102.779 | 102.025 | 102.5923 | 102.76 | 102.6571 | 102.0667 | 102.2 | 102.3 | 102.14 | 101.716 | 101.8 | 101.86 | 101.75
-0.6 | 102.9007 | 102.9 | 102.755 | 102.58 | 102.366 | 102.626 | 102.54 | 102.399 | 101.7333 | 102.299 | 102.4 | 102.2 | 101.775 | 101.8667 | 101.95 | 101.7943 | 101.591 | 101.54 | 101.46 | 101.5571
-0.5 | 102.484 | 102.2333 | 102.1293 | 102 | 102.191 | 102.076 | 102.06 | 101.959 | 101.4 | 101.859 | 102.1 | 101.66 | 101.5 | 101.4 | 101.5667 | 101.5133 | 101.366 | 101.2 | 101 | 101.35
-0.4 | 101.834 | 101.9 | 101.7721 | 101.65 | 101.766 | 101.6 | 101.64 | 101.719 | 101.2 | 101.484 | 101.525 | 101.4 | 101.3 | 101.2667 | 101.14 | 101.18 | 101.191 | 100.9 | 100.7 | 101.05
-0.3 | 101.3007 | 101.25 | 101.3542 | 101.125 | 101.366 | 101.275 | 101.24 | 101.359 | 100.9333 | 100.91 | 101.24 | 101.14 | 101 | 101 | 100.7 | 100.98 | 100.966 | 100.7 | 100.64 | 100.86
-0.2 | 100.9007 | 101 | 100.864 | 100.7333 | 101.066 | 100.825 | 100.86 | 100.7933 | 100.5333 | 100.58 | 100.86 | 100.7 | 100.6 | 100.7 | 100.425 | 100.58 | 100.7327 | 100.5 | 100.5 | 100.55
-0.1 | 100.4835 | 100.5 | 100.414 | 100.44 | 100.467 | 100.2917 | 100.54 | 100.46 | 100.2 | 100.36 | 100.25 | 100.15 | 100.3 | 100.4 | 100.2556 | 100.28 | 100.367 | 100.35 | 100.2 | 100.3
0 | 99.9998 | 100 | 99.9998 | 100 | 100.0001 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0001 | 100 | 100 | 100
0.1 | 99.6832 | 99.6 | 99.6857 | 99.56 | 99.6168 | 99.6916 | 99.52 | 99.5602 | 99.5 | 99.6268 | 99.58 | 99.66 | 99.7 | 99.8 | 99.7143 | 99.68 | 99.7667 | 99.75 | 99.48 | 99.65
0.2 | 99.1664 | 99.0667 | 99.1392 | 99.0667 | 99.2669 | 99.3649 | 99 | 99.2803 | 99.05 | 99.3603 | 99.3 | 99.15 | 99.3 | 99.4 | 99.4 | 99.4051 | 99.5001 | 99.44 | 99.1667 | 99.54
0.3 | 98.8329 | 98.65 | 98.7891 | 98.78 | 98.7671 | 98.8747 | 98.7 | 98.8747 | 98.6 | 98.8605 | 99 | 98.8 | 98.9 | 99.1 | 98.9429 | 99.0802 | 99.1669 | 99.2667 | 98.94 | 99.26
0.4 | 98.4328 | 98.35 | 98.4516 | 98.4 | 98.3005 | 98.5246 | 98.3 | 98.5106 | 98.1 | 98.4606 | 98.76 | 98.65 | 98.5 | 98.8 | 98.6667 | 98.7469 | 98.8004 | 99.025 | 98.7667 | 98.9
0.5 | 97.7993 | 97.9667 | 98.014 | 98.08 | 97.9673 | 97.8745 | 98 | 98.1321 | 97.55 | 98.0608 | 98.4 | 98.2 | 98.3 | 98.5 | 98.36 | 98.3803 | 98.5671 | 98.74 | 98.54 | 98.6
0.6 | 97.4075 | 97.54 | 97.6639 | 97.65 | 97.4175 | 97.6244 | 97.6 | 97.9108 | 97.2 | 97.8109 | 97.9 | 97.8429 | 97.9 | 98.3 | 98.0714 | 98.2553 | 98.3672 | 98.275 | 98.2 | 98.4
0.7 | 97.1324 | 97.0333 | 97.0939 | 97.3 | 96.8677 | 97.1643 | 97.25 | 97.4468 | 96.9 | 97.461 | 97.5 | 97.4 | 97.6 | 97.95 | 97.6429 | 97.9304 | 97.9173 | 98.025 | 97.9 | 98.12
0.8 | 96.4321 | 96.6 | 96.6888 | 96.82 | 96.4345 | 96.6909 | 96.84 | 97.1112 | 96.4 | 97.0945 | 97.16 | 97.0857 | 97.3 | 97.5333 | 97.3 | 97.6138 | 97.5175 | 97.725 | 97.52 | 97.92
0.9 | 95.932 | 96.3 | 96.4138 | 96.5 | 95.918 | 96.4241 | 96.46 | 96.7613 | 95.9 | 96.6613 | 96.8333 | 96.8857 | 96.65 | 97.15 | 97.05 | 97.4005 | 97.2176 | 97.44 | 97.25 | 97.78
1 | 95.6319 | 95.925 | 95.9137 | 96.2 | 95.4182 | 96.2241 | 96.22 | 96.3864 | 95.4333 | 96.3215 | 96.44 | 96.58 | 96.325 | 96.9 | 96.75 | 96.9606 | 96.8344 | 97.2333 | 96.96 | 97.35
1.1 | 94.9316 | 95.4333 | 95.647 | 95.9 | 94.8017 | 95.924 | 95.9333 | 95.9616 | 94.7 | 95.9616 | 96.2 | 96.2 | 95.8 | 96.55 | 96.48 | 96.6807 | 96.4179 | 96.9333 | 96.6667 | 97.12
1.2 | 94.4315 | 95.125 | 95.4136 | 95.6 | 94.0686 | 95.3238 | 95.65 | 95.6331 | 94.2 | 95.5368 | 95.86 | 95.9 | 95.325 | 96.1 | 96.2 | 96.3607 | 96.1346 | 96.6667 | 96.46 | 96.92
1.3 | 93.5312 | 94.775 | 94.885 | 95.1 | 93.2689 | 95.0238 | 95.25 | 95.2904 | 93.4 | 95.2619 | 95.4 | 95.55 | 94.6 | 95.9 | 95.75 | 96.1558 | 95.6681 | 96.3 | 96.175 | 96.68
1.4 | 92.4308 | 94.25 | 94.5885 | 94.74 | 92.1443 | 94.6237 | 94.84 | 95.012 | 92.5 | 94.962 | 95.06 | 95.12 | 94.15 | 95.6 | 95.4333 | 95.8608 | 94.9683 | 95.9 | 96.025 | 96.475
1.5 | 89.7299 | 93.975 | 94.3579 | 94.34 | 90.3699 | 93.9235 | 94.54 | 94.5177 | 91.6 | 94.7621 | 94.7667 | 94.88 | 93.2 | 95.3 | 95.18 | 95.5809 | 94.4352 | 95.5 | 95.68 | 96.34
1.6 | 85.9286 | 93.64 | 93.8384 | 94.1 | 87.7707 | 93.6234 | 94.14 | 94.2623 | 89.7 | 93.9624 | 94.56 | 94.6 | 92.15 | 95 | 94.9 | 95.081 | 94.0686 | 95.2833 | 95.45 | 96.14
1.7 | 79.5265 | 93.2 | 93.3733 | 93.7 | 82.6225 | 93.2233 | 93.68 | 93.8958 | 87.16 | 93.6625 | 94.175 | 94.2429 | 91.4333 | 94.4333 | 94.64 | 94.856 | 93.7687 | 95.0667 | 95.16 | 95.9
1.8 | 69.2564 | 92.8333 | 93.0847 | 93.3333 | 75.0749 | 93.0233 | 93.38 | 93.6625 | 82.9 | 93.2127 | 93.78 | 93.96 | 89.3 | 94.1 | 94.3875 | 94.701 | 92.5691 | 94.7 | 94.86 | 95.48
1.9 | 58.5195 | 92.48 | 92.7132 | 92.9 | 65.3282 | 92.6232 | 93.06 | 93.3769 | 76 | 92.7629 | 93.425 | 93.575 | 86.5333 | 93.9 | 94.075 | 94.2811 | 91.3695 | 94.275 | 94.54 | 95
2 | 45.9653 | 92.05 | 92.2132 | 92.54 | 53.7571 | 92.2231 | 92.725 | 92.9629 | 66.4 | 92.423 | 93.1 | 93.22 | 81.6333 | 93.4 | 93.7 | 94.1012 | 89.2702 | 94 | 94.38 | 94.75
2.1 | 33.7862 | 91.5 | 91.8381 | 92.08 | 42.3859 | 91.823 | 92.4667 | 92.588 | 56.0333 | 92.1631 | 92.78 | 92.8571 | 75.0333 | 93.1 | 93.42 | 93.8312 | 86.6378 | 93.7 | 94 | 94.5
2.2 | 23.2411 | 91.1 | 91.5731 | 91.74 | 31.7894 | 91.3228 | 91.975 | 92.3202 | 44.9 | 91.6133 | 92.3 | 92.55 | 66.0333 | 92.7 | 93.0714 | 93.3813 | 82.1726 | 93.6 | 93.625 | 94.16
2.3 | 14.2297 | 90.75 | 91.293 | 91.48 | 21.9927 | 91.0478 | 91.6 | 91.9133 | 34.6 | 91.2635 | 91.975 | 92.25 | 56.2 | 92.3 | 92.8 | 93.1564 | 75.9413 | 93.2 | 93.4 | 93.92
2.4 | 8.4528 | 90.6 | 90.7629 | 91.08 | 14.4952 | 90.5726 | 91.34 | 91.4063 | 24.95 | 90.8137 | 91.75 | 91.95 | 45.75 | 91.9 | 92.5 | 92.9014 | 68.1773 | 92.72 | 93.2167 | 93.8167
2.5 | 5.4351 | 89.9 | 90.4129 | 90.84 | 9.6968 | 90.2476 | 90.9 | 91.1235 | 17.3333 | 90.3439 | 91.52 | 91.6 | 35.94 | 91.425 | 92.24 | 92.5148 | 58.6471 | 92.45 | 92.975 | 93.56
2.6 | 4.1014 | 89.55 | 90.1129 | 90.45 | 7.1976 | 89.7224 | 90.6 | 90.7237 | 12.3 | 89.964 | 91.1 | 91.18 | 26.65 | 91.2 | 92 | 92.2565 | 48.917 | 92.04 | 92.66 | 93.3
2.7 | 3.2678 | 89.1 | 89.7328 | 90 | 5.7981 | 89.3557 | 90.0667 | 90.4638 | 9.5 | 89.6641 | 90.62 | 90.8 | 19.55 | 90.7667 | 91.6 | 91.9149 | 38.5872 | 91.72 | 92.4333 | 93.08
2.8 | 2.851 | 88.8 | 89.3628 | 89.75 | 4.9983 | 88.8722 | 89.6 | 90.0212 | 7.9667 | 89.2643 | 90.32 | 90.475 | 14.9 | 90.2 | 91.25 | 91.6617 | 29.5902 | 91.45 | 92.16 | 92.7333
2.9 | 2.5008 | 88.4 | 89.1127 | 89.46 | 4.4485 | 88.4554 | 89.32 | 89.5642 | 6.9667 | 88.5646 | 90.025 | 90.0333 | 11.975 | 89.8 | 90.9 | 91.4484 | 22.7424 | 90.975 | 91.9333 | 92.42
3 | 2.3008 | 87.8 | 88.5698 | 89 | 4.0986 | 87.8887 | 89 | 89.2421 | 6.3667 | 88.1647 | 89.8 | 89.8571 | 10.6 | 89.4667 | 90.72 | 91.0418 | 17.5941 | 90.725 | 91.56 | 92.14
3.1 | 2.2007 | 87.4 | 88.2326 | 88.48 | 3.7987 | 87.4552 | 88.64 | 88.9844 | 5.95 | 87.6649 | 89.45 | 89.5 | 9.4667 | 89 | 90.2889 | 90.7485 | 14.3952 | 90.3667 | 91.26 | 91.78
3.2 | 2.1007 | 86.9 | 87.9697 | 88.16 | 3.4988 | 86.9884 | 88.36 | 88.8045 | 5.55 | 87.465 | 88.96 | 89.24 | 8.7 | 88.625 | 89.92 | 90.3819 | 12.8957 | 89.94 | 90.775 | 91.48
3.3 | 2.0007 | 86.375 | 87.5325 | 87.925 | 3.3739 | 86.5717 | 88.1 | 88.3846 | 5.2333 | 86.9652 | 88.48 | 88.95 | 8.28 | 88.25 | 89.64 | 90.182 | 11.5961 | 89.5667 | 90.64 | 91.2667
3.4 | 1.9006 | 86.1 | 87.2125 | 87.62 | 3.099 | 86.1549 | 87.76 | 88.0362 | 4.95 | 86.2655 | 88.22 | 88.48 | 7.8667 | 87.6667 | 89.42 | 89.882 | 10.4466 | 89.16 | 90.34 | 90.9
3.5 | 1.8006 | 85.6 | 86.8524 | 87.18 | 2.899 | 85.5881 | 87.36 | 87.6649 | 4.8 | 85.9156 | 87.86 | 88.14 | 7.4 | 87.1333 | 89.0429 | 89.4571 | 9.9467 | 88.78 | 90.0833 | 90.7
3.6 | 1.7006 | 85.2 | 86.4923 | 86.94 | 2.7991 | 85.188 | 86.96 | 87.3793 | 4.575 | 85.5158 | 87.475 | 87.8667 | 7.1 | 86.8 | 88.56 | 89.1322 | 9.5635 | 88.5 | 89.78 | 90.5
3.7 | 1.7006 | 84.8 | 86.1123 | 86.46 | 2.6991 | 84.5212 | 86.5 | 87.0319 | 4.35 | 84.8661 | 87.24 | 87.52 | 6.9 | 86.45 | 88.2429 | 88.9072 | 9.2303 | 87.975 | 89.5 | 90.2
3.8 | 1.6005 | 84.4 | 85.7522 | 86.075 | 2.5991 | 84.1544 | 86.1333 | 86.6653 | 4.2 | 84.4063 | 86.82 | 87.1 | 6.7 | 86.0333 | 88 | 88.6823 | 8.8637 | 87.4667 | 89.3 | 89.9
3.9 | 1.6005 | 83.7 | 85.3122 | 85.85 | 2.4992 | 83.5709 | 85.86 | 86.2322 | 3.94 | 83.6665 | 86.4 | 86.9 | 6.4 | 85.6 | 87.7667 | 88.2823 | 8.6221 | 87.0667 | 89 | 89.6
4 | 1.5005 | 83.3 | 84.9835 | 85.64 | 2.3992 | 83.0541 | 85.56 | 85.8657 | 3.8 | 83.4067 | 86.2 | 86.62 | 6.2 | 84.9333 | 87.44 | 87.9024 | 8.3972 | 86.6 | 88.7 | 89.25
4.1 | 1.4005 | 82.7 | 84.7121 | 85.2 | 2.1993 | 82.4206 | 85.12 | 85.5658 | 3.65 | 82.7169 | 85.96 | 86.2 | 6 | 84.5667 | 87.25 | 87.6825 | 8.0773 | 86.1 | 88.4 | 89.02
4.2 | 1.4005 | 82.08 | 84.1977 | 84.7 | 2.0993 | 82.0205 | 84.84 | 85.323 | 3.4333 | 81.9673 | 85.7 | 85.72 | 5.8 | 84 | 86.9 | 87.3025 | 7.8974 | 85.85 | 88.1 | 88.72
4.3 | 1.3004 | 81.3 | 83.912 | 84.4 | 2.0993 | 81.2703 | 84.5 | 84.986 | 3.36 | 81.5674 | 85.32 | 85.4 | 5.6667 | 83.4333 | 86.5571 | 86.9826 | 7.6475 | 85.36 | 87.6667 | 88.46
4.4 | 1.3004 | 80.275 | 83.6286 | 84.0667 | 1.9993 | 80.5535 | 84.3 | 84.4662 | 3.2333 | 80.9676 | 85 | 85.225 | 5.525 | 83 | 86.1 | 86.716 | 7.4975 | 85.0333 | 87.3 | 88.16
4.5 | 1.2004 | 78.8667 | 83.2119 | 83.7 | 1.8994 | 79.92 | 83.96 | 84.1663 | 3.1 | 80.1679 | 84.66 | 84.76 | 5.3 | 82.4667 | 85.775 | 86.4494 | 7.2476 | 84.48 | 87.0333 | 87.8667
4.6 | 1.2004 | 76.25 | 82.9118 | 83.4 | 1.7994 | 78.5196 | 83.6 | 83.8522 | 3 | 79.5682 | 84.4 | 84.38 | 5.125 | 82.1 | 85.5667 | 86.1628 | 7.0976 | 84.04 | 86.92 | 87.66
4.7 | 1.2004 | 72.1 | 82.4718 | 83.12 | 1.7494 | 77.0193 | 83.22 | 83.5999 | 2.88 | 78.6885 | 83.98 | 84.14 | 5.1 | 81.3 | 85.35 | 85.7578 | 7.0576 | 83.12 | 86.6 | 87.4333
4.8 | 1.2004 | 64.75 | 82.0117 | 82.825 | 1.6994 | 74.2936 | 82.7667 | 83.2667 | 2.8 | 77.8689 | 83.52 | 83.925 | 4.925 | 80.6667 | 85 | 85.4629 | 6.8311 | 82.75 | 86.225 | 87.04
4.9 | 1.1004 | 56.16 | 81.6517 | 82.3333 | 1.5995 | 69.8175 | 82.42 | 82.9169 | 2.7 | 76.4694 | 83.16 | 83.55 | 4.7 | 80.12 | 84.575 | 85.2579 | 6.5728 | 82.25 | 86 | 86.68
5 | 1.1004 | 46 | 81.3687 | 82.0167 | 1.5995 | 62.9157 | 82 | 82.587 | 2.6 | 74.3703 | 82.92 | 83.18 | 4.6 | 79.4 | 84.3429 | 84.6431 | 6.3979 | 81.74 | 85.76 | 86.4667
5.1 | 1.1004 | 35.9 | 81.0716 | 81.64 | 1.4995 | 55.0138 | 81.775 | 82.0672 | 2.5 | 71.0716 | 82.54 | 82.9 | 4.5 | 78.6 | 83.9857 | 84.3081 | 6.2779 | 81.3 | 85.5 | 86.24
5.2 | 1.1004 | 26.9 | 80.749 | 81.075 | 1.4995 | 45.6864 | 81.5 | 81.9006 | 2.5 | 65.7737 | 82.08 | 82.54 | 4.4 | 78 | 83.54 | 84.1832 | 6.098 | 80.86 | 85.16 | 85.85
5.3 | 1.0003 | 18.975 | 80.3715 | 80.86 | 1.4995 | 36.5091 | 81.06 | 81.6673 | 2.4 | 58.9764 | 81.88 | 82.28 | 4.3 | 77 | 83.2 | 83.8832 | 5.978 | 80.3667 | 84.7667 | 85.6
5.4 | 1.0003 | 12.8 | 79.9914 | 80.58 | 1.3995 | 27.807 | 80.625 | 81.0476 | 2.3 | 50.8796 | 81.55 | 81.68 | 4.1 | 75.5 | 82.85 | 83.5833 | 5.7981 | 79.76 | 84.38 | 85.24
5.5 | 1.0003 | 9.35 | 79.6364 | 80.24 | 1.3995 | 20.5051 | 80.26 | 80.8077 | 2.2 | 42.2831 | 81.2667 | 81.4 | 4 | 73.5 | 82.5375 | 83.2833 | 5.7981 | 79.1 | 84.075 | 85
5.6 | 1.0003 | 7.5 | 79.3446 | 79.78 | 1.2996 | 15.0038 | 79.78 | 80.4079 | 2.2 | 33.8864 | 80.8 | 81.2 | 3.9 | 70.6 | 82.28 | 82.9834 | 5.6981 | 78.52 | 83.9 | 84.8
5.7 | 1.0003 | 6.525 | 79.1113 | 79.55 | 1.2996 | 11.803 | 79.55 | 80.1279 | 2.1 | 26.8892 | 80.46 | 80.85 | 3.8 | 66.1 | 81.92 | 82.6835 | 5.5981 | 77.6 | 83.6 | 84.4
5.8 | 0.9003 | 5.95 | 78.6487 | 79.18 | 1.2996 | 10.0025 | 79.14 | 79.7431 | 2 | 20.925 | 80.04 | 80.42 | 3.75 | 59.7 | 81.58 | 82.3335 | 5.4982 | 76.55 | 83.26 | 84.0333
5.9 | 0.9003 | 5.575 | 78.1512 | 78.94 | 1.2996 | 8.9022 | 78.9 | 79.5282 | 2 | 16.3934 | 79.64 | 80 | 3.6 | 52.9 | 81.25 | 82.0036 | 5.1983 | 75.2 | 83.04 | 83.75
6 | 0.9003 | 5.3 | 77.7968 | 78.5333 | 1.1996 | 8.202 | 78.5 | 79.1683 | 1.9 | 14.061 | 79.2667 | 79.82 | 3.6 | 44.9 | 80.925 | 81.617 | 5.1983 | 73.36 | 82.775 | 83.46
6.1 | 0.9003 | 5.1 | 77.5111 | 78.04 | 1.1996 | 7.7353 | 78.16 | 78.6486 | 1.9 | 12.495 | 78.95 | 79.38 | 3.4667 | 37.7 | 80.7 | 81.3037 | 5.0983 | 70.85 | 82.4 | 83.2333
6.2 | 0.9003 | 4.9 | 77.2485 | 77.7 | 1.1996 | 7.4018 | 77.85 | 78.3487 | 1.8 | 11.662 | 78.7 | 79.02 | 3.3667 | 30.8 | 80.4 | 80.9171 | 4.9983 | 66.3 | 82.075 | 82.92
6.3 | 0.9003 | 4.8 | 76.6909 | 77.3 | 1.1663 | 7.0351 | 77.6 | 77.9288 | 1.8 | 10.8956 | 78.3 | 78.72 | 3.3 | 25.2 | 79.94 | 80.7039 | 4.8984 | 61 | 81.7375 | 82.5333
6.4 | 0.9003 | 4.6 | 76.3609 | 77 | 1.0996 | 6.7767 | 77.1333 | 77.549 | 1.7 | 10.5208 | 77.96 | 78.4 | 3.2667 | 21.35 | 79.6333 | 80.264 | 4.7484 | 54.6 | 81.35 | 82.14
6.5 | 0.9003 | 4.525 | 76.0109 | 76.6333 | 1.0996 | 6.5516 | 76.82 | 77.1691 | 1.7 | 10.096 | 77.62 | 77.98 | 3.2 | 18.975 | 79.175 | 79.884 | 4.6984 | 47.6 | 81.04 | 82
6.6 | 0.82528 | 4.3667 | 75.8108 | 76.24 | 1.0996 | 6.3816 | 76.42 | 76.8493 | 1.7 | 9.6961 | 77.24 | 77.58 | 3.1 | 17.3 | 78.9 | 79.5441 | 4.5985 | 40.8 | 80.68 | 81.58
6.7 | 0.80027 | 4.225 | 75.4394 | 76 | 1.0996 | 6.1015 | 75.9333 | 76.4894 | 1.7 | 9.3962 | 76.84 | 77.28 | 3.05 | 16.2333 | 78.44 | 79.2341 | 4.4985 | 34.4 | 80.325 | 81.3333
6.8 | 0.80027 | 4.2 | 75.0107 | 75.58 | 0.99967 | 5.8815 | 75.62 | 76.1195 | 1.6 | 8.9964 | 76.45 | 77 | 3 | 15.5 | 78.2286 | 78.8985 | 4.3985 | 28.95 | 79.925 | 81.0571
6.9 | 0.80027 | 4.1 | 74.6106 | 75.1333 | 0.99967 | 5.6681 | 75.26 | 75.8297 | 1.6 | 8.7965 | 76.12 | 76.55 | 2.9 | 14.9667 | 77.925 | 78.4843 | 4.3485 | 24.94 | 79.65 | 80.56
7 | 0.80027 | 4 | 74.3249 | 74.8333 | 0.99967 | 5.5014 | 74.9 | 75.4499 | 1.5 | 8.5966 | 75.725 | 76.1 | 2.9 | 14.5 | 77.5333 | 78.0844 | 4.2986 | 22.5 | 79.325 | 80.2143
7.1 |  | 3.8333 | 73.9306 | 74.5 | 0.99967 | 5.4013 | 74.45 | 75.0899 | 1.5 | 8.33 | 75.32 | 75.9 | 2.8 | 14.1 | 76.98 | 77.6845 | 4.1986 | 20.8 | 79.08 | 79.9286
7.2 |  | 3.8 | 73.5438 | 74.06 |  | 5.2013 | 74.1667 | 74.7951 | 1.5 | 8.0568 | 75 | 75.56 | 2.8 | 13.8 | 76.7625 | 77.4095 | 4.0986 | 19.8 | 78.8333 | 79.58
7.3 |  | 3.7 | 73.2305 | 73.7667 |  | 5.1013 | 73.94 | 74.4202 |  | 7.8469 | 74.64 | 75.2 | 2.8 | 13.45 | 76.38 | 77.0846 | 4.0986 | 18.95 | 78.46 | 79.1
7.4 |  | 3.675 | 72.7704 | 73.34 |  | 4.9012 | 73.46 | 74.0704 |  | 7.597 | 74.15 | 74.7 | 2.8 | 12.9667 | 76.12 | 76.8096 | 3.8987 | 18.4 | 78.08 | 78.9
7.5 |  | 3.55 | 72.4704 | 73.04 |  | 4.8012 | 72.98 | 73.6505 |  | 7.3471 | 73.9 | 74.32 |  | 12.675 | 75.82 | 76.4847 | 3.7987 | 17.8 | 77.8 | 78.66
7.6 |  | 3.5 | 72.0503 | 72.66 |  | 4.7012 | 72.7 | 73.3207 |  | 7.1571 | 73.5 | 74.02 |  | 12.4 | 75.42 | 75.9648 | 3.8987 | 17.46 | 77.48 | 78.34
7.7 |  | 3.4 | 71.5102 | 72.28 |  | 4.5011 | 72.3 | 72.9509 |  | 6.9472 | 73.22 | 73.7 |  | 12.18 | 74.88 | 75.6099 |  | 17.125 | 77.1333 | 77.95
7.8 |  | 3.375 | 71.2102 | 71.96 |  | 4.4011 | 71.9 | 72.571 |  | 6.7723 | 72.78 | 73.225 |  | 11.9 | 74.5286 | 75.2849 |  | 16.7 | 76.8 | 77.62
7.9 |  | 3.3 | 70.7851 | 71.48 |  | 4.3011 | 71.5 | 72.1311 |  | 6.5974 | 72.3 | 72.88 |  | 11.58 | 74.2 | 74.9707 |  | 16.4 | 76.5 | 77.26
8 |  | 3.3 | 70.5244 | 71.24 |  | 4.2261 | 71.02 | 71.5714 |  | 6.3974 | 71.9 | 72.5 |  | 11.4 | 73.9625 | 74.4708 |  | 16.2 | 76.125 | 77.075
8.1 |  | 3.2 | 70.0767 | 70.7 |  | 4.1344 | 70.7333 | 71.2715 |  | 6.2975 | 71.54 | 72.08 |  | 11.1 | 73.64 | 74.2102 |  | 15.6667 | 75.625 | 76.9
8.2 |  | 3.1 | 69.6499 | 70.26 |  | 4.101 | 70.3333 | 71.0716 |  | 6.0976 | 71.22 | 71.72 |  | 10.9 | 73.3 | 73.8852 |  | 15.3 | 75.3286 | 76.38
8.3 |  | 3.1 | 69.4099 | 69.88 |  | 4.001 | 69.94 | 70.7146 |  | 5.9976 | 70.72 | 71.18 |  | 10.6 | 72.88 | 73.4853 |  | 15.1 | 75.1 | 76.02
8.4 |  | 3 | 68.8898 | 69.525 |  | 3.901 | 69.575 | 70.2119 |  | 5.8976 | 70.36 | 70.925 |  | 10.3 | 72.3571 | 73.1354 |  | 14.76 | 74.84 | 75.5667
8.5 |  | 3 | 68.5098 | 69.14 |  | 3.8009 | 69.22 | 69.7221 |  | 5.6977 | 69.95 | 70.54 |  | 10.1 | 71.86 | 72.7854 |  | 14.45 | 74.4 | 75.14
8.6 |  | 2.9 | 68.1597 | 68.72 |  | 3.8009 | 68.78 | 69.4222 |  | 5.5978 | 69.52 | 70.1 |  | 9.9333 | 71.4778 | 72.3455 |  | 14.225 | 74.1333 | 74.7667
8.7 |  | 2.9 | 67.7897 | 68.325 |  | 3.7009 | 68.4 | 69.039 |  | 5.3978 | 69.12 | 69.6143 |  | 9.725 | 71.26 | 71.8856 |  | 13.94 | 73.9 | 74.56
8.8 |  | 2.8 | 67.2846 | 67.9667 |  | 3.6009 | 68.05 | 68.6226 |  | 5.2979 | 68.6875 | 69.2 |  | 9.6 | 70.8625 | 71.4457 |  | 13.72 | 73.35 | 74.25
8.9 |  | 2.8 | 66.8896 | 67.5714 |  | 3.5009 | 67.475 | 68.2727 |  | 5.1979 | 68.26 | 68.76 |  | 9.4 | 70.45 | 71.1144 |  | 13.54 | 73 | 73.8714
9 |  | 2.8 | 66.4845 | 67.16 |  | 3.5009 | 67 | 67.7229 |  | 5.1979 | 67.92 | 68.4333 |  | 9.2 | 70 | 70.7358 |  | 13.28 | 72.66 | 73.54
9.1 |  | 2.7 | 66.0694 | 66.82 |  | 3.4009 | 66.475 | 67.2731 |  | 5.0313 | 67.4 | 68.025 |  | 9 | 69.58 | 70.4192 |  | 12.98 | 72.2 | 73.375
9.2 |  | 2.7 | 65.6344 | 66.1 |  | 3.3008 | 66.14 | 66.7233 |  | 4.998 | 67.02 | 67.56 |  | 8.8 | 69.32 | 70.086 |  | 12.78 | 71.825 | 72.9714
9.3 |  | 2.7 | 65.0293 | 65.7429 |  | 3.3008 | 65.7333 | 66.3135 |  | 4.7981 | 66.5889 | 67.16 |  | 8.6 | 68.86 | 69.7003 |  | 12.58 | 71.475 | 72.52
9.4 |  | 2.6 | 64.3892 | 65.2 |  | 3.2008 | 65.26 | 65.8537 |  | 4.6981 | 65.975 | 66.66 |  | 8.5 | 68.4143 | 69.3861 |  | 12.3833 | 71.25 | 72.05
9.5 |  | 2.6 | 63.7091 | 64.2857 |  | 3.2008 | 64.6 | 65.3024 |  | 4.6981 | 65.5667 | 66.26 |  | 8.3 | 67.9857 | 68.7663 |  | 12.2 | 70.9 | 71.62
9.6 |  | 2.6 | 62.4089 | 63.1 |  | 3.2008 | 64.15 | 64.8629 |  | 4.5982 | 65.18 | 65.85 |  | 8.15 | 67.625 | 68.4263 |  | 11.9333 | 70.5333 | 71.2571
9.7 |  | 2.5 | 60.0586 | 60.8571 |  | 3.1008 | 63.56 | 64.2743 |  | 4.4982 | 64.75 | 65.4 |  | 8 | 67.2714 | 67.9664 |  | 11.8 | 70.2 | 71.025
9.8 |  | 2.5 | 56.268 | 57.4667 |  | 3.1008 | 62.82 | 63.5317 |  | 4.4182 | 64.2333 | 64.84 |  | 7.8 | 66.8 | 67.6865 |  | 11.5667 | 69.76 | 70.84
9.9 |  | 2.5 | 51.0073 | 52.12 |  | 3.0008 | 61.675 | 62.6749 |  | 4.2983 | 63.675 | 64.275 |  | 7.7 | 66.44 | 67.2866 |  | 11.4 | 69.45 | 70.3
10 |  | 2.4 | 44.5064 | 45.7333 |  | 3.0008 | 60.2 | 61.1613 |  | 4.2233 | 63.25 | 63.9 |  | 7.6 | 65.9333 | 66.8067 |  | 11.2 | 69.1 | 70
10.1 |  |  | 37.3053 | 38.76 |  | 2.9007 | 57.78 | 58.8564 |  | 4.1983 | 62.7111 | 63.25 |  | 7.4 | 65.475 | 66.3867 |  | 11.06 | 68.6143 | 69.5
10.2 |  |  | 29.9043 | 31.86 |  | 2.9007 | 53.75 | 55.3612 |  | 4.0984 | 62.025 | 62.7667 |  | 7.3 | 65.1 | 65.8868 |  | 10.9 | 68.25 | 69.075
10.3 |  |  | 23.2833 | 25.4857 |  |  | 48.7333 | 50.08 |  | 3.9984 | 61.18 | 61.975 |  | 7.2 | 64.68 | 65.3869 |  | 10.8 | 67.875 | 68.7
10.4 |  |  | 17.6425 | 20.3 |  |  | 42.46 | 44.1823 |  | 3.9234 | 60.24 | 61 |  | 7.1 | 64.22 | 64.887 |  | 10.56 | 67.45 | 68.45
10.5 |  |  | 13.5019 | 16.46 |  |  | 35.7 | 38.0098 |  | 3.8984 | 58.7714 | 59.525 |  | 7 | 63.6857 | 64.6071 |  | 10.325 | 67.06 | 68.02
10.6 |  |  | 11.3016 | 14.3 |  |  | 29.4667 | 32.0072 |  |  | 56.68 | 57.42 |  | 6.9 | 63.26 | 64.0872 |  | 10.1429 | 66.66 | 67.625
10.7 |  |  | 10.1814 | 13.14 |  |  | 23.26 | 26.2095 |  |  | 53.2 | 54.26 |  | 6.8 | 62.8 | 63.7872 |  | 10 | 66.22 | 67.26
10.8 |  |  | 9.6014 | 12.58 |  |  | 18.42 | 21.5485 |  |  | 48.6 | 49.7 |  | 6.6 | 62.25 | 63.1673 |  | 9.8429 | 65.8375 | 66.9333
10.9 |  |  | 9.1413 | 12.1375 |  |  | 15.2 | 18.1784 |  |  | 43.2143 | 44.5286 |  | 6.5 | 61.8 | 62.7474 |  | 9.75 | 65.44 | 66.6
11 |  |  | 8.8013 | 11.8667 |  |  | 13.28 | 16.3334 |  |  | 37 | 39.1 |  |  | 61.34 | 62.2675 |  | 9.5857 | 64.9857 | 66.075
11.1 |  |  | 8.7012 | 11.68 |  |  | 12.2 | 15.094 |  |  | 31.1714 | 33.5286 |  |  | 60.725 | 61.6676 |  | 9.5 | 64.64 | 65.6
11.2 |  |  | 8.5212 | 11.55 |  |  | 11.45 | 14.5371 |  |  | 25.8 | 28.475 |  |  | 60.12 | 61.0478 |  | 9.4 | 64.1857 | 65.25
11.3 |  |  | 8.4012 | 11.4 |  |  | 10.98 | 13.9944 |  |  | 21.3556 | 23.9125 |  |  | 59.4 | 60.4129 |  | 9.2 | 63.66 | 64.7
11.4 |  |  | 8.3012 | 11.2333 |  |  | 10.68 | 13.707 |  |  | 18.1429 | 21.075 |  |  | 58.325 | 59.5481 |  | 9.08 | 63.325 | 64.3
11.5 |  |  | 8.2012 | 11.125 |  |  | 10.4 | 13.3447 |  |  | 16.2714 | 19.2 |  |  | 57.0857 | 58.3217 |  |  | 62.8375 | 64
11.6 |  |  | 8.1012 | 11.0714 |  |  | 10.14 | 13.1547 |  |  | 15.1 | 18.2 |  |  | 55.1 | 56.3721 |  |  | 62.5 | 63.38
11.7 |  |  | 8.0261 | 11 |  |  | 9.95 | 12.9548 |  |  | 14.4 | 17.46 |  |  | 52.15 | 53.5693 |  |  | 62.0444 | 63.0333
11.8 |  |  | 8.0011 | 10.9 |  |  | 9.74 | 12.7806 |  |  | 13.8 | 16.98 |  |  | 48.34 | 49.7901 |  |  | 61.64 | 62.7
11.9 |  |  | 7.9011 | 10.84 |  |  | 9.6 | 12.615 |  |  | 13.45 | 16.5 |  |  | 43.4 | 45.6309 |  |  | 61.0286 | 62.2333
12 |  |  | 7.9011 | 10.8 |  |  | 9.5 | 12.4325 |  |  | 13.1 | 16.2143 |  |  | 38.6 | 41.1668 |  |  | 60.56 | 61.8
12.1 |  |  | 7.8761 | 10.8 |  |  | 9.4 | 12.2618 |  |  | 12.8 | 15.9667 |  |  | 33.6571 | 36.6927 |  |  | 59.9714 | 61.2667
12.2 |  |  | 7.8011 | 10.7 |  |  | 9.2 | 12.1095 |  |  | 12.5 | 15.7 |  |  | 29.1 | 32.3435 |  |  | 59.475 | 60.5
12.3 |  |  | 7.8011 | 10.7 |  |  | 9.1 | 12.0396 |  |  | 12.32 | 15.4667 |  |  | 25.4 | 28.7443 |  |  | 58.9286 | 60
12.4 |  |  | 7.7811 | 10.6 |  |  | 9 | 11.8952 |  |  | 12.12 | 15.175 |  |  | 22.6 | 26.1947 |  |  | 57.8 | 59.1143
12.5 |  |  | 7.7011 | 10.55 |  |  | 8.9 | 11.8619 |  |  | 11.92 | 14.96 |  |  | 20.9 | 24.3618 |  |  | 56.5111 | 57.925
12.6 |  |  | 7.7011 | 10.5 |  |  | 8.8 | 11.7096 |  |  | 11.675 | 14.8 |  |  | 19.6143 | 23.1954 |  |  | 54.84 | 56.4
12.7 |  |  | 7.6011 | 10.5 |  |  | 8.72 | 11.6154 |  |  | 11.525 | 14.66 |  |  | 18.8 | 22.4755 |  |  | 52.2714 | 53.8333
12.8 |  |  | 7.5868 | 10.4 |  |  | 8.68 | 11.5954 |  |  | 11.32 | 14.425 |  |  | 18.35 | 21.8956 |  |  | 49.2 | 51.05
12.9 |  |  | 7.5011 | 10.35 |  |  | 8.5 | 11.4954 |  |  | 11.22 | 14.2 |  |  | 17.9 | 21.4457 |  |  | 45.05 | 47.46
13 |  |  | 7.5011 | 10.3 |  |  | 8.5 | 11.3954 |  |  | 11.0333 | 14.04 |  |  | 17.4714 | 21.1758 |  |  | 40.9 | 43.325
13.1 |  |  | 7.4011 | 10.2 |  |  | 8.4 | 11.2955 |  |  | 10.88 | 13.98 |  |  | 17.22 | 20.8758 |  |  | 36.4222 | 39.42
13.2 |  |  | 7.341 | 10.175 |  |  | 8.3333 | 11.2241 |  |  | 10.7 | 13.8 |  |  | 16.8714 | 20.5359 |  |  | 32.5 | 35.575
13.3 |  |  | 7.241 | 10.1 |  |  | 8.3 | 11.1755 |  |  | 10.6 | 13.7 |  |  | 16.62 | 20.196 |  |  | 29.05 | 32.46
13.4 |  |  | 7.201 | 10 |  |  | 8.2 | 11.0706 |  |  | 10.4 | 13.55 |  |  | 16.3 | 19.996 |  |  | 26.2625 | 29.9667
13.5 |  |  | 7.101 | 9.9 |  |  | 8.1 | 10.9956 |  |  | 10.3 | 13.4 |  |  | 16.125 | 19.6961 |  |  | 24.4 | 28.52
13.6 |  |  | 7.101 | 9.825 |  |  | 8 | 10.8956 |  |  | 10.1 | 13.3 |  |  | 15.9 | 19.4961 |  |  | 23.3 | 27.2
13.7 |  |  | 7.001 | 9.78 |  |  | 7.9 | 10.7957 |  |  | 10.0667 | 13.05 |  |  | 15.62 | 19.2362 |  |  | 22.32 | 26.46
13.8 |  |  | 6.941 | 9.7 |  |  | 7.8 | 10.7207 |  |  | 9.925 | 12.94 |  |  | 15.4 | 18.9962 |  |  | 21.8571 | 26
13.9 |  |  | 6.8176 | 9.6 |  |  | 7.7143 | 10.5958 |  |  | 9.8 | 12.86 |  |  | 15.1571 | 18.6963 |  |  | 21.4 | 25.46
14 |  |  | 6.761 | 9.5167 |  |  | 7.7 | 10.5291 |  |  | 9.7 | 12.75 |  |  | 14.92 | 18.4963 |  |  | 21.0571 | 25.1
14.1 |  |  | 6.701 | 9.42 |  |  | 7.58 | 10.4358 |  |  | 9.525 | 12.6 |  |  | 14.7444 | 18.2963 |  |  | 20.72 | 24.72
14.2 |  |  | 6.6009 | 9.3778 |  |  | 7.5 | 10.3387 |  |  | 9.4714 | 12.45 |  |  | 14.55 | 18.1564 |  |  | 20.4 | 24.34
14.3 |  |  | 6.6009 | 9.3 |  |  | 7.4 | 10.1959 |  |  | 9.38 | 12.36 |  |  | 14.32 | 17.8214 |  |  | 20.125 | 24.1333
14.4 |  |  | 6.5009 | 9.2 |  |  | 7.4 | 10.096 |  |  | 9.2 | 12.22 |  |  | 14.2 | 17.6965 |  |  | 19.9 | 23.82
14.5 |  |  | 6.4009 | 9.1 |  |  | 7.24 | 10.046 |  |  | 9.1 | 12.1 |  |  | 14.025 | 17.4965 |  |  | 19.525 | 23.6
14.6 |  |  | 6.3809 | 9.08 |  |  | 7.2 | 9.916 |  |  | 9.0857 | 12.0167 |  |  | 13.82 | 17.2965 |  |  | 19.4 | 23.28
14.7 |  |  | 6.3009 | 9 |  |  | 7.1 | 9.8461 |  |  | 8.9 | 11.82 |  |  | 13.62 | 17.0966 |  |  | 19.025 | 23.05
14.8 |  |  | 6.2009 | 8.9 |  |  | 7 | 9.7532 |  |  | 8.7857 | 11.75 |  |  | 13.42 | 16.8966 |  |  | 18.8 | 22.88
14.9 |  |  | 6.1009 | 8.8 |  |  | 6.9 | 9.6961 |  |  | 8.7 | 11.66 |  |  | 13.2 | 16.7966 |  |  | 18.6 | 22.5667
15 |  |  | 6.1009 | 8.8 |  |  | 6.875 | 9.5962 |  |  | 8.5571 | 11.5 |  |  | 13.0667 | 16.5567 |  |  | 18.375 | 22.2
15.1 |  |  |  | 8.6 |  |  | 6.7 | 9.5462 |  |  | 8.5 | 11.4 |  |  | 12.9 | 16.3467 |  |  | 18.1 | 22.08
15.2 | [6 empty] | 6.675 | 9.4212 |  |  | 8.4 | 11.3 |  |  | 12.82 | 16.2168 |  |  | 17.76 | 21.9
15.3 | [6 empty] | 6.6 | 9.2963 |  |  | 8.32 | 11.18 |  |  | 12.62 | 16.1135 |  |  | 17.5714 | 21.6
15.4 | [6 empty] | 6.5 | 9.1963 |  |  | 8.2 | 11.0714 |  |  | 12.5 | 15.9218 |  |  | 17.4286 | 21.44
15.5 | [10 empty] | 8.1 | 10.9 |  |  | 12.32 | 15.7968 |  |  | 17.2 | 21.25
15.6 | [10 empty] | 8 | 10.8 |  |  | 12.18 | 15.5969 |  |  | 16.98 | 21.05
15.7 | [10 empty] | 7.925 | 10.7 |  |  | 12 | 15.4969 |  |  | 16.8 | 20.8
15.8 | [10 empty] | 7.85 | 10.6 |  |  | 11.9 | 15.217 |  |  | 16.54 | 20.5
15.9 | [10 empty] | 7.6714 | 10.55 |  |  | 11.8 | 15.122 |  |  | 16.4 | 20.38
16 | [14 empty] | 11.68 | 15.017 |  |  | 16.2 | 20.2143
16.1 | [14 empty] | 11.45 | 14.777 |  |  | 16 | 20
16.2 | [14 empty] | 11.34 | 14.6971 |  |  | 15.86 | 19.8667
16.3 | [14 empty] | 11.16 | 14.5471 |  |  | 15.75 | 19.54
16.4 | [14 empty] | 11.0667 | 14.3771 |  |  | 15.5 | 19.35
16.5 | [14 empty] | 10.9 | 14.2721 |  |  | 15.36 | 19.26
16.6 | [14 empty] | 10.8 | 14.0972 |  |  | 15.125 | 19.05
16.7 | [14 empty] | 10.64 | 13.9772 |  |  | 15.06 | 18.86
16.8 | [14 empty] | 10.45 | 13.8472 |  |  | 14.86 | 18.6875
16.9 | [14 empty] | 10.4 | 13.6773 |  |  | 14.6 | 18.525
17 | [18 empty] | 14.48 | 18.3375
17.1 | [18 empty] | 14.4 | 18.125
17.2 | [18 empty] | 14.2571 | 17.98
17.3 | [18 empty] | 14.075 | 17.8
17.4 | [18 empty] | 13.84 | 17.625
17.5 | [18 empty] | 13.725 | 17.5
17.6 | [18 empty] | 13.58 | 17.3667
17.7 | [18 empty] | 13.4 | 17.1889
17.8 | [18 empty] | 13.3 | 17.05
17.9 | [18 empty] | 13.1889 | 16.84

## Wedge 45 Longitudinal Profiles
Longitudinal Profiles
Wedge 45
SSD=100cm
Field Size=40x20cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-34.9 |  |  |  |  | 10.9222
-34.8 |  |  |  |  | 11
-34.7 |  |  |  |  | 11.1
-34.6 |  |  |  |  | 11.1286
-34.5 |  |  |  |  | 11.2
-34.4 |  |  |  |  | 11.3
-34.3 |  |  |  |  | 11.425
-34.2 |  |  |  |  | 11.5571
-34.1 |  |  |  |  | 11.6
-34 |  |  |  |  | 11.725
-33.9 |  |  |  |  | 11.85
-33.8 |  |  |  |  | 11.9
-33.7 |  |  |  |  | 12.1
-33.6 |  |  |  |  | 12.1833
-33.5 |  |  |  |  | 12.3
-33.4 |  |  |  |  | 12.3857
-33.3 |  |  |  |  | 12.5429
-33.2 |  |  |  |  | 12.675
-33.1 |  |  |  |  | 12.72
-33 |  |  |  |  | 12.8857
-32.9 |  |  |  |  | 12.98
-32.8 |  |  |  |  | 13.1
-32.7 |  |  |  |  | 13.2
-32.6 |  |  |  |  | 13.35
-32.5 |  |  |  |  | 13.5
-32.4 |  |  |  |  | 13.6
-32.3 |  |  |  | 9.3538 | 13.74
-32.2 |  |  |  | 9.425 | 13.82
-32.1 |  |  |  | 9.4875 | 13.92
-32 |  |  |  | 9.5667 | 14.1714
-31.9 |  |  |  | 9.6 | 14.2833
-31.8 |  |  |  | 9.69 | 14.4
-31.7 |  |  |  | 9.8143 | 14.5
-31.6 |  |  |  | 9.94 | 14.6167
-31.5 |  |  |  | 10 | 14.7625
-31.4 |  |  |  | 10.0857 | 14.9
-31.3 |  |  |  | 10.2 | 15.1571
-31.2 |  |  |  | 10.3 | 15.3
-31.1 |  |  |  | 10.3857 | 15.42
-31 |  |  |  | 10.5 | 15.5167
-30.9 |  |  |  | 10.58 | 15.7
-30.8 |  |  |  | 10.68 | 15.9
-30.7 |  |  |  | 10.7 | 16.1
-30.6 |  |  |  | 10.8 | 16.3
-30.5 |  |  |  | 10.98 | 16.4
-30.4 |  |  |  | 11.0571 | 16.6
-30.3 |  |  |  | 11.1875 | 16.7
-30.2 |  |  |  | 11.3 | 16.9
-30.1 |  |  |  | 11.3571 | 17.0625
-30 |  |  |  | 11.4875 | 17.2
-29.9 |  |  |  | 11.6143 | 17.4571
-29.8 |  |  |  | 11.75 | 17.6625
-29.7 |  |  |  | 11.8857 | 17.875
-29.6 |  |  | 7.8 | 12.025 | 18.025
-29.5 |  |  | 7.8889 | 12.15 | 18.2143
-29.4 |  |  | 7.95 | 12.2 | 18.475
-29.3 |  |  | 8 | 12.4125 | 18.65
-29.2 |  |  | 8.0375 | 12.5429 | 18.85
-29.1 |  |  | 8.1625 | 12.675 | 19.0714
-29 |  |  | 8.2 | 12.8111 | 19.3429
-28.9 |  |  | 8.2143 | 12.9222 | 19.65
-28.8 |  |  | 8.3571 | 13.06 | 19.775
-28.7 |  |  | 8.4 | 13.26 | 20.0125
-28.6 |  |  | 8.525 | 13.4067 | 20.2286
-28.5 |  |  | 8.6 | 13.54 | 20.55
-28.4 |  |  | 8.6 | 13.6733 | 20.8125
-28.3 |  | 7.1125 | 8.7 | 13.86 | 20.9857
-28.2 |  | 7.2 | 8.74 | 14.0429 | 21.325
-28.1 |  | 7.2 | 8.84 | 14.1857 | 21.5375
-28 |  | 7.3 | 8.94 | 14.3636 | 21.8857
-27.9 |  | 7.3 | 9 | 14.5067 | 22.15
-27.8 |  | 7.3 | 9.1 | 14.66 | 22.425
-27.7 |  | 7.4 | 9.2 | 14.86 | 22.6429
-27.6 |  | 7.4625 | 9.28 | 15.06 | 22.925
-27.5 |  | 7.5 | 9.38 | 15.26 | 23.325
-27.4 |  | 7.5333 | 9.4 | 15.46 | 23.6286
-27.3 |  | 7.6 | 9.5 | 15.6143 | 23.95
-27.2 | 6.9 | 7.6571 | 9.6 | 15.8091 | 24.3
-27.1 | 6.9 | 7.7 | 9.7 | 16.06 | 24.675
-27 | 7 | 7.78 | 9.88 | 16.275 | 25.05
-26.9 | 7 | 7.8 | 9.9 | 16.5 | 25.5
-26.8 | 7 | 7.85 | 10.0667 | 16.7429 | 26.0857
-26.7 | 7.1 | 7.9 | 10.1 | 16.9 | 26.8
-26.6 | 7.1 | 8 | 10.24 | 17.175 | 27.8
-26.5 | 7.1 | 8 | 10.425 | 17.4429 | 29.0429
-26.4 | 7.1167 | 8.1 | 10.5 | 17.7 | 31.025
-26.3 | 7.2 | 8.1857 | 10.5875 | 17.975 | 33.4
-26.2 | 7.2 | 8.2429 | 10.7143 | 18.2429 | 36.4571
-26.1 | 7.3 | 8.3 | 10.8 | 18.5 | 40.1
-26 | 7.3 | 8.3714 | 10.9714 | 18.7286 | 44.4
-25.9 | 7.3 | 8.4286 | 11.0143 | 18.925 | 48.85
-25.8 | 7.4 | 8.5 | 11.15 | 19.3 | 53.4
-25.7 | 7.4273 | 8.6 | 11.3714 | 19.5857 | 57.7
-25.6 | 7.5 | 8.6286 | 11.44 | 19.925 | 61.75
-25.5 | 7.5 | 8.7 | 11.6333 | 20.2 | 64.95
-25.4 | 7.6 | 8.8 | 11.775 | 20.575 | 67.7429
-25.3 | 7.6333 | 8.9 | 11.9125 | 20.9143 | 69.5167
-25.2 | 7.7 | 8.9625 | 12.0375 | 21.2625 | 71.0167
-25.1 | 7.7 | 9.1 | 12.1714 | 21.575 | 72.2571
-25 | 7.8 | 9.125 | 12.4 | 22.0286 | 72.84
-24.9 | 7.8 | 9.2571 | 12.575 | 22.55 | 73.3714
-24.8 | 7.8875 | 9.3 | 12.7714 | 23.15 | 73.8
-24.7 | 7.925 | 9.425 | 13 | 23.9 | 74.3
-24.6 | 8 | 9.5571 | 13.2 | 24.8857 | 74.75
-24.5 | 8.0875 | 9.6 | 13.3429 | 26.3286 | 75.15
-24.4 | 8.1 | 9.725 | 13.6 | 28.6 | 75.6
-24.3 | 8.15 | 9.8 | 13.775 | 31.3857 | 76.0286
-24.2 | 8.2167 | 9.975 | 14.025 | 34.9571 | 76.3667
-24.1 | 8.3 | 10.025 | 14.3 | 39.55 | 76.575
-24 | 8.4 | 10.1571 | 14.4857 | 44.7429 | 76.84
-23.9 | 8.4375 | 10.2875 | 14.7625 | 50.1 | 77.14
-23.8 | 8.5 | 10.425 | 15 | 55.4429 | 77.64
-23.7 | 8.5 | 10.5571 | 15.2571 | 60.675 | 78.02
-23.6 | 8.6 | 10.775 | 15.525 | 65 | 78.34
-23.5 | 8.6714 | 10.925 | 15.9 | 69.2143 | 78.64
-23.4 | 8.7 | 11.0571 | 16.15 | 72.125 | 78.9714
-23.3 | 8.8375 | 11.1875 | 16.4143 | 74.3 | 79.2286
-23.2 | 8.9 | 11.35 | 16.7625 | 75.5 | 79.625
-23.1 | 9 | 11.5571 | 17.175 | 76.5714 | 79.925
-23 | 9.1 | 11.6875 | 17.6286 | 77.5 | 80.175
-22.9 | 9.1571 | 11.925 | 18.0625 | 78.0333 | 80.4429
-22.8 | 9.2875 | 12.1 | 18.75 | 78.5286 | 80.7
-22.7 | 9.4 | 12.3714 | 19.6 | 79.1375 | 80.9375
-22.6 | 9.4625 | 12.5286 | 20.8 | 79.4 | 81.1429
-22.5 | 9.5875 | 12.8 | 22.6833 | 79.8143 | 81.5
-22.4 | 9.75 | 13.05 | 25.2333 | 80.25 | 81.775
-22.3 | 9.9 | 13.325 | 28.9 | 80.56 | 82
-22.2 | 10 | 13.5857 | 33.9 | 80.875 | 82.2571
-22.1 | 10.2286 | 13.85 | 39.5 | 81.25 | 82.4714
-22 | 10.3 | 14.2375 | 45.94 | 81.55 | 82.7167
-21.9 | 10.4 | 14.6714 | 53.1 | 81.8333 | 82.8625
-21.8 | 10.525 | 15.275 | 60.1 | 82.1667 | 83.2
-21.7 | 10.7143 | 16.1125 | 66.62 | 82.5154 | 83.3571
-21.6 | 10.8875 | 17.35 | 72.26 | 82.9 | 83.625
-21.5 | 11.125 | 19.225 | 76.6 | 83.1667 | 83.9
-21.4 | 11.2571 | 22 | 79.6333 | 83.425 | 84.1571
-21.3 | 11.475 | 26.1 | 81.6 | 83.675 | 84.425
-21.2 | 11.775 | 31.6 | 82.8 | 83.94 | 84.675
-21.1 | 12.15 | 38.0667 | 83.6 | 84.2429 | 84.925
-21 | 12.7286 | 45.8727 | 84.4 | 84.6 | 85.0571
-20.9 | 13.5286 | 54.2 | 85 | 84.85 | 85.275
-20.8 | 15.05 | 62.3429 | 85.3286 | 85.1067 | 85.5286
-20.7 | 17.4286 | 69.575 | 85.74 | 85.3733 | 85.7571
-20.6 | 21.1 | 76.2 | 86.125 | 85.58 | 85.9143
-20.5 | 26.45 | 80.75 | 86.4375 | 85.8333 | 86.05
-20.4 | 33.1714 | 83.8111 | 86.7857 | 86.1667 | 86.2714
-20.3 | 41.3286 | 85.8667 | 87.05 | 86.38 | 86.6143
-20.2 | 50.55 | 87.0857 | 87.45 | 86.6091 | 86.75
-20.1 | 60.1143 | 88.06 | 87.8375 | 86.8643 | 86.8857
-20 | 68.9429 | 88.6444 | 88.0429 | 87.0786 | 87.1
-19.9 | 76.7 | 89.1 | 88.4 | 87.25 | 87.375
-19.8 | 82.8286 | 89.6 | 88.575 | 87.3875 | 87.5714
-19.7 | 86.5714 | 89.9 | 88.8429 | 87.6 | 87.7
-19.6 | 89 | 90.1714 | 89.1 | 87.9 | 87.875
-19.5 | 90.4857 | 90.425 | 89.375 | 88.1143 | 88.0625
-19.4 | 91.25 | 90.6667 | 89.5714 | 88.3 | 88.2273
-19.3 | 91.7286 | 90.95 | 89.8 | 88.4857 | 88.5
-19.2 | 91.975 | 91.2714 | 90.075 | 88.7143 | 88.8133
-19.1 | 92.2429 | 91.4222 | 90.325 | 88.85 | 88.9467
-19 | 92.55 | 91.65 | 90.6 | 89.0375 | 89.08
-18.9 | 92.6857 | 91.9 | 90.8286 | 89.3429 | 89.2067
-18.8 | 92.9222 | 92.1714 | 90.9625 | 89.55 | 89.2733
-18.7 | 93.1222 | 92.3143 | 91.2 | 89.725 | 89.45
-18.6 | 93.2375 | 92.45 | 91.34 | 89.975 | 89.7
-18.5 | 93.3714 | 92.6714 | 91.5143 | 90.1714 | 89.9
-18.4 | 93.5 | 92.9125 | 91.7143 | 90.3 | 90.0667
-18.3 | 93.6375 | 93.0429 | 91.925 | 90.475 | 90.2
-18.2 | 93.8429 | 93.175 | 92.0571 | 90.6714 | 90.4
-18.1 | 94 | 93.3 | 92.275 | 90.8 | 90.5667
-18 | 94.1375 | 93.5 | 92.4571 | 91.075 | 90.7
-17.9 | 94.2625 | 93.5375 | 92.6 | 91.2625 | 90.8333
-17.8 | 94.3 | 93.6714 | 92.7571 | 91.4 | 90.9333
-17.7 | 94.4571 | 93.9 | 92.975 | 91.5286 | 91
-17.6 | 94.6625 | 94.075 | 93.125 | 91.6625 | 91.2
-17.5 | 94.8 | 94.2833 | 93.2889 | 92 | 91.3385
-17.4 | 94.9286 | 94.46 | 93.4938 | 92.1286 | 91.4308
-17.3 | 95 | 94.6 | 93.6813 | 92.2625 | 91.5846
-17.2 | 95.0875 | 94.6714 | 93.82 | 92.4 | 91.66
-17.1 | 95.225 | 94.8333 | 93.9533 | 92.4333 | 91.7571
-17 | 95.3571 | 94.9923 | 94.0867 | 92.5 | 91.9
-16.9 | 95.4875 | 95.1 | 94.22 | 92.7143 | 92.1727
-16.8 | 95.625 | 95.2429 | 94.3533 | 92.9 | 92.3286
-16.7 | 95.7 | 95.4 | 94.4867 | 93.075 | 92.4625
-16.6 | 95.7875 | 95.5375 | 94.6125 | 93.1333 | 92.5
-16.5 | 95.98 | 95.6714 | 94.76 | 93.3 | 92.6
-16.4 | 96.025 | 95.8 | 94.96 | 93.4429 | 92.6625
-16.3 | 96.1571 | 95.975 | 95.1067 | 93.5 | 92.8
-16.2 | 96.375 | 96.1 | 95.24 | 93.7 | 92.925
-16.1 | 96.525 | 96.1875 | 95.3733 | 93.7286 | 93.0667
-16 | 96.6667 | 96.325 | 95.4533 | 93.8625 | 93.1286
-15.9 | 96.7143 | 96.4 | 95.52 | 93.9 | 93.2
-15.8 | 96.8 | 96.4875 | 95.5867 | 94 | 93.2875
-15.7 | 96.8875 | 96.525 | 95.7067 | 94.0625 | 93.44
-15.6 | 96.9143 | 96.6571 | 95.84 | 94.275 | 93.5167
-15.5 | 97.1 | 96.7875 | 95.9733 | 94.325 | 93.7
-15.4 | 97.2 | 96.9143 | 96.0533 | 94.4571 | 93.8
-15.3 | 97.3 | 97 | 96.14 | 94.5875 | 93.9
-15.2 | 97.3429 | 97.175 | 96.2733 | 94.7 | 94
-15.1 | 97.475 | 97.2 | 96.3533 | 94.7875 | 94.2
-15 | 97.6 | 97.25 | 96.42 | 94.825 | 94.3
-14.9 | 97.6429 | 97.375 | 96.4867 | 94.98 | 94.4
-14.8 | 97.775 | 97.4125 | 96.6 | 95.075 | 94.5
-14.7 | 97.8125 | 97.5 | 96.7133 | 95.1 | 94.6
-14.6 | 97.9 | 97.575 | 96.78 | 95.2 | 94.7429
-14.5 | 97.975 | 97.6 | 96.8467 | 95.3 | 94.95
-14.4 | 98.0125 | 97.63 | 96.9143 | 95.4 | 95.0125
-14.3 | 98.1375 | 97.7 | 96.9857 | 95.55 | 95.1
-14.2 | 98.2 | 97.775 | 97 | 95.7 | 95.175
-14.1 | 98.3 | 97.8125 | 97.1375 | 95.72 | 95.3125
-14 | 98.3 | 97.9429 | 97.2 | 95.9 | 95.4375
-13.9 | 98.3 | 98 | 97.2875 | 96 | 95.5714
-13.8 | 98.4 | 98.0125 | 97.3 | 96.0833 | 95.6
-13.7 | 98.4 | 98.1 | 97.3571 | 96.2 | 95.7
-13.6 | 98.4 | 98.1 | 97.4 | 96.25 | 95.7625
-13.5 | 98.5 | 98.2 | 97.5 | 96.3 | 95.9
-13.4 | 98.5 | 98.3375 | 97.55 | 96.3857 | 96
-13.3 | 98.6 | 98.4 | 97.6857 | 96.5 | 96.05
-13.2 | 98.6625 | 98.4 | 97.7 | 96.6 | 96.1
-13.1 | 98.7 | 98.4333 | 97.7 | 96.7857 | 96.2
-13 | 98.8 | 98.5 | 97.7 | 96.8111 | 96.34
-12.9 | 98.8 | 98.5 | 97.8125 | 96.925 | 96.4286
-12.8 | 98.8 | 98.5 | 97.9429 | 97 | 96.6
-12.7 | 98.8 | 98.6 | 98.075 | 97.0667 | 96.7
-12.6 | 98.8571 | 98.6 | 98.1125 | 97.1467 | 96.7
-12.5 | 98.9 | 98.6556 | 98.2 | 97.2133 | 96.725
-12.4 | 99 | 98.7857 | 98.275 | 97.28 | 96.8
-12.3 | 99 | 98.825 | 98.4 | 97.3 | 96.8875
-12.2 | 99.0857 | 98.9 | 98.4375 | 97.4 | 96.9
-12.1 | 99.1143 | 98.9 | 98.5 | 97.4375 | 96.95
-12 | 99.25 | 99 | 98.6 | 97.5 | 96.9143
-11.9 | 99.3 | 99 | 98.6375 | 97.6 | 97
-11.8 | 99.3143 | 99.0857 | 98.7 | 97.6 | 97
-11.7 | 99.4 | 99.1 | 98.7875 | 97.6571 | 97
-11.6 | 99.475 | 99.15 | 98.8 | 97.7875 | 97.1286
-11.5 | 99.5 | 99.2 | 98.8 | 97.8 | 97.4
-11.4 | 99.5429 | 99.25 | 98.875 | 97.85 | 97.5
-11.3 | 99.6 | 99.4 | 98.9125 | 97.9857 | 97.5
-11.2 | 99.62 | 99.4 | 99 | 98.0125 | 97.5
-11.1 | 99.7 | 99.4429 | 99.075 | 98.1 | 97.5
-11 | 99.7 | 99.5 | 99.1 | 98.175 | 97.5714
-10.9 | 99.7714 | 99.5 | 99.1 | 98.2125 | 97.7
-10.8 | 99.8 | 99.5 | 99.1 | 98.3 | 97.7375
-10.7 | 99.9 | 99.5333 | 99.1 | 98.3 | 97.8
-10.6 | 99.9083 | 99.6133 | 99.1429 | 98.4 | 97.9
-10.5 | 99.9917 | 99.68 | 99.2 | 98.4375 | 98.05
-10.4 | 100 | 99.7438 | 99.2 | 98.5714 | 98.14
-10.3 | 100 | 99.8067 | 99.3 | 98.6 | 98.2286
-10.2 | 100.06 | 99.8733 | 99.3625 | 98.6333 | 98.3
-10.1 | 100.1 | 99.9 | 99.4 | 98.7 | 98.44
-10 | 100.1 | 99.9067 | 99.4 | 98.7455 | 98.5143
-9.9 | 100.13 | 99.9733 | 99.4714 | 98.8571 | 98.6
-9.8 | 100.2 | 100 | 99.56 | 98.9 | 98.7
-9.7 | 100.2 | 100 | 99.65 | 98.9727 | 98.7
-9.6 | 100.2 | 100 | 99.7 | 99 | 98.7
-9.5 | 100.2857 | 100 | 99.7 | 99 | 98.7
-9.4 | 100.3 | 100.0067 | 99.7625 | 99 | 98.7182
-9.3 | 100.3 | 100.0733 | 99.8 | 99 | 98.8083
-9.2 | 100.3 | 100.1429 | 99.9 | 99.1231 | 98.8917
-9.1 | 100.3 | 100.2167 | 99.9 | 99.2357 | 98.9
-9 | 100.3 | 100.3 | 99.9 | 99.3 | 99
-8.9 | 100.4 | 100.3 | 100 | 99.3333 | 99
-8.8 | 100.4 | 100.3 | 100 | 99.4308 | 99.0462
-8.7 | 100.4714 | 100.3 | 100 | 99.5 | 99.1231
-8.6 | 100.5 | 100.3 | 100 | 99.5 | 99.2
-8.5 | 100.5 | 100.4 | 100.05 | 99.5 | 99.2125
-8.4 | 100.5889 | 100.4 | 100.1 | 99.5 | 99.3
-8.3 | 100.6 | 100.4 | 100.2 | 99.6 | 99.3714
-8.2 | 100.6 | 100.525 | 100.2 | 99.6 | 99.4
-8.1 | 100.6 | 100.6 | 100.2 | 99.6286 | 99.4286
-8 | 100.6 | 100.6 | 100.2857 | 99.7 | 99.5
-7.9 | 100.6375 | 100.6 | 100.3 | 99.7 | 99.5
-7.8 | 100.7 | 100.6 | 100.3 | 99.8 | 99.5
-7.7 | 100.7 | 100.6714 | 100.3 | 99.8 | 99.5571
-7.6 | 100.725 | 100.7 | 100.38 | 99.8 | 99.6875
-7.5 | 100.8 | 100.7 | 100.4 | 99.9 | 99.7
-7.4 | 100.8 | 100.7 | 100.45 | 99.9 | 99.7
-7.3 | 100.8 | 100.7 | 100.5 | 100 | 99.7
-7.2 | 100.8571 | 100.7455 | 100.5 | 100 | 99.7
-7.1 | 100.9875 | 100.8 | 100.6 | 100 | 99.7
-7 | 101 | 100.8 | 100.6 | 100.1 | 99.8
-6.9 | 101 | 100.8 | 100.6 | 100.1 | 99.8
-6.8 | 101 | 100.8 | 100.6 | 100.0429 | 99.8667
-6.7 | 101 | 100.8 | 100.6 | 100.0875 | 99.9
-6.6 | 101 | 100.8 | 100.6 | 100.1 | 99.9286
-6.5 | 101 | 100.8429 | 100.6 | 100.15 | 100
-6.4 | 101.1 | 100.9 | 100.6 | 100.2 | 100
-6.3 | 101.1 | 100.9143 | 100.6167 | 100.2 | 100
-6.2 | 101.1 | 101 | 100.7 | 100.2 | 100
-6.1 | 101.1 | 101 | 100.7 | 100.2857 | 100.1
-6 | 101.1875 | 101 | 100.8 | 100.3 | 100.1286
-5.9 | 101.2 | 101 | 100.8 | 100.3429 | 100.2
-5.8 | 101.2 | 101 | 100.8 | 100.4 | 100.2
-5.7 | 101.2 | 101.1 | 100.8 | 100.4 | 100.2857
-5.6 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
-5.5 | 101.2 | 101.1 | 100.8 | 100.4 | 100.2143
-5.4 | 101.2 | 101.06 | 100.8 | 100.4 | 100.2143
-5.3 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
-5.2 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3667
-5.1 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
-5 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
-4.9 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
-4.8 | 101.2 | 101.0857 | 100.8 | 100.4 | 100.3
-4.7 | 101.2 | 101 | 100.7 | 100.4 | 100.3
-4.6 | 101.2 | 101 | 100.7 | 100.4 | 100.3
-4.5 | 101.2 | 100.9 | 100.7 | 100.4 | 100.4
-4.4 | 101.2 | 100.9 | 100.7 | 100.3714 | 100.4
-4.3 | 101.2 | 100.9 | 100.7 | 100.3 | 100.4
-4.2 | 101.2 | 100.9 | 100.7 | 100.3167 | 100.3
-4.1 | 101.1 | 100.9 | 100.7714 | 100.4 | 100.3
-4 | 101.1 | 100.9 | 100.8 | 100.4 | 100.3
-3.9 | 101 | 100.9 | 100.8 | 100.4 | 100.4
-3.8 | 101 | 100.9 | 100.8 | 100.4875 | 100.4
-3.7 | 101 | 100.9 | 100.8 | 100.5 | 100.5
-3.6 | 100.9 | 100.9 | 100.8 | 100.5 | 100.4714
-3.5 | 100.9 | 100.8714 | 100.8 | 100.5 | 100.4
-3.4 | 100.9 | 100.8 | 100.8 | 100.5 | 100.4
-3.3 | 100.9 | 100.8 | 100.7 | 100.5 | 100.4
-3.2 | 100.9 | 100.8 | 100.7 | 100.4143 | 100.3429
-3.1 | 100.9 | 100.8 | 100.625 | 100.4143 | 100.3
-3 | 100.8125 | 100.8 | 100.6 | 100.5 | 100.3
-2.9 | 100.7857 | 100.8 | 100.6 | 100.5 | 100.3
-2.8 | 100.7 | 100.8 | 100.6 | 100.4875 | 100.3
-2.7 | 100.7 | 100.7125 | 100.6 | 100.3625 | 100.3
-2.6 | 100.7 | 100.6857 | 100.5625 | 100.3 | 100.3571
-2.5 | 100.7 | 100.6 | 100.5 | 100.3 | 100.4
-2.4 | 100.6143 | 100.6 | 100.5 | 100.3 | 100.4571
-2.3 | 100.6 | 100.5 | 100.4625 | 100.3133 | 100.4125
-2.2 | 100.6 | 100.5 | 100.4 | 100.38 | 100.5
-2.1 | 100.6 | 100.5 | 100.4 | 100.3533 | 100.5
-2 | 100.5143 | 100.5 | 100.3714 | 100.3 | 100.425
-1.9 | 100.5 | 100.5 | 100.3 | 100.3 | 100.4
-1.8 | 100.5 | 100.5 | 100.3875 | 100.2562 | 100.4
-1.7 | 100.5 | 100.4 | 100.3 | 100.2 | 100.3143
-1.6 | 100.4 | 100.4 | 100.3 | 100.2 | 100.3
-1.5 | 100.4 | 100.4 | 100.3 | 100.24 | 100.3
-1.4 | 100.4 | 100.4 | 100.2 | 100.3 | 100.2
-1.3 | 100.3625 | 100.4 | 100.2 | 100.3 | 100.2
-1.2 | 100.2286 | 100.3 | 100.2 | 100.3 | 100.2
-1.1 | 100.2 | 100.3 | 100.2 | 100.3 | 100.1143
-1 | 100.2 | 100.3 | 100.2 | 100.2 | 100.1
-0.9 | 100.2 | 100.22 | 100.2 | 100.2 | 100.1
-0.8 | 100.1 | 100.2 | 100.2 | 100.1286 | 100.1
-0.7 | 100.1 | 100.1 | 100.1733 | 100 | 100.1
-0.6 | 100.0714 | 100.1 | 100.1067 | 100 | 100.1
-0.5 | 100 | 100 | 100.1 | 100 | 100.1
-0.4 | 100 | 100 | 100.1 | 99.9111 | 100.1
-0.3 | 100 | 100.075 | 100.1 | 99.925 | 100.06
-0.2 | 100 | 100.0625 | 100.04 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100.0286 | 100.0067 | 100 | 100
0.3 | 100 | 100.1 | 100.0733 | 99.9875 | 100.1
0.4 | 100 | 100 | 100.1 | 99.9375 | 100.1
0.5 | 100 | 100.08 | 100.1067 | 100 | 100.1
0.6 | 100.0375 | 100.1 | 100.1733 | 100 | 100.1
0.7 | 100.1 | 100.1 | 100.2 | 100.0286 | 100.1
0.8 | 100.1 | 100.2 | 100.2 | 100.1625 | 100.1
0.9 | 100.2 | 100.2 | 100.2 | 100.2 | 100.1
1 | 100.2 | 100.3 | 100.2 | 100.2 | 100.1
1.1 | 100.2 | 100.3 | 100.2 | 100.25 | 100.1857
1.2 | 100.3 | 100.3 | 100.2 | 100.3 | 100.2
1.3 | 100.3286 | 100.4 | 100.2 | 100.3 | 100.2
1.4 | 100.4 | 100.4 | 100.2429 | 100.2733 | 100.2143
1.5 | 100.4 | 100.4 | 100.3 | 100.2067 | 100.3
1.6 | 100.4 | 100.4 | 100.3 | 100.2 | 100.3
1.7 | 100.4875 | 100.4 | 100.3 | 100.225 | 100.3
1.8 | 100.5 | 100.5 | 100.375 | 100.2875 | 100.4
1.9 | 100.5 | 100.5 | 100.3 | 100.3 | 100.4
2 | 100.5 | 100.5 | 100.3286 | 100.32 | 100.4
2.1 | 100.6 | 100.5 | 100.4 | 100.3867 | 100.5
2.2 | 100.6 | 100.5 | 100.4 | 100.3467 | 100.5
2.3 | 100.6 | 100.5 | 100.4286 | 100.3 | 100.425
2.4 | 100.6857 | 100.5857 | 100.5 | 100.3 | 100.4875
2.5 | 100.7 | 100.6 | 100.5 | 100.3 | 100.4
2.6 | 100.7 | 100.65 | 100.5286 | 100.3 | 100.3875
2.7 | 100.7 | 100.7 | 100.6 | 100.3875 | 100.3
2.8 | 100.7143 | 100.8 | 100.6 | 100.425 | 100.3
2.9 | 100.8 | 100.8 | 100.6 | 100.5 | 100.3
3 | 100.8857 | 100.8 | 100.6 | 100.5 | 100.3
3.1 | 100.9 | 100.8 | 100.6875 | 100.4857 | 100.3
3.2 | 100.9 | 100.8 | 100.7 | 100.45 | 100.3125
3.3 | 100.9 | 100.8 | 100.7 | 100.5 | 100.4
3.4 | 100.9 | 100.8 | 100.7857 | 100.5 | 100.4
3.5 | 100.9 | 100.8375 | 100.8 | 100.5 | 100.4
3.6 | 100.9714 | 100.9 | 100.8 | 100.5 | 100.5
3.7 | 101 | 100.9 | 100.8 | 100.5 | 100.4375
3.8 | 101 | 100.9 | 100.8 | 100.5 | 100.4
3.9 | 101 | 100.9 | 100.8 | 100.4286 | 100.4
4 | 101.1 | 100.9 | 100.8 | 100.4 | 100.34
4.1 | 101.1 | 100.9 | 100.8 | 100.4 | 100.3
4.2 | 101.1625 | 100.9 | 100.7 | 100.4 | 100.3
4.3 | 101.2 | 100.9 | 100.7 | 100.3286 | 100.4
4.4 | 101.2 | 100.9 | 100.7 | 100.4 | 100.4
4.5 | 101.2 | 100.9 | 100.7 | 100.4 | 100.4
4.6 | 101.2 | 100.9875 | 100.7 | 100.4 | 100.3
4.7 | 101.2 | 101 | 100.7 | 100.4 | 100.3
4.8 | 101.2 | 101.0571 | 100.7625 | 100.4 | 100.3
4.9 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
5 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
5.1 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
5.2 | 101.2 | 101.1 | 100.8 | 100.4 | 100.36
5.3 | 101.2 | 101.1 | 100.8 | 100.4 | 100.2857
5.4 | 101.2 | 101.04 | 100.8 | 100.4 | 100.2
5.5 | 101.2 | 101.1 | 100.8 | 100.4 | 100.2857
5.6 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
5.7 | 101.2 | 101.1 | 100.8 | 100.4 | 100.3
5.8 | 101.2 | 101.0143 | 100.8 | 100.4 | 100.2
5.9 | 101.2 | 101 | 100.8 | 100.3125 | 100.2
6 | 101.125 | 101 | 100.75 | 100.3 | 100.1625
6.1 | 101.1 | 101 | 100.7 | 100.25 | 100.0286
6.2 | 101.1 | 100.9857 | 100.7 | 100.2 | 100
6.3 | 101.1 | 100.9 | 100.64 | 100.2 | 100
6.4 | 101.1 | 100.9 | 100.6 | 100.2 | 100
6.5 | 101.0125 | 100.8833 | 100.6 | 100.1143 | 99.94
6.6 | 101 | 100.8 | 100.6 | 100.1 | 99.9
6.7 | 101 | 100.8 | 100.6 | 100.05 | 99.8909
6.8 | 101 | 100.8 | 100.6 | 100.075 | 99.8
6.9 | 101 | 100.8 | 100.6 | 100.1 | 99.8
7 | 101 | 100.8667 | 100.6 | 100.1 | 99.8
7.1 | 101 | 100.8538 | 100.6 | 100.04 | 99.7364
7.2 | 100.8875 | 100.775 | 100.5286 | 100 | 99.7
7.3 | 100.8 | 100.7 | 100.5 | 100 | 99.7
7.4 | 100.8 | 100.7 | 100.4857 | 99.9 | 99.7
7.5 | 100.8 | 100.7 | 100.4 | 99.9 | 99.7
7.6 | 100.7625 | 100.7 | 100.4 | 99.86 | 99.625
7.7 | 100.7 | 100.7 | 100.3 | 99.8 | 99.5875
7.8 | 100.7 | 100.6 | 100.3 | 99.8 | 99.5
7.9 | 100.6778 | 100.6 | 100.3 | 99.7714 | 99.5
8 | 100.6 | 100.6 | 100.22 | 99.7 | 99.5
8.1 | 100.6 | 100.6 | 100.2 | 99.6 | 99.4
8.2 | 100.6 | 100.5 | 100.2 | 99.6 | 99.4
8.3 | 100.6 | 100.4375 | 100.15 | 99.6 | 99.4
8.4 | 100.6 | 100.4 | 100.1 | 99.5625 | 99.275
8.5 | 100.5111 | 100.4 | 100.0857 | 99.5 | 99.2
8.6 | 100.5 | 100.3125 | 100 | 99.5 | 99.1417
8.7 | 100.5 | 100.3 | 100 | 99.4833 | 99.0583
8.8 | 100.4 | 100.3 | 100 | 99.4 | 99
8.9 | 100.35 | 100.3 | 99.9571 | 99.3091 | 99
9 | 100.3 | 100.3 | 99.9 | 99.2833 | 98.9308
9.1 | 100.3 | 100.25 | 99.9 | 99.2 | 98.975
9.2 | 100.3 | 100.1688 | 99.8625 | 99.1333 | 98.9273
9.3 | 100.3 | 100.1062 | 99.8 | 99.0375 | 98.8462
9.4 | 100.3 | 100.04 | 99.7 | 99 | 98.75
9.5 | 100.3 | 100 | 99.7 | 99 | 98.7
9.6 | 100.2 | 100 | 99.7 | 99 | 98.7
9.7 | 100.2 | 100 | 99.6833 | 98.9417 | 98.7
9.8 | 100.1583 | 99.9733 | 99.5 | 98.9 | 98.7
9.9 | 100.1 | 99.9067 | 99.5 | 98.9 | 98.6
10 | 100.1 | 99.9 | 99.4 | 98.7222 | 98.5
10.1 | 100.0643 | 99.8733 | 99.4 | 98.7 | 98.4
10.2 | 100 | 99.8067 | 99.4 | 98.7 | 98.3143
10.3 | 100 | 99.74 | 99.3 | 98.6 | 98.2571
10.4 | 100 | 99.7 | 99.2375 | 98.5375 | 98.18
10.5 | 100 | 99.7 | 99.2 | 98.5 | 98
10.6 | 99.9333 | 99.5933 | 99.175 | 98.4 | 97.9
10.7 | 99.9 | 99.5 | 99.1 | 98.3375 | 97.8
10.8 | 99.8286 | 99.5 | 99.1 | 98.3 | 97.7714
10.9 | 99.8 | 99.5 | 99.1 | 98.2 | 97.6375
11 | 99.7 | 99.4467 | 99.1 | 98.2 | 97.5
11.1 | 99.7 | 99.4 | 99.0125 | 98.1125 | 97.5
11.2 | 99.65 | 99.4 | 98.975 | 98 | 97.5
11.3 | 99.6 | 99.35 | 98.9 | 97.95 | 97.5
11.4 | 99.5 | 99.2571 | 98.9 | 97.8143 | 97.5
11.5 | 99.5 | 99.2 | 98.8 | 97.8 | 97.3714
11.6 | 99.5 | 99.1857 | 98.8 | 97.7571 | 97.1
11.7 | 99.4 | 99.1 | 98.7222 | 97.7 | 97
11.8 | 99.35 | 99.1 | 98.7 | 97.6 | 97
11.9 | 99.3 | 99 | 98.6714 | 97.5625 | 97
12 | 99.2 | 98.95 | 98.6 | 97.5 | 96.9143
12.1 | 99.15 | 98.9 | 98.5 | 97.5 | 96.9857
12.2 | 99.1 | 98.9 | 98.4714 | 97.3714 | 96.9
12.3 | 99 | 98.8625 | 98.4 | 97.3 | 96.9
12.4 | 99 | 98.8 | 98.3125 | 97.25 | 96.8
12.5 | 98.925 | 98.6875 | 98.2 | 97.18 | 96.7571
12.6 | 98.8875 | 98.6 | 98.1429 | 97.1133 | 96.7
12.7 | 98.8 | 98.6 | 98.0125 | 97.0333 | 96.6571
12.8 | 98.8 | 98.5 | 97.9714 | 97 | 96.6
12.9 | 98.8 | 98.5 | 97.8 | 96.99 | 96.4571
13 | 98.8 | 98.5 | 97.7143 | 96.8857 | 96.3
13.1 | 98.725 | 98.4714 | 97.7 | 96.75 | 96.2
13.2 | 98.6 | 98.4 | 97.7 | 96.6143 | 96.1
13.3 | 98.6 | 98.4 | 97.6143 | 96.4857 | 96
13.4 | 98.5 | 98.3714 | 97.5857 | 96.35 | 96
13.5 | 98.5 | 98.2286 | 97.5 | 96.3 | 95.86
13.6 | 98.4167 | 98.1 | 97.425 | 96.2143 | 95.8
13.7 | 98.4 | 98.1 | 97.3875 | 96.1857 | 95.6714
13.8 | 98.3375 | 98.0375 | 97.3 | 96.1 | 95.6
13.9 | 98.3 | 98 | 97.3 | 96 | 95.5
14 | 98.3 | 97.9 | 97.1875 | 95.9 | 95.375
14.1 | 98.2375 | 97.8429 | 97.1 | 95.7 | 95.3
14.2 | 98.2 | 97.8 | 97.0375 | 95.7 | 95.2125
14.3 | 98.1 | 97.7 | 97 | 95.6 | 95.1
14.4 | 98.0429 | 97.66 | 96.8933 | 95.4333 | 95.0429
14.5 | 98 | 97.6 | 96.8 | 95.3 | 94.9125
14.6 | 97.9 | 97.6 | 96.8 | 95.2 | 94.775
14.7 | 97.8375 | 97.475 | 96.6933 | 95.1333 | 94.62
14.8 | 97.8 | 97.4 | 96.5812 | 95.04 | 94.5
14.9 | 97.675 | 97.4 | 96.5187 | 94.94 | 94.32
15 | 97.6 | 97.275 | 96.4533 | 94.8 | 94.3
15.1 | 97.5125 | 97.2 | 96.3733 | 94.7571 | 94.2
15.2 | 97.375 | 97.1143 | 96.24 | 94.7 | 94.04
15.3 | 97.3 | 96.9875 | 96.1533 | 94.5571 | 93.94
15.4 | 97.2143 | 96.9 | 96.0867 | 94.425 | 93.84
15.5 | 97.0857 | 96.75 | 96.02 | 94.3 | 93.74
15.6 | 96.95 | 96.6875 | 95.9067 | 94.2625 | 93.55
15.7 | 96.9 | 96.5571 | 95.7733 | 94.1286 | 93.4
15.8 | 96.8 | 96.5 | 95.64 | 94 | 93.3143
15.9 | 96.75 | 96.4 | 95.5533 | 93.9 | 93.2
16 | 96.7 | 96.3571 | 95.4733 | 93.8286 | 93.1571
16.1 | 96.5571 | 96.225 | 95.34 | 93.7 | 93.0875
16.2 | 96.35 | 96.0889 | 95.2067 | 93.6625 | 92.9625
16.3 | 96.1875 | 95.9714 | 95.0733 | 93.5375 | 92.8286
16.4 | 96.0571 | 95.8286 | 94.94 | 93.4125 | 92.7
16.5 | 95.9333 | 95.7 | 94.7687 | 93.2667 | 92.5625
16.6 | 95.8222 | 95.5714 | 94.5867 | 93.1 | 92.5
16.7 | 95.6875 | 95.375 | 94.4533 | 93.0429 | 92.4
16.8 | 95.5625 | 95.2 | 94.32 | 92.9125 | 92.28
16.9 | 95.4286 | 95.075 | 94.1867 | 92.6875 | 92.08
17 | 95.2875 | 94.9467 | 94.0533 | 92.5571 | 91.88
17.1 | 95.1571 | 94.8133 | 93.96 | 92.5 | 91.76
17.2 | 95.1 | 94.7 | 93.88 | 92.3625 | 91.6867
17.3 | 95 | 94.575 | 93.68 | 92.2286 | 91.5533
17.4 | 94.8625 | 94.4273 | 93.4714 | 92.2 | 91.45
17.5 | 94.7286 | 94.2455 | 93.2778 | 91.9875 | 91.3429
17.6 | 94.7 | 94.0714 | 93.1571 | 91.7286 | 91.2
17.7 | 94.4625 | 93.875 | 93.025 | 91.6 | 91.0667
17.8 | 94.3286 | 93.7 | 92.7875 | 91.3625 | 90.9333
17.9 | 94.3 | 93.5714 | 92.64 | 91.2286 | 90.8
18 | 94.1714 | 93.45 | 92.4143 | 91.1 | 90.6667
18.1 | 94.0375 | 93.3375 | 92.225 | 90.8714 | 90.5333
18.2 | 93.8 | 93.2125 | 92.075 | 90.675 | 90.4
18.3 | 93.6429 | 93.05 | 91.8571 | 90.5 | 90.2
18.4 | 93.5 | 92.85 | 91.675 | 90.3714 | 90.0333
18.5 | 93.4 | 92.6286 | 91.45 | 90.175 | 89.9
18.6 | 93.2714 | 92.4857 | 91.2833 | 90.0125 | 89.7
18.7 | 93.1545 | 92.35 | 91.1286 | 89.75 | 89.5062
18.8 | 92.95 | 92.2143 | 91 | 89.5429 | 89.3187
18.9 | 92.7143 | 91.9714 | 90.825 | 89.325 | 89.24
19 | 92.5714 | 91.68 | 90.5571 | 89.06 | 89.1467
19.1 | 92.25 | 91.45 | 90.3 | 88.8143 | 89.0133
19.2 | 92.0143 | 91.2286 | 90.025 | 88.6875 | 88.82
19.3 | 91.725 | 90.9714 | 89.775 | 88.5143 | 88.5667
19.4 | 91.2143 | 90.72 | 89.6 | 88.3143 | 88.2333
19.5 | 90.475 | 90.5 | 89.3429 | 88.0875 | 88
19.6 | 89.0571 | 90.2286 | 89.075 | 87.9 | 87.8714
19.7 | 86.5 | 89.9615 | 88.8 | 87.58 | 87.675
19.8 | 82.6571 | 89.5769 | 88.5714 | 87.3556 | 87.5
19.9 | 76.8286 | 89.1571 | 88.375 | 87.2444 | 87.4429
20 | 68.95 | 88.7 | 88 | 87.08 | 87.1125
20.1 | 59.8714 | 88.08 | 87.85 | 86.8727 | 86.9125
20.2 | 50.6429 | 87.1143 | 87.4714 | 86.6 | 86.7857
20.3 | 41.45 | 85.9111 | 87.0375 | 86.3857 | 86.6
20.4 | 33.0286 | 83.8 | 86.75 | 86.1 | 86.2429
20.5 | 26.5 | 80.5778 | 86.4286 | 85.78 | 86.0857
20.6 | 21.2 | 75.9429 | 86.125 | 85.58 | 85.95
20.7 | 17.625 | 69.8 | 85.75 | 85.38 | 85.7714
20.8 | 15.05 | 62.275 | 85.3333 | 85.1667 | 85.5
20.9 | 13.6571 | 54.1143 | 84.9429 | 84.8333 | 85.25
21 | 12.6857 | 45.8 | 84.4 | 84.54 | 85.0875
21.1 | 12.15 | 38.3833 | 83.6143 | 84.2667 | 84.9143
21.2 | 11.7143 | 31.6 | 82.7 | 83.9333 | 84.65
21.3 | 11.45 | 26.15 | 81.54 | 83.65 | 84.4625
21.4 | 11.275 | 22.0429 | 79.65 | 83.4 | 84.1143
21.5 | 11.0571 | 19.1 | 76.6167 | 83.1333 | 83.85
21.6 | 10.925 | 17.4 | 72.5 | 82.8333 | 83.675
21.7 | 10.6875 | 16.1 | 66.9 | 82.5 | 83.425
21.8 | 10.5625 | 15.275 | 60.4 | 82.14 | 83.1571
21.9 | 10.4286 | 14.7 | 53.3 | 81.825 | 82.9
22 | 10.2875 | 14.2286 | 46.1 | 81.5429 | 82.75
22.1 | 10.1625 | 13.8375 | 39.5 | 81.2375 | 82.5
22.2 | 10.0286 | 13.625 | 33.9667 | 80.8667 | 82.225
22.3 | 9.8875 | 13.2857 | 29.08 | 80.5333 | 82.0167
22.4 | 9.7571 | 13.025 | 25.3571 | 80.1714 | 81.7429
22.5 | 9.625 | 12.7714 | 22.75 | 79.775 | 81.475
22.6 | 9.4875 | 12.55 | 20.82 | 79.375 | 81.2
22.7 | 9.3571 | 12.3286 | 19.6125 | 79.0714 | 80.9429
22.8 | 9.3 | 12.1714 | 18.7429 | 78.525 | 80.675
22.9 | 9.1875 | 11.9 | 18.1 | 78 | 80.5
23 | 9.0625 | 11.7143 | 17.5625 | 77.4571 | 80.1714
23.1 | 9 | 11.4875 | 17.1714 | 76.5375 | 79.9125
23.2 | 8.9 | 11.3571 | 16.775 | 75.5 | 79.625
23.3 | 8.7714 | 11.15 | 16.4625 | 74.2286 | 79.25
23.4 | 8.7 | 11.075 | 16.125 | 72.05 | 78.94
23.5 | 8.6 | 10.8571 | 15.8571 | 69.375 | 78.66
23.6 | 8.5714 | 10.7143 | 15.5 | 65.2 | 78.38
23.7 | 8.5 | 10.575 | 15.225 | 60.7429 | 77.96
23.8 | 8.5 | 10.4 | 14.9571 | 55.6125 | 77.68
23.9 | 8.4 | 10.325 | 14.8 | 49.875 | 77.16
24 | 8.3375 | 10.1875 | 14.525 | 44.7 | 76.86
24.1 | 8.3 | 10.0571 | 14.2571 | 39.4286 | 76.625
24.2 | 8.24 | 9.925 | 14 | 35.0857 | 76.36
24.3 | 8.1875 | 9.7875 | 13.7714 | 31.4 | 76.0429
24.4 | 8.1 | 9.7 | 13.575 | 28.5714 | 75.5667
24.5 | 8.1 | 9.625 | 13.4 | 26.4143 | 75.1714
24.6 | 8 | 9.4875 | 13.1714 | 24.9 | 74.7429
24.7 | 7.9571 | 9.4 | 12.975 | 23.9 | 74.28
24.8 | 7.9 | 9.325 | 12.8 | 23.0625 | 73.82
24.9 | 7.8 | 9.2 | 12.5714 | 22.5857 | 73.35
25 | 7.8 | 9.1571 | 12.375 | 22 | 72.8
25.1 | 7.725 | 9.025 | 12.2 | 21.575 | 72.2
25.2 | 7.7 | 9 | 12.05 | 21.2714 | 70.9714
25.3 | 7.6625 | 8.9 | 11.8429 | 20.95 | 69.4571
25.4 | 7.5444 | 8.8286 | 11.8 | 20.5625 | 67.7143
25.5 | 7.5 | 8.7 | 11.6667 | 20.1714 | 65.0875
25.6 | 7.5 | 8.6625 | 11.48 | 19.85 | 61.7
25.7 | 7.4556 | 8.6 | 11.3143 | 19.6 | 57.85
25.8 | 7.4 | 8.5 | 11.1857 | 19.3143 | 53.275
25.9 | 7.3143 | 8.4 | 11.05 | 18.9571 | 48.9875
26 | 7.3 | 8.4 | 10.9143 | 18.7 | 44.3714
26.1 | 7.3 | 8.3 | 10.8 | 18.4625 | 40.1
26.2 | 7.225 | 8.2714 | 10.75 | 18.2125 | 36.4625
26.3 | 7.2 | 8.2 | 10.625 | 18 | 33.4
26.4 | 7.15 | 8.1 | 10.4875 | 17.7429 | 31
26.5 | 7.1 | 8.0333 | 10.35 | 17.475 | 29.0875
26.6 | 7.1 | 7.9571 | 10.28 | 17.2 | 27.8286
26.7 | 7.1 | 7.9 | 10.1 | 16.9714 | 26.8
26.8 | 7 | 7.8857 | 10.1 | 16.775 | 26.0875
26.9 | 7 | 7.8 | 9.9 | 16.4667 | 25.5429
27 | 6.94 | 7.8 | 9.8 | 16.3 | 25
27.1 | 6.9 | 7.7 | 9.7 | 16.075 | 24.6875
27.2 |  | 7.6875 | 9.6167 | 15.825 | 24.3125
27.3 |  | 7.6 | 9.52 | 15.6364 | 23.9375
27.4 |  | 7.5625 | 9.42 | 15.45 | 23.625
27.5 |  | 7.5 | 9.32 | 15.26 | 23.2857
27.6 |  | 7.4 | 9.22 | 15.06 | 22.9375
27.7 |  | 7.4 | 9.12 | 14.86 | 22.65
27.8 |  | 7.3 | 9.1 | 14.6733 | 22.4429
27.9 |  | 7.3 | 9 | 14.54 | 22.1375
28 |  | 7.2375 | 8.9 | 14.3385 | 21.85
28.1 |  | 7.2 | 8.8 | 14.1333 | 21.5857
28.2 |  | 7.175 | 8.7667 | 13.9733 | 21.325
28.3 |  |  | 8.7 | 13.84 | 20.9714
28.4 |  |  | 8.6 | 13.7067 | 20.7429
28.5 |  |  | 8.5857 | 13.5733 | 20.525
28.6 |  |  | 8.5 | 13.44 | 20.2714
28.7 |  |  | 8.425 | 13.26 | 20.05
28.8 |  |  | 8.3875 | 13.0667 | 19.8286
28.9 |  |  | 8.2571 | 12.9 | 19.5857
29 |  |  | 8.2 | 12.8 | 19.4
29.1 |  |  | 8.1 | 12.7429 | 19.0429
29.2 |  |  | 8.07 | 12.525 | 18.8875
29.3 |  |  | 8 | 12.375 | 18.7143
29.4 |  |  | 7.98 | 12.2429 | 18.45
29.5 |  |  | 7.9 | 12.1125 | 18.1875
29.6 |  |  | 7.8111 | 12 | 18.0571
29.7 |  |  |  | 11.9429 | 17.85
29.8 |  |  |  | 11.8222 | 17.675
29.9 |  |  |  | 11.675 | 17.425
30 |  |  |  | 11.4571 | 17.2286
30.1 |  |  |  | 11.4 | 17.1
30.2 |  |  |  | 11.2875 | 16.82
30.3 |  |  |  | 11.1571 | 16.72
30.4 |  |  |  | 11.025 | 16.54
30.5 |  |  |  | 10.94 | 16.4
30.6 |  |  |  | 10.8 | 16.3
30.7 |  |  |  | 10.74 | 16.08
30.8 |  |  |  | 10.7 | 15.9
30.9 |  |  |  | 10.54 | 15.74
31 |  |  |  | 10.5 | 15.54
31.1 |  |  |  | 10.3571 | 15.4
31.2 |  |  |  | 10.28 | 15.3
31.3 |  |  |  | 10.2 | 15.1625
31.4 |  |  |  | 10.05 | 14.9571
31.5 |  |  |  | 10 | 14.7
31.6 |  |  |  | 9.91 | 14.65
31.7 |  |  |  | 9.7857 | 14.5
31.8 |  |  |  | 9.7 | 14.42
31.9 |  |  |  | 9.65 | 14.32
32 |  |  |  | 9.5333 | 14.12
32.1 |  |  |  | 9.4533 | 13.96
32.2 |  |  |  | 9.3875 | 13.8
32.3 |  |  |  | 9.325 | 13.7
32.4 |  |  |  |  | 13.6
32.5 |  |  |  |  | 13.48
32.6 |  |  |  |  | 13.3833
32.7 |  |  |  |  | 13.2
32.8 |  |  |  |  | 13.1
32.9 |  |  |  |  | 13
33 |  |  |  |  | 12.8167
33.1 |  |  |  |  | 12.7571
33.2 |  |  |  |  | 12.7
33.3 |  |  |  |  | 12.575
33.4 |  |  |  |  | 12.42
33.5 |  |  |  |  | 12.3
33.6 |  |  |  |  | 12.22
33.7 |  |  |  |  | 12.0571
33.8 |  |  |  |  | 11.925
33.9 |  |  |  |  | 11.7875
34 |  |  |  |  | 11.7
34.1 |  |  |  |  | 11.625
34.2 |  |  |  |  | 11.5
34.3 |  |  |  |  | 11.4625
34.4 |  |  |  |  | 11.3286
34.5 |  |  |  |  | 11.2
34.6 |  |  |  |  | 11.1625
34.7 |  |  |  |  | 11.0167
34.8 |  |  |  |  | 10.975
34.9 |  |  |  |  | 10.9

## Wedge 45 Output factors
Output factors
Wedge 45 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20
| 3 | 0.87 | 0.882 | 0.891 | 0.902 | 0.909 | 0.917 | 0.921
| 4 | 0.881 | 0.896 | 0.907 | 0.92 | 0.931 | 0.94 | 0.947
| 5 | 0.891 | 0.907 | 0.92 | 0.935 | 0.948 | 0.96 | 0.968
| 7 | 0.904 | 0.925 | 0.94 | 0.959 | 0.975 | 0.991 | 1.002
| 10 | 0.915 | 0.939 | 0.956 | 0.98 | 1 | 1.023 | 1.038
| 15 | 0.923 | 0.95 | 0.97 | 0.998 | 1.026 | 1.055 | 1.076
| 20 | 0.929 | 0.957 | 0.979 | 1.011 | 1.042 | 1.076 | 1.103
| 30 | 0.936 | 0.967 | 0.99 | 1.026 | 1.063 | 1.107 | 1.14
| 40 | 0.941 | 0.973 | 0.997 | 1.035 | 1.075 | 1.123 | 1.16

## Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 48.4 | 53.6 | 57.4 | 63.4
0.1 | 53.3 | 58.15 | 61.78 | 67.74
0.2 | 61.92 | 66.275 | 69.9 | 73.95
0.3 | 72.625 | 75.92 | 78.66 | 82.25
0.4 | 82.425 | 84.7143 | 86.15 | 89
0.5 | 88.45 | 90.075 | 91.3 | 93.05
0.6 | 92.975 | 94.05 | 94.8 | 96.08
0.7 | 95.7 | 96.66 | 97.05 | 97.875
0.8 | 97.825 | 98.3714 | 98.55 | 99
0.9 | 99.2 | 99.36 | 99.4333 | 99.72
1 | 99.675 | 99.94 | 99.9 | 99.95
1.1 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 100 | 100
1.3 | 99.95 | 100 | 99.7857 | 99.9
1.4 | 99.7 | 99.875 | 99.54 | 99.7
1.5 | 99.5 | 99.65 | 99.4 | 99.2
1.6 | 99.04 | 99.2 | 99.15 | 99
1.7 | 98.48 | 98.9 | 98.6 | 98.66
1.8 | 98.1 | 98.6 | 98.25 | 98.25
1.9 | 97.7 | 98.2 | 97.94 | 97.8667
2 | 97.2 | 97.6 | 97.6 | 97.62
2.1 | 96.6333 | 97.2 | 97.18 | 97.1667
2.2 | 96.12 | 96.82 | 96.76 | 96.725
2.3 | 95.7 | 96.4857 | 96.48 | 96.3667
2.4 | 95.2667 | 95.96 | 96.0143 | 96.025
2.5 | 94.75 | 95.525 | 95.7 | 95.7
2.6 | 94.075 | 95.22 | 95.3125 | 95.26
2.7 | 93.65 | 94.7 | 94.85 | 94.9
2.8 | 93.15 | 94.25 | 94.3857 | 94.55
2.9 | 92.525 | 93.75 | 94.1 | 94.15
3 | 92.1333 | 93.36 | 93.64 | 93.8
3.1 | 91.575 | 92.92 | 93.2 | 93.38
3.2 | 91.075 | 92.5 | 92.95 | 93
3.3 | 90.7 | 92.14 | 92.52 | 92.65
3.4 | 90 | 91.7 | 92.1 | 92.3
3.5 | 89.7 | 91.3 | 91.55 | 91.9
3.6 | 89.3 | 90.9 | 91.2 | 91.6
3.7 | 88.6 | 90.45 | 91.02 | 91.42
3.8 | 88.15 | 90.05 | 90.65 | 91.1
3.9 | 87.66 | 89.6 | 90.175 | 90.44
4 | 87.2 | 89.18 | 89.75 | 90.06
4.1 | 86.55 | 88.6714 | 89.38 | 89.7375
4.2 | 86.1 | 88.42 | 88.975 | 89.44
4.3 | 85.6 | 87.95 | 88.62 | 89.025
4.4 | 85.32 | 87.5 | 88.2 | 88.76
4.5 | 84.95 | 87.3 | 87.85 | 88.225
4.6 | 84.275 | 86.875 | 87.5 | 87.88
4.7 | 83.65 | 86.2333 | 87.1 | 87.52
4.8 | 83.22 | 85.88 | 86.4 | 87.15
4.9 | 82.75 | 85.48 | 86.28 | 86.7286
5 | 82.2 | 85.04 | 85.85 | 86.3
5.1 | 81.86 | 84.6 | 85.56 | 86
5.2 | 81.5286 | 84.2 | 85.0667 | 85.56
5.3 | 80.9 | 83.7667 | 84.65 | 85.32
5.4 | 80.35 | 83.425 | 84.3 | 84.925
5.5 | 79.8 | 83.15 | 83.85 | 84.7333
5.6 | 79.35 | 82.52 | 83.3 | 84.4
5.7 | 78.925 | 82.1 | 83.0625 | 83.7
5.8 | 78.4 | 81.62 | 82.8 | 83.4
5.9 | 78.04 | 81.34 | 82.54 | 83.22
6 | 77.475 | 81.0333 | 82.1 | 82.76
6.1 | 76.975 | 80.6 | 81.7 | 82.3429
6.2 | 76.3 | 80.0667 | 81.375 | 81.95
6.3 | 75.95 | 79.82 | 81.125 | 81.3833
6.4 | 75.5 | 79.4667 | 80.65 | 81.2375
6.5 | 75.08 | 79.175 | 80.18 | 81
6.6 | 74.75 | 78.75 | 79.85 | 80.65
6.7 | 74.2143 | 78.25 | 79.5 | 80.25
6.8 | 73.72 | 77.8 | 79.22 | 79.76
6.9 | 73.2 | 77.3667 | 78.84 | 79.4
7 | 72.8333 | 76.8 | 78.5333 | 79
7.1 | 72.24 | 76.44 | 77.98 | 78.8
7.2 | 71.775 | 76.0857 | 77.6 | 78.5286
7.3 | 71.35 | 75.84 | 77.25 | 77.8
7.4 | 71 | 75.35 | 76.85 | 77.6167
7.5 | 70.5 | 74.84 | 76.56 | 77.3
7.6 | 69.95 | 74.5333 | 76.2 | 76.9
7.7 | 69.62 | 74.1 | 75.6714 | 76.7
7.8 | 69.3 | 73.75 | 75.4429 | 76.325
7.9 | 68.7857 | 73.46 | 75.08 | 75.875
8 | 68.375 | 73.1 | 74.7 | 75.625
8.1 | 67.95 | 72.66 | 74.3 | 75.26
8.2 | 67.4667 | 72.26 | 73.875 | 74.92
8.3 | 67.1 | 71.925 | 73.6 | 74.6
8.4 | 66.775 | 71.42 | 73.3 | 74.2429
8.5 | 66.35 | 71.16 | 72.85 | 74
8.6 | 65.8667 | 70.8333 | 72.62 | 73.5667
8.7 | 65.5 | 70.475 | 72.2667 | 73.3
8.8 | 65.2 | 70.12 | 71.88 | 73.06
8.9 | 64.7 | 69.65 | 71.66 | 72.55
9 | 64.28 | 69.2 | 71.275 | 72.2
9.1 | 63.9714 | 68.85 | 70.76 | 71.9
9.2 | 63.5 | 68.7 | 70.45 | 71.62
9.3 | 63 | 68.42 | 70.2 | 71.28
9.4 | 62.5 | 67.92 | 69.65 | 70.9
9.5 | 62.3 | 67.55 | 69.3667 | 70.44
9.6 | 61.95 | 67.18 | 69.14 | 70.175
9.7 | 61.65 | 66.8 | 68.8429 | 69.94
9.8 | 61.18 | 66.5 | 68.4 | 69.6
9.9 | 60.74 | 66.05 | 67.925 | 69.2
10 | 60.44 | 65.55 | 67.675 | 68.925
10.1 | 59.95 | 65.3 | 67.35 | 68.72
10.2 | 59.46 | 64.95 | 67.1 | 68.26
10.3 | 59.1333 | 64.5 | 66.8286 | 67.9429
10.4 | 58.85 | 64.2333 | 66.5 | 67.6571
10.5 | 58.5 | 63.9 | 66.075 | 67.24
10.6 | 58.18 | 63.6 | 65.8 | 67
10.7 | 57.74 | 63.22 | 65.4 | 66.6
10.8 | 57.36 | 62.75 | 65.025 | 66.4
10.9 | 57.05 | 62.48 | 64.9 | 66.15
11 | 56.7333 | 62.2 | 64.6 | 65.8
11.1 | 56.42 | 61.84 | 64.2 | 65.3667
11.2 | 55.9333 | 61.4667 | 63.92 | 65.2
11.3 | 55.55 | 61.2 | 63.625 | 64.88
11.4 | 55.22 | 60.66 | 63.22 | 64.4333
11.5 | 54.84 | 60.5 | 62.88 | 64.2
11.6 | 54.6 | 59.9333 | 62.56 | 63.9
11.7 | 54.1 | 59.8333 | 62.3 | 63.6
11.8 | 53.775 | 59.6 | 61.9 | 63.26
11.9 | 53.54 | 59.15 | 61.62 | 62.95
12 | 53.3 | 58.82 | 61.3667 | 62.62
12.1 | 53 | 58.55 | 61.12 | 62.3333
12.2 | 52.6 | 58.25 | 60.74 | 62.1429
12.3 | 52.25 | 57.9 | 60.38 | 61.8
12.4 | 51.925 | 57.54 | 60 | 61.6111
12.5 | 51.54 | 57 | 59.8 | 61.2
12.6 | 51.18 | 56.7 | 59.4 | 60.78
12.7 | 50.92 | 56.5667 | 59.0286 | 60.6
12.8 | 50.625 | 56.2667 | 58.8 | 60.36
12.9 | 50.34 | 55.95 | 58.575 | 60
13 | 49.9333 | 55.6 | 58.15 | 59.6
13.1 | 49.72 | 55.2375 | 57.825 | 59.4
13.2 | 49.3333 | 54.9667 | 57.5167 | 59.05
13.3 | 49.075 | 54.66 | 57.26 | 58.82
13.4 | 48.75 | 54.4 | 56.98 | 58.5
13.5 | 48.3 | 54.12 | 56.58 | 58.175
13.6 | 48.15 | 53.75 | 56.3 | 57.925
13.7 | 47.8 | 53.45 | 56.15 | 57.7
13.8 | 47.45 | 53.1 | 55.86 | 57.35
13.9 | 47.125 | 52.9 | 55.7125 | 57.1
14 | 46.8 | 52.65 | 55.4 | 56.95
14.1 | 46.5 | 52.3 | 55.0333 | 56.6
14.2 | 46.24 | 52.02 | 54.74 | 56.4
14.3 | 46.1 | 51.68 | 54.4571 | 56
14.4 | 45.6714 | 51.3667 | 54.15 | 55.74
14.5 | 45.3286 | 51.1 | 53.825 | 55.5
14.6 | 45.1 | 50.7667 | 53.58 | 55.125
14.7 | 44.7857 | 50.4571 | 53.3 | 54.95
14.8 | 44.6 | 50.2 | 53.1 | 54.7
14.9 | 44.375 | 50.05 | 52.8 | 54.42
15 | 44.05 | 49.7 | 52.48 | 54.1
15.1 | 43.72 | 49.4667 | 52.15 | 53.8167
15.2 | 43.4333 | 49.225 | 51.8667 | 53.6333
15.3 | 43.1 | 48.8333 | 51.6 | 53.25
15.4 | 42.8 | 48.475 | 51.3 | 52.96
15.5 | 42.575 | 48.3 | 51.1 | 52.7
15.6 | 42.42 | 48.04 | 50.85 | 52.6
15.7 | 42.075 | 47.875 | 50.575 | 52.25
15.8 | 41.7 | 47.55 | 50.4 | 52
15.9 | 41.5 | 47.26 | 50.1125 | 51.76
16 | 41.3333 | 47.02 | 49.85 | 51.45
16.1 | 41.08 | 46.625 | 49.4857 | 51.2
16.2 | 40.7333 | 46.425 | 49.3 | 51.025
16.3 | 40.5 | 46.175 | 49.0333 | 50.8
16.4 | 40.3143 | 45.8333 | 48.775 | 50.55
16.5 | 40.1 | 45.58 | 48.6 | 50.2833
16.6 | 39.8 | 45.425 | 48.3667 | 49.98
16.7 | 39.65 | 45.12 | 48 | 49.8
16.8 | 39.3667 | 44.84 | 47.775 | 49.475
16.9 | 39.125 | 44.525 | 47.525 | 49.35
17 | 38.85 | 44.38 | 47.36 | 48.9444
17.1 | 38.55 | 44.1333 | 47.05 | 48.6333
17.2 | 38.3667 | 43.95 | 46.78 | 48.5
17.3 | 38.1 | 43.64 | 46.5 | 48.2
17.4 | 37.86 | 43.375 | 46.2714 | 47.95
17.5 | 37.6 | 43.1 | 45.9667 | 47.7
17.6 | 37.35 | 42.825 | 45.675 | 47.575
17.7 | 37.1625 | 42.6 | 45.525 | 47.325
17.8 | 37 | 42.3875 | 45.375 | 46.975
17.9 | 36.7167 | 42.24 | 45.15 | 46.7
18 | 36.4667 | 42 | 44.7857 | 46.35
18.1 | 36.2667 | 41.65 | 44.54 | 46.2
18.2 | 36 | 41.4 | 44.3 | 46.06
18.3 | 35.74 | 41.0667 | 44.1 | 45.9
18.4 | 35.5667 | 40.875 | 43.9 | 45.675
18.5 | 35.35 | 40.7 | 43.75 | 45.46
18.6 | 35 | 40.5 | 43.5 | 45.18
18.7 | 34.8333 | 40.24 | 43.1 | 44.95
18.8 | 34.725 | 40 | 42.95 | 44.8
18.9 | 34.4714 | 39.74 | 42.7 | 44.55
19 | 34.28 | 39.6 | 42.44 | 44.3
19.1 | 34.1 | 39.32 | 42.28 | 44.0667
19.2 | 33.7667 | 39.12 | 42.15 | 43.7875
19.3 | 33.6 | 38.875 | 41.775 | 43.58
19.4 | 33.35 | 38.625 | 41.68 | 43.4
19.5 | 33.1 | 38.52 | 41.45 | 43.275
19.6 | 32.95 | 38.3 | 41.2 | 43.05
19.7 | 32.76 | 38.025 | 41.1 | 42.8
19.8 | 32.6 | 37.75 | 40.8 | 42.55
19.9 | 32.4143 | 37.5 | 40.5 | 42.3
20 | 32.2 | 37.38 | 40.24 | 42.0333
20.1 | 32.0375 | 37.1 | 40.1 | 41.85
20.2 | 31.76 | 36.975 | 39.9 | 41.72
20.3 | 31.6 | 36.825 | 39.68 | 41.4
20.4 | 31.4 | 36.6 | 39.4 | 41.275
20.5 | 31.15 | 36.3857 | 39.15 | 41.1
20.6 | 30.92 | 36.1857 | 39 | 40.9
20.7 | 30.7667 | 35.94 | 38.825 | 40.6571
20.8 | 30.6 | 35.8 | 38.7167 | 40.475
20.9 | 30.38 | 35.46 | 38.52 | 40.3
21 | 30.275 | 35.325 | 38.2857 | 40.2
21.1 | 30.0667 | 35.175 | 38.075 | 39.95
21.2 | 29.88 | 34.95 | 37.9 | 39.7111
21.3 | 29.675 | 34.65 | 37.6 | 39.48
21.4 | 29.525 | 34.4429 | 37.46 | 39.2667
21.5 | 29.3 | 34.25 | 37.2 | 39.0667
21.6 | 29.1667 | 34.1 | 37 | 38.9
21.7 | 29 | 33.96 | 36.8 | 38.725
21.8 | 28.7286 | 33.725 | 36.7 | 38.4
21.9 | 28.525 | 33.5 | 36.4833 | 38.3
22 | 28.4 | 33.25 | 36.1667 | 38.12
22.1 | 28.2 | 33.2 | 35.98 | 37.8
22.2 | 28.1 | 33.025 | 35.8 | 37.6833
22.3 | 27.88 | 32.9 | 35.6 | 37.5
22.4 | 27.7111 | 32.65 | 35.44 | 37.2833
22.5 | 27.5 | 32.5 | 35.2429 | 37.1
22.6 | 27.36 | 32.25 | 35.04 | 36.8
22.7 | 27.2 | 32 | 34.7667 | 36.625
22.8 | 27.1 | 31.88 | 34.6 | 36.5
22.9 | 26.9 | 31.7 | 34.5 | 36.275
23 | 26.725 | 31.525 | 34.4 | 36.125
23.1 | 26.525 | 31.2833 | 34.16 | 36
23.2 | 26.38 | 31.08 | 33.925 | 35.8143
23.3 | 26.15 | 30.975 | 33.7 | 35.6
23.4 | 26.1 | 30.7 | 33.5 | 35.35
23.5 | 25.9 | 30.6 | 33.375 | 35.18
23.6 | 25.7714 | 30.4286 | 33.2 | 35.05
23.7 | 25.52 | 30.2 | 33.02 | 34.9
23.8 | 25.4 | 30.1 | 32.875 | 34.675
23.9 | 25.275 | 29.9 | 32.675 | 34.5333
24 | 25.2 | 29.7 | 32.425 | 34.34
24.1 | 25 | 29.475 | 32.3 | 34.12
24.2 | 24.85 | 29.3 | 32.14 | 33.975
24.3 | 24.75 | 29.2167 | 31.98 | 33.84
24.4 | 24.55 | 29.0333 | 31.85 | 33.65
24.5 | 24.34 | 28.8 | 31.6 | 33.45
24.6 | 24.1714 | 28.8 | 31.5 | 33.3143
24.7 | 24.1 | 28.7 | 31.3333 | 33.2
24.8 | 23.9375 | 28.4667 | 31.1 | 33
24.9 | 23.75 | 28.225 | 30.9667 | 32.8286
25 | 23.6571 | 28.1 | 30.8 | 32.7
25.1 | 23.5 | 27.94 | 30.65 | 32.45
25.2 | 23.3 | 27.7 | 30.44 | 32.275
25.3 | 23.2 | 27.6 | 30.2 | 32.025
25.4 | 23.025 | 27.48 | 30 | 31.9
25.5 | 22.9 | 27.3 | 29.9 | 31.8333
25.6 | 22.76 | 27.1 | 29.8 | 31.65
25.7 | 22.62 | 26.95 | 29.6143 | 31.5
25.8 | 22.5 | 26.8 | 29.5 | 31.3667
25.9 | 22.3 | 26.7 | 29.4 | 31.225
26 | 22.2 | 26.5143 | 29.1333 | 31.075
26.1 | 22.1444 | 26.3333 | 29 | 30.8
26.2 | 21.975 | 26.175 | 28.8444 | 30.65
26.3 | 21.8667 | 26.1 | 28.7 | 30.5
26.4 | 21.7 | 26 | 28.5667 | 30.35
26.5 | 21.52 | 25.85 | 28.4 | 30.2667
26.6 | 21.42 | 25.7 | 28.3 | 30.05
26.7 | 21.38 | 25.5 | 28.06 | 29.88
26.8 | 21.2 | 25.4 | 27.825 | 29.7833
26.9 | 21.125 | 25.22 | 27.75 | 29.68
27 | 21 | 25.1 | 27.7 | 29.48
27.1 | 20.88 | 25 | 27.5 | 29.3
27.2 | 20.8 | 24.74 | 27.4286 | 29.1667
27.3 | 20.6 | 24.6 | 27.25 | 28.925
27.4 | 20.36 | 24.5 | 27.1 | 28.775
27.5 | 20.3 | 24.4 | 26.9 | 28.7
27.6 | 20.1625 | 24.2 | 26.7 | 28.6
27.7 | 20 | 24.0875 | 26.5833 | 28.46
27.8 | 19.9167 | 24 | 26.4 | 28.3
27.9 | 19.8 | 23.9 | 26.3 | 28.175
28 | 19.7 | 23.7 | 26.16 | 28
28.1 | 19.52 | 23.5 | 26.04 | 27.925
28.2 | 19.5 | 23.3889 | 25.9 | 27.7
28.3 | 19.3571 | 23.26 | 25.8 | 27.6
28.4 | 19.2333 | 23.04 | 25.675 | 27.5
28.5 | 19.15 | 22.92 | 25.5 | 27.3
28.6 | 19 | 22.8 | 25.34 | 27.1667
28.7 | 18.875 | 22.725 | 25.2 | 27
28.8 | 18.7 | 22.6 | 25.1 | 26.86
28.9 | 18.6 | 22.54 | 24.975 | 26.8
29 | 18.56 | 22.375 | 24.8 | 26.5875
29.1 | 18.425 | 22.225 | 24.65 | 26.4571
29.2 | 18.325 | 22.1 | 24.6 | 26.3
29.3 | 18.18 | 21.975 | 24.4 | 26.2
29.4 | 18.0714 | 21.86 | 24.3 | 26.1
29.5 | 17.98 | 21.7 | 24.125 | 25.9
29.6 | 17.8 | 21.6 | 24.02 | 25.8
29.7 | 17.7 | 21.5 | 23.9 | 25.6667
29.8 | 17.7 | 21.4 | 23.8 | 25.45
29.9 | 17.5 | 21.26 | 23.6714 | 25.35
30 | 17.48 | 21.1 | 23.48 | 25.25

## SSD80 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 48.3 | 54.5 | 60.3 | 68.3
0.1 | 53.54 | 59.7 | 64.9 | 72.3333
0.2 | 61.3333 | 66.62 | 71.1 | 77.92
0.3 | 72.5 | 77.2 | 80.3429 | 85
0.4 | 81.8 | 85.1 | 87.5 | 90.82
0.5 | 88.12 | 90.6 | 92.4 | 94.4333
0.6 | 92.72 | 94.48 | 95.6 | 97.125
0.7 | 96 | 97.08 | 97.8 | 98.65
0.8 | 97.98 | 98.74 | 98.9 | 99.475
0.9 | 99.18 | 99.64 | 99.7 | 99.95
1 | 99.82 | 99.98 | 100 | 100
1.1 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 100 | 99.92
1.3 | 99.9 | 100 | 99.82 | 99.64
1.4 | 99.6 | 99.6 | 99.44 | 99.16
1.5 | 99.18 | 99.05 | 99.0333 | 98.78
1.6 | 98.75 | 98.6 | 98.5667 | 98.4
1.7 | 98.28 | 98.125 | 98.15 | 97.96
1.8 | 97.75 | 97.7 | 97.8 | 97.5333
1.9 | 97.04 | 97.3333 | 97.34 | 97.1
2 | 96.45 | 96.8667 | 96.94 | 96.6
2.1 | 95.9818 | 96.25 | 96.36 | 96.3
2.2 | 95.4 | 95.95 | 95.74 | 95.8
2.3 | 94.7545 | 95.42 | 95.2333 | 95.28
2.4 | 94.2 | 94.86 | 94.7 | 94.7333
2.5 | 93.58 | 94.4 | 94.2 | 94.325
2.6 | 92.94 | 93.85 | 93.9 | 93.9667
2.7 | 92.4 | 93.5 | 93.3 | 93.375
2.8 | 91.84 | 93 | 93.05 | 93.0333
2.9 | 91.24 | 92.5 | 92.56 | 92.55
3 | 90.7273 | 91.9 | 92.16 | 92.14
3.1 | 90.15 | 91.5 | 91.66 | 91.6667
3.2 | 89.55 | 91 | 91.14 | 91.34
3.3 | 89 | 90.5 | 90.66 | 90.8333
3.4 | 88.45 | 90 | 90.26 | 90.44
3.5 | 87.9 | 89.6 | 89.84 | 89.96
3.6 | 87.34 | 89.02 | 89.36 | 89.46
3.7 | 86.8 | 88.5 | 88.86 | 89
3.8 | 86.18 | 88.1 | 88.4 | 88.48
3.9 | 85.66 | 87.72 | 87.9 | 88.1
4 | 85.16 | 87.16 | 87.5 | 87.7143
4.1 | 84.7 | 86.66 | 87.1 | 87.2857
4.2 | 84.16 | 86.14 | 86.6 | 86.9
4.3 | 83.58 | 85.72 | 86.1 | 86.4875
4.4 | 82.88 | 85.24 | 85.7 | 86.05
4.5 | 82.38 | 84.8667 | 85.1333 | 85.6167
4.6 | 81.82 | 84.3 | 84.8 | 85.28
4.7 | 81.35 | 83.86 | 84.4333 | 84.8333
4.8 | 80.9 | 83.36 | 84.0333 | 84.3667
4.9 | 80.38 | 82.86 | 83.5667 | 83.85
5 | 79.8 | 82.4667 | 83.2 | 83.45
5.1 | 79.38 | 81.96 | 82.76 | 83.0667
5.2 | 78.68 | 81.56 | 82.24 | 82.76
5.3 | 78.22 | 81.0333 | 81.86 | 82.16
5.4 | 77.58 | 80.46 | 81.36 | 81.76
5.5 | 77.08 | 80.08 | 80.86 | 81.46
5.6 | 76.5 | 79.68 | 80.55 | 80.96
5.7 | 76.16 | 79.28 | 80.26 | 80.6
5.8 | 75.75 | 78.88 | 79.72 | 80.3
5.9 | 75.125 | 78.375 | 79.28 | 79.7
6 | 74.6 | 77.95 | 78.9091 | 79.3
6.1 | 74.16 | 77.5333 | 78.45 | 79
6.2 | 73.7429 | 77 | 78.02 | 78.6
6.3 | 73.1333 | 76.5 | 77.62 | 78.24
6.4 | 72.5667 | 76.1 | 77.15 | 77.74
6.5 | 72.2 | 75.7 | 76.79 | 77.34
6.6 | 71.68 | 75.2 | 76.42 | 76.94
6.7 | 71.28 | 74.8 | 75.95 | 76.52
6.8 | 70.775 | 74.3 | 75.52 | 76.18
6.9 | 70.2 | 73.9 | 75.05 | 75.68
7 | 69.9 | 73.475 | 74.5556 | 75.32
7.1 | 69.3 | 73.06 | 74.15 | 74.925
7.2 | 68.8 | 72.58 | 73.85 | 74.5143
7.3 | 68.42 | 72.22 | 73.46 | 73.98
7.4 | 67.94 | 71.76 | 73.06 | 73.7
7.5 | 67.35 | 71.275 | 72.7 | 73.25
7.6 | 66.96 | 70.94 | 72.3 | 72.9143
7.7 | 66.4667 | 70.4667 | 71.8 | 72.4
7.8 | 65.92 | 70 | 71.425 | 72.14
7.9 | 65.58 | 69.675 | 71 | 71.7
8 | 65.14 | 69.2 | 70.675 | 71.44
8.1 | 64.78 | 68.75 | 70.15 | 70.95
8.2 | 64.2 | 68.3 | 69.85 | 70.64
8.3 | 63.82 | 68.04 | 69.5 | 70.2667
8.4 | 63.42 | 67.6 | 69.1 | 69.85
8.5 | 63.02 | 67.2 | 68.75 | 69.54
8.6 | 62.72 | 66.7 | 68.2 | 69.18
8.7 | 62.08 | 66.4 | 67.9167 | 68.88
8.8 | 61.62 | 65.9667 | 67.6 | 68.54
8.9 | 61.26 | 65.6 | 67.32 | 68.18
9 | 60.88 | 65.32 | 66.84 | 67.74
9.1 | 60.4333 | 64.88 | 66.52 | 67.38
9.2 | 60.06 | 64.42 | 66.06 | 66.98
9.3 | 59.5571 | 64.02 | 65.78 | 66.64
9.4 | 59.175 | 63.58 | 65.44 | 66.42
9.5 | 58.8286 | 63.26 | 64.92 | 65.95
9.6 | 58.36 | 62.875 | 64.58 | 65.52
9.7 | 57.96 | 62.44 | 64.24 | 65.22
9.8 | 57.65 | 62.12 | 63.88 | 64.96
9.9 | 57.2667 | 61.74 | 63.48 | 64.6
10 | 56.8 | 61.24 | 63.14 | 64.1
10.1 | 56.44 | 61.02 | 62.88 | 63.9
10.2 | 56.02 | 60.68 | 62.48 | 63.5333
10.3 | 55.64 | 60.28 | 62.08 | 63.1
10.4 | 55.2667 | 59.8667 | 61.76 | 62.8167
10.5 | 54.9 | 59.525 | 61.42 | 62.46
10.6 | 54.5 | 59.18 | 61 | 62.15
10.7 | 54.15 | 58.78 | 60.65 | 61.76
10.8 | 53.85 | 58.48 | 60.32 | 61.45
10.9 | 53.5333 | 58.08 | 60.06 | 61.16
11 | 53 | 57.78 | 59.72 | 60.76
11.1 | 52.7286 | 57.38 | 59.36 | 60.46
11.2 | 52.3333 | 56.98 | 59.02 | 60.08
11.3 | 52.0286 | 56.64 | 58.66 | 59.86
11.4 | 51.64 | 56.28 | 58.32 | 59.5
11.5 | 51.4 | 56 | 57.96 | 59.1
11.6 | 51.04 | 55.6 | 57.66 | 58.8667
11.7 | 50.66 | 55.3 | 57.32 | 58.5
11.8 | 50.28 | 54.95 | 57.02 | 58.2
11.9 | 49.88 | 54.52 | 56.72 | 58
12 | 49.56 | 54.16 | 56.42 | 57.6
12.1 | 49.1429 | 53.9 | 56.0667 | 57.2
12.2 | 48.88 | 53.575 | 55.7333 | 56.9
12.3 | 48.54 | 53.1667 | 55.36 | 56.7
12.4 | 48.18 | 52.95 | 55.08 | 56.3
12.5 | 47.78 | 52.58 | 54.76 | 55.9
12.6 | 47.44 | 52.28 | 54.46 | 55.72
12.7 | 47.24 | 51.88 | 54.08 | 55.3
12.8 | 47 | 51.58 | 53.76 | 55.1167
12.9 | 46.6 | 51.36 | 53.38 | 54.84
13 | 46.3 | 51 | 53.18 | 54.44
13.1 | 45.9 | 50.64 | 52.86 | 54.2
13.2 | 45.64 | 50.34 | 52.56 | 53.8286
13.3 | 45.3 | 50.08 | 52.2 | 53.52
13.4 | 45 | 49.68 | 52.0333 | 53.2333
13.5 | 44.7 | 49.44 | 51.7 | 52.98
13.6 | 44.4 | 49.15 | 51.3667 | 52.7286
13.7 | 44.1 | 48.82 | 51.0333 | 52.42
13.8 | 43.8 | 48.42 | 50.7667 | 52.1667
13.9 | 43.6 | 48.16 | 50.4 | 51.8333
14 | 43.3 | 47.92 | 50.1333 | 51.5714
14.1 | 42.9 | 47.58 | 49.8 | 51.32
14.2 | 42.72 | 47.36 | 49.5 | 51.07
14.3 | 42.36 | 47.08 | 49.3 | 50.77
14.4 | 42.14 | 46.76 | 48.98 | 50.46
14.5 | 41.82 | 46.48 | 48.8 | 50.16
14.6 | 41.44 | 46.06 | 48.5 | 49.88
14.7 | 41.24 | 45.86 | 48.3 | 49.575
14.8 | 40.94 | 45.5 | 47.9 | 49.3143
14.9 | 40.6667 | 45.38 | 47.7 | 49.05
15 | 40.32 | 45.06 | 47.3 | 48.77
15.1 | 40.14 | 44.78 | 47.1 | 48.4833
15.2 | 39.88 | 44.46 | 46.8 | 48.1857
15.3 | 39.64 | 44.18 | 46.54 | 47.975
15.4 | 39.28 | 43.96 | 46.3 | 47.725
15.5 | 39.08 | 43.58 | 45.9 | 47.45
15.6 | 38.8833 | 43.3667 | 45.675 | 47.15
15.7 | 38.6 | 43.14 | 45.46 | 46.875
15.8 | 38.2333 | 42.825 | 45.16 | 46.7
15.9 | 37.98 | 42.525 | 44.86 | 46.4
16 | 37.75 | 42.25 | 44.72 | 46.1
16.1 | 37.5 | 42.04 | 44.42 | 45.9
16.2 | 37.3 | 41.825 | 44.12 | 45.6
16.3 | 37.1 | 41.5 | 43.86 | 45.5
16.4 | 36.9 | 41.225 | 43.6 | 45.1
16.5 | 36.5 | 41.0333 | 43.3 | 44.9182
16.6 | 36.3 | 40.7333 | 43.1 | 44.63
16.7 | 36.1 | 40.55 | 43 | 44.42
16.8 | 35.9 | 40.18 | 42.6 | 44.1714
16.9 | 35.6 | 40 | 42.32 | 43.88
17 | 35.34 | 39.7667 | 42.12 | 43.62
17.1 | 35.12 | 39.5333 | 41.8625 | 43.38
17.2 | 34.82 | 39.24 | 41.54 | 43.1333
17.3 | 34.62 | 39.04 | 41.36 | 42.9143
17.4 | 34.32 | 38.78 | 41.08 | 42.68
17.5 | 34.02 | 38.38 | 40.88 | 42.3667
17.6 | 33.92 | 38.14 | 40.6727 | 42.1
17.7 | 33.62 | 38.04 | 40.4 | 42
17.8 | 33.44 | 37.84 | 40.1727 | 41.68
17.9 | 33.22 | 37.58 | 39.9 | 41.5
18 | 33.02 | 37.24 | 39.7 | 41.3
18.1 | 32.82 | 37.04 | 39.48 | 41.1
18.2 | 32.54 | 36.86 | 39.2667 | 40.8
18.3 | 32.4 | 36.72 | 39 | 40.6
18.4 | 32.2333 | 36.36 | 38.7 | 40.3
18.5 | 31.94 | 36.12 | 38.5 | 40
18.6 | 31.78 | 35.96 | 38.3111 | 39.9
18.7 | 31.48 | 35.66 | 38.1 | 39.7
18.8 | 31.28 | 35.52 | 37.875 | 39.4
18.9 | 31.08 | 35.26 | 37.66 | 39.2333
19 | 30.8833 | 35.1 | 37.42 | 39
19.1 | 30.7286 | 34.96 | 37.26 | 38.8167
19.2 | 30.55 | 34.72 | 36.975 | 38.52
19.3 | 30.3 | 34.45 | 36.78 | 38.3
19.4 | 30.1 | 34.14 | 36.58 | 38.14
19.5 | 29.8714 | 33.96 | 36.3 | 37.9
19.6 | 29.7 | 33.76 | 36.1286 | 37.74
19.7 | 29.5 | 33.6 | 35.92 | 37.44
19.8 | 29.3 | 33.425 | 35.74 | 37.3
19.9 | 29.1 | 33.1667 | 35.54 | 37.12
20 | 28.96 | 33 | 35.34 | 36.9
20.1 | 28.7 | 32.8 | 35.04 | 36.7429
20.2 | 28.5 | 32.575 | 34.8 | 36.49
20.3 | 28.3 | 32.42 | 34.6857 | 36.2667
20.4 | 28.2 | 32.18 | 34.4667 | 36.1714
20.5 | 28.0182 | 31.98 | 34.24 | 35.92
20.6 | 27.82 | 31.78 | 34.04 | 35.6857
20.7 | 27.62 | 31.58 | 33.94 | 35.4667
20.8 | 27.51 | 31.5 | 33.74 | 35.28
20.9 | 27.32 | 31.3 | 33.44 | 35.1
21 | 27.12 | 30.98 | 33.24 | 34.96
21.1 | 26.92 | 30.7833 | 33.1 | 34.68
21.2 | 26.72 | 30.6 | 32.94 | 34.48
21.3 | 26.52 | 30.44 | 32.65 | 34.28
21.4 | 26.41 | 30.34 | 32.46 | 34.16
21.5 | 26.22 | 30.1 | 32.36 | 33.98
21.6 | 25.93 | 29.9 | 32.16 | 33.68
21.7 | 25.9 | 29.7 | 31.96 | 33.48
21.8 | 25.72 | 29.5 | 31.76 | 33.3
21.9 | 25.61 | 29.3 | 31.56 | 33.175
22 | 25.42 | 29.12 | 31.36 | 32.94
22.1 | 25.2364 | 28.92 | 31.16 | 32.8
22.2 | 25.04 | 28.8 | 31 | 32.64
22.3 | 24.84 | 28.6 | 30.86 | 32.4
22.4 | 24.72 | 28.4 | 30.66 | 32.2
22.5 | 24.5 | 28.2 | 30.425 | 32.0667
22.6 | 24.35 | 28.1 | 30.32 | 31.84
22.7 | 24.2 | 27.9667 | 30.1 | 31.7
22.8 | 24.06 | 27.72 | 30 | 31.54
22.9 | 23.86 | 27.5333 | 29.76 | 31.34
23 | 23.7 | 27.4 | 29.58 | 31.2
23.1 | 23.6 | 27.2 | 29.38 | 30.94
23.2 | 23.4 | 27.1 | 29.28 | 30.86
23.3 | 23.26 | 26.9 | 29.08 | 30.66
23.4 | 23.06 | 26.775 | 28.88 | 30.5
23.5 | 22.96 | 26.5667 | 28.78 | 30.36
23.6 | 22.86 | 26.44 | 28.58 | 30.1
23.7 | 22.66 | 26.24 | 28.38 | 30
23.8 | 22.58 | 26.14 | 28.28 | 29.82
23.9 | 22.4 | 25.94 | 28.08 | 29.6667
24 | 22.28 | 25.74 | 27.9 | 29.52
24.1 | 22.08 | 25.6 | 27.78 | 29.36
24.2 | 21.98 | 25.44 | 27.6 | 29.2
24.3 | 21.8667 | 25.3 | 27.4 | 28.975
24.4 | 21.74 | 25.14 | 27.3 | 28.82
24.5 | 21.54 | 24.94 | 27.1 | 28.6
24.6 | 21.34 | 24.84 | 26.9 | 28.5
24.7 | 21.225 | 24.64 | 26.8 | 28.3333
24.8 | 21.1 | 24.5 | 26.6 | 28.2
24.9 | 21 | 24.34 | 26.4 | 28.0333
25 | 20.8 | 24.24 | 26.3 | 27.9
25.1 | 20.7 | 24.04 | 26.2 | 27.7333
25.2 | 20.6 | 23.96 | 26 | 27.52
25.3 | 20.4 | 23.76 | 25.8167 | 27.4
25.4 | 20.3 | 23.6 | 25.7167 | 27.3
25.5 | 20.2 | 23.46 | 25.6 | 27.125
25.6 | 20.1 | 23.36 | 25.32 | 26.92
25.7 | 19.96 | 23.26 | 25.22 | 26.8
25.8 | 19.8 | 23.1 | 25.12 | 26.6
25.9 | 19.7 | 22.96 | 25 | 26.5
26 | 19.55 | 22.8 | 24.8 | 26.325
26.1 | 19.4 | 22.66 | 24.7 | 26.2
26.2 | 19.25 | 22.46 | 24.52 | 26.075
26.3 | 19.1 | 22.36 | 24.42 | 25.925
26.4 | 19 | 22.16 | 24.22 | 25.78
26.5 | 18.925 | 22.06 | 24.125 | 25.58
26.6 | 18.825 | 21.9667 | 23.92 | 25.48
26.7 | 18.725 | 21.78 | 23.8 | 25.38
26.8 | 18.6 | 21.6667 | 23.7333 | 25.2
26.9 | 18.475 | 21.5 | 23.6 | 25.08
27 | 18.3 | 21.4 | 23.4 | 24.9
27.1 | 18.2 | 21.3143 | 23.24 | 24.8
27.2 | 18.1 | 21.15 | 23.14 | 24.7
27.3 | 18 | 21.0455 | 23 | 24.5
27.4 | 17.9 | 20.9286 | 22.85 | 24.3
27.5 | 17.8 | 20.8 | 22.7 | 24.2
27.6 | 17.6 | 20.6 | 22.6 | 24.1
27.7 | 17.5 | 20.5 | 22.4375 | 24
27.8 | 17.5 | 20.4 | 22.35 | 23.8
27.9 | 17.3167 | 20.3 | 22.2714 | 23.7
28 | 17.2 | 20.24 | 22.1167 | 23.6
28.1 | 17.02 | 20.04 | 22 | 23.4
28.2 | 17 | 19.94 | 21.8 | 23.3
28.3 | 16.9 | 19.84 | 21.7 | 23.175
28.4 | 16.72 | 19.64 | 21.6 | 23.0667
28.5 | 16.62 | 19.54 | 21.4 | 22.9
28.6 | 16.52 | 19.4 | 21.3 | 22.7667
28.7 | 16.42 | 19.34 | 21.2 | 22.6
28.8 | 16.32 | 19.2 | 21.1091 | 22.6
28.9 | 16.2333 | 19.1 | 20.9 | 22.4
29 | 16.1333 | 18.96 | 20.9 | 22.28
29.1 | 16.04 | 18.86 | 20.72 | 22.18
29.2 | 15.9 | 18.75 | 20.61 | 22.0667
29.3 | 15.8 | 18.7 | 20.51 | 21.9
29.4 | 15.7 | 18.5 | 20.32 | 21.8
29.5 | 15.6 | 18.4 | 20.21 | 21.64
29.6 | 15.5 | 18.3167 | 20.11 | 21.54
29.7 | 15.4 | 18.22 | 20.01 | 21.4667
29.8 | 15.3 | 18.1 | 19.9182 | 21.3
29.9 | 15.2 | 17.92 | 19.74 | 21.2
30 | 15.2 | 17.85 | 19.62 | 21.1

## SSD120 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 49.9 | 55.6 | 59.4 | 64.5
0.1 | 56.1 | 61.08 | 65.08 | 69.5
0.2 | 65.5 | 70.25 | 73.56 | 76.88
0.3 | 76.16 | 79.76 | 82.14 | 84.34
0.4 | 84.34 | 86.9 | 88.08 | 90.26
0.5 | 90.3333 | 91.8 | 92.36 | 93.95
0.6 | 94.1714 | 95.1 | 95.78 | 96.45
0.7 | 96.62 | 97.2 | 97.56 | 98.16
0.8 | 98.2143 | 98.5 | 99 | 99.22
0.9 | 99.15 | 99.5 | 99.78 | 99.68
1 | 99.7429 | 99.9 | 99.94 | 99.94
1.1 | 100 | 100 | 100 | 100
1.2 | 100 | 100 | 99.96 | 100
1.3 | 99.91 | 99.82 | 99.8667 | 99.88
1.4 | 99.7364 | 99.72 | 99.68 | 99.66
1.5 | 99.38 | 99.52 | 99.44 | 99.34
1.6 | 98.98 | 99.24 | 99.075 | 99.04
1.7 | 98.61 | 98.76 | 98.66 | 98.84
1.8 | 98.28 | 98.42 | 98.38 | 98.52
1.9 | 97.8875 | 98.1667 | 98.1333 | 98.12
2 | 97.4 | 97.78 | 97.88 | 97.7556
2.1 | 96.95 | 97.3875 | 97.64 | 97.45
2.2 | 96.54 | 97 | 97.16 | 97.09
2.3 | 95.92 | 96.65 | 96.66 | 96.7375
2.4 | 95.56 | 96.1 | 96.38 | 96.42
2.5 | 95.1333 | 95.7 | 96.26 | 96.22
2.6 | 94.575 | 95.5 | 95.86 | 95.92
2.7 | 94.26 | 95.2 | 95.38 | 95.4286
2.8 | 93.84 | 94.6 | 95 | 95.1
2.9 | 93.32 | 94.34 | 94.6667 | 94.96
3 | 92.7667 | 94.08 | 94.4 | 94.56
3.1 | 92.425 | 93.56 | 94 | 94.1714
3.2 | 92.06 | 93.3 | 93.7 | 93.8
3.3 | 91.5 | 92.82 | 93.3 | 93.3714
3.4 | 91.12 | 92.42 | 93.0333 | 92.98
3.5 | 90.38 | 92.09 | 92.5 | 92.77
3.6 | 90 | 91.7455 | 92.2 | 92.4667
3.7 | 89.6 | 91.3429 | 91.8 | 92.2
3.8 | 89.12 | 91.05 | 91.5333 | 91.84
3.9 | 88.66 | 90.46 | 91.1 | 91.5625
4 | 88.28 | 90.15 | 90.92 | 91.12
4.1 | 87.8333 | 89.66 | 90.56 | 90.7625
4.2 | 87.22 | 89.3 | 90.125 | 90.35
4.3 | 86.68 | 88.9857 | 89.725 | 90
4.4 | 86.44 | 88.66 | 89.325 | 89.6
4.5 | 85.86 | 88.3 | 88.94 | 89.2
4.6 | 85.48 | 87.72 | 88.72 | 88.9
4.7 | 84.94 | 87.32 | 88.34 | 88.64
4.8 | 84.5667 | 86.92 | 88.04 | 88.3333
4.9 | 84.2 | 86.62 | 87.54 | 88.02
5 | 83.65 | 86.22 | 87.3 | 87.74
5.1 | 83.2 | 85.96 | 86.95 | 87.4
5.2 | 82.75 | 85.56 | 86.52 | 87.04
5.3 | 82.1667 | 85.28 | 86.0667 | 86.72
5.4 | 81.68 | 85.02 | 85.82 | 86.44
5.5 | 81.34 | 84.36 | 85.425 | 86.1
5.6 | 80.86 | 83.86 | 85 | 85.74
5.7 | 80.4429 | 83.66 | 84.825 | 85.4
5.8 | 79.9 | 83.24 | 84.42 | 84.92
5.9 | 79.4 | 82.86 | 83.88 | 84.7
6 | 78.9 | 82.5833 | 83.58 | 84.3667
6.1 | 78.4 | 82.3 | 83.34 | 84.05
6.2 | 77.9667 | 81.8833 | 83.14 | 83.55
6.3 | 77.68 | 81.6 | 82.82 | 83.2
6.4 | 77.22 | 81.3 | 82.44 | 83
6.5 | 76.81 | 80.7 | 82.1 | 82.6
6.6 | 76.48 | 80.4 | 81.68 | 82.2
6.7 | 76.08 | 79.9 | 81.32 | 81.975
6.8 | 75.65 | 79.5 | 81.02 | 81.65
6.9 | 75.0667 | 79.2 | 80.66 | 81.3
7 | 74.7 | 78.9 | 80.42 | 81.0333
7.1 | 74.35 | 78.65 | 80.08 | 80.8
7.2 | 73.92 | 78.22 | 79.62 | 80.45
7.3 | 73.44 | 77.8 | 79.28 | 80.0333
7.4 | 73.02 | 77.36 | 78.92 | 79.6333
7.5 | 72.6 | 76.96 | 78.5333 | 79.3
7.6 | 72.18 | 76.58 | 78.3 | 79.1
7.7 | 71.82 | 76.3333 | 77.9667 | 78.8667
7.8 | 71.16 | 75.84 | 77.5667 | 78.5
7.9 | 70.86 | 75.625 | 77.1667 | 78.14
8 | 70.3833 | 75.0833 | 76.84 | 77.8
8.1 | 70.06 | 74.8 | 76.6333 | 77.5667
8.2 | 69.7 | 74.4 | 76.325 | 77
8.3 | 69.35 | 74 | 76.0333 | 76.6667
8.4 | 68.88 | 73.8 | 75.5333 | 76.475
8.5 | 68.5 | 73.4 | 75.34 | 76.16
8.6 | 68.1 | 73.0857 | 75.02 | 75.8
8.7 | 67.78 | 72.72 | 74.58 | 75.4333
8.8 | 67.3 | 72.3375 | 74.18 | 75.22
8.9 | 66.9333 | 72.02 | 73.94 | 75.0333
9 | 66.52 | 71.56 | 73.64 | 74.55
9.1 | 66.08 | 71.26 | 73.28 | 74.2
9.2 | 65.8 | 70.9 | 72.88 | 74
9.3 | 65.28 | 70.48 | 72.6 | 73.7
9.4 | 64.94 | 70.1857 | 72.05 | 73.4
9.5 | 64.44 | 69.94 | 71.85 | 73
9.6 | 64.04 | 69.5429 | 71.7143 | 72.6
9.7 | 63.7 | 69.1 | 71.3 | 72.2667
9.8 | 63.55 | 68.775 | 71 | 72.08
9.9 | 63 | 68.4143 | 70.6 | 71.76
10 | 62.65 | 68.175 | 70.4167 | 71.48
10.1 | 62.3333 | 67.8714 | 70.1 | 71.25
10.2 | 61.96 | 67.6 | 69.7667 | 70.9571
10.3 | 61.58 | 67.16 | 69.5667 | 70.675
10.4 | 61.34 | 66.76 | 69.1333 | 70.28
10.5 | 60.78 | 66.36 | 68.7667 | 69.94
10.6 | 60.55 | 66.0875 | 68.4667 | 69.58
10.7 | 60.1625 | 65.82 | 68.1333 | 69.18
10.8 | 59.78 | 65.4143 | 67.7667 | 68.94
10.9 | 59.4571 | 65.16 | 67.4667 | 68.68
11 | 59.1 | 64.8 | 67.22 | 68.44
11.1 | 58.7 | 64.475 | 66.925 | 68.08
11.2 | 58.275 | 64.18 | 66.6333 | 67.74
11.3 | 58.04 | 63.7667 | 66.3 | 67.5143
11.4 | 57.7667 | 63.5 | 66.0333 | 67.25
11.5 | 57.3 | 63.2 | 65.72 | 66.92
11.6 | 57 | 62.8429 | 65.38 | 66.72
11.7 | 56.6875 | 62.55 | 65.16 | 66.42
11.8 | 56.4 | 62.25 | 64.76 | 66.05
11.9 | 55.9 | 61.95 | 64.38 | 65.6
12 | 55.675 | 61.65 | 64.18 | 65.36
12.1 | 55.32 | 61.375 | 63.88 | 65.2333
12.2 | 54.96 | 61.1 | 63.56 | 64.86
12.3 | 54.65 | 60.62 | 63.28 | 64.575
12.4 | 54.32 | 60.2429 | 63.06 | 64.275
12.5 | 53.96 | 60.05 | 62.68 | 64.075
12.6 | 53.76 | 59.8 | 62.46 | 63.775
12.7 | 53.35 | 59.475 | 62.06 | 63.5
12.8 | 52.9333 | 59.08 | 61.78 | 63.3
12.9 | 52.875 | 58.775 | 61.56 | 63
13 | 52.52 | 58.4667 | 61.26 | 62.6
13.1 | 52.175 | 58.26 | 60.96 | 62.36
13.2 | 51.82 | 57.98 | 60.6 | 62.1
13.3 | 51.575 | 57.66 | 60.26 | 61.86
13.4 | 51.14 | 57.2714 | 60 | 61.66
13.5 | 50.92 | 57.02 | 59.7833 | 61.26
13.6 | 50.62 | 56.78 | 59.6 | 60.92
13.7 | 50.24 | 56.4667 | 59.3 | 60.675
13.8 | 49.9667 | 56.24 | 58.9667 | 60.475
13.9 | 49.54 | 55.88 | 58.74 | 60.08
14 | 49.32 | 55.48 | 58.38 | 59.88
14.1 | 48.96 | 55.24 | 58.14 | 59.68
14.2 | 48.7 | 54.98 | 57.84 | 59.36
14.3 | 48.4 | 54.62 | 57.44 | 59.08
14.4 | 48.0667 | 54.36 | 57.24 | 58.6667
14.5 | 47.725 | 54.06 | 56.94 | 58.44
14.6 | 47.6 | 53.76 | 56.74 | 58.2333
14.7 | 47.3333 | 53.46 | 56.46 | 57.98
14.8 | 46.92 | 53.25 | 56.12 | 57.68
14.9 | 46.6333 | 52.92 | 55.82 | 57.38
15 | 46.475 | 52.6333 | 55.56 | 57.16
15.1 | 46.16 | 52.475 | 55.32 | 56.85
15.2 | 45.86 | 52.22 | 55.16 | 56.58
15.3 | 45.4857 | 51.84 | 54.86 | 56.34
15.4 | 45.2 | 51.68 | 54.56 | 56.08
15.5 | 44.98 | 51.34 | 54.36 | 55.84
15.6 | 44.78 | 51.05 | 53.92 | 55.64
15.7 | 44.56 | 50.76 | 53.7 | 55.38
15.8 | 44.1667 | 50.48 | 53.46 | 55.18
15.9 | 44.04 | 50.28 | 53.2 | 54.84
16 | 43.78 | 50.08 | 52.9 | 54.58
16.1 | 43.48 | 49.66 | 52.8 | 54.3
16.2 | 43.26 | 49.4 | 52.6 | 54.2
16.3 | 43.04 | 49.18 | 52.3 | 53.9
16.4 | 42.7143 | 48.8 | 51.95 | 53.6
16.5 | 42.4 | 48.58 | 51.76 | 53.3
16.6 | 42.1 | 48.38 | 51.55 | 53.1
16.7 | 41.8 | 48.2 | 51.32 | 52.9
16.8 | 41.6 | 47.9 | 51.0667 | 52.625
16.9 | 41.46 | 47.58 | 50.74 | 52.4
17 | 41.2 | 47.4 | 50.38 | 52.175
17.1 | 41.1 | 47.18 | 50.2 | 51.95
17.2 | 40.825 | 46.9 | 49.96 | 51.6333
17.3 | 40.425 | 46.6 | 49.68 | 51.375
17.4 | 40.2667 | 46.3 | 49.48 | 51.1333
17.5 | 39.98 | 46.2 | 49.28 | 50.9
17.6 | 39.7667 | 45.9 | 49.06 | 50.66
17.7 | 39.5833 | 45.6 | 48.78 | 50.5
17.8 | 39.4 | 45.3 | 48.58 | 50.32
17.9 | 39 | 45.1 | 48.38 | 50.02
18 | 38.825 | 44.9 | 48.08 | 49.8
18.1 | 38.6 | 44.8 | 47.88 | 49.6333
18.2 | 38.4 | 44.4 | 47.58 | 49.32
18.3 | 38.16 | 44.3 | 47.38 | 49
18.4 | 37.92 | 43.94 | 46.9875 | 48.76
18.5 | 37.62 | 43.7 | 46.84 | 48.6333
18.6 | 37.4 | 43.4286 | 46.64 | 48.3667
18.7 | 37.2 | 43.22 | 46.44 | 48.175
18.8 | 36.86 | 43.0429 | 46.15 | 47.98
18.9 | 36.68 | 42.8167 | 45.94 | 47.7333
19 | 36.5 | 42.6 | 45.8 | 47.42
19.1 | 36.38 | 42.3667 | 45.6 | 47.225
19.2 | 36.1286 | 42.16 | 45.3 | 47.04
19.3 | 35.8889 | 41.97 | 45.1 | 46.825
19.4 | 35.68 | 41.68 | 44.88 | 46.68
19.5 | 35.54 | 41.48 | 44.575 | 46.34
19.6 | 35.32 | 41.27 | 44.38 | 46.125
19.7 | 35.08 | 40.97 | 44.18 | 46
19.8 | 34.7875 | 40.69 | 44 | 45.78
19.9 | 34.6143 | 40.57 | 43.74 | 45.38
20 | 34.425 | 40.29 | 43.5 | 45.14
20.1 | 34.24 | 40.18 | 43.28 | 45
20.2 | 33.9833 | 39.97 | 43.06 | 44.78
20.3 | 33.9 | 39.6818 | 42.8 | 44.54
20.4 | 33.6 | 39.5 | 42.64 | 44.4
20.5 | 33.3571 | 39.2818 | 42.44 | 44.2
20.6 | 33.2 | 39.1 | 42.24 | 44.04
20.7 | 32.94 | 38.8 | 42.04 | 43.86
20.8 | 32.8 | 38.56 | 41.8 | 43.66
20.9 | 32.64 | 38.36 | 41.6 | 43.38
21 | 32.45 | 38.2 | 41.375 | 43.2
21.1 | 32.22 | 38 | 41.125 | 43.02
21.2 | 31.96 | 37.8 | 40.9667 | 42.7
21.3 | 31.8 | 37.6 | 40.725 | 42.6
21.4 | 31.66 | 37.42 | 40.5333 | 42.4
21.5 | 31.45 | 37.1111 | 40.3333 | 42.1
21.6 | 31.3 | 36.95 | 40.1333 | 42
21.7 | 31.1 | 36.775 | 39.85 | 41.82
21.8 | 30.9 | 36.6 | 39.7 | 41.6
21.9 | 30.775 | 36.46 | 39.44 | 41.42
22 | 30.6 | 36.2333 | 39.4 | 41.22
22.1 | 30.4 | 35.98 | 39.2 | 41.02
22.2 | 30.2333 | 35.78 | 38.9 | 40.82
22.3 | 30.075 | 35.5667 | 38.8 | 40.62
22.4 | 29.9 | 35.38 | 38.6 | 40.42
22.5 | 29.725 | 35.2143 | 38.4 | 40.22
22.6 | 29.5333 | 35.1 | 38.2 | 40.02
22.7 | 29.3333 | 34.9 | 38 | 39.84
22.8 | 29.16 | 34.72 | 37.8 | 39.62
22.9 | 28.98 | 34.46 | 37.54 | 39.5
23 | 28.8333 | 34.32 | 37.4 | 39.22
23.1 | 28.7 | 34.21 | 37.3 | 39.12
23.2 | 28.55 | 33.93 | 37.0667 | 38.92
23.3 | 28.35 | 33.72 | 36.8 | 38.72
23.4 | 28.1 | 33.5571 | 36.6 | 38.5
23.5 | 27.94 | 33.3286 | 36.4 | 38.32
23.6 | 27.84 | 33.12 | 36.2667 | 38.12
23.7 | 27.5667 | 33 | 36.075 | 37.92
23.8 | 27.5 | 32.84 | 35.875 | 37.82
23.9 | 27.3 | 32.74 | 35.7 | 37.6111
24 | 27.1 | 32.54 | 35.6 | 37.4167
24.1 | 27 | 32.3333 | 35.35 | 37.25
24.2 | 26.84 | 32.16 | 35.1333 | 37.1143
24.3 | 26.7 | 32 | 35.0333 | 36.92
24.4 | 26.56 | 31.8 | 34.8 | 36.81
24.5 | 26.3 | 31.5 | 34.6 | 36.62
24.6 | 26.1571 | 31.4 | 34.46 | 36.33
24.7 | 26.075 | 31.3 | 34.36 | 36.16
24.8 | 25.9 | 31.1 | 34.1 | 36
24.9 | 25.76 | 31 | 34 | 35.86
25 | 25.56 | 30.8333 | 33.8 | 35.675
25.1 | 25.4333 | 30.6333 | 33.6 | 35.55
25.2 | 25.36 | 30.475 | 33.5 | 35.36
25.3 | 25.2 | 30.275 | 33.2 | 35.16
25.4 | 25 | 30.1 | 33.1 | 35.02
25.5 | 24.86 | 29.975 | 33 | 34.8
25.6 | 24.76 | 29.775 | 32.8333 | 34.62
25.7 | 24.5333 | 29.6 | 32.6333 | 34.46
25.8 | 24.5 | 29.5 | 32.5 | 34.26
25.9 | 24.32 | 29.3333 | 32.3 | 34.16
26 | 24.12 | 29.125 | 32.1 | 33.96
26.1 | 24 | 29.04 | 31.925 | 33.86
26.2 | 23.9 | 28.8 | 31.825 | 33.76
26.3 | 23.8 | 28.64 | 31.6 | 33.46
26.4 | 23.66 | 28.5 | 31.4333 | 33.26
26.5 | 23.46 | 28.44 | 31.2 | 33.18
26.6 | 23.3 | 28.2 | 31.1 | 32.98
26.7 | 23.16 | 28.0667 | 31 | 32.88
26.8 | 23.06 | 27.9 | 30.8 | 32.74
26.9 | 22.9 | 27.74 | 30.7 | 32.52
27 | 22.8 | 27.6 | 30.5 | 32.44
27.1 | 22.65 | 27.5 | 30.3 | 32.24
27.2 | 22.6 | 27.3 | 30.2 | 32.04
27.3 | 22.425 | 27.2 | 30 | 31.94
27.4 | 22.225 | 27 | 29.9 | 31.74
27.5 | 22.025 | 26.9 | 29.7 | 31.54
27.6 | 21.98 | 26.8 | 29.6 | 31.44
27.7 | 21.85 | 26.6667 | 29.4667 | 31.24
27.8 | 21.74 | 26.4 | 29.2 | 31.1
27.9 | 21.6 | 26.3 | 29.1 | 30.94
28 | 21.44 | 26.125 | 29 | 30.8
28.1 | 21.3 | 26 | 28.8667 | 30.7
28.2 | 21.2 | 25.925 | 28.7 | 30.54
28.3 | 21.1 | 25.8 | 28.58 | 30.34
28.4 | 21 | 25.6667 | 28.5 | 30.2
28.5 | 20.8667 | 25.48 | 28.3 | 30.1
28.6 | 20.725 | 25.3 | 28.1 | 30
28.7 | 20.625 | 25.18 | 27.9 | 29.825
28.8 | 20.5 | 25.08 | 27.8 | 29.7
28.9 | 20.4 | 24.98 | 27.7 | 29.5
29 | 20.3 | 24.8 | 27.5333 | 29.4
29.1 | 20.2 | 24.7 | 27.4 | 29.2
29.2 | 20.1 | 24.48 | 27.2333 | 29.1
29.3 | 19.95 | 24.3333 | 27.1 | 29
29.4 | 19.8 | 24.2 | 27 | 28.9
29.5 | 19.7 | 24.1 | 26.8429 | 28.6444
29.6 | 19.55 | 24 | 26.7857 | 28.4667
29.7 | 19.4 | 23.9 | 26.54 | 28.3857
29.8 | 19.3 | 23.8 | 26.3875 | 28.22
29.9 | 19.2 | 23.575 | 26.3 | 28.1
30 | 19.1 | 23.45 | 26.15 | 27.95

## Wedge 60 Lateral Profiles
Lateral Profiles
Wedge 60
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:1.2 cm FieldSize:4x4 cm2 | Depth:1.2 cm FieldSize:10x10 cm2 | Depth:1.2 cm FieldSize:15x15 cm2 | Depth:1.2 cm FieldSize:40x15 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:15x15 cm2 | Depth:5 cm FieldSize:40x15 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:15x15 cm2 | Depth:10 cm FieldSize:40x15 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:15x15 cm2 | Depth:20 cm FieldSize:40x15 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:15x15 cm2 | Depth:30 cm FieldSize:40x15 cm2
-14.7 | [18 empty] | 17.52 | 24.4895
-14.6 | [18 empty] | 17.85 | 24.5894
-14.5 | [18 empty] | 18.0333 | 24.7894
-14.4 | [18 empty] | 18.3 | 25.0559
-14.3 | [18 empty] | 18.6 | 25.3225
-14.2 | [18 empty] | 18.9 | 25.589
-14.1 | [18 empty] | 19.1 | 25.9289
-14 | [18 empty] | 19.4143 | 26.2387
-13.9 | [14 empty] | 13.9056 | 19.72 |  |  | 19.65 | 26.5553
-13.8 | [14 empty] | 14.0723 | 19.95 |  |  | 19.8 | 26.7885
-13.7 | [14 empty] | 14.3457 | 20.2 |  |  | 20.2 | 26.9599
-13.6 | [14 empty] | 14.5725 | 20.425 |  |  | 20.38 | 27.3383
-13.5 | [14 empty] | 14.8059 | 20.7 |  |  | 20.7 | 27.6881
-13.4 | [14 empty] | 15.0727 | 20.95 |  |  | 20.98 | 27.8381
-13.3 | [14 empty] | 15.3311 | 21.16 |  |  | 21.35 | 28.1879
-13.2 | [10 empty] | 9.5 | 14.4 |  |  | 15.6062 | 21.45 |  |  | 21.825 | 28.4128
-13.1 | [10 empty] | 9.6 | 14.5 |  |  | 15.8063 | 21.82 |  |  | 21.975 | 28.8476
-13 | [10 empty] | 9.68 | 14.7 |  |  | 16.0564 | 22.1 |  |  | 22.425 | 29.3374
-12.9 | [10 empty] | 9.975 | 14.95 |  |  | 16.2565 | 22.42 |  |  | 22.9 | 29.6873
-12.8 | [6 empty] | 7.325 | 11.9 |  |  | 10.1 | 15.1 |  |  | 16.4816 | 22.65 |  |  | 23.2 | 30.1246
-12.7 | [6 empty] | 7.5 | 12 |  |  | 10.26 | 15.2571 |  |  | 16.9401 | 22.9 |  |  | 23.4 | 30.5069
-12.6 | [6 empty] | 7.6143 | 12.2 |  |  | 10.4 | 15.475 |  |  | 17.2069 | 23.2 |  |  | 23.74 | 30.8618
-12.5 |  |  | 6.4857 | 10.775 |  |  | 7.75 | 12.24 |  |  | 10.6 | 15.6 |  |  | 17.607 | 23.56 |  |  | 24.2 | 31.2723
-12.4 |  |  | 6.5667 | 10.9 |  |  | 7.88 | 12.425 |  |  | 10.825 | 15.7833 |  |  | 17.8672 | 23.7714 |  |  | 24.75 | 31.5865
-12.3 |  |  | 6.6 | 10.9857 |  |  | 8 | 12.54 |  |  | 11 | 16.1 |  |  | 18.0072 | 24.125 |  |  | 25.1 | 32.0149
-12.2 |  |  | 6.7 | 11.125 |  |  | 8.125 | 12.775 |  |  | 11.1833 | 16.2833 |  |  | 18.4074 | 24.5 |  |  | 25.55 | 32.3576
-12.1 |  |  | 6.8 | 11.2 |  |  | 8.2 | 12.9 |  |  | 11.44 | 16.48 |  |  | 18.7075 | 24.8 |  |  | 25.9 | 32.8059
-12 |  |  | 6.9 | 11.325 |  |  | 8.325 | 13.1 |  |  | 11.7 | 16.65 |  |  | 19.1826 | 25.2667 |  |  | 26.35 | 33.2857
-11.9 |  |  | 7 | 11.4 |  |  | 8.55 | 13.1 |  |  | 11.975 | 16.925 |  |  | 19.6412 | 25.6667 |  |  | 26.9333 | 33.6856
-11.8 |  |  | 7.04 | 11.5714 |  |  | 8.625 | 13.4 |  |  | 12.125 | 17.28 |  |  | 19.8479 | 26 |  |  | 27.25 | 34.152
-11.7 |  |  | 7.2 | 11.7 |  |  | 8.8333 | 13.5143 |  |  | 12.3 | 17.5667 |  |  | 20.1795 | 26.4286 |  |  | 27.85 | 34.6185
-11.6 |  |  | 7.3 | 11.8 |  |  | 8.925 | 13.7143 |  |  | 12.55 | 17.85 |  |  | 20.7283 | 26.7143 |  |  | 28.225 | 35.01
-11.5 |  |  | 7.4 | 11.875 |  |  | 9.175 | 13.96 |  |  | 12.9625 | 18.1 |  |  | 21.1656 | 27.15 |  |  | 28.6 | 35.5347
-11.4 |  |  | 7.5 | 12.0571 |  |  | 9.325 | 14.1571 |  |  | 13.18 | 18.4 |  |  | 21.5086 | 27.7 |  | 11.5 | 29.05 | 35.8846
-11.3 |  |  | 7.6 | 12.2 |  |  | 9.575 | 14.3 |  |  | 13.42 | 18.6 |  |  | 21.9088 | 28.1 |  | 11.7 | 29.6 | 36.4344
-11.2 |  |  | 7.7833 | 12.325 |  |  | 9.7 | 14.48 |  |  | 13.7 | 18.9333 |  |  | 22.2088 | 28.44 |  | 11.96 | 30.25 | 36.9841
-11.1 |  |  | 7.9 | 12.5 |  |  | 9.85 | 14.6667 |  |  | 14 | 19.2857 |  |  | 22.9092 | 28.9 |  | 12.1 | 30.7667 | 37.384
-11 |  |  | 8 | 12.65 |  |  | 10.1 | 14.85 |  |  | 14.36 | 19.6 |  |  | 23.3593 | 29.4 |  | 12.2333 | 31.4 | 37.9838
-10.9 |  |  | 8.1 | 12.8 |  |  | 10.3143 | 15.075 |  |  | 14.6 | 19.8333 |  | 8.76 | 23.8095 | 29.8 |  | 12.42 | 31.9 | 38.5235
-10.8 |  |  | 8.2 | 12.95 |  |  | 10.5 | 15.3 |  |  | 15.1 | 20.2 |  | 9 | 24.2297 | 30.3 |  | 12.6 | 32.5 | 39.0975
-10.7 |  |  | 8.4 | 13.1 |  |  | 10.725 | 15.56 |  |  | 15.45 | 20.525 |  | 9 | 24.7699 | 30.8 |  | 12.7 | 33.3 | 39.8429
-10.6 |  |  | 8.6 | 13.25 |  |  | 10.9 | 15.8333 |  |  | 15.8 | 21 |  | 9.1 | 25.16 | 31.3 |  | 12.925 | 34.15 | 40.6112
-10.5 |  |  | 8.65 | 13.4333 |  |  | 11.2286 | 16 |  |  | 16.15 | 21.42 |  | 9.25 | 25.7503 | 31.85 |  | 13.2 | 35.2667 | 41.5822
-10.4 |  |  | 8.775 | 13.6 |  |  | 11.48 | 16.325 |  | 5.6 | 16.44 | 21.7 |  | 9.4 | 26.3706 | 32.4 |  | 13.4 | 36.5 | 42.8442
-10.3 |  |  | 9 | 13.7333 |  |  | 11.7 | 16.575 |  | 5.725 | 16.85 | 22.15 |  | 9.6 | 26.9108 | 32.95 |  | 13.625 | 38.6 | 44.6809
-10.2 |  |  | 9.22 | 13.9714 |  | 4.3 | 11.9 | 16.925 |  | 5.8 | 17.45 | 22.6 |  | 9.8 | 27.471 | 33.42 |  | 13.8 | 41.4 | 47.2422
-10.1 |  |  | 9.3 | 14.125 |  | 4.3 | 12.25 | 17.25 |  | 5.9333 | 17.9 | 23.08 |  | 10 | 28.2613 | 34.14 |  | 13.95 | 46.4 | 51.4779
-10 |  | 3.7 | 9.55 | 14.3 |  | 4.4 | 12.6 | 17.5667 |  | 6 | 18.3333 | 23.52 |  | 10.2 | 29.0116 | 34.8 |  | 14.42 | 53.7 | 57.6586
-9.9 |  | 3.7 | 9.675 | 14.525 |  | 4.5 | 12.9 | 17.8333 |  | 6.2 | 18.8 | 23.975 |  | 10.35 | 29.7119 | 35.6625 |  | 14.7 | 64.15 | 66.6214
-9.8 |  | 3.8 | 9.9 | 14.7 |  | 4.5 | 13.25 | 18.175 |  | 6.3 | 19.2 | 24.45 |  | 10.5 | 30.6456 | 36.5 |  | 15.1 | 76.6 | 77.7667
-9.7 |  | 3.8 | 10.05 | 14.95 |  | 4.6667 | 13.6 | 18.5667 |  | 6.3 | 19.8 | 24.95 |  | 10.7 | 32.0128 | 37.6 |  | 15.2 | 89.175 | 89.8115
-9.6 |  | 3.9 | 10.28 | 15.1667 |  | 4.7 | 14.06 | 18.98 |  | 6.4 | 20.35 | 25.6 |  | 10.95 | 33.6468 | 39.4 |  | 15.38 | 103.45 | 102.3314
-9.5 |  | 4 | 10.54 | 15.4 |  | 4.775 | 14.4429 | 19.35 |  | 6.6 | 21.02 | 26.08 |  | 11.175 | 36.4146 | 41.6 |  | 15.625 | 116.275 | 113.3515
-9.4 |  | 4 | 10.76 | 15.58 |  | 4.8 | 14.8667 | 19.725 |  | 6.8 | 21.5875 | 26.72 |  | 11.3 | 40.5662 | 45.725 |  | 15.95 | 127.275 | 122.7225
-9.3 |  | 4.1 | 11.025 | 15.86 |  | 4.9 | 15.3 | 20.14 |  | 6.9 | 22.15 | 27.325 |  | 11.425 | 47.519 | 52.04 |  | 16.175 | 135.125 | 129.6197
-9.2 |  | 4.1 | 11.28 | 16.16 |  | 5 | 15.7 | 20.58 |  | 7.02 | 23 | 28.0333 |  | 11.66 | 57.583 | 61.975 |  | 16.55 | 140.925 | 134.0673
-9.1 |  | 4.2 | 11.6 | 16.4 |  | 5.1 | 16.25 | 21.1 |  | 7.2 | 24 | 28.98 |  | 12.0714 | 71.3285 | 73.825 |  | 16.9 | 143.475 | 136.6167
-9 |  | 4.2 | 11.85 | 16.7 |  | 5.1 | 16.8 | 21.6 |  | 7.3 | 25.06 | 30.075 |  | 12.4 | 86.6347 | 87.2 |  | 17.1333 | 144.6 | 137.816
-8.9 |  | 4.3 | 12.1 | 17.025 |  | 5.25 | 17.375 | 22.2 |  | 7.5 | 26.525 | 31.46 |  | 12.725 | 103.0128 | 101.9 |  | 17.5 | 145.15 | 138.441
-8.8 |  | 4.3 | 12.45 | 17.35 |  | 5.3 | 18.05 | 22.8857 |  | 7.6 | 28.95 | 33.7 |  | 12.975 | 118.747 | 116.6429 |  | 17.7 | 145.46 | 138.441
-8.7 |  | 4.4 | 12.76 | 17.76 |  | 5.5 | 18.9 | 23.7 |  | 7.8 | 32.7333 | 37.36 |  | 13.1333 | 132.753 | 128.4 |  | 18.1 | 145.5 | 138.541
-8.6 |  | 4.5 | 13.1429 | 18.05 |  | 5.6 | 20.1 | 24.9143 |  | 8 | 39.2 | 43.5429 |  | 13.42 | 142.9372 | 138.34 |  | 18.6333 | 145.44 | 138.516
-8.5 |  | 4.5 | 13.64 | 18.64 |  | 5.7 | 21.9 | 26.575 |  | 8.2 | 50.3429 | 53.74 |  | 13.7 | 149.7095 | 143.8 |  | 18.8 | 145.2 | 138.281
-8.4 |  | 4.5333 | 14.325 | 19.2429 |  | 5.8 | 24.7 | 29.05 |  | 8.4 | 66.16 | 67.9333 |  | 14 | 152.841 | 146.8 |  | 19.2 | 145 | 138.1124
-8.3 |  | 4.7 | 15.3667 | 20.22 |  | 5.9 | 29.45 | 33.4 |  | 8.5 | 84.4625 | 86.08 |  | 14.2667 | 154.362 | 147.96 |  | 19.55 | 144.65 | 138.041
-8.2 |  | 4.7 | 16.925 | 21.675 |  | 6 | 37.7 | 41.9429 |  | 8.9 | 104.8667 | 104.9 |  | 14.68 | 154.762 | 148.42 |  | 19.775 | 144.3 | 137.791
-8.1 |  | 4.8 | 19.8 | 24.35 |  | 6.2 | 52.55 | 55.8 |  | 9 | 124.68 | 121.7 |  | 15.04 | 154.762 | 148.6 |  | 20.26 | 143.96 | 137.581
-8 |  | 4.9 | 25.2 | 29.66 |  | 6.35 | 72.3 | 74.15 |  | 9.4 | 141.72 | 137.7 |  | 15.3667 | 154.762 | 148.6 |  | 20.7333 | 143.55 | 137.291
-7.9 |  | 5 | 35.925 | 40.75 |  | 6.5 | 94.24 | 95 |  | 9.6 | 155.58 | 151.4 |  | 15.7 | 154.612 | 148.5 |  | 21.4 | 143.0333 | 136.941
-7.8 |  | 5.1 | 53.6 | 57.56 |  | 6.65 | 117.525 | 115.9625 |  | 9.8 | 164.36 | 158.9889 |  | 15.95 | 154.162 | 148.15 |  | 21.7 | 142.8 | 136.666
-7.7 |  | 5.2 | 76.8 | 77.94 |  | 6.8 | 139.22 | 136.3 |  | 10.0667 | 168.66 | 162.7 |  | 16.36 | 153.7115 | 147.65 |  | 22 | 142.5125 | 136.402
-7.6 |  | 5.3 | 102.1 | 102.45 |  | 7.05 | 157.1 | 153.4833 |  | 10.35 | 169.6222 | 164.05 |  | 16.85 | 153.061 | 147.25 |  | 22.58 | 141.9 | 135.992
-7.5 |  | 5.4 | 127.86 | 126.04 |  | 7.25 | 171.3 | 166 |  | 10.6667 | 169.7 | 164.2 |  | 17.26 | 152.661 | 146.9 | 4.45 | 23.1 | 141.44 | 135.492
-7.4 |  | 5.55 | 152.2 | 149.3 |  | 7.4333 | 178.1571 | 172.8 |  | 11 | 169.7 | 164.2 | 3 | 17.65 | 151.661 | 146.3 | 4.6 | 23.76 | 141.1667 | 135.0795
-7.3 |  | 5.7 | 171.575 | 166.8 |  | 7.6333 | 180.8 | 175.3 |  | 11.3667 | 169.425 | 163.9 | 3.1 | 18.22 | 151.261 | 145.75 | 4.7 | 24.6 | 140.5667 | 134.792
-7.2 |  | 5.8 | 183.85 | 179.075 |  | 7.9 | 180.9 | 175.5 | 1.9 | 11.7333 | 168.925 | 163.26 | 3.2 | 18.7 | 150.66 | 145.3 | 4.7 | 25.45 | 140.0333 | 134.542
-7.1 |  | 5.9 | 190.12 | 184.6 | 1.5 | 8 | 180.9 | 175.5 | 2 | 11.975 | 167.85 | 162.7 | 3.2 | 19.2333 | 150.135 | 144.55 | 4.7 | 27.05 | 139.66 | 134.043
-7 | 1.1994 | 6.1 | 190.7 | 185.1 | 1.5 | 8.4 | 180.55 | 175.4 | 2 | 12.45 | 167.025 | 161.68 | 3.3 | 19.85 | 149.48 | 144.02 | 4.7 | 29.2 | 139.25 | 133.543
-6.9 | 1.1994 | 6.3 | 190.7 | 185.1 | 1.5 | 8.7667 | 179.9 | 174.75 | 2 | 12.8667 | 166.1 | 160.7667 | 3.3 | 20.3857 | 148.3813 | 143.45 | 4.8 | 32.55 | 138.5 | 132.943
-6.8 | 1.1994 | 6.4 | 190.275 | 184.66 | 1.5 | 8.975 | 178.5 | 173.25 | 2 | 13.25 | 165.6 | 160.2 | 3.475 | 21.2 | 147.859 | 142.9 | 5 | 38 | 138.1 | 132.593
-6.7 | 1.1994 | 6.6 | 188.74 | 183.28 | 1.5 | 9.3333 | 177.4 | 172.3 | 2.1 | 13.75 | 164.3 | 159.275 | 3.5 | 22.36 | 147.399 | 142.2 | 5 | 45.575 | 137.5 | 132.293
-6.6 | 1.1994 | 6.8 | 187.25 | 181.6 | 1.5 | 9.6667 | 176.08 | 171.1 | 2.1 | 14.15 | 163.4 | 158.525 | 3.6 | 23.74 | 146.7304 | 141.6 | 5.2 | 55.5 | 136.95 | 131.6937
-6.5 | 1.1994 | 6.975 | 185.5 | 179.9 | 1.5 | 9.9 | 174.6667 | 169.775 | 2.2 | 14.75 | 162.2167 | 157.2 | 3.7 | 25.88 | 145.833 | 141.1 | 5.3 | 67.1286 | 136.4 | 131.2773
-6.4 | 1.1994 | 7.2 | 183.5 | 178.38 | 1.5 | 10.3333 | 172.975 | 168.425 | 2.2 | 15.4 | 161.4 | 156.075 | 3.7 | 29.4 | 145.2247 | 140.6 | 5.4 | 79.9 | 135.825 | 130.944
-6.3 | 1.2494 | 7.4 | 181.925 | 176.725 | 1.5 | 10.8 | 171.8 | 167.125 | 2.2 | 16.2 | 160.1857 | 155.1 | 3.75 | 35.3 | 144.583 | 139.8 | 5.5 | 91.95 | 135.1 | 130.4107
-6.2 | 1.2994 | 7.65 | 179.9 | 174.8 | 1.6 | 11.325 | 170.35 | 165.64 | 2.3 | 17.2 | 158.9 | 154.4 | 3.9 | 44.4 | 143.6327 | 139.1 | 5.6 | 102.14 | 134.7 | 129.944
-6.1 | 1.2994 | 7.9 | 178.1 | 173.125 | 1.6 | 12.05 | 169.2 | 164.2667 | 2.3 | 18.75 | 158.15 | 153.5333 | 3.9 | 55.95 | 142.717 | 138.25 | 5.5 | 110.9286 | 134.1 | 129.544
-6 | 1.2994 | 8.25 | 176.06 | 171.56 | 1.6 | 12.8 | 167.5167 | 162.8 | 2.3 | 21.025 | 156.85 | 152.25 | 4.0667 | 69.6 | 142.357 | 137.55 | 5.5 | 117 | 133.4 | 128.9117
-5.9 | 1.2994 | 8.65 | 174.44 | 169.8143 | 1.6 | 13.9333 | 166.12 | 161.45 | 2.4 | 25.05 | 155.45 | 151.26 | 4.1 | 83.4 | 141.4568 | 136.88 | 5.6 | 120.5889 | 132.7 | 128.395
-5.8 | 1.2994 | 9.4 | 172.7 | 168.06 | 1.625 | 15.9 | 164.4714 | 160.34 | 2.4 | 32 | 154.3 | 150.2333 | 4.2 | 95.9 | 140.581 | 136.3 | 6 | 122.58 | 132.3 | 127.745
-5.7 | 1.2994 | 10.3 | 171.025 | 166.225 | 1.7 | 18.9 | 163.125 | 158.8 | 2.5 | 42.7333 | 153.275 | 149.24 | 4.3 | 107.2 | 139.831 | 135.7 | 5.85 | 123.3714 | 131.64 | 127.445
-5.6 | 1.2994 | 12.15 | 169.3 | 164.75 | 1.7 | 24.4 | 161.7444 | 157.5 | 2.6 | 55.65 | 151.8 | 148.3714 | 4.5 | 116.9 | 139.306 | 134.88 | 6.25 | 123.7 | 131.0222 | 127.086
-5.5 | 1.2994 | 15.3 | 167.4333 | 163 | 1.8 | 34.45 | 159.9 | 156.2667 | 2.6 | 71 | 150.95 | 147.2 | 4.5 | 122.26 | 138.555 | 134.3429 | 6.3 | 123.7 | 130.32 | 126.466
-5.4 | 1.3993 | 21.9 | 165.45 | 161.375 | 1.8 | 47.7 | 158.2333 | 155.2 | 2.7667 | 87.1 | 149.8 | 146.15 | 4.6 | 125.5 | 137.5383 | 133.84 | 6.5 | 123.7 | 129.9667 | 126.106
-5.3 | 1.3993 | 32.9 | 163.7 | 159.8 | 1.8 | 64.3 | 156.92 | 153.5667 | 2.9 | 102 | 148.6333 | 145.06 | 4.7 | 126.8 | 136.7217 | 133.0625 | 6.8 | 123.38 | 129.2714 | 125.586
-5.2 | 1.3993 | 49.4 | 162.1667 | 158.2333 | 1.9 | 81.4 | 155.4 | 152.2 | 2.9 | 115.6 | 147.575 | 144.1 | 4.9 | 127.18 | 135.979 | 132.24 | 6.98 | 123.1 | 128.56 | 124.9086
-5.1 | 1.3993 | 67.3 | 160.4714 | 156.32 | 2 | 99.05 | 154.12 | 150.8 | 3 | 124.8 | 146.34 | 143 | 5 | 127.2 | 135.114 | 131.4 | 7.075 | 122.925 | 128.025 | 124.422
-5 | 1.4992 | 86.3 | 158.72 | 155.025 | 2 | 114.3 | 152.7143 | 149.3 | 3.0667 | 130.2 | 145.2857 | 141.85 | 5.3 | 127.2 | 134.374 | 130.5667 | 7.1 | 122.5667 | 127.65 | 123.647
-4.9 | 1.4992 | 105.6 | 156.7667 | 153.575 | 2.1 | 126.4 | 151.26 | 148.2714 | 3.1 | 133 | 144.5 | 140.8 | 5.3 | 127.05 | 133.453 | 129.95 | 7.2 | 122.35 | 127.025 | 123.227
-4.8 | 1.4992 | 121.1 | 155.1667 | 151.95 | 2.1 | 134.5 | 149.7 | 146.7667 | 3.3 | 133.6 | 143.275 | 139.7 | 5.4 | 126.8 | 132.503 | 129.34 | 7.3 | 122 | 126.375 | 122.847
-4.7 | 1.5992 | 134.3 | 153.525 | 150.42 | 2.2 | 137.8667 | 148.3 | 145.5429 | 3.4 | 133.6 | 141.98 | 138.7 | 5.5 | 126.4 | 131.753 | 128.6 | 7.5 | 121.3833 | 125.55 | 122.348
-4.6 | 1.5992 | 140.9 | 151.85 | 148.7 | 2.25 | 139.1 | 147.0333 | 144.1 | 3.5 | 133.6 | 140.4 | 137.6 | 5.65 | 125.7 | 130.952 | 127.9333 | 7.7 | 120.86 | 124.9 | 121.548
-4.5 | 1.6325 | 143.4 | 150.225 | 147.35 | 2.3 | 139.1 | 145.62 | 142.7 | 3.6 | 133.4 | 139.2 | 136.8 | 5.9 | 125.2 | 130.102 | 127.25 | 8 | 120.6333 | 124.3 | 121.223
-4.4 | 1.6991 | 143.6 | 148.575 | 145.6 | 2.4 | 139.1 | 144.125 | 141.45 | 3.7 | 132.8 | 137.7 | 135.8 | 6.125 | 124.65 | 129.352 | 126.575 | 8.2 | 120.2 | 123.6 | 120.873
-4.3 | 1.6991 | 143.6 | 146.95 | 144.2 | 2.5 | 138.4 | 142.525 | 140.1 | 3.8 | 132.1 | 137.1 | 134.85 | 6.3 | 124.1714 | 128.401 | 125.625 | 8.35 | 120.0286 | 122.8 | 120.523
-4.2 | 1.7991 | 143.3 | 145.4 | 142.68 | 2.6 | 137.45 | 141.3714 | 138.5 | 4 | 131.2333 | 136.1 | 133.5 | 6.425 | 123.82 | 127.801 | 124.85 | 8.6 | 119.64 | 122.2286 | 119.749
-4.1 | 1.8491 | 141.95 | 143.9 | 141.2 | 2.675 | 136.7 | 140.1333 | 137.4 | 4.2 | 130.4667 | 134.85 | 132.3 | 6.6 | 123.125 | 127.131 | 124.35 | 8.8 | 119.2 | 121.9333 | 119.369
-4 | 1.899 | 140.75 | 142.5 | 139.86 | 2.8 | 135.6 | 138.1 | 136.25 | 4.3 | 129.8 | 133.8 | 131.4 | 6.72 | 122.575 | 126.2503 | 123.725 | 8.95 | 118.68 | 121.375 | 118.649
-3.9 | 1.999 | 139.4 | 140.82 | 138.2 | 2.9667 | 134.4 | 137.2833 | 135.0667 | 4.5333 | 128.9667 | 132.7375 | 130.48 | 7.05 | 122.05 | 125.25 | 123.22 | 9.25 | 118.36 | 121.12 | 118.1919
-3.8 | 2.049 | 138.35 | 139.325 | 137 | 3.1 | 133.325 | 135.8333 | 133.7 | 4.7333 | 128.3 | 131.52 | 129.55 | 7.3 | 121.5667 | 124.45 | 122.46 | 9.6 | 117.925 | 120.35 | 117.6781
-3.7 | 2.0989 | 136.6 | 138.02 | 135.45 | 3.2 | 132.4 | 134.58 | 132.7 | 4.9 | 127.68 | 130.44 | 128.325 | 7.75 | 121.1 | 123.7371 | 122.08 | 9.7 | 117.62 | 119.7 | 117.27
-3.6 | 2.1989 | 135.45 | 136.7333 | 134.4 | 3.4 | 131.7 | 133.5333 | 131.6571 | 5.1 | 126.9 | 129.775 | 127.3 | 8 | 120.66 | 123.249 | 121.36 | 10.0667 | 117.2 | 119.1714 | 116.875
-3.5 | 2.2988 | 134.525 | 135.05 | 133.1 | 3.55 | 130.7 | 132.4857 | 130.4 | 5.375 | 125.85 | 128.475 | 126.76 | 8.2 | 120.2 | 122.7776 | 120.6 | 10.6667 | 116.875 | 118.725 | 116.45
-3.4 | 2.3988 | 133.4 | 133.675 | 131.7 | 3.7 | 129.9 | 131.1286 | 129.3 | 5.6 | 125.3 | 127.6 | 126.1 | 8.5 | 119.75 | 122.349 | 119.76 | 11.25 | 116.55 | 118.125 | 115.89
-3.3 | 2.4987 | 131.85 | 132.4 | 130.48 | 3.9 | 129 | 129.9 | 128.3 | 5.875 | 124.725 | 126.65 | 124.7167 | 8.9333 | 119.4143 | 121.7633 | 119.225 | 11.8 | 116.22 | 117.74 | 115.65
-3.2 | 2.5987 | 130.75 | 131.1286 | 129.3 | 4.1 | 127.8 | 128.7667 | 127.2 | 6.15 | 124 | 125.8 | 124.1 | 9.45 | 118.74 | 121.088 | 118.6 | 12.8667 | 115.75 | 117.2333 | 115.211
-3.1 | 2.6987 | 129.875 | 130.0333 | 128.15 | 4.3 | 127.1 | 127.85 | 126.425 | 6.66 | 123.225 | 125.2 | 123.2 | 10.3667 | 118.1167 | 120.348 | 118 | 14.4 | 115.4 | 116.9 | 114.671
-3 | 2.8486 | 128.9 | 129.08 | 127.3333 | 4.7333 | 126 | 127.075 | 125.4 | 7.1 | 122.6 | 124 | 122.475 | 11.3 | 117.3667 | 119.628 | 117.6 | 17.2 | 114.95 | 116.3625 | 114.351
-2.9 | 3.0651 | 127.8 | 127.7 | 126.26 | 5.3 | 125.2 | 125.8667 | 124.5 | 7.7 | 121.95 | 123.34 | 121.82 | 13.0333 | 117.05 | 118.6901 | 117.05 | 21.55 | 114.3 | 115.98 | 113.776
-2.8 | 3.3983 | 126.7 | 126.6 | 125.1 | 5.9 | 124.3 | 125 | 123.7 | 8.8 | 121.3333 | 122.6 | 120.9 | 15.7333 | 116.65 | 117.987 | 116.7 | 27.4 | 113.95 | 115.3 | 113.301
-2.7 | 3.8981 | 125.4 | 125.58 | 124.3 | 7 | 123.3 | 124 | 122.6 | 10.475 | 120.5 | 121.6 | 120.1 | 20.2 | 116.1667 | 117.247 | 115.925 | 37.25 | 113.675 | 114.5 | 112.692
-2.6 | 4.7976 | 124.2 | 124.6143 | 123.25 | 8.5 | 122.575 | 122.7 | 121.775 | 13.3 | 119.7333 | 120.7 | 119.25 | 27.775 | 115.78 | 116.647 | 115.35 | 47.6 | 113.34 | 114.04 | 112.002
-2.5 | 6.6467 | 123.4 | 123.5857 | 122.2 | 11.65 | 121.625 | 122 | 120.7 | 18.5 | 118.8 | 120.22 | 118.4143 | 37.9 | 115.4 | 116.2467 | 114.65 | 59.95 | 112.825 | 113.55 | 111.7187
-2.4 | 10.1949 | 122.5 | 122.5 | 121.05 | 17.3 | 120.9 | 121 | 119.9 | 26.55 | 118.3 | 119.1 | 117.75 | 49.3 | 114.55 | 115.471 | 114.05 | 71.9 | 112.25 | 113 | 111.292
-2.3 | 16.5251 | 121.4333 | 121.375 | 120.2 | 25.725 | 119.8 | 119.85 | 118.98 | 37.3333 | 117.3 | 118.2667 | 116.9 | 61.8 | 113.9 | 114.796 | 113.425 | 83.1333 | 111.9 | 112.4 | 110.7724
-2.2 | 26.9865 | 120.4 | 120.34 | 119.3 | 37.25 | 118.9 | 119.06 | 118.3 | 50.7 | 116.5 | 117.225 | 116.2 | 74.1 | 113.3 | 114.296 | 112.8 | 92.1 | 111.45 | 112.02 | 110.353
-2.1 | 39.7801 | 119.4 | 119.5 | 118.275 | 51.4 | 118.2333 | 118.125 | 117.26 | 64.3 | 115.6 | 116.575 | 115.2 | 86.4667 | 112.6 | 113.6654 | 112.2 | 98.3 | 111.0143 | 111.4333 | 109.903
-2 | 54.4728 | 118.4 | 118.36 | 117.3 | 65.3 | 117.45 | 117.26 | 116.3 | 77.2667 | 115.1 | 115.7 | 114.5 | 95 | 112.14 | 112.9164 | 111.5 | 102.9 | 110.3 | 110.85 | 109.453
-1.9 | 69.7651 | 117.4333 | 117.38 | 116.3 | 79.3 | 116.3 | 116.3 | 115.6 | 88.4 | 114.4 | 114.85 | 113.95 | 101.125 | 111.425 | 112.045 | 111.2 | 105 | 109.9 | 110.35 | 108.9863
-1.8 | 83.1834 | 116.2 | 116.35 | 115.6 | 90.8 | 115.35 | 115.2 | 114.7 | 97.2667 | 113.6 | 114 | 113.225 | 104.9 | 110.7857 | 111.4648 | 110.4667 | 105.975 | 109.3857 | 109.825 | 108.354
-1.7 | 95.1774 | 115.4333 | 115.275 | 114.5 | 99.8 | 114.5667 | 114.3 | 113.7 | 103.475 | 113 | 113.2 | 112.44 | 106.5 | 110.2 | 110.844 | 110.02 | 106.475 | 108.96 | 109.15 | 107.8207
-1.6 | 103.648 | 114.4 | 114.3667 | 113.58 | 104.95 | 113.3667 | 113.5 | 112.78 | 106.6 | 112 | 112.42 | 111.6286 | 107.1667 | 109.6143 | 109.944 | 109.2 | 106.7 | 108.3286 | 108.5 | 107.354
-1.5 | 108.371 | 113.2333 | 113.275 | 112.66 | 107.6 | 112.8333 | 112.5333 | 111.825 | 107.4333 | 111.3 | 111.725 | 110.94 | 107.2 | 109.2125 | 109.569 | 108.65 | 106.64 | 107.82 | 107.9333 | 106.954
-1.4 | 109.745 | 112.4 | 112.25 | 111.75 | 108 | 111.8 | 111.7 | 111.1 | 107.725 | 110.5667 | 111.06 | 110.0429 | 107.2 | 108.56 | 108.844 | 108.18 | 106.3 | 107.2143 | 107.5 | 106.454
-1.3 | 109.845 | 111.6 | 111.5 | 110.8 | 108.1 | 110.9 | 110.8333 | 110.25 | 107.8 | 109.65 | 109.95 | 109.22 | 107.1667 | 107.8 | 108.293 | 107.525 | 106.0667 | 106.65 | 107.0333 | 106.075
-1.2 | 109.745 | 110.5 | 110.4 | 109.8333 | 108.05 | 109.9 | 109.8 | 109.525 | 107.6 | 108.95 | 109.2 | 108.6 | 106.75 | 107.26 | 107.693 | 106.88 | 105.7 | 106.0667 | 106.5 | 105.555
-1.1 | 109.1783 | 109.7 | 109.5167 | 109 | 107.8333 | 109 | 108.925 | 108.5333 | 106.9 | 108.3 | 108.475 | 108.1 | 106.4 | 106.6714 | 107.093 | 106.275 | 105.25 | 105.7667 | 106.125 | 105.105
-1 | 108.546 | 108.6 | 108.5 | 108.15 | 107.425 | 108.3 | 108.08 | 107.7 | 106.3 | 107.4667 | 107.725 | 107.0333 | 105.9 | 106.05 | 106.3826 | 105.5667 | 104.8 | 105.3 | 105.38 | 104.515
-0.9 | 107.646 | 107.6 | 107.5714 | 107.4333 | 106.65 | 107.4 | 107.3714 | 106.875 | 105.75 | 106.8 | 107 | 106.1 | 105.45 | 105.45 | 105.5087 | 105.04 | 104.5 | 104.7667 | 104.76 | 103.915
-0.8 | 106.797 | 106.6667 | 107.04 | 106.7 | 106.1 | 106.5 | 106.6 | 106.2 | 105.1 | 105.8333 | 105.9 | 105.42 | 104.675 | 105 | 104.8134 | 104.55 | 104.25 | 104.075 | 104.2667 | 103.536
-0.7 | 106.047 | 105.7 | 106 | 105.6 | 105.3 | 105.7 | 105.525 | 105.3 | 104.5667 | 105.2 | 105.16 | 104.85 | 104.075 | 104.4 | 104.342 | 103.9 | 103.5667 | 103.475 | 103.9 | 103.016
-0.6 | 105.247 | 105 | 104.95 | 104.75 | 104.4667 | 104.8333 | 104.76 | 104.4 | 104.1 | 104.2 | 104.52 | 104.16 | 103.85 | 103.68 | 103.7216 | 103.45 | 103.2 | 102.95 | 103.58 | 102.676
-0.5 | 104.348 | 104.1 | 104.0667 | 104 | 103.6 | 103.975 | 104.15 | 103.55 | 103.5 | 103.6 | 103.725 | 103.5 | 103.4 | 103.05 | 103.1743 | 103.0143 | 102.8 | 102.5 | 102.9 | 102.156
-0.4 | 103.548 | 103.1 | 103.1833 | 102.9 | 102.65 | 102.9 | 103.2 | 102.85 | 102.7 | 102.9 | 103.04 | 102.78 | 102.7 | 102.45 | 102.4981 | 102.35 | 102 | 102.05 | 102.3429 | 101.756
-0.3 | 102.749 | 102.5333 | 102.4 | 102.24 | 102.05 | 102.2 | 102.4 | 102.16 | 101.9667 | 102.1 | 102.26 | 102.2 | 102.1 | 101.65 | 102.081 | 101.65 | 101.65 | 101.6 | 101.7714 | 101.5006
-0.2 | 101.8157 | 101.5333 | 101.5286 | 101.55 | 101.5 | 101.5 | 101.3 | 101.5333 | 101.4 | 101.6 | 101.4571 | 101.25 | 101.7 | 101.0667 | 101.761 | 101 | 101.3 | 101 | 101.28 | 100.907
-0.1 | 100.7 | 100.625 | 100.725 | 100.8 | 100.7 | 100.7333 | 100.6 | 100.74 | 100.5 | 100.8 | 100.7 | 100.72 | 100.9 | 100.4 | 100.78 | 100.7 | 100.8 | 100.6 | 100.8 | 100.4284
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.1505 | 98.9 | 99.38 | 99.3 | 99.125 | 99.05 | 99.2 | 99.1667 | 99.45 | 99.275 | 99.4 | 99.3333 | 99.2333 | 99.62 | 99.4147 | 99.45 | 99.2667 | 99.4 | 99.3875 | 99.4574
0.2 | 98.2509 | 98.2 | 98.3857 | 98.4 | 98.475 | 98.25 | 98.38 | 98.6143 | 98.55 | 98.6667 | 98.6833 | 98.6333 | 98.7 | 98.975 | 98.8996 | 99.04 | 98.9 | 98.9 | 99.1 | 99.1775
0.3 | 97.3264 | 97.4 | 97.5 | 97.7 | 97.5 | 97.6 | 97.66 | 97.7571 | 97.7 | 97.95 | 98.16 | 97.975 | 98.1 | 98.14 | 98.3994 | 98.38 | 98.5 | 98.45 | 98.7333 | 98.6577
0.4 | 96.9515 | 96.5 | 96.775 | 97.025 | 96.9667 | 96.8 | 96.85 | 97.06 | 97 | 97.3 | 97.55 | 97.4286 | 97.5 | 97.625 | 97.7892 | 97.8857 | 97.9 | 97.925 | 98.4 | 98.3078
0.5 | 96.1519 | 95.55 | 96.02 | 96.275 | 96.2 | 96.1667 | 96.2 | 96.4286 | 96.2 | 96.6 | 96.7 | 96.9 | 97.1333 | 96.8 | 97.3189 | 97.1 | 97.05 | 97.56 | 97.7833 | 97.7081
0.6 | 95.2524 | 94.7 | 95.325 | 95.725 | 95.25 | 95.42 | 95.52 | 95.65 | 95.5 | 96.05 | 95.85 | 96.3 | 96.5333 | 96.325 | 96.7387 | 96.6375 | 96.7 | 97.05 | 97.2 | 97.2583
0.7 | 94.2196 | 94.15 | 94.6 | 94.95 | 94.45 | 94.7 | 94.74 | 95 | 94.5 | 95.2 | 95.44 | 95.48 | 95.85 | 95.65 | 96.0384 | 96.02 | 96.15 | 96.65 | 96.85 | 96.7385
0.8 | 93.2534 | 93.3333 | 93.82 | 94.2 | 93.55 | 94.1 | 94.0857 | 94.4667 | 93.75 | 94.4 | 94.75 | 94.96 | 95.0667 | 95.16 | 95.5582 | 95.58 | 95.65 | 96.3 | 96.5 | 96.1874
0.9 | 92.5037 | 92.6 | 93.1 | 93.625 | 92.8 | 93.6 | 93.4429 | 93.75 | 93.25 | 93.75 | 93.9667 | 94.5 | 94.4 | 94.575 | 94.8379 | 95.2 | 95 | 95.3667 | 95.9 | 95.7018
1 | 91.8541 | 92.2 | 92.5333 | 92.8667 | 92 | 92.8 | 92.95 | 93.2333 | 92.5333 | 93.05 | 93.4 | 93.55 | 93.7 | 94.16 | 94.4378 | 94.62 | 94.3 | 94.9143 | 95.46 | 95.3591
1.1 | 90.9046 | 91.3 | 91.975 | 92.1286 | 90.75 | 91.8 | 91.98 | 92.625 | 91.75 | 92.3 | 92.85 | 93 | 93.1 | 93.5 | 93.8175 | 94.1 | 93.68 | 94.26 | 94.925 | 94.9593
1.2 | 90.2549 | 90.4 | 91.0333 | 91.4 | 89.9 | 91.18 | 91.54 | 91.85 | 90.9333 | 91.65 | 92.3 | 92.58 | 92.1 | 92.9 | 93.2373 | 93.65 | 93.05 | 93.85 | 94.25 | 94.5594
1.3 | 88.9555 | 89.9 | 90.45 | 90.9 | 88.55 | 90.55 | 90.6 | 91.08 | 89.7 | 91.25 | 91.45 | 91.9571 | 91.275 | 92.32 | 92.7514 | 93.32 | 92.64 | 93.48 | 93.78 | 93.9598
1.4 | 87.0065 | 89.1 | 89.55 | 90.18 | 86.9 | 90 | 90 | 90.4667 | 88.7 | 90.75 | 90.875 | 91.2 | 90.44 | 91.7286 | 92.2169 | 92.75 | 91.8 | 93.025 | 93.1714 | 93.5599
1.5 | 84.3578 | 88.2333 | 89.025 | 89.5 | 85.4 | 89.2333 | 89.32 | 89.96 | 87.3 | 90 | 90.18 | 90.625 | 89.7 | 91.1571 | 91.508 | 92.2 | 91.1333 | 92.45 | 92.7 | 92.9602
1.6 | 79.9101 | 87.6 | 88.1286 | 88.94 | 82.1 | 88.5 | 88.7 | 89.3143 | 85.5 | 89.2 | 89.6 | 90.05 | 88.4 | 90.5 | 90.9864 | 91.48 | 90.3 | 91.85 | 92.18 | 92.6103
1.7 | 72.8302 | 87.1 | 87.5 | 88.1667 | 77.2 | 87.8 | 88.1 | 88.68 | 82.1 | 88.65 | 88.975 | 89.45 | 87 | 89.95 | 90.5162 | 91.04 | 89.5 | 91.28 | 91.86 | 92.0605
1.8 | 63.3184 | 86.4 | 86.8667 | 87.5571 | 70 | 87 | 87.3 | 88.1 | 77.5667 | 88.2 | 88.4 | 89.04 | 84.8 | 89.3857 | 90.1761 | 90.5 | 88.15 | 90.76 | 91.38 | 91.8106
1.9 | 52.9736 | 85.7 | 86.3 | 86.9 | 61.7 | 86.4667 | 86.7 | 87.4 | 70.85 | 87.5 | 87.7333 | 88.3 | 81.875 | 88.7667 | 89.4024 | 90 | 86.3667 | 90.24 | 91 | 91.4808
2 | 41.8457 | 84.9 | 85.65 | 86.2 | 50.4333 | 85.6 | 86.18 | 86.7667 | 62.3 | 86.8 | 87 | 87.8333 | 77.3 | 88.3333 | 88.8355 | 89.6429 | 84.775 | 89.5 | 90.35 | 90.881
2.1 | 30.6597 | 84.3 | 84.9 | 85.7 | 39.6 | 85.25 | 85.5 | 86.18 | 51.65 | 86.2 | 86.3 | 87.2 | 71 | 87.8857 | 88.2603 | 89.1 | 82.0333 | 89.22 | 89.7 | 90.2813
2.2 | 21.3893 | 83.45 | 84.14 | 85.0667 | 29.6 | 84.5 | 84.8143 | 85.4 | 42.1 | 85.5 | 85.94 | 86.55 | 63.05 | 87.2 | 87.495 | 88.3 | 77.4 | 88.8143 | 89.125 | 89.9115
2.3 | 13.4433 | 82.9 | 83.375 | 84.4571 | 21.18 | 83.75 | 84.12 | 85.075 | 31.8667 | 84.6 | 85.2429 | 85.78 | 53.5 | 86.3 | 87.1348 | 87.9333 | 72 | 88.2429 | 88.7 | 89.4117
2.4 | 8.1959 | 82.15 | 82.825 | 83.7 | 13.95 | 82.95 | 83.6143 | 84.46 | 23.4667 | 83.6 | 84.76 | 85.38 | 44.0333 | 85.7 | 86.7147 | 87.42 | 64.3 | 87.62 | 88.45 | 89.0118
2.5 | 5.364 | 81.6 | 82.15 | 83.225 | 9.7333 | 82.1333 | 83.18 | 84 | 16.6667 | 83.2 | 84.24 | 84.85 | 34.7 | 85.2 | 86.2059 | 87 | 55 | 87.32 | 87.8 | 88.5121
2.6 | 4.0979 | 80.8 | 81.7 | 82.7 | 7.2 | 81.6333 | 82.5625 | 83.5 | 11.7667 | 82.7333 | 83.4667 | 84.28 | 26.1667 | 84.575 | 85.5342 | 86.5 | 45.8 | 86.48 | 87.15 | 87.9623
2.7 | 3.3317 | 80.2333 | 81.15 | 82.15 | 5.7 | 81.1 | 81.84 | 82.7 | 9.4 | 82.3 | 83.04 | 83.8 | 19.8 | 84.0333 | 84.9506 | 85.9 | 36.625 | 85.92 | 86.75 | 87.4625
2.8 | 2.7486 | 79.7 | 80.65 | 81.5 | 4.85 | 80.4 | 81.1857 | 82.1667 | 7.7 | 81.5 | 82.35 | 83.1 | 15.2 | 83.475 | 84.4338 | 85.42 | 29.3 | 85.25 | 86.3667 | 87.2126
2.9 | 2.5654 | 78.95 | 80.1 | 81.14 | 4.3 | 79.9 | 80.75 | 81.84 | 6.9 | 80.7 | 81.95 | 82.6 | 11.9 | 82.75 | 84.0336 | 84.86 | 22.075 | 85 | 85.8333 | 87.0627
3 | 2.3988 | 78.5333 | 79.4429 | 80.46 | 4 | 79.3333 | 80.2889 | 81.36 | 6.3 | 80.3 | 81.45 | 82.25 | 10.25 | 82.25 | 83.3733 | 84.375 | 17.5 | 84.5 | 85.375 | 86.4963
3.1 | 2.2489 | 78.1 | 78.9 | 80 | 3.6 | 78.6 | 79.84 | 80.72 | 5.7333 | 79.6 | 80.7 | 81.8 | 9.3 | 81.65 | 83.0832 | 83.9 | 14.8 | 84 | 84.975 | 85.8382
3.2 | 2.0989 | 77.25 | 78.4 | 79.54 | 3.4667 | 78.1 | 79.4143 | 80.1857 | 5.4 | 78.95 | 80.25 | 81.35 | 8.8 | 81.16 | 82.533 | 83.575 | 12.6 | 83.5 | 84.54 | 85.4134
3.3 | 1.999 | 76.7 | 78.1 | 79 | 3.2 | 77.5333 | 78.9 | 79.7 | 5.1667 | 78.5333 | 79.74 | 80.78 | 8.1 | 80.72 | 82.0928 | 83.1857 | 11.3 | 82.7714 | 84.1 | 85.0885
3.4 | 1.899 | 76.25 | 77.4667 | 78.55 | 3.1 | 77.1 | 78.225 | 79.26 | 4.8 | 78 | 79.26 | 80.15 | 7.7 | 79.9 | 81.4325 | 82.5667 | 10.4 | 82.42 | 83.6 | 84.7037
3.5 | 1.7991 | 75.7 | 76.96 | 78.04 | 3 | 76.1667 | 77.74 | 78.6 | 4.65 | 77.3 | 78.62 | 79.65 | 7.4 | 79.275 | 81.0324 | 82 | 10.04 | 81.9 | 83.1667 | 84.2472
3.6 | 1.7991 | 75 | 76.5 | 77.475 | 2.8 | 75.6 | 77.15 | 78.14 | 4.4 | 76.7 | 78.175 | 79.16 | 6.975 | 79 | 80.5822 | 81.5857 | 9.56 | 81.34 | 82.7333 | 83.914
3.7 | 1.6991 | 74.5667 | 76 | 76.95 | 2.7 | 75.2 | 76.7 | 77.6429 | 4.25 | 75.9333 | 77.8667 | 78.8333 | 6.8 | 78.45 | 79.907 | 81.22 | 9.35 | 80.9143 | 82.325 | 83.4043
3.8 | 1.5992 | 73.95 | 75.3571 | 76.475 | 2.6 | 74.3 | 76.35 | 77.2 | 4.1 | 75.3 | 77.3 | 78.4333 | 6.575 | 77.75 | 79.5819 | 80.82 | 8.7 | 80.24 | 82.05 | 83.0644
3.9 | 1.5992 | 73.3 | 75 | 76.05 | 2.4 | 73.85 | 75.75 | 76.8 | 3.975 | 74.7667 | 76.94 | 77.84 | 6.325 | 77.325 | 79.2567 | 80.1333 | 8.45 | 79.7 | 81.55 | 82.7245
4 | 1.4992 | 72.64 | 74.5 | 75.65 | 2.3 | 73.15 | 75.1875 | 76.2 | 3.8 | 74.2 | 76.3 | 77.5 | 6.18 | 76.8 | 78.8566 | 79.78 | 8.3 | 79.06 | 81.1667 | 82.2647
4.1 | 1.4992 | 72 | 73.9 | 75.16 | 2.2 | 72.6 | 74.625 | 75.85 | 3.6 | 73.72 | 75.84 | 76.85 | 6.1 | 76.0667 | 78.3813 | 79.35 | 8.1 | 78.5 | 80.65 | 81.8249
4.2 | 1.3993 | 71.4 | 73.4857 | 74.6 | 2.1 | 71.9333 | 74.125 | 75.3 | 3.5 | 72.8667 | 75.3 | 76.35 | 5.9 | 75.46 | 77.9646 | 78.9 | 7.9 | 78.2143 | 80.18 | 81.2652
4.3 | 1.3993 | 70.375 | 73 | 74.1 | 2.1 | 71.2 | 73.625 | 74.85 | 3.3 | 72.22 | 74.85 | 75.86 | 5.65 | 74.8667 | 77.331 | 78.5 | 7.6667 | 77.72 | 79.74 | 80.832
4.4 | 1.3193 | 69.3 | 72.35 | 73.64 | 2 | 70.375 | 73.2 | 74.25 | 3.2667 | 71.3 | 74.4 | 75.5 | 5.4 | 74.425 | 76.8736 | 78 | 7.4 | 77.2714 | 79.56 | 80.5155
4.5 | 1.2994 | 67.85 | 71.925 | 73.1 | 1.9 | 69.3 | 72.65 | 73.85 | 3.1 | 70.7 | 73.8571 | 75.02 | 5.3 | 73.575 | 76.3591 | 77.55 | 7.3 | 76.725 | 79.1 | 79.999
4.6 | 1.2994 | 65.25 | 71.35 | 72.74 | 1.8 | 68.25 | 72.2 | 73.4 | 2.9 | 69.8667 | 73.375 | 74.5 | 5.15 | 72.96 | 75.8503 | 77.15 | 7.2 | 76.14 | 78.5143 | 79.7408
4.7 | 1.1994 | 61.0667 | 71 | 72.2 | 1.8 | 66.5 | 71.7 | 73.1 | 2.9 | 69 | 72.7 | 74.06 | 5 | 72.54 | 75.2301 | 76.65 | 6.9 | 75.525 | 78.1 | 79.3659
4.8 | 1.1994 | 54.95 | 70.45 | 71.6 | 1.7 | 64 | 71.2 | 72.4 | 2.85 | 67.8 | 72.38 | 73.65 | 5 | 71.8 | 74.7299 | 76.26 | 6.9 | 74.9333 | 77.7333 | 78.8662
4.9 | 1.1994 | 47.8 | 70 | 71.2 | 1.7 | 60 | 70.55 | 71.86 | 2.8 | 66.6 | 71.88 | 73.02 | 4.8 | 71.0286 | 74.3726 | 75.7333 | 6.7 | 74.325 | 77.26 | 78.4664
5 | 1.1994 | 39.6 | 69.45 | 70.7667 | 1.7 | 54.3 | 70.2 | 71.42 | 2.6 | 64.6 | 71.5 | 72.62 | 4.6 | 70.36 | 74.0096 | 75.3 | 6.6333 | 73.66 | 76.6714 | 78.0165
5.1 | 1.1994 | 31 | 68.94 | 70.35 | 1.6 | 47.3 | 69.7 | 70.92 | 2.5 | 61.3 | 70.9 | 72.1286 | 4.5 | 69.8 | 73.5169 | 74.76 | 6.5 | 73 | 76.375 | 77.6668
5.2 | 1.0995 | 23.4 | 68.58 | 69.75 | 1.525 | 39.5 | 69.28 | 70.6 | 2.4 | 56.9 | 70.36 | 71.7 | 4.3 | 68.72 | 73.1292 | 74.35 | 6.1 | 72.5 | 76.2 | 77.2419
5.3 | 1.0995 | 16.9 | 68.1286 | 69.2333 | 1.5 | 31.7 | 68.84 | 70.18 | 2.4 | 51.4 | 70.1 | 71.02 | 4.3 | 67.825 | 72.5124 | 73.92 | 6 | 72 | 75.6 | 76.8671
5.4 | 1.0995 | 11.8 | 67.52 | 68.75 | 1.4333 | 24.7 | 68.32 | 69.52 | 2.4 | 44.7 | 69.68 | 70.6 | 4.2 | 66.275 | 72.0888 | 73.44 | 5.9 | 71.3 | 74.9429 | 76.4673
5.5 | 1.0995 | 8.9 | 67 | 68.26 | 1.4 | 18.3 | 67.85 | 68.9429 | 2.3 | 37.2 | 69.075 | 70.02 | 4.1 | 64.75 | 71.6787 | 73.1 | 5.8 | 70.65 | 74.5 | 75.9874
5.6 | 0.9995 | 7.2 | 66.3833 | 67.8 | 1.4 | 13.8 | 67.4 | 68.5 | 2.2 | 30.25 | 68.4 | 69.55 | 4 | 62 | 71.084 | 72.65 | 5.8 | 70.1333 | 74.075 | 75.4927
5.7 | 0.9995 | 6.3667 | 65.9667 | 67.3 | 1.4 | 10.9 | 66.85 | 68.05 | 2.1 | 24.2 | 67.9 | 69.075 | 3.925 | 58.05 | 70.5282 | 72.06 | 5.7 | 69.2 | 73.75 | 75.0678
5.8 | 0.9995 | 5.84 | 65.48 | 66.82 | 1.4 | 9.3 | 66.16 | 67.45 | 2.0333 | 19 | 67.4286 | 68.725 | 3.8 | 53.1 | 70.208 | 71.48 | 5.5 | 68.2 | 73.3286 | 74.618
5.9 | 0.9995 | 5.5 | 65 | 66.3 | 1.3 | 8.4 | 65.82 | 66.9667 | 2 | 15.1 | 67.02 | 68.15 | 3.7 | 46.7 | 69.6678 | 71.06 | 5.5 | 67.1 | 72.7 | 74.2182
6 | 0.9995 | 5.2333 | 64.45 | 65.9 | 1.3 | 7.8 | 65.22 | 66.6 | 2 | 13.1 | 66.15 | 67.6 | 3.7 | 40.15 | 69.1991 | 70.68 | 5.3 | 65.275 | 72.1625 | 73.7684
6.1 | 0.9995 | 5.05 | 64 | 65.35 | 1.3 | 7.3 | 64.7 | 66.0333 | 1.9 | 11.8 | 65.72 | 67.12 | 3.5333 | 33.9 | 68.7775 | 70.2 | 5.1 | 62.75 | 71.72 | 73.4019
6.2 | 0.9995 | 5 | 63.5333 | 64.8 | 1.3 | 7 | 64.0667 | 65.5143 | 1.9 | 10.9333 | 65.2143 | 66.65 | 3.5 | 28.1 | 68.3162 | 69.7 | 5.0333 | 59.0667 | 71.28 | 72.9687
6.3 | 0.89955 | 4.8 | 62.9333 | 64.2 | 1.2 | 6.7 | 63.6333 | 64.96 | 1.8 | 10.34 | 64.925 | 66.15 | 3.4 | 23.2 | 67.8021 | 69.3222 | 4.8 | 54.4 | 70.9 | 72.5689
6.4 | 0.89955 | 4.7 | 62.35 | 63.7 | 1.2 | 6.5 | 63.0429 | 64.4 | 1.8 | 9.85 | 64.22 | 65.6667 | 3.3 | 19.9 | 67.1769 | 68.8 | 4.7 | 48.775 | 70.5 | 72.0691
6.5 | 0.9995 | 4.6 | 61.675 | 63.32 | 1.2 | 6.3 | 62.6 | 63.975 | 1.8 | 9.45 | 63.85 | 65.14 | 3.3 | 17.65 | 66.7067 | 68.2667 | 4.7 | 43.2 | 70.08 | 71.6693
6.6 | 0.89955 | 4.5 | 61.225 | 62.85 | 1.2 | 6.1 | 62.02 | 63.4 | 1.75 | 9.2 | 63.22 | 64.475 | 3.2 | 16.1 | 66.2837 | 67.86 | 4.6 | 36.925 | 69.58 | 71.2361
6.7 | 0.89955 | 4.4 | 60.64 | 62.18 | 1.1 | 5.8 | 61.58 | 62.925 | 1.7 | 8.85 | 62.56 | 64 | 3.1 | 15.26 | 66.0264 | 67.32 | 4.6 | 31.5 | 69.15 | 70.5697
6.8 | 0.89955 | 4.225 | 60.1 | 61.375 | 1.1 | 5.7 | 60.84 | 62.2 | 1.7 | 8.6 | 61.98 | 63.44 | 3.0333 | 14.68 | 65.3512 | 66.825 | 4.6 | 27.15 | 68.82 | 70.2699
6.9 | 0.89955 | 4.2 | 59.1333 | 60.6 | 1.1 | 5.6 | 60.24 | 61.7333 | 1.7 | 8.35 | 61.5 | 62.9857 | 3 | 14.06 | 64.7259 | 66.3 | 4.6 | 23.7 | 68.4286 | 69.8451
7 | 0.89955 | 4.0333 | 57.875 | 59.58 | 1.1 | 5.5 | 59.5333 | 61.025 | 1.6 | 8.15 | 60.925 | 62.4 | 2.85 | 13.725 | 64.2757 | 65.8 | 4.5 | 21.35 | 67.68 | 69.3702
7.1 |  | 4 | 56.3 | 58 | 1.1 | 5.26 | 58.84 | 60.475 | 1.6 | 8 | 60.2 | 61.8143 | 2.9 | 13.3 | 63.8255 | 65.4 | 4.3 | 19.9 | 67.4 | 68.9704
7.2 |  | 3.9 | 53.825 | 55.34 |  | 5.1 | 57.8875 | 59.5 | 1.6 | 7.7 | 59.72 | 61.2556 | 2.8 | 12.95 | 63.1252 | 64.8 | 4.3 | 18.74 | 67.05 | 68.5956
7.3 |  | 3.8 | 49.4 | 51.4 |  | 5 | 56.65 | 58.2333 |  | 7.525 | 59.05 | 60.6 | 2.7 | 12.725 | 62.575 | 64.35 | 4.1333 | 18.125 | 66.55 | 67.9309
7.4 |  | 3.8 | 44.1 | 46.075 |  | 4.9 | 54.75 | 56.575 |  | 7.275 | 58.42 | 59.94 | 2.7 | 12.4 | 62.1248 | 63.85 | 3.95 | 17.5 | 65.975 | 67.691
7.5 |  | 3.7 | 37.7 | 40.2667 |  | 4.8 | 52.1 | 54 |  | 7.125 | 57.6 | 59.1 |  | 12.0333 | 61.5913 | 63.26 | 3.95 | 17.075 | 65.54 | 67.2512
7.6 |  | 3.6 | 30.7 | 33.975 |  | 4.7 | 47.7625 | 50.05 |  | 6.9 | 56.7 | 58.325 |  | 11.8667 | 61.0244 | 62.75 | 3.8 | 16.7667 | 65.04 | 66.6428
7.7 |  | 3.6 | 24.15 | 27.475 |  | 4.6 | 42.75 | 45.25 |  | 6.8 | 55.35 | 57 |  | 11.5667 | 60.4842 | 62.2889 |  | 16.38 | 64.5 | 66.0717
7.8 |  | 3.6 | 18.6 | 22.175 |  | 4.5 | 36.9333 | 39.725 |  | 6.6 | 53.48 | 55.36 |  | 11.3 | 60.074 | 61.74 |  | 15.9375 | 64.04 | 65.6385
7.9 |  | 3.5 | 13.9857 | 17.9 |  | 4.4 | 30.65 | 33.7333 |  | 6.4 | 50.575 | 52.6 |  | 11.025 | 59.4838 | 61.1 |  | 15.8 | 63.65 | 65.2221
8 |  | 3.5 | 11 | 15.14 |  | 4.3 | 24.5714 | 28.15 |  | 6.3 | 46.325 | 48.84 |  | 10.9 | 58.9903 | 60.56 |  | 15.44 | 63.0333 | 64.9388
8.1 |  | 3.5 | 9.35 | 13.52 |  | 4.2 | 19.25 | 23.5 |  | 6.2 | 41.55 | 44.5 |  | 10.675 | 58.3567 | 60.0667 |  | 14.975 | 62.4667 | 64.3474
8.2 |  | 3.4 | 8.46 | 12.7111 |  | 4.2 | 15.52 | 19.5 |  | 6 | 36.12 | 39.1429 |  | 10.45 | 57.423 | 59.3 |  | 14.8333 | 61.9333 | 63.806
8.3 |  | 3.4 | 7.9714 | 12.3 |  | 4.1 | 12.86 | 16.8 |  | 5.8 | 30.34 | 33.85 |  | 10.2 | 56.6227 | 58.5286 |  | 14.55 | 61.35 | 63.2529
8.4 |  | 3.4 | 7.6875 | 11.94 |  | 4.1 | 11.1 | 15.4 |  | 5.8 | 25.35 | 28.75 |  | 10 | 55.7223 | 57.4 |  | 14.16 | 60.85 | 62.8016
8.5 |  | 3.4 | 7.5143 | 11.8 |  | 4 | 10.34 | 14.5 |  | 5.7 | 20.65 | 24.76 |  | 9.85 | 54.2883 | 56.25 |  | 14.0125 | 60.35 | 62.2534
8.6 |  | 3.4 | 7.3 | 11.575 |  | 4 | 9.7 | 14 |  | 5.6 | 17.3 | 21.55 |  | 9.6 | 52.1208 | 54.3333 |  | 13.9 | 60 | 61.8235
8.7 |  | 3.4 | 7.2125 | 11.5 |  | 4 | 9.34 | 13.7 |  | 5.4 | 14.9 | 19.425 |  | 9.36 | 49.2397 | 51.7429 |  | 13.5 | 59.34 | 61.3238
8.8 |  | 3.3 | 7.1 | 11.4 |  | 4 | 9.0857 | 13.45 |  | 5.4 | 13.6 | 18.125 |  | 9.2 | 45.4581 | 48.3 |  | 13.2429 | 58.9 | 60.724
8.9 |  | 3.3 | 7.1 | 11.3 |  | 3.9 | 8.84 | 13.2 |  | 5.3 | 12.82 | 17.375 |  | 9.1 | 40.9664 | 44.05 |  | 13.1333 | 58.15 | 60.2742
9 |  | 3.3 | 7 | 11.2 |  | 3.8 | 8.65 | 13 |  | 5.3 | 12.3 | 16.86 |  | 8.8333 | 36.4146 | 39.7 |  | 12.9429 | 57.4 | 59.5869
9.1 |  | 3.2 | 7 | 11.18 |  | 3.8 | 8.48 | 12.8167 |  | 5.2 | 11.85 | 16.5 |  | 8.7 | 31.4126 | 35.55 |  | 12.74 | 56.6 | 58.6149
9.2 |  | 3.2 | 6.9 | 11.1 |  | 3.8 | 8.3667 | 12.6 |  | 5.1 | 11.5 | 16.1571 |  | 8.6 | 26.8908 | 31.5 |  | 12.5 | 55.3 | 57.5754
9.3 |  | 3.2 | 6.9 | 11.1 |  | 3.7 | 8.2 | 12.5 |  | 5 | 11.2333 | 15.95 |  | 8.38 | 23.3093 | 28.3 |  | 12.4 | 53.6 | 56.3883
9.4 |  | 3.2 | 6.9 | 11 |  | 3.7 | 8.0833 | 12.3667 |  | 5 | 11.02 | 15.55 |  | 8.2333 | 20.6583 | 25.66 |  | 12.1286 | 51.575 | 54.1546
9.5 |  | 3.1 | 6.825 | 10.95 |  | 3.6 | 8 | 12.3 |  | 4.9 | 10.7667 | 15.375 |  | 8.1 | 18.9325 | 24.15 |  | 11.82 | 48.6333 | 51.7112
9.6 |  | 3.1 | 6.8 | 10.9 |  | 3.6 | 7.8833 | 12.1 |  | 4.9 | 10.55 | 15.2 |  | 8 | 17.707 | 23.1 |  | 11.525 | 44.9 | 48.3793
9.7 |  | 3.1 | 6.8 | 10.9 |  | 3.6 | 7.8 | 12 |  | 4.8 | 10.35 | 15 |  | 7.9 | 16.9318 | 22.4 |  | 11.4 | 40.5 | 44.6609
9.8 |  | 3.1 | 6.8 | 10.9 |  | 3.5 | 7.7 | 11.96 |  | 4.7 | 10.1 | 14.8 |  | 7.8 | 16.4566 | 21.85 |  | 11.2143 | 36.25 | 40.6111
9.9 |  | 3 | 6.8 | 10.9 |  | 3.5 | 7.7 | 11.9 |  | 4.6 | 10.1 | 14.6 |  | 7.7 | 16.0564 | 21.5 |  | 11.06 | 32.1 | 36.9841
10 |  | 3 | 6.8 | 10.8667 |  | 3.4 | 7.6 | 11.9 |  | 4.5 | 9.9 | 14.4571 |  | 7.6 | 15.6562 | 21.12 |  | 10.9 | 28.5 | 33.8569
10.1 |  |  | 6.8 | 10.8 |  | 3.4 | 7.6 | 11.8 |  | 4.5 | 9.725 | 14.3143 |  | 7.525 | 15.3061 | 20.8143 |  | 10.9 | 25.45 | 31.1724
10.2 |  |  | 6.7 | 10.8 |  | 3.4 | 7.5 | 11.7 |  | 4.4 | 9.675 | 14.1 |  | 7.4 | 15.106 | 20.4429 |  | 10.5857 | 23.2714 | 29.1875
10.3 |  |  | 6.7 | 10.8 |  |  | 7.4 | 11.6 |  | 4.4 | 9.475 | 14 |  | 7.325 | 14.8559 | 20.2714 |  | 10.375 | 21.86 | 27.988
10.4 |  |  | 6.6 | 10.7 |  |  | 7.35 | 11.6 |  | 4.3 | 9.375 | 13.9 |  | 7.2 | 14.5458 | 19.9 |  | 10.2 | 20.9 | 27.055
10.5 |  |  | 6.6 | 10.625 |  |  | 7.3 | 11.5 |  | 4.3 | 9.22 | 13.7 |  | 7.15 | 14.4058 | 19.7 |  | 10.1 | 20.24 | 26.4386
10.6 |  |  | 6.5 | 10.5 |  |  | 7.2 | 11.4 |  |  | 9.1 | 13.6 |  | 7 | 14.1056 | 19.4 |  | 10.08 | 19.65 | 25.9888
10.7 |  |  | 6.4 | 10.5 |  |  | 7.2 | 11.3 |  |  | 9.025 | 13.4857 |  | 6.9 | 13.8556 | 19.2 |  | 10 | 19.42 | 25.649
10.8 |  |  | 6.4 | 10.4 |  |  | 7.0667 | 11.25 |  |  | 8.9 | 13.4 |  | 6.9 | 13.6555 | 19 |  | 10 | 19.12 | 25.2392
10.9 |  |  | 6.3 | 10.4 |  |  | 7 | 11.2 |  |  | 8.8 | 13.3 |  | 6.775 | 13.5054 | 18.78 |  | 9.875 | 18.775 | 24.8643
11 |  |  | 6.2571 | 10.3 |  |  | 6.9 | 11.1 |  |  | 8.7 | 13.14 |  |  | 13.3053 | 18.5286 |  | 9.7 | 18.5 | 24.5609
11.1 |  |  | 6.2 | 10.2 |  |  | 6.825 | 11 |  |  | 8.6 | 13.1 |  |  | 13.1052 | 18.3 |  | 9.64 | 18.05 | 24.3467
11.2 |  |  | 6.1 | 10.1 |  |  | 6.8 | 10.9 |  |  | 8.5 | 12.8778 |  |  | 12.9052 | 18.175 |  | 9.425 | 17.8 | 24.1097
11.3 |  |  | 6.1 | 10 |  |  | 6.7333 | 10.9 |  |  | 8.4 | 12.8 |  |  | 12.6384 | 17.9571 |  | 9.4 | 17.6333 | 23.8498
11.4 |  |  | 5.9833 | 9.95 |  |  | 6.65 | 10.7 |  |  | 8.3 | 12.7 |  |  | 12.53 | 17.76 |  | 9.3 | 17.3333 | 23.5399
11.5 |  |  | 5.9 | 9.9 |  |  | 6.6 | 10.6 |  |  | 8.2 | 12.5625 |  |  | 12.405 | 17.6125 |  |  | 17.2 | 23.29
11.6 |  |  | 5.82 | 9.74 |  |  | 6.5 | 10.55 |  |  | 8.0625 | 12.5 |  |  | 12.3049 | 17.4 |  |  | 16.9571 | 23.0501
11.7 |  |  | 5.7667 | 9.7 |  |  | 6.4 | 10.4625 |  |  | 8 | 12.3833 |  |  | 12.1048 | 17.14 |  |  | 16.65 | 22.7402
11.8 |  |  | 5.7 | 9.65 |  |  | 6.3 | 10.3 |  |  | 7.9 | 12.3 |  |  | 11.9173 | 17.02 |  |  | 16.375 | 22.4904
11.9 |  |  | 5.64 | 9.6 |  |  | 6.22 | 10.3 |  |  | 7.8 | 12.1 |  |  | 11.8047 | 16.8571 |  |  | 16.16 | 22.3104
12 |  |  | 5.6 | 9.4286 |  |  | 6.2 | 10.16 |  |  | 7.8 | 11.94 |  |  | 11.6713 | 16.7333 |  |  | 16 | 22.1155
12.1 |  |  | 5.5 | 9.4 |  |  | 6.1 | 10.05 |  |  | 7.6 | 11.9 |  |  | 11.6046 | 16.6 |  |  | 15.975 | 21.9156
12.2 |  |  | 5.5 | 9.3 |  |  | 6 | 10 |  |  | 7.5 | 11.8 |  |  | 11.4474 | 16.48 |  |  | 15.8 | 21.7157
12.3 |  |  | 5.4 | 9.3 |  |  | 5.9571 | 9.875 |  |  | 7.4125 | 11.6 |  |  | 11.1845 | 16.25 |  |  | 15.48 | 21.5907
12.4 |  |  | 5.35 | 9.2 |  |  | 5.9 | 9.8 |  |  | 7.4 | 11.54 |  |  | 11.0544 | 16.1 |  |  | 15.3 | 21.4908
12.5 |  |  | 5.3 | 9.2 |  |  | 5.8 | 9.8 |  |  | 7.2571 | 11.5 |  |  | 10.9044 | 15.95 |  |  | 15.08 | 21.071
12.6 | [6 empty] | 5.8 | 9.65 |  |  | 7.1667 | 11.3167 |  |  | 10.7376 | 15.8 |  |  | 14.78 | 20.8577
12.7 | [6 empty] | 5.74 | 9.6 |  |  | 7.05 | 11.25 |  |  | 10.6042 | 15.7 |  |  | 14.65 | 20.5912
12.8 | [6 empty] | 5.7 | 9.5125 |  |  | 6.92 | 11.1833 |  |  | 10.5792 | 15.5 |  |  | 14.55 | 20.4912
12.9 | [10 empty] | 6.875 | 11.06 |  |  | 10.3441 | 15.3 |  |  | 14.28 | 20.3056
13 | [10 empty] | 6.72 | 10.95 |  |  | 10.2841 | 15.175 |  |  | 14.1 | 20.0664
13.1 | [10 empty] | 6.7 | 10.86 |  |  | 10.129 | 15.025 |  |  | 14 | 19.9315
13.2 | [10 empty] | 6.6 | 10.7429 |  |  | 10.054 | 14.925 |  |  | 13.84 | 19.7915
13.3 | [14 empty] | 9.8039 | 14.9 |  |  | 13.7 | 19.6416
13.4 | [14 empty] | 9.7039 | 14.6667 |  |  | 13.6 | 19.4416
13.5 | [14 empty] | 9.6639 | 14.56 |  |  | 13.5 | 19.1918
13.6 | [14 empty] | 9.5038 | 14.4 |  |  | 13.25 | 18.9919
13.7 | [14 empty] | 9.4038 | 14.2 |  |  | 13.2 | 18.8169
13.8 | [14 empty] | 9.3037 | 14.1 |  |  | 13.075 | 18.6919
13.9 | [14 empty] | 9.1236 | 13.9667 |  |  | 12.825 | 18.4921
14 | [18 empty] | 12.7 | 18.3921
14.1 | [18 empty] | 12.6 | 18.2172
14.2 | [18 empty] | 12.48 | 18.1422
14.3 | [18 empty] | 12.3 | 17.8923
14.4 | [18 empty] | 12.1 | 17.6924
14.5 | [18 empty] | 12 | 17.4758
14.6 | [18 empty] | 11.9 | 17.3925
14.7 | [18 empty] | 11.925 | 17.1926

## Wedge 60 Longitudinal Profiles
Longitudinal Profiles
Wedge 60
SSD=100cm
Field Size=40x15cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-32.9 |  |  |  |  | 11.9
-32.8 |  |  |  |  | 12.0375
-32.7 |  |  |  |  | 12.2
-32.6 |  |  |  |  | 12.38
-32.5 |  |  |  |  | 12.5
-32.4 |  |  |  |  | 12.6
-32.3 |  |  |  |  | 12.7
-32.2 |  |  |  |  | 12.8
-32.1 |  |  |  |  | 12.8714
-32 |  |  |  |  | 13
-31.9 |  |  |  |  | 13.125
-31.8 |  |  |  |  | 13.2571
-31.7 |  |  |  |  | 13.3875
-31.6 |  |  |  |  | 13.52
-31.5 |  |  |  |  | 13.6857
-31.4 |  |  |  |  | 13.8143
-31.3 |  |  |  |  | 13.98
-31.2 |  |  |  |  | 14.0571
-31.1 |  |  |  |  | 14.3
-31 |  |  |  |  | 14.3714
-30.9 |  |  |  | 9.8 | 14.5125
-30.8 |  |  |  | 9.86 | 14.6429
-30.7 |  |  |  | 10 | 14.8143
-30.6 |  |  |  | 10.1 | 14.95
-30.5 |  |  |  | 10.175 | 15.075
-30.4 |  |  |  | 10.3 | 15.225
-30.3 |  |  |  | 10.4375 | 15.4429
-30.2 |  |  |  | 10.5 | 15.65
-30.1 |  |  |  | 10.6 | 15.8125
-30 |  |  |  | 10.7 | 15.9429
-29.9 |  |  |  | 10.8286 | 16.075
-29.8 |  |  |  | 10.9 | 16.3
-29.7 |  |  |  | 11 | 16.475
-29.6 |  |  |  | 11.1286 | 16.6714
-29.5 |  |  |  | 11.2625 | 16.8
-29.4 |  |  |  | 11.3 | 17.0375
-29.3 |  |  |  | 11.425 | 17.2429
-29.2 |  |  |  | 11.5571 | 17.5
-29.1 |  |  |  | 11.6875 | 17.65
-29 |  |  |  | 11.925 | 17.9
-28.9 |  |  | 7.8 | 12.0571 | 18.2286
-28.8 |  |  | 7.8 | 12.1875 | 18.3625
-28.7 |  |  | 7.9 | 12.22 | 18.62
-28.6 |  |  | 7.94 | 12.4 | 18.8167
-28.5 |  |  | 8.04 | 12.52 | 19.0667
-28.4 |  |  | 8.1 | 12.64 | 19.28
-28.3 |  |  | 8.2 | 12.8167 | 19.56
-28.2 |  |  | 8.2 | 13 | 19.78
-28.1 |  |  | 8.3 | 13.1 | 19.98
-28 |  |  | 8.4 | 13.2 | 20.2714
-27.9 |  | 6.97 | 8.4 | 13.4 | 20.6
-27.8 |  | 7 | 8.5 | 13.6 | 20.85
-27.7 |  | 7.0714 | 8.6 | 13.7667 | 21.1143
-27.6 |  | 7.1 | 8.7 | 13.96 | 21.5
-27.5 |  | 7.15 | 8.7 | 14.06 | 21.675
-27.4 |  | 7.2 | 8.8 | 14.26 | 22.0143
-27.3 |  | 7.2 | 8.9 | 14.46 | 22.45
-27.2 | 6.7 | 7.2429 | 9.02 | 14.72 | 22.8286
-27.1 | 6.7 | 7.3 | 9.1 | 14.9 | 23.15
-27 | 6.7 | 7.4 | 9.2 | 15.06 | 23.5571
-26.9 | 6.8 | 7.4625 | 9.28 | 15.26 | 23.9714
-26.8 | 6.8 | 7.5 | 9.36 | 15.46 | 24.6
-26.7 | 6.8167 | 7.5 | 9.5 | 15.82 | 25.3143
-26.6 | 6.9 | 7.5571 | 9.56 | 16.06 | 26.2429
-26.5 | 6.9 | 7.6 | 9.65 | 16.2429 | 27.65
-26.4 | 6.9 | 7.7 | 9.7875 | 16.4714 | 29.325
-26.3 | 6.9 | 7.7571 | 9.825 | 16.7429 | 31.8833
-26.2 | 7 | 7.8 | 9.9571 | 16.9714 | 35.05
-26.1 | 7 | 7.8143 | 10.0875 | 17.18 | 38.6429
-26 | 7.0833 | 7.9 | 10.125 | 17.4667 | 43.1
-25.9 | 7.1 | 7.9857 | 10.2571 | 17.725 | 47.45
-25.8 | 7.1143 | 8.0143 | 10.3875 | 18.0286 | 52.625
-25.7 | 7.2 | 8.1 | 10.525 | 18.35 | 56.875
-25.6 | 7.2 | 8.1857 | 10.6571 | 18.6 | 61.0714
-25.5 | 7.2 | 8.2143 | 10.7 | 18.875 | 64.8
-25.4 | 7.2429 | 8.3 | 10.825 | 19.2857 | 67.2125
-25.3 | 7.3 | 8.3857 | 11.0429 | 19.6 | 69.4
-25.2 | 7.4 | 8.4143 | 11.1 | 19.975 | 71
-25.1 | 7.4 | 8.5 | 11.3125 | 20.3143 | 71.925
-25 | 7.4 | 8.575 | 11.4857 | 20.8 | 72.6571
-24.9 | 7.5 | 8.6125 | 11.675 | 21.225 | 73.3
-24.8 | 7.5 | 8.7375 | 11.8125 | 21.7714 | 73.85
-24.7 | 7.5714 | 8.8 | 11.9429 | 22.6 | 74.2875
-24.6 | 7.6125 | 8.9 | 12.15 | 23.675 | 74.8
-24.5 | 7.7 | 8.9133 | 12.4125 | 25.3 | 75.3143
-24.4 | 7.7 | 8.98 | 12.5857 | 27.5625 | 75.7667
-24.3 | 7.8 | 9.0933 | 12.85 | 30.65 | 76.0571
-24.2 | 7.8 | 9.2133 | 13.025 | 34.3714 | 76.4
-24.1 | 7.9 | 9.28 | 13.2857 | 38.9375 | 76.6875
-24 | 7.9 | 9.3933 | 13.52 | 44.225 | 77.2
-23.9 | 7.9625 | 9.5286 | 13.75 | 49.9429 | 77.4286
-23.8 | 8 | 9.6714 | 14 | 55.6375 | 77.6875
-23.7 | 8 | 9.8 | 14.2286 | 60.875 | 78
-23.6 | 8.0625 | 9.9375 | 14.4875 | 65.6857 | 78.2857
-23.5 | 8.1 | 10.0714 | 14.8 | 69.375 | 78.625
-23.4 | 8.1286 | 10.2 | 15.0857 | 72.3 | 78.875
-23.3 | 8.2 | 10.3571 | 15.425 | 74.3714 | 79.275
-23.2 | 8.2875 | 10.5625 | 15.7 | 75.85 | 79.6143
-23.1 | 8.3286 | 10.6875 | 16.0857 | 76.75 | 79.9625
-23 | 8.4 | 10.85 | 16.4875 | 77.4286 | 80.325
-22.9 | 8.5 | 11.0571 | 17 | 78.0375 | 80.5143
-22.8 | 8.6375 | 11.275 | 17.6429 | 78.3625 | 80.6875
-22.7 | 8.7625 | 11.45 | 18.5 | 78.9714 | 80.9286
-22.6 | 8.8 | 11.7143 | 19.7 | 79.425 | 81.2
-22.5 | 8.9286 | 11.8875 | 21.5857 | 79.8375 | 81.45
-22.4 | 9.0625 | 12.15 | 24.3 | 80.1857 | 81.725
-22.3 | 9.1875 | 12.4143 | 28 | 80.525 | 82.0286
-22.2 | 9.325 | 12.675 | 32.7857 | 80.95 | 82.35
-22.1 | 9.4571 | 13.0333 | 38.3333 | 81.3857 | 82.525
-22 | 9.5875 | 13.3667 | 44.6286 | 81.65 | 82.76
-21.9 | 9.7 | 13.8 | 51.6429 | 82.0125 | 83.0125
-21.8 | 9.8 | 14.35 | 58.75 | 82.2125 | 83.1429
-21.7 | 9.9857 | 15.2 | 65.2571 | 82.6143 | 83.4571
-21.6 | 10.1286 | 16.5333 | 70.9286 | 82.8 | 83.7125
-21.5 | 10.35 | 18.5 | 75.55 | 83.1875 | 84.0429
-21.4 | 10.4857 | 21.4875 | 78.7143 | 83.5 | 84.2333
-21.3 | 10.7286 | 25.6 | 80.9286 | 83.8143 | 84.5429
-21.2 | 11 | 30.9857 | 82.1 | 83.9875 | 84.8125
-21.1 | 11.325 | 38.0429 | 83.0143 | 84.3 | 84.9429
-21 | 11.8625 | 45.91 | 83.7714 | 84.7 | 85.15
-20.9 | 12.6857 | 53.8 | 84.4 | 84.8667 | 85.2286
-20.8 | 14.1 | 61.6625 | 84.8571 | 85.1333 | 85.55
-20.7 | 16.3571 | 69.1 | 85.3429 | 85.42 | 85.7857
-20.6 | 20.2 | 75.2273 | 85.7 | 85.6 | 86.0143
-20.5 | 25.125 | 79.875 | 86 | 85.78 | 86.2
-20.4 | 32.1375 | 83.22 | 86.4 | 86.06 | 86.4714
-20.3 | 40 | 85.3 | 86.6 | 86.36 | 86.7125
-20.2 | 48.7 | 86.5875 | 86.8833 | 86.5571 | 86.875
-20.1 | 58.4625 | 87.38 | 87.2 | 86.825 | 87.0714
-20 | 67.0125 | 87.95 | 87.7 | 87.0286 | 87.2
-19.9 | 74.9143 | 88.44 | 87.88 | 87.1625 | 87.375
-19.8 | 80.9 | 88.825 | 88.08 | 87.4 | 87.6429
-19.7 | 85.3 | 89.12 | 88.38 | 87.6571 | 87.8
-19.6 | 87.9429 | 89.4857 | 88.76 | 87.925 | 87.975
-19.5 | 89.4 | 89.75 | 89 | 88.0875 | 88.1714
-19.4 | 90.3875 | 89.9375 | 89.24 | 88.35 | 88.3
-19.3 | 90.9143 | 90.2857 | 89.3667 | 88.55 | 88.4286
-19.2 | 91.2 | 90.55 | 89.5167 | 88.7714 | 88.5625
-19.1 | 91.475 | 90.7 | 89.7 | 89.0125 | 88.8
-19 | 91.7429 | 90.8571 | 90.0667 | 89.1375 | 89.0571
-18.9 | 91.9 | 91.0625 | 90.2 | 89.2714 | 89.2625
-18.8 | 92.175 | 91.3 | 90.38 | 89.5 | 89.3875
-18.7 | 92.3625 | 91.4857 | 90.58 | 89.75 | 89.525
-18.6 | 92.5 | 91.825 | 90.96 | 90 | 89.7143
-18.5 | 92.7375 | 92 | 91.18 | 90.1286 | 89.8875
-18.4 | 92.8 | 92.1286 | 91.2667 | 90.3429 | 90.1
-18.3 | 92.9 | 92.325 | 91.46 | 90.56 | 90.2
-18.2 | 93.1 | 92.5 | 91.66 | 90.66 | 90.2667
-18.1 | 93.1625 | 92.6286 | 91.95 | 90.85 | 90.46
-18 | 93.2875 | 92.7625 | 92.2 | 91 | 90.65
-17.9 | 93.55 | 93 | 92.3 | 91.14 | 90.7875
-17.8 | 93.7571 | 93.1286 | 92.4 | 91.34 | 91.0333
-17.7 | 93.8167 | 93.325 | 92.5 | 91.54 | 91.2714
-17.6 | 93.9 | 93.5 | 92.675 | 91.7 | 91.3143
-17.5 | 94.0125 | 93.625 | 92.8714 | 91.78 | 91.45
-17.4 | 94.1429 | 93.7571 | 92.9 | 91.94 | 91.5
-17.3 | 94.35 | 93.8875 | 93.1571 | 92.04 | 91.5143
-17.2 | 94.4125 | 94.025 | 93.425 | 92.2286 | 91.7
-17.1 | 94.5375 | 94.1571 | 93.6 | 92.4 | 91.875
-17 | 94.6714 | 94.2875 | 93.7286 | 92.54 | 92.0125
-16.9 | 94.7 | 94.4143 | 93.8625 | 92.7 | 92.1
-16.8 | 94.8286 | 94.55 | 94 | 92.85 | 92.175
-16.7 | 94.9625 | 94.675 | 94.1286 | 92.9875 | 92.3
-16.6 | 95.1 | 94.8143 | 94.2625 | 93 | 92.4
-16.5 | 95.2286 | 94.95 | 94.3875 | 93.1 | 92.46
-16.4 | 95.3 | 95.1125 | 94.44 | 93.2857 | 92.6
-16.3 | 95.3778 | 95.2 | 94.62 | 93.4286 | 92.74
-16.2 | 95.5143 | 95.375 | 94.8 | 93.65 | 92.84
-16.1 | 95.6615 | 95.5143 | 94.9 | 93.7 | 93.04
-16 | 95.8222 | 95.6571 | 95 | 93.9 | 93.24
-15.9 | 96.0182 | 95.7 | 95.1833 | 93.9 | 93.4
-15.8 | 96.1083 | 95.8125 | 95.24 | 94 | 93.48
-15.7 | 96.1917 | 95.9 | 95.34 | 94.1 | 93.5667
-15.6 | 96.32 | 95.9333 | 95.4286 | 94.2 | 93.68
-15.5 | 96.4333 | 96.0889 | 95.6 | 94.34 | 93.76
-15.4 | 96.5333 | 96.2455 | 95.7 | 94.54 | 93.8
-15.3 | 96.7 | 96.3333 | 95.8 | 94.6467 | 93.96
-15.2 | 96.825 | 96.4286 | 95.96 | 94.725 | 94.06
-15.1 | 96.9444 | 96.5 | 96.0143 | 94.85 | 94.16
-15 | 97.0455 | 96.6 | 96.1 | 94.9 | 94.36
-14.9 | 97.1 | 96.7083 | 96.2 | 94.9545 | 94.46
-14.8 | 97.1625 | 96.7917 | 96.2429 | 95.1182 | 94.5
-14.7 | 97.2538 | 96.9 | 96.375 | 95.2133 | 94.6
-14.6 | 97.3444 | 97.0154 | 96.5 | 95.3467 | 94.6333
-14.5 | 97.4455 | 97.0923 | 96.5 | 95.4429 | 94.8143
-14.4 | 97.5444 | 97.1 | 96.5714 | 95.5 | 94.9571
-14.3 | 97.6 | 97.1 | 96.6 | 95.5067 | 95.1
-14.2 | 97.6 | 97.2067 | 96.7 | 95.5733 | 95.1857
-14.1 | 97.7 | 97.3333 | 96.7 | 95.65 | 95.2
-14 | 97.7286 | 97.4 | 96.8375 | 95.7571 | 95.2571
-13.9 | 97.8 | 97.4417 | 96.9625 | 95.8778 | 95.3875
-13.8 | 97.8875 | 97.52 | 97 | 95.9125 | 95.5
-13.7 | 97.9 | 97.5867 | 97.1 | 96.0429 | 95.55
-13.6 | 97.9286 | 97.6 | 97.1625 | 96.175 | 95.6
-13.5 | 98.0625 | 97.6214 | 97.2 | 96.3 | 95.7
-13.4 | 98.1 | 97.6929 | 97.325 | 96.3 | 95.9
-13.3 | 98.1125 | 97.7111 | 97.4 | 96.375 | 95.9
-13.2 | 98.225 | 97.8222 | 97.4 | 96.5125 | 95.95
-13.1 | 98.3 | 97.9 | 97.425 | 96.675 | 96.0267
-13 | 98.3143 | 97.9133 | 97.5 | 96.8 | 96.0933
-12.9 | 98.3857 | 97.98 | 97.5 | 96.9 | 96.16
-12.8 | 98.4 | 98.0467 | 97.6333 | 96.9286 | 96.2533
-12.7 | 98.4333 | 98.1286 | 97.7 | 97.0625 | 96.3867
-12.6 | 98.5 | 98.2714 | 97.7833 | 97.1 | 96.52
-12.5 | 98.5 | 98.3 | 97.9 | 97.1375 | 96.6
-12.4 | 98.5467 | 98.4 | 98 | 97.2 | 96.6
-12.3 | 98.6 | 98.4 | 98.1 | 97.2 | 96.6562
-12.2 | 98.6 | 98.5 | 98.1333 | 97.3 | 96.72
-12.1 | 98.6467 | 98.5 | 98.2 | 97.3625 | 96.7867
-12 | 98.7133 | 98.5273 | 98.2571 | 97.4 | 96.9333
-11.9 | 98.78 | 98.6133 | 98.3 | 97.525 | 97.05
-11.8 | 98.8583 | 98.68 | 98.4 | 97.6 | 97.1
-11.7 | 98.9714 | 98.7 | 98.475 | 97.6875 | 97.1
-11.6 | 99 | 98.7 | 98.5 | 97.8 | 97.13
-11.5 | 99.1 | 98.7133 | 98.5 | 97.8 | 97.2
-11.4 | 99.1625 | 98.8467 | 98.5 | 97.8 | 97.2714
-11.3 | 99.2875 | 98.94 | 98.6 | 97.8125 | 97.4
-11.2 | 99.3 | 99.0067 | 98.7 | 97.9 | 97.5
-11.1 | 99.3 | 99.0733 | 98.7077 | 97.975 | 97.5625
-11 | 99.3875 | 99.1 | 98.7846 | 98.1 | 97.6
-10.9 | 99.4 | 99.1 | 98.8 | 98.1286 | 97.6
-10.8 | 99.4 | 99.1667 | 98.8 | 98.2 | 97.6
-10.7 | 99.4 | 99.2 | 98.9 | 98.3 | 97.7
-10.6 | 99.4571 | 99.2 | 98.9 | 98.3286 | 97.8
-10.5 | 99.5 | 99.2769 | 98.96 | 98.4 | 97.9143
-10.4 | 99.5 | 99.3583 | 99.04 | 98.47 | 98.175
-10.3 | 99.6 | 99.4 | 99.1067 | 98.5875 | 98.3143
-10.2 | 99.6 | 99.4 | 99.1733 | 98.66 | 98.4
-10.1 | 99.7 | 99.4714 | 99.24 | 98.7 | 98.4
-10 | 99.7 | 99.5 | 99.3 | 98.7 | 98.4
-9.9 | 99.8 | 99.6 | 99.3 | 98.7533 | 98.4
-9.8 | 99.8 | 99.6 | 99.3417 | 98.82 | 98.4
-9.7 | 99.8714 | 99.6231 | 99.4 | 98.8867 | 98.4583
-9.6 | 99.9125 | 99.7 | 99.4167 | 98.9533 | 98.5714
-9.5 | 100 | 99.7 | 99.5 | 99.0214 | 98.6727
-9.4 | 100 | 99.7 | 99.5 | 99.0929 | 98.7
-9.3 | 100 | 99.775 | 99.5 | 99.1 | 98.7083
-9.2 | 100 | 99.8 | 99.575 | 99.1909 | 98.7917
-9.1 | 100 | 99.8 | 99.6091 | 99.3 | 98.8
-9 | 100.1 | 99.8625 | 99.7 | 99.3 | 98.8889
-8.9 | 100.1625 | 99.9 | 99.7625 | 99.3714 | 98.925
-8.8 | 100.2 | 100 | 99.8 | 99.4 | 99
-8.7 | 100.3 | 100 | 99.8154 | 99.4 | 99.0875
-8.6 | 100.3 | 100.1 | 99.8923 | 99.4 | 99.1
-8.5 | 100.3 | 100.1 | 100 | 99.4 | 99.0429
-8.4 | 100.4 | 100.1625 | 100.1 | 99.5286 | 99.0875
-8.3 | 100.4 | 100.2 | 100.1625 | 99.6625 | 99.1
-8.2 | 100.4 | 100.3 | 100.2 | 99.7 | 99.1857
-8.1 | 100.4 | 100.3556 | 100.22 | 99.7 | 99.3333
-8 | 100.4 | 100.4 | 100.3 | 99.725 | 99.4
-7.9 | 100.48 | 100.4 | 100.35 | 99.8 | 99.4
-7.8 | 100.525 | 100.4 | 100.4 | 99.84 | 99.4
-7.7 | 100.6 | 100.4 | 100.4 | 99.9 | 99.4
-7.6 | 100.6875 | 100.4 | 100.4 | 99.9286 | 99.5
-7.5 | 100.7 | 100.45 | 100.4 | 100 | 99.5
-7.4 | 100.7 | 100.5 | 100.4 | 100 | 99.6
-7.3 | 100.7857 | 100.6 | 100.4 | 100 | 99.7
-7.2 | 100.7143 | 100.6 | 100.4 | 100 | 99.7
-7.1 | 100.8 | 100.6 | 100.4 | 100 | 99.75
-7 | 100.8857 | 100.6 | 100.4 | 100.1 | 99.8
-6.9 | 100.9 | 100.6 | 100.4143 | 100.1 | 99.8
-6.8 | 100.9429 | 100.6 | 100.5 | 100.1 | 99.8
-6.7 | 101 | 100.7 | 100.5 | 100.125 | 99.8
-6.6 | 101 | 100.7 | 100.5 | 100.2 | 99.8
-6.5 | 101.0429 | 100.7 | 100.5429 | 100.2875 | 99.9
-6.4 | 101.1 | 100.7 | 100.6 | 100.3 | 99.9
-6.3 | 101.0875 | 100.7 | 100.7 | 100.25 | 99.9
-6.2 | 101 | 100.7 | 100.7 | 100.2 | 99.9
-6.1 | 100.9 | 100.72 | 100.7 | 100.3 | 99.9
-6 | 100.9714 | 100.8 | 100.7 | 100.2571 | 99.9
-5.9 | 101 | 100.8714 | 100.7375 | 100.275 | 99.9
-5.8 | 101 | 100.9 | 100.8 | 100.4 | 100
-5.7 | 101.0714 | 100.9 | 100.9 | 100.4 | 100
-5.6 | 101.1 | 100.9 | 100.9 | 100.475 | 100
-5.5 | 101.1286 | 100.9 | 100.9 | 100.5 | 100.0875
-5.4 | 101.2 | 100.9 | 100.8 | 100.5 | 100.125
-5.3 | 101.2 | 100.9 | 100.8 | 100.425 | 100.1429
-5.2 | 101.2 | 100.9 | 100.8 | 100.4 | 100.1
-5.1 | 101.2 | 100.9 | 100.8 | 100.4 | 100.2
-5 | 101.2 | 100.9 | 100.9 | 100.475 | 100.2
-4.9 | 101.1 | 100.9 | 100.9 | 100.5 | 100.1
-4.8 | 101.1 | 100.9 | 100.9 | 100.5 | 100.1667
-4.7 | 101.1 | 100.9 | 100.9 | 100.5 | 100.3
-4.6 | 101.1 | 100.9 | 100.9 | 100.5 | 100.3
-4.5 | 101.1 | 100.8 | 100.8286 | 100.5 | 100.3
-4.4 | 101 | 100.8 | 100.8 | 100.4833 | 100.3
-4.3 | 101 | 100.8 | 100.8 | 100.4 | 100.3
-4.2 | 101 | 100.8 | 100.8 | 100.4 | 100.3
-4.1 | 101 | 100.8 | 100.8 | 100.5 | 100.3
-4 | 101 | 100.8 | 100.8 | 100.5 | 100.3
-3.9 | 101 | 100.7857 | 100.8 | 100.5 | 100.3
-3.8 | 101 | 100.7 | 100.8 | 100.5 | 100.4
-3.7 | 101 | 100.7 | 100.7 | 100.5 | 100.4
-3.6 | 101 | 100.7 | 100.7 | 100.5 | 100.4
-3.5 | 100.9857 | 100.7 | 100.7 | 100.5 | 100.4
-3.4 | 100.9 | 100.7 | 100.7 | 100.525 | 100.4
-3.3 | 100.8333 | 100.625 | 100.7 | 100.55 | 100.3
-3.2 | 100.8636 | 100.6 | 100.7 | 100.5 | 100.3
-3.1 | 100.9 | 100.6 | 100.7 | 100.5 | 100.3
-3 | 100.8769 | 100.6 | 100.7 | 100.5 | 100.3
-2.9 | 100.8 | 100.6 | 100.7 | 100.4857 | 100.3
-2.8 | 100.8 | 100.6 | 100.7 | 100.4 | 100.3
-2.7 | 100.8 | 100.6 | 100.7 | 100.4 | 100.4
-2.6 | 100.7167 | 100.5 | 100.6429 | 100.4 | 100.4
-2.5 | 100.7 | 100.5 | 100.6 | 100.4 | 100.4
-2.4 | 100.8 | 100.5 | 100.55 | 100.4 | 100.32
-2.3 | 100.8 | 100.5 | 100.5 | 100.4 | 100.2
-2.2 | 100.8 | 100.5 | 100.5 | 100.4 | 100.2
-2.1 | 100.8 | 100.5 | 100.44 | 100.4 | 100.2
-2 | 100.8 | 100.5 | 100.45 | 100.5 | 100.1
-1.9 | 100.7125 | 100.5 | 100.5 | 100.425 | 100.1
-1.8 | 100.7 | 100.4 | 100.5 | 100.3 | 100.1
-1.7 | 100.65 | 100.4 | 100.5 | 100.3 | 100
-1.6 | 100.6 | 100.3 | 100.5 | 100.3 | 100.1
-1.5 | 100.5875 | 100.3 | 100.4625 | 100.2 | 100.1
-1.4 | 100.4571 | 100.3 | 100.4 | 100.2 | 100.2
-1.3 | 100.4 | 100.3 | 100.4 | 100.2 | 100.2
-1.2 | 100.3 | 100.3 | 100.3714 | 100.2 | 100.2
-1.1 | 100.2625 | 100.2125 | 100.3 | 100.2 | 100.1857
-1 | 100.2 | 100.2 | 100.3 | 100.2 | 100.15
-0.9 | 100.2 | 100.1429 | 100.2 | 100.1714 | 100.15
-0.8 | 100.2375 | 100.1 | 100.2 | 100.1 | 100.05
-0.7 | 100.3 | 100.1 | 100.2 | 100 | 99.9333
-0.6 | 100.2 | 100 | 100.1714 | 100.0143 | 99.9
-0.5 | 100.2 | 100 | 100.1 | 100.05 | 99.9
-0.4 | 100.2 | 100 | 100.1 | 100 | 99.9
-0.3 | 100.1125 | 100 | 100.1 | 100 | 99.9438
-0.2 | 100.075 | 100 | 100.0429 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100.0429 | 100 | 100 | 100 | 100
0.3 | 100.1 | 100 | 100.1 | 100 | 99.9438
0.4 | 100.2 | 100 | 100.1 | 100 | 99.9
0.5 | 100.2 | 100 | 100.1 | 100.0857 | 99.9
0.6 | 100.2714 | 100 | 100.2 | 100 | 99.9
0.7 | 100.3 | 100.1 | 100.2 | 100.0375 | 99.9333
0.8 | 100.2714 | 100.1 | 100.2 | 100.1 | 100.05
0.9 | 100.2 | 100.1125 | 100.2 | 100.1 | 100.15
1 | 100.2 | 100.2 | 100.2375 | 100.2 | 100.15
1.1 | 100.225 | 100.2 | 100.3 | 100.2 | 100.1857
1.2 | 100.3 | 100.3 | 100.4 | 100.2 | 100.2
1.3 | 100.3875 | 100.3 | 100.4 | 100.2 | 100.2
1.4 | 100.425 | 100.3 | 100.4 | 100.2 | 100.2
1.5 | 100.55 | 100.3 | 100.5 | 100.2 | 100.1
1.6 | 100.6 | 100.3 | 100.5 | 100.25 | 100.1
1.7 | 100.7 | 100.3625 | 100.5 | 100.3 | 100
1.8 | 100.7 | 100.4 | 100.5 | 100.3286 | 100.1
1.9 | 100.775 | 100.4625 | 100.5 | 100.4625 | 100.1
2 | 100.8 | 100.5 | 100.42 | 100.5 | 100.1
2.1 | 100.8 | 100.5 | 100.48 | 100.4429 | 100.2
2.2 | 100.8 | 100.5 | 100.5 | 100.4 | 100.2
2.3 | 100.8 | 100.5 | 100.5 | 100.4 | 100.2
2.4 | 100.8 | 100.5 | 100.5 | 100.4 | 100.32
2.5 | 100.7375 | 100.5 | 100.6 | 100.4 | 100.4
2.6 | 100.7 | 100.525 | 100.6 | 100.4 | 100.4
2.7 | 100.7727 | 100.6 | 100.7 | 100.4 | 100.4
2.8 | 100.8 | 100.6 | 100.7 | 100.4 | 100.3
2.9 | 100.8 | 100.6 | 100.7 | 100.4143 | 100.3
3 | 100.8429 | 100.6 | 100.7 | 100.5 | 100.3
3.1 | 100.9 | 100.6 | 100.7 | 100.5 | 100.3
3.2 | 100.8917 | 100.6 | 100.7 | 100.5 | 100.3
3.3 | 100.8083 | 100.6875 | 100.7 | 100.575 | 100.3
3.4 | 100.9 | 100.7 | 100.7 | 100.5 | 100.4
3.5 | 100.95 | 100.7 | 100.7 | 100.5 | 100.4
3.6 | 101 | 100.7 | 100.7 | 100.5 | 100.4
3.7 | 101 | 100.7 | 100.7 | 100.5 | 100.4
3.8 | 101 | 100.7143 | 100.8 | 100.5 | 100.4
3.9 | 101 | 100.8 | 100.8 | 100.5 | 100.3
4 | 101 | 100.8 | 100.8 | 100.5 | 100.3
4.1 | 101 | 100.8 | 100.8 | 100.5 | 100.3
4.2 | 101 | 100.8 | 100.8 | 100.4 | 100.3
4.3 | 101 | 100.8 | 100.8 | 100.4 | 100.3
4.4 | 101.05 | 100.8 | 100.8 | 100.5 | 100.3
4.5 | 101.1 | 100.8429 | 100.8625 | 100.5 | 100.3
4.6 | 101.1 | 100.9 | 100.9 | 100.5 | 100.3
4.7 | 101.1 | 100.9 | 100.9 | 100.5 | 100.3
4.8 | 101.1 | 100.9 | 100.9 | 100.5 | 100.1667
4.9 | 101.1 | 100.9 | 100.9 | 100.5 | 100.1
5 | 101.1667 | 100.9 | 100.8857 | 100.4429 | 100.2
5.1 | 101.2 | 100.9 | 100.8 | 100.4 | 100.2
5.2 | 101.2 | 100.9 | 100.8 | 100.4 | 100.1
5.3 | 101.2 | 100.9 | 100.8 | 100.4 | 100.1429
5.4 | 101.2 | 100.9 | 100.8375 | 100.4875 | 100.125
5.5 | 101.1625 | 100.9 | 100.9 | 100.5 | 100.0875
5.6 | 101.1 | 100.9 | 100.9 | 100.4429 | 100
5.7 | 101.1 | 100.9 | 100.8625 | 100.4 | 100
5.8 | 101 | 100.9 | 100.8 | 100.375 | 100
5.9 | 101 | 100.9 | 100.7 | 100.2429 | 99.9
6 | 101 | 100.8 | 100.7 | 100.2875 | 99.9
6.1 | 100.9 | 100.7 | 100.7 | 100.2857 | 99.9
6.2 | 101 | 100.7 | 100.7 | 100.2 | 99.9
6.3 | 101.0571 | 100.7 | 100.6714 | 100.2857 | 99.9
6.4 | 101.1 | 100.7 | 100.6 | 100.3 | 99.9
6.5 | 101.075 | 100.7 | 100.5125 | 100.2571 | 99.9
6.6 | 101 | 100.7 | 100.5 | 100.2 | 99.8
6.7 | 101 | 100.7 | 100.5 | 100.2 | 99.8
6.8 | 100.9 | 100.6 | 100.5 | 100.1 | 99.8
6.9 | 100.9 | 100.6 | 100.4 | 100.1 | 99.8
7 | 100.9 | 100.6 | 100.4 | 100.1 | 99.8
7.1 | 100.7857 | 100.6 | 100.4 | 100 | 99.75
7.2 | 100.7 | 100.6 | 100.4 | 100 | 99.7
7.3 | 100.7857 | 100.55 | 100.4 | 100 | 99.7
7.4 | 100.7 | 100.5 | 100.4 | 100 | 99.6
7.5 | 100.7 | 100.4857 | 100.4 | 100 | 99.5
7.6 | 100.7 | 100.4 | 100.4 | 100 | 99.5
7.7 | 100.5875 | 100.4 | 100.4 | 99.9 | 99.4
7.8 | 100.5 | 100.4 | 100.4 | 99.8 | 99.4
7.9 | 100.5 | 100.4 | 100.32 | 99.8 | 99.4
8 | 100.4 | 100.4 | 100.3 | 99.7 | 99.4
8.1 | 100.4 | 100.3 | 100.2875 | 99.7 | 99.3333
8.2 | 100.4 | 100.3 | 100.2 | 99.7 | 99.1857
8.3 | 100.4 | 100.2286 | 100.1286 | 99.6286 | 99.1
8.4 | 100.4 | 100.2 | 100.1 | 99.6 | 99.0875
8.5 | 100.3286 | 100.1625 | 99.9625 | 99.4714 | 99.0429
8.6 | 100.3 | 100.1 | 99.8643 | 99.4 | 99.1
8.7 | 100.2625 | 100 | 99.8 | 99.4 | 99.0875
8.8 | 100.2 | 100 | 99.7933 | 99.4 | 99
8.9 | 100.1 | 99.9286 | 99.7267 | 99.3375 | 98.925
9 | 100.1 | 99.9 | 99.6571 | 99.3 | 98.8889
9.1 | 100.0375 | 99.8 | 99.6 | 99.2857 | 98.8
9.2 | 100 | 99.8 | 99.5429 | 99.1636 | 98.7917
9.3 | 100 | 99.8 | 99.5 | 99.1 | 98.7083
9.4 | 100 | 99.7 | 99.5 | 99.1 | 98.7
9.5 | 100 | 99.7 | 99.5 | 99.0733 | 98.6727
9.6 | 99.9429 | 99.6385 | 99.4833 | 98.94 | 98.5714
9.7 | 99.9 | 99.6 | 99.4 | 98.8533 | 98.4583
9.8 | 99.8 | 99.6 | 99.3091 | 98.7867 | 98.4
9.9 | 99.8 | 99.6 | 99.3 | 98.72 | 98.4
10 | 99.7 | 99.5125 | 99.2733 | 98.7 | 98.4
10.1 | 99.7 | 99.5 | 99.2067 | 98.6875 | 98.4
10.2 | 99.6143 | 99.4333 | 99.2 | 98.625 | 98.4
10.3 | 99.6 | 99.4 | 99.1467 | 98.5455 | 98.3143
10.4 | 99.54 | 99.4 | 99.0133 | 98.5 | 98.175
10.5 | 99.5 | 99.3 | 98.925 | 98.425 | 97.9143
10.6 | 99.4 | 99.2333 | 98.9 | 98.4 | 97.8
10.7 | 99.4 | 99.1667 | 98.8636 | 98.3 | 97.7
10.8 | 99.4 | 99.1 | 98.8 | 98.3 | 97.6
10.9 | 99.4 | 99.1 | 98.8 | 98.2 | 97.6
11 | 99.4 | 99.0687 | 98.8 | 98.0667 | 97.6
11.1 | 99.3 | 99.0062 | 98.7727 | 97.9429 | 97.5625
11.2 | 99.3 | 99 | 98.675 | 97.9 | 97.5
11.3 | 99.3 | 98.9467 | 98.6 | 97.8 | 97.4
11.4 | 99.2 | 98.8133 | 98.5375 | 97.8 | 97.2714
11.5 | 99.0714 | 98.74 | 98.5 | 97.8 | 97.2
11.6 | 99 | 98.7 | 98.5 | 97.7857 | 97.13
11.7 | 99 | 98.6917 | 98.45 | 97.7 | 97.1
11.8 | 98.8857 | 98.6083 | 98.4 | 97.625 | 97.1
11.9 | 98.8143 | 98.54 | 98.3571 | 97.4875 | 97.05
12 | 98.7467 | 98.5 | 98.225 | 97.4 | 96.9333
12.1 | 98.68 | 98.5 | 98.2 | 97.325 | 96.7867
12.2 | 98.6133 | 98.5 | 98.1 | 97.3 | 96.72
12.3 | 98.6 | 98.43 | 98.1 | 97.2625 | 96.6562
12.4 | 98.58 | 98.4 | 98 | 97.2 | 96.6
12.5 | 98.5133 | 98.3375 | 97.94 | 97.1 | 96.6
12.6 | 98.4467 | 98.2222 | 97.8 | 97.1 | 96.52
12.7 | 98.4 | 98.1467 | 97.7 | 97.0286 | 96.3867
12.8 | 98.4 | 98.08 | 97.6 | 96.9 | 96.2533
12.9 | 98.4 | 98.0133 | 97.55 | 96.8625 | 96.16
13 | 98.3462 | 97.9467 | 97.5 | 96.8 | 96.0933
13.1 | 98.3 | 97.8667 | 97.4 | 96.7 | 96.0267
13.2 | 98.25 | 97.8 | 97.4 | 96.575 | 95.95
13.3 | 98.1556 | 97.75 | 97.4 | 96.4429 | 95.9
13.4 | 98.1 | 97.7 | 97.3 | 96.3125 | 95.9
13.5 | 98 | 97.6533 | 97.2625 | 96.275 | 95.7
13.6 | 97.9 | 97.5867 | 97.2 | 96.1429 | 95.6
13.7 | 97.9 | 97.52 | 97.1 | 96.0125 | 95.55
13.8 | 97.9 | 97.5 | 97 | 95.975 | 95.5
13.9 | 97.8 | 97.48 | 96.9286 | 95.85 | 95.3875
14 | 97.7625 | 97.375 | 96.8 | 95.725 | 95.2571
14.1 | 97.7 | 97.2667 | 96.7571 | 95.7 | 95.2
14.2 | 97.6222 | 97.1857 | 96.7 | 95.64 | 95.1857
14.3 | 97.6 | 97.1143 | 96.6714 | 95.5692 | 95.1
14.4 | 97.5727 | 97.1 | 96.6 | 95.4917 | 94.9571
14.5 | 97.4778 | 97.0333 | 96.5 | 95.4083 | 94.8143
14.6 | 97.375 | 96.9455 | 96.4714 | 95.34 | 94.6333
14.7 | 97.2818 | 96.9 | 96.3375 | 95.2333 | 94.6
14.8 | 97.1 | 96.8333 | 96.2125 | 95.0833 | 94.5
14.9 | 97.0727 | 96.68 | 96.175 | 95 | 94.46
15 | 97 | 96.525 | 96.1 | 94.9091 | 94.36
15.1 | 96.97 | 96.5 | 96 | 94.8438 | 94.16
15.2 | 96.825 | 96.4667 | 95.92 | 94.76 | 94.06
15.3 | 96.6444 | 96.3385 | 95.8 | 94.6267 | 93.96
15.4 | 96.5455 | 96.1889 | 95.76 | 94.4933 | 93.8
15.5 | 96.4286 | 96.0556 | 95.66 | 94.3455 | 93.76
15.6 | 96.2933 | 95.9 | 95.4833 | 94.1778 | 93.68
15.7 | 96.2267 | 95.86 | 95.4 | 94.07 | 93.5667
15.8 | 96.14 | 95.7429 | 95.2 | 94 | 93.48
15.9 | 96 | 95.7 | 95.15 | 93.9583 | 93.4
16 | 95.8556 | 95.6857 | 95 | 93.8667 | 93.24
16.1 | 95.7286 | 95.5 | 94.925 | 93.75 | 93.04
16.2 | 95.5846 | 95.325 | 94.78 | 93.6143 | 92.84
16.3 | 95.4308 | 95.2 | 94.68 | 93.4714 | 92.74
16.4 | 95.3 | 95.0857 | 94.5143 | 93.25 | 92.6
16.5 | 95.2625 | 94.9714 | 94.3625 | 93.1143 | 92.46
16.6 | 95.0571 | 94.8 | 94.2286 | 93 | 92.4
16.7 | 95 | 94.7125 | 94.1 | 92.95 | 92.3
16.8 | 94.8625 | 94.5857 | 93.9625 | 92.8143 | 92.175
16.9 | 94.7286 | 94.45 | 93.9 | 92.6571 | 92.1
17 | 94.6 | 94.3143 | 93.8 | 92.5 | 92.0125
17.1 | 94.5 | 94.1875 | 93.625 | 92.35 | 91.875
17.2 | 94.4429 | 94.0571 | 93.4286 | 92.2 | 91.7
17.3 | 94.3 | 93.925 | 93.2 | 92 | 91.5143
17.4 | 94.175 | 93.6875 | 92.9714 | 91.9 | 91.5
17.5 | 94.0429 | 93.5625 | 92.8375 | 91.8 | 91.45
17.6 | 93.9125 | 93.4286 | 92.7 | 91.6625 | 91.3143
17.7 | 93.8 | 93.3 | 92.5 | 91.5 | 91.2714
17.8 | 93.7875 | 93.0625 | 92.4 | 91.3167 | 91.0333
17.9 | 93.5714 | 92.9286 | 92.3 | 91.12 | 90.7875
18 | 93.325 | 92.8 | 92.2 | 91.02 | 90.65
18.1 | 93.2 | 92.625 | 92 | 90.82 | 90.46
18.2 | 93.0625 | 92.4286 | 91.6571 | 90.7 | 90.2667
18.3 | 92.9375 | 92.3 | 91.44 | 90.52 | 90.2
18.4 | 92.8 | 92.1625 | 91.34 | 90.3333 | 90.1
18.5 | 92.6714 | 91.9571 | 91.14 | 90.2 | 89.8875
18.6 | 92.4571 | 91.8 | 90.94 | 89.9625 | 89.7143
18.7 | 92.3 | 91.4875 | 90.58 | 89.8 | 89.525
18.8 | 92.1429 | 91.2286 | 90.44 | 89.4714 | 89.3875
18.9 | 91.875 | 91 | 90.24 | 89.2571 | 89.2625
19 | 91.7 | 90.8625 | 90.05 | 89.2 | 89.0571
19.1 | 91.4429 | 90.675 | 89.76 | 88.975 | 88.8
19.2 | 91.175 | 90.5 | 89.58 | 88.8 | 88.5625
19.3 | 90.9 | 90.3 | 89.4 | 88.5286 | 88.4286
19.4 | 90.3286 | 89.8857 | 89.22 | 88.2875 | 88.3
19.5 | 89.4 | 89.7125 | 89.02 | 88.1143 | 88.1714
19.6 | 88.05 | 89.425 | 88.74 | 87.925 | 87.975
19.7 | 85.3571 | 89.08 | 88.42 | 87.7 | 87.8
19.8 | 80.9375 | 88.8 | 88.12 | 87.425 | 87.6429
19.9 | 74.9 | 88.4273 | 87.86 | 87.2286 | 87.375
20 | 67.375 | 87.9273 | 87.62 | 87 | 87.2
20.1 | 58.2 | 87.36 | 87.12 | 86.7571 | 87.0714
20.2 | 48.9875 | 86.5556 | 86.86 | 86.5 | 86.875
20.3 | 39.95 | 85.25 | 86.66 | 86.34 | 86.7125
20.4 | 32.0429 | 83.39 | 86.36 | 86.04 | 86.4714
20.5 | 25.225 | 80.1615 | 85.98 | 85.76 | 86.2
20.6 | 19.95 | 75.7 | 85.625 | 85.5667 | 86.0143
20.7 | 16.4 | 69.15 | 85.3429 | 85.4 | 85.7857
20.8 | 14.1143 | 61.6571 | 84.85 | 85.1 | 85.55
20.9 | 12.7 | 53.42 | 84.4429 | 84.9333 | 85.2286
21 | 11.9 | 45.5333 | 83.8143 | 84.6333 | 85.15
21.1 | 11.3571 | 38.16 | 83 | 84.2429 | 84.9429
21.2 | 10.95 | 31.05 | 82.1 | 84.0143 | 84.8125
21.3 | 10.7 | 25.5429 | 80.9 | 83.75 | 84.5429
21.4 | 10.5143 | 21.5111 | 78.7 | 83.5333 | 84.2333
21.5 | 10.2857 | 18.675 | 75.6429 | 83.175 | 84.0429
21.6 | 10.15 | 16.5857 | 70.9429 | 82.8429 | 83.7125
21.7 | 10.0143 | 15.225 | 65.2 | 82.6375 | 83.4571
21.8 | 9.7857 | 14.35 | 59.2 | 82.2 | 83.1429
21.9 | 9.65 | 13.8286 | 51.85 | 81.95 | 83.0125
22 | 9.5143 | 13.3429 | 45.0286 | 81.6857 | 82.76
22.1 | 9.3875 | 13.0143 | 38.475 | 81.425 | 82.525
22.2 | 9.3 | 12.6429 | 32.3625 | 81 | 82.35
22.3 | 9.225 | 12.3875 | 28.0625 | 80.5286 | 82.0286
22.4 | 9 | 12.2143 | 24.2857 | 80.225 | 81.725
22.5 | 8.9625 | 11.95 | 21.6 | 79.825 | 81.45
22.6 | 8.8286 | 11.675 | 19.7875 | 79.4286 | 81.2
22.7 | 8.7 | 11.4571 | 18.4571 | 79.0375 | 80.9286
22.8 | 8.5625 | 11.25 | 17.6 | 78.425 | 80.6875
22.9 | 8.5 | 11.075 | 16.95 | 78 | 80.5143
23 | 8.4 | 10.8571 | 16.4857 | 77.3714 | 80.325
23.1 | 8.3571 | 10.725 | 16.1 | 76.725 | 79.9625
23.2 | 8.3 | 10.575 | 15.6875 | 75.8143 | 79.6143
23.3 | 8.2 | 10.3625 | 15.4286 | 74.4 | 79.275
23.4 | 8.1625 | 10.2286 | 15.1 | 72.325 | 78.875
23.5 | 8.1 | 10 | 14.7875 | 69.4143 | 78.625
23.6 | 8.1 | 9.9429 | 14.4571 | 65.7 | 78.2857
23.7 | 8 | 9.7375 | 14.2 | 60.8875 | 78
23.8 | 8 | 9.6222 | 13.925 | 55.6286 | 77.6875
23.9 | 8 | 9.4933 | 13.675 | 49.95 | 77.4286
24 | 7.9 | 9.38 | 13.46 | 44.175 | 77.2
24.1 | 7.9 | 9.3133 | 13.225 | 39.0429 | 76.6875
24.2 | 7.8 | 9.1933 | 13.05 | 34.325 | 76.4
24.3 | 7.8 | 9.0812 | 12.8429 | 30.5875 | 76.0571
24.4 | 7.7286 | 9.0188 | 12.5375 | 27.5857 | 75.7667
24.5 | 7.6857 | 8.9417 | 12.375 | 25.3 | 75.3143
24.6 | 7.6 | 8.8375 | 12.1857 | 23.675 | 74.8
24.7 | 7.6 | 8.8 | 11.925 | 22.5 | 74.2875
24.8 | 7.5 | 8.7 | 11.85 | 21.8 | 73.85
24.9 | 7.5 | 8.6375 | 11.6429 | 21.3 | 73.3
25 | 7.4125 | 8.6 | 11.5125 | 20.8143 | 72.6571
25.1 | 7.4 | 8.5 | 11.275 | 20.35 | 71.925
25.2 | 7.3429 | 8.45 | 11.1429 | 19.9 | 71
25.3 | 7.3 | 8.4 | 11.0125 | 19.6429 | 69.4
25.4 | 7.275 | 8.3 | 10.8778 | 19.25 | 67.2125
25.5 | 7.2 | 8.25 | 10.7571 | 18.9 | 64.8
25.6 | 7.2 | 8.1143 | 10.625 | 18.575 | 61.0714
25.7 | 7.1857 | 8.1 | 10.5 | 18.3857 | 56.875
25.8 | 7.1 | 8.05 | 10.4143 | 18.0375 | 52.625
25.9 | 7.1 | 8 | 10.225 | 17.75 | 47.45
26 | 7.0286 | 7.9 | 10.1 | 17.5333 | 43.1
26.1 | 7 | 7.85 | 10.0571 | 17.18 | 38.6429
26.2 | 7 | 7.8 | 9.925 | 16.9571 | 35.05
26.3 | 6.94 | 7.7 | 9.8875 | 16.7286 | 31.8833
26.4 | 6.9 | 7.6556 | 9.7571 | 16.45 | 29.325
26.5 | 6.9 | 7.6 | 9.7 | 16.2143 | 27.65
26.6 | 6.9 | 7.6 | 9.52 | 16.02 | 26.2429
26.7 | 6.8571 | 7.5 | 9.5 | 15.82 | 25.3143
26.8 | 6.8 | 7.5 | 9.4 | 15.52 | 24.6
26.9 | 6.8 | 7.5 | 9.26 | 15.32 | 23.9714
27 | 6.7125 | 7.4 | 9.1714 | 15.02 | 23.5571
27.1 | 6.7 | 7.3167 | 9.1 | 14.9 | 23.15
27.2 |  | 7.275 | 9 | 14.7333 | 22.8286
27.3 |  | 7.2 | 8.9 | 14.4333 | 22.45
27.4 |  | 7.2 | 8.8 | 14.24 | 22.0143
27.5 |  | 7.1 | 8.7 | 14.14 | 21.675
27.6 |  | 7.1 | 8.7 | 13.94 | 21.5
27.7 |  | 7.1 | 8.6 | 13.74 | 21.1143
27.8 |  | 7 | 8.5 | 13.6 | 20.85
27.9 |  | 6.9 | 8.4375 | 13.38 | 20.6
28 |  |  | 8.4 | 13.2 | 20.2714
28.1 |  |  | 8.3 | 13.1 | 19.98
28.2 |  |  | 8.2 | 13 | 19.78
28.3 |  |  | 8.2 | 12.88 | 19.56
28.4 |  |  | 8.1 | 12.68 | 19.28
28.5 |  |  | 8.02 | 12.5 | 19.0667
28.6 |  |  | 7.92 | 12.4 | 18.8167
28.7 |  |  | 7.9 | 12.2875 | 18.62
28.8 |  |  | 7.8625 | 12.1571 | 18.3625
28.9 |  |  | 7.8 | 12.025 | 18.2286
29 |  |  |  | 11.8875 | 17.9
29.1 |  |  |  | 11.7571 | 17.65
29.2 |  |  |  | 11.55 | 17.5
29.3 |  |  |  | 11.3875 | 17.2429
29.4 |  |  |  | 11.3 | 17.0375
29.5 |  |  |  | 11.225 | 16.8
29.6 |  |  |  | 11.1875 | 16.6714
29.7 |  |  |  | 11.0625 | 16.475
29.8 |  |  |  | 10.9286 | 16.3
29.9 |  |  |  | 10.8 | 16.075
30 |  |  |  | 10.6625 | 15.9429
30.1 |  |  |  | 10.6 | 15.8125
30.2 |  |  |  | 10.5 | 15.65
30.3 |  |  |  | 10.4 | 15.4429
30.4 |  |  |  | 10.3 | 15.225
30.5 |  |  |  | 10.2375 | 15.075
30.6 |  |  |  | 10.1 | 14.95
30.7 |  |  |  | 9.975 | 14.8143
30.8 |  |  |  | 9.9 | 14.6429
30.9 |  |  |  | 9.7625 | 14.5125
31 |  |  |  |  | 14.3714
31.1 |  |  |  |  | 14.3
31.2 |  |  |  |  | 14.0571
31.3 |  |  |  |  | 13.98
31.4 |  |  |  |  | 13.8143
31.5 |  |  |  |  | 13.6857
31.6 |  |  |  |  | 13.52
31.7 |  |  |  |  | 13.3875
31.8 |  |  |  |  | 13.2571
31.9 |  |  |  |  | 13.125
32 |  |  |  |  | 13
32.1 |  |  |  |  | 12.8714
32.2 |  |  |  |  | 12.8
32.3 |  |  |  |  | 12.7
32.4 |  |  |  |  | 12.6
32.5 |  |  |  |  | 12.5
32.6 |  |  |  |  | 12.38
32.7 |  |  |  |  | 12.2
32.8 |  |  |  |  | 12.0375
32.9 |  |  |  |  | 11.9

## Wedge 60 Output factors
Output factors
Wedge 60 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15
| 3 | 0.867 | 0.878 | 0.888 | 0.898 | 0.906 | 0.914
| 4 | 0.878 | 0.893 | 0.903 | 0.917 | 0.928 | 0.938
| 5 | 0.888 | 0.905 | 0.917 | 0.933 | 0.946 | 0.96
| 7 | 0.901 | 0.922 | 0.936 | 0.956 | 0.973 | 0.992
| 10 | 0.912 | 0.935 | 0.954 | 0.978 | 1 | 1.026
| 15 | 0.922 | 0.949 | 0.97 | 0.999 | 1.028 | 1.062
| 20 | 0.928 | 0.955 | 0.978 | 1.011 | 1.045 | 1.086
| 30 | 0.936 | 0.967 | 0.991 | 1.029 | 1.071 | 1.121
| 40 | 0.941 | 0.972 | 0.998 | 1.038 | 1.084 | 1.14