## Title
Varian TrueBeam
6FFF Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables
Open field
Open Field Depth Doses
Open Field Profiles at 1.5cm
Open Field Profiles at 5cm
Open Field Profiles at 10cm
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
2. Plots
Open Field

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 57.1 | 57.7 | 58.8 | 60.1 | 61.2 | 66.2 | 69.5 | 72
0.1 | 61.18 | 61.72 | 62.92 | 63.8667 | 65.2125 | 69.975 | 73.1 | 75.5
0.2 | 67.7 | 67.9 | 69.1 | 70.4 | 71.275 | 75.4 | 78.35 | 80.9
0.3 | 76.8333 | 77 | 77.95 | 78.9 | 79.4 | 82.5 | 84.8 | 87.025
0.4 | 83.6 | 84.2 | 84.7 | 84.7 | 85.66 | 88.2 | 89.75 | 91.2
0.5 | 88.9 | 89.1 | 89.4333 | 89.625 | 90.075 | 92.2 | 93.25 | 94.4
0.6 | 92.375 | 92.6 | 92.82 | 93.0429 | 93.35 | 94.9 | 95.6 | 96.2667
0.7 | 95.1 | 95.1 | 95.25 | 95.4 | 95.6 | 96.65 | 97.3167 | 97.8
0.8 | 96.925 | 96.9 | 96.95 | 97.16 | 97.25 | 98.0375 | 98.4 | 98.75
0.9 | 98.175 | 98.2 | 98.2 | 98.3 | 98.4 | 98.975 | 99.18 | 99.35
1 | 99 | 99.0333 | 99.025 | 99.15 | 99.225 | 99.55 | 99.7 | 99.7667
1.1 | 99.65 | 99.7 | 99.65 | 99.7 | 99.725 | 99.9 | 99.9429 | 100
1.2 | 99.9 | 99.9 | 99.9667 | 100 | 99.96 | 100 | 100 | 100
1.3 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
1.4 | 100 | 100 | 100 | 100 | 100 | 99.9571 | 99.925 | 99.9
1.5 | 99.775 | 99.84 | 99.875 | 99.85 | 99.875 | 99.85 | 99.7 | 99.7
1.6 | 99.525 | 99.625 | 99.7 | 99.62 | 99.7 | 99.5 | 99.55 | 99.45
1.7 | 99.16 | 99.2625 | 99.36 | 99.35 | 99.4 | 99.25 | 99.3 | 99.2333
1.8 | 98.75 | 98.8875 | 99.05 | 99.025 | 99.05 | 98.95 | 98.95 | 98.86
1.9 | 98.28 | 98.45 | 98.6 | 98.58 | 98.6 | 98.6 | 98.62 | 98.55
2 | 97.8 | 97.95 | 98.2333 | 98.25 | 98.25 | 98.24 | 98.2333 | 98.175
2.1 | 97.2 | 97.4556 | 97.76 | 97.9 | 97.9 | 97.9 | 97.875 | 97.8667
2.2 | 96.7 | 96.9667 | 97.28 | 97.32 | 97.4667 | 97.5 | 97.55 | 97.375
2.3 | 96.05 | 96.45 | 96.7 | 96.9 | 97 | 97.1 | 97.1 | 97.1
2.4 | 95.6 | 95.9 | 96.25 | 96.45 | 96.625 | 96.7 | 96.7 | 96.6667
2.5 | 94.95 | 95.42 | 95.9 | 96 | 96.15 | 96.2857 | 96.4 | 96.25
2.6 | 94.38 | 94.86 | 95.25 | 95.65 | 95.65 | 95.9 | 96.025 | 95.85
2.7 | 93.8 | 94.25 | 94.825 | 95.05 | 95.25 | 95.52 | 95.6 | 95.4667
2.8 | 93.25 | 93.8 | 94.35 | 94.58 | 94.75 | 95.15 | 95.3429 | 95.1
2.9 | 92.62 | 93.18 | 93.85 | 94.1286 | 94.35 | 94.675 | 94.9 | 94.75
3 | 92.05 | 92.6 | 93.3125 | 93.7 | 93.925 | 94.3333 | 94.45 | 94.4
3.1 | 91.425 | 92 | 92.75 | 93.1714 | 93.425 | 93.9 | 94.05 | 93.95
3.2 | 90.775 | 91.4 | 92.25 | 92.7 | 92.95 | 93.4 | 93.65 | 93.6
3.3 | 90.25 | 90.98 | 91.74 | 92.2 | 92.5 | 93.05 | 93.225 | 93.175
3.4 | 89.68 | 90.3714 | 91.2 | 91.7667 | 92.02 | 92.65 | 92.9 | 92.8333
3.5 | 89.1 | 89.85 | 90.78 | 91.2714 | 91.56 | 92.25 | 92.5 | 92.4
3.6 | 88.5 | 89.3375 | 90.175 | 90.725 | 91.1667 | 91.775 | 92.04 | 92.05
3.7 | 87.9 | 88.7125 | 89.7 | 90.28 | 90.6 | 91.425 | 91.65 | 91.7
3.8 | 87.35 | 88.15 | 89.1571 | 89.75 | 90.15 | 90.975 | 91.25 | 91.3
3.9 | 86.8 | 87.5625 | 88.675 | 89.3 | 89.75 | 90.55 | 90.875 | 90.8333
4 | 86.25 | 87.1 | 88.15 | 88.8 | 89.2333 | 90.15 | 90.4333 | 90.5333
4.1 | 85.7 | 86.5333 | 87.74 | 88.4 | 88.8333 | 89.725 | 90 | 90.1667
4.2 | 85.06 | 86.05 | 87.2 | 87.9 | 88.35 | 89.3 | 89.6667 | 89.6667
4.3 | 84.5 | 85.425 | 86.62 | 87.44 | 87.9 | 88.9 | 89.2 | 89.35
4.4 | 83.9 | 84.8667 | 86.15 | 86.95 | 87.4 | 88.5 | 88.8667 | 88.8333
4.5 | 83.35 | 84.25 | 85.6 | 86.4286 | 86.9 | 88 | 88.5 | 88.4667
4.6 | 82.8 | 83.7 | 85.15 | 85.94 | 86.4 | 87.5667 | 88.05 | 88.0667
4.7 | 82.1 | 83.2 | 84.6 | 85.45 | 85.95 | 87.16 | 87.625 | 87.75
4.8 | 81.55 | 82.66 | 84.0667 | 84.94 | 85.525 | 86.725 | 87.3 | 87.35
4.9 | 80.975 | 82.0333 | 83.5667 | 84.475 | 85 | 86.35 | 86.85 | 86.9
5 | 80.4 | 81.5 | 83 | 84.05 | 84.6 | 85.94 | 86.4 | 86.5333
5.1 | 79.8667 | 81 | 82.55 | 83.5333 | 84.08 | 85.5667 | 86.02 | 86.2333
5.2 | 79.4 | 80.5 | 82 | 82.9 | 83.7 | 85.02 | 85.64 | 85.7
5.3 | 78.85 | 79.9 | 81.55 | 82.5 | 83.2 | 84.7 | 85.15 | 85.35
5.4 | 78.34 | 79.36 | 81.025 | 82 | 82.7333 | 84.175 | 84.75 | 84.9
5.5 | 77.65 | 78.9 | 80.5667 | 81.5 | 82.25 | 83.925 | 84.35 | 84.6
5.6 | 77.15 | 78.4 | 80.0333 | 81.0667 | 81.85 | 83.375 | 84.05 | 84.2
5.7 | 76.65 | 77.825 | 79.55 | 80.64 | 81.35 | 83 | 83.58 | 83.7667
5.8 | 76.1 | 77.3 | 79.05 | 80.175 | 80.85 | 82.7 | 83.3 | 83.4
5.9 | 75.7 | 76.8 | 78.55 | 79.75 | 80.46 | 82.25 | 82.88 | 83.05
6 | 75.15 | 76.32 | 78.0667 | 79.26 | 80.08 | 81.84 | 82.5 | 82.625
6.1 | 74.64 | 75.85 | 77.66 | 78.8 | 79.65 | 81.5 | 82.1 | 82.3333
6.2 | 74.075 | 75.4333 | 77.2 | 78.3 | 79.2 | 81.1 | 81.7 | 81.9
6.3 | 73.575 | 74.85 | 76.6667 | 77.94 | 78.8 | 80.55 | 81.3 | 81.5
6.4 | 73.1 | 74.375 | 76.15 | 77.45 | 78.3 | 80.25 | 80.92 | 81.0667
6.5 | 72.58 | 73.95 | 75.6 | 76.95 | 77.85 | 79.8667 | 80.55 | 80.7667
6.6 | 72.1 | 73.4 | 75.25 | 76.55 | 77.45 | 79.5 | 80.15 | 80.3667
6.7 | 71.6 | 72.9 | 74.82 | 76.125 | 77.025 | 79 | 79.75 | 80
6.8 | 71.2 | 72.4 | 74.35 | 75.7125 | 76.6333 | 78.65 | 79.45 | 79.6667
6.9 | 70.65 | 71.9 | 73.875 | 75.225 | 76.1125 | 78.2 | 79 | 79.1667
7 | 70.15 | 71.55 | 73.45 | 74.8 | 75.7375 | 77.9 | 78.7 | 78.9
7.1 | 69.75 | 71.05 | 72.98 | 74.375 | 75.2625 | 77.52 | 78.3 | 78.4
7.2 | 69.28 | 70.55 | 72.5 | 73.9 | 74.8875 | 77.075 | 77.85 | 78.125
7.3 | 68.8 | 70.075 | 72.1 | 73.5 | 74.4857 | 76.74 | 77.5 | 77.7
7.4 | 68.25 | 69.7 | 71.65 | 73 | 73.98 | 76.35 | 77.1 | 77.3
7.5 | 67.8 | 69.18 | 71.25 | 72.64 | 73.65 | 75.95 | 76.8 | 76.975
7.6 | 67.375 | 68.6714 | 70.75 | 72.15 | 73.2 | 75.5 | 76.45 | 76.625
7.7 | 66.8 | 68.3 | 70.275 | 71.825 | 72.8 | 75.15 | 76 | 76.25
7.8 | 66.45 | 67.8 | 69.9 | 71.4 | 72.4429 | 74.8 | 75.6667 | 75.9
7.9 | 66 | 67.3 | 69.4 | 70.925 | 71.94 | 74.4571 | 75.275 | 75.5667
8 | 65.5 | 66.9 | 69 | 70.5 | 71.5625 | 74.05 | 74.94 | 75.2
8.1 | 65.05 | 66.5375 | 68.55 | 70.0333 | 71.1667 | 73.65 | 74.6 | 74.8
8.2 | 64.625 | 65.975 | 68.15 | 69.65 | 70.7286 | 73.275 | 74.15 | 74.5
8.3 | 64.15 | 65.575 | 67.68 | 69.2 | 70.325 | 72.85 | 73.84 | 74.1
8.4 | 63.825 | 65.2 | 67.3 | 68.8 | 69.9286 | 72.475 | 73.4 | 73.8
8.5 | 63.35 | 64.7 | 66.9 | 68.4 | 69.5 | 72.12 | 73.075 | 73.4333
8.6 | 62.9 | 64.3 | 66.4 | 68 | 69.1 | 71.75 | 72.8 | 73.1
8.7 | 62.4 | 63.8 | 65.92 | 67.5625 | 68.7 | 71.4286 | 72.4 | 72.65
8.8 | 61.9667 | 63.4 | 65.5 | 67.15 | 68.3 | 71.0429 | 72.05 | 72.3
8.9 | 61.6 | 62.95 | 65.15 | 66.7333 | 67.9 | 70.6778 | 71.6 | 71.975
9 | 61.1333 | 62.5 | 64.68 | 66.3 | 67.52 | 70.275 | 71.3 | 71.65
9.1 | 60.75 | 62.0667 | 64.3 | 65.9 | 67.05 | 69.9 | 70.85 | 71.3333
9.2 | 60.2333 | 61.75 | 63.82 | 65.5 | 66.65 | 69.575 | 70.575 | 70.95
9.3 | 59.9 | 61.25 | 63.4143 | 65.1 | 66.25 | 69.1 | 70.2 | 70.525
9.4 | 59.5 | 60.85 | 63.05 | 64.66 | 65.85 | 68.7857 | 69.8 | 70.2333
9.5 | 59 | 60.35 | 62.625 | 64.25 | 65.45 | 68.4333 | 69.4 | 69.85
9.6 | 58.6 | 59.96 | 62.175 | 63.8125 | 65.1 | 68.0625 | 69.1 | 69.45
9.7 | 58.15 | 59.6 | 61.825 | 63.44 | 64.6667 | 67.65 | 68.82 | 69.15
9.8 | 57.74 | 59.15 | 61.375 | 63.1 | 64.2 | 67.3 | 68.4 | 68.825
9.9 | 57.35 | 58.7 | 60.95 | 62.65 | 63.875 | 66.9 | 68.075 | 68.525
10 | 57 | 58.35 | 60.5333 | 62.25 | 63.5286 | 66.5 | 67.7333 | 68.15
10.1 | 56.525 | 57.95 | 60.2 | 61.9 | 63.1286 | 66.225 | 67.3 | 67.8
10.2 | 56.2 | 57.575 | 59.8 | 61.4667 | 62.6875 | 65.8 | 66.95 | 67.475
10.3 | 55.8 | 57.15 | 59.3 | 61.0667 | 62.3125 | 65.4667 | 66.5667 | 67.1
10.4 | 55.4 | 56.775 | 59 | 60.7 | 62.025 | 65.1 | 66.3 | 66.8333
10.5 | 55.02 | 56.4 | 58.6 | 60.34 | 61.55 | 64.7 | 65.9333 | 66.4
10.6 | 54.65 | 56.1 | 58.25 | 59.92 | 61.225 | 64.45 | 65.5667 | 66.0333
10.7 | 54.26 | 55.7 | 57.9 | 59.55 | 60.86 | 64.125 | 65.3 | 65.8
10.8 | 53.925 | 55.325 | 57.55 | 59.24 | 60.5429 | 63.75 | 65 | 65.4
10.9 | 53.55 | 54.9333 | 57.1625 | 58.8875 | 60.15 | 63.44 | 64.62 | 65.125
11 | 53.2333 | 54.5 | 56.8 | 58.475 | 59.825 | 63.125 | 64.3 | 64.72
11.1 | 52.8 | 54.225 | 56.5 | 58.225 | 59.45 | 62.75 | 64.025 | 64.4
11.2 | 52.425 | 53.9 | 56.1 | 57.78 | 59.1 | 62.45 | 63.7 | 64.1
11.3 | 52.1 | 53.4667 | 55.6667 | 57.475 | 58.75 | 62.1333 | 63.375 | 63.825
11.4 | 51.8 | 53.1 | 55.4 | 57.1 | 58.4333 | 61.775 | 63 | 63.4667
11.5 | 51.44 | 52.7333 | 55.0667 | 56.8 | 58.075 | 61.4875 | 62.68 | 63.2
11.6 | 51 | 52.45 | 54.6333 | 56.45 | 57.7667 | 61.1125 | 62.3667 | 62.85
11.7 | 50.7333 | 52.05 | 54.28 | 56 | 57.35 | 60.825 | 62.075 | 62.54
11.8 | 50.4333 | 51.72 | 53.975 | 55.7 | 57 | 60.475 | 61.7375 | 62.2
11.9 | 50.0625 | 51.425 | 53.65 | 55.35 | 56.7 | 60.15 | 61.4429 | 61.85
12 | 49.725 | 50.975 | 53.3 | 55.05 | 56.35 | 59.775 | 61.1375 | 61.5333
12.1 | 49.4125 | 50.725 | 52.9 | 54.675 | 56.05 | 59.4375 | 60.75 | 61.25
12.2 | 49.0375 | 50.3333 | 52.6 | 54.325 | 55.725 | 59.1 | 60.4429 | 60.96
12.3 | 48.76 | 50.05 | 52.26 | 53.975 | 55.4 | 58.825 | 60.1 | 60.65
12.4 | 48.45 | 49.7 | 51.9 | 53.65 | 55.05 | 58.55 | 59.85 | 60.3
12.5 | 48.1 | 49.34 | 51.625 | 53.35 | 54.7 | 58.2 | 59.5 | 60
12.6 | 47.7 | 49.05 | 51.25 | 53 | 54.36 | 57.875 | 59.2 | 59.7
12.7 | 47.4 | 48.72 | 50.925 | 52.7 | 54.04 | 57.55 | 58.8667 | 59.4
12.8 | 47.05 | 48.375 | 50.575 | 52.35 | 53.75 | 57.2 | 58.6 | 59.1
12.9 | 46.7714 | 48.05 | 50.3 | 52.1 | 53.375 | 56.95 | 58.35 | 58.82
13 | 46.4286 | 47.775 | 49.925 | 51.76 | 53.05 | 56.6 | 58.025 | 58.4667
13.1 | 46.14 | 47.44 | 49.66 | 51.35 | 52.75 | 56.35 | 57.675 | 58.2333
13.2 | 45.85 | 47.1 | 49.35 | 51.075 | 52.425 | 56 | 57.425 | 57.8667
13.3 | 45.525 | 46.8 | 49 | 50.725 | 52.075 | 55.6667 | 57.08 | 57.6
13.4 | 45.2 | 46.5 | 48.6667 | 50.38 | 51.84 | 55.375 | 56.8333 | 57.3
13.5 | 44.925 | 46.1429 | 48.375 | 50.14 | 51.5 | 55.075 | 56.5667 | 57
13.6 | 44.7 | 45.9 | 48.025 | 49.85 | 51.25 | 54.725 | 56.25 | 56.775
13.7 | 44.3 | 45.55 | 47.8 | 49.55 | 50.9 | 54.5 | 56 | 56.5
13.8 | 44.06 | 45.25 | 47.5 | 49.2286 | 50.6 | 54.2 | 55.65 | 56.1333
13.9 | 43.725 | 44.95 | 47.1 | 48.8667 | 50.2571 | 53.875 | 55.34 | 55.86
14 | 43.5 | 44.7 | 46.85 | 48.5571 | 49.95 | 53.625 | 55.075 | 55.625
14.1 | 43.225 | 44.4 | 46.52 | 48.3222 | 49.6 | 53.3 | 54.725 | 55.3
14.2 | 42.9 | 44.0375 | 46.25 | 47.975 | 49.325 | 53.05 | 54.475 | 55.1
14.3 | 42.6 | 43.775 | 45.9667 | 47.6875 | 49.05 | 52.7 | 54.125 | 54.75
14.4 | 42.25 | 43.5 | 45.6375 | 47.375 | 48.725 | 52.45 | 53.9 | 54.4667
14.5 | 42 | 43.18 | 45.35 | 47.125 | 48.375 | 52.1 | 53.66 | 54.2
14.6 | 41.7 | 42.9 | 45.0333 | 46.78 | 48.125 | 51.82 | 53.35 | 53.825
14.7 | 41.4 | 42.55 | 44.8 | 46.45 | 47.9 | 51.525 | 53.1 | 53.56
14.8 | 41.15 | 42.3 | 44.5 | 46.15 | 47.5 | 51.275 | 52.775 | 53.325
14.9 | 40.84 | 42 | 44.2 | 45.8286 | 47.25 | 50.9667 | 52.475 | 53.0333
15 | 40.575 | 41.8 | 43.85 | 45.58 | 47 | 50.675 | 52.125 | 52.75
15.1 | 40.325 | 41.55 | 43.5 | 45.3 | 46.675 | 50.44 | 51.875 | 52.5
15.2 | 39.975 | 41.15 | 43.3 | 45.02 | 46.425 | 50.1 | 51.65 | 52.18
15.3 | 39.725 | 40.9 | 43.0667 | 44.775 | 46.1333 | 49.85 | 51.375 | 51.925
15.4 | 39.5333 | 40.625 | 42.775 | 44.4667 | 45.875 | 49.6 | 51.08 | 51.65
15.5 | 39.3 | 40.375 | 42.45 | 44.1333 | 45.55 | 49.35 | 50.8 | 51.46
15.6 | 38.9333 | 40.125 | 42.16 | 43.875 | 45.34 | 49 | 50.55 | 51.2
15.7 | 38.725 | 39.8667 | 41.925 | 43.65 | 45 | 48.8 | 50.275 | 50.9
15.8 | 38.475 | 39.55 | 41.7 | 43.35 | 44.7333 | 48.4 | 50.025 | 50.65
15.9 | 38.18 | 39.32 | 41.35 | 43.1 | 44.45 | 48.2 | 49.8 | 50.4
16 | 37.9 | 39.075 | 41.1 | 42.85 | 44.2333 | 48 | 49.5 | 50.125
16.1 | 37.75 | 38.78 | 40.9 | 42.62 | 44 | 47.75 | 49.2 | 49.9
16.2 | 37.4 | 38.5667 | 40.68 | 42.3 | 43.7143 | 47.45 | 48.9667 | 49.625
16.3 | 37.25 | 38.275 | 40.425 | 42.075 | 43.45 | 47.15 | 48.7 | 49.4
16.4 | 37 | 38.125 | 40.175 | 41.825 | 43.2 | 46.92 | 48.45 | 49.05
16.5 | 36.76 | 37.875 | 39.8714 | 41.575 | 42.95 | 46.675 | 48.2 | 48.85
16.6 | 36.525 | 37.575 | 39.6 | 41.325 | 42.65 | 46.425 | 47.9667 | 48.575
16.7 | 36.275 | 37.3 | 39.34 | 41.075 | 42.35 | 46.175 | 47.7 | 48.3
16.8 | 36.025 | 37.05 | 39.075 | 40.825 | 42.1 | 45.925 | 47.475 | 48.05
16.9 | 35.775 | 36.9 | 38.825 | 40.58 | 41.8556 | 45.675 | 47.225 | 47.775
17 | 35.525 | 36.575 | 38.575 | 40.25 | 41.625 | 45.4333 | 46.975 | 47.54
17.1 | 35.28 | 36.425 | 38.325 | 40 | 41.375 | 45.2 | 46.7 | 47.3
17.2 | 35.1333 | 36.175 | 38.075 | 39.76 | 41.125 | 44.95 | 46.5 | 47.1
17.3 | 34.9 | 35.9 | 37.925 | 39.525 | 40.875 | 44.7 | 46.15 | 46.775
17.4 | 34.6429 | 35.7 | 37.675 | 39.275 | 40.65 | 44.375 | 46 | 46.525
17.5 | 34.4 | 35.5 | 37.325 | 39.025 | 40.38 | 44.1 | 45.7 | 46.4
17.6 | 34.2 | 35.225 | 37.175 | 38.8375 | 40.15 | 43.9 | 45.4667 | 46.1333
17.7 | 34 | 35.0375 | 36.9333 | 38.6 | 39.9 | 43.7 | 45.225 | 45.8667
17.8 | 33.75 | 34.74 | 36.7 | 38.34 | 39.6333 | 43.375 | 45.025 | 45.6667
17.9 | 33.55 | 34.55 | 36.44 | 38.14 | 39.475 | 43.225 | 44.775 | 45.4
18 | 33.3 | 34.38 | 36.2 | 37.9 | 39.225 | 42.975 | 44.5333 | 45.1
18.1 | 33.1 | 34.1 | 36.05 | 37.65 | 38.95 | 42.75 | 44.325 | 44.9
18.2 | 32.92 | 33.9 | 35.8 | 37.4 | 38.8 | 42.5 | 44.075 | 44.65
18.3 | 32.7 | 33.675 | 35.6 | 37.15 | 38.5 | 42.25 | 43.8333 | 44.5
18.4 | 32.425 | 33.425 | 35.35 | 37 | 38.3 | 42.0667 | 43.575 | 44.325
18.5 | 32.2778 | 33.275 | 35.125 | 36.75 | 38.0333 | 41.8 | 43.425 | 44.0333
18.6 | 32.0429 | 33 | 34.875 | 36.5 | 37.85 | 41.575 | 43.2 | 43.775
18.7 | 31.8444 | 32.8333 | 34.7 | 36.25 | 37.6 | 41.32 | 42.96 | 43.55
18.8 | 31.6286 | 32.6167 | 34.475 | 36.12 | 37.35 | 41.1 | 42.725 | 43.32
18.9 | 31.4 | 32.3857 | 34.225 | 35.86 | 37.1 | 40.86 | 42.5 | 43.1
19 | 31.2 | 32.2 | 33.98 | 35.6 | 36.95 | 40.625 | 42.2333 | 42.9333
19.1 | 31 | 32 | 33.825 | 35.4 | 36.7 | 40.3833 | 42 | 42.6333
19.2 | 30.75 | 31.8 | 33.575 | 35.2 | 36.4333 | 40.1889 | 41.8 | 42.4667
19.3 | 30.6 | 31.55 | 33.325 | 34.9333 | 36.2 | 39.9286 | 41.5667 | 42.2
19.4 | 30.375 | 31.3 | 33.1667 | 34.7 | 35.975 | 39.68 | 41.36 | 41.975
19.5 | 30.2125 | 31.15 | 32.95 | 34.5 | 35.825 | 39.5 | 41.08 | 41.8
19.6 | 29.98 | 30.9 | 32.725 | 34.275 | 35.575 | 39.25 | 40.8714 | 41.525
19.7 | 29.8 | 30.75 | 32.55 | 34.025 | 35.4 | 39 | 40.675 | 41.2667
19.8 | 29.6 | 30.5667 | 32.3 | 33.875 | 35.1 | 38.8 | 40.45 | 41.15
19.9 | 29.4 | 30.3 | 32.1 | 33.625 | 34.95 | 38.6 | 40.2 | 40.9
20 | 29.2 | 30.125 | 31.9 | 33.4667 | 34.725 | 38.375 | 39.975 | 40.6333
20.1 | 29.05 | 29.875 | 31.65 | 33.26 | 34.475 | 38.2 | 39.725 | 40.5
20.2 | 28.825 | 29.75 | 31.525 | 33.025 | 34.3 | 37.975 | 39.6 | 40.3
20.3 | 28.675 | 29.6 | 31.3 | 32.775 | 34.06 | 37.725 | 39.34 | 40.1
20.4 | 28.425 | 29.35 | 31.075 | 32.6 | 33.825 | 37.5 | 39.2 | 39.85
20.5 | 28.275 | 29.2 | 30.9 | 32.4 | 33.675 | 37.34 | 38.9 | 39.6667
20.6 | 28.05 | 28.96 | 30.68 | 32.175 | 33.425 | 37.1429 | 38.7 | 39.475
20.7 | 27.9 | 28.825 | 30.55 | 32 | 33.2167 | 36.95 | 38.475 | 39.1333
20.8 | 27.7 | 28.575 | 30.3 | 31.775 | 33 | 36.7 | 38.325 | 39
20.9 | 27.55 | 28.45 | 30.1 | 31.6 | 32.9 | 36.475 | 38.1 | 38.8
21 | 27.35 | 28.3 | 29.95 | 31.4 | 32.6667 | 36.325 | 37.9333 | 38.6
21.1 | 27.2 | 28.0667 | 29.7333 | 31.26 | 32.5 | 36.1375 | 37.7 | 38.4
21.2 | 27.025 | 27.9 | 29.6 | 31.1 | 32.2333 | 35.875 | 37.55 | 38.2333
21.3 | 26.9 | 27.725 | 29.4 | 30.9 | 32.05 | 35.7 | 37.3333 | 38
21.4 | 26.675 | 27.6 | 29.2 | 30.7 | 31.9 | 35.5 | 37.1 | 37.8
21.5 | 26.54 | 27.4 | 29.05 | 30.5 | 31.7 | 35.35 | 36.9429 | 37.6
21.6 | 26.42 | 27.25 | 28.9 | 30.35 | 31.55 | 35.1 | 36.6889 | 37.4
21.7 | 26.1667 | 27.1 | 28.7 | 30.1 | 31.3667 | 34.975 | 36.525 | 37.2
21.8 | 26.025 | 26.85 | 28.5 | 29.9571 | 31.2 | 34.725 | 36.4 | 37
21.9 | 25.9 | 26.7 | 28.4 | 29.8 | 30.975 | 34.575 | 36.1 | 36.8667
22 | 25.675 | 26.55 | 28.18 | 29.6 | 30.8 | 34.4 | 35.9333 | 36.6
22.1 | 25.55 | 26.4 | 28 | 29.4 | 30.625 | 34.16 | 35.8 | 36.4
22.2 | 25.375 | 26.225 | 27.8 | 29.25 | 30.5 | 34.025 | 35.575 | 36.275
22.3 | 25.2 | 26.05 | 27.65 | 29.1 | 30.275 | 33.82 | 35.4 | 36.075
22.4 | 25.075 | 25.9 | 27.5 | 28.8667 | 30.1 | 33.6 | 35.2 | 35.85
22.5 | 24.925 | 25.7 | 27.3 | 28.7 | 29.85 | 33.4167 | 35 | 35.7
22.6 | 24.8 | 25.6 | 27.1 | 28.5 | 29.725 | 33.275 | 34.8 | 35.4667
22.7 | 24.55 | 25.4 | 26.975 | 28.4 | 29.575 | 33.025 | 34.6333 | 35.2667
22.8 | 24.45 | 25.225 | 26.825 | 28.275 | 29.325 | 32.875 | 34.5 | 35.1
22.9 | 24.3 | 25.075 | 26.675 | 28.02 | 29.2 | 32.725 | 34.28 | 34.9
23 | 24.15 | 24.925 | 26.44 | 27.85 | 29 | 32.5 | 34.05 | 34.7333
23.1 | 24 | 24.7375 | 26.325 | 27.7 | 28.8 | 32.325 | 33.9 | 34.6
23.2 | 23.875 | 24.6125 | 26.175 | 27.55 | 28.7 | 32.1333 | 33.75 | 34.35
23.3 | 23.6625 | 24.4833 | 26 | 27.4 | 28.5 | 32 | 33.6 | 34.2
23.4 | 23.5286 | 24.3 | 25.875 | 27.2 | 28.3667 | 31.8 | 33.35 | 34
23.5 | 23.4 | 24.14 | 25.625 | 27 | 28.15 | 31.625 | 33.2 | 33.9
23.6 | 23.2 | 24 | 25.5 | 26.85 | 28 | 31.475 | 33.05 | 33.7
23.7 | 23.1 | 23.85 | 25.35 | 26.7 | 27.8333 | 31.325 | 32.8 | 33.55
23.8 | 22.9 | 23.7 | 25.2 | 26.575 | 27.7 | 31.1 | 32.65 | 33.4
23.9 | 22.8 | 23.55 | 25.05 | 26.4 | 27.5667 | 30.94 | 32.425 | 33.2
24 | 22.6333 | 23.42 | 24.9 | 26.2 | 27.3 | 30.8 | 32.32 | 33
24.1 | 22.4667 | 23.275 | 24.775 | 26.04 | 27.2 | 30.6 | 32.15 | 32.8
24.2 | 22.4 | 23.125 | 24.6 | 25.9 | 27 | 30.4 | 31.98 | 32.7
24.3 | 22.2 | 22.975 | 24.46 | 25.7333 | 26.9 | 30.2667 | 31.84 | 32.5
24.4 | 22.06 | 22.825 | 24.32 | 25.6 | 26.75 | 30.1 | 31.6 | 32.3
24.5 | 21.9 | 22.675 | 24.1 | 25.475 | 26.54 | 29.9 | 31.45 | 32.2
24.6 | 21.8 | 22.54 | 24.025 | 25.3 | 26.4 | 29.7667 | 31.3 | 31.96
24.7 | 21.625 | 22.4 | 23.8 | 25.14 | 26.2 | 29.58 | 31.1 | 31.8
24.8 | 21.5 | 22.3 | 23.7 | 24.96 | 26.075 | 29.4 | 30.9833 | 31.6
24.9 | 21.4 | 22.1 | 23.575 | 24.825 | 25.9 | 29.225 | 30.775 | 31.5
25 | 21.25 | 22 | 23.4 | 24.675 | 25.7 | 29.1 | 30.6 | 31.3
25.1 | 21.1 | 21.85 | 23.2 | 24.55 | 25.55 | 28.9 | 30.42 | 31.1
25.2 | 21 | 21.7 | 23.1 | 24.4 | 25.5 | 28.74 | 30.3333 | 31
25.3 | 20.825 | 21.575 | 23 | 24.3 | 25.3 | 28.575 | 30.18 | 30.8
25.4 | 20.7 | 21.4 | 22.8 | 24.1 | 25.125 | 28.42 | 30 | 30.6
25.5 | 20.56 | 21.3 | 22.7 | 23.92 | 25.025 | 28.275 | 29.775 | 30.5
25.6 | 20.425 | 21.2 | 22.6 | 23.8 | 24.85 | 28.125 | 29.625 | 30.3
25.7 | 20.36 | 21 | 22.44 | 23.7 | 24.7 | 28 | 29.5 | 30.2
25.8 | 20.225 | 20.95 | 22.25 | 23.55 | 24.6 | 27.8 | 29.35 | 30
25.9 | 20.1 | 20.8 | 22.1 | 23.4 | 24.425 | 27.7 | 29.2125 | 29.9
26 | 19.975 | 20.65 | 22 | 23.3 | 24.275 | 27.55 | 29.0286 | 29.7
26.1 | 19.875 | 20.52 | 21.9 | 23.1 | 24.125 | 27.4 | 28.9 | 29.55
26.2 | 19.725 | 20.4 | 21.76 | 23 | 24 | 27.275 | 28.725 | 29.4
26.3 | 19.6 | 20.2875 | 21.625 | 22.8667 | 23.825 | 27.1 | 28.6 | 29.3
26.4 | 19.5 | 20.1625 | 21.45 | 22.7 | 23.7375 | 26.9 | 28.46 | 29.1
26.5 | 19.375 | 20.0375 | 21.3 | 22.5833 | 23.6125 | 26.8 | 28.3125 | 28.9
26.6 | 19.25 | 19.9125 | 21.225 | 22.4375 | 23.475 | 26.7 | 28.1 | 28.8
26.7 | 19.1 | 19.775 | 21.1111 | 22.3222 | 23.3 | 26.5 | 27.95 | 28.6
26.8 | 19 | 19.675 | 20.9778 | 22.2 | 23.1333 | 26.3333 | 27.8 | 28.5
26.9 | 18.9 | 19.525 | 20.85 | 22.05 | 23 | 26.2 | 27.7 | 28.3
27 | 18.775 | 19.4 | 20.725 | 21.9 | 22.9 | 26.075 | 27.525 | 28.2
27.1 | 18.625 | 19.3 | 20.6 | 21.8 | 22.8 | 25.9 | 27.4 | 28.075
27.2 | 18.5 | 19.2 | 20.5 | 21.7 | 22.65 | 25.8 | 27.24 | 27.95
27.3 | 18.4 | 19.05 | 20.4 | 21.55 | 22.52 | 25.64 | 27.1 | 27.75
27.4 | 18.3 | 19 | 20.25 | 21.4 | 22.375 | 25.5 | 26.95 | 27.6
27.5 | 18.2 | 18.875 | 20.1 | 21.275 | 22.225 | 25.4 | 26.8 | 27.5
27.6 | 18.1 | 18.725 | 20 | 21.1333 | 22.1 | 25.225 | 26.66 | 27.3
27.7 | 17.9667 | 18.6 | 19.9 | 21 | 22 | 25.075 | 26.5 | 27.2
27.8 | 17.85 | 18.5 | 19.775 | 20.9 | 21.875 | 24.925 | 26.35 | 27.05
27.9 | 17.72 | 18.4 | 19.6 | 20.78 | 21.725 | 24.8 | 26.225 | 26.9
28 | 17.675 | 18.3 | 19.5 | 20.65 | 21.575 | 24.625 | 26.1 | 26.8
28.1 | 17.52 | 18.1889 | 19.425 | 20.5 | 21.425 | 24.5 | 26 | 26.6667
28.2 | 17.45 | 18.0714 | 19.3 | 20.4 | 21.3 | 24.4 | 25.8 | 26.5
28.3 | 17.3 | 17.9286 | 19.2 | 20.3 | 21.225 | 24.2667 | 25.7 | 26.4
28.4 | 17.2 | 17.8 | 19 | 20.15 | 21.1 | 24.1667 | 25.6 | 26.2
28.5 | 17.1 | 17.725 | 18.95 | 20 | 20.95 | 24.05 | 25.4 | 26.1
28.6 | 17 | 17.6 | 18.8 | 19.9 | 20.9 | 23.925 | 25.3 | 26
28.7 | 16.9 | 17.525 | 18.7 | 19.8 | 20.7 | 23.8 | 25.125 | 25.8
28.8 | 16.8 | 17.4 | 18.575 | 19.7 | 20.625 | 23.575 | 25 | 25.7
28.9 | 16.7 | 17.3 | 18.5 | 19.6 | 20.5 | 23.4333 | 24.9 | 25.55
29 | 16.6 | 17.2 | 18.35 | 19.475 | 20.4 | 23.4 | 24.775 | 25.4
29.1 | 16.5 | 17.1 | 18.3 | 19.3333 | 20.25 | 23.275 | 24.625 | 25.3
29.2 | 16.4 | 17 | 18.1333 | 19.2 | 20.175 | 23.1 | 24.475 | 25.15
29.3 | 16.3 | 16.9 | 18.1 | 19.1 | 20 | 22.975 | 24.3 | 25
29.4 | 16.2 | 16.8 | 17.95 | 19 | 19.9 | 22.825 | 24.2333 | 24.9
29.5 | 16.1 | 16.7 | 17.8 | 18.9 | 19.8 | 22.7 | 24.125 | 24.7667
29.6 | 16 | 16.5571 | 17.75 | 18.8 | 19.7 | 22.6 | 24 | 24.6
29.7 | 15.9 | 16.5 | 17.62 | 18.7 | 19.6 | 22.5 | 23.9 | 24.5
29.8 | 15.8 | 16.4 | 17.5 | 18.6 | 19.4333 | 22.3 | 23.7 | 24.325
29.9 | 15.7 | 16.3 | 17.4 | 18.4667 | 19.4 | 22.25 | 23.6571 | 24.2667
30 | 15.625 | 16.2 | 17.3 | 18.3667 | 19.2167 | 22.1 | 23.48 | 24.1

## Open Field Profiles at 1.5cm
Lateral Profiles
Open field
SSD=100cm
Depth=1.5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.2 | [7 empty] | 2.6
-25.1 | [7 empty] | 2.7
-25 | [7 empty] | 2.7
-24.9 | [7 empty] | 2.7
-24.8 | [7 empty] | 2.8
-24.7 | [7 empty] | 2.8
-24.6 | [7 empty] | 2.8
-24.5 | [7 empty] | 2.9
-24.4 | [7 empty] | 2.9
-24.3 | [7 empty] | 3
-24.2 | [7 empty] | 3
-24.1 | [7 empty] | 3.0667
-24 | [7 empty] | 3.1
-23.9 | [7 empty] | 3.1
-23.8 | [7 empty] | 3.2
-23.7 | [7 empty] | 3.2571
-23.6 | [7 empty] | 3.3
-23.5 | [7 empty] | 3.3
-23.4 | [7 empty] | 3.4
-23.3 | [7 empty] | 3.4
-23.2 | [7 empty] | 3.5
-23.1 | [7 empty] | 3.6
-23 | [7 empty] | 3.6333
-22.9 | [7 empty] | 3.7
-22.8 | [7 empty] | 3.8
-22.7 | [7 empty] | 3.8667
-22.6 | [7 empty] | 3.9
-22.5 | [7 empty] | 4
-22.4 | [7 empty] | 4.0833
-22.3 | [7 empty] | 4.1
-22.2 | [7 empty] | 4.25
-22.1 | [7 empty] | 4.34
-22 | [7 empty] | 4.4333
-21.9 | [7 empty] | 4.5167
-21.8 | [7 empty] | 4.6
-21.7 | [7 empty] | 4.7
-21.6 | [7 empty] | 4.8667
-21.5 | [7 empty] | 5
-21.4 | [7 empty] | 5.1333
-21.3 | [7 empty] | 5.34
-21.2 | [7 empty] | 5.5333
-21.1 | [7 empty] | 5.8667
-21 | [7 empty] | 6.25
-20.9 | [7 empty] | 6.8333
-20.8 | [7 empty] | 7.8
-20.7 | [7 empty] | 9.5667
-20.6 | [7 empty] | 12.9
-20.5 | [7 empty] | 17.6143
-20.4 | [7 empty] | 23.3
-20.3 | [7 empty] | 29.4333
-20.2 | [6 empty] | 2.5 | 35.8667
-20.1 | [6 empty] | 2.6 | 42.1
-20 | [6 empty] | 2.6 | 46.8333
-19.9 | [6 empty] | 2.625 | 50.0167
-19.8 | [6 empty] | 2.7 | 51.85
-19.7 | [6 empty] | 2.7 | 52.9
-19.6 | [6 empty] | 2.7667 | 53.55
-19.5 | [6 empty] | 2.8 | 54.05
-19.4 | [6 empty] | 2.8667 | 54.4
-19.3 | [6 empty] | 2.9 | 54.8
-19.2 | [6 empty] | 2.9 | 55.1
-19.1 | [6 empty] | 3 | 55.3
-19 | [6 empty] | 3 | 55.5667
-18.9 | [6 empty] | 3.1 | 55.8333
-18.8 | [6 empty] | 3.1 | 56
-18.7 | [6 empty] | 3.2 | 56.325
-18.6 | [6 empty] | 3.2 | 56.56
-18.5 | [6 empty] | 3.3 | 56.7167
-18.4 | [6 empty] | 3.34 | 56.9667
-18.3 | [6 empty] | 3.4 | 57.1833
-18.2 | [6 empty] | 3.4 | 57.4
-18.1 | [6 empty] | 3.5 | 57.6167
-18 | [6 empty] | 3.6 | 57.8667
-17.9 | [6 empty] | 3.68 | 58.0429
-17.8 | [6 empty] | 3.7167 | 58.32
-17.7 | [6 empty] | 3.8 | 58.5
-17.6 | [6 empty] | 3.85 | 58.725
-17.5 | [6 empty] | 3.95 | 58.9333
-17.4 | [6 empty] | 4 | 59.2
-17.3 | [6 empty] | 4.1 | 59.3667
-17.2 | [6 empty] | 4.2 | 59.62
-17.1 | [6 empty] | 4.3 | 59.8167
-17 | [6 empty] | 4.4 | 60.0667
-16.9 | [6 empty] | 4.5 | 60.25
-16.8 | [6 empty] | 4.6 | 60.48
-16.7 | [6 empty] | 4.76 | 60.7143
-16.6 | [6 empty] | 4.875 | 60.9
-16.5 | [6 empty] | 5 | 61.2
-16.4 | [6 empty] | 5.1167 | 61.4
-16.3 | [6 empty] | 5.3667 | 61.58
-16.2 | [6 empty] | 5.6333 | 61.8
-16.1 | [6 empty] | 5.88 | 62.0143
-16 | [6 empty] | 6.32 | 62.22
-15.9 | [6 empty] | 6.84 | 62.5
-15.8 | [6 empty] | 7.8 | 62.65
-15.7 | [6 empty] | 9.275 | 62.9143
-15.6 | [6 empty] | 12.4333 | 63.08
-15.5 | [6 empty] | 16.825 | 63.32
-15.4 | [6 empty] | 23.3 | 63.5667
-15.3 | [6 empty] | 30.8 | 63.8
-15.2 | [6 empty] | 38.62 | 64.02
-15.1 | [5 empty] | 2.2 | 45.8 | 64.2
-15 | [5 empty] | 2.2 | 52.1429 | 64.5333
-14.9 | [5 empty] | 2.2 | 57.1 | 64.7429
-14.8 | [5 empty] | 2.2667 | 60 | 64.9
-14.7 | [5 empty] | 2.3 | 61.6429 | 65.1
-14.6 | [5 empty] | 2.3 | 62.5667 | 65.34
-14.5 | [5 empty] | 2.3 | 63.2333 | 65.6167
-14.4 | [5 empty] | 2.4 | 63.8 | 65.8667
-14.3 | [5 empty] | 2.4 | 64.2125 | 66.04
-14.2 | [5 empty] | 2.5 | 64.65 | 66.25
-14.1 | [5 empty] | 2.5 | 64.9857 | 66.52
-14 | [5 empty] | 2.5 | 65.3 | 66.775
-13.9 | [5 empty] | 2.6 | 65.66 | 67.0167
-13.8 | [5 empty] | 2.6 | 65.9667 | 67.1833
-13.7 | [5 empty] | 2.7 | 66.2333 | 67.45
-13.6 | [5 empty] | 2.7 | 66.5667 | 67.74
-13.5 | [5 empty] | 2.8 | 66.8 | 67.9
-13.4 | [5 empty] | 2.8 | 67.0667 | 68.1833
-13.3 | [5 empty] | 2.9 | 67.3429 | 68.3667
-13.2 | [5 empty] | 2.9667 | 67.66 | 68.6667
-13.1 | [5 empty] | 3 | 67.9286 | 68.9
-13 | [5 empty] | 3.1 | 68.1429 | 69.12
-12.9 | [5 empty] | 3.1 | 68.42 | 69.38
-12.8 | [5 empty] | 3.2 | 68.74 | 69.66
-12.7 | [5 empty] | 3.3 | 69 | 69.88
-12.6 | [5 empty] | 3.3 | 69.2714 | 70.08
-12.5 | [5 empty] | 3.4 | 69.5 | 70.3667
-12.4 | [5 empty] | 3.5 | 69.82 | 70.6
-12.3 | [5 empty] | 3.6 | 70.0714 | 70.7833
-12.2 | [5 empty] | 3.7 | 70.35 | 71.05
-12.1 | [5 empty] | 3.8 | 70.575 | 71.3333
-12 | [5 empty] | 3.9 | 70.8667 | 71.6
-11.9 | [5 empty] | 4 | 71.125 | 71.8429
-11.8 | [5 empty] | 4.1 | 71.425 | 72.14
-11.7 | [5 empty] | 4.2 | 71.7 | 72.3333
-11.6 | [5 empty] | 4.3 | 71.95 | 72.5
-11.5 | [5 empty] | 4.475 | 72.2286 | 72.8333
-11.4 | [5 empty] | 4.6 | 72.55 | 73.04
-11.3 | [5 empty] | 4.8 | 72.7857 | 73.3333
-11.2 | [5 empty] | 5 | 73.06 | 73.5
-11.1 | [5 empty] | 5.2 | 73.2667 | 73.8
-11 | [5 empty] | 5.6 | 73.65 | 74.1167
-10.9 | [5 empty] | 6.05 | 73.85 | 74.3667
-10.8 | [5 empty] | 6.725 | 74.1333 | 74.65
-10.7 | [5 empty] | 7.8 | 74.425 | 74.85
-10.6 | [5 empty] | 9.9 | 74.7667 | 75.1714
-10.5 | [5 empty] | 13.525 | 75 | 75.4333
-10.4 | [5 empty] | 20.0333 | 75.3 | 75.7
-10.3 | [5 empty] | 28.0333 | 75.6 | 75.8667
-10.2 | [5 empty] | 37.275 | 75.9 | 76.14
-10.1 | [5 empty] | 46.9 | 76.14 | 76.4
-10 |  |  |  |  | 1.3 | 54.9 | 76.425 | 76.64
-9.9 |  |  |  |  | 1.3 | 63.075 | 76.66 | 77
-9.8 |  |  |  |  | 1.4 | 68.56 | 76.96 | 77.3
-9.7 |  |  |  |  | 1.4 | 71.7 | 77.2333 | 77.5
-9.6 |  |  |  |  | 1.4 | 73.4333 | 77.54 | 77.78
-9.5 |  |  |  |  | 1.4 | 74.5 | 77.78 | 78.0333
-9.4 |  |  |  |  | 1.5 | 75.2 | 78.1286 | 78.3
-9.3 |  |  |  |  | 1.5 | 76 | 78.4143 | 78.5667
-9.2 |  |  |  |  | 1.5 | 76.54 | 78.725 | 78.84
-9.1 |  |  |  |  | 1.5 | 76.95 | 78.9333 | 79.12
-9 |  |  |  | 1 | 1.6 | 77.3667 | 79.32 | 79.475
-8.9 |  |  |  | 1.1 | 1.6 | 77.8 | 79.5286 | 79.725
-8.8 |  |  |  | 1.1 | 1.6 | 78.1667 | 79.85 | 79.94
-8.7 |  |  |  | 1.1 | 1.6833 | 78.55 | 80.16 | 80.2
-8.6 |  |  |  | 1.1 | 1.7 | 79 | 80.4 | 80.4
-8.5 |  |  |  | 1.1 | 1.7 | 79.3333 | 80.6714 | 80.8
-8.4 |  |  |  | 1.2 | 1.8 | 79.7 | 80.9429 | 81.04
-8.3 |  |  |  | 1.2 | 1.8 | 80.075 | 81.28 | 81.3167
-8.2 |  |  |  | 1.2 | 1.8 | 80.34 | 81.5286 | 81.5667
-8.1 |  |  |  | 1.2 | 1.9 | 80.675 | 81.8 | 81.86
-8 |  |  | 0.8 | 1.3 | 1.9 | 81.0667 | 82.2 | 82.16
-7.9 |  |  | 0.8 | 1.3 | 1.975 | 81.4333 | 82.4429 | 82.44
-7.8 |  |  | 0.8 | 1.3 | 2 | 81.7667 | 82.7429 | 82.8
-7.7 |  |  | 0.8 | 1.3 | 2.1 | 82.05 | 83.0333 | 83.12
-7.6 |  |  | 0.8 | 1.35 | 2.1 | 82.375 | 83.3 | 83.3571
-7.5 |  |  | 0.8 | 1.4 | 2.2 | 82.8 | 83.6333 | 83.5833
-7.4 |  |  | 0.9 | 1.4 | 2.3 | 83.025 | 83.95 | 83.9
-7.3 |  |  | 0.9 | 1.4 | 2.3 | 83.3 | 84.16 | 84.14
-7.2 |  |  | 0.9 | 1.5 | 2.4 | 83.65 | 84.5 | 84.52
-7.1 |  |  | 0.9 | 1.5 | 2.5 | 84.0333 | 84.8571 | 84.725
-7 |  | 0.5 | 0.9 | 1.575 | 2.6 | 84.325 | 85.0714 | 85.1
-6.9 |  | 0.5 | 0.9 | 1.6 | 2.6333 | 84.68 | 85.5 | 85.325
-6.8 |  | 0.5 | 1 | 1.6333 | 2.7333 | 85.0333 | 85.76 | 85.675
-6.7 |  | 0.5 | 1 | 1.7 | 2.8833 | 85.375 | 86.06 | 85.875
-6.6 |  | 0.5 | 1 | 1.7667 | 3 | 85.7 | 86.3 | 86.2
-6.5 | 0.4 | 0.5 | 1 | 1.8 | 3.1 | 85.925 | 86.54 | 86.5
-6.4 | 0.4 | 0.6 | 1 | 1.85 | 3.2 | 86.325 | 86.925 | 86.8333
-6.3 | 0.4 | 0.6 | 1.1 | 1.9 | 3.375 | 86.6 | 87.2333 | 87.1333
-6.2 | 0.4 | 0.6 | 1.1 | 2 | 3.525 | 86.875 | 87.4 | 87.4
-6.1 | 0.4 | 0.6 | 1.1 | 2.04 | 3.825 | 87.2667 | 87.8 | 87.8
-6 | 0.4 | 0.6 | 1.2 | 2.1333 | 4.1 | 87.55 | 88.125 | 88.02
-5.9 | 0.4 | 0.6 | 1.2 | 2.2 | 4.4333 | 87.8667 | 88.4333 | 88.22
-5.8 | 0.4 | 0.6 | 1.2 | 2.3 | 5 | 88.1 | 88.7 | 88.525
-5.7 | 0.4 | 0.6 | 1.3 | 2.4 | 5.8 | 88.4333 | 89 | 88.8333
-5.6 | 0.4 | 0.7 | 1.3 | 2.5 | 7.1 | 88.7333 | 89.2 | 89.0833
-5.5 | 0.4 | 0.7 | 1.4 | 2.6 | 9.425 | 89.1 | 89.56 | 89.4
-5.4 | 0.5 | 0.7 | 1.4 | 2.7857 | 14 | 89.4 | 89.8 | 89.65
-5.3 | 0.5 | 0.7 | 1.5 | 2.9 | 21.8 | 89.7333 | 90.075 | 89.92
-5.2 | 0.5 | 0.7 | 1.5 | 3 | 32.225 | 90.05 | 90.4333 | 90.3
-5.1 | 0.5 | 0.7 | 1.6 | 3.3 | 43.7333 | 90.3 | 90.675 | 90.5
-5 | 0.5 | 0.8 | 1.6 | 3.5 | 55.25 | 90.5667 | 91 | 90.8333
-4.9 | 0.5 | 0.8 | 1.7 | 3.8667 | 66.02 | 90.9 | 91.2429 | 91.1667
-4.8 | 0.5 | 0.8 | 1.8 | 4.3 | 75.375 | 91.2 | 91.56 | 91.38
-4.7 | 0.5 | 0.8 | 1.9 | 5 | 82.4 | 91.475 | 91.8 | 91.6
-4.6 | 0.5 | 0.9 | 2 | 6.2333 | 86.075 | 91.775 | 92.12 | 91.9
-4.5 | 0.6 | 0.9 | 2.1 | 8.2 | 88 | 92.0667 | 92.4 | 92.1
-4.4 | 0.6 | 0.9 | 2.2 | 12.4 | 89.3 | 92.35 | 92.6571 | 92.4667
-4.3 | 0.6 | 1 | 2.3 | 19.3 | 90.175 | 92.7 | 92.9286 | 92.7
-4.2 | 0.6 | 1 | 2.4 | 29.2 | 90.8667 | 92.9 | 93.26 | 92.95
-4.1 | 0.6 | 1 | 2.6 | 40.9333 | 91.3667 | 93.1667 | 93.5 | 93.2333
-4 | 0.7 | 1.1 | 2.8667 | 53.1 | 91.9 | 93.5 | 93.7143 | 93.5667
-3.9 | 0.7 | 1.1 | 3.2 | 64.6 | 92.3333 | 93.75 | 93.98 | 93.85
-3.8 | 0.7 | 1.2 | 3.6 | 75.275 | 92.72 | 94.0333 | 94.34 | 94.1
-3.7 | 0.7 | 1.2 | 4.2667 | 83 | 93.06 | 94.2667 | 94.5 | 94.3286
-3.6 | 0.8 | 1.3 | 5.3 | 87.9667 | 93.48 | 94.6 | 94.7571 | 94.6333
-3.5 | 0.8 | 1.4 | 7 | 90.4 | 93.9 | 94.8333 | 95.04 | 94.85
-3.4 | 0.8 | 1.5 | 10.5 | 91.8 | 94.2333 | 95.12 | 95.3 | 95.1
-3.3 | 0.9 | 1.6 | 16.9 | 92.8667 | 94.5 | 95.3667 | 95.56 | 95.3667
-3.2 | 0.9 | 1.7 | 27.3 | 93.55 | 94.9 | 95.62 | 95.8 | 95.6
-3.1 | 1 | 1.8333 | 38.45 | 94.1 | 95.14 | 95.84 | 95.9333 | 95.82
-3 | 1 | 2 | 51.4 | 94.6 | 95.4 | 96.06 | 96.2667 | 96.075
-2.9 | 1.1 | 2.3 | 63.7 | 95.025 | 95.7 | 96.35 | 96.45 | 96.26
-2.8 | 1.2 | 2.7 | 75.4667 | 95.3667 | 95.9667 | 96.55 | 96.62 | 96.54
-2.7 | 1.3 | 3.2667 | 84.0667 | 95.7 | 96.35 | 96.775 | 96.85 | 96.75
-2.6 | 1.4 | 4.25 | 89.8 | 96 | 96.5 | 97.0167 | 97.1 | 96.86
-2.5 | 1.6 | 5.9 | 92.6 | 96.35 | 96.8333 | 97.2667 | 97.2667 | 97.06
-2.4 | 1.9 | 9 | 94.4 | 96.725 | 97.075 | 97.4333 | 97.56 | 97.3
-2.3 | 2.1667 | 15 | 95.3 | 97 | 97.3 | 97.6667 | 97.7 | 97.5
-2.2 | 2.8 | 25.2 | 96.05 | 97.2 | 97.475 | 97.8 | 97.9 | 97.7333
-2.1 | 3.6333 | 36.95 | 96.5 | 97.4333 | 97.7 | 98.0333 | 98.1 | 97.9333
-2 | 5.2 | 50.3 | 96.9667 | 97.725 | 97.9 | 98.2 | 98.24 | 98.0667
-1.9 | 7.9667 | 63.8 | 97.4 | 97.9 | 98.175 | 98.375 | 98.4 | 98.3
-1.8 | 13.5 | 75.8 | 97.7 | 98.2 | 98.3333 | 98.5 | 98.6 | 98.4
-1.7 | 23.7 | 85.2 | 97.95 | 98.3 | 98.475 | 98.65 | 98.7 | 98.5333
-1.6 | 35.45 | 91.6 | 98.2 | 98.6 | 98.7 | 98.86 | 98.8667 | 98.7167
-1.5 | 49.5 | 94.9 | 98.45 | 98.75 | 98.8667 | 98.98 | 98.96 | 98.8833
-1.4 | 62.3 | 96.4 | 98.7 | 98.88 | 99 | 99.1 | 99.1286 | 99.05
-1.3 | 74.6667 | 97.6 | 98.9 | 99.0667 | 99.2 | 99.2667 | 99.3 | 99.1167
-1.2 | 85.3 | 98.1 | 99 | 99.2 | 99.2 | 99.4 | 99.4 | 99.2833
-1.1 | 92.1 | 98.65 | 99.2 | 99.3 | 99.4 | 99.5 | 99.5 | 99.35
-1 | 95.45 | 98.95 | 99.3 | 99.4667 | 99.4333 | 99.6 | 99.5429 | 99.4
-0.9 | 97.2 | 99.2 | 99.45 | 99.525 | 99.5667 | 99.7 | 99.6667 | 99.6
-0.8 | 98.2 | 99.4 | 99.6 | 99.7 | 99.6333 | 99.76 | 99.7875 | 99.6
-0.7 | 98.9 | 99.6 | 99.7 | 99.7 | 99.7 | 99.8 | 99.8 | 99.7
-0.6 | 99.3 | 99.7 | 99.8 | 99.8 | 99.8 | 99.825 | 99.88 | 99.7
-0.5 | 99.6 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9 | 99.85
-0.4 | 99.7 | 99.9 | 99.9 | 99.9 | 100 | 100 | 99.9 | 99.9
-0.3 | 99.8667 | 100 | 99.9 | 100 | 100 | 100 | 100 | 99.9
-0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9333
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 99.96
0.3 | 99.8333 | 99.9 | 99.9 | 99.9 | 99.9667 | 99.95 | 99.9125 | 99.9
0.4 | 99.7 | 99.9 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9
0.5 | 99.5 | 99.8 | 99.8 | 99.8 | 99.9 | 99.9 | 99.8 | 99.8
0.6 | 99.3 | 99.6667 | 99.7 | 99.8 | 99.8 | 99.8 | 99.8 | 99.7167
0.7 | 98.8 | 99.5 | 99.6 | 99.7 | 99.7 | 99.7 | 99.74 | 99.7
0.8 | 98.1 | 99.3 | 99.6 | 99.6 | 99.6 | 99.7 | 99.7 | 99.6
0.9 | 97.1 | 99.1 | 99.4 | 99.5 | 99.5 | 99.6 | 99.5857 | 99.525
1 | 95.3 | 98.9 | 99.25 | 99.4 | 99.4 | 99.5 | 99.5 | 99.4333
1.1 | 92 | 98.5 | 99.1 | 99.275 | 99.3 | 99.4 | 99.34 | 99.3833
1.2 | 84.95 | 98.05 | 98.9 | 99.0667 | 99.2 | 99.2667 | 99.3 | 99.3
1.3 | 74.5 | 97.4 | 98.7 | 98.9333 | 99.08 | 99.1 | 99.22 | 99.15
1.4 | 62.15 | 96.3 | 98.5 | 98.8 | 98.9 | 99 | 99 | 99.0833
1.5 | 48.7 | 94.7 | 98.3 | 98.6 | 98.7 | 98.8667 | 98.94 | 98.9167
1.6 | 35.45 | 91.35 | 98 | 98.45 | 98.66 | 98.6 | 98.7833 | 98.75
1.7 | 23.5 | 85.1 | 97.8 | 98.2 | 98.4 | 98.5 | 98.6167 | 98.5
1.8 | 13.8 | 75.5333 | 97.5 | 98.05 | 98.15 | 98.4 | 98.4 | 98.4
1.9 | 8 | 63.4 | 97.15 | 97.7667 | 98 | 98.2333 | 98.3167 | 98.3
2 | 5.1 | 50.5 | 96.7 | 97.62 | 97.8 | 98 | 98.14 | 98.06
2.1 | 3.6 | 37.0667 | 96.4 | 97.35 | 97.5333 | 97.875 | 97.9 | 97.9
2.2 | 2.75 | 24.9 | 95.8 | 97 | 97.2833 | 97.7 | 97.72 | 97.76
2.3 | 2.15 | 14.8 | 95 | 96.8 | 97.1 | 97.4667 | 97.6 | 97.54
2.4 | 1.85 | 8.95 | 94 | 96.5 | 96.9 | 97.3 | 97.28 | 97.32
2.5 | 1.6 | 5.85 | 92.5 | 96.3 | 96.65 | 97.1 | 97.1 | 97.1
2.6 | 1.4 | 4.2 | 89.5 | 95.925 | 96.3333 | 96.8333 | 96.8833 | 96.9
2.7 | 1.3 | 3.3 | 84.2 | 95.5667 | 96.1 | 96.575 | 96.7167 | 96.675
2.8 | 1.2 | 2.65 | 74.5 | 95.225 | 95.8 | 96.3 | 96.5 | 96.5333
2.9 | 1.1 | 2.3 | 64.4 | 94.9 | 95.5 | 96.1667 | 96.2 | 96.3
3 | 1 | 2 | 51.6333 | 94.35 | 95.225 | 95.9 | 96 | 96.1
3.1 | 1 | 1.8 | 38.45 | 93.8333 | 94.9 | 95.6 | 95.7667 | 95.7833
3.2 | 0.9 | 1.7 | 27 | 93.2333 | 94.6 | 95.35 | 95.5167 | 95.5571
3.3 | 0.9 | 1.6 | 16.5 | 92.6 | 94.26 | 95.1 | 95.2857 | 95.3667
3.4 | 0.8 | 1.5 | 10.5 | 91.6 | 93.925 | 94.8 | 95.05 | 95.08
3.5 | 0.8 | 1.4 | 7 | 90.1667 | 93.5667 | 94.5667 | 94.8 | 94.88
3.6 | 0.8 | 1.3 | 5.15 | 87.8 | 93.2 | 94.34 | 94.55 | 94.6
3.7 | 0.7 | 1.25 | 4.2 | 83.15 | 92.9 | 94 | 94.26 | 94.3857
3.8 | 0.7 | 1.2 | 3.6 | 75.0333 | 92.4 | 93.7667 | 93.98 | 94.0833
3.9 | 0.7 | 1.1 | 3.1 | 65.0333 | 92.15 | 93.55 | 93.725 | 93.8333
4 | 0.7 | 1.1 | 2.85 | 53.1 | 91.6 | 93.3 | 93.44 | 93.5333
4.1 | 0.6 | 1 | 2.6 | 40.3 | 91.2 | 92.9 | 93.2 | 93.25
4.2 | 0.6 | 1 | 2.4 | 28.5 | 90.5667 | 92.6 | 92.9333 | 92.9667
4.3 | 0.6 | 0.9 | 2.3 | 19.0333 | 89.8667 | 92.4 | 92.6667 | 92.7167
4.4 | 0.6 | 0.9 | 2.1 | 11.8 | 88.95 | 92.1 | 92.36 | 92.5
4.5 | 0.6 | 0.9 | 2 | 8.2 | 87.725 | 91.8 | 92.05 | 92.14
4.6 | 0.6 | 0.9 | 1.9 | 6.1667 | 85.7 | 91.52 | 91.75 | 91.9
4.7 | 0.5 | 0.8 | 1.8 | 5 | 82.1 | 91.2 | 91.55 | 91.7
4.8 | 0.5 | 0.8 | 1.8 | 4.32 | 75.1 | 90.9 | 91.3 | 91.36
4.9 | 0.5 | 0.8 | 1.7 | 3.85 | 65.7 | 90.6333 | 90.92 | 91.1143
5 | 0.5 | 0.8 | 1.6 | 3.45 | 55.3833 | 90.34 | 90.6667 | 90.8
5.1 | 0.5 | 0.7 | 1.6 | 3.2667 | 43.225 | 89.975 | 90.46 | 90.5667
5.2 | 0.5 | 0.7 | 1.5 | 2.9833 | 32.3 | 89.7 | 90.1 | 90.225
5.3 | 0.5 | 0.7 | 1.5 | 2.9 | 21.9667 | 89.4 | 89.7667 | 89.96
5.4 | 0.5 | 0.7 | 1.4 | 2.7 | 14.24 | 89.08 | 89.5 | 89.66
5.5 | 0.5 | 0.7 | 1.4 | 2.6 | 9.575 | 88.7667 | 89.275 | 89.3833
5.6 | 0.4 | 0.7 | 1.3 | 2.5 | 7.05 | 88.5 | 88.96 | 89.1333
5.7 | 0.4 | 0.6 | 1.3 | 2.38 | 5.7 | 88.2 | 88.5833 | 88.8
5.8 | 0.4 | 0.6 | 1.2 | 2.3 | 5 | 87.9 | 88.3 | 88.46
5.9 | 0.4 | 0.6 | 1.2 | 2.2 | 4.45 | 87.575 | 88.025 | 88.26
6 | 0.4 | 0.6 | 1.2 | 2.1 | 4.0667 | 87.3 | 87.7 | 88
6.1 | 0.4 | 0.6 | 1.1 | 2.075 | 3.75 | 86.98 | 87.4333 | 87.7333
6.2 | 0.4 | 0.6 | 1.1 | 1.9667 | 3.5667 | 86.64 | 87.1667 | 87.4
6.3 | 0.4 | 0.6 | 1.1 | 1.9 | 3.4 | 86.325 | 86.8333 | 87.0571
6.4 | 0.4 | 0.6 | 1.1 | 1.825 | 3.2 | 86 | 86.6 | 86.8
6.5 | 0.4 | 0.5 | 1 | 1.8 | 3.1 | 85.6333 | 86.28 | 86.4667
6.6 |  | 0.5 | 1 | 1.7 | 3 | 85.3667 | 86.02 | 86.22
6.7 |  | 0.5 | 1 | 1.7 | 2.82 | 85.0333 | 85.7 | 85.9
6.8 |  | 0.5 | 1 | 1.6 | 2.7667 | 84.7 | 85.46 | 85.6
6.9 |  | 0.5 | 0.9 | 1.6 | 2.6 | 84.3667 | 85.0714 | 85.32
7 |  | 0.5 | 0.9 | 1.5 | 2.55 | 84.06 | 84.7857 | 85.04
7.1 |  |  | 0.9 | 1.5 | 2.46 | 83.75 | 84.54 | 84.75
7.2 |  |  | 0.9 | 1.5 | 2.4 | 83.425 | 84.2 | 84.48
7.3 |  |  | 0.9 | 1.4 | 2.3 | 83.0667 | 83.9667 | 84.2
7.4 |  |  | 0.9 | 1.4 | 2.2333 | 82.825 | 83.58 | 83.8833
7.5 |  |  | 0.8 | 1.4 | 2.2 | 82.4 | 83.3286 | 83.6333
7.6 |  |  | 0.8 | 1.3 | 2.1 | 82.1 | 83.0333 | 83.3667
7.7 |  |  | 0.8 | 1.3 | 2.1 | 81.8 | 82.66 | 83.05
7.8 |  |  | 0.8 | 1.3 | 2 | 81.45 | 82.4 | 82.8
7.9 |  |  | 0.8 | 1.3 | 1.9 | 81.15 | 82.2 | 82.4429
8 |  |  | 0.8 | 1.2 | 1.9 | 80.8333 | 81.86 | 82.1
8.1 |  |  |  | 1.2 | 1.9 | 80.4667 | 81.56 | 81.8833
8.2 |  |  |  | 1.2 | 1.8 | 80.1333 | 81.28 | 81.55
8.3 |  |  |  | 1.2 | 1.8 | 79.7 | 80.92 | 81.35
8.4 |  |  |  | 1.1 | 1.7 | 79.375 | 80.65 | 81.05
8.5 |  |  |  | 1.1 | 1.7 | 79 | 80.34 | 80.78
8.6 |  |  |  | 1.1 | 1.7 | 78.6667 | 80.0667 | 80.44
8.7 |  |  |  | 1.1 | 1.6 | 78.3333 | 79.7333 | 80.24
8.8 |  |  |  | 1.1 | 1.6 | 77.9 | 79.5333 | 79.95
8.9 |  |  |  | 1.1 | 1.6 | 77.575 | 79.2667 | 79.75
9 |  |  |  | 1 | 1.525 | 77.15 | 78.9714 | 79.4
9.1 |  |  |  |  | 1.5 | 76.7 | 78.6667 | 79.1
9.2 |  |  |  |  | 1.5 | 76.1667 | 78.325 | 78.8333
9.3 |  |  |  |  | 1.5 | 75.7 | 78.04 | 78.6
9.4 |  |  |  |  | 1.4 | 75.06 | 77.84 | 78.3333
9.5 |  |  |  |  | 1.4 | 74.2667 | 77.5125 | 78.05
9.6 |  |  |  |  | 1.4 | 73.1333 | 77.2778 | 77.775
9.7 |  |  |  |  | 1.4 | 71.6 | 77.0333 | 77.54
9.8 |  |  |  |  | 1.3 | 68.4 | 76.7 | 77.24
9.9 |  |  |  |  | 1.3 | 62.85 | 76.4 | 76.98
10 |  |  |  |  | 1.3 | 55.4 | 76.1667 | 76.7
10.1 | [5 empty] | 46.8 | 75.8333 | 76.4
10.2 | [5 empty] | 37.3667 | 75.5571 | 76.1
10.3 | [5 empty] | 28.1 | 75.3 | 75.9
10.4 | [5 empty] | 19.8 | 75.04 | 75.6667
10.5 | [5 empty] | 13.65 | 74.72 | 75.3667
10.6 | [5 empty] | 9.72 | 74.425 | 75.2
10.7 | [5 empty] | 7.85 | 74.2 | 74.9
10.8 | [5 empty] | 6.7333 | 73.9 | 74.6667
10.9 | [5 empty] | 6.0333 | 73.65 | 74.3333
11 | [5 empty] | 5.56 | 73.375 | 74.1143
11.1 | [5 empty] | 5.2 | 73.06 | 73.85
11.2 | [5 empty] | 5 | 72.82 | 73.5286
11.3 | [5 empty] | 4.7667 | 72.5 | 73.3333
11.4 | [5 empty] | 4.6 | 72.225 | 73.08
11.5 | [5 empty] | 4.425 | 71.9571 | 72.8667
11.6 | [5 empty] | 4.3 | 71.7 | 72.55
11.7 | [5 empty] | 4.2 | 71.4667 | 72.36
11.8 | [5 empty] | 4.1 | 71.2 | 72.08
11.9 | [5 empty] | 3.92 | 70.8667 | 71.8286
12 | [5 empty] | 3.8 | 70.6667 | 71.5667
12.1 | [5 empty] | 3.7 | 70.3333 | 71.2667
12.2 | [5 empty] | 3.7 | 70.0714 | 71.075
12.3 | [5 empty] | 3.5333 | 69.7857 | 70.825
12.4 | [5 empty] | 3.5 | 69.56 | 70.625
12.5 | [5 empty] | 3.4 | 69.275 | 70.3667
12.6 | [5 empty] | 3.3 | 68.98 | 70.04
12.7 | [5 empty] | 3.2 | 68.66 | 69.92
12.8 | [5 empty] | 3.2 | 68.425 | 69.6333
12.9 | [5 empty] | 3.1 | 68.1667 | 69.32
13 | [5 empty] | 3.02 | 67.9167 | 69.1
13.1 | [5 empty] | 3 | 67.62 | 68.9333
13.2 | [5 empty] | 2.9 | 67.3571 | 68.6333
13.3 | [5 empty] | 2.9 | 67.0667 | 68.4
13.4 | [5 empty] | 2.8 | 66.8 | 68.2
13.5 | [5 empty] | 2.775 | 66.5 | 67.95
13.6 | [5 empty] | 2.7 | 66.28 | 67.675
13.7 | [5 empty] | 2.65 | 65.96 | 67.4833
13.8 | [5 empty] | 2.6 | 65.7 | 67.2333
13.9 | [5 empty] | 2.6 | 65.36 | 67
14 | [5 empty] | 2.5 | 65.08 | 66.8
14.1 | [5 empty] | 2.5 | 64.74 | 66.5
14.2 | [5 empty] | 2.4 | 64.38 | 66.26
14.3 | [5 empty] | 2.4 | 64.02 | 66.05
14.4 | [5 empty] | 2.4 | 63.5333 | 65.82
14.5 | [5 empty] | 2.3 | 63 | 65.6
14.6 | [5 empty] | 2.3 | 62.3625 | 65.375
14.7 | [5 empty] | 2.275 | 61.425 | 65.125
14.8 | [5 empty] | 2.2 | 59.88 | 64.9333
14.9 | [5 empty] | 2.2 | 57.06 | 64.7
15 | [5 empty] | 2.1 | 52.2 | 64.5
15.1 | [5 empty] | 2.1 | 45.85 | 64.2429
15.2 | [6 empty] | 38.7 | 64.02
15.3 | [6 empty] | 31.4 | 63.8
15.4 | [6 empty] | 23.8 | 63.6
15.5 | [6 empty] | 17.1 | 63.28
15.6 | [6 empty] | 12.4 | 63.1143
15.7 | [6 empty] | 9.3 | 62.9
15.8 | [6 empty] | 7.7 | 62.6833
15.9 | [6 empty] | 6.78 | 62.4571
16 | [6 empty] | 6.24 | 62.22
16.1 | [6 empty] | 5.8286 | 62
16.2 | [6 empty] | 5.5429 | 61.8333
16.3 | [6 empty] | 5.3167 | 61.6
16.4 | [6 empty] | 5.12 | 61.32
16.5 | [6 empty] | 4.95 | 61.1333
16.6 | [6 empty] | 4.8 | 60.925
16.7 | [6 empty] | 4.7 | 60.75
16.8 | [6 empty] | 4.56 | 60.48
16.9 | [6 empty] | 4.4571 | 60.2833
17 | [6 empty] | 4.4 | 60.0333
17.1 | [6 empty] | 4.3 | 59.85
17.2 | [6 empty] | 4.2 | 59.58
17.3 | [6 empty] | 4.1 | 59.4
17.4 | [6 empty] | 4 | 59.2333
17.5 | [6 empty] | 3.9167 | 58.9667
17.6 | [6 empty] | 3.8 | 58.7
17.7 | [6 empty] | 3.7333 | 58.5333
17.8 | [6 empty] | 3.675 | 58.25
17.9 | [6 empty] | 3.6 | 58.0667
18 | [6 empty] | 3.5 | 57.8286
18.1 | [6 empty] | 3.46 | 57.65
18.2 | [6 empty] | 3.4 | 57.4667
18.3 | [6 empty] | 3.32 | 57.2167
18.4 | [6 empty] | 3.3 | 56.925
18.5 | [6 empty] | 3.2 | 56.75
18.6 | [6 empty] | 3.2 | 56.5
18.7 | [6 empty] | 3.1 | 56.25
18.8 | [6 empty] | 3.1 | 56.04
18.9 | [6 empty] | 3 | 55.82
19 | [6 empty] | 2.98 | 55.5667
19.1 | [6 empty] | 2.9 | 55.3
19.2 | [6 empty] | 2.9 | 55.0667
19.3 | [6 empty] | 2.8 | 54.75
19.4 | [6 empty] | 2.8 | 54.35
19.5 | [6 empty] | 2.7286 | 54.04
19.6 | [6 empty] | 2.7 | 53.6
19.7 | [6 empty] | 2.7 | 52.9167
19.8 | [6 empty] | 2.6111 | 51.9286
19.9 | [6 empty] | 2.6 | 50
20 | [6 empty] | 2.6 | 46.9
20.1 | [6 empty] | 2.5 | 42
20.2 | [7 empty] | 35.8333
20.3 | [7 empty] | 29.4
20.4 | [7 empty] | 23.2333
20.5 | [7 empty] | 17.3857
20.6 | [7 empty] | 12.85
20.7 | [7 empty] | 9.46
20.8 | [7 empty] | 7.7833
20.9 | [7 empty] | 6.8667
21 | [7 empty] | 6.2
21.1 | [7 empty] | 5.8333
21.2 | [7 empty] | 5.5571
21.3 | [7 empty] | 5.375
21.4 | [7 empty] | 5.16
21.5 | [7 empty] | 5
21.6 | [7 empty] | 4.9
21.7 | [7 empty] | 4.7333
21.8 | [7 empty] | 4.6333
21.9 | [7 empty] | 4.5429
22 | [7 empty] | 4.4
22.1 | [7 empty] | 4.3
22.2 | [7 empty] | 4.2
22.3 | [7 empty] | 4.1167
22.4 | [7 empty] | 4.02
22.5 | [7 empty] | 3.95
22.6 | [7 empty] | 3.9
22.7 | [7 empty] | 3.8
22.8 | [7 empty] | 3.7333
22.9 | [7 empty] | 3.7
23 | [7 empty] | 3.6
23.1 | [7 empty] | 3.575
23.2 | [7 empty] | 3.5
23.3 | [7 empty] | 3.4
23.4 | [7 empty] | 3.4
23.5 | [7 empty] | 3.3167
23.6 | [7 empty] | 3.3
23.7 | [7 empty] | 3.2
23.8 | [7 empty] | 3.2
23.9 | [7 empty] | 3.1
24 | [7 empty] | 3.1
24.1 | [7 empty] | 3.1
24.2 | [7 empty] | 3
24.3 | [7 empty] | 3
24.4 | [7 empty] | 2.9
24.5 | [7 empty] | 2.9
24.6 | [7 empty] | 2.8
24.7 | [7 empty] | 2.8
24.8 | [7 empty] | 2.8
24.9 | [7 empty] | 2.7
25 | [7 empty] | 2.7
25.1 | [7 empty] | 2.7
25.2 | [7 empty] | 2.6

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-25.9 | [7 empty] | 3.5143
-25.8 | [7 empty] | 3.6
-25.7 | [7 empty] | 3.6
-25.6 | [7 empty] | 3.7
-25.5 | [7 empty] | 3.7375
-25.4 | [7 empty] | 3.8
-25.3 | [7 empty] | 3.8375
-25.2 | [7 empty] | 3.9
-25.1 | [7 empty] | 3.9286
-25 | [7 empty] | 4
-24.9 | [7 empty] | 4.1
-24.8 | [7 empty] | 4.1571
-24.7 | [7 empty] | 4.2
-24.6 | [7 empty] | 4.2857
-24.5 | [7 empty] | 4.3333
-24.4 | [7 empty] | 4.4
-24.3 | [7 empty] | 4.5
-24.2 | [7 empty] | 4.5625
-24.1 | [7 empty] | 4.6125
-24 | [7 empty] | 4.7429
-23.9 | [7 empty] | 4.8
-23.8 | [7 empty] | 4.925
-23.7 | [7 empty] | 5
-23.6 | [7 empty] | 5.1
-23.5 | [7 empty] | 5.1833
-23.4 | [7 empty] | 5.325
-23.3 | [7 empty] | 5.4
-23.2 | [7 empty] | 5.5143
-23.1 | [7 empty] | 5.6667
-23 | [7 empty] | 5.8125
-22.9 | [7 empty] | 5.9
-22.8 | [7 empty] | 6.075
-22.7 | [7 empty] | 6.2
-22.6 | [7 empty] | 6.3143
-22.5 | [7 empty] | 6.4667
-22.4 | [7 empty] | 6.64
-22.3 | [7 empty] | 6.9
-22.2 | [7 empty] | 7.0429
-22.1 | [7 empty] | 7.3
-22 | [7 empty] | 7.52
-21.9 | [7 empty] | 7.84
-21.8 | [7 empty] | 8.2
-21.7 | [7 empty] | 8.6857
-21.6 | [7 empty] | 9.3429
-21.5 | [7 empty] | 10.2833
-21.4 | [7 empty] | 12.1143
-21.3 | [7 empty] | 15.1143
-21.2 | [7 empty] | 19.38
-21.1 | [7 empty] | 24.6
-21 | [7 empty] | 30.0571
-20.9 | [7 empty] | 35.7833
-20.8 | [7 empty] | 40.7
-20.7 | [6 empty] | 3.4 | 45.1667
-20.6 | [6 empty] | 3.4 | 48.2333
-20.5 | [6 empty] | 3.5 | 50.0125
-20.4 | [6 empty] | 3.6 | 51.0625
-20.3 | [6 empty] | 3.6 | 51.8125
-20.2 | [6 empty] | 3.7 | 52.35
-20.1 | [6 empty] | 3.7 | 52.8143
-20 | [6 empty] | 3.8 | 53.2333
-19.9 | [6 empty] | 3.8714 | 53.65
-19.8 | [6 empty] | 3.9 | 53.96
-19.7 | [6 empty] | 4 | 54.32
-19.6 | [6 empty] | 4.1 | 54.6
-19.5 | [6 empty] | 4.125 | 54.8571
-19.4 | [6 empty] | 4.2 | 55.125
-19.3 | [6 empty] | 4.3 | 55.425
-19.2 | [6 empty] | 4.375 | 55.6857
-19.1 | [6 empty] | 4.4286 | 56
-19 | [6 empty] | 4.5 | 56.18
-18.9 | [6 empty] | 4.6 | 56.46
-18.8 | [6 empty] | 4.7 | 56.7
-18.7 | [6 empty] | 4.8 | 56.9833
-18.6 | [6 empty] | 4.9 | 57.2
-18.5 | [6 empty] | 5 | 57.45
-18.4 | [6 empty] | 5.1 | 57.7
-18.3 | [6 empty] | 5.2 | 57.95
-18.2 | [6 empty] | 5.3571 | 58.2143
-18.1 | [6 empty] | 5.45 | 58.3875
-18 | [6 empty] | 5.6 | 58.6286
-17.9 | [6 empty] | 5.7 | 58.9143
-17.8 | [6 empty] | 5.825 | 59.1
-17.7 | [6 empty] | 6 | 59.3667
-17.6 | [6 empty] | 6.15 | 59.55
-17.5 | [6 empty] | 6.3 | 59.8333
-17.4 | [6 empty] | 6.5 | 60.0571
-17.3 | [6 empty] | 6.6 | 60.3667
-17.2 | [6 empty] | 6.84 | 60.55
-17.1 | [6 empty] | 7 | 60.8143
-17 | [6 empty] | 7.2429 | 61.0143
-16.9 | [6 empty] | 7.4286 | 61.2143
-16.8 | [6 empty] | 7.8 | 61.425
-16.7 | [6 empty] | 8.1 | 61.65
-16.6 | [6 empty] | 8.5 | 61.9286
-16.5 | [6 empty] | 9.04 | 62.125
-16.4 | [6 empty] | 9.72 | 62.375
-16.3 | [6 empty] | 10.94 | 62.6222
-16.2 | [6 empty] | 12.86 | 62.85
-16.1 | [6 empty] | 15.975 | 63.05
-16 | [6 empty] | 20.74 | 63.325
-15.9 | [6 empty] | 26.8667 | 63.55
-15.8 | [6 empty] | 32.9 | 63.775
-15.7 | [6 empty] | 39.84 | 64.0167
-15.6 | [6 empty] | 46.52 | 64.2429
-15.5 | [6 empty] | 51.7125 | 64.4286
-15.4 | [5 empty] | 3 | 55.6143 | 64.6571
-15.3 | [5 empty] | 3 | 58.16 | 64.875
-15.2 | [5 empty] | 3.1 | 59.7 | 65.125
-15.1 | [5 empty] | 3.12 | 60.675 | 65.375
-15 | [5 empty] | 3.2 | 61.425 | 65.625
-14.9 | [5 empty] | 3.3 | 62.02 | 65.875
-14.8 | [5 empty] | 3.3 | 62.55 | 66.1111
-14.7 | [5 empty] | 3.3833 | 62.975 | 66.25
-14.6 | [5 empty] | 3.4 | 63.4286 | 66.5
-14.5 | [5 empty] | 3.5 | 63.84 | 66.75
-14.4 | [5 empty] | 3.6 | 64.2286 | 66.94
-14.3 | [5 empty] | 3.66 | 64.55 | 67.175
-14.2 | [5 empty] | 3.7 | 65 | 67.4286
-14.1 | [5 empty] | 3.8 | 65.3143 | 67.6889
-14 | [5 empty] | 3.9 | 65.6 | 67.9667
-13.9 | [5 empty] | 3.9 | 65.95 | 68.2286
-13.8 | [5 empty] | 4 | 66.28 | 68.4111
-13.7 | [5 empty] | 4.0875 | 66.575 | 68.65
-13.6 | [5 empty] | 4.2 | 66.9 | 68.9333
-13.5 | [5 empty] | 4.3 | 67.26 | 69.15
-13.4 | [5 empty] | 4.4 | 67.55 | 69.34
-13.3 | [5 empty] | 4.5 | 67.8143 | 69.6
-13.2 | [5 empty] | 4.6 | 68.15 | 69.86
-13.1 | [5 empty] | 4.7 | 68.48 | 70.0857
-13 | [5 empty] | 4.825 | 68.6875 | 70.3
-12.9 | [5 empty] | 4.975 | 69 | 70.5667
-12.8 | [5 empty] | 5.1 | 69.3 | 70.7857
-12.7 | [5 empty] | 5.2 | 69.6143 | 71.0429
-12.6 | [5 empty] | 5.3167 | 69.9 | 71.2714
-12.5 | [5 empty] | 5.5 | 70.1833 | 71.525
-12.4 | [5 empty] | 5.7 | 70.5 | 71.8125
-12.3 | [5 empty] | 5.8 | 70.8 | 72.0625
-12.2 | [5 empty] | 6 | 71.0667 | 72.275
-12.1 | [5 empty] | 6.2 | 71.28 | 72.55
-12 | [5 empty] | 6.4 | 71.6143 | 72.8
-11.9 | [5 empty] | 6.625 | 71.84 | 72.975
-11.8 | [5 empty] | 6.8 | 72.14 | 73.2
-11.7 | [5 empty] | 7.12 | 72.4 | 73.6
-11.6 | [5 empty] | 7.38 | 72.7 | 73.6833
-11.5 | [5 empty] | 7.75 | 73.0667 | 74.05
-11.4 | [5 empty] | 8.15 | 73.4 | 74.3143
-11.3 | [5 empty] | 8.66 | 73.64 | 74.5
-11.2 | [5 empty] | 9.35 | 73.9 | 74.825
-11.1 | [5 empty] | 10.3 | 74.08 | 75.0571
-11 | [5 empty] | 12 | 74.46 | 75.325
-10.9 | [5 empty] | 14.4 | 74.7429 | 75.525
-10.8 | [5 empty] | 19.1333 | 75 | 75.775
-10.7 | [5 empty] | 25.8333 | 75.3167 | 76.0111
-10.6 | [5 empty] | 33.5 | 75.6 | 76.2714
-10.5 | [5 empty] | 41.625 | 75.8 | 76.55
-10.4 | [5 empty] | 50.1 | 76.175 | 76.84
-10.3 | [5 empty] | 57.55 | 76.5 | 77.0857
-10.2 |  |  |  |  | 1.8667 | 64 | 76.7 | 77.35
-10.1 |  |  |  |  | 1.9 | 67.975 | 77 | 77.6286
-10 |  |  |  |  | 1.9 | 70.325 | 77.25 | 77.88
-9.9 |  |  |  |  | 2 | 71.7 | 77.6 | 78.16
-9.8 |  |  |  |  | 2 | 72.78 | 77.8 | 78.4333
-9.7 |  |  |  |  | 2 | 73.62 | 78.1714 | 78.65
-9.6 |  |  |  |  | 2.1 | 74.34 | 78.4667 | 78.9
-9.5 |  |  |  |  | 2.1 | 74.92 | 78.65 | 79.175
-9.4 |  |  |  |  | 2.2 | 75.525 | 79.025 | 79.425
-9.3 |  |  |  |  | 2.2 | 76.04 | 79.275 | 79.7625
-9.2 |  |  |  |  | 2.3 | 76.4667 | 79.6 | 79.95
-9.1 |  |  |  | 1.5 | 2.3 | 76.9333 | 79.8 | 80.2333
-9 |  |  |  | 1.6 | 2.4 | 77.36 | 80.1167 | 80.5
-8.9 |  |  |  | 1.6 | 2.5 | 77.775 | 80.3286 | 80.75
-8.8 |  |  |  | 1.6 | 2.5 | 78.2 | 80.7 | 81.0333
-8.7 |  |  |  | 1.6667 | 2.6 | 78.6333 | 81 | 81.3429
-8.6 |  |  |  | 1.7 | 2.7 | 79.05 | 81.3 | 81.5857
-8.5 |  |  |  | 1.7 | 2.7 | 79.3667 | 81.58 | 81.8625
-8.4 |  |  |  | 1.8 | 2.8 | 79.8 | 81.88 | 82.2286
-8.3 |  |  |  | 1.8 | 2.875 | 80.1 | 82.14 | 82.4429
-8.2 |  |  |  | 1.9 | 2.9333 | 80.48 | 82.45 | 82.74
-8.1 |  |  | 1.1 | 1.9 | 3.0333 | 80.9 | 82.6714 | 82.975
-8 |  |  | 1.1 | 1.9 | 3.1 | 81.3 | 83 | 83.275
-7.9 |  |  | 1.2 | 2 | 3.2 | 81.6857 | 83.3 | 83.525
-7.8 |  |  | 1.2 | 2.075 | 3.3 | 82.04 | 83.5714 | 83.775
-7.7 |  |  | 1.2 | 2.1 | 3.4 | 82.3 | 83.9571 | 84.0429
-7.6 |  |  | 1.3 | 2.1667 | 3.525 | 82.72 | 84.3 | 84.325
-7.5 |  |  | 1.3 | 2.2 | 3.7 | 83.1 | 84.5 | 84.6
-7.4 |  |  | 1.3 | 2.3 | 3.8 | 83.375 | 84.76 | 84.85
-7.3 |  |  | 1.3667 | 2.4 | 3.9667 | 83.7 | 85.1 | 85.1143
-7.2 |  |  | 1.4 | 2.4 | 4.1 | 84.05 | 85.28 | 85.4625
-7.1 |  |  | 1.4 | 2.5 | 4.2667 | 84.45 | 85.65 | 85.7333
-7 |  | 0.7 | 1.5 | 2.6 | 4.5 | 84.8 | 85.8857 | 86.0333
-6.9 |  | 0.8 | 1.5 | 2.7 | 4.6 | 85.05 | 86.2 | 86.3
-6.8 |  | 0.8 | 1.5 | 2.8 | 4.8 | 85.4 | 86.52 | 86.6143
-6.7 |  | 0.8 | 1.6 | 2.9 | 5 | 85.74 | 86.8 | 86.8
-6.6 |  | 0.8 | 1.6 | 3 | 5.25 | 86.1 | 86.98 | 87.15
-6.5 | 0.5 | 0.8 | 1.7 | 3.1 | 5.5 | 86.425 | 87.3667 | 87.38
-6.4 | 0.6 | 0.9 | 1.7 | 3.2 | 5.8 | 86.65 | 87.675 | 87.6667
-6.3 | 0.6 | 0.9 | 1.8 | 3.35 | 6.2 | 87 | 87.8667 | 87.9286
-6.2 | 0.6 | 0.9 | 1.8 | 3.5 | 6.55 | 87.375 | 88.2143 | 88.2333
-6.1 | 0.6 | 0.9 | 1.9 | 3.6 | 7.05 | 87.7 | 88.52 | 88.45
-6 | 0.6 | 0.95 | 2 | 3.8 | 7.7 | 88.0667 | 88.8 | 88.8167
-5.9 | 0.6 | 1 | 2.05 | 3.9 | 8.6 | 88.325 | 88.9857 | 89.0429
-5.8 | 0.6 | 1 | 2.1 | 4.1 | 9.9 | 88.6 | 89.32 | 89.3571
-5.7 | 0.7 | 1 | 2.2 | 4.3 | 12.25 | 88.95 | 89.6571 | 89.6429
-5.6 | 0.7 | 1.05 | 2.3 | 4.5 | 16 | 89.2333 | 89.8833 | 89.8571
-5.5 | 0.7 | 1.1 | 2.4 | 4.8 | 22.7 | 89.525 | 90.125 | 90.1429
-5.4 | 0.7 | 1.1 | 2.5 | 5 | 31.6 | 89.8143 | 90.4333 | 90.44
-5.3 | 0.7 | 1.2 | 2.6 | 5.3 | 41.8 | 90.16 | 90.64 | 90.7
-5.2 | 0.7 | 1.2 | 2.7 | 5.7 | 51.7 | 90.4667 | 91.02 | 90.9714
-5.1 | 0.8 | 1.3 | 2.8 | 6.1333 | 62.4 | 90.7 | 91.2667 | 91.2556
-5 | 0.8 | 1.3 | 3 | 6.6 | 71.625 | 91.05 | 91.55 | 91.4
-4.9 | 0.8 | 1.35 | 3.1 | 7.4333 | 78.2 | 91.325 | 91.76 | 91.75
-4.8 | 0.85 | 1.4 | 3.3 | 8.5 | 82.7 | 91.6 | 92.0857 | 92
-4.7 | 0.9 | 1.5 | 3.45 | 10.15 | 85.2 | 91.9 | 92.38 | 92.3
-4.6 | 0.9 | 1.5 | 3.6333 | 12.9667 | 86.775 | 92.1 | 92.64 | 92.5333
-4.5 | 0.9 | 1.6 | 3.8 | 18.1 | 87.825 | 92.475 | 92.9333 | 92.825
-4.4 | 1 | 1.7 | 4.1 | 25.55 | 88.725 | 92.7667 | 93.12 | 93.075
-4.3 | 1 | 1.7 | 4.3 | 35.1 | 89.475 | 93 | 93.4167 | 93.3143
-4.2 | 1.1 | 1.8 | 4.6 | 46.05 | 90.12 | 93.3 | 93.7 | 93.6
-4.1 | 1.1 | 1.9 | 5 | 57.3 | 90.7 | 93.6 | 94 | 93.8
-4 | 1.1333 | 2 | 5.45 | 67.55 | 91.2 | 93.9 | 94.1 | 94.1
-3.9 | 1.2 | 2.1 | 6.075 | 76.6 | 91.7 | 94.0667 | 94.4 | 94.2857
-3.8 | 1.3 | 2.25 | 7 | 83.15 | 92.1333 | 94.35 | 94.68 | 94.54
-3.7 | 1.3 | 2.4 | 8.2 | 86.6 | 92.5667 | 94.6286 | 94.9 | 94.8
-3.6 | 1.4 | 2.5 | 10.3333 | 88.8 | 93.05 | 94.9 | 95.15 | 95.05
-3.5 | 1.5 | 2.7 | 13.7333 | 90.3 | 93.4 | 95.175 | 95.46 | 95.3
-3.4 | 1.6 | 2.8 | 20.2 | 91.2 | 93.8 | 95.3429 | 95.65 | 95.5333
-3.3 | 1.65 | 3.1 | 29.4667 | 92.02 | 94.2 | 95.6 | 95.9 | 95.7571
-3.2 | 1.8 | 3.3 | 40.55 | 92.725 | 94.5 | 95.8 | 96.04 | 95.9625
-3.1 | 1.9 | 3.6 | 51.8 | 93.35 | 94.9 | 96.1 | 96.2571 | 96.2143
-3 | 2 | 4 | 63.2 | 93.85 | 95.2 | 96.3 | 96.5429 | 96.3
-2.9 | 2.15 | 4.5333 | 74 | 94.4 | 95.55 | 96.55 | 96.78 | 96.5125
-2.8 | 2.3 | 5.2 | 82.5 | 94.775 | 95.8 | 96.74 | 96.94 | 96.675
-2.7 | 2.5 | 6.15 | 87.8 | 95.15 | 96.15 | 97 | 97.2 | 96.8556
-2.6 | 2.8 | 7.7 | 90.7 | 95.6 | 96.45 | 97.1667 | 97.4 | 97.05
-2.5 | 3.125 | 10.5 | 92.6 | 95.95 | 96.7 | 97.3667 | 97.5286 | 97.4
-2.4 | 3.5 | 15.8 | 93.7333 | 96.3 | 96.9 | 97.6 | 97.7 | 97.5667
-2.3 | 4.1333 | 23.8 | 94.7 | 96.6 | 97.24 | 97.8 | 97.8625 | 97.8
-2.2 | 5.05 | 34.1 | 95.3 | 96.85 | 97.4 | 98 | 98.14 | 97.95
-2.1 | 6.4333 | 46.4 | 95.9333 | 97.225 | 97.7 | 98.1 | 98.3 | 98.0875
-2 | 8.9 | 59.5 | 96.45 | 97.5 | 97.9 | 98.3167 | 98.44 | 98.3125
-1.9 | 13.3 | 70.5 | 96.7333 | 97.7 | 98.1333 | 98.4714 | 98.6375 | 98.4
-1.8 | 21 | 81.1 | 97.2 | 97.875 | 98.3 | 98.65 | 98.7833 | 98.5
-1.7 | 30.85 | 88.3 | 97.55 | 98.1667 | 98.5 | 98.8143 | 98.8 | 98.6714
-1.6 | 43.4333 | 92.55 | 97.8 | 98.4333 | 98.7 | 98.98 | 98.9 | 98.8
-1.5 | 55.9 | 94.7667 | 98.1667 | 98.6 | 98.8 | 99.1 | 99.12 | 98.8714
-1.4 | 68.95 | 96.2 | 98.4 | 98.8 | 99 | 99.2 | 99.2 | 99.0286
-1.3 | 79.9 | 97.05 | 98.7 | 99 | 99.1 | 99.3333 | 99.35 | 99.1714
-1.2 | 87.5 | 97.7333 | 98.9 | 99.1 | 99.2667 | 99.4 | 99.45 | 99.325
-1.1 | 92.9 | 98.2 | 99.05 | 99.2 | 99.3667 | 99.5 | 99.525 | 99.4
-1 | 95.4 | 98.625 | 99.3 | 99.4 | 99.5 | 99.6 | 99.7 | 99.4857
-0.9 | 97.1 | 98.975 | 99.4333 | 99.4 | 99.6 | 99.6667 | 99.7 | 99.6
-0.8 | 97.9667 | 99.225 | 99.6 | 99.6 | 99.7 | 99.7375 | 99.8 | 99.6143
-0.7 | 98.7 | 99.5 | 99.7 | 99.7 | 99.7 | 99.8 | 99.9 | 99.7
-0.6 | 99.1 | 99.6 | 99.8 | 99.8 | 99.8 | 99.9 | 99.9 | 99.7
-0.5 | 99.4 | 99.7 | 99.9 | 99.85 | 99.9 | 99.9 | 99.9 | 99.85
-0.4 | 99.6 | 99.8 | 99.9 | 99.9 | 99.95 | 100 | 100 | 99.9
-0.3 | 99.775 | 99.9 | 100 | 99.9 | 100 | 100 | 100 | 100
-0.2 | 99.9 | 99.9667 | 100 | 100 | 100 | 100 | 100 | 99.9667
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.85 | 99.9 | 100 | 99.95 | 100 | 100 | 100 | 99.9625
0.4 | 99.7 | 99.8 | 99.9 | 99.9 | 99.9 | 99.975 | 99.95 | 99.9
0.5 | 99.4 | 99.7 | 99.8333 | 99.82 | 99.9 | 99.9 | 99.9 | 99.8
0.6 | 99.1 | 99.5667 | 99.7 | 99.7 | 99.8 | 99.88 | 99.9 | 99.7
0.7 | 98.6 | 99.4 | 99.6 | 99.6 | 99.7 | 99.78 | 99.8 | 99.7
0.8 | 97.9 | 99.15 | 99.5 | 99.5667 | 99.7 | 99.7 | 99.72 | 99.6625
0.9 | 97 | 98.85 | 99.4 | 99.45 | 99.6 | 99.6 | 99.7 | 99.6
1 | 95.3 | 98.6 | 99.2 | 99.3 | 99.45 | 99.6 | 99.6 | 99.4333
1.1 | 92.7 | 98.1 | 99.0333 | 99.2 | 99.3 | 99.46 | 99.4833 | 99.4
1.2 | 87.55 | 97.65 | 98.9 | 99 | 99.2 | 99.4 | 99.4 | 99.3
1.3 | 80 | 97 | 98.7 | 98.9 | 99.1 | 99.26 | 99.26 | 99.1571
1.4 | 68.7 | 95.925 | 98.4 | 98.7 | 98.9 | 99.1 | 99.06 | 99.0833
1.5 | 55.9 | 94.5 | 98.15 | 98.5 | 98.8 | 98.9667 | 99 | 98.9286
1.6 | 43.3 | 92.35 | 97.8 | 98.35 | 98.6 | 98.875 | 98.9 | 98.8
1.7 | 31.3 | 88.3 | 97.5 | 98.1 | 98.4 | 98.68 | 98.76 | 98.6286
1.8 | 21.2 | 81.3 | 97.0667 | 97.875 | 98.2 | 98.56 | 98.5333 | 98.5143
1.9 | 13 | 70.7 | 96.7 | 97.6 | 98.05 | 98.3 | 98.4 | 98.4
2 | 8.8 | 59.5 | 96.4 | 97.35 | 97.8 | 98.1833 | 98.28 | 98.2625
2.1 | 6.4333 | 46.6 | 95.7667 | 97.0333 | 97.6 | 98.0286 | 98.1143 | 98.1375
2.2 | 5 | 34.5 | 95.2 | 96.75 | 97.3 | 97.8571 | 97.94 | 97.925
2.3 | 4.1333 | 24.0333 | 94.5667 | 96.4 | 97.1 | 97.66 | 97.74 | 97.8
2.4 | 3.55 | 15.8 | 93.6 | 96.0333 | 96.9 | 97.4 | 97.56 | 97.6143
2.5 | 3.1333 | 10.45 | 92.4667 | 95.75 | 96.6333 | 97.2667 | 97.4 | 97.4143
2.6 | 2.8 | 7.7 | 90.5 | 95.4 | 96.3 | 97.025 | 97.2143 | 97.075
2.7 | 2.5 | 6.15 | 87.6 | 95 | 96.025 | 96.8333 | 96.9667 | 96.9111
2.8 | 2.3 | 5.2 | 82.5 | 94.62 | 95.7 | 96.6 | 96.7333 | 96.675
2.9 | 2.2 | 4.4667 | 73.9 | 94.1 | 95.4333 | 96.35 | 96.5 | 96.5
3 | 2 | 4 | 63.2 | 93.65 | 95.1 | 96.2 | 96.3333 | 96.25
3.1 | 1.9 | 3.6 | 51.8 | 93.16 | 94.75 | 95.9 | 96.1 | 96.1714
3.2 | 1.8 | 3.3 | 40.6 | 92.6 | 94.4 | 95.65 | 95.95 | 95.9143
3.3 | 1.6 | 3.1 | 29.5333 | 91.8 | 94 | 95.42 | 95.625 | 95.7833
3.4 | 1.6 | 2.8 | 20.5 | 91 | 93.6 | 95.18 | 95.44 | 95.5333
3.5 | 1.5 | 2.7 | 13.8 | 90.1 | 93.3 | 94.94 | 95.1833 | 95.2
3.6 | 1.4 | 2.55 | 10.1333 | 88.7 | 92.8 | 94.675 | 95 | 95
3.7 | 1.3 | 2.4 | 8.1 | 86.5 | 92.35 | 94.44 | 94.75 | 94.8143
3.8 | 1.3 | 2.25 | 7 | 82.88 | 92 | 94.16 | 94.4143 | 94.5714
3.9 | 1.2 | 2.1 | 6.1 | 76.7 | 91.5 | 93.925 | 94.1 | 94.2667
4 | 1.1 | 2 | 5.5 | 67.66 | 91 | 93.62 | 94 | 94.1
4.1 | 1.1 | 1.9 | 5 | 57.4 | 90.475 | 93.34 | 93.775 | 93.8571
4.2 | 1.1 | 1.8 | 4.6333 | 45.95 | 89.9 | 93.125 | 93.4333 | 93.6286
4.3 | 1 | 1.7 | 4.3 | 35.2 | 89.3 | 92.8 | 93.16 | 93.3429
4.4 | 1 | 1.7 | 4.0667 | 25.9 | 88.45 | 92.5 | 92.8667 | 93.075
4.5 | 0.9 | 1.6 | 3.8 | 18.1 | 87.7 | 92.2333 | 92.6 | 92.825
4.6 | 0.9 | 1.5 | 3.6 | 12.8667 | 86.45 | 91.95 | 92.3714 | 92.5667
4.7 | 0.9 | 1.5 | 3.425 | 10.05 | 84.9 | 91.6333 | 92.1571 | 92.35
4.8 | 0.85 | 1.4 | 3.3 | 8.4667 | 82.5 | 91.4 | 91.9 | 92
4.9 | 0.8 | 1.3 | 3.1 | 7.4 | 78.35 | 91.1 | 91.6 | 91.75
5 | 0.8 | 1.3 | 3 | 6.6 | 71.2667 | 90.8 | 91.32 | 91.45
5.1 | 0.8 | 1.3 | 2.8 | 6.1 | 62.25 | 90.48 | 91.06 | 91.2333
5.2 | 0.8 | 1.2 | 2.7 | 5.7 | 51.7 | 90.2667 | 90.75 | 91
5.3 | 0.7 | 1.2 | 2.6 | 5.3 | 41.3333 | 89.85 | 90.46 | 90.65
5.4 | 0.7 | 1.1 | 2.5 | 5 | 31.8 | 89.58 | 90.1667 | 90.4
5.5 | 0.7 | 1.1 | 2.4 | 4.7 | 22.45 | 89.28 | 89.925 | 90.1333
5.6 | 0.7 | 1.0667 | 2.3 | 4.5 | 15.9 | 89 | 89.62 | 89.8429
5.7 | 0.7 | 1 | 2.2 | 4.3 | 12.1 | 88.68 | 89.32 | 89.6167
5.8 | 0.6 | 1 | 2.1 | 4.1333 | 9.9 | 88.4 | 89.1 | 89.3667
5.9 | 0.6 | 1 | 2.05 | 3.9 | 8.6 | 88.125 | 88.8 | 89.0333
6 | 0.6 | 0.9 | 2 | 3.8 | 7.75 | 87.8 | 88.5571 | 88.7667
6.1 | 0.6 | 0.9 | 1.9 | 3.6 | 7.05 | 87.4 | 88.2667 | 88.5
6.2 | 0.6 | 0.9 | 1.9 | 3.46 | 6.6 | 87.125 | 87.9 | 88.2333
6.3 | 0.6 | 0.9 | 1.8 | 3.3 | 6.2 | 86.85 | 87.6286 | 87.95
6.4 | 0.6 | 0.9 | 1.7 | 3.2 | 5.8 | 86.475 | 87.325 | 87.65
6.5 | 0.6 | 0.8 | 1.7 | 3.1 | 5.5 | 86.125 | 87.0625 | 87.34
6.6 |  | 0.8 | 1.6 | 3 | 5.25 | 85.85 | 86.8571 | 87.125
6.7 |  | 0.8 | 1.6 | 2.9 | 5 | 85.48 | 86.56 | 86.8429
6.8 |  | 0.8 | 1.5 | 2.8 | 4.8 | 85.2 | 86.2714 | 86.625
6.9 |  | 0.8 | 1.5 | 2.7 | 4.6 | 84.8 | 85.95 | 86.35
7 |  | 0.7 | 1.5 | 2.6 | 4.45 | 84.525 | 85.6667 | 86.0333
7.1 |  |  | 1.4 | 2.5 | 4.3 | 84.2 | 85.4 | 85.7667
7.2 |  |  | 1.4 | 2.4 | 4.1 | 83.75 | 85.0429 | 85.4125
7.3 |  |  | 1.3 | 2.4 | 3.9667 | 83.475 | 84.7333 | 85.125
7.4 |  |  | 1.3 | 2.3 | 3.8 | 83.2 | 84.52 | 84.85
7.5 |  |  | 1.3 | 2.2 | 3.7 | 82.8333 | 84.22 | 84.6
7.6 |  |  | 1.3 | 2.175 | 3.525 | 82.46 | 83.9 | 84.35
7.7 |  |  | 1.2 | 2.1 | 3.5 | 82.14 | 83.5875 | 84.1
7.8 |  |  | 1.2 | 2.0667 | 3.3 | 81.78 | 83.3 | 83.7857
7.9 |  |  | 1.2 | 2 | 3.2333 | 81.52 | 83.0333 | 83.525
8 |  |  | 1.15 | 1.9 | 3.1 | 81.08 | 82.74 | 83.2714
8.1 |  |  | 1.1 | 1.9 | 3 | 80.6 | 82.44 | 82.9857
8.2 |  |  |  | 1.8 | 3 | 80.32 | 82.14 | 82.8
8.3 |  |  |  | 1.8 | 2.9 | 79.92 | 81.88 | 82.4286
8.4 |  |  |  | 1.75 | 2.8 | 79.6 | 81.6 | 82.16
8.5 |  |  |  | 1.7 | 2.7 | 79.2 | 81.3 | 81.9143
8.6 |  |  |  | 1.7 | 2.6 | 78.7667 | 81.02 | 81.5667
8.7 |  |  |  | 1.6667 | 2.6 | 78.4333 | 80.64 | 81.3167
8.8 |  |  |  | 1.6 | 2.5 | 78 | 80.4286 | 81.1
8.9 |  |  |  | 1.6 | 2.5 | 77.55 | 80.175 | 80.75
9 |  |  |  | 1.5 | 2.4 | 77.1 | 79.8667 | 80.5
9.1 |  |  |  | 1.5 | 2.3 | 76.65 | 79.54 | 80.2286
9.2 |  |  |  |  | 2.3 | 76.25 | 79.25 | 79.9429
9.3 |  |  |  |  | 2.2 | 75.75 | 79.04 | 79.675
9.4 |  |  |  |  | 2.2 | 75.2143 | 78.675 | 79.425
9.5 |  |  |  |  | 2.1 | 74.74 | 78.4 | 79.175
9.6 |  |  |  |  | 2.1 | 74.025 | 78.16 | 78.9143
9.7 |  |  |  |  | 2 | 73.3333 | 77.96 | 78.65
9.8 |  |  |  |  | 2 | 72.5 | 77.625 | 78.4667
9.9 |  |  |  |  | 2 | 71.5 | 77.3 | 78.1
10 |  |  |  |  | 1.9 | 70.1 | 77 | 77.84
10.1 |  |  |  |  | 1.9 | 67.6333 | 76.8 | 77.6143
10.2 |  |  |  |  | 1.88 | 63.675 | 76.48 | 77.3
10.3 | [5 empty] | 57.7 | 76.2143 | 77.0857
10.4 | [5 empty] | 50.35 | 75.9 | 76.86
10.5 | [5 empty] | 41.775 | 75.62 | 76.5
10.6 | [5 empty] | 33.4 | 75.2833 | 76.2857
10.7 | [5 empty] | 26 | 75.1 | 76.1125
10.8 | [5 empty] | 19.4143 | 74.8 | 75.775
10.9 | [5 empty] | 14.75 | 74.5 | 75.5571
11 | [5 empty] | 11.82 | 74.1429 | 75.3286
11.1 | [5 empty] | 10.3 | 73.9 | 75.08
11.2 | [5 empty] | 9.375 | 73.5833 | 74.85
11.3 | [5 empty] | 8.6 | 73.3333 | 74.5
11.4 | [5 empty] | 8.1 | 73.06 | 74.2714
11.5 | [5 empty] | 7.725 | 72.7714 | 74
11.6 | [5 empty] | 7.3667 | 72.5 | 73.7667
11.7 | [5 empty] | 7.04 | 72.2 | 73.55
11.8 | [5 empty] | 6.825 | 71.925 | 73.2667
11.9 | [5 empty] | 6.575 | 71.6333 | 73.0286
12 | [5 empty] | 6.4 | 71.3333 | 72.8
12.1 | [5 empty] | 6.2 | 71.08 | 72.5667
12.2 | [5 empty] | 6 | 70.75 | 72.275
12.3 | [5 empty] | 5.85 | 70.525 | 72.025
12.4 | [5 empty] | 5.675 | 70.2429 | 71.775
12.5 | [5 empty] | 5.45 | 69.9 | 71.525
12.6 | [5 empty] | 5.3 | 69.675 | 71.275
12.7 | [5 empty] | 5.2 | 69.3667 | 70.9857
12.8 | [5 empty] | 5.1 | 69.0667 | 70.7667
12.9 | [5 empty] | 4.9 | 68.8 | 70.5714
13 | [5 empty] | 4.8 | 68.4875 | 70.3429
13.1 | [5 empty] | 4.7 | 68.1333 | 70.0714
13.2 | [5 empty] | 4.6 | 67.8714 | 69.82
13.3 | [5 empty] | 4.4429 | 67.62 | 69.5857
13.4 | [5 empty] | 4.325 | 67.3 | 69.38
13.5 | [5 empty] | 4.26 | 66.96 | 69.15
13.6 | [5 empty] | 4.2 | 66.6667 | 68.8833
13.7 | [5 empty] | 4.0857 | 66.4 | 68.6571
13.8 | [5 empty] | 4 | 66.0333 | 68.3333
13.9 | [5 empty] | 3.9 | 65.74 | 68.1833
14 | [5 empty] | 3.8333 | 65.3857 | 67.9571
14.1 | [5 empty] | 3.775 | 65.0667 | 67.675
14.2 | [5 empty] | 3.7 | 64.7 | 67.425
14.3 | [5 empty] | 3.6 | 64.4 | 67.175
14.4 | [5 empty] | 3.5667 | 64 | 66.9833
14.5 | [5 empty] | 3.5 | 63.6143 | 66.7571
14.6 | [5 empty] | 3.4 | 63.2 | 66.5
14.7 | [5 empty] | 3.38 | 62.75 | 66.2556
14.8 | [5 empty] | 3.3 | 62.34 | 66.0625
14.9 | [5 empty] | 3.2 | 61.775 | 65.875
15 | [5 empty] | 3.2 | 61.16 | 65.625
15.1 | [5 empty] | 3.1 | 60.45 | 65.375
15.2 | [5 empty] | 3.1 | 59.5 | 65.125
15.3 | [5 empty] | 3 | 57.9 | 64.8889
15.4 | [5 empty] | 2.96 | 55.45 | 64.7
15.5 | [6 empty] | 51.7 | 64.4429
15.6 | [6 empty] | 46.3 | 64.2167
15.7 | [6 empty] | 39.9571 | 63.9714
15.8 | [6 empty] | 33.4 | 63.775
15.9 | [6 empty] | 26.6 | 63.5
16 | [6 empty] | 20.8571 | 63.2571
16.1 | [6 empty] | 15.825 | 63
16.2 | [6 empty] | 12.6 | 62.85
16.3 | [6 empty] | 10.75 | 62.6111
16.4 | [6 empty] | 9.675 | 62.375
16.5 | [6 empty] | 8.95 | 62.125
16.6 | [6 empty] | 8.4 | 61.9667
16.7 | [6 empty] | 8.0286 | 61.7
16.8 | [6 empty] | 7.725 | 61.4625
16.9 | [6 empty] | 7.4 | 61.2
17 | [6 empty] | 7.15 | 61.0333
17.1 | [6 empty] | 6.9714 | 60.8143
17.2 | [6 empty] | 6.72 | 60.575
17.3 | [6 empty] | 6.6 | 60.3667
17.4 | [6 empty] | 6.44 | 60.0429
17.5 | [6 empty] | 6.225 | 59.8167
17.6 | [6 empty] | 6.1 | 59.5714
17.7 | [6 empty] | 5.94 | 59.4167
17.8 | [6 empty] | 5.8 | 59.1
17.9 | [6 empty] | 5.625 | 58.825
18 | [6 empty] | 5.5 | 58.625
18.1 | [6 empty] | 5.3667 | 58.4375
18.2 | [6 empty] | 5.2333 | 58.225
18.3 | [6 empty] | 5.125 | 57.95
18.4 | [6 empty] | 5.1 | 57.7
18.5 | [6 empty] | 4.9 | 57.45
18.6 | [6 empty] | 4.8333 | 57.2
18.7 | [6 empty] | 4.8 | 56.98
18.8 | [6 empty] | 4.6833 | 56.7143
18.9 | [6 empty] | 4.6 | 56.4286
19 | [6 empty] | 4.5 | 56.1667
19.1 | [6 empty] | 4.4 | 55.95
19.2 | [6 empty] | 4.3 | 55.675
19.3 | [6 empty] | 4.2 | 55.4143
19.4 | [6 empty] | 4.2 | 55.1286
19.5 | [6 empty] | 4.1 | 54.86
19.6 | [6 empty] | 4 | 54.6
19.7 | [6 empty] | 3.9286 | 54.3333
19.8 | [6 empty] | 3.9 | 53.9
19.9 | [6 empty] | 3.8 | 53.6
20 | [6 empty] | 3.75 | 53.2333
20.1 | [6 empty] | 3.7 | 52.8
20.2 | [6 empty] | 3.6 | 52.3
20.3 | [6 empty] | 3.6 | 51.7571
20.4 | [6 empty] | 3.5 | 51.1
20.5 | [6 empty] | 3.5 | 49.9875
20.6 | [6 empty] | 3.4 | 48.2167
20.7 | [6 empty] | 3.4 | 45.32
20.8 | [7 empty] | 40.6375
20.9 | [7 empty] | 35.5571
21 | [7 empty] | 30.3143
21.1 | [7 empty] | 24.75
21.2 | [7 empty] | 19.2333
21.3 | [7 empty] | 14.9833
21.4 | [7 empty] | 12
21.5 | [7 empty] | 10.3571
21.6 | [7 empty] | 9.3333
21.7 | [7 empty] | 8.7143
21.8 | [7 empty] | 8.2333
21.9 | [7 empty] | 7.8667
22 | [7 empty] | 7.5
22.1 | [7 empty] | 7.3
22.2 | [7 empty] | 7.0714
22.3 | [7 empty] | 6.85
22.4 | [7 empty] | 6.7
22.5 | [7 empty] | 6.5
22.6 | [7 empty] | 6.3571
22.7 | [7 empty] | 6.2
22.8 | [7 empty] | 6.0429
22.9 | [7 empty] | 5.9
23 | [7 empty] | 5.7625
23.1 | [7 empty] | 5.6167
23.2 | [7 empty] | 5.5
23.3 | [7 empty] | 5.4167
23.4 | [7 empty] | 5.275
23.5 | [7 empty] | 5.2
23.6 | [7 empty] | 5.1
23.7 | [7 empty] | 5.0143
23.8 | [7 empty] | 4.9
23.9 | [7 empty] | 4.8286
24 | [7 empty] | 4.7
24.1 | [7 empty] | 4.6571
24.2 | [7 empty] | 4.6
24.3 | [7 empty] | 4.4857
24.4 | [7 empty] | 4.4
24.5 | [7 empty] | 4.3714
24.6 | [7 empty] | 4.2375
24.7 | [7 empty] | 4.2
24.8 | [7 empty] | 4.1
24.9 | [7 empty] | 4.1
25 | [7 empty] | 4.025
25.1 | [7 empty] | 3.975
25.2 | [7 empty] | 3.9
25.3 | [7 empty] | 3.8875
25.4 | [7 empty] | 3.8
25.5 | [7 empty] | 3.7
25.6 | [7 empty] | 3.7
25.7 | [7 empty] | 3.6375
25.8 | [7 empty] | 3.6
25.9 | [7 empty] | 3.5

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-26.9 | [7 empty] | 5.1
-26.8 | [7 empty] | 5.2
-26.7 | [7 empty] | 5.2429
-26.6 | [7 empty] | 5.3857
-26.5 | [7 empty] | 5.4
-26.4 | [7 empty] | 5.5
-26.3 | [7 empty] | 5.6
-26.2 | [7 empty] | 5.68
-26.1 | [7 empty] | 5.75
-26 | [7 empty] | 5.8
-25.9 | [7 empty] | 5.9429
-25.8 | [7 empty] | 6
-25.7 | [7 empty] | 6.1
-25.6 | [7 empty] | 6.2286
-25.5 | [7 empty] | 6.3625
-25.4 | [7 empty] | 6.4
-25.3 | [7 empty] | 6.5333
-25.2 | [7 empty] | 6.7
-25.1 | [7 empty] | 6.7286
-25 | [7 empty] | 6.8714
-24.9 | [7 empty] | 7
-24.8 | [7 empty] | 7.1
-24.7 | [7 empty] | 7.24
-24.6 | [7 empty] | 7.4
-24.5 | [7 empty] | 7.4833
-24.4 | [7 empty] | 7.6429
-24.3 | [7 empty] | 7.7857
-24.2 | [7 empty] | 7.9167
-24.1 | [7 empty] | 8.0714
-24 | [7 empty] | 8.2571
-23.9 | [7 empty] | 8.4714
-23.8 | [7 empty] | 8.6
-23.7 | [7 empty] | 8.8333
-23.6 | [7 empty] | 9
-23.5 | [7 empty] | 9.175
-23.4 | [7 empty] | 9.38
-23.3 | [7 empty] | 9.62
-23.2 | [7 empty] | 9.8167
-23.1 | [7 empty] | 10.0667
-23 | [7 empty] | 10.3857
-22.9 | [7 empty] | 10.725
-22.8 | [7 empty] | 11.1
-22.7 | [7 empty] | 11.525
-22.6 | [7 empty] | 12.1857
-22.5 | [7 empty] | 13.1
-22.4 | [7 empty] | 14.7
-22.3 | [7 empty] | 17
-22.2 | [7 empty] | 20.7
-22.1 | [7 empty] | 24.9
-22 | [7 empty] | 29.8
-21.9 | [7 empty] | 34.6857
-21.8 | [7 empty] | 39.3
-21.7 | [7 empty] | 43.2125
-21.6 | [7 empty] | 46
-21.5 | [7 empty] | 47.8286
-21.4 | [6 empty] | 5.0714 | 48.975
-21.3 | [6 empty] | 5.1 | 49.8333
-21.2 | [6 empty] | 5.2 | 50.4429
-21.1 | [6 empty] | 5.3 | 50.8714
-21 | [6 empty] | 5.4 | 51.4
-20.9 | [6 empty] | 5.5 | 51.7667
-20.8 | [6 empty] | 5.58 | 52.1571
-20.7 | [6 empty] | 5.66 | 52.4429
-20.6 | [6 empty] | 5.76 | 52.8143
-20.5 | [6 empty] | 5.9 | 53.1625
-20.4 | [6 empty] | 6 | 53.4375
-20.3 | [6 empty] | 6.1 | 53.7857
-20.2 | [6 empty] | 6.1857 | 54.0714
-20.1 | [6 empty] | 6.3 | 54.425
-20 | [6 empty] | 6.4 | 54.6333
-19.9 | [6 empty] | 6.525 | 54.95
-19.8 | [6 empty] | 6.625 | 55.2143
-19.7 | [6 empty] | 6.8 | 55.5
-19.6 | [6 empty] | 6.9 | 55.7857
-19.5 | [6 empty] | 7.0333 | 56.05
-19.4 | [6 empty] | 7.14 | 56.3
-19.3 | [6 empty] | 7.32 | 56.65
-19.2 | [6 empty] | 7.5 | 56.9143
-19.1 | [6 empty] | 7.625 | 57.1333
-19 | [6 empty] | 7.8 | 57.3833
-18.9 | [6 empty] | 7.975 | 57.675
-18.8 | [6 empty] | 8.1 | 57.9111
-18.7 | [6 empty] | 8.2571 | 58.2
-18.6 | [6 empty] | 8.4833 | 58.4857
-18.5 | [6 empty] | 8.625 | 58.7
-18.4 | [6 empty] | 8.8 | 58.95
-18.3 | [6 empty] | 9.05 | 59.2
-18.2 | [6 empty] | 9.25 | 59.45
-18.1 | [6 empty] | 9.4833 | 59.7333
-18 | [6 empty] | 9.675 | 59.9714
-17.9 | [6 empty] | 9.95 | 60.1778
-17.8 | [6 empty] | 10.2286 | 60.4286
-17.7 | [6 empty] | 10.5143 | 60.68
-17.6 | [6 empty] | 10.75 | 60.9143
-17.5 | [6 empty] | 11.075 | 61.1333
-17.4 | [6 empty] | 11.5143 | 61.4
-17.3 | [6 empty] | 12.05 | 61.6333
-17.2 | [6 empty] | 12.5857 | 61.9571
-17.1 | [6 empty] | 13.5571 | 62.16
-17 | [6 empty] | 14.9 | 62.4
-16.9 | [6 empty] | 17.1 | 62.6333
-16.8 | [6 empty] | 20.6143 | 62.8714
-16.7 | [6 empty] | 25.55 | 63.2
-16.6 | [6 empty] | 31 | 63.3333
-16.5 | [6 empty] | 36.875 | 63.625
-16.4 | [6 empty] | 42.7 | 63.8
-16.3 | [6 empty] | 48 | 64.0667
-16.2 | [6 empty] | 52.2 | 64.3
-16.1 | [6 empty] | 55.1 | 64.475
-16 | [6 empty] | 56.925 | 64.7111
-15.9 | [5 empty] | 4.6 | 58 | 64.95
-15.8 | [5 empty] | 4.7143 | 58.8667 | 65.2
-15.7 | [5 empty] | 4.8 | 59.65 | 65.35
-15.6 | [5 empty] | 4.875 | 60.1714 | 65.5571
-15.5 | [5 empty] | 4.95 | 60.6571 | 65.775
-15.4 | [5 empty] | 5.1 | 61.175 | 65.9444
-15.3 | [5 empty] | 5.175 | 61.6 | 66.175
-15.2 | [5 empty] | 5.2 | 62.08 | 66.425
-15.1 | [5 empty] | 5.3 | 62.5 | 66.7333
-15 | [5 empty] | 5.45 | 62.9 | 66.9167
-14.9 | [5 empty] | 5.54 | 63.26 | 67.1429
-14.8 | [5 empty] | 5.7 | 63.7 | 67.4667
-14.7 | [5 empty] | 5.75 | 64 | 67.6857
-14.6 | [5 empty] | 5.9 | 64.3875 | 67.975
-14.5 | [5 empty] | 6.025 | 64.8 | 68.2667
-14.4 | [5 empty] | 6.1 | 65.125 | 68.5667
-14.3 | [5 empty] | 6.25 | 65.3857 | 68.775
-14.2 | [5 empty] | 6.4 | 65.7857 | 69.1
-14.1 | [5 empty] | 6.52 | 66.0833 | 69.325
-14 | [5 empty] | 6.7 | 66.475 | 69.5143
-13.9 | [5 empty] | 6.85 | 66.8143 | 69.775
-13.8 | [5 empty] | 6.98 | 67.1667 | 70.025
-13.7 | [5 empty] | 7.1 | 67.44 | 70.275
-13.6 | [5 empty] | 7.3333 | 67.7625 | 70.4833
-13.5 | [5 empty] | 7.5 | 68.1 | 70.7667
-13.4 | [5 empty] | 7.68 | 68.3571 | 70.9571
-13.3 | [5 empty] | 7.8 | 68.6429 | 71.2
-13.2 | [5 empty] | 8.06 | 69.025 | 71.48
-13.1 | [5 empty] | 8.25 | 69.4 | 71.775
-13 | [5 empty] | 8.4667 | 69.6857 | 71.9286
-12.9 | [5 empty] | 8.625 | 69.9714 | 72.2
-12.8 | [5 empty] | 8.875 | 70.28 | 72.3667
-12.7 | [5 empty] | 9.15 | 70.5286 | 72.5857
-12.6 | [5 empty] | 9.4 | 70.8333 | 72.925
-12.5 | [5 empty] | 9.5857 | 71.2 | 73.175
-12.4 | [5 empty] | 9.8714 | 71.5167 | 73.3625
-12.3 | [5 empty] | 10.1667 | 71.8 | 73.7
-12.2 | [5 empty] | 10.5143 | 72.0429 | 73.85
-12.1 | [5 empty] | 10.8 | 72.34 | 74.1
-12 | [5 empty] | 11.25 | 72.6375 | 74.375
-11.9 | [5 empty] | 11.7 | 73.02 | 74.65
-11.8 | [5 empty] | 12.28 | 73.32 | 74.85
-11.7 | [5 empty] | 13.025 | 73.575 | 75.2
-11.6 | [5 empty] | 14.06 | 73.85 | 75.35
-11.5 | [5 empty] | 15.7 | 74.2 | 75.74
-11.4 | [5 empty] | 18.2714 | 74.4 | 75.96
-11.3 | [5 empty] | 22.4 | 74.74 | 76.2
-11.2 | [5 empty] | 28.2 | 75.0714 | 76.425
-11.1 | [5 empty] | 34.76 | 75.3 | 76.7
-11 | [5 empty] | 42.56 | 75.5667 | 76.95
-10.9 | [5 empty] | 49.9 | 75.84 | 77.2143
-10.8 | [5 empty] | 56.2333 | 76.225 | 77.475
-10.7 | [5 empty] | 61.875 | 76.4333 | 77.7
-10.6 | [5 empty] | 65.675 | 76.76 | 77.9
-10.5 | [5 empty] | 67.92 | 77.075 | 78.1857
-10.4 |  |  |  |  | 3.1 | 69.4 | 77.34 | 78.4
-10.3 |  |  |  |  | 3.2 | 70.5 | 77.58 | 78.7
-10.2 |  |  |  |  | 3.2 | 71.4333 | 77.95 | 79
-10.1 |  |  |  |  | 3.3 | 72.075 | 78.16 | 79.1667
-10 |  |  |  |  | 3.35 | 72.8 | 78.48 | 79.4667
-9.9 |  |  |  |  | 3.4 | 73.3 | 78.75 | 79.7
-9.8 |  |  |  |  | 3.5 | 73.85 | 79.06 | 79.9
-9.7 |  |  |  |  | 3.6 | 74.375 | 79.34 | 80.1857
-9.6 |  |  |  |  | 3.7 | 74.88 | 79.62 | 80.4
-9.5 |  |  |  |  | 3.8 | 75.3667 | 79.84 | 80.75
-9.4 |  |  |  |  | 3.8 | 75.8714 | 80.1667 | 81
-9.3 |  |  |  | 2.5 | 3.9667 | 76.325 | 80.4375 | 81.25
-9.2 |  |  |  | 2.6 | 4 | 76.8 | 80.775 | 81.4778
-9.1 |  |  |  | 2.6667 | 4.1333 | 77.24 | 81.0333 | 81.725
-9 |  |  |  | 2.7 | 4.22 | 77.64 | 81.3571 | 81.9143
-8.9 |  |  |  | 2.8 | 4.4 | 78.0667 | 81.6 | 82.2625
-8.8 |  |  |  | 2.85 | 4.5 | 78.5 | 81.9 | 82.525
-8.7 |  |  |  | 2.9 | 4.6 | 78.9667 | 82.1714 | 82.775
-8.6 |  |  |  | 3 | 4.7 | 79.35 | 82.475 | 83.0111
-8.5 |  |  |  | 3.1 | 4.8 | 79.8 | 82.7 | 83.325
-8.4 |  |  |  | 3.1 | 5 | 80.0857 | 82.9857 | 83.6
-8.3 |  |  |  | 3.2 | 5.1 | 80.5333 | 83.3429 | 83.875
-8.2 |  |  | 1.9 | 3.3 | 5.3 | 80.9286 | 83.64 | 84.15
-8.1 |  |  | 2 | 3.4 | 5.475 | 81.2714 | 83.8375 | 84.3714
-8 |  |  | 2 | 3.5 | 5.6 | 81.58 | 84.1857 | 84.6375
-7.9 |  |  | 2.1 | 3.6 | 5.8 | 82 | 84.45 | 84.8875
-7.8 |  |  | 2.1 | 3.7 | 5.9333 | 82.46 | 84.775 | 85.2
-7.7 |  |  | 2.2 | 3.8 | 6.2 | 82.75 | 85.0333 | 85.4429
-7.6 |  |  | 2.2 | 3.9 | 6.3 | 83.12 | 85.3333 | 85.6
-7.5 |  |  | 2.3 | 4 | 6.575 | 83.4667 | 85.625 | 85.96
-7.4 |  |  | 2.4 | 4.1 | 6.8 | 83.85 | 85.96 | 86.2286
-7.3 |  |  | 2.4 | 4.2333 | 7.075 | 84.2333 | 86.14 | 86.55
-7.2 |  |  | 2.5 | 4.4 | 7.2 | 84.52 | 86.38 | 86.8222
-7.1 |  | 1.3 | 2.6 | 4.525 | 7.5 | 84.86 | 86.82 | 87.0667
-7 |  | 1.3 | 2.6 | 4.7 | 7.8 | 85.24 | 86.9857 | 87.3
-6.9 |  | 1.3 | 2.7 | 4.8 | 8.125 | 85.6 | 87.32 | 87.56
-6.8 |  | 1.4 | 2.8 | 5 | 8.4 | 85.925 | 87.5286 | 87.84
-6.7 |  | 1.4 | 2.9 | 5.15 | 8.7 | 86.26 | 87.8 | 88.15
-6.6 | 0.9 | 1.4 | 2.95 | 5.325 | 9.1333 | 86.58 | 88.1167 | 88.4
-6.5 | 0.9 | 1.5 | 3.1 | 5.5 | 9.5 | 86.875 | 88.3 | 88.5857
-6.4 | 0.9 | 1.5 | 3.14 | 5.7 | 10.0667 | 87.2333 | 88.6571 | 88.925
-6.3 | 1 | 1.5333 | 3.3 | 5.95 | 10.775 | 87.625 | 88.95 | 89.0875
-6.2 | 1 | 1.6 | 3.4 | 6.2 | 11.6 | 87.92 | 89.25 | 89.35
-6.1 | 1 | 1.6 | 3.48 | 6.4 | 12.7 | 88.2 | 89.44 | 89.7
-6 | 1 | 1.7 | 3.6 | 6.7 | 14.5 | 88.45 | 89.7333 | 89.85
-5.9 | 1.1 | 1.7 | 3.66 | 6.9 | 17.5 | 88.8 | 89.96 | 90.1714
-5.8 | 1.1 | 1.8 | 3.85 | 7.18 | 22.45 | 89.1 | 90.2571 | 90.3333
-5.7 | 1.1 | 1.8667 | 4 | 7.4667 | 29.1 | 89.38 | 90.475 | 90.6
-5.6 | 1.2 | 1.9 | 4.1 | 7.8 | 38.3 | 89.75 | 90.74 | 90.8833
-5.5 | 1.2 | 2 | 4.3 | 8.1667 | 47.95 | 90.075 | 91 | 91.1857
-5.4 | 1.2 | 2 | 4.4667 | 8.7 | 57 | 90.325 | 91.2875 | 91.4625
-5.3 | 1.3 | 2.1 | 4.6 | 9.1 | 65.8 | 90.6 | 91.6 | 91.7
-5.2 | 1.3 | 2.2 | 4.8 | 9.7 | 73.55 | 90.95 | 91.78 | 91.9
-5.1 | 1.3 | 2.2333 | 5 | 10.5333 | 78.7 | 91.225 | 92.08 | 92.1714
-5 | 1.4 | 2.3 | 5.2 | 11.8 | 82.1 | 91.5 | 92.3 | 92.4143
-4.9 | 1.4 | 2.4 | 5.4 | 13.4333 | 84.14 | 91.8 | 92.54 | 92.6143
-4.8 | 1.5 | 2.5 | 5.65 | 16.2 | 85.45 | 92.1 | 92.8333 | 92.925
-4.7 | 1.5 | 2.6 | 5.9 | 20.95 | 86.6 | 92.3 | 93.14 | 93.1375
-4.6 | 1.6 | 2.7 | 6.225 | 28.2333 | 87.3667 | 92.55 | 93.36 | 93.3429
-4.5 | 1.65 | 2.8 | 6.5 | 37 | 88.1 | 92.88 | 93.6 | 93.5
-4.4 | 1.7 | 2.9333 | 6.825 | 47.3 | 88.8 | 93.15 | 93.85 | 93.68
-4.3 | 1.75 | 3.1 | 7.25 | 57.7 | 89.4 | 93.5 | 94.05 | 93.92
-4.2 | 1.8 | 3.2 | 7.75 | 66.8 | 89.95 | 93.66 | 94.2429 | 94.2
-4.1 | 1.9333 | 3.3 | 8.3333 | 75.15 | 90.45 | 93.95 | 94.45 | 94.4333
-4 | 2 | 3.5 | 9.1 | 81.1 | 90.95 | 94.25 | 94.74 | 94.7
-3.9 | 2.1 | 3.7 | 10.3 | 85.05 | 91.4 | 94.55 | 94.96 | 94.9571
-3.8 | 2.2 | 3.8 | 12 | 87.25 | 91.8667 | 94.8 | 95.12 | 95.1625
-3.7 | 2.3 | 4 | 14.6 | 88.7 | 92.2667 | 94.98 | 95.4333 | 95.375
-3.6 | 2.4 | 4.2 | 19.4 | 89.8 | 92.8 | 95.2 | 95.62 | 95.65
-3.5 | 2.5 | 4.4 | 26.8333 | 90.6 | 93.2 | 95.5 | 95.8333 | 95.8167
-3.4 | 2.6 | 4.7 | 36.7667 | 91.4667 | 93.6 | 95.775 | 96.12 | 96.0111
-3.3 | 2.8 | 5.02 | 47.4 | 92 | 93.9 | 95.9333 | 96.2333 | 96.25
-3.2 | 2.925 | 5.35 | 57.7 | 92.675 | 94.2 | 96.25 | 96.5 | 96.4
-3.1 | 3.075 | 5.8 | 68.8 | 93.2 | 94.6667 | 96.4167 | 96.7 | 96.725
-3 | 3.3 | 6.35 | 77.4 | 93.6 | 95.0667 | 96.55 | 96.88 | 96.9
-2.9 | 3.4333 | 7.1 | 83.9667 | 94.1 | 95.3 | 96.85 | 97.0571 | 97.0857
-2.8 | 3.65 | 8.3 | 88.3 | 94.5333 | 95.6 | 97.02 | 97.2167 | 97.24
-2.7 | 3.9 | 9.9 | 90.6 | 94.9333 | 96 | 97.22 | 97.4 | 97.5167
-2.6 | 4.3 | 12.8 | 92.1 | 95.3 | 96.2667 | 97.4 | 97.64 | 97.6
-2.5 | 4.8 | 17.95 | 93.3 | 95.68 | 96.6 | 97.5833 | 97.78 | 97.7
-2.4 | 5.4 | 25.65 | 94 | 95.975 | 96.8 | 97.8 | 97.8286 | 97.9125
-2.3 | 6.3 | 36.2 | 94.7 | 96.425 | 97.1 | 97.96 | 98.0667 | 98.0375
-2.2 | 7.6 | 47 | 95.3333 | 96.675 | 97.4 | 98.1 | 98.25 | 98.1625
-2.1 | 9.7 | 59.9 | 95.8 | 97.05 | 97.6 | 98.3 | 98.4167 | 98.375
-2 | 13.7 | 70.6333 | 96.3 | 97.4 | 97.85 | 98.46 | 98.6 | 98.6143
-1.9 | 20 | 80.5 | 96.7667 | 97.6 | 98 | 98.625 | 98.675 | 98.7
-1.8 | 29.1 | 87.2 | 97 | 97.8333 | 98.2 | 98.8 | 98.84 | 98.8
-1.7 | 40.5333 | 91.55 | 97.4667 | 98.1 | 98.5 | 98.9667 | 99 | 98.925
-1.6 | 53.4333 | 94.1 | 97.8 | 98.325 | 98.7 | 99.075 | 99.1 | 99.05
-1.5 | 64.7 | 95.6 | 98.1 | 98.5667 | 98.7 | 99.2 | 99.1667 | 99.1
-1.4 | 76.6 | 96.6 | 98.2 | 98.7 | 98.9 | 99.25 | 99.2375 | 99.2714
-1.3 | 85.3 | 97.35 | 98.6 | 98.8667 | 99.1 | 99.3 | 99.3714 | 99.4
-1.2 | 91.3 | 97.8667 | 98.8333 | 99.1 | 99.2 | 99.5 | 99.425 | 99.45
-1.1 | 94.4 | 98.3 | 98.95 | 99.3 | 99.35 | 99.5667 | 99.5571 | 99.5
-1 | 96.2 | 98.65 | 99.1 | 99.35 | 99.5 | 99.6 | 99.6 | 99.6
-0.9 | 97.4 | 98.95 | 99.4 | 99.5 | 99.6 | 99.7 | 99.6875 | 99.7
-0.8 | 98.2 | 99.2 | 99.5667 | 99.5 | 99.7 | 99.8 | 99.7667 | 99.7429
-0.7 | 98.9 | 99.45 | 99.7 | 99.7 | 99.725 | 99.8 | 99.8 | 99.8
-0.6 | 99.3 | 99.6 | 99.8 | 99.8 | 99.8 | 99.9 | 99.9 | 99.8
-0.5 | 99.5 | 99.7 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9125
-0.4 | 99.7 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9 | 100 | 100
-0.3 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100
-0.2 | 100 | 99.9667 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 99.9 | 100 | 100 | 100 | 100
0.3 | 99.8 | 99.9 | 99.9 | 99.9 | 100 | 99.95 | 99.9 | 99.9429
0.4 | 99.7 | 99.8 | 99.8667 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9875
0.5 | 99.4 | 99.7 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9
0.6 | 99.2 | 99.5 | 99.7 | 99.7 | 99.8 | 99.8 | 99.8 | 99.8429
0.7 | 98.7667 | 99.3333 | 99.625 | 99.7 | 99.7 | 99.8 | 99.8 | 99.8
0.8 | 98.2 | 99.1 | 99.5 | 99.5 | 99.6 | 99.8 | 99.7 | 99.7143
0.9 | 97.4 | 98.9 | 99.3 | 99.4 | 99.5 | 99.7 | 99.7 | 99.6143
1 | 96.2 | 98.6 | 99.1 | 99.3667 | 99.4 | 99.6 | 99.6 | 99.6333
1.1 | 94.2 | 98.2 | 98.975 | 99.2 | 99.3 | 99.5667 | 99.5 | 99.575
1.2 | 91.2 | 97.775 | 98.725 | 99.05 | 99.1333 | 99.5 | 99.4 | 99.4333
1.3 | 85.1 | 97.25 | 98.5 | 98.8667 | 99 | 99.3167 | 99.4 | 99.375
1.4 | 76.7 | 96.5 | 98.1 | 98.7 | 98.9 | 99.14 | 99.1625 | 99.3
1.5 | 64.9 | 95.45 | 98 | 98.4 | 98.8 | 99.1 | 99.0167 | 99.18
1.6 | 52.8 | 94 | 97.7 | 98.225 | 98.6 | 98.9667 | 98.9286 | 99.1
1.7 | 40.7 | 91.3 | 97.4 | 97.98 | 98.375 | 98.8 | 98.84 | 98.9
1.8 | 29.6 | 87.1 | 96.95 | 97.75 | 98.1 | 98.7143 | 98.7333 | 98.7714
1.9 | 20.2 | 80.5 | 96.5667 | 97.5 | 97.875 | 98.525 | 98.55 | 98.7
2 | 13.52 | 70.5667 | 96.08 | 97.25 | 97.625 | 98.38 | 98.4 | 98.575
2.1 | 9.7 | 59.5667 | 95.7667 | 96.95 | 97.5 | 98.2 | 98.3 | 98.3625
2.2 | 7.6 | 47.4 | 95.2 | 96.56 | 97.2333 | 98.05 | 98.12 | 98.175
2.3 | 6.3 | 35.7 | 94.55 | 96.25 | 97 | 97.8571 | 98 | 98.0125
2.4 | 5.3667 | 25.9 | 93.9 | 95.95 | 96.75 | 97.7 | 97.9 | 97.8889
2.5 | 4.8 | 17.8 | 93 | 95.55 | 96.5 | 97.46 | 97.65 | 97.775
2.6 | 4.3 | 12.6 | 92 | 95.15 | 96.1667 | 97.26 | 97.4167 | 97.6429
2.7 | 3.9 | 9.9 | 90.5 | 94.85 | 95.825 | 97.12 | 97.26 | 97.5
2.8 | 3.7 | 8.2 | 87.8 | 94.3667 | 95.6 | 96.84 | 97.1 | 97.24
2.9 | 3.5 | 7.1 | 83.9 | 94 | 95.15 | 96.68 | 96.9429 | 97.0571
3 | 3.2 | 6.3 | 76.95 | 93.5 | 94.8 | 96.44 | 96.7 | 96.85
3.1 | 3.1 | 5.8 | 68.2667 | 93.1 | 94.4667 | 96.2667 | 96.6 | 96.7
3.2 | 2.9 | 5.3333 | 58 | 92.4667 | 94.1 | 96.1 | 96.2429 | 96.45
3.3 | 2.75 | 5 | 47.4 | 91.95 | 93.8 | 95.8333 | 96.0167 | 96.25
3.4 | 2.6 | 4.7 | 36.4 | 91.3 | 93.4 | 95.6333 | 95.8857 | 96.0667
3.5 | 2.5 | 4.4 | 26.9 | 90.45 | 93 | 95.3 | 95.7333 | 95.9
3.6 | 2.4 | 4.2 | 19.4 | 89.6 | 92.6 | 95.0333 | 95.5333 | 95.7
3.7 | 2.3 | 4 | 14.4 | 88.5 | 92.25 | 94.875 | 95.2667 | 95.425
3.8 | 2.2 | 3.8 | 11.9 | 87 | 91.7667 | 94.7 | 95.0286 | 95.175
3.9 | 2.1 | 3.7 | 10.2 | 84.85 | 91.2 | 94.4 | 94.8 | 95.0125
4 | 2 | 3.5 | 9.1 | 81.1 | 90.8 | 94.1 | 94.58 | 94.75
4.1 | 1.9 | 3.3 | 8.3667 | 74.8 | 90.3 | 93.7667 | 94.3 | 94.4
4.2 | 1.8 | 3.2 | 7.7 | 66.6333 | 89.7 | 93.5 | 94.075 | 94.26
4.3 | 1.8 | 3 | 7.2667 | 56.975 | 89.1667 | 93.3 | 93.85 | 94
4.4 | 1.7 | 2.9 | 6.85 | 47.2 | 88.6 | 92.95 | 93.6 | 93.7167
4.5 | 1.7 | 2.8 | 6.475 | 36.7 | 87.925 | 92.75 | 93.3286 | 93.5
4.6 | 1.6 | 2.7 | 6.2 | 28.1 | 87.1 | 92.44 | 93.14 | 93.3857
4.7 | 1.55 | 2.6 | 5.9 | 20.75 | 86.4 | 92.18 | 92.94 | 93.125
4.8 | 1.5 | 2.5 | 5.625 | 16.2 | 85.25 | 91.9 | 92.675 | 92.9714
4.9 | 1.4667 | 2.4 | 5.4 | 13.3 | 83.9 | 91.58 | 92.425 | 92.6667
5 | 1.4 | 2.3 | 5.2333 | 11.6 | 81.9 | 91.35 | 92.1333 | 92.375
5.1 | 1.3 | 2.22 | 5 | 10.5 | 78.5 | 91.02 | 91.9143 | 92.1625
5.2 | 1.3 | 2.2 | 4.8 | 9.75 | 73.4 | 90.65 | 91.625 | 91.9571
5.3 | 1.3 | 2.1 | 4.6 | 9.1 | 65.7 | 90.4333 | 91.3667 | 91.6667
5.4 | 1.2 | 2 | 4.45 | 8.6 | 57 | 90.1667 | 91.0667 | 91.4375
5.5 | 1.2 | 1.95 | 4.3 | 8.15 | 47.8 | 89.825 | 90.78 | 91.2
5.6 | 1.2 | 1.9 | 4.16 | 7.8 | 38 | 89.525 | 90.5571 | 90.85
5.7 | 1.1 | 1.8667 | 3.9667 | 7.5 | 29.1 | 89.18 | 90.3 | 90.58
5.8 | 1.1 | 1.8 | 3.8 | 7.1667 | 22.4 | 88.94 | 90.0333 | 90.3
5.9 | 1.1 | 1.7 | 3.7 | 6.9333 | 17.2 | 88.58 | 89.6857 | 90.1667
6 | 1 | 1.7 | 3.6 | 6.6 | 14.4667 | 88.25 | 89.52 | 89.8
6.1 | 1 | 1.6 | 3.4667 | 6.4 | 12.7 | 88 | 89.22 | 89.7
6.2 | 1 | 1.6 | 3.4 | 6.2 | 11.5333 | 87.6667 | 88.96 | 89.375
6.3 | 1 | 1.52 | 3.275 | 5.9333 | 10.7 | 87.325 | 88.7 | 89.125
6.4 | 0.9 | 1.5 | 3.125 | 5.7 | 10.1 | 86.96 | 88.325 | 88.875
6.5 | 0.9 | 1.4333 | 3 | 5.56 | 9.6 | 86.675 | 88.1167 | 88.625
6.6 | 0.9 | 1.4 | 2.925 | 5.3 | 9.1 | 86.35 | 87.875 | 88.35
6.7 |  | 1.4 | 2.9 | 5.175 | 8.7 | 86.05 | 87.55 | 88.1
6.8 |  | 1.3 | 2.8 | 5.025 | 8.4 | 85.675 | 87.3 | 87.8
6.9 |  | 1.3 | 2.7 | 4.9 | 8.1 | 85.3333 | 87.0667 | 87.6
7 |  | 1.3 | 2.6 | 4.7 | 7.7667 | 84.98 | 86.7 | 87.3429
7.1 |  | 1.3 | 2.6 | 4.5 | 7.5 | 84.6333 | 86.4571 | 87.0333
7.2 |  |  | 2.5 | 4.4 | 7.3 | 84.24 | 86.15 | 86.7714
7.3 |  |  | 2.4 | 4.3 | 7 | 83.88 | 85.92 | 86.5
7.4 |  |  | 2.3 | 4.18 | 6.8 | 83.6 | 85.58 | 86.2286
7.5 |  |  | 2.3 | 4.025 | 6.6 | 83.3 | 85.3 | 85.96
7.6 |  |  | 2.2 | 3.9 | 6.3333 | 82.9 | 85.12 | 85.6429
7.7 |  |  | 2.2 | 3.8 | 6.1667 | 82.5 | 84.75 | 85.4167
7.8 |  |  | 2.1 | 3.7 | 5.95 | 82.1 | 84.5167 | 85.1714
7.9 |  |  | 2.1 | 3.6 | 5.8 | 81.75 | 84.18 | 84.9125
8 |  |  | 2 | 3.5 | 5.6 | 81.42 | 83.9 | 84.625
8.1 |  |  | 2 | 3.4 | 5.45 | 81.0333 | 83.6714 | 84.3667
8.2 |  |  | 1.9 | 3.3 | 5.3 | 80.6429 | 83.3 | 84.15
8.3 |  |  |  | 3.2 | 5.1 | 80.25 | 83 | 83.8
8.4 |  |  |  | 3.1 | 5 | 79.8714 | 82.7333 | 83.575
8.5 |  |  |  | 3.1 | 4.85 | 79.5 | 82.4667 | 83.35
8.6 |  |  |  | 3 | 4.7 | 79.075 | 82.2429 | 82.975
8.7 |  |  |  | 2.9 | 4.6 | 78.7333 | 82 | 82.8111
8.8 |  |  |  | 2.9 | 4.5 | 78.2 | 81.6667 | 82.575
8.9 |  |  |  | 2.8 | 4.4 | 77.7857 | 81.3 | 82.2875
9 |  |  |  | 2.7 | 4.3 | 77.4 | 81.1 | 81.975
9.1 |  |  |  | 2.6667 | 4.15 | 77.02 | 80.8143 | 81.7
9.2 |  |  |  | 2.6 | 4.1 | 76.55 | 80.475 | 81.4667
9.3 |  |  |  | 2.575 | 3.95 | 76.1 | 80.18 | 81.25
9.4 |  |  |  |  | 3.82 | 75.65 | 79.85 | 80.95
9.5 |  |  |  |  | 3.8 | 75.2 | 79.6 | 80.8
9.6 |  |  |  |  | 3.7 | 74.7 | 79.3571 | 80.4429
9.7 |  |  |  |  | 3.6 | 74.14 | 79.075 | 80.2167
9.8 |  |  |  |  | 3.5 | 73.7 | 78.75 | 79.98
9.9 |  |  |  |  | 3.4 | 73.1 | 78.5 | 79.6833
10 |  |  |  |  | 3.35 | 72.5 | 78.175 | 79.4
10.1 |  |  |  |  | 3.3 | 71.8714 | 77.9 | 79.1857
10.2 |  |  |  |  | 3.2 | 71.125 | 77.6 | 78.9714
10.3 |  |  |  |  | 3.125 | 70.3 | 77.3286 | 78.6857
10.4 |  |  |  |  | 3.1 | 69.1 | 77.075 | 78.4833
10.5 | [5 empty] | 67.65 | 76.85 | 78.2333
10.6 | [5 empty] | 65.3333 | 76.5 | 77.92
10.7 | [5 empty] | 61.8 | 76.2 | 77.72
10.8 | [5 empty] | 56.15 | 75.9 | 77.4714
10.9 | [5 empty] | 49.9 | 75.68 | 77.2571
11 | [5 empty] | 42.3 | 75.3429 | 76.9
11.1 | [5 empty] | 35 | 75.06 | 76.675
11.2 | [5 empty] | 28.2286 | 74.76 | 76.45
11.3 | [5 empty] | 22.7 | 74.5 | 76.2
11.4 | [5 empty] | 18.26 | 74.1 | 75.9667
11.5 | [5 empty] | 15.6 | 73.9167 | 75.7
11.6 | [5 empty] | 14.04 | 73.66 | 75.4
11.7 | [5 empty] | 13 | 73.32 | 75.15
11.8 | [5 empty] | 12.28 | 72.9333 | 74.9
11.9 | [5 empty] | 11.7 | 72.74 | 74.65
12 | [5 empty] | 11.25 | 72.4714 | 74.4
12.1 | [5 empty] | 10.8333 | 72.14 | 74.1429
12.2 | [5 empty] | 10.46 | 71.78 | 73.8571
12.3 | [5 empty] | 10.15 | 71.55 | 73.6429
12.4 | [5 empty] | 9.84 | 71.2 | 73.375
12.5 | [5 empty] | 9.55 | 70.9 | 73.125
12.6 | [5 empty] | 9.35 | 70.6 | 72.9625
12.7 | [5 empty] | 9.1 | 70.26 | 72.6143
12.8 | [5 empty] | 8.875 | 70 | 72.4143
12.9 | [5 empty] | 8.62 | 69.7333 | 72.1667
13 | [5 empty] | 8.4 | 69.36 | 71.9286
13.1 | [5 empty] | 8.25 | 69.1 | 71.725
13.2 | [5 empty] | 8.04 | 68.8 | 71.5167
13.3 | [5 empty] | 7.8 | 68.48 | 71.25
13.4 | [5 empty] | 7.625 | 68.2 | 70.9667
13.5 | [5 empty] | 7.46 | 67.8 | 70.8111
13.6 | [5 empty] | 7.2714 | 67.5 | 70.46
13.7 | [5 empty] | 7.1375 | 67.2 | 70.225
13.8 | [5 empty] | 6.9667 | 66.9333 | 69.9875
13.9 | [5 empty] | 6.85 | 66.58 | 69.825
14 | [5 empty] | 6.625 | 66.2667 | 69.575
14.1 | [5 empty] | 6.5 | 65.9 | 69.3
14.2 | [5 empty] | 6.3714 | 65.5875 | 69.05
14.3 | [5 empty] | 6.2 | 65.2 | 68.8
14.4 | [5 empty] | 6.1 | 64.8571 | 68.5286
14.5 | [5 empty] | 6 | 64.5 | 68.2571
14.6 | [5 empty] | 5.9 | 64.14 | 68.0286
14.7 | [5 empty] | 5.75 | 63.825 | 67.7333
14.8 | [5 empty] | 5.62 | 63.5 | 67.5143
14.9 | [5 empty] | 5.5 | 63.1 | 67.15
15 | [5 empty] | 5.4333 | 62.7125 | 66.8833
15.1 | [5 empty] | 5.325 | 62.2 | 66.7
15.2 | [5 empty] | 5.2 | 61.8286 | 66.4857
15.3 | [5 empty] | 5.125 | 61.4375 | 66.1375
15.4 | [5 empty] | 5.025 | 61 | 65.8889
15.5 | [5 empty] | 4.92 | 60.54 | 65.7714
15.6 | [5 empty] | 4.875 | 60 | 65.5571
15.7 | [5 empty] | 4.8 | 59.4333 | 65.4
15.8 | [5 empty] | 4.7 | 58.7 | 65.15
15.9 | [5 empty] | 4.6 | 57.8 | 64.9
16 | [6 empty] | 56.72 | 64.6667
16.1 | [6 empty] | 54.9 | 64.425
16.2 | [6 empty] | 52.04 | 64.2667
16.3 | [6 empty] | 48.2 | 64
16.4 | [6 empty] | 42.9 | 63.85
16.5 | [6 empty] | 37.3 | 63.6
16.6 | [6 empty] | 31.34 | 63.4
16.7 | [6 empty] | 25.5667 | 63.15
16.8 | [6 empty] | 20.7333 | 62.95
16.9 | [6 empty] | 17.14 | 62.7
17 | [6 empty] | 14.95 | 62.4167
17.1 | [6 empty] | 13.5 | 62.14
17.2 | [6 empty] | 12.64 | 61.9333
17.3 | [6 empty] | 12 | 61.7143
17.4 | [6 empty] | 11.4667 | 61.45
17.5 | [6 empty] | 11.0875 | 61.1857
17.6 | [6 empty] | 10.7 | 60.875
17.7 | [6 empty] | 10.4 | 60.65
17.8 | [6 empty] | 10.15 | 60.475
17.9 | [6 empty] | 9.875 | 60.2333
18 | [6 empty] | 9.625 | 59.9857
18.1 | [6 empty] | 9.44 | 59.6714
18.2 | [6 empty] | 9.225 | 59.45
18.3 | [6 empty] | 8.94 | 59.25
18.4 | [6 empty] | 8.7667 | 58.9
18.5 | [6 empty] | 8.56 | 58.6667
18.6 | [6 empty] | 8.4 | 58.4286
18.7 | [6 empty] | 8.22 | 58.2429
18.8 | [6 empty] | 8.0429 | 57.9125
18.9 | [6 empty] | 7.9 | 57.7125
19 | [6 empty] | 7.7167 | 57.4333
19.1 | [6 empty] | 7.52 | 57.2
19.2 | [6 empty] | 7.4 | 56.9333
19.3 | [6 empty] | 7.25 | 56.6
19.4 | [6 empty] | 7.125 | 56.35
19.5 | [6 empty] | 6.9714 | 56.1
19.6 | [6 empty] | 6.8833 | 55.85
19.7 | [6 empty] | 6.7 | 55.5429
19.8 | [6 empty] | 6.6 | 55.2571
19.9 | [6 empty] | 6.5 | 54.9
20 | [6 empty] | 6.34 | 54.6667
20.1 | [6 empty] | 6.25 | 54.4
20.2 | [6 empty] | 6.16 | 54.1143
20.3 | [6 empty] | 6.06 | 53.7429
20.4 | [6 empty] | 5.9 | 53.4625
20.5 | [6 empty] | 5.825 | 53.125
20.6 | [6 empty] | 5.75 | 52.8125
20.7 | [6 empty] | 5.6714 | 52.4625
20.8 | [6 empty] | 5.5 | 52.1143
20.9 | [6 empty] | 5.4 | 51.8
21 | [6 empty] | 5.3429 | 51.325
21.1 | [6 empty] | 5.275 | 50.9
21.2 | [6 empty] | 5.2 | 50.4429
21.3 | [6 empty] | 5.1 | 49.7667
21.4 | [6 empty] | 5.025 | 49
21.5 | [6 empty] | 4.9 | 47.85
21.6 | [7 empty] | 46.0167
21.7 | [7 empty] | 43.3
21.8 | [7 empty] | 39.22
21.9 | [7 empty] | 34.7125
22 | [7 empty] | 29.6
22.1 | [7 empty] | 24.95
22.2 | [7 empty] | 20.5333
22.3 | [7 empty] | 17.08
22.4 | [7 empty] | 14.675
22.5 | [7 empty] | 13.1833
22.6 | [7 empty] | 12.1714
22.7 | [7 empty] | 11.6
22.8 | [7 empty] | 11.025
22.9 | [7 empty] | 10.65
23 | [7 empty] | 10.3286
23.1 | [7 empty] | 10.1333
23.2 | [7 empty] | 9.7833
23.3 | [7 empty] | 9.6
23.4 | [7 empty] | 9.4167
23.5 | [7 empty] | 9.14
23.6 | [7 empty] | 8.9714
23.7 | [7 empty] | 8.8
23.8 | [7 empty] | 8.575
23.9 | [7 empty] | 8.4429
24 | [7 empty] | 8.3
24.1 | [7 empty] | 8.1143
24.2 | [7 empty] | 8
24.3 | [7 empty] | 7.7571
24.4 | [7 empty] | 7.7
24.5 | [7 empty] | 7.4571
24.6 | [7 empty] | 7.4
24.7 | [7 empty] | 7.22
24.8 | [7 empty] | 7.1333
24.9 | [7 empty] | 6.9857
25 | [7 empty] | 6.85
25.1 | [7 empty] | 6.7
25.2 | [7 empty] | 6.6714
25.3 | [7 empty] | 6.5
25.4 | [7 empty] | 6.4
25.5 | [7 empty] | 6.3375
25.6 | [7 empty] | 6.2
25.7 | [7 empty] | 6.15
25.8 | [7 empty] | 6
25.9 | [7 empty] | 5.9167
26 | [7 empty] | 5.8714
26.1 | [7 empty] | 5.8
26.2 | [7 empty] | 5.7
26.3 | [7 empty] | 5.5833
26.4 | [7 empty] | 5.5
26.5 | [7 empty] | 5.4333
26.6 | [7 empty] | 5.3714
26.7 | [7 empty] | 5.3
26.8 | [7 empty] | 5.2
26.9 | [7 empty] | 5.15

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 8.1167
-28.8 | [7 empty] | 8.2
-28.7 | [7 empty] | 8.3125
-28.6 | [7 empty] | 8.4375
-28.5 | [7 empty] | 8.5625
-28.4 | [7 empty] | 8.6875
-28.3 | [7 empty] | 8.8
-28.2 | [7 empty] | 8.85
-28.1 | [7 empty] | 9
-28 | [7 empty] | 9.15
-27.9 | [7 empty] | 9.2
-27.8 | [7 empty] | 9.35
-27.7 | [7 empty] | 9.475
-27.6 | [7 empty] | 9.6
-27.5 | [7 empty] | 9.725
-27.4 | [7 empty] | 9.95
-27.3 | [7 empty] | 10.0143
-27.2 | [7 empty] | 10.1571
-27.1 | [7 empty] | 10.2875
-27 | [7 empty] | 10.4375
-26.9 | [7 empty] | 10.6429
-26.8 | [7 empty] | 10.8
-26.7 | [7 empty] | 10.9
-26.6 | [7 empty] | 11.125
-26.5 | [7 empty] | 11.25
-26.4 | [7 empty] | 11.4
-26.3 | [7 empty] | 11.6125
-26.2 | [7 empty] | 11.7375
-26.1 | [7 empty] | 11.925
-26 | [7 empty] | 12.125
-25.9 | [7 empty] | 12.375
-25.8 | [7 empty] | 12.5625
-25.7 | [7 empty] | 12.775
-25.6 | [7 empty] | 12.925
-25.5 | [7 empty] | 13.1667
-25.4 | [7 empty] | 13.35
-25.3 | [7 empty] | 13.5714
-25.2 | [7 empty] | 13.85
-25.1 | [7 empty] | 14.1143
-25 | [7 empty] | 14.4
-24.9 | [7 empty] | 14.775
-24.8 | [7 empty] | 15.1333
-24.7 | [7 empty] | 15.6571
-24.6 | [7 empty] | 16.2667
-24.5 | [7 empty] | 17.1
-24.4 | [7 empty] | 18.5
-24.3 | [7 empty] | 20.6667
-24.2 | [7 empty] | 23.55
-24.1 | [7 empty] | 27
-24 | [7 empty] | 30.65
-23.9 | [7 empty] | 34.55
-23.8 | [7 empty] | 38.2
-23.7 | [7 empty] | 41.45
-23.6 | [7 empty] | 43.9
-23.5 | [7 empty] | 45.525
-23.4 | [7 empty] | 46.6
-23.3 | [7 empty] | 47.4667
-23.2 | [7 empty] | 48.05
-23.1 | [7 empty] | 48.5857
-23 | [7 empty] | 48.9429
-22.9 | [6 empty] | 8.1 | 49.3333
-22.8 | [6 empty] | 8.2 | 49.7667
-22.7 | [6 empty] | 8.3 | 50.2
-22.6 | [6 empty] | 8.4429 | 50.475
-22.5 | [6 empty] | 8.6 | 50.78
-22.4 | [6 empty] | 8.7286 | 51.1286
-22.3 | [6 empty] | 8.8625 | 51.4714
-22.2 | [6 empty] | 8.975 | 51.8
-22.1 | [6 empty] | 9.1 | 52
-22 | [6 empty] | 9.25 | 52.325
-21.9 | [6 empty] | 9.4 | 52.6286
-21.8 | [6 empty] | 9.55 | 52.9143
-21.7 | [6 empty] | 9.625 | 53.25
-21.6 | [6 empty] | 9.8 | 53.5375
-21.5 | [6 empty] | 10 | 53.875
-21.4 | [6 empty] | 10.1 | 54.125
-21.3 | [6 empty] | 10.2429 | 54.4333
-21.2 | [6 empty] | 10.45 | 54.7
-21.1 | [6 empty] | 10.6 | 54.95
-21 | [6 empty] | 10.8143 | 55.2143
-20.9 | [6 empty] | 10.9571 | 55.5
-20.8 | [6 empty] | 11.1667 | 55.85
-20.7 | [6 empty] | 11.3 | 56.1333
-20.6 | [6 empty] | 11.54 | 56.3667
-20.5 | [6 empty] | 11.7 | 56.5857
-20.4 | [6 empty] | 11.9143 | 56.875
-20.3 | [6 empty] | 12.0667 | 57.125
-20.2 | [6 empty] | 12.3 | 57.375
-20.1 | [6 empty] | 12.5429 | 57.6333
-20 | [6 empty] | 12.75 | 57.95
-19.9 | [6 empty] | 12.98 | 58.1667
-19.8 | [6 empty] | 13.1286 | 58.5
-19.7 | [6 empty] | 13.4 | 58.75
-19.6 | [6 empty] | 13.6286 | 58.975
-19.5 | [6 empty] | 13.9 | 59.2
-19.4 | [6 empty] | 14.18 | 59.55
-19.3 | [6 empty] | 14.475 | 59.7444
-19.2 | [6 empty] | 14.775 | 59.9875
-19.1 | [6 empty] | 15.075 | 60.3
-19 | [6 empty] | 15.4714 | 60.5667
-18.9 | [6 empty] | 15.9 | 60.75
-18.8 | [6 empty] | 16.4 | 60.975
-18.7 | [6 empty] | 17.02 | 61.22
-18.6 | [6 empty] | 17.825 | 61.4714
-18.5 | [6 empty] | 19.1 | 61.75
-18.4 | [6 empty] | 21.075 | 62.0143
-18.3 | [6 empty] | 23.8857 | 62.2
-18.2 | [6 empty] | 27.8 | 62.4833
-18.1 | [6 empty] | 32.4 | 62.68
-18 | [6 empty] | 37.3 | 62.8857
-17.9 | [6 empty] | 42 | 63.1833
-17.8 | [6 empty] | 46.4 | 63.44
-17.7 | [6 empty] | 49.7 | 63.7286
-17.6 | [6 empty] | 52.6 | 63.98
-17.5 | [6 empty] | 54.2 | 64.1667
-17.4 | [6 empty] | 55.4 | 64.475
-17.3 | [6 empty] | 56.26 | 64.725
-17.2 | [6 empty] | 57.02 | 64.9333
-17.1 | [6 empty] | 57.6 | 65.225
-17 | [6 empty] | 58.14 | 65.4
-16.9 | [5 empty] | 7.8 | 58.55 | 65.575
-16.8 | [5 empty] | 7.975 | 59.1 | 65.9
-16.7 | [5 empty] | 8.02 | 59.54 | 66.125
-16.6 | [5 empty] | 8.2 | 59.84 | 66.375
-16.5 | [5 empty] | 8.3 | 60.3857 | 66.625
-16.4 | [5 empty] | 8.44 | 60.7429 | 66.9333
-16.3 | [5 empty] | 8.6143 | 61.1 | 67.1
-16.2 | [5 empty] | 8.7 | 61.5375 | 67.3571
-16.1 | [5 empty] | 8.8667 | 61.9 | 67.5833
-16 | [5 empty] | 9 | 62.25 | 67.84
-15.9 | [5 empty] | 9.2167 | 62.62 | 68.0333
-15.8 | [5 empty] | 9.3 | 62.9429 | 68.3111
-15.7 | [5 empty] | 9.5 | 63.3 | 68.55
-15.6 | [5 empty] | 9.7 | 63.7 | 68.8
-15.5 | [5 empty] | 9.82 | 64.0286 | 69.05
-15.4 | [5 empty] | 9.95 | 64.3667 | 69.25
-15.3 | [5 empty] | 10.1667 | 64.66 | 69.5
-15.2 | [5 empty] | 10.375 | 65.0429 | 69.7
-15.1 | [5 empty] | 10.525 | 65.35 | 69.8375
-15 | [5 empty] | 10.775 | 65.725 | 70.16
-14.9 | [5 empty] | 10.925 | 66 | 70.4333
-14.8 | [5 empty] | 11.12 | 66.3 | 70.7429
-14.7 | [5 empty] | 11.325 | 66.675 | 70.9333
-14.6 | [5 empty] | 11.56 | 66.975 | 71.2
-14.5 | [5 empty] | 11.76 | 67.3333 | 71.3667
-14.4 | [5 empty] | 11.95 | 67.625 | 71.6286
-14.3 | [5 empty] | 12.18 | 67.9 | 71.9143
-14.2 | [5 empty] | 12.45 | 68.2857 | 72.1
-14.1 | [5 empty] | 12.68 | 68.6 | 72.3333
-14 | [5 empty] | 12.95 | 68.9714 | 72.625
-13.9 | [5 empty] | 13.2 | 69.3167 | 72.8
-13.8 | [5 empty] | 13.45 | 69.525 | 73.1125
-13.7 | [5 empty] | 13.7 | 69.9 | 73.275
-13.6 | [5 empty] | 14.05 | 70.2 | 73.5667
-13.5 | [5 empty] | 14.25 | 70.4714 | 73.7571
-13.4 | [5 empty] | 14.6 | 70.78 | 74.025
-13.3 | [5 empty] | 14.95 | 71.075 | 74.2556
-13.2 | [5 empty] | 15.2667 | 71.4143 | 74.5
-13.1 | [5 empty] | 15.65 | 71.675 | 74.7571
-13 | [5 empty] | 16.05 | 72 | 74.9714
-12.9 | [5 empty] | 16.575 | 72.3571 | 75.2
-12.8 | [5 empty] | 17.16 | 72.575 | 75.4571
-12.7 | [5 empty] | 17.98 | 72.875 | 75.7667
-12.6 | [5 empty] | 19.05 | 73.2333 | 75.95
-12.5 | [5 empty] | 20.575 | 73.475 | 76.2
-12.4 | [5 empty] | 23.04 | 73.8 | 76.3857
-12.3 | [5 empty] | 26.5 | 74.1 | 76.6
-12.2 | [5 empty] | 31.45 | 74.375 | 77.0143
-12.1 | [5 empty] | 37.35 | 74.65 | 77.12
-12 | [5 empty] | 43.3333 | 75 | 77.4167
-11.9 | [5 empty] | 49.6 | 75.35 | 77.6667
-11.8 | [5 empty] | 55.3 | 75.5857 | 77.9
-11.7 | [5 empty] | 60.1 | 75.8667 | 78.05
-11.6 | [5 empty] | 63.4 | 76.1286 | 78.3
-11.5 | [5 empty] | 65.58 | 76.4143 | 78.55
-11.4 | [5 empty] | 67.05 | 76.72 | 78.9
-11.3 | [5 empty] | 68.1 | 77.05 | 79.1143
-11.2 | [5 empty] | 68.975 | 77.2667 | 79.3143
-11.1 | [5 empty] | 69.6667 | 77.6 | 79.575
-11 | [5 empty] | 70.3333 | 77.8714 | 79.825
-10.9 |  |  |  |  | 5.5 | 70.95 | 78.14 | 80.0333
-10.8 |  |  |  |  | 5.7 | 71.52 | 78.45 | 80.2
-10.7 |  |  |  |  | 5.75 | 72 | 78.72 | 80.525
-10.6 |  |  |  |  | 5.86 | 72.5 | 79.025 | 80.8
-10.5 |  |  |  |  | 6 | 73 | 79.24 | 80.95
-10.4 |  |  |  |  | 6.1 | 73.5 | 79.6 | 81.2143
-10.3 |  |  |  |  | 6.225 | 73.96 | 79.95 | 81.4333
-10.2 |  |  |  |  | 6.3667 | 74.45 | 80.16 | 81.7333
-10.1 |  |  |  |  | 6.4 | 74.9 | 80.38 | 81.9714
-10 |  |  |  |  | 6.6 | 75.3429 | 80.75 | 82.25
-9.9 |  |  |  |  | 6.7 | 75.7333 | 80.9429 | 82.4
-9.8 |  |  |  |  | 6.85 | 76.22 | 81.15 | 82.7333
-9.7 |  |  |  | 4.65 | 6.98 | 76.6714 | 81.5667 | 82.9625
-9.6 |  |  |  | 4.7 | 7.14 | 77.05 | 81.8333 | 83.175
-9.5 |  |  |  | 4.8 | 7.3 | 77.45 | 82.1167 | 83.4375
-9.4 |  |  |  | 4.9 | 7.475 | 77.9 | 82.3333 | 83.6875
-9.3 |  |  |  | 5 | 7.575 | 78.25 | 82.6 | 83.975
-9.2 |  |  |  | 5.1 | 7.8 | 78.6 | 82.9 | 84.1333
-9.1 |  |  |  | 5.2 | 7.9 | 79.04 | 83.2 | 84.3714
-9 |  |  |  | 5.35 | 8.14 | 79.45 | 83.48 | 84.6571
-8.9 |  |  |  | 5.5 | 8.325 | 79.8 | 83.75 | 84.8556
-8.8 |  |  |  | 5.6 | 8.5 | 80.2 | 84.06 | 85.0714
-8.7 |  |  |  | 5.7 | 8.75 | 80.6 | 84.28 | 85.3
-8.6 |  |  |  | 5.8 | 8.98 | 81 | 84.48 | 85.6
-8.5 |  |  | 3.5 | 5.9 | 9.14 | 81.35 | 84.8167 | 85.9
-8.4 |  |  | 3.6 | 6.1 | 9.375 | 81.7 | 85 | 86.08
-8.3 |  |  | 3.7 | 6.2 | 9.62 | 82.12 | 85.425 | 86.4
-8.2 |  |  | 3.8 | 6.325 | 9.8 | 82.4333 | 85.7 | 86.625
-8.1 |  |  | 3.9 | 6.5 | 10.025 | 82.8 | 85.84 | 86.8286
-8 |  |  | 3.9 | 6.6 | 10.35 | 83.14 | 86.175 | 87
-7.9 |  |  | 4 | 6.85 | 10.6286 | 83.525 | 86.46 | 87.2375
-7.8 |  |  | 4.1 | 7 | 10.88 | 83.85 | 86.68 | 87.5857
-7.7 |  |  | 4.2 | 7.2 | 11.1 | 84.2 | 86.86 | 87.775
-7.6 |  |  | 4.3 | 7.3 | 11.4 | 84.55 | 87.175 | 88
-7.5 |  |  | 4.4 | 7.5 | 11.7667 | 84.9 | 87.4 | 88.25
-7.4 |  |  | 4.5333 | 7.7 | 12.05 | 85.225 | 87.7 | 88.5
-7.3 |  | 2.3 | 4.7 | 7.9 | 12.4333 | 85.5 | 88 | 88.75
-7.2 |  | 2.3 | 4.7 | 8.1 | 12.8143 | 85.86 | 88.2667 | 89.025
-7.1 |  | 2.4 | 4.8 | 8.35 | 13.2 | 86.3 | 88.5 | 89.275
-7 |  | 2.4333 | 5 | 8.6 | 13.65 | 86.5857 | 88.7714 | 89.5429
-6.9 |  | 2.5 | 5.1 | 8.8 | 14.24 | 86.9875 | 89.02 | 89.6571
-6.8 |  | 2.6 | 5.2 | 9 | 15.06 | 87.24 | 89.2667 | 89.8833
-6.7 | 1.6 | 2.6 | 5.4 | 9.3 | 15.9 | 87.4667 | 89.5 | 90.1333
-6.6 | 1.7 | 2.7 | 5.5 | 9.55 | 17.125 | 87.88 | 89.8 | 90.4167
-6.5 | 1.7 | 2.7 | 5.6 | 9.7333 | 19.14 | 88.16 | 89.9833 | 90.5833
-6.4 | 1.75 | 2.8 | 5.8 | 10.1 | 22.1143 | 88.42 | 90.18 | 90.88
-6.3 | 1.8 | 2.9 | 6 | 10.3333 | 26.9 | 88.8 | 90.56 | 91.0429
-6.2 | 1.8 | 2.9 | 6.15 | 10.65 | 33.175 | 89.1 | 90.78 | 91.3167
-6.1 | 1.9 | 3 | 6.3 | 11 | 40.8 | 89.4 | 91.02 | 91.4833
-6 | 1.9 | 3.1 | 6.45 | 11.4 | 49.6 | 89.72 | 91.175 | 91.6333
-5.9 | 1.9 | 3.2 | 6.6 | 11.8333 | 57.52 | 90.025 | 91.5 | 91.925
-5.8 | 2 | 3.275 | 6.8333 | 12.3 | 65.7167 | 90.26 | 91.7 | 92.2429
-5.7 | 2.075 | 3.325 | 7.05 | 12.8 | 72.4 | 90.475 | 91.9667 | 92.3667
-5.6 | 2.1 | 3.4 | 7.2 | 13.5 | 77.275 | 90.775 | 92.25 | 92.575
-5.5 | 2.1 | 3.5 | 7.4 | 14.4 | 80.5 | 91.1 | 92.46 | 92.7375
-5.4 | 2.2 | 3.6 | 7.7 | 15.65 | 82.45 | 91.45 | 92.62 | 93.0375
-5.3 | 2.3 | 3.7 | 7.9 | 17.6 | 83.9333 | 91.675 | 92.8 | 93.2875
-5.2 | 2.3 | 3.825 | 8.2 | 20.3333 | 85 | 91.925 | 93.14 | 93.4778
-5.1 | 2.4 | 3.95 | 8.4 | 24.9 | 85.925 | 92.22 | 93.2875 | 93.6
-5 | 2.4 | 4.05 | 8.7333 | 31.7667 | 86.6833 | 92.525 | 93.525 | 93.85
-4.9 | 2.5 | 4.2 | 9.05 | 40 | 87.3 | 92.775 | 93.775 | 94.1
-4.8 | 2.6 | 4.3 | 9.4 | 48.6 | 87.9 | 93.05 | 94 | 94.275
-4.7 | 2.7 | 4.425 | 9.7 | 58.6 | 88.48 | 93.3 | 94.18 | 94.4333
-4.6 | 2.7333 | 4.6 | 10.2 | 66.95 | 88.98 | 93.52 | 94.4143 | 94.6714
-4.5 | 2.8 | 4.7667 | 10.7 | 74.1333 | 89.6 | 93.8 | 94.675 | 94.9286
-4.4 | 2.9 | 4.9 | 11.4667 | 79.7 | 90 | 93.95 | 94.85 | 95.1667
-4.3 | 3 | 5 | 12.25 | 83.55 | 90.5 | 94.2 | 95.06 | 95.3
-4.2 | 3.1 | 5.25 | 13.475 | 85.8 | 90.95 | 94.45 | 95.3 | 95.425
-4.1 | 3.2 | 5.4 | 15.4 | 87.5 | 91.35 | 94.725 | 95.46 | 95.7
-4 | 3.3 | 5.6 | 18.3 | 88.7 | 91.84 | 94.98 | 95.6 | 95.8143
-3.9 | 3.4 | 5.8 | 22.8 | 89.6 | 92.25 | 95.2 | 95.8333 | 95.9875
-3.8 | 3.5667 | 6.05 | 30.1 | 90.2 | 92.575 | 95.5 | 96.02 | 96.3125
-3.7 | 3.7 | 6.325 | 38.7 | 90.95 | 93 | 95.7 | 96.25 | 96.4375
-3.6 | 3.8 | 6.6 | 48.4 | 91.5 | 93.35 | 95.9 | 96.36 | 96.5714
-3.5 | 3.95 | 6.9 | 58.18 | 92.1 | 93.74 | 96 | 96.6125 | 96.82
-3.4 | 4.1 | 7.3 | 68.2 | 92.55 | 94.05 | 96.275 | 96.76 | 96.9
-3.3 | 4.25 | 7.85 | 76.0667 | 93.1 | 94.45 | 96.475 | 97 | 97.0375
-3.2 | 4.425 | 8.5 | 82.9 | 93.4 | 94.78 | 96.6 | 97.08 | 97.225
-3.1 | 4.675 | 9.3 | 86.9 | 93.85 | 95.1333 | 96.9 | 97.325 | 97.4333
-3 | 4.9 | 10.5 | 89.4 | 94.4 | 95.4 | 97.1 | 97.4333 | 97.575
-2.9 | 5.15 | 12.3 | 91.15 | 94.6 | 95.7667 | 97.3 | 97.625 | 97.8
-2.8 | 5.5 | 15.4333 | 92.4 | 95 | 96 | 97.45 | 97.8 | 97.95
-2.7 | 5.95 | 20.2 | 93.2 | 95.5 | 96.2333 | 97.64 | 97.9429 | 98.1
-2.6 | 6.54 | 27.9 | 94.0333 | 95.7 | 96.55 | 97.7333 | 98.075 | 98.15
-2.5 | 7.4 | 37.1 | 94.6 | 96.1 | 96.85 | 97.975 | 98.25 | 98.3
-2.4 | 8.55 | 48 | 95.2 | 96.4333 | 97.1 | 98.125 | 98.34 | 98.3857
-2.3 | 10.4 | 58.725 | 95.6 | 96.7 | 97.3 | 98.25 | 98.5 | 98.5143
-2.2 | 13.6 | 69.7 | 96.1 | 97 | 97.575 | 98.4 | 98.6 | 98.7143
-2.1 | 18.4 | 78.8 | 96.4 | 97.25 | 97.82 | 98.45 | 98.7 | 98.8
-2 | 26.3 | 86.05 | 96.7 | 97.5667 | 98 | 98.7 | 98.8667 | 98.95
-1.9 | 36.35 | 90.4 | 97.05 | 97.8333 | 98.2 | 98.8 | 98.98 | 99.1
-1.8 | 47.1 | 93.1667 | 97.4 | 98.1 | 98.375 | 98.9 | 99.06 | 99.1571
-1.7 | 59.1 | 94.9 | 97.6 | 98.2 | 98.625 | 99.025 | 99.1 | 99.2
-1.6 | 70.5 | 95.95 | 98 | 98.35 | 98.7 | 99.125 | 99.2 | 99.3
-1.5 | 79.9333 | 96.8 | 98.3 | 98.7 | 98.8 | 99.2 | 99.26 | 99.4
-1.4 | 87.1333 | 97.5 | 98.5 | 98.8 | 99 | 99.4 | 99.4 | 99.4
-1.3 | 92 | 97.95 | 98.8 | 99 | 99.1333 | 99.475 | 99.5 | 99.4333
-1.2 | 94.7 | 98.325 | 98.9 | 99.2 | 99.325 | 99.5333 | 99.6 | 99.5857
-1.1 | 96.4 | 98.7333 | 99.1 | 99.3 | 99.5 | 99.6 | 99.6286 | 99.7
-1 | 97.4 | 98.9 | 99.3 | 99.4 | 99.6 | 99.7 | 99.7 | 99.7167
-0.9 | 98.2 | 99.15 | 99.4 | 99.5 | 99.7 | 99.74 | 99.72 | 99.8
-0.8 | 98.7667 | 99.35 | 99.5 | 99.6 | 99.7833 | 99.8 | 99.8167 | 99.8
-0.7 | 99.0333 | 99.5 | 99.7 | 99.8 | 99.8 | 99.8 | 99.9 | 99.9
-0.6 | 99.4 | 99.6 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9
-0.5 | 99.725 | 99.8 | 99.8 | 99.8333 | 99.9 | 99.9 | 99.9 | 99.9
-0.4 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.92 | 99.9167 | 99.9
-0.3 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100 | 100 | 99.9
-0.2 | 99.96 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9833
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.9 | 99.9 | 99.9 | 99.9 | 100 | 100 | 99.9 | 99.9
0.4 | 99.875 | 99.9 | 99.9 | 99.9 | 100 | 99.925 | 99.9 | 99.9
0.5 | 99.72 | 99.8 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9
0.6 | 99.4 | 99.6 | 99.8 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9333
0.7 | 99.125 | 99.5 | 99.6333 | 99.7 | 99.8 | 99.9 | 99.8 | 99.8833
0.8 | 98.7667 | 99.3 | 99.5 | 99.6 | 99.7 | 99.8 | 99.8 | 99.8
0.9 | 98.0667 | 99.1 | 99.4 | 99.56 | 99.64 | 99.8 | 99.7 | 99.8
1 | 97.35 | 98.9 | 99.1667 | 99.5 | 99.56 | 99.7 | 99.7 | 99.7
1.1 | 96.2667 | 98.6 | 99 | 99.275 | 99.4 | 99.6 | 99.6 | 99.7
1.2 | 94.5 | 98.2 | 98.9 | 99.125 | 99.275 | 99.5 | 99.5 | 99.6
1.3 | 91.8 | 97.8333 | 98.65 | 99 | 99.1 | 99.4 | 99.4 | 99.4
1.4 | 87.0667 | 97.4 | 98.3 | 98.75 | 98.95 | 99.3 | 99.3 | 99.4
1.5 | 79.6 | 96.7 | 98.1667 | 98.6 | 98.8 | 99.2 | 99.2833 | 99.3875
1.6 | 70.4 | 95.9 | 98 | 98.4 | 98.675 | 99.1 | 99.2 | 99.3
1.7 | 59.1 | 94.7 | 97.4667 | 98.2 | 98.52 | 99 | 99.1 | 99.3
1.8 | 47.1 | 92.8333 | 97.2 | 98 | 98.325 | 98.85 | 99.0125 | 99.125
1.9 | 36.45 | 90.2 | 97 | 97.7 | 98.05 | 98.74 | 98.88 | 99.05
2 | 26.1333 | 85.7667 | 96.7 | 97.5 | 97.9 | 98.6 | 98.8 | 99
2.1 | 18.4 | 79.2 | 96.3 | 97.25 | 97.7 | 98.45 | 98.65 | 98.8571
2.2 | 13.55 | 69.8 | 95.9333 | 96.9 | 97.45 | 98.3 | 98.4429 | 98.7
2.3 | 10.325 | 58.9 | 95.4 | 96.6 | 97.26 | 98.2 | 98.225 | 98.5
2.4 | 8.5667 | 48.1333 | 95 | 96.3 | 97 | 98 | 98.2 | 98.3875
2.5 | 7.4 | 37.7 | 94.5 | 96 | 96.75 | 97.9 | 98.05 | 98.2571
2.6 | 6.55 | 27.7 | 93.8333 | 95.7 | 96.475 | 97.7 | 97.9 | 98.2
2.7 | 5.95 | 20.2 | 93 | 95.3333 | 96.125 | 97.525 | 97.8 | 98.0571
2.8 | 5.55 | 15.2 | 92.2 | 94.975 | 95.9 | 97.3286 | 97.6 | 97.9
2.9 | 5.2 | 12.2 | 90.9 | 94.6 | 95.6 | 97.14 | 97.5 | 97.75
3 | 4.9 | 10.45 | 89.3 | 94.2 | 95.25 | 96.975 | 97.325 | 97.625
3.1 | 4.65 | 9.3 | 86.6333 | 93.6667 | 95.02 | 96.7333 | 97.2 | 97.3833
3.2 | 4.45 | 8.5 | 82.8 | 93.4 | 94.6667 | 96.5 | 97.05 | 97.2167
3.3 | 4.275 | 7.75 | 76.2 | 93 | 94.35 | 96.3 | 96.86 | 97.0857
3.4 | 4.1 | 7.3 | 68.4 | 92.4 | 93.95 | 96.2 | 96.6 | 96.9429
3.5 | 4 | 6.9 | 58.7 | 91.9 | 93.55 | 96 | 96.4 | 96.78
3.6 | 3.8 | 6.65 | 48.4 | 91.3 | 93.2 | 95.825 | 96.24 | 96.55
3.7 | 3.7 | 6.3 | 38.8 | 90.8 | 92.85 | 95.5 | 96.04 | 96.3778
3.8 | 3.5333 | 6 | 30.1 | 90.1 | 92.5 | 95.4 | 95.86 | 96.2625
3.9 | 3.4 | 5.8 | 23 | 89.4 | 92.0333 | 95.05 | 95.66 | 96.0125
4 | 3.3 | 5.6667 | 18.2667 | 88.5 | 91.68 | 94.775 | 95.46 | 95.8125
4.1 | 3.2 | 5.4 | 15.3 | 87.2333 | 91.24 | 94.625 | 95.275 | 95.6571
4.2 | 3.1 | 5.2 | 13.4 | 85.7 | 90.825 | 94.36 | 95.08 | 95.4286
4.3 | 3 | 5.025 | 12.2333 | 83.2 | 90.35 | 94.1 | 94.9 | 95.2857
4.4 | 2.9333 | 4.9 | 11.4 | 79.6 | 89.78 | 93.8 | 94.7 | 95.2
4.5 | 2.8 | 4.7 | 10.7 | 74.1 | 89.3 | 93.65 | 94.475 | 94.9714
4.6 | 2.8 | 4.575 | 10.2 | 67 | 88.78 | 93.3 | 94.36 | 94.6857
4.7 | 2.7 | 4.4 | 9.7 | 58.6 | 88.2 | 93.05 | 94.06 | 94.4667
4.8 | 2.6 | 4.3 | 9.4 | 48.6 | 87.68 | 92.75 | 93.88 | 94.3
4.9 | 2.5 | 4.2 | 9 | 39.9 | 87.1 | 92.6 | 93.6 | 94.1
5 | 2.4 | 4.1 | 8.7 | 31.85 | 86.46 | 92.25 | 93.46 | 93.85
5.1 | 2.4 | 3.95 | 8.45 | 24.7 | 85.7 | 92.1 | 93.14 | 93.65
5.2 | 2.3 | 3.8 | 8.2 | 20.1333 | 84.8 | 91.775 | 92.9333 | 93.525
5.3 | 2.3 | 3.7 | 7.9 | 17.5 | 83.7 | 91.525 | 92.74 | 93.2333
5.4 | 2.2 | 3.6 | 7.7 | 15.55 | 82.125 | 91.28 | 92.48 | 92.975
5.5 | 2.2 | 3.5 | 7.45 | 14.4 | 80.125 | 90.94 | 92.1833 | 92.7625
5.6 | 2.1 | 3.45 | 7.2 | 13.5 | 76.9 | 90.525 | 91.9333 | 92.575
5.7 | 2.0667 | 3.35 | 7.0667 | 12.8 | 71.95 | 90.28 | 91.7 | 92.4125
5.8 | 2 | 3.3 | 6.8 | 12.25 | 65.775 | 90.0333 | 91.4286 | 92.1833
5.9 | 1.95 | 3.2 | 6.6 | 11.8333 | 57.6 | 89.82 | 91.2 | 91.9571
6 | 1.9 | 3.1 | 6.5 | 11.4 | 49.0286 | 89.45 | 91.1 | 91.6
6.1 | 1.8667 | 3 | 6.3 | 11.0667 | 41.1333 | 89.2 | 90.8375 | 91.5333
6.2 | 1.8 | 2.9 | 6.1 | 10.675 | 33.2667 | 88.8857 | 90.625 | 91.2667
6.3 | 1.8 | 2.9 | 5.95 | 10.4 | 26.5667 | 88.55 | 90.2667 | 91
6.4 | 1.7333 | 2.8 | 5.8 | 10.075 | 22.05 | 88.3143 | 90.0714 | 90.9333
6.5 | 1.7 | 2.75 | 5.7 | 9.75 | 19.05 | 88.06 | 89.8333 | 90.58
6.6 | 1.6 | 2.7 | 5.5 | 9.475 | 17.2 | 87.65 | 89.55 | 90.3667
6.7 | 1.6 | 2.6 | 5.4 | 9.3 | 15.9 | 87.3333 | 89.2857 | 90.1286
6.8 |  | 2.6 | 5.2667 | 9 | 15.0714 | 86.98 | 88.92 | 89.94
6.9 |  | 2.5 | 5.1 | 8.7667 | 14.325 | 86.7 | 88.7125 | 89.6857
7 |  | 2.5 | 4.95 | 8.52 | 13.72 | 86.3333 | 88.44 | 89.52
7.1 |  | 2.4 | 4.8667 | 8.36 | 13.22 | 86 | 88.24 | 89.275
7.2 |  | 2.35 | 4.7 | 8.1333 | 12.8 | 85.7 | 88.04 | 89.025
7.3 |  | 2.3 | 4.6 | 7.9 | 12.42 | 85.375 | 87.84 | 88.7333
7.4 |  |  | 4.5 | 7.7 | 12.1143 | 85.025 | 87.54 | 88.5
7.5 |  |  | 4.4 | 7.5 | 11.7333 | 84.68 | 87.32 | 88.25
7.6 |  |  | 4.3 | 7.35 | 11.4333 | 84.35 | 87 | 88
7.7 |  |  | 4.2 | 7.1 | 11.2 | 84 | 86.68 | 87.75
7.8 |  |  | 4.1 | 7 | 10.9 | 83.65 | 86.5 | 87.5625
7.9 |  |  | 4 | 6.8 | 10.58 | 83.2667 | 86.2 | 87.225
8 |  |  | 3.9 | 6.6 | 10.4 | 82.9333 | 85.9 | 87.0286
8.1 |  |  | 3.9 | 6.5 | 10.08 | 82.66 | 85.6333 | 86.8714
8.2 |  |  | 3.8 | 6.32 | 9.85 | 82.25 | 85.3667 | 86.6143
8.3 |  |  | 3.7 | 6.2 | 9.62 | 81.9 | 85.1 | 86.3571
8.4 |  |  | 3.6 | 6.1 | 9.375 | 81.52 | 84.8 | 86.1
8.5 |  |  | 3.6 | 6 | 9.14 | 81.15 | 84.55 | 85.9
8.6 |  |  |  | 5.8 | 8.98 | 80.75 | 84.275 | 85.6
8.7 |  |  |  | 5.7 | 8.7 | 80.34 | 83.96 | 85.3
8.8 |  |  |  | 5.6 | 8.55 | 80.1 | 83.7 | 85.05
8.9 |  |  |  | 5.5 | 8.3 | 79.62 | 83.42 | 84.9111
9 |  |  |  | 5.3 | 8.1 | 79.25 | 83.22 | 84.5833
9.1 |  |  |  | 5.2 | 7.975 | 78.82 | 82.9667 | 84.4167
9.2 |  |  |  | 5.1 | 7.8 | 78.425 | 82.7 | 84.0714
9.3 |  |  |  | 5 | 7.6 | 78.05 | 82.4286 | 83.9375
9.4 |  |  |  | 4.9 | 7.475 | 77.65 | 82.125 | 83.6375
9.5 |  |  |  | 4.8 | 7.35 | 77.175 | 81.8333 | 83.4714
9.6 |  |  |  | 4.7 | 7.125 | 76.82 | 81.5333 | 83.1857
9.7 |  |  |  | 4.6 | 7 | 76.3857 | 81.32 | 82.925
9.8 |  |  |  |  | 6.9 | 75.98 | 81.0286 | 82.7143
9.9 |  |  |  |  | 6.75 | 75.5286 | 80.84 | 82.4167
10 |  |  |  |  | 6.6 | 75.1 | 80.4667 | 82.2429
10.1 |  |  |  |  | 6.4333 | 74.65 | 80.2 | 81.9333
10.2 |  |  |  |  | 6.32 | 74.2 | 79.9667 | 81.7333
10.3 |  |  |  |  | 6.2 | 73.76 | 79.66 | 81.4667
10.4 |  |  |  |  | 6.1 | 73.32 | 79.34 | 81.2
10.5 |  |  |  |  | 5.95 | 72.8 | 79.0833 | 80.975
10.6 |  |  |  |  | 5.8 | 72.3 | 78.75 | 80.75
10.7 |  |  |  |  | 5.7 | 71.85 | 78.54 | 80.525
10.8 |  |  |  |  | 5.65 | 71.35 | 78.2857 | 80.2333
10.9 |  |  |  |  | 5.5 | 70.775 | 77.9375 | 80.075
11 | [5 empty] | 70.12 | 77.58 | 79.825
11.1 | [5 empty] | 69.475 | 77.325 | 79.575
11.2 | [5 empty] | 68.8333 | 77.0333 | 79.325
11.3 | [5 empty] | 67.9 | 76.8 | 79.075
11.4 | [5 empty] | 66.8 | 76.5 | 78.85
11.5 | [5 empty] | 65.3 | 76.2333 | 78.55
11.6 | [5 empty] | 63.14 | 75.9 | 78.3
11.7 | [5 empty] | 59.86 | 75.5857 | 78.075
11.8 | [5 empty] | 55.2667 | 75.3667 | 77.9
11.9 | [5 empty] | 49.675 | 75.04 | 77.65
12 | [5 empty] | 43.8 | 74.8143 | 77.42
12.1 | [5 empty] | 37.3 | 74.5 | 77.1667
12.2 | [5 empty] | 31.55 | 74.15 | 77.0143
12.3 | [5 empty] | 26.625 | 73.92 | 76.65
12.4 | [5 empty] | 22.95 | 73.5667 | 76.375
12.5 | [5 empty] | 20.45 | 73.28 | 76.15
12.6 | [5 empty] | 19.06 | 72.925 | 75.9
12.7 | [5 empty] | 18 | 72.72 | 75.7333
12.8 | [5 empty] | 17.2 | 72.46 | 75.45
12.9 | [5 empty] | 16.6 | 72.1333 | 75.2
13 | [5 empty] | 16.05 | 71.775 | 75.025
13.1 | [5 empty] | 15.625 | 71.5 | 74.825
13.2 | [5 empty] | 15.2333 | 71.2 | 74.5
13.3 | [5 empty] | 14.9 | 70.88 | 74.2667
13.4 | [5 empty] | 14.6 | 70.6286 | 74.025
13.5 | [5 empty] | 14.3 | 70.3 | 73.7667
13.6 | [5 empty] | 13.9333 | 69.95 | 73.5286
13.7 | [5 empty] | 13.7 | 69.72 | 73.2875
13.8 | [5 empty] | 13.45 | 69.375 | 73.0875
13.9 | [5 empty] | 13.18 | 69 | 72.7889
14 | [5 empty] | 12.925 | 68.7 | 72.55
14.1 | [5 empty] | 12.6 | 68.4333 | 72.3857
14.2 | [5 empty] | 12.4 | 68.125 | 72.1286
14.3 | [5 empty] | 12.2 | 67.7667 | 71.9125
14.4 | [5 empty] | 11.975 | 67.5 | 71.6
14.5 | [5 empty] | 11.7 | 67.175 | 71.3143
14.6 | [5 empty] | 11.5 | 66.8375 | 71.15
14.7 | [5 empty] | 11.2571 | 66.52 | 70.9
14.8 | [5 empty] | 11.05 | 66.1667 | 70.7333
14.9 | [5 empty] | 10.9 | 65.825 | 70.4429
15 | [5 empty] | 10.675 | 65.4857 | 70.2
15.1 | [5 empty] | 10.525 | 65.1714 | 69.9
15.2 | [5 empty] | 10.3 | 64.8667 | 69.65
15.3 | [5 empty] | 10.125 | 64.5 | 69.45
15.4 | [5 empty] | 10 | 64.1667 | 69.3
15.5 | [5 empty] | 9.82 | 63.8 | 69.05
15.6 | [5 empty] | 9.62 | 63.425 | 68.8
15.7 | [5 empty] | 9.5 | 63.12 | 68.55
15.8 | [5 empty] | 9.2833 | 62.7857 | 68.3222
15.9 | [5 empty] | 9.2 | 62.4 | 68.0333
16 | [5 empty] | 9 | 62.0667 | 67.8667
16.1 | [5 empty] | 8.82 | 61.675 | 67.575
16.2 | [5 empty] | 8.7 | 61.25 | 67.3667
16.3 | [5 empty] | 8.5667 | 60.9 | 67.1286
16.4 | [5 empty] | 8.425 | 60.475 | 66.9333
16.5 | [5 empty] | 8.34 | 60.1667 | 66.6286
16.6 | [5 empty] | 8.2 | 59.7714 | 66.375
16.7 | [5 empty] | 8.05 | 59.34 | 66.1333
16.8 | [5 empty] | 7.875 | 58.8429 | 65.85
16.9 | [5 empty] | 7.8 | 58.3667 | 65.625
17 | [6 empty] | 57.95 | 65.4
17.1 | [6 empty] | 57.4 | 65.15
17.2 | [6 empty] | 56.76 | 64.9
17.3 | [6 empty] | 56.08 | 64.7143
17.4 | [6 empty] | 55.2 | 64.4667
17.5 | [6 empty] | 54.04 | 64.1333
17.6 | [6 empty] | 52.3 | 64
17.7 | [6 empty] | 49.9 | 63.7333
17.8 | [6 empty] | 46.5 | 63.4714
17.9 | [6 empty] | 41.98 | 63.2333
18 | [6 empty] | 37.42 | 62.9714
18.1 | [6 empty] | 32.45 | 62.64
18.2 | [6 empty] | 27.86 | 62.5
18.3 | [6 empty] | 24.1333 | 62.25
18.4 | [6 empty] | 21.12 | 62
18.5 | [6 empty] | 19.15 | 61.675
18.6 | [6 empty] | 17.9 | 61.4667
18.7 | [6 empty] | 16.98 | 61.2667
18.8 | [6 empty] | 16.3857 | 60.975
18.9 | [6 empty] | 15.85 | 60.7
19 | [6 empty] | 15.45 | 60.54
19.1 | [6 empty] | 15.0125 | 60.26
19.2 | [6 empty] | 14.7333 | 60.0444
19.3 | [6 empty] | 14.4143 | 59.7778
19.4 | [6 empty] | 14.2 | 59.525
19.5 | [6 empty] | 13.88 | 59.2
19.6 | [6 empty] | 13.64 | 59.025
19.7 | [6 empty] | 13.375 | 58.7
19.8 | [6 empty] | 13.1 | 58.4143
19.9 | [6 empty] | 12.95 | 58.2143
20 | [6 empty] | 12.75 | 57.9556
20.1 | [6 empty] | 12.48 | 57.6875
20.2 | [6 empty] | 12.26 | 57.375
20.3 | [6 empty] | 12.0667 | 57.125
20.4 | [6 empty] | 11.84 | 56.8714
20.5 | [6 empty] | 11.7 | 56.5857
20.6 | [6 empty] | 11.525 | 56.3667
20.7 | [6 empty] | 11.3 | 56.1167
20.8 | [6 empty] | 11.1286 | 55.825
20.9 | [6 empty] | 10.96 | 55.5
21 | [6 empty] | 10.76 | 55.2
21.1 | [6 empty] | 10.6125 | 55.025
21.2 | [6 empty] | 10.3667 | 54.7
21.3 | [6 empty] | 10.2 | 54.4667
21.4 | [6 empty] | 10.075 | 54.1375
21.5 | [6 empty] | 9.96 | 53.8625
21.6 | [6 empty] | 9.78 | 53.525
21.7 | [6 empty] | 9.6 | 53.3
21.8 | [6 empty] | 9.5 | 52.9333
21.9 | [6 empty] | 9.3 | 52.58
22 | [6 empty] | 9.2143 | 52.3286
22.1 | [6 empty] | 9.1 | 52.0667
22.2 | [6 empty] | 8.9286 | 51.8
22.3 | [6 empty] | 8.82 | 51.4857
22.4 | [6 empty] | 8.6333 | 51.2
22.5 | [6 empty] | 8.54 | 50.8
22.6 | [6 empty] | 8.4 | 50.42
22.7 | [6 empty] | 8.3 | 50.1286
22.8 | [7 empty] | 49.75
22.9 | [7 empty] | 49.3333
23 | [7 empty] | 49.0143
23.1 | [7 empty] | 48.55
23.2 | [7 empty] | 48.0667
23.3 | [7 empty] | 47.4571
23.4 | [7 empty] | 46.55
23.5 | [7 empty] | 45.525
23.6 | [7 empty] | 44
23.7 | [7 empty] | 41.475
23.8 | [7 empty] | 38.2
23.9 | [7 empty] | 34.55
24 | [7 empty] | 30.6
24.1 | [7 empty] | 26.8333
24.2 | [7 empty] | 23.5625
24.3 | [7 empty] | 20.5714
24.4 | [7 empty] | 18.45
24.5 | [7 empty] | 17.1286
24.6 | [7 empty] | 16.2667
24.7 | [7 empty] | 15.5714
24.8 | [7 empty] | 15.1556
24.9 | [7 empty] | 14.725
25 | [7 empty] | 14.4
25.1 | [7 empty] | 14.1
25.2 | [7 empty] | 13.85
25.3 | [7 empty] | 13.5667
25.4 | [7 empty] | 13.4
25.5 | [7 empty] | 13.175
25.6 | [7 empty] | 12.9625
25.7 | [7 empty] | 12.775
25.8 | [7 empty] | 12.525
25.9 | [7 empty] | 12.375
26 | [7 empty] | 12.1571
26.1 | [7 empty] | 11.9286
26.2 | [7 empty] | 11.775
26.3 | [7 empty] | 11.5667
26.4 | [7 empty] | 11.4429
26.5 | [7 empty] | 11.3
26.6 | [7 empty] | 11.0714
26.7 | [7 empty] | 10.9286
26.8 | [7 empty] | 10.7625
26.9 | [7 empty] | 10.6167
27 | [7 empty] | 10.4714
27.1 | [7 empty] | 10.3
27.2 | [7 empty] | 10.125
27.3 | [7 empty] | 10.0571
27.4 | [7 empty] | 9.9
27.5 | [7 empty] | 9.7625
27.6 | [7 empty] | 9.6444
27.7 | [7 empty] | 9.525
27.8 | [7 empty] | 9.4
27.9 | [7 empty] | 9.26
28 | [7 empty] | 9.2
28.1 | [7 empty] | 9.0571
28.2 | [7 empty] | 8.9
28.3 | [7 empty] | 8.7625
28.4 | [7 empty] | 8.6375
28.5 | [7 empty] | 8.6
28.6 | [7 empty] | 8.4875
28.7 | [7 empty] | 8.3625
28.8 | [7 empty] | 8.2286
28.9 | [7 empty] | 8.1667

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-30.9 | [7 empty] | 10.84
-30.8 | [7 empty] | 10.9875
-30.7 | [7 empty] | 11.1
-30.6 | [7 empty] | 11.1125
-30.5 | [7 empty] | 11.2429
-30.4 | [7 empty] | 11.3857
-30.3 | [7 empty] | 11.5125
-30.2 | [7 empty] | 11.6429
-30.1 | [7 empty] | 11.775
-30 | [7 empty] | 11.9
-29.9 | [7 empty] | 12.025
-29.8 | [7 empty] | 12.15
-29.7 | [7 empty] | 12.35
-29.6 | [7 empty] | 12.5
-29.5 | [7 empty] | 12.5714
-29.4 | [7 empty] | 12.7125
-29.3 | [7 empty] | 12.875
-29.2 | [7 empty] | 13
-29.1 | [7 empty] | 13.2125
-29 | [7 empty] | 13.3375
-28.9 | [7 empty] | 13.4625
-28.8 | [7 empty] | 13.675
-28.7 | [7 empty] | 13.825
-28.6 | [7 empty] | 13.9571
-28.5 | [7 empty] | 14.2
-28.4 | [7 empty] | 14.325
-28.3 | [7 empty] | 14.5
-28.2 | [7 empty] | 14.7333
-28.1 | [7 empty] | 14.8875
-28 | [7 empty] | 15.1125
-27.9 | [7 empty] | 15.2375
-27.8 | [7 empty] | 15.425
-27.7 | [7 empty] | 15.675
-27.6 | [7 empty] | 15.9
-27.5 | [7 empty] | 16.0667
-27.4 | [7 empty] | 16.2571
-27.3 | [7 empty] | 16.5
-27.2 | [7 empty] | 16.75
-27.1 | [7 empty] | 17
-27 | [7 empty] | 17.325
-26.9 | [7 empty] | 17.5778
-26.8 | [7 empty] | 18.0375
-26.7 | [7 empty] | 18.45
-26.6 | [7 empty] | 19.0333
-26.5 | [7 empty] | 19.8
-26.4 | [7 empty] | 21.0857
-26.3 | [7 empty] | 22.84
-26.2 | [7 empty] | 25.3429
-26.1 | [7 empty] | 27.9667
-26 | [7 empty] | 31.4
-25.9 | [7 empty] | 34.475
-25.8 | [7 empty] | 37.5375
-25.7 | [7 empty] | 40.425
-25.6 | [7 empty] | 42.66
-25.5 | [7 empty] | 44.22
-25.4 | [7 empty] | 45.4
-25.3 | [7 empty] | 46.2
-25.2 | [7 empty] | 46.7333
-25.1 | [7 empty] | 47.3
-25 | [7 empty] | 47.7125
-24.9 | [7 empty] | 48.15
-24.8 | [7 empty] | 48.4
-24.7 | [7 empty] | 48.84
-24.6 | [7 empty] | 49.2
-24.5 | [7 empty] | 49.5667
-24.4 | [6 empty] | 11.28 | 49.7667
-24.3 | [6 empty] | 11.4 | 50.1
-24.2 | [6 empty] | 11.58 | 50.4286
-24.1 | [6 empty] | 11.7 | 50.68
-24 | [6 empty] | 11.8 | 51
-23.9 | [6 empty] | 11.9333 | 51.3286
-23.8 | [6 empty] | 12.05 | 51.55
-23.7 | [6 empty] | 12.2333 | 51.8
-23.6 | [6 empty] | 12.3857 | 52.1833
-23.5 | [6 empty] | 12.56 | 52.35
-23.4 | [6 empty] | 12.68 | 52.7286
-23.3 | [6 empty] | 12.85 | 53.0333
-23.2 | [6 empty] | 13.0167 | 53.22
-23.1 | [6 empty] | 13.1833 | 53.5
-23 | [6 empty] | 13.3 | 53.7714
-22.9 | [6 empty] | 13.4714 | 54.1167
-22.8 | [6 empty] | 13.68 | 54.3667
-22.7 | [6 empty] | 13.8 | 54.6143
-22.6 | [6 empty] | 13.98 | 54.86
-22.5 | [6 empty] | 14.2 | 55.1333
-22.4 | [6 empty] | 14.42 | 55.3833
-22.3 | [6 empty] | 14.6 | 55.6333
-22.2 | [6 empty] | 14.78 | 55.9333
-22.1 | [6 empty] | 14.925 | 56.1833
-22 | [6 empty] | 15.15 | 56.4333
-21.9 | [6 empty] | 15.35 | 56.7
-21.8 | [6 empty] | 15.5167 | 56.95
-21.7 | [6 empty] | 15.7429 | 57.2143
-21.6 | [6 empty] | 15.9333 | 57.475
-21.5 | [6 empty] | 16.2 | 57.7286
-21.4 | [6 empty] | 16.4333 | 57.9571
-21.3 | [6 empty] | 16.68 | 58.175
-21.2 | [6 empty] | 16.9 | 58.425
-21.1 | [6 empty] | 17.15 | 58.7
-21 | [6 empty] | 17.3667 | 58.9571
-20.9 | [6 empty] | 17.64 | 59.2667
-20.8 | [6 empty] | 17.85 | 59.4333
-20.7 | [6 empty] | 18.26 | 59.7
-20.6 | [6 empty] | 18.5 | 59.9857
-20.5 | [6 empty] | 18.84 | 60.175
-20.4 | [6 empty] | 19.2667 | 60.4
-20.3 | [6 empty] | 19.7 | 60.725
-20.2 | [6 empty] | 20.3 | 60.9143
-20.1 | [6 empty] | 21.1 | 61.2167
-20 | [6 empty] | 22.14 | 61.4
-19.9 | [6 empty] | 23.8 | 61.6714
-19.8 | [6 empty] | 26.25 | 61.94
-19.7 | [6 empty] | 29.375 | 62.1571
-19.6 | [6 empty] | 32.8833 | 62.3625
-19.5 | [6 empty] | 37.3 | 62.6333
-19.4 | [6 empty] | 41.15 | 62.8714
-19.3 | [6 empty] | 44.975 | 63.125
-19.2 | [6 empty] | 48.55 | 63.36
-19.1 | [6 empty] | 50.96 | 63.5333
-19 | [6 empty] | 52.74 | 63.775
-18.9 | [6 empty] | 54 | 64.0571
-18.8 | [6 empty] | 54.95 | 64.2
-18.7 | [6 empty] | 55.6667 | 64.5286
-18.6 | [6 empty] | 56.3 | 64.7667
-18.5 | [6 empty] | 56.775 | 65.025
-18.4 | [6 empty] | 57.28 | 65.2375
-18.3 | [6 empty] | 57.75 | 65.4667
-18.2 | [6 empty] | 58.2 | 65.65
-18.1 | [6 empty] | 58.55 | 65.95
-18 | [6 empty] | 58.8667 | 66.25
-17.9 | [5 empty] | 10.4333 | 59.3 | 66.4
-17.8 | [5 empty] | 10.5833 | 59.64 | 66.7
-17.7 | [5 empty] | 10.78 | 60.05 | 66.8
-17.6 | [5 empty] | 10.8714 | 60.42 | 67.0333
-17.5 | [5 empty] | 11.0667 | 60.82 | 67.35
-17.4 | [5 empty] | 11.2 | 61.175 | 67.5444
-17.3 | [5 empty] | 11.4 | 61.5 | 67.7875
-17.2 | [5 empty] | 11.5 | 61.8667 | 67.9875
-17.1 | [5 empty] | 11.68 | 62.1667 | 68.1375
-17 | [5 empty] | 11.8 | 62.5 | 68.45
-16.9 | [5 empty] | 12.0429 | 62.78 | 68.675
-16.8 | [5 empty] | 12.225 | 63.2 | 68.92
-16.7 | [5 empty] | 12.4 | 63.52 | 69.1167
-16.6 | [5 empty] | 12.575 | 63.8333 | 69.45
-16.5 | [5 empty] | 12.78 | 64.15 | 69.6286
-16.4 | [5 empty] | 12.92 | 64.5 | 69.8
-16.3 | [5 empty] | 13.1333 | 64.85 | 70.0857
-16.2 | [5 empty] | 13.3571 | 65.2 | 70.4167
-16.1 | [5 empty] | 13.54 | 65.45 | 70.625
-16 | [5 empty] | 13.8 | 65.8 | 70.7667
-15.9 | [5 empty] | 14 | 66.125 | 71.0714
-15.8 | [5 empty] | 14.2 | 66.48 | 71.25
-15.7 | [5 empty] | 14.4 | 66.65 | 71.5
-15.6 | [5 empty] | 14.6333 | 66.9667 | 71.66
-15.5 | [5 empty] | 14.875 | 67.32 | 71.9571
-15.4 | [5 empty] | 15.1 | 67.6 | 72.1
-15.3 | [5 empty] | 15.375 | 67.95 | 72.35
-15.2 | [5 empty] | 15.5714 | 68.2667 | 72.6
-15.1 | [5 empty] | 15.875 | 68.5 | 72.85
-15 | [5 empty] | 16.075 | 68.8 | 73.1
-14.9 | [5 empty] | 16.325 | 69.18 | 73.3
-14.8 | [5 empty] | 16.575 | 69.46 | 73.55
-14.7 | [5 empty] | 16.86 | 69.7333 | 73.6875
-14.6 | [5 empty] | 17.1833 | 70.0667 | 73.9333
-14.5 | [5 empty] | 17.4 | 70.4 | 74.15
-14.4 | [5 empty] | 17.75 | 70.5571 | 74.4333
-14.3 | [5 empty] | 18.0667 | 70.8333 | 74.6833
-14.2 | [5 empty] | 18.45 | 71 | 74.9571
-14.1 | [5 empty] | 18.9 | 71.45 | 75.1167
-14 | [5 empty] | 19.3 | 71.7429 | 75.3667
-13.9 | [5 empty] | 19.8 | 72.2 | 75.4857
-13.8 | [5 empty] | 20.5 | 72.3667 | 75.7714
-13.7 | [5 empty] | 21.175 | 72.6333 | 76.0857
-13.6 | [5 empty] | 22.2 | 72.92 | 76.275
-13.5 | [5 empty] | 23.9 | 73.15 | 76.5143
-13.4 | [5 empty] | 26.1 | 73.4 | 76.6667
-13.3 | [5 empty] | 29.42 | 73.825 | 76.9571
-13.2 | [5 empty] | 33.775 | 74.15 | 77.225
-13.1 | [5 empty] | 38.9 | 74.4333 | 77.475
-13 | [5 empty] | 44.2 | 74.6667 | 77.6333
-12.9 | [5 empty] | 49.7286 | 74.9 | 77.8714
-12.8 | [5 empty] | 54.95 | 75.1 | 78.16
-12.7 | [5 empty] | 58.8714 | 75.5167 | 78.3571
-12.6 | [5 empty] | 62.1333 | 75.7667 | 78.5714
-12.5 | [5 empty] | 64.2 | 75.95 | 78.8333
-12.4 | [5 empty] | 65.7 | 76.3333 | 79.14
-12.3 | [5 empty] | 66.875 | 76.5714 | 79.3
-12.2 | [5 empty] | 67.8 | 76.84 | 79.4571
-12.1 | [5 empty] | 68.4 | 77.08 | 79.7333
-12 | [5 empty] | 69.0333 | 77.38 | 80
-11.9 | [5 empty] | 69.65 | 77.7 | 80.225
-11.8 | [5 empty] | 70.2 | 77.95 | 80.3875
-11.7 | [5 empty] | 70.7 | 78.2 | 80.7125
-11.6 | [5 empty] | 71.2 | 78.56 | 80.875
-11.5 | [5 empty] | 71.6667 | 78.8 | 81.1875
-11.4 |  |  |  |  | 7.6 | 72.1 | 79 | 81.3
-11.3 |  |  |  |  | 7.7 | 72.6 | 79.325 | 81.5286
-11.2 |  |  |  |  | 7.8 | 73 | 79.5 | 81.7429
-11.1 |  |  |  |  | 7.9 | 73.425 | 79.76 | 82
-11 |  |  |  |  | 8.15 | 73.85 | 80.1429 | 82.3429
-10.9 |  |  |  |  | 8.3 | 74.3 | 80.3 | 82.4
-10.8 |  |  |  |  | 8.4 | 74.725 | 80.5667 | 82.7429
-10.7 |  |  |  |  | 8.575 | 75.12 | 80.8667 | 82.8333
-10.6 |  |  |  |  | 8.7 | 75.6 | 81.2 | 83.0833
-10.5 |  |  |  |  | 8.8 | 75.95 | 81.3833 | 83.325
-10.4 |  |  |  |  | 9 | 76.35 | 81.675 | 83.4571
-10.3 |  |  |  |  | 9.2 | 76.675 | 81.9333 | 83.8143
-10.2 |  |  |  |  | 9.3 | 77.1 | 82.2 | 83.9571
-10.1 |  |  |  | 6.4 | 9.5 | 77.55 | 82.4 | 84.2333
-10 |  |  |  | 6.5 | 9.7 | 77.875 | 82.66 | 84.4714
-9.9 |  |  |  | 6.6 | 9.85 | 78.25 | 82.9 | 84.65
-9.8 |  |  |  | 6.7 | 10.05 | 78.6 | 83.3333 | 84.9
-9.7 |  |  |  | 6.9 | 10.3 | 78.95 | 83.5833 | 85.2
-9.6 |  |  |  | 7 | 10.4 | 79.3571 | 83.7667 | 85.3
-9.5 |  |  |  | 7.1 | 10.6 | 79.68 | 83.9667 | 85.5429
-9.4 |  |  |  | 7.3 | 10.8 | 80.15 | 84.2 | 85.82
-9.3 |  |  |  | 7.4 | 11 | 80.55 | 84.5 | 86.1
-9.2 |  |  |  | 7.5333 | 11.2333 | 80.9 | 84.775 | 86.2333
-9.1 |  |  |  | 7.6 | 11.4333 | 81.2 | 85.14 | 86.3667
-9 |  |  |  | 7.85 | 11.65 | 81.4 | 85.3 | 86.575
-8.9 |  |  |  | 8 | 11.9 | 81.875 | 85.5333 | 86.725
-8.8 |  |  | 5 | 8.1 | 12.2 | 82.25 | 85.7833 | 87.0125
-8.7 |  |  | 5 | 8.25 | 12.5 | 82.6 | 85.95 | 87.3429
-8.6 |  |  | 5.1 | 8.4 | 12.65 | 82.8667 | 86.2286 | 87.525
-8.5 |  |  | 5.2 | 8.6 | 12.9 | 83.225 | 86.4667 | 87.7
-8.4 |  |  | 5.3 | 8.8 | 13.15 | 83.5286 | 86.75 | 87.95
-8.3 |  |  | 5.4 | 8.9 | 13.5 | 84 | 86.9667 | 88.1
-8.2 |  |  | 5.5 | 9.0667 | 13.75 | 84.4 | 87.3 | 88.425
-8.1 |  |  | 5.6 | 9.3 | 14.075 | 84.5833 | 87.4 | 88.65
-8 |  |  | 5.7 | 9.5333 | 14.4 | 84.9 | 87.75 | 88.875
-7.9 |  |  | 5.9 | 9.7 | 14.6333 | 85.2 | 87.9 | 88.9889
-7.8 |  |  | 5.9 | 9.9 | 15 | 85.5167 | 88.2 | 89.2375
-7.7 |  |  | 6.1 | 10.1 | 15.4 | 85.825 | 88.52 | 89.575
-7.6 |  |  | 6.2 | 10.35 | 15.8 | 86.15 | 88.72 | 89.7625
-7.5 |  | 3.2 | 6.35 | 10.6 | 16.3 | 86.4 | 88.9 | 89.8778
-7.4 |  | 3.3 | 6.4 | 10.8 | 17 | 86.7 | 89.12 | 90.1
-7.3 |  | 3.3 | 6.6 | 11.05 | 17.7 | 87.1286 | 89.35 | 90.425
-7.2 |  | 3.4 | 6.7 | 11.3 | 18.65 | 87.3 | 89.56 | 90.55
-7.1 |  | 3.5 | 6.85 | 11.5 | 20 | 87.6 | 89.9 | 90.75
-7 |  | 3.6 | 7 | 11.8 | 22 | 87.8667 | 90 | 90.8889
-6.9 | 2.2 | 3.6 | 7.1 | 12.05 | 24.9 | 88.22 | 90.3333 | 91.1667
-6.8 | 2.2 | 3.7 | 7.3 | 12.4 | 29.6 | 88.5 | 90.5333 | 91.3286
-6.7 | 2.3 | 3.7 | 7.5 | 12.7 | 35.3333 | 88.8333 | 90.8 | 91.6143
-6.6 | 2.3 | 3.8 | 7.7 | 13 | 42.6667 | 89.15 | 90.9667 | 91.8
-6.5 | 2.4 | 3.9 | 7.9 | 13.36 | 50.35 | 89.4667 | 91.2 | 91.9667
-6.4 | 2.4 | 4 | 8.1 | 13.7 | 58.2 | 89.7 | 91.5167 | 92.1286
-6.3 | 2.5 | 4.1 | 8.2667 | 14.15 | 65.8 | 90.025 | 91.6833 | 92.325
-6.2 | 2.5 | 4.1 | 8.4 | 14.6 | 71.6 | 90.3 | 91.9 | 92.575
-6.1 | 2.6 | 4.2 | 8.6333 | 15.1667 | 76.1667 | 90.5333 | 92.1 | 92.8125
-6 | 2.7 | 4.3 | 8.85 | 16 | 79.4 | 90.8286 | 92.3 | 92.95
-5.9 | 2.725 | 4.4 | 9.1 | 16.9 | 81.55 | 91.12 | 92.5333 | 93.1333
-5.8 | 2.8 | 4.6 | 9.3 | 18.3 | 83.1 | 91.35 | 92.68 | 93.275
-5.7 | 2.8 | 4.6 | 9.5 | 20.2 | 84.18 | 91.6 | 93.05 | 93.55
-5.6 | 2.9 | 4.7 | 9.7 | 23.1 | 85.025 | 91.875 | 93.22 | 93.9
-5.5 | 2.9 | 4.8333 | 10 | 27.6333 | 85.86 | 92.1333 | 93.45 | 93.98
-5.4 | 3 | 5 | 10.3 | 33.9 | 86.55 | 92.475 | 93.525 | 94.1429
-5.3 | 3.1 | 5.1 | 10.6333 | 41.725 | 87.05 | 92.675 | 93.78 | 94.275
-5.2 | 3.1 | 5.3 | 11 | 49.8 | 87.7 | 92.92 | 94.0333 | 94.6
-5.1 | 3.2 | 5.375 | 11.3 | 58.7 | 88.2333 | 93.06 | 94.3 | 94.675
-5 | 3.3 | 5.5 | 11.65 | 66.7 | 88.65 | 93.3 | 94.4 | 94.9444
-4.9 | 3.3 | 5.6 | 12.1 | 73.4 | 89.05 | 93.6 | 94.55 | 95.125
-4.8 | 3.4 | 5.8 | 12.72 | 78.9 | 89.6 | 93.9 | 94.8 | 95.2
-4.7 | 3.6 | 6 | 13.4 | 82.55 | 90 | 94.125 | 94.9667 | 95.325
-4.6 | 3.7 | 6.1 | 14.325 | 85.14 | 90.45 | 94.275 | 95.2 | 95.5375
-4.5 | 3.8 | 6.3 | 15.66 | 86.8 | 90.9 | 94.55 | 95.3 | 95.725
-4.4 | 3.9 | 6.4 | 17.7 | 87.9 | 91.3 | 94.76 | 95.5667 | 95.8778
-4.3 | 4 | 6.65 | 20.55 | 88.9 | 91.65 | 95 | 95.6667 | 96.1
-4.2 | 4.1 | 6.875 | 25.45 | 89.7 | 92.05 | 95.1667 | 95.94 | 96.175
-4.1 | 4.2 | 7.1 | 31.75 | 90.3 | 92.35 | 95.3667 | 96.1857 | 96.45
-4 | 4.3 | 7.35 | 40.1 | 90.9 | 92.75 | 95.65 | 96.32 | 96.5
-3.9 | 4.4 | 7.6 | 48.8 | 91.5 | 93.2 | 95.8 | 96.4 | 96.75
-3.8 | 4.55 | 7.9333 | 59 | 91.8 | 93.45 | 96 | 96.54 | 96.8857
-3.7 | 4.7 | 8.3 | 67.45 | 92.3 | 93.8333 | 96.175 | 96.775 | 97
-3.6 | 4.8333 | 8.6 | 75.5667 | 92.75 | 94.2 | 96.375 | 96.925 | 97.2
-3.5 | 5.0667 | 9.1 | 81.7 | 93.1667 | 94.45 | 96.64 | 97.1143 | 97.2
-3.4 | 5.2 | 9.8 | 86 | 93.58 | 94.8 | 96.76 | 97.28 | 97.4625
-3.3 | 5.45 | 10.75 | 88.7 | 94 | 95.05 | 96.9333 | 97.45 | 97.525
-3.2 | 5.7 | 12.05 | 90.4 | 94.3 | 95.3 | 97.1 | 97.5 | 97.7375
-3.1 | 5.94 | 14 | 91.6 | 94.75 | 95.625 | 97.24 | 97.725 | 97.8625
-3 | 6.3 | 17.0333 | 92.65 | 95.2 | 95.8667 | 97.325 | 97.76 | 97.9875
-2.9 | 6.74 | 22.1 | 93.6 | 95.45 | 96.3 | 97.525 | 97.9 | 98.1
-2.8 | 7.2333 | 29.2 | 94.2 | 95.7 | 96.4667 | 97.675 | 98.1667 | 98.1833
-2.7 | 8.1 | 37.875 | 94.6 | 96.05 | 96.7 | 97.9 | 98.2333 | 98.325
-2.6 | 9.16 | 48.4 | 95.1 | 96.3 | 96.9 | 98 | 98.3 | 98.5143
-2.5 | 10.7 | 59 | 95.5 | 96.6 | 97.2 | 98.1333 | 98.3667 | 98.6
-2.4 | 13.2333 | 69.4 | 95.85 | 96.9 | 97.4 | 98.2333 | 98.5 | 98.6
-2.3 | 17 | 77.95 | 96.2 | 97.2 | 97.5667 | 98.4333 | 98.6 | 98.7429
-2.2 | 23.2 | 84.4 | 96.6333 | 97.4333 | 97.8 | 98.6 | 98.7667 | 99
-2.1 | 32.3 | 89.5 | 96.9 | 97.7333 | 98 | 98.7167 | 98.76 | 99
-2 | 41.8667 | 92.35 | 97.2 | 97.9 | 98.2 | 98.8 | 99 | 99.0571
-1.9 | 52.7 | 94.1333 | 97.5 | 98 | 98.3 | 98.9 | 99.02 | 99.1125
-1.8 | 63.8 | 95.5 | 97.8 | 98.3 | 98.5 | 99.0429 | 99.15 | 99.2375
-1.7 | 74.5 | 96.35 | 98 | 98.5 | 98.7 | 99.2 | 99.2 | 99.3
-1.6 | 82.8 | 97 | 98.2 | 98.7 | 98.85 | 99.2 | 99.3 | 99.3875
-1.5 | 88.5 | 97.6 | 98.3 | 98.8 | 99 | 99.3 | 99.4 | 99.5
-1.4 | 92.4 | 98.0333 | 98.6667 | 98.9 | 99.1 | 99.46 | 99.5 | 99.5
-1.3 | 94.85 | 98.3 | 98.9 | 99.1 | 99.2 | 99.5 | 99.55 | 99.5
-1.2 | 96.24 | 98.5 | 99.05 | 99.3 | 99.4 | 99.6 | 99.65 | 99.6
-1.1 | 97.4 | 98.92 | 99.2 | 99.4 | 99.4 | 99.7 | 99.675 | 99.7
-1 | 98.04 | 99.0667 | 99.34 | 99.5 | 99.5 | 99.7 | 99.66 | 99.8
-0.9 | 98.58 | 99.3667 | 99.475 | 99.6 | 99.6 | 99.7 | 99.8 | 99.7429
-0.8 | 99 | 99.45 | 99.7 | 99.6667 | 99.7 | 99.8 | 99.8 | 99.9
-0.7 | 99.26 | 99.6 | 99.7 | 99.8 | 99.7 | 99.9 | 99.825 | 99.82
-0.6 | 99.5 | 99.625 | 99.8 | 99.9 | 99.8 | 99.9 | 100 | 99.95
-0.5 | 99.7 | 99.7 | 99.8 | 99.9 | 99.9 | 99.96 | 99.95 | 99.9
-0.4 | 99.775 | 99.85 | 99.9 | 100 | 99.9 | 100 | 99.95 | 99.9
-0.3 | 99.825 | 99.9 | 99.9 | 100 | 99.9 | 100 | 99.96 | 99.9
-0.2 | 99.9 | 99.925 | 100 | 100 | 100 | 100 | 100 | 99.8125
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9625
0.2 | 99.975 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9
0.3 | 99.8 | 99.9 | 100 | 100 | 99.9 | 99.95 | 99.92 | 99.9875
0.4 | 99.7 | 99.825 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9
0.5 | 99.6 | 99.7 | 99.8 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9429
0.6 | 99.4 | 99.7 | 99.8 | 99.8667 | 99.8 | 99.85 | 99.9833 | 99.9
0.7 | 99.2 | 99.5 | 99.7 | 99.75 | 99.8 | 99.8 | 99.8 | 99.8
0.8 | 98.92 | 99.3 | 99.6 | 99.6 | 99.75 | 99.7 | 99.8 | 99.85
0.9 | 98.52 | 99.2 | 99.4 | 99.6 | 99.7 | 99.7 | 99.8 | 99.8
1 | 98 | 98.95 | 99.3 | 99.5 | 99.6 | 99.7 | 99.625 | 99.75
1.1 | 97.4 | 98.7667 | 99.1 | 99.4 | 99.4 | 99.6333 | 99.64 | 99.7
1.2 | 96.3 | 98.5 | 98.95 | 99.2 | 99.325 | 99.6 | 99.68 | 99.6
1.3 | 94.8 | 98.2 | 98.9 | 99.1 | 99.2 | 99.5 | 99.5167 | 99.5857
1.4 | 92.2 | 97.8 | 98.7 | 98.9 | 99.0333 | 99.4 | 99.5 | 99.5
1.5 | 88.1333 | 97.45 | 98.4 | 98.725 | 98.9 | 99.3 | 99.4 | 99.4625
1.6 | 82.3667 | 96.9 | 98.2 | 98.5 | 98.7 | 99.2 | 99.3667 | 99.3375
1.7 | 74.3 | 96.2667 | 97.9333 | 98.4333 | 98.6 | 99.16 | 99.2 | 99.3
1.8 | 63.5 | 95.4 | 97.6667 | 98.3 | 98.4667 | 99 | 99.125 | 99.2
1.9 | 52.7 | 94 | 97.3 | 98 | 98.3 | 98.9 | 99 | 99.1625
2 | 41.8 | 92.1 | 97.1 | 97.8 | 98.1 | 98.7889 | 99 | 99.1
2.1 | 32.4 | 89.25 | 96.8 | 97.5333 | 97.925 | 98.6714 | 98.725 | 99
2.2 | 23.38 | 84.4 | 96.5333 | 97.3333 | 97.75 | 98.5 | 98.75 | 98.9429
2.3 | 17.1 | 78 | 96.2 | 97.1 | 97.55 | 98.3286 | 98.5143 | 98.8
2.4 | 13.0667 | 69.0333 | 95.8333 | 96.8667 | 97.4 | 98.2 | 98.52 | 98.6
2.5 | 10.65 | 59.1333 | 95.4333 | 96.5 | 97.1 | 98.08 | 98.4 | 98.5429
2.6 | 8.9667 | 48.35 | 95 | 96.2 | 96.82 | 97.96 | 98.3 | 98.5
2.7 | 7.9 | 37.975 | 94.5 | 95.94 | 96.5 | 97.825 | 98.2 | 98.35
2.8 | 7.2333 | 29 | 94 | 95.6 | 96.325 | 97.675 | 98.1333 | 98.2
2.9 | 6.75 | 22 | 93.35 | 95.3333 | 96.1 | 97.52 | 97.96 | 98.1
3 | 6.3 | 17.1 | 92.55 | 95 | 95.8 | 97.375 | 97.8 | 98.0375
3.1 | 5.9 | 13.9 | 91.6 | 94.5 | 95.6 | 97.125 | 97.7 | 97.825
3.2 | 5.65 | 11.9 | 90.3 | 94.2 | 95.25 | 97.04 | 97.5 | 97.775
3.3 | 5.45 | 10.7 | 88.4 | 93.85 | 94.9333 | 96.8 | 97.42 | 97.5571
3.4 | 5.2 | 9.8 | 85.7 | 93.45 | 94.6 | 96.66 | 97.24 | 97.4143
3.5 | 5 | 9.1 | 81.4 | 92.95 | 94.4 | 96.4667 | 97.0833 | 97.3
3.6 | 4.8 | 8.6 | 75.45 | 92.6 | 94.1 | 96.2333 | 97 | 97.15
3.7 | 4.7 | 8.3 | 67.5 | 92.2 | 93.7 | 96.075 | 96.725 | 97.075
3.8 | 4.6 | 7.9 | 58.9 | 91.85 | 93.4 | 95.9 | 96.5 | 96.9333
3.9 | 4.4333 | 7.55 | 49 | 91.3 | 93 | 95.7 | 96.4 | 96.7
4 | 4.3 | 7.35 | 40.1 | 90.65 | 92.5333 | 95.55 | 96.275 | 96.54
4.1 | 4.3 | 7.1333 | 31.5667 | 90.1 | 92.1 | 95.25 | 96.16 | 96.4
4.2 | 4.1 | 6.85 | 25 | 89.5 | 91.9 | 95.025 | 95.98 | 96.25
4.3 | 4 | 6.6 | 20.35 | 88.7 | 91.525 | 94.875 | 95.7 | 96.05
4.4 | 3.9 | 6.5 | 17.4 | 87.625 | 91.1 | 94.625 | 95.5 | 95.9333
4.5 | 3.8 | 6.3 | 15.65 | 86.4 | 90.8 | 94.4 | 95.3 | 95.725
4.6 | 3.6 | 6.1 | 14.3 | 84.7667 | 90.3 | 94.2333 | 95.1286 | 95.4875
4.7 | 3.6 | 5.9333 | 13.325 | 82.3 | 89.8 | 94 | 94.95 | 95.325
4.8 | 3.5 | 5.8 | 12.66 | 78.75 | 89.4 | 93.775 | 94.7833 | 95.2
4.9 | 3.35 | 5.6667 | 12.1 | 73.2 | 88.9333 | 93.5 | 94.5333 | 95.1125
5 | 3.3 | 5.5 | 11.55 | 66.6667 | 88.4 | 93.2 | 94.5 | 95
5.1 | 3.2 | 5.3 | 11.3 | 58.9 | 88 | 93 | 94.2714 | 94.75
5.2 | 3.1 | 5.2 | 10.9 | 50 | 87.5 | 92.8333 | 94 | 94.55
5.3 | 3.1 | 5.1 | 10.6 | 41.58 | 86.9 | 92.575 | 93.8 | 94.275
5.4 | 3 | 5 | 10.25 | 33.8333 | 86.2 | 92.25 | 93.5 | 94.2
5.5 | 3 | 4.8333 | 9.9667 | 27.55 | 85.6 | 92.075 | 93.45 | 93.9333
5.6 | 2.9 | 4.7 | 9.75 | 23.0333 | 84.8667 | 91.74 | 93.14 | 93.9
5.7 | 2.9 | 4.7 | 9.5 | 20.1333 | 83.95 | 91.4667 | 93.02 | 93.6
5.8 | 2.8 | 4.6 | 9.2 | 18.2 | 82.75 | 91.2 | 92.7 | 93.325
5.9 | 2.8 | 4.4 | 9.1 | 16.9 | 81.32 | 90.95 | 92.6 | 93.08
6 | 2.7 | 4.3 | 8.9 | 16 | 79.3 | 90.76 | 92.275 | 93
6.1 | 2.6 | 4.2 | 8.7 | 15.2 | 75.9 | 90.4 | 92.1 | 92.7625
6.2 | 2.6 | 4.1 | 8.5 | 14.6 | 71.4 | 90.2333 | 91.9 | 92.575
6.3 | 2.5 | 4 | 8.25 | 14.2 | 65.5 | 89.95 | 91.65 | 92.325
6.4 | 2.4 | 3.9 | 8.1 | 13.76 | 58.3 | 89.55 | 91.4833 | 92.1857
6.5 | 2.4 | 3.9 | 7.8667 | 13.3667 | 50.35 | 89.3 | 91.2333 | 91.9667
6.6 | 2.3 | 3.8 | 7.7 | 13 | 42.7667 | 89.025 | 90.9667 | 91.8
6.7 | 2.3 | 3.7 | 7.5 | 12.7 | 35.5 | 88.675 | 90.7667 | 91.6
6.8 | 2.2667 | 3.6 | 7.3 | 12.5 | 29.3333 | 88.4 | 90.6 | 91.3714
6.9 | 2.2 | 3.6 | 7.2 | 12.1 | 24.9 | 88.0857 | 90.3333 | 91.1333
7 |  | 3.5 | 7 | 11.9 | 21.8667 | 87.85 | 90.0333 | 90.9429
7.1 |  | 3.475 | 6.9 | 11.6 | 20 | 87.5 | 89.9 | 90.7444
7.2 |  | 3.4 | 6.8 | 11.32 | 18.62 | 87.2333 | 89.6 | 90.5
7.3 |  | 3.3 | 6.6 | 11.1 | 17.675 | 86.9333 | 89.325 | 90.35
7.4 |  | 3.3 | 6.5 | 10.8333 | 16.95 | 86.6 | 89.1 | 90.05
7.5 |  | 3.2 | 6.35 | 10.6 | 16.4 | 86.3 | 88.88 | 89.825
7.6 |  |  | 6.2 | 10.4 | 15.9 | 86.05 | 88.6833 | 89.8
7.7 |  |  | 6.1 | 10.2 | 15.4333 | 85.725 | 88.55 | 89.575
7.8 |  |  | 5.95 | 9.9 | 15 | 85.3333 | 88.24 | 89.2875
7.9 |  |  | 5.9 | 9.72 | 14.7 | 85.06 | 87.8 | 89.0375
8 |  |  | 5.8 | 9.6 | 14.425 | 84.76 | 87.7333 | 88.9222
8.1 |  |  | 5.7 | 9.325 | 14.1 | 84.4 | 87.4 | 88.7
8.2 |  |  | 5.5 | 9.15 | 13.7333 | 84.1 | 87.2667 | 88.35
8.3 |  |  | 5.4 | 9 | 13.5 | 83.86 | 87 | 88.15
8.4 |  |  | 5.3 | 8.8 | 13.2 | 83.45 | 86.72 | 87.95
8.5 |  |  | 5.2 | 8.6 | 12.9 | 83.125 | 86.5 | 87.7
8.6 |  |  | 5.1 | 8.4 | 12.65 | 82.8 | 86.1833 | 87.55
8.7 |  |  | 5 | 8.25 | 12.45 | 82.4 | 85.9333 | 87.32
8.8 |  |  | 5 | 8.1333 | 12.2 | 82.04 | 85.75 | 86.975
8.9 |  |  |  | 7.9667 | 11.9 | 81.7 | 85.4833 | 86.7625
9 |  |  |  | 7.8 | 11.7 | 81.275 | 85.3167 | 86.575
9.1 |  |  |  | 7.7 | 11.5 | 81.1 | 85.02 | 86.3444
9.2 |  |  |  | 7.5 | 11.2333 | 80.7 | 84.74 | 86.2
9.3 |  |  |  | 7.4 | 11 | 80.42 | 84.5 | 86.0286
9.4 |  |  |  | 7.3 | 10.85 | 80 | 84.2333 | 85.76
9.5 |  |  |  | 7.1 | 10.62 | 79.6125 | 84 | 85.4857
9.6 |  |  |  | 7 | 10.4 | 79.18 | 83.8 | 85.32
9.7 |  |  |  | 6.9 | 10.2 | 78.7667 | 83.55 | 85.1857
9.8 |  |  |  | 6.7 | 10 | 78.4 | 83.36 | 84.9
9.9 |  |  |  | 6.6 | 9.8667 | 78.1 | 82.95 | 84.7143
10 |  |  |  | 6.5 | 9.7 | 77.75 | 82.625 | 84.4286
10.1 |  |  |  | 6.4 | 9.5 | 77.2667 | 82.4 | 84.18
10.2 |  |  |  |  | 9.3 | 76.9 | 82.22 | 83.9286
10.3 |  |  |  |  | 9.2 | 76.475 | 81.96 | 83.8143
10.4 |  |  |  |  | 9 | 76.15 | 81.7 | 83.4429
10.5 |  |  |  |  | 8.9 | 75.9 | 81.4 | 83.3429
10.6 |  |  |  |  | 8.7 | 75.4 | 81.1667 | 83.0571
10.7 |  |  |  |  | 8.6 | 74.98 | 80.9 | 82.86
10.8 |  |  |  |  | 8.4333 | 74.5667 | 80.5667 | 82.7625
10.9 |  |  |  |  | 8.3 | 74.15 | 80.3333 | 82.44
11 |  |  |  |  | 8.1 | 73.65 | 80.1 | 82.3
11.1 |  |  |  |  | 8 | 73.3 | 79.74 | 82.0714
11.2 |  |  |  |  | 7.9 | 72.78 | 79.52 | 81.675
11.3 |  |  |  |  | 7.7 | 72.4 | 79.28 | 81.475
11.4 |  |  |  |  | 7.6 | 71.9 | 78.92 | 81.2857
11.5 | [5 empty] | 71.5 | 78.8 | 81.2
11.6 | [5 empty] | 71 | 78.54 | 80.8
11.7 | [5 empty] | 70.5 | 78.22 | 80.6875
11.8 | [5 empty] | 70 | 77.95 | 80.375
11.9 | [5 empty] | 69.45 | 77.6857 | 80.2125
12 | [5 empty] | 68.875 | 77.36 | 80
12.1 | [5 empty] | 68.25 | 77.1 | 79.8
12.2 | [5 empty] | 67.5 | 76.8 | 79.5
12.3 | [5 empty] | 66.6 | 76.5857 | 79.34
12.4 | [5 empty] | 65.575 | 76.35 | 79.1333
12.5 | [5 empty] | 63.9714 | 75.9333 | 78.8667
12.6 | [5 empty] | 61.8167 | 75.7 | 78.6167
12.7 | [5 empty] | 58.8 | 75.4667 | 78.3714
12.8 | [5 empty] | 54.65 | 75.1333 | 78.1
12.9 | [5 empty] | 49.6667 | 74.8571 | 77.8667
13 | [5 empty] | 44.38 | 74.6 | 77.6833
13.1 | [5 empty] | 38.8 | 74.3667 | 77.475
13.2 | [5 empty] | 33.8 | 74.22 | 77.225
13.3 | [5 empty] | 29.45 | 73.8 | 76.9667
13.4 | [5 empty] | 26 | 73.4 | 76.7167
13.5 | [5 empty] | 23.8 | 73.1 | 76.5143
13.6 | [5 empty] | 22.225 | 72.88 | 76.325
13.7 | [5 empty] | 21.2 | 72.6 | 76.1
13.8 | [5 empty] | 20.375 | 72.3333 | 75.7857
13.9 | [5 empty] | 19.8 | 72.15 | 75.6
14 | [5 empty] | 19.2667 | 71.7667 | 75.3333
14.1 | [5 empty] | 18.8 | 71.4333 | 75.1333
14.2 | [5 empty] | 18.45 | 71.05 | 74.9143
14.3 | [5 empty] | 18.1 | 70.8667 | 74.7333
14.4 | [5 empty] | 17.8333 | 70.5571 | 74.5
14.5 | [5 empty] | 17.4714 | 70.3143 | 74.15
14.6 | [5 empty] | 17.1667 | 70.1 | 73.9667
14.7 | [5 empty] | 16.88 | 69.7333 | 73.675
14.8 | [5 empty] | 16.575 | 69.48 | 73.4875
14.9 | [5 empty] | 16.3 | 69.14 | 73.35
15 | [5 empty] | 16 | 68.7 | 73.1
15.1 | [5 empty] | 15.775 | 68.58 | 72.85
15.2 | [5 empty] | 15.52 | 68.3 | 72.5857
15.3 | [5 empty] | 15.275 | 67.9 | 72.4
15.4 | [5 empty] | 15.1 | 67.5667 | 72.1
15.5 | [5 empty] | 14.8333 | 67.3286 | 71.9833
15.6 | [5 empty] | 14.575 | 66.9667 | 71.6
15.7 | [5 empty] | 14.34 | 66.7 | 71.4571
15.8 | [5 empty] | 14.18 | 66.44 | 71.3
15.9 | [5 empty] | 13.95 | 66.1 | 71.0429
16 | [5 empty] | 13.7 | 65.8333 | 70.8
16.1 | [5 empty] | 13.55 | 65.46 | 70.6333
16.2 | [5 empty] | 13.3143 | 65.16 | 70.36
16.3 | [5 empty] | 13.1 | 64.8 | 70.0625
16.4 | [5 empty] | 12.92 | 64.5 | 69.8
16.5 | [5 empty] | 12.76 | 64.2 | 69.675
16.6 | [5 empty] | 12.5333 | 63.8 | 69.4667
16.7 | [5 empty] | 12.4 | 63.54 | 69.1667
16.8 | [5 empty] | 12.24 | 63.1667 | 68.9286
16.9 | [5 empty] | 12.02 | 62.85 | 68.7429
17 | [5 empty] | 11.86 | 62.48 | 68.4
17.1 | [5 empty] | 11.7 | 62.15 | 68.125
17.2 | [5 empty] | 11.55 | 61.82 | 68
17.3 | [5 empty] | 11.3 | 61.5 | 67.7375
17.4 | [5 empty] | 11.2 | 61.14 | 67.4889
17.5 | [5 empty] | 11 | 60.82 | 67.3429
17.6 | [5 empty] | 10.82 | 60.4 | 67.0333
17.7 | [5 empty] | 10.78 | 60.1 | 66.85
17.8 | [5 empty] | 10.5833 | 59.7 | 66.6857
17.9 | [5 empty] | 10.4 | 59.32 | 66.45
18 | [6 empty] | 58.9 | 66.225
18.1 | [6 empty] | 58.52 | 65.9
18.2 | [6 empty] | 58.15 | 65.65
18.3 | [6 empty] | 57.78 | 65.4333
18.4 | [6 empty] | 57.24 | 65.1875
18.5 | [6 empty] | 56.8 | 65.025
18.6 | [6 empty] | 56.2667 | 64.7333
18.7 | [6 empty] | 55.6 | 64.55
18.8 | [6 empty] | 54.96 | 64.2444
18.9 | [6 empty] | 54 | 64
19 | [6 empty] | 52.72 | 63.8429
19.1 | [6 empty] | 50.925 | 63.5333
19.2 | [6 empty] | 48.52 | 63.32
19.3 | [6 empty] | 45.1 | 63.0714
19.4 | [6 empty] | 41.24 | 62.875
19.5 | [6 empty] | 37.22 | 62.7
19.6 | [6 empty] | 33.06 | 62.3286
19.7 | [6 empty] | 29.4 | 62.1333
19.8 | [6 empty] | 26.28 | 61.9667
19.9 | [6 empty] | 23.8 | 61.6571
20 | [6 empty] | 22.1 | 61.45
20.1 | [6 empty] | 21 | 61.1667
20.2 | [6 empty] | 20.25 | 60.9286
20.3 | [6 empty] | 19.75 | 60.675
20.4 | [6 empty] | 19.3 | 60.45
20.5 | [6 empty] | 18.86 | 60.15
20.6 | [6 empty] | 18.56 | 60
20.7 | [6 empty] | 18.18 | 59.7
20.8 | [6 empty] | 17.92 | 59.4286
20.9 | [6 empty] | 17.6 | 59.1667
21 | [6 empty] | 17.3 | 58.98
21.1 | [6 empty] | 17.12 | 58.8
21.2 | [6 empty] | 16.92 | 58.4625
21.3 | [6 empty] | 16.7 | 58.2125
21.4 | [6 empty] | 16.4 | 58
21.5 | [6 empty] | 16.2 | 57.725
21.6 | [6 empty] | 16 | 57.4889
21.7 | [6 empty] | 15.75 | 57.2
21.8 | [6 empty] | 15.5667 | 56.95
21.9 | [6 empty] | 15.3167 | 56.6857
22 | [6 empty] | 15.125 | 56.4714
22.1 | [6 empty] | 14.975 | 56.1667
22.2 | [6 empty] | 14.76 | 55.92
22.3 | [6 empty] | 14.5667 | 55.6714
22.4 | [6 empty] | 14.4 | 55.4333
22.5 | [6 empty] | 14.1667 | 55.125
22.6 | [6 empty] | 14.05 | 54.8333
22.7 | [6 empty] | 13.7667 | 54.5714
22.8 | [6 empty] | 13.65 | 54.3571
22.9 | [6 empty] | 13.44 | 54.0667
23 | [6 empty] | 13.275 | 53.7429
23.1 | [6 empty] | 13.15 | 53.52
23.2 | [6 empty] | 12.9833 | 53.2429
23.3 | [6 empty] | 12.8167 | 53.0167
23.4 | [6 empty] | 12.65 | 52.72
23.5 | [6 empty] | 12.525 | 52.4
23.6 | [6 empty] | 12.4 | 52.1667
23.7 | [6 empty] | 12.2 | 51.85
23.8 | [6 empty] | 12.02 | 51.6
23.9 | [6 empty] | 11.9 | 51.3143
24 | [6 empty] | 11.78 | 51
24.1 | [6 empty] | 11.7 | 50.7
24.2 | [6 empty] | 11.6 | 50.4143
24.3 | [6 empty] | 11.4 | 50.1
24.4 | [6 empty] | 11.24 | 49.74
24.5 | [7 empty] | 49.58
24.6 | [7 empty] | 49.22
24.7 | [7 empty] | 48.86
24.8 | [7 empty] | 48.5
24.9 | [7 empty] | 48.0667
25 | [7 empty] | 47.6625
25.1 | [7 empty] | 47.25
25.2 | [7 empty] | 46.7333
25.3 | [7 empty] | 46.15
25.4 | [7 empty] | 45.4
25.5 | [7 empty] | 44.1714
25.6 | [7 empty] | 42.7667
25.7 | [7 empty] | 40.4286
25.8 | [7 empty] | 37.6
25.9 | [7 empty] | 34.8
26 | [7 empty] | 31.35
26.1 | [7 empty] | 27.9167
26.2 | [7 empty] | 25.1429
26.3 | [7 empty] | 22.7667
26.4 | [7 empty] | 21
26.5 | [7 empty] | 19.8
26.6 | [7 empty] | 19
26.7 | [7 empty] | 18.45
26.8 | [7 empty] | 17.9875
26.9 | [7 empty] | 17.6333
27 | [7 empty] | 17.35
27.1 | [7 empty] | 17
27.2 | [7 empty] | 16.75
27.3 | [7 empty] | 16.4857
27.4 | [7 empty] | 16.3
27.5 | [7 empty] | 16.05
27.6 | [7 empty] | 15.85
27.7 | [7 empty] | 15.6667
27.8 | [7 empty] | 15.425
27.9 | [7 empty] | 15.275
28 | [7 empty] | 15.0625
28.1 | [7 empty] | 14.875
28.2 | [7 empty] | 14.7125
28.3 | [7 empty] | 14.4889
28.4 | [7 empty] | 14.35
28.5 | [7 empty] | 14.15
28.6 | [7 empty] | 14
28.7 | [7 empty] | 13.8429
28.8 | [7 empty] | 13.6375
28.9 | [7 empty] | 13.5125
29 | [7 empty] | 13.375
29.1 | [7 empty] | 13.1625
29.2 | [7 empty] | 13.0286
29.3 | [7 empty] | 12.8875
29.4 | [7 empty] | 12.725
29.5 | [7 empty] | 12.6
29.6 | [7 empty] | 12.4667
29.7 | [7 empty] | 12.325
29.8 | [7 empty] | 12.2
29.9 | [7 empty] | 12
30 | [7 empty] | 11.95
30.1 | [7 empty] | 11.825
30.2 | [7 empty] | 11.6875
30.3 | [7 empty] | 11.5625
30.4 | [7 empty] | 11.375
30.5 | [7 empty] | 11.2875
30.6 | [7 empty] | 11.15
30.7 | [7 empty] | 11.0833
30.8 | [7 empty] | 11
30.9 | [7 empty] | 10.8833

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:1.5 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-43.7 |  |  |  |  | 4.9
-43.6 |  |  |  |  | 4.9288
-43.5 |  |  |  |  | 5
-43.4 |  |  |  |  | 5.0308
-43.3 |  |  |  |  | 5.1
-43.2 |  |  |  |  | 5.1
-43.1 |  |  |  |  | 5.1
-43 |  |  |  |  | 5.1492
-42.9 |  |  |  |  | 5.2
-42.8 |  |  |  |  | 5.2
-42.7 |  |  |  |  | 5.2665
-42.6 |  |  |  |  | 5.3
-42.5 |  |  |  |  | 5.3114
-42.4 |  |  |  |  | 5.4
-42.3 |  |  |  |  | 5.4117
-42.2 |  |  |  |  | 5.5
-42.1 |  |  |  |  | 5.5
-42 |  |  |  |  | 5.5358
-41.9 |  |  |  |  | 5.6
-41.8 |  |  |  |  | 5.6042
-41.7 |  |  |  |  | 5.7
-41.6 |  |  |  |  | 5.7
-41.5 |  |  |  |  | 5.7929
-41.4 |  |  |  |  | 5.8
-41.3 |  |  |  |  | 5.8
-41.2 |  |  |  |  | 5.8545
-41.1 |  |  |  |  | 5.9
-41 |  |  |  |  | 5.9837
-40.9 |  |  |  | 3.4617 | 6
-40.8 |  |  |  | 3.5 | 6.0002
-40.7 |  |  |  | 3.5 | 6.1
-40.6 |  |  |  | 3.5 | 6.1025
-40.5 |  |  |  | 3.5603 | 6.2
-40.4 |  |  |  | 3.6 | 6.2
-40.3 |  |  |  | 3.6 | 6.2727
-40.2 |  |  |  | 3.6554 | 6.3072
-40.1 |  |  |  | 3.7 | 6.4
-40 |  |  |  | 3.7 | 6.4
-39.9 |  |  |  | 3.7949 | 6.5
-39.8 |  |  |  | 3.8 | 6.5343
-39.7 |  |  |  | 3.8 | 6.6
-39.6 |  |  |  | 3.8836 | 6.6809
-39.5 |  |  |  | 3.9 | 6.7
-39.4 |  |  |  | 3.9 | 6.8
-39.3 |  |  |  | 3.9345 | 6.8178
-39.2 |  |  |  | 4 | 6.9
-39.1 |  |  |  | 4 | 6.9443
-39 |  |  |  | 4.0547 | 7
-38.9 |  |  |  | 4.1 | 7.1
-38.8 |  |  |  | 4.1 | 7.1485
-38.7 |  |  |  | 4.15 | 7.2
-38.6 |  |  |  | 4.2 | 7.2938
-38.5 |  |  |  | 4.2 | 7.3091
-38.4 |  |  |  | 4.253 | 7.4101
-38.3 |  |  |  | 4.3 | 7.5
-38.2 |  |  |  | 4.3481 | 7.5371
-38.1 |  |  | 2.2 | 4.4 | 7.6
-38 |  |  | 2.2 | 4.4427 | 7.6571
-37.9 |  |  | 2.2723 | 4.5 | 7.7677
-37.8 |  |  | 2.3 | 4.5 | 7.8733
-37.7 |  |  | 2.3 | 4.5 | 7.9743
-37.6 |  |  | 2.3 | 4.6039 | 8.0754
-37.5 |  |  | 2.3049 | 4.7 | 8.1
-37.4 |  |  | 2.4 | 4.7 | 8.1917
-37.3 |  |  | 2.4 | 4.7071 | 8.3
-37.2 |  |  | 2.4 | 4.8 | 8.3366
-37.1 |  |  | 2.4 | 4.8233 | 8.4376
-37 |  |  | 2.4 | 4.9 | 8.5387
-36.9 |  |  | 2.4755 | 4.9254 | 8.6397
-36.8 |  |  | 2.5 | 5 | 8.7407
-36.7 |  | 1.6983 | 2.5 | 5.024 | 8.8417
-36.6 |  | 1.6983 | 2.5 | 5.1 | 8.9
-36.5 |  | 1.6983 | 2.5009 | 5.1294 | 8.9588
-36.4 |  | 1.6983 | 2.6 | 5.2 | 9.1304
-36.3 |  | 1.6983 | 2.6 | 5.2314 | 9.2314
-36.2 |  | 1.6983 | 2.6 | 5.3 | 9.3651
-36.1 |  | 1.7603 | 2.6 | 5.3469 | 9.5
-36 |  | 1.7982 | 2.7 | 5.4 | 9.5631
-35.9 |  | 1.7982 | 2.7 | 5.5 | 9.6641
-35.8 |  | 1.7982 | 2.7 | 5.564 | 9.7652
-35.7 | 1.4 | 1.7982 | 2.7 | 5.6105 | 9.9323
-35.6 | 1.4 | 1.7982 | 2.7617 | 5.7 | 10.0587
-35.5 | 1.4 | 1.8092 | 2.8 | 5.7629 | 10.1538
-35.4 | 1.4 | 1.8981 | 2.8 | 5.8 | 10.3097
-35.3 | 1.4 | 1.8981 | 2.8782 | 5.8614 | 10.4559
-35.2 | 1.4498 | 1.8981 | 2.9 | 5.9569 | 10.5664
-35.1 | 1.5 | 1.8981 | 2.9 | 6.058 | 10.6722
-35 | 1.5 | 1.8981 | 2.9266 | 6.1 | 10.7732
-34.9 | 1.5 | 1.922 | 3 | 6.1839 | 10.9485
-34.8 | 1.5 | 1.998 | 3 | 6.2896 | 11.0753
-34.7 | 1.5 | 1.998 | 3.0868 | 6.3267 | 11.189
-34.6 | 1.5 | 1.998 | 3.1 | 6.4 | 11.3832
-34.5 | 1.5 | 1.998 | 3.1 | 6.5069 | 11.5853
-34.4 | 1.5 | 2.063 | 3.171 | 6.6 | 11.6819
-34.3 | 1.577 | 2.0979 | 3.2 | 6.6525 | 11.9124
-34.2 | 1.6 | 2.0979 | 3.2283 | 6.7671 | 12.1078
-34.1 | 1.6 | 2.0979 | 3.3 | 6.9 | 12.2953
-34 | 1.6 | 2.0979 | 3.3 | 6.9119 | 12.4358
-33.9 | 1.6 | 2.0979 | 3.3 | 7.0129 | 12.7388
-33.8 | 1.6 | 2.1978 | 3.4 | 7.114 | 13.0558
-33.7 | 1.6176 | 2.1978 | 3.4 | 7.2176 | 13.4602
-33.6 | 1.7 | 2.1978 | 3.5 | 7.3303 | 13.8803
-33.5 | 1.7 | 2.201 | 3.5 | 7.4313 | 14.4024
-33.4 | 1.7 | 2.2977 | 3.571 | 7.5323 | 14.9949
-33.3 | 1.7 | 2.2977 | 3.6 | 7.6667 | 15.5427
-33.2 | 1.7 | 2.2977 | 3.6234 | 7.8401 | 16.4633
-33.1 | 1.7 | 2.2977 | 3.7 | 7.958 | 17.5812
-33 | 1.7 | 2.3626 | 3.7221 | 8.0649 | 18.8089
-32.9 | 1.8 | 2.3976 | 3.8 | 8.177 | 20.2672
-32.8 | 1.8 | 2.3976 | 3.8 | 8.3115 | 21.9598
-32.7 | 1.8 | 2.3976 | 3.8952 | 8.4352 | 23.8641
-32.6 | 1.8 | 2.4665 | 3.9167 | 8.5767 | 25.95
-32.5 | 1.8 | 2.4975 | 4 | 8.6738 | 27.849
-32.4 | 1.8329 | 2.4975 | 4.0994 | 8.842 | 29.8912
-32.3 | 1.9 | 2.5838 | 4.1 | 8.9841 | 31.7977
-32.2 | 1.9 | 2.5974 | 4.1823 | 9.0874 | 33.397
-32.1 | 1.9 | 2.5974 | 4.2238 | 9.2768 | 34.7638
-32 | 1.9 | 2.6154 | 4.3 | 9.4075 | 35.8515
-31.9 | 1.9 | 2.6973 | 4.3866 | 9.6055 | 36.6331
-31.8 | 1.99 | 2.6973 | 4.4233 | 9.7467 | 37.1087
-31.7 | 2 | 2.7218 | 4.5 | 9.9494 | 37.5647
-31.6 | 2 | 2.7972 | 4.5424 | 10.1514 | 37.9737
-31.5 | 2 | 2.7972 | 4.6 | 10.3604 | 38.2307
-31.4 | 2.0337 | 2.8355 | 4.6936 | 10.6436 | 38.5561
-31.3 | 2.1 | 2.8971 | 4.8 | 10.9252 | 38.8655
-31.2 | 2.1 | 2.8971 | 4.8457 | 11.2095 | 39.1609
-31.1 | 2.1 | 2.9527 | 4.9816 | 11.6783 | 39.418
-31 | 2.1495 | 2.997 | 5.0162 | 12.1646 | 39.6188
-30.9 | 2.2 | 2.9978 | 5.134 | 12.8041 | 39.8416
-30.8 | 2.2 | 3.0969 | 5.2518 | 13.5961 | 40.1245
-30.7 | 2.2 | 3.1423 | 5.3 | 14.5294 | 40.3074
-30.6 | 2.2 | 3.1968 | 5.425 | 15.7189 | 40.5751
-30.5 | 2.3 | 3.2586 | 5.5554 | 17.1038 | 40.8108
-30.4 | 2.3 | 3.2967 | 5.6628 | 18.8057 | 41.0465
-30.3 | 2.3 | 3.2967 | 5.7894 | 20.691 | 41.2187
-30.2 | 2.3908 | 3.3472 | 5.8307 | 22.762 | 41.4846
-30.1 | 2.4 | 3.4481 | 5.9721 | 25.2259 | 41.6721
-30 | 2.4 | 3.4965 | 6.1136 | 27.8424 | 41.9339
-29.9 | 2.4 | 3.5714 | 6.255 | 30.3699 | 42.1625
-29.8 | 2.4971 | 3.5964 | 6.3803 | 32.6528 | 42.3929
-29.7 | 2.5 | 3.6945 | 6.4178 | 34.5895 | 42.5446
-29.6 | 2.5 | 3.7203 | 6.5988 | 36.0921 | 42.8655
-29.5 | 2.6 | 3.7962 | 6.7808 | 37.335 | 43.0607
-29.4 | 2.6 | 3.8976 | 6.9021 | 38.0343 | 43.303
-29.3 | 2.6436 | 3.996 | 7.0436 | 38.6316 | 43.5051
-29.2 | 2.7 | 4.0009 | 7.2416 | 39.1141 | 43.7093
-29.1 | 2.7 | 4.1422 | 7.3887 | 39.5136 | 43.9745
-29 | 2.8 | 4.2689 | 7.6097 | 39.813 | 44.1826
-28.9 | 2.8 | 4.3049 | 7.8731 | 40.0975 | 44.3846
-28.8 | 2.8307 | 4.4462 | 8.1227 | 40.4723 | 44.68
-28.7 | 2.9 | 4.4955 | 8.4361 | 40.7115 | 44.82
-28.6 | 3 | 4.6088 | 8.8604 | 40.9892 | 45.0479
-28.5 | 3 | 4.741 | 9.3956 | 41.2606 | 45.2748
-28.4 | 3.1 | 4.8715 | 9.9988 | 41.5272 | 45.5519
-28.3 | 3.1481 | 4.9129 | 10.8187 | 41.8094 | 45.7629
-28.2 | 3.2 | 5.0541 | 11.8263 | 42.0416 | 45.9591
-28.1 | 3.3 | 5.1667 | 13.1561 | 42.3009 | 46.2005
-28 | 3.362 | 5.2972 | 14.8247 | 42.5315 | 46.3551
-27.9 | 3.4 | 5.4399 | 16.6543 | 42.7503 | 46.5596
-27.8 | 3.5 | 5.6025 | 19.0519 | 43.0356 | 46.7818
-27.7 | 3.5386 | 5.838 | 21.8869 | 43.2438 | 47.0232
-27.6 | 3.6678 | 6.0286 | 24.9741 | 43.4681 | 47.2096
-27.5 | 3.7 | 6.2846 | 28.5702 | 43.6818 | 47.3988
-27.4 | 3.7306 | 6.6519 | 31.7282 | 43.9008 | 47.6008
-27.3 | 3.86 | 7.0556 | 34.6388 | 44.1542 | 47.8028
-27.2 | 3.9934 | 7.5816 | 37.0009 | 44.4048 | 48.0377
-27.1 | 4.029 | 8.3478 | 38.8484 | 44.6869 | 48.2457
-27 | 4.1469 | 9.321 | 40.1861 | 44.9375 | 48.4271
-26.9 | 4.2777 | 10.5819 | 40.988 | 45.1396 | 48.6628
-26.8 | 4.4239 | 12.0535 | 41.5636 | 45.3416 | 48.7992
-26.7 | 4.6006 | 14.0659 | 42.0519 | 45.6155 | 49.1009
-26.6 | 4.8256 | 16.5769 | 42.4548 | 45.8274 | 49.3438
-26.5 | 5.1666 | 19.7907 | 42.7642 | 46.0389 | 49.5634
-26.4 | 5.5825 | 23.3712 | 43.0495 | 46.308 | 49.7497
-26.3 | 6.1311 | 27.3828 | 43.3324 | 46.5187 | 49.9517
-26.2 | 6.9223 | 31.3665 | 43.7076 | 46.6552 | 50.1537
-26.1 | 7.861 | 34.9907 | 43.9226 | 46.9479 | 50.3363
-26 | 9.2748 | 37.8337 | 44.2175 | 47.2006 | 50.5939
-25.9 | 11.1177 | 40.2042 | 44.4531 | 47.4026 | 50.8313
-25.8 | 13.4548 | 41.8909 | 44.6266 | 47.6046 | 51.0333
-25.7 | 16.6552 | 42.8943 | 44.9367 | 47.86 | 51.2578
-25.6 | 20.246 | 43.6372 | 45.2522 | 48.1269 | 51.4659
-25.5 | 24.5028 | 44.0686 | 45.496 | 48.2697 | 51.6469
-25.4 | 28.6876 | 44.4922 | 45.7316 | 48.4414 | 51.8414
-25.3 | 33.6106 | 44.813 | 45.9405 | 48.6883 | 52.0673
-25.2 | 37.1085 | 45.1157 | 46.1636 | 48.917 | 52.3182
-25.1 | 40.3955 | 45.36 | 46.376 | 49.1664 | 52.4433
-25 | 42.8994 | 45.6071 | 46.6067 | 49.3352 | 52.6411
-24.9 | 44.3729 | 45.9098 | 46.8101 | 49.5521 | 52.8101
-24.8 | 45.203 | 46.193 | 47.0458 | 49.7675 | 53.0549
-24.7 | 45.7559 | 46.4018 | 47.3048 | 49.9816 | 53.2574
-24.6 | 46.0607 | 46.6252 | 47.5719 | 50.2172 | 53.3753
-24.5 | 46.3434 | 46.8179 | 47.7739 | 50.453 | 53.6235
-24.4 | 46.6132 | 47.0194 | 47.976 | 50.6532 | 53.8887
-24.3 | 46.8364 | 47.2771 | 48.2024 | 50.8893 | 54.1067
-24.2 | 47.0451 | 47.4485 | 48.4515 | 51.112 | 54.3087
-24.1 | 47.2348 | 47.7343 | 48.6535 | 51.3392 | 54.5661
-24 | 47.4588 | 47.9012 | 48.8556 | 51.5412 | 54.7564
-23.9 | 47.7004 | 48.1239 | 49.0576 | 51.7006 | 54.9147
-23.8 | 47.8617 | 48.3684 | 49.2695 | 51.9522 | 55.1168
-23.7 | 48.029 | 48.5702 | 49.4902 | 52.1331 | 55.3188
-23.6 | 48.2045 | 48.7358 | 49.6923 | 52.3615 | 55.5409
-23.5 | 48.3919 | 48.931 | 49.9319 | 52.6383 | 55.7767
-23.4 | 48.5874 | 49.1329 | 50.1105 | 52.8124 | 55.891
-23.3 | 48.811 | 49.3348 | 50.3175 | 53.0269 | 56.1761
-23.2 | 48.9703 | 49.5649 | 50.5419 | 53.1902 | 56.3861
-23.1 | 49.1116 | 49.7599 | 50.7434 | 53.4634 | 56.6264
-23 | 49.3327 | 49.8974 | 50.9885 | 53.6522 | 56.7884
-22.9 | 49.5182 | 50.1074 | 51.1921 | 53.8397 | 56.9621
-22.8 | 49.6949 | 50.3614 | 51.3941 | 54.0799 | 57.1599
-22.7 | 49.8717 | 50.5306 | 51.6248 | 54.282 | 57.4174
-22.6 | 50.0979 | 50.7627 | 51.8269 | 54.484 | 57.6261
-22.5 | 50.2502 | 50.8896 | 52.0144 | 54.7004 | 57.8603
-22.4 | 50.3816 | 51.1516 | 52.1464 | 54.9166 | 58.0453
-22.3 | 50.6192 | 51.318 | 52.4329 | 55.1186 | 58.2473
-22.2 | 50.7741 | 51.5834 | 52.6349 | 55.3206 | 58.5239
-22.1 | 51.0098 | 51.7668 | 52.837 | 55.561 | 58.6757
-22 | 51.1873 | 51.8823 | 53.0455 | 55.7962 | 58.8533
-21.9 | 51.3146 | 52.0746 | 53.2406 | 55.9982 | 59.0812
-21.8 | 51.5251 | 52.298 | 53.4168 | 56.1702 | 59.2859
-21.7 | 51.6516 | 52.5571 | 53.6526 | 56.3193 | 59.5217
-21.6 | 51.855 | 52.6944 | 53.9108 | 56.5471 | 59.7867
-21.5 | 51.9953 | 52.8962 | 54.0344 | 56.786 | 59.9919
-21.4 | 52.2358 | 53.0917 | 54.2516 | 57.0465 | 60.0971
-21.3 | 52.4288 | 53.2591 | 54.4287 | 57.1966 | 60.337
-21.2 | 52.5822 | 53.3733 | 54.6386 | 57.3978 | 60.6267
-21.1 | 52.8201 | 53.5752 | 54.8335 | 57.5801 | 60.8167
-21 | 52.9769 | 53.7821 | 55.0903 | 57.9215 | 60.9297
-20.9 | 53.203 | 53.9819 | 55.2829 | 58.0525 | 61.086
-20.8 | 53.3444 | 54.1532 | 55.4554 | 58.311 | 61.411
-20.7 | 53.5715 | 54.3886 | 55.7267 | 58.5097 | 61.5756
-20.6 | 53.7272 | 54.5558 | 55.8766 | 58.7226 | 61.8682
-20.5 | 53.8686 | 54.7576 | 56.1327 | 58.8973 | 62.0144
-20.4 | 54.13 | 54.9594 | 56.3143 | 59.1501 | 62.275
-20.3 | 54.2714 | 55.1592 | 56.5164 | 59.3393 | 62.5368
-20.2 | 54.3941 | 55.3359 | 56.7215 | 59.5327 | 62.6441
-20.1 | 54.6343 | 55.5364 | 56.9343 | 59.7347 | 62.7743
-20 | 54.8513 | 55.7382 | 57.1513 | 59.9595 | 63.0479
-19.9 | 54.9836 | 55.9401 | 57.3143 | 60.1953 | 63.1976
-19.8 | 55.1643 | 56.2408 | 57.4772 | 60.3772 | 63.4265
-19.7 | 55.36 | 56.3686 | 57.6801 | 60.6143 | 63.6399
-19.6 | 55.5691 | 56.5455 | 58.0214 | 60.8629 | 63.8614
-19.5 | 55.8285 | 56.7473 | 58.1524 | 61.0326 | 64.1572
-19.4 | 55.9053 | 56.9502 | 58.3404 | 61.2203 | 64.2888
-19.3 | 56.1642 | 57.1686 | 58.6094 | 61.4141 | 64.5694
-19.2 | 56.3471 | 57.3098 | 58.8226 | 61.6959 | 64.6794
-19.1 | 56.4885 | 57.5566 | 58.9809 | 61.9143 | 64.9324
-19 | 56.6875 | 57.7672 | 59.2 | 62.1799 | 65.157
-18.9 | 56.9027 | 57.9448 | 59.402 | 62.3448 | 65.3856
-18.8 | 57.0928 | 58.1193 | 59.604 | 62.5469 | 65.6845
-18.7 | 57.2903 | 58.3558 | 59.7743 | 62.757 | 65.7959
-18.6 | 57.4757 | 58.5342 | 59.9927 | 63.0112 | 66.0157
-18.5 | 57.6741 | 58.7514 | 60.2371 | 63.2213 | 66.2427
-18.4 | 57.8654 | 58.9532 | 60.457 | 63.4121 | 66.4463
-18.3 | 58.0665 | 59.155 | 60.6332 | 63.6598 | 66.7333
-18.2 | 58.2689 | 59.322 | 60.859 | 63.7867 | 66.8511
-18.1 | 58.4627 | 59.5424 | 61.0427 | 64.063 | 67.0655
-18 | 58.6736 | 59.6472 | 61.2403 | 64.3059 | 67.3403
-17.9 | 58.8512 | 59.8914 | 61.5092 | 64.5046 | 67.5455
-17.8 | 59.1117 | 60.107 | 61.7386 | 64.6579 | 67.687
-17.7 | 59.2737 | 60.3634 | 61.9406 | 64.9406 | 67.8834
-17.6 | 59.483 | 60.5018 | 62.1298 | 65.1426 | 68.1165
-17.5 | 59.7261 | 60.7619 | 62.3187 | 65.3446 | 68.3625
-17.4 | 59.8439 | 60.9917 | 62.5323 | 65.5467 | 68.6038
-17.3 | 60.0482 | 61.133 | 62.7568 | 65.7568 | 68.8059
-17.2 | 60.2925 | 61.3542 | 62.9651 | 66.0112 | 69.0618
-17.1 | 60.5141 | 61.5606 | 63.1671 | 66.1969 | 69.2769
-17 | 60.648 | 61.78 | 63.3691 | 66.4964 | 69.3845
-16.9 | 60.8663 | 61.9377 | 63.6567 | 66.6598 | 69.6396
-16.8 | 61.102 | 62.1538 | 63.7866 | 66.917 | 69.8612
-16.7 | 61.2826 | 62.4146 | 64.1017 | 67.0733 | 70.0067
-16.6 | 61.53 | 62.6424 | 64.19 | 67.323 | 70.3343
-16.5 | 61.7709 | 62.8463 | 64.4688 | 67.5507 | 70.5364
-16.4 | 61.9173 | 62.9991 | 64.6812 | 67.7781 | 70.7537
-16.3 | 62.1208 | 63.2262 | 64.9119 | 67.9179 | 71.0366
-16.2 | 62.3697 | 63.4649 | 65.1163 | 68.2139 | 71.1
-16.1 | 62.5389 | 63.6124 | 65.3259 | 68.416 | 71.3519
-16 | 62.7853 | 63.8614 | 65.5127 | 68.6252 | 71.6052
-15.9 | 62.954 | 64.0809 | 65.79 | 68.845 | 71.8881
-15.8 | 63.1909 | 64.2865 | 65.9754 | 69.0615 | 72.0506
-15.7 | 63.414 | 64.5303 | 66.221 | 69.3281 | 72.2369
-15.6 | 63.5766 | 64.7081 | 66.4564 | 69.4702 | 72.3652
-15.5 | 63.8328 | 64.906 | 66.5998 | 69.6994 | 72.6992
-15.4 | 64.0684 | 65.1377 | 66.9351 | 69.9588 | 72.9211
-15.3 | 64.2042 | 65.4533 | 67.1608 | 70.1709 | 73.1339
-15.2 | 64.4239 | 65.5784 | 67.344 | 70.3639 | 73.3533
-15.1 | 64.6307 | 65.7765 | 67.4854 | 70.6107 | 73.5423
-15 | 64.8781 | 66.0408 | 67.7718 | 70.878 | 73.7525
-14.9 | 65.1236 | 66.2371 | 67.9832 | 71.1137 | 73.9804
-14.8 | 65.324 | 66.4956 | 68.243 | 71.3281 | 74.2779
-14.7 | 65.5851 | 66.6852 | 68.4586 | 71.5301 | 74.3936
-14.6 | 65.8208 | 66.9081 | 68.6606 | 71.785 | 74.5893
-14.5 | 66.0283 | 67.1483 | 68.8626 | 72.0283 | 74.8848
-14.4 | 66.2106 | 67.4116 | 69.0648 | 72.2185 | 75
-14.3 | 66.3967 | 67.6134 | 69.2669 | 72.4405 | 75.2278
-14.2 | 66.6362 | 67.8291 | 69.5533 | 72.6245 | 75.4818
-14.1 | 66.8993 | 68.0117 | 69.7709 | 72.8993 | 75.7848
-14 | 67.135 | 68.2911 | 70.0017 | 73.081 | 76
-13.9 | 67.4024 | 68.5363 | 70.2375 | 73.3624 | 76.1338
-13.8 | 67.6095 | 68.7678 | 70.4877 | 73.577 | 76.4065
-13.7 | 67.8087 | 68.9616 | 70.6853 | 73.8685 | 76.6218
-13.6 | 68.0083 | 69.2087 | 70.9715 | 74.081 | 76.8084
-13.5 | 68.2703 | 69.4467 | 71.183 | 74.3134 | 77.0461
-13.4 | 68.5296 | 69.7293 | 71.4158 | 74.5492 | 77.2497
-13.3 | 68.7774 | 69.9059 | 71.6774 | 74.7509 | 77.4521
-13.2 | 69.0047 | 70.1708 | 71.9747 | 74.9847 | 77.7034
-13.1 | 69.1938 | 70.3923 | 72.1768 | 75.1346 | 77.9477
-13 | 69.444 | 70.6214 | 72.3552 | 75.4253 | 78.1306
-12.9 | 69.7099 | 70.8597 | 72.5932 | 75.6932 | 78.3566
-12.8 | 69.9975 | 71.0742 | 72.8302 | 75.938 | 78.596
-12.7 | 70.1658 | 71.3422 | 73.0658 | 76.1135 | 78.8325
-12.6 | 70.4015 | 71.544 | 73.3015 | 76.3156 | 79.0341
-12.5 | 70.6045 | 71.7826 | 73.4823 | 76.5544 | 79.1519
-12.4 | 70.8437 | 72.0342 | 73.7396 | 76.7911 | 79.5055
-12.3 | 71.113 | 72.3101 | 73.9902 | 77.0397 | 79.7075
-12.2 | 71.4111 | 72.5511 | 74.2731 | 77.2476 | 79.9111
-12.1 | 71.6467 | 72.8032 | 74.5134 | 77.4458 | 80.1467
-12 | 71.8824 | 73.09 | 74.763 | 77.6868 | 80.3707
-11.9 | 72.1182 | 73.3524 | 75.044 | 77.9528 | 80.5727
-11.8 | 72.3845 | 73.5386 | 75.2461 | 78.1577 | 80.7747
-11.7 | 72.6075 | 73.8157 | 75.4561 | 78.4105 | 80.9895
-11.6 | 72.8585 | 74.0512 | 75.7378 | 78.7073 | 81.2609
-11.5 | 73.1331 | 74.2063 | 75.9331 | 78.9093 | 81.4478
-11.4 | 73.416 | 74.4223 | 76.216 | 79.2338 | 81.6657
-11.3 | 73.6242 | 74.7041 | 76.4656 | 79.3995 | 81.9424
-11.2 | 73.8615 | 75.027 | 76.7297 | 79.6017 | 82.1303
-11.1 | 74.1444 | 75.327 | 76.9317 | 79.8845 | 82.3622
-11 | 74.4074 | 75.4683 | 77.1394 | 80.0836 | 82.6037
-10.9 | 74.6451 | 75.737 | 77.4127 | 80.3679 | 82.7902
-10.8 | 74.9593 | 76.0321 | 77.7424 | 80.5806 | 82.9721
-10.7 | 75.1465 | 76.2595 | 77.997 | 80.8398 | 83.2699
-10.6 | 75.3987 | 76.539 | 78.2733 | 80.9781 | 83.4187
-10.5 | 75.6816 | 76.7744 | 78.5297 | 81.277 | 83.7269
-10.4 | 75.9304 | 77.0348 | 78.7317 | 81.4889 | 83.9
-10.3 | 76.192 | 77.3375 | 78.917 | 81.7864 | 84.1008
-10.2 | 76.4702 | 77.601 | 79.1608 | 81.9929 | 84.4499
-10.1 | 76.7528 | 77.8163 | 79.4639 | 82.1949 | 84.6164
-10 | 76.9964 | 78.0888 | 79.7298 | 82.4947 | 84.7826
-9.9 | 77.2386 | 78.3915 | 80.0321 | 82.7418 | 85.027
-9.8 | 77.5214 | 78.6559 | 80.2678 | 82.9439 | 85.2867
-9.7 | 77.8042 | 78.8257 | 80.4621 | 83.2553 | 85.4888
-9.6 | 78.0435 | 79.128 | 80.6726 | 83.4765 | 85.6908
-9.5 | 78.3099 | 79.3956 | 80.9678 | 83.6624 | 85.9392
-9.4 | 78.5928 | 79.6863 | 81.2137 | 83.8864 | 86.1948
-9.3 | 78.8756 | 79.9664 | 81.513 | 84.1397 | 86.3678
-9.2 | 79.1365 | 80.1476 | 81.7487 | 84.3974 | 86.6341
-9.1 | 79.4178 | 80.4618 | 81.9844 | 84.7507 | 86.7
-9 | 79.6842 | 80.6395 | 82.3059 | 84.8658 | 86.9802
-8.9 | 79.9225 | 80.8944 | 82.5919 | 85.1004 | 87.1669
-8.8 | 80.1898 | 81.2061 | 82.7785 | 85.3356 | 87.4149
-8.7 | 80.509 | 81.5249 | 83.0207 | 85.5376 | 87.5563
-8.6 | 80.7778 | 81.7579 | 83.2963 | 85.8933 | 87.7629
-8.5 | 81.0979 | 82.1154 | 83.5697 | 86.0411 | 87.9703
-8.4 | 81.3412 | 82.2935 | 83.8727 | 86.3294 | 88.2012
-8.3 | 81.6034 | 82.6203 | 84.1025 | 86.522 | 88.5261
-8.2 | 81.9086 | 82.8258 | 84.336 | 86.7057 | 88.7835
-8.1 | 82.2061 | 83.1666 | 84.6297 | 86.9697 | 89.0035
-8 | 82.4657 | 83.4271 | 84.8157 | 87.2263 | 89.109
-7.9 | 82.7 | 83.6625 | 85.1231 | 87.4129 | 89.3145
-7.8 | 82.9728 | 83.9344 | 85.3228 | 87.6486 | 89.5796
-7.7 | 83.2811 | 84.2001 | 85.6512 | 87.9764 | 89.6933
-7.6 | 83.6039 | 84.4955 | 85.9028 | 88.0799 | 89.9828
-7.5 | 83.9167 | 84.802 | 86.1573 | 88.389 | 90.0953
-7.4 | 84.1247 | 85.1646 | 86.4403 | 88.6247 | 90.2685
-7.3 | 84.3925 | 85.375 | 86.6937 | 88.8604 | 90.5491
-7.2 | 84.6942 | 85.6297 | 86.9664 | 89.1314 | 90.8521
-7.1 | 84.9591 | 85.93 | 87.2318 | 89.3182 | 91.0477
-7 | 85.3006 | 86.1354 | 87.4439 | 89.5675 | 91.3253
-6.9 | 85.5239 | 86.3955 | 87.7365 | 89.8027 | 91.4313
-6.8 | 85.9042 | 86.6777 | 87.9722 | 90.0048 | 91.6233
-6.7 | 86.1295 | 87.0542 | 88.208 | 90.2068 | 91.8295
-6.6 | 86.4124 | 87.3117 | 88.4724 | 90.4133 | 91.9885
-6.5 | 86.719 | 87.6017 | 88.8345 | 90.7063 | 92.2127
-6.4 | 87.0151 | 87.8847 | 89.0151 | 90.8726 | 92.429
-6.3 | 87.3714 | 88.1479 | 89.2609 | 91.1721 | 92.5881
-6.2 | 87.6238 | 88.2982 | 89.5198 | 91.3741 | 92.7356
-6.1 | 87.9066 | 88.6208 | 89.7999 | 91.5761 | 93.1066
-6 | 88.1579 | 88.9711 | 90.0353 | 91.8094 | 93.3579
-5.9 | 88.4903 | 89.2196 | 90.2803 | 92.0602 | 93.483
-5.8 | 88.7527 | 89.4767 | 90.5108 | 92.3351 | 93.5
-5.7 | 89.0983 | 89.7583 | 90.7692 | 92.559 | 93.7557
-5.6 | 89.3608 | 90.061 | 91.0723 | 92.7673 | 93.9577
-5.5 | 89.6546 | 90.2915 | 91.3364 | 92.903 | 94.1523
-5.4 | 89.9721 | 90.4856 | 91.5721 | 93.1232 | 94.3332
-5.3 | 90.2078 | 90.8209 | 91.7781 | 93.3078 | 94.5411
-5.2 | 90.4722 | 91.1523 | 91.9801 | 93.5722 | 94.7329
-5.1 | 90.7425 | 91.3977 | 92.2732 | 93.8732 | 94.9425
-5 | 91.0482 | 91.6497 | 92.5763 | 94.0289 | 95.1
-4.9 | 91.3173 | 91.8732 | 92.7862 | 94.2173 | 95.2291
-4.8 | 91.5794 | 92.1608 | 93.0236 | 94.4147 | 95.4311
-4.7 | 91.8887 | 92.4216 | 93.2887 | 94.622 | 95.5666
-4.6 | 92.2238 | 92.6854 | 93.5492 | 94.8352 | 95.7892
-4.5 | 92.3967 | 92.9005 | 93.7401 | 94.996 | 95.8972
-4.4 | 92.6524 | 93.1728 | 94.0291 | 95.2146 | 96.0964
-4.3 | 92.9589 | 93.4756 | 94.1824 | 95.3648 | 96.2984
-4.2 | 93.2508 | 93.721 | 94.4004 | 95.5503 | 96.4
-4.1 | 93.5362 | 93.9423 | 94.6037 | 95.7635 | 96.5029
-4 | 93.7719 | 94.2268 | 94.9026 | 95.9289 | 96.6831
-3.9 | 94.0076 | 94.4798 | 95.0409 | 96.1065 | 96.8065
-3.8 | 94.272 | 94.6853 | 95.292 | 96.3085 | 97.01
-3.7 | 94.5124 | 94.9781 | 95.4874 | 96.5053 | 97.2196
-3.6 | 94.7481 | 95.1889 | 95.7126 | 96.6126 | 97.3206
-3.5 | 95.0256 | 95.4217 | 95.9146 | 96.8073 | 97.4378
-3.4 | 95.2717 | 95.636 | 96.1167 | 96.9146 | 97.6083
-3.3 | 95.5414 | 95.7773 | 96.3187 | 97.0914 | 97.7218
-3.2 | 95.789 | 96.0282 | 96.529 | 97.2922 | 97.9247
-3.1 | 96.0359 | 96.2807 | 96.7466 | 97.4449 | 98.0257
-3 | 96.1774 | 96.5346 | 96.9547 | 97.6947 | 98.1
-2.9 | 96.4376 | 96.7013 | 97.099 | 97.8157 | 98.1776
-2.8 | 96.6753 | 96.9177 | 97.3202 | 97.9804 | 98.3
-2.7 | 96.8694 | 97.0612 | 97.5027 | 98.032 | 98.4
-2.6 | 97.063 | 97.2412 | 97.7 | 98.1717 | 98.5076
-2.5 | 97.2741 | 97.5229 | 97.7408 | 98.3074 | 98.7
-2.4 | 97.4259 | 97.6788 | 97.9765 | 98.5 | 98.8049
-2.3 | 97.6122 | 97.8582 | 98.1673 | 98.5296 | 98.9273
-2.2 | 97.8775 | 98.0287 | 98.2777 | 98.6411 | 99.0573
-2.1 | 98.0169 | 98.1662 | 98.4102 | 98.8215 | 99.1
-2 | 98.2263 | 98.3848 | 98.5368 | 98.9226 | 99.1797
-1.9 | 98.3883 | 98.4965 | 98.6883 | 99.0236 | 99.2
-1.8 | 98.5744 | 98.5974 | 98.8465 | 99.1246 | 99.2674
-1.7 | 98.6965 | 98.7953 | 98.9399 | 99.2 | 99.3685
-1.6 | 98.8345 | 98.9134 | 99.0409 | 99.3104 | 99.4695
-1.5 | 99 | 99.0166 | 99.1838 | 99.4 | 99.5705
-1.4 | 99.1201 | 99.1343 | 99.3 | 99.4668 | 99.6
-1.3 | 99.2769 | 99.252 | 99.3 | 99.5 | 99.6582
-1.2 | 99.3229 | 99.3698 | 99.4471 | 99.6358 | 99.6408
-1.1 | 99.4055 | 99.4006 | 99.5 | 99.7 | 99.7205
-1 | 99.5822 | 99.5801 | 99.5258 | 99.7858 | 99.7387
-0.9 | 99.634 | 99.7003 | 99.648 | 99.8 | 99.7872
-0.8 | 99.7 | 99.7065 | 99.7 | 99.8286 | 99.9
-0.7 | 99.7626 | 99.8002 | 99.775 | 99.9 | 99.9215
-0.6 | 99.8 | 99.8002 | 99.8082 | 99.9 | 100
-0.5 | 99.8911 | 99.8276 | 99.9 | 99.9 | 100
-0.4 | 99.9453 | 99.9001 | 99.9 | 99.9 | 100
-0.3 | 100.0757 | 99.9001 | 99.9 | 99.9684 | 100
-0.2 | 100 | 99.9552 | 99.9095 | 100 | 99.9306
-0.1 | 100 | 100 | 99.9881 | 100 | 99.9448
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 99.999 | 99.9991 | 100 | 99.9418
0.2 | 100.0524 | 99.9001 | 99.9107 | 100 | 99.9518
0.3 | 99.9964 | 99.9001 | 99.9 | 99.9827 | 100
0.4 | 99.9 | 99.9001 | 99.9 | 99.9 | 100
0.5 | 99.9 | 99.873 | 99.9 | 99.9 | 100
0.6 | 99.8 | 99.8002 | 99.8082 | 99.9 | 100
0.7 | 99.7876 | 99.7503 | 99.775 | 99.9 | 99.9929
0.8 | 99.7 | 99.7003 | 99.7 | 99.8486 | 99.8905
0.9 | 99.6 | 99.7003 | 99.648 | 99.8 | 99.7672
1 | 99.6 | 99.5658 | 99.5258 | 99.7048 | 99.7618
1.1 | 99.4305 | 99.4323 | 99.5 | 99.7703 | 99.7402
1.2 | 99.343 | 99.4006 | 99.4471 | 99.5859 | 99.645
1.3 | 99.3 | 99.2008 | 99.3 | 99.5 | 99.6
1.4 | 99.0401 | 99.1344 | 99.3 | 99.5 | 99.6571
1.5 | 98.9984 | 99.001 | 99.1838 | 99.3822 | 99.5838
1.6 | 98.8477 | 98.9115 | 99.0409 | 99.2695 | 99.4483
1.7 | 98.7158 | 98.7478 | 98.9399 | 99.1632 | 99.3542
1.8 | 98.5907 | 98.6402 | 98.8465 | 99.1 | 99.2532
1.9 | 98.3775 | 98.4772 | 98.6883 | 99.0379 | 99.2
2 | 98.2516 | 98.4016 | 98.5368 | 98.9447 | 99.1511
2.1 | 98.0503 | 98.2187 | 98.4102 | 98.8002 | 99.0501
2.2 | 97.8145 | 98.0145 | 98.2777 | 98.6491 | 99
2.3 | 97.6746 | 97.9073 | 98.1673 | 98.5481 | 98.9073
2.4 | 97.4459 | 97.6739 | 97.9765 | 98.4382 | 98.7716
2.5 | 97.3204 | 97.4914 | 97.7408 | 98.3318 | 98.7
2.6 | 97.0717 | 97.3057 | 97.7 | 98.1061 | 98.4576
2.7 | 96.8616 | 97.08 | 97.5027 | 98 | 98.4
2.8 | 96.7002 | 96.9033 | 97.3202 | 97.9602 | 98.3
2.9 | 96.449 | 96.7023 | 97.099 | 97.842 | 98.1376
3 | 96.1956 | 96.5582 | 96.9547 | 97.7174 | 98.1
3.1 | 95.9699 | 96.2665 | 96.7466 | 97.4699 | 98.0114
3.2 | 95.729 | 96.0494 | 96.529 | 97.3288 | 97.9104
3.3 | 95.5552 | 95.8304 | 96.3187 | 97.1473 | 97.7901
3.4 | 95.2861 | 95.5761 | 96.1167 | 96.9453 | 97.5941
3.5 | 95.0256 | 95.3884 | 95.9146 | 96.7432 | 97.4861
3.6 | 94.7574 | 95.1529 | 95.7126 | 96.6206 | 97.292
3.7 | 94.4822 | 94.9643 | 95.4874 | 96.4593 | 97.1895
3.8 | 94.305 | 94.7038 | 95.292 | 96.2657 | 97.0575
3.9 | 94.0614 | 94.4692 | 95.0409 | 96.1318 | 96.8743
4 | 93.7631 | 94.2443 | 94.9026 | 95.9616 | 96.7
4.1 | 93.5043 | 94.0089 | 94.6037 | 95.7035 | 96.5434
4.2 | 93.3169 | 93.7069 | 94.4004 | 95.5669 | 96.4145
4.3 | 92.9578 | 93.5041 | 94.1824 | 95.3981 | 96.2698
4.4 | 92.6437 | 93.2014 | 94.0291 | 95.1624 | 96.0678
4.5 | 92.416 | 92.9005 | 93.7401 | 95.0134 | 95.8829
4.6 | 92.191 | 92.6711 | 93.5492 | 94.791 | 95.7492
4.7 | 91.933 | 92.4406 | 93.2887 | 94.5809 | 95.6
4.8 | 91.6035 | 92.139 | 93.0236 | 94.4395 | 95.4538
4.9 | 91.3506 | 91.9017 | 92.7862 | 94.1753 | 95.2503
5 | 90.9815 | 91.6354 | 92.5763 | 94.0489 | 95.1492
5.1 | 90.735 | 91.3961 | 92.2732 | 93.8458 | 94.8825
5.2 | 90.4152 | 91.1319 | 91.9801 | 93.6102 | 94.7186
5.3 | 90.1411 | 90.8708 | 91.7781 | 93.2706 | 94.5078
5.4 | 89.9054 | 90.5176 | 91.5721 | 93.1432 | 94.3046
5.5 | 89.6697 | 90.3279 | 91.3364 | 92.9364 | 94.1026
5.6 | 89.3012 | 90.0679 | 91.0723 | 92.7007 | 93.9006
5.7 | 89.0316 | 89.7708 | 90.7692 | 92.558 | 93.7237
5.8 | 88.7439 | 89.5053 | 90.5108 | 92.3376 | 93.5
5.9 | 88.4602 | 89.2053 | 90.2803 | 91.9968 | 93.4687
6 | 88.1868 | 88.9959 | 90.0353 | 91.8494 | 93.3245
6.1 | 87.8555 | 88.5971 | 89.7999 | 91.6047 | 93.0666
6.2 | 87.5797 | 88.3049 | 89.5198 | 91.354 | 92.8198
6.3 | 87.3174 | 88.1626 | 89.2609 | 91.1503 | 92.6631
6.4 | 87.0726 | 87.8665 | 89.0151 | 90.9226 | 92.4484
6.5 | 86.7352 | 87.5376 | 88.8345 | 90.723 | 92.269
6.6 | 86.477 | 87.259 | 88.4724 | 90.4561 | 91.9718
6.7 | 86.1695 | 86.967 | 88.208 | 90.256 | 91.7895
6.8 | 85.8584 | 86.7217 | 87.9722 | 89.9929 | 91.6034
6.9 | 85.6032 | 86.4501 | 87.7365 | 89.7599 | 91.5014
7 | 85.281 | 86.1814 | 87.4439 | 89.501 | 91.3007
7.1 | 84.9651 | 85.9323 | 87.2318 | 89.3651 | 91.0982
7.2 | 84.7153 | 85.7085 | 86.9664 | 89.1153 | 90.8395
7.3 | 84.4325 | 85.375 | 86.6937 | 88.8325 | 90.6062
7.4 | 84.058 | 85.102 | 86.4403 | 88.558 | 90.3355
7.5 | 83.8418 | 84.8486 | 86.1573 | 88.3223 | 90.1417
7.6 | 83.58 | 84.5454 | 85.9028 | 88.0866 | 89.96
7.7 | 83.2264 | 84.2874 | 85.6512 | 87.9106 | 89.7579
7.8 | 83.0228 | 83.9144 | 85.3228 | 87.5909 | 89.5559
7.9 | 82.6677 | 83.7394 | 85.1231 | 87.4462 | 89.3539
8 | 82.4157 | 83.527 | 84.8157 | 87.1926 | 89.1579
8.1 | 82.1897 | 83.1466 | 84.6297 | 87.0049 | 88.9748
8.2 | 81.8724 | 82.8531 | 84.336 | 86.6695 | 88.7634
8.3 | 81.6051 | 82.5875 | 84.1025 | 86.4841 | 88.5441
8.4 | 81.301 | 82.3047 | 83.8727 | 86.258 | 88.1677
8.5 | 81.0384 | 82.0739 | 83.5697 | 86.056 | 87.9417
8.6 | 80.7955 | 81.7627 | 83.2963 | 85.8296 | 87.7648
8.7 | 80.5127 | 81.5418 | 83.0207 | 85.5662 | 87.5727
8.8 | 80.2582 | 81.2079 | 82.7785 | 85.3642 | 87.3949
8.9 | 79.9558 | 80.9098 | 82.5919 | 85.1243 | 87.2004
9 | 79.6302 | 80.708 | 82.3059 | 84.8 | 86.8973
9.1 | 79.4511 | 80.5008 | 81.9844 | 84.7413 | 86.7
9.2 | 79.1615 | 80.2183 | 81.7487 | 84.4484 | 86.5989
9.3 | 78.8235 | 79.9664 | 81.513 | 84.1797 | 86.3797
9.4 | 78.566 | 79.7309 | 81.2137 | 83.9128 | 86.144
9.5 | 78.25 | 79.433 | 80.9678 | 83.6562 | 85.9357
9.6 | 78.0271 | 79.1268 | 80.6726 | 83.5035 | 85.7544
9.7 | 77.7443 | 78.8265 | 80.4621 | 83.303 | 85.4903
9.8 | 77.5614 | 78.6221 | 80.2678 | 82.9587 | 85.2581
9.9 | 77.2988 | 78.3828 | 80.0321 | 82.6704 | 85.0561
10 | 76.9947 | 78.1166 | 79.7298 | 82.4723 | 84.8541
10.1 | 76.7929 | 77.7829 | 79.4639 | 82.2328 | 84.5956
10.2 | 76.4875 | 77.5381 | 79.1608 | 81.9321 | 84.3917
10.3 | 76.217 | 77.285 | 78.917 | 81.742 | 84.1272
10.4 | 75.9522 | 77.0831 | 78.7317 | 81.5461 | 83.9
10.5 | 75.6322 | 76.8646 | 78.5297 | 81.2269 | 83.6768
10.6 | 75.4386 | 76.5468 | 78.2733 | 81.0133 | 83.3912
10.7 | 75.1799 | 76.2396 | 77.997 | 80.8113 | 83.2358
10.8 | 74.913 | 76.0171 | 77.7424 | 80.6093 | 82.9949
10.9 | 74.6301 | 75.7345 | 77.4127 | 80.4073 | 82.825
11 | 74.3472 | 75.5081 | 77.1394 | 80.0105 | 82.5836
11.1 | 74.0845 | 75.3092 | 76.9317 | 79.9245 | 82.3845
11.2 | 73.9008 | 75.0267 | 76.7297 | 79.6681 | 82.0876
11.3 | 73.6493 | 74.7047 | 76.4656 | 79.3485 | 81.9282
11.4 | 73.4338 | 74.5221 | 76.216 | 79.196 | 81.6813
11.5 | 73.1365 | 74.2171 | 75.9331 | 78.938 | 81.4521
11.6 | 72.8378 | 74.018 | 75.7378 | 78.7159 | 81.2501
11.7 | 72.5737 | 73.7824 | 75.4561 | 78.4436 | 81.0481
11.8 | 72.4154 | 73.5244 | 75.2461 | 78.1604 | 80.8205
11.9 | 72.0847 | 73.2858 | 75.044 | 77.9514 | 80.6385
12 | 71.8694 | 73.0757 | 74.763 | 77.6915 | 80.363
12.1 | 71.6801 | 72.7833 | 74.5134 | 77.4257 | 80.1134
12.2 | 71.3665 | 72.5386 | 74.2731 | 77.2332 | 79.9285
12.3 | 71.1087 | 72.3352 | 73.9902 | 77.0503 | 79.6503
12.4 | 70.8396 | 72.0048 | 73.7396 | 76.8483 | 79.4396
12.5 | 70.6279 | 71.803 | 73.4823 | 76.6194 | 79.2423
12.6 | 70.3682 | 71.5926 | 73.3015 | 76.3442 | 79.1174
12.7 | 70.1595 | 71.3279 | 73.0658 | 76.1422 | 78.8986
12.8 | 69.9362 | 71.1249 | 72.8302 | 75.9161 | 78.5781
12.9 | 69.7133 | 70.8569 | 72.5932 | 75.6334 | 78.3733
13 | 69.469 | 70.5999 | 72.3552 | 75.3752 | 78.1904
13.1 | 69.2277 | 70.3779 | 72.1768 | 75.1676 | 77.9075
13.2 | 68.9905 | 70.1761 | 71.9747 | 74.9207 | 77.6539
13.3 | 68.7273 | 69.9482 | 71.6774 | 74.771 | 77.4637
13.4 | 68.5159 | 69.7094 | 71.4158 | 74.5238 | 77.2707
13.5 | 68.3101 | 69.4441 | 71.183 | 74.2469 | 77.003
13.6 | 68.0334 | 69.2087 | 70.9715 | 74.1055 | 76.7667
13.7 | 67.8316 | 68.9416 | 70.6853 | 73.8112 | 76.5646
13.8 | 67.5799 | 68.771 | 70.4877 | 73.5448 | 76.3626
13.9 | 67.345 | 68.5164 | 70.2375 | 73.3321 | 76.1375
14 | 67.1017 | 68.2903 | 70.0017 | 73.1017 | 76.0008
14.1 | 66.9389 | 68.0647 | 69.7709 | 72.8593 | 75.7368
14.2 | 66.6485 | 67.8855 | 69.5533 | 72.6227 | 75.5259
14.3 | 66.4334 | 67.5993 | 69.2669 | 72.4382 | 75.2536
14.4 | 66.1795 | 67.3974 | 69.0648 | 72.2362 | 75.0538
14.5 | 66.0232 | 67.1955 | 68.8626 | 72.0078 | 74.9173
14.6 | 65.7651 | 66.974 | 68.6606 | 71.7249 | 74.6541
14.7 | 65.5184 | 66.6851 | 68.4586 | 71.5873 | 74.4388
14.8 | 65.2871 | 66.4617 | 68.243 | 71.3853 | 74.2566
14.9 | 65.0803 | 66.2598 | 67.9832 | 71.1648 | 74.0205
15 | 64.9223 | 66.0276 | 67.7718 | 70.8557 | 73.724
15.1 | 64.709 | 65.8553 | 67.4854 | 70.6506 | 73.5545
15.2 | 64.4533 | 65.6025 | 67.344 | 70.3867 | 73.3733
15.3 | 64.2709 | 65.3389 | 67.1608 | 70.1876 | 73.1608
15.4 | 64.0023 | 65.1035 | 66.9351 | 69.8873 | 72.9023
15.5 | 63.8491 | 64.8919 | 66.5998 | 69.7197 | 72.7389
15.6 | 63.5652 | 64.7626 | 66.4564 | 69.4976 | 72.3971
15.7 | 63.3737 | 64.5146 | 66.221 | 69.2948 | 72.2168
15.8 | 63.1257 | 64.2615 | 65.9754 | 69.0385 | 72.022
15.9 | 62.9926 | 64.124 | 65.79 | 68.754 | 71.8479
16 | 62.8157 | 63.8888 | 65.5127 | 68.6271 | 71.5651
16.1 | 62.5279 | 63.6587 | 65.3259 | 68.3667 | 71.3186
16.2 | 62.4 | 63.4505 | 65.1163 | 68.1424 | 71.1
16.3 | 62.1354 | 63.2487 | 64.9119 | 67.9404 | 70.9806
16.4 | 61.8836 | 62.9519 | 64.6812 | 67.7115 | 70.8069
16.5 | 61.7069 | 62.8164 | 64.4688 | 67.4792 | 70.4909
16.6 | 61.5281 | 62.6146 | 64.19 | 67.2772 | 70.4026
16.7 | 61.3034 | 62.4159 | 64.1017 | 67.0783 | 70
16.8 | 61.0687 | 62.2394 | 63.7866 | 66.9016 | 69.8823
16.9 | 60.8598 | 62.0091 | 63.6567 | 66.6798 | 69.5999
17 | 60.6784 | 61.8073 | 63.3691 | 66.4459 | 69.4406
17.1 | 60.4949 | 61.5809 | 63.1671 | 66.2242 | 69.257
17.2 | 60.3156 | 61.3017 | 62.9651 | 65.9511 | 69.0794
17.3 | 60.0901 | 61.173 | 62.7568 | 65.7451 | 68.8774
17.4 | 59.8727 | 61.0243 | 62.5323 | 65.5754 | 68.6754
17.5 | 59.6854 | 60.7905 | 62.3187 | 65.3732 | 68.4226
17.6 | 59.4497 | 60.5593 | 62.1298 | 65.1309 | 68.0915
17.7 | 59.3071 | 60.3365 | 61.9406 | 64.9489 | 67.9474
17.8 | 59.0725 | 60.0538 | 61.7386 | 64.7436 | 67.7117
17.9 | 58.8911 | 59.9058 | 61.5092 | 64.4759 | 67.5256
18 | 58.7041 | 59.732 | 61.2403 | 64.3201 | 67.3068
18.1 | 58.4856 | 59.5302 | 61.0427 | 64.0488 | 67.1254
18.2 | 58.3226 | 59.362 | 60.859 | 63.8305 | 66.8427
18.3 | 58.1199 | 59.1407 | 60.6332 | 63.5999 | 66.7142
18.4 | 57.8969 | 58.9389 | 60.457 | 63.4221 | 66.4642
18.5 | 57.657 | 58.7371 | 60.2371 | 63.1836 | 66.257
18.6 | 57.4963 | 58.5371 | 59.9927 | 63.0513 | 65.9927
18.7 | 57.3485 | 58.35 | 59.7743 | 62.7903 | 65.7785
18.8 | 57.1328 | 58.0674 | 59.604 | 62.5365 | 65.641
18.9 | 56.8714 | 57.9335 | 59.402 | 62.3428 | 65.4285
19 | 56.7124 | 57.7589 | 59.2 | 62.1198 | 65.1285
19.1 | 56.5142 | 57.4997 | 58.9809 | 61.9549 | 64.9265
19.2 | 56.367 | 57.2546 | 58.8226 | 61.7294 | 64.7244
19.3 | 56.1141 | 57.1086 | 58.6094 | 61.4428 | 64.5224
19.4 | 55.9304 | 56.9348 | 58.3404 | 61.1885 | 64.3204
19.5 | 55.8256 | 56.733 | 58.1524 | 61.0306 | 64.1184
19.6 | 55.6018 | 56.5312 | 58.0214 | 60.8028 | 63.8829
19.7 | 55.3834 | 56.3294 | 57.6801 | 60.6428 | 63.6
19.8 | 55.2171 | 56.176 | 57.4772 | 60.4171 | 63.3693
19.9 | 55.0171 | 55.9238 | 57.3143 | 60.162 | 63.1809
20 | 54.7958 | 55.714 | 57.1513 | 59.9929 | 63.0262
20.1 | 54.6086 | 55.5221 | 56.9343 | 59.7632 | 62.7905
20.2 | 54.4274 | 55.3203 | 56.7215 | 59.5491 | 62.6274
20.3 | 54.2915 | 55.1342 | 56.5164 | 59.2879 | 62.5096
20.4 | 54.1001 | 54.9451 | 56.3143 | 59.0835 | 62.2215
20.5 | 53.9086 | 54.7433 | 56.1327 | 58.8686 | 61.9906
20.6 | 53.7472 | 54.5848 | 55.8766 | 58.6787 | 61.8453
20.7 | 53.5116 | 54.3573 | 55.7267 | 58.5215 | 61.6226
20.8 | 53.3608 | 54.1522 | 55.4554 | 58.2739 | 61.4073
20.9 | 53.1229 | 54.0151 | 55.2829 | 58.0691 | 61.0687
21 | 52.9026 | 53.8487 | 55.0903 | 57.8831 | 61
21.1 | 52.74 | 53.5966 | 54.8335 | 57.6003 | 60.8002
21.2 | 52.6155 | 53.4028 | 54.6386 | 57.441 | 60.5695
21.3 | 52.3946 | 53.2571 | 54.4287 | 57.2085 | 60.3431
21.4 | 52.1559 | 53.0553 | 54.2516 | 56.9685 | 60.0724
21.5 | 52.0287 | 52.8652 | 54.0344 | 56.8241 | 59.9634
21.6 | 51.8882 | 52.6803 | 53.9108 | 56.5883 | 59.7421
21.7 | 51.6763 | 52.4992 | 53.6526 | 56.3526 | 59.4594
21.8 | 51.5168 | 52.2681 | 53.4168 | 56.1802 | 59.2502
21.9 | 51.3287 | 52.0605 | 53.2406 | 55.9359 | 59.0145
22 | 51.1122 | 51.9209 | 53.0455 | 55.7533 | 58.8481
22.1 | 50.9764 | 51.7581 | 52.837 | 55.5513 | 58.7073
22.2 | 50.7645 | 51.5406 | 52.6349 | 55.3493 | 58.5206
22.3 | 50.5717 | 51.2562 | 52.4329 | 55.1209 | 58.2779
22.4 | 50.4016 | 51.1183 | 52.1464 | 54.8453 | 58.0166
22.5 | 50.2337 | 50.916 | 52.0144 | 54.6432 | 57.8402
22.6 | 50.0388 | 50.7473 | 51.8269 | 54.441 | 57.6554
22.7 | 49.8933 | 50.5668 | 51.6248 | 54.239 | 57.3973
22.8 | 49.62 | 50.3415 | 51.3941 | 54.037 | 57.1898
22.9 | 49.5431 | 50.1174 | 51.1921 | 53.8349 | 56.9454
23 | 49.2931 | 49.9494 | 50.9885 | 53.6665 | 56.7551
23.1 | 49.1396 | 49.71 | 50.7434 | 53.5033 | 56.6097
23.2 | 49.0085 | 49.5342 | 50.5419 | 53.2204 | 56.3505
23.3 | 48.7976 | 49.3204 | 50.3175 | 53.0556 | 56.1269
23.4 | 48.6074 | 49.1185 | 50.1105 | 52.8535 | 55.9248
23.5 | 48.4319 | 48.9168 | 49.9319 | 52.666 | 55.755
23.6 | 48.2246 | 48.6923 | 49.6923 | 52.4385 | 55.5385
23.7 | 48.0539 | 48.554 | 49.4902 | 52.1902 | 55.336
23.8 | 47.8034 | 48.3541 | 49.2695 | 51.9881 | 55.1596
23.9 | 47.6254 | 48.1522 | 49.0576 | 51.7407 | 54.9753
24 | 47.4579 | 47.9219 | 48.8556 | 51.4849 | 54.784
24.1 | 47.2375 | 47.7147 | 48.6535 | 51.3518 | 54.523
24.2 | 47.0701 | 47.4975 | 48.4515 | 51.0761 | 54.34
24.3 | 46.7933 | 47.2437 | 48.2024 | 50.8365 | 54.117
24.4 | 46.5665 | 47.0082 | 47.976 | 50.6463 | 53.8554
24.5 | 46.3196 | 46.7843 | 47.7739 | 50.4432 | 53.6754
24.6 | 46.0504 | 46.5825 | 47.5719 | 50.2506 | 53.3586
24.7 | 45.7267 | 46.3806 | 47.3048 | 50.0414 | 53.3
24.8 | 45.317 | 46.1788 | 47.0458 | 49.7751 | 53.0225
24.9 | 44.3426 | 45.9386 | 46.8101 | 49.5461 | 52.8435
25 | 42.7708 | 45.6621 | 46.6067 | 49.3638 | 52.5894
25.1 | 40.4344 | 45.3627 | 46.376 | 49.1065 | 52.4193
25.2 | 37.2945 | 45.0912 | 46.1636 | 48.9609 | 52.3015
25.3 | 33.4907 | 44.7988 | 45.9405 | 48.6719 | 52.034
25.4 | 28.8717 | 44.4961 | 45.7316 | 48.4699 | 51.8127
25.5 | 24.3303 | 44.0884 | 45.496 | 48.2679 | 51.6107
25.6 | 20.1407 | 43.6006 | 45.2522 | 48.0602 | 51.4087
25.7 | 16.409 | 42.9303 | 44.9367 | 47.8354 | 51.2067
25.8 | 13.6436 | 41.9953 | 44.6266 | 47.6125 | 51.0046
25.9 | 11.0781 | 40.3179 | 44.4531 | 47.3598 | 50.8026
26 | 9.2944 | 37.849 | 44.2175 | 47.1578 | 50.5509
26.1 | 7.8417 | 34.5787 | 43.9226 | 46.9484 | 50.2986
26.2 | 6.8458 | 31.3893 | 43.7076 | 46.6953 | 50.1483
26.3 | 6.131 | 27.5199 | 43.3324 | 46.533 | 50.0077
26.4 | 5.5494 | 23.4836 | 43.0495 | 46.2619 | 49.7815
26.5 | 5.1102 | 19.8635 | 42.7642 | 45.9881 | 49.5433
26.6 | 4.8366 | 16.6687 | 42.4548 | 45.8612 | 49.304
26.7 | 4.6256 | 14.1374 | 42.0519 | 45.5721 | 49.1582
26.8 | 4.4488 | 12.0607 | 41.5636 | 45.3701 | 48.8651
26.9 | 4.298 | 10.5657 | 40.988 | 45.1681 | 48.6577
27 | 4.1802 | 9.301 | 40.1861 | 44.9491 | 48.4804
27.1 | 4.0548 | 8.3153 | 38.8484 | 44.658 | 48.258
27.2 | 3.9278 | 7.6354 | 37.0009 | 44.4333 | 48.0477
27.3 | 3.8767 | 7.043 | 34.6388 | 44.2099 | 47.8457
27.4 | 3.7588 | 6.6237 | 31.7282 | 43.958 | 47.5947
27.5 | 3.7 | 6.3134 | 28.5702 | 43.7486 | 47.3701
27.6 | 3.6 | 6.0104 | 24.9741 | 43.5128 | 47.1681
27.7 | 3.5719 | 5.8031 | 21.8869 | 43.2771 | 47.033
27.8 | 3.5 | 5.6013 | 19.0519 | 43.0216 | 46.8216
27.9 | 3.4196 | 5.4807 | 16.6543 | 42.7361 | 46.5724
28 | 3.3 | 5.3471 | 14.8247 | 42.47 | 46.304
28.1 | 3.3 | 5.212 | 13.1561 | 42.3172 | 46.1612
28.2 | 3.2 | 5.0941 | 11.8263 | 42.0555 | 45.9392
28.3 | 3.1778 | 4.9528 | 10.8187 | 41.7524 | 45.7968
28.4 | 3.1 | 4.8115 | 9.9988 | 41.4606 | 45.5921
28.5 | 3.0149 | 4.6953 | 9.3956 | 41.2356 | 45.2936
28.6 | 3 | 4.6566 | 8.8604 | 40.9837 | 45.1168
28.7 | 2.912 | 4.5276 | 8.4361 | 40.703 | 44.7934
28.8 | 2.8706 | 4.3956 | 8.1227 | 40.4515 | 44.637
28.9 | 2.8 | 4.3449 | 7.8731 | 40.1231 | 44.378
29 | 2.8 | 4.219 | 7.6097 | 39.8024 | 44.2539
29.1 | 2.7 | 4.1822 | 7.3887 | 39.4661 | 44.0279
29.2 | 2.7 | 4.0501 | 7.2416 | 39.0715 | 43.7499
29.3 | 2.6635 | 3.9414 | 7.0436 | 38.6118 | 43.5479
29.4 | 2.6 | 3.8961 | 6.9021 | 38.1076 | 43.3369
29.5 | 2.6 | 3.8134 | 6.7808 | 37.304 | 43.1407
29.6 | 2.5 | 3.7703 | 6.5988 | 35.9929 | 42.9321
29.7 | 2.5 | 3.6963 | 6.4178 | 34.5703 | 42.5958
29.8 | 2.5 | 3.6224 | 6.3803 | 32.5825 | 42.3637
29.9 | 2.4 | 3.509 | 6.255 | 30.2655 | 42.1916
30 | 2.4 | 3.492 | 6.1136 | 27.9378 | 41.8936
30.1 | 2.4 | 3.3966 | 5.9721 | 25.4691 | 41.7044
30.2 | 2.4 | 3.39 | 5.8307 | 22.715 | 41.4512
30.3 | 2.3 | 3.2967 | 5.7894 | 20.6585 | 41.2788
30.4 | 2.3 | 3.2891 | 5.6628 | 18.7722 | 41.0132
30.5 | 2.3 | 3.2009 | 5.5554 | 17.0205 | 40.8388
30.6 | 2.225 | 3.1968 | 5.425 | 15.6576 | 40.5626
30.7 | 2.2 | 3.0969 | 5.3 | 14.5653 | 40.3672
30.8 | 2.2 | 3.0791 | 5.2518 | 13.6427 | 40.0845
30.9 | 2.2 | 2.997 | 5.134 | 12.9042 | 39.9008
31 | 2.1 | 2.997 | 5.0162 | 12.1981 | 39.6393
31.1 | 2.1 | 2.9832 | 4.9816 | 11.7265 | 39.396
31.2 | 2.1 | 2.8971 | 4.8457 | 11.2914 | 39.1132
31.3 | 2.1 | 2.8971 | 4.8 | 10.9702 | 38.8585
31.4 | 2.0671 | 2.8783 | 4.6936 | 10.6343 | 38.5343
31.5 | 2 | 2.7972 | 4.6 | 10.4176 | 38.2307
31.6 | 2 | 2.7972 | 4.5424 | 10.1847 | 37.8945
31.7 | 2 | 2.7611 | 4.5 | 9.9694 | 37.5554
31.8 | 2 | 2.6973 | 4.4233 | 9.78 | 37.1266
31.9 | 1.9 | 2.6973 | 4.3866 | 9.6332 | 36.5855
32 | 1.9 | 2.6581 | 4.3 | 9.4419 | 35.7657
32.1 | 1.9 | 2.5974 | 4.2238 | 9.317 | 34.7281
32.2 | 1.9 | 2.5974 | 4.1823 | 9.1319 | 33.4818
32.3 | 1.9 | 2.5315 | 4.1 | 8.9014 | 31.8822
32.4 | 1.8496 | 2.4975 | 4.0994 | 8.7996 | 29.9921
32.5 | 1.8 | 2.4975 | 4 | 8.6986 | 27.8709
32.6 | 1.8 | 2.4975 | 3.9167 | 8.5966 | 25.7497
32.7 | 1.8 | 2.4085 | 3.8952 | 8.4552 | 23.721
32.8 | 1.8 | 2.3976 | 3.8 | 8.3282 | 21.8722
32.9 | 1.8 | 2.3976 | 3.8 | 8.1936 | 20.2506
33 | 1.7 | 2.3911 | 3.7221 | 8.0637 | 18.8095
33.1 | 1.7 | 2.2977 | 3.7 | 7.8895 | 17.4909
33.2 | 1.7 | 2.2977 | 3.6234 | 7.7568 | 16.5236
33.3 | 1.7 | 2.2977 | 3.6 | 7.6477 | 15.6335
33.4 | 1.7 | 2.2586 | 3.571 | 7.5467 | 14.937
33.5 | 1.7 | 2.1978 | 3.5 | 7.4456 | 14.3424
33.6 | 1.6687 | 2.1978 | 3.5 | 7.3515 | 13.8452
33.7 | 1.6 | 2.1978 | 3.4 | 7.241 | 13.4031
33.8 | 1.6 | 2.1978 | 3.4 | 7.1372 | 13.0991
33.9 | 1.6 | 2.1463 | 3.3 | 7.0415 | 12.7961
34 | 1.6 | 2.0979 | 3.3 | 6.9405 | 12.4946
34.1 | 1.6 | 2.0979 | 3.3 | 6.9 | 12.2753
34.2 | 1.6 | 2.0979 | 3.2283 | 6.7813 | 12.0679
34.3 | 1.6 | 2.0932 | 3.2 | 6.6724 | 11.8925
34.4 | 1.5 | 2.0049 | 3.171 | 6.6 | 11.7301
34.5 | 1.5 | 1.998 | 3.1 | 6.5212 | 11.564
34.6 | 1.5 | 1.998 | 3.1 | 6.4 | 11.4261
34.7 | 1.5 | 1.998 | 3.0868 | 6.3468 | 11.2113
34.8 | 1.5 | 1.998 | 3 | 6.2211 | 11.0534
34.9 | 1.5 | 1.962 | 3 | 6.2 | 10.96
35 | 1.5 | 1.8981 | 2.9266 | 6.1 | 10.818
35.1 | 1.5 | 1.8981 | 2.9 | 6.0757 | 10.658
35.2 | 1.4664 | 1.8981 | 2.9 | 5.9855 | 10.5498
35.3 | 1.4 | 1.8981 | 2.8782 | 5.8844 | 10.4416
35.4 | 1.4 | 1.8981 | 2.8 | 5.8 | 10.3406
35.5 | 1.4 | 1.8443 | 2.8 | 5.7795 | 10.1792
35.6 | 1.4 | 1.7982 | 2.7617 | 5.7 | 10.0386
35.7 | 1.4 | 1.7982 | 2.7 | 5.6271 | 9.9376
35.8 |  | 1.7982 | 2.7 | 5.5911 | 9.7889
35.9 |  | 1.7982 | 2.7 | 5.5 | 9.6498
36 |  | 1.7982 | 2.7 | 5.4083 | 9.5488
36.1 |  | 1.7982 | 2.6 | 5.3724 | 9.5
36.2 |  | 1.7022 | 2.6 | 5.3 | 9.4046
36.3 |  | 1.6983 | 2.6 | 5.26 | 9.2301
36.4 |  | 1.6983 | 2.6 | 5.2 | 9.1162
36.5 |  | 1.6983 | 2.5009 | 5.151 | 9.0152
36.6 |  | 1.6983 | 2.5 | 5.1 | 8.9141
36.7 |  |  | 2.5 | 5.0417 | 8.8131
36.8 |  |  | 2.5 | 5 | 8.7231
36.9 |  |  | 2.4755 | 4.9472 | 8.6254
37 |  |  | 2.4 | 4.9 | 8.5243
37.1 |  |  | 2.4 | 4.8439 | 8.4233
37.2 |  |  | 2.4 | 4.8 | 8.4
37.3 |  |  | 2.4 | 4.7356 | 8.2915
37.4 |  |  | 2.4 | 4.7 | 8.1775
37.5 |  |  | 2.3049 | 4.6225 | 8.1
37.6 |  |  | 2.3 | 4.6 | 8.0547
37.7 |  |  | 2.3 | 4.5034 | 7.9459
37.8 |  |  | 2.3 | 4.5 | 7.8517
37.9 |  |  | 2.2723 | 4.4844 | 7.8
38 |  |  | 2.2 | 4.4 | 7.7427
38.1 |  |  | 2.2 | 4.4 | 7.6417
38.2 |  |  |  | 4.3693 | 7.5171
38.3 |  |  |  | 4.3 | 7.4963
38.4 |  |  |  | 4.2673 | 7.4
38.5 |  |  |  | 4.2 | 7.3806
38.6 |  |  |  | 4.2 | 7.2761
38.7 |  |  |  | 4.1785 | 7.2
38.8 |  |  |  | 4.1 | 7.1286
38.9 |  |  |  | 4.0765 | 7.0893
39 |  |  |  | 4 | 7
39.1 |  |  |  | 4 | 7
39.2 |  |  |  | 4 | 6.9163
39.3 |  |  |  | 3.9512 | 6.9
39.4 |  |  |  | 3.9 | 6.7833
39.5 |  |  |  | 3.9 | 6.7
39.6 |  |  |  | 3.8122 | 6.6476
39.7 |  |  |  | 3.8 | 6.6
39.8 |  |  |  | 3.8 | 6.5143
39.9 |  |  |  | 3.7092 | 6.5
40 |  |  |  | 3.7 | 6.4
40.1 |  |  |  | 3.7 | 6.4
40.2 |  |  |  | 3.6776 | 6.3
40.3 |  |  |  | 3.6 | 6.3
40.4 |  |  |  | 3.6 | 6.2327
40.5 |  |  |  | 3.6 | 6.2
40.6 |  |  |  | 3.5768 | 6.1
40.7 |  |  |  | 3.5 | 6.1
40.8 |  |  |  | 3.5 | 6
40.9 |  |  |  |  | 6
41 |  |  |  |  | 5.9695
41.1 |  |  |  |  | 5.9
41.2 |  |  |  |  | 5.9
41.3 |  |  |  |  | 5.8
41.4 |  |  |  |  | 5.8
41.5 |  |  |  |  | 5.7787
41.6 |  |  |  |  | 5.7
41.7 |  |  |  |  | 5.6767
41.8 |  |  |  |  | 5.6
41.9 |  |  |  |  | 5.6
42 |  |  |  |  | 5.5192
42.1 |  |  |  |  | 5.5
42.2 |  |  |  |  | 5.4859
42.3 |  |  |  |  | 5.4
42.4 |  |  |  |  | 5.3838
42.5 |  |  |  |  | 5.3
42.6 |  |  |  |  | 5.3
42.7 |  |  |  |  | 5.2522
42.8 |  |  |  |  | 5.2
42.9 |  |  |  |  | 5.2
43 |  |  |  |  | 5.2
43.1 |  |  |  |  | 5.1229
43.2 |  |  |  |  | 5.1
43.3 |  |  |  |  | 5.1
43.4 |  |  |  |  | 5.0022
43.5 |  |  |  |  | 5
43.6 |  |  |  |  | 4.9126
43.7 |  |  |  |  | 4.9

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.896 | 0.907 | 0.914 | 0.922 | 0.928 | 0.932 | 0.934 | 0.937 | 0.937
| 4 | 0.907 | 0.921 | 0.929 | 0.94 | 0.947 | 0.953 | 0.956 | 0.958 | 0.959
| 5 | 0.914 | 0.93 | 0.94 | 0.952 | 0.962 | 0.969 | 0.972 | 0.975 | 0.976
| 7 | 0.925 | 0.943 | 0.955 | 0.97 | 0.982 | 0.991 | 0.996 | 1 | 1.001
| 10 | 0.933 | 0.953 | 0.967 | 0.985 | 1 | 1.012 | 1.018 | 1.024 | 1.025
| 15 | 0.938 | 0.96 | 0.975 | 0.996 | 1.013 | 1.029 | 1.038 | 1.046 | 1.048
| 20 | 0.941 | 0.963 | 0.979 | 1.001 | 1.02 | 1.038 | 1.049 | 1.058 | 1.06
| 30 | 0.944 | 0.967 | 0.983 | 1.006 | 1.028 | 1.048 | 1.061 | 1.072 | 1.075
| 40 | 0.945 | 0.968 | 0.985 | 1.009 | 1.031 | 1.052 | 1.065 | 1.078 | 1.08