## Title
Varian TrueBeam
8MV Eclipse data tables
Revision 1, November 2012

## Table of Contents
Table of contents
1. Data tables |  | 2. Plots
| Open field |  | Open Field
| Open Field Depth Doses |  | Wedge 15
| Open Field Profiles at 2cm |  | Wedge 30 Plots
| Open Field Profiles at 5cm |  | Wedge 45 Plots
| Open Field Profiles at 10cm |  | Wedge 60 Plots
Open Field Profiles at 20cm
Open Field Profiles at 30cm
Diagonal Profiles
Open field Output Factors
Wedge 15
Wedge 15 Depth Doses
SSD80 Wedge 15 Depth Doses
SSD120 Wedge 15 Depth Doses
Wedge 15 Lateral Profiles
Wedge 15 Longitudinal Profiles
Wedge 15 Output Factors
Wedge 30
Wedge 30 Depth Doses
SSD80 Wedge 30 Depth Doses
SSD120 Wedge 30 Depth Doses
Wedge 30 Lateral Profiles
Wedge 30 Longitudinal Profiles
Wedge 30 Output Factors
Wedge 45
Wedge 45 Depth Doses
SSD80 Wedge 45 Depth Doses
Wedge 45 Longitudinal Profiles
Wedge 45 Lateral Profiles
Wedge 45 Longitudinal Profiles
Wedge 45 Output Factors
Wedge 60
Wedge 60 Depth Doses
SSD80 Wedge 60 Depth Doses
SSD120 Wedge 60 Depth Doses
Wedge 60 Lateral Profiles
Wedge 60 Longitudinal Profiles
Wedge 60 Output Factors

## Open Field Depth Dose
Percentage Depth Doses
Open field
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization Chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 3x3cm2 | 4x4cm2 | 6x6cm2 | 8x8cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x40cm2
0 | 36.3 | 37 | 39.2 | 41.2 | 43.3 | 53 | 60.2 | 63.7
0.1 | 40.175 | 41.08 | 43.2667 | 45.2667 | 47.2 | 56.9 | 64.06 | 67.55
0.2 | 47.7 | 47.9 | 49.85 | 51.7333 | 53.75 | 62.725 | 69.2 | 72.56
0.3 | 57.46 | 57.8143 | 59.2 | 61.1333 | 62.56 | 70.5667 | 76.16 | 78.6857
0.4 | 65.3889 | 66.1 | 67.7 | 68.5 | 70.1333 | 77.225 | 81.45 | 83.6
0.5 | 72.78 | 72.7667 | 73.9 | 75.15 | 76.76 | 82.4286 | 86.26 | 87.675
0.6 | 78.125 | 78.3667 | 79.65 | 80.5 | 81.775 | 86.5 | 89.9 | 90.74
0.7 | 82.96 | 83.18 | 83.9 | 84.8 | 85.84 | 89.7714 | 92.55 | 93.0667
0.8 | 86.525 | 86.5 | 87.32 | 88.0333 | 88.925 | 92.5 | 94.54 | 95.025
0.9 | 89.6 | 89.58 | 90.1333 | 90.9286 | 91.6 | 94.4571 | 96.225 | 96.4333
1 | 92.3 | 92.2 | 92.56 | 93.1 | 93.6857 | 96.1333 | 97.4 | 97.54
1.1 | 94.16 | 94.08 | 94.45 | 94.9571 | 95.4 | 97.35 | 98.325 | 98.3857
1.2 | 95.825 | 95.6143 | 96.0333 | 96.38 | 96.8 | 98.25 | 98.92 | 99
1.3 | 97 | 96.96 | 97.2 | 97.4667 | 97.74 | 98.9429 | 99.45 | 99.3375
1.4 | 98.1 | 98 | 98.14 | 98.325 | 98.5833 | 99.44 | 99.7 | 99.72
1.5 | 98.84 | 98.7333 | 98.9 | 98.9667 | 99.2 | 99.7 | 99.92 | 99.85
1.6 | 99.3286 | 99.3429 | 99.4 | 99.425 | 99.6286 | 99.9 | 100 | 100
1.7 | 99.6333 | 99.7 | 99.75 | 99.78 | 99.8 | 100 | 100 | 100
1.8 | 99.88 | 99.8857 | 99.9571 | 100 | 100 | 100 | 100 | 100
1.9 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9 | 99.8571
2 | 100 | 100 | 100 | 100 | 100 | 99.8875 | 99.76 | 99.725
2.1 | 99.9333 | 99.9 | 99.94 | 99.9 | 99.9 | 99.74 | 99.6667 | 99.52
2.2 | 99.575 | 99.8 | 99.7429 | 99.8 | 99.8 | 99.525 | 99.45 | 99.4
2.3 | 99.3714 | 99.6 | 99.6 | 99.5667 | 99.625 | 99.34 | 99.3 | 99.04
2.4 | 99.16 | 99.4 | 99.28 | 99.225 | 99.32 | 99 | 98.96 | 98.8333
2.5 | 98.76 | 99.04 | 99 | 99 | 98.95 | 98.65 | 98.76 | 98.58
2.6 | 98.4 | 98.7333 | 98.68 | 98.6333 | 98.625 | 98.4 | 98.475 | 98.2429
2.7 | 98.0857 | 98.25 | 98.42 | 98.3 | 98.28 | 98.05 | 98.2 | 97.98
2.8 | 97.55 | 97.825 | 97.94 | 98.05 | 97.95 | 97.8 | 97.875 | 97.6429
2.9 | 97.1 | 97.375 | 97.62 | 97.6 | 97.5 | 97.45 | 97.5 | 97.5
3 | 96.65 | 96.95 | 97.2 | 97.2 | 97.18 | 97.1 | 97.25 | 97.02
3.1 | 96.3 | 96.575 | 96.82 | 96.75 | 96.825 | 96.75 | 96.9 | 96.7
3.2 | 95.9 | 96.16 | 96.3333 | 96.35 | 96.4 | 96.42 | 96.6 | 96.3667
3.3 | 95.3 | 95.72 | 95.8571 | 95.95 | 96.0571 | 96.0333 | 96.3 | 95.9833
3.4 | 94.85 | 95.2 | 95.44 | 95.56 | 95.6 | 95.675 | 95.8714 | 95.7571
3.5 | 94.3 | 94.8 | 95.08 | 95.0667 | 95.25 | 95.275 | 95.5286 | 95.4
3.6 | 93.8 | 94.3 | 94.64 | 94.6857 | 94.8 | 94.925 | 95.2 | 95.1
3.7 | 93.32 | 93.8 | 94.18 | 94.25 | 94.4429 | 94.6 | 94.94 | 94.82
3.8 | 92.94 | 93.3 | 93.75 | 93.8 | 94.02 | 94.18 | 94.5667 | 94.5
3.9 | 92.32 | 92.85 | 93.25 | 93.4667 | 93.5875 | 93.8333 | 94.275 | 94.04
4 | 91.84 | 92.42 | 92.88 | 93.1 | 93.22 | 93.5333 | 93.9 | 93.775
4.1 | 91.42 | 91.95 | 92.4 | 92.6 | 92.7571 | 93.15 | 93.5571 | 93.45
4.2 | 90.825 | 91.5 | 91.8667 | 92.1714 | 92.34 | 92.76 | 93.2 | 93.125
4.3 | 90.45 | 90.96 | 91.5 | 91.78 | 92 | 92.45 | 92.94 | 92.7667
4.4 | 89.875 | 90.525 | 91.05 | 91.35 | 91.55 | 92.2 | 92.46 | 92.375
4.5 | 89.45 | 90.15 | 90.62 | 90.95 | 91.2333 | 91.75 | 92.18 | 92.08
4.6 | 88.95 | 89.64 | 90.2667 | 90.55 | 90.775 | 91.3667 | 91.8667 | 91.7
4.7 | 88.525 | 89.2 | 89.76 | 90.1333 | 90.4 | 91 | 91.5571 | 91.44
4.8 | 88.025 | 88.76 | 89.32 | 89.675 | 89.98 | 90.7 | 91.15 | 91.1
4.9 | 87.375 | 88.2 | 88.95 | 89.2667 | 89.65 | 90.2 | 90.85 | 90.75
5 | 87.075 | 87.8 | 88.5 | 88.8 | 89.26 | 89.9 | 90.48 | 90.5
5.1 | 86.58 | 87.3 | 88.075 | 88.5 | 88.8667 | 89.575 | 90.1333 | 90.1286
5.2 | 86.15 | 86.9 | 87.65 | 88.14 | 88.46 | 89.3 | 89.85 | 89.82
5.3 | 85.6 | 86.36 | 87.22 | 87.7143 | 88.05 | 88.85 | 89.5 | 89.4333
5.4 | 85.2 | 85.88 | 86.8 | 87.28 | 87.7333 | 88.5 | 89.175 | 89.1
5.5 | 84.7 | 85.4714 | 86.38 | 86.975 | 87.3 | 88.25 | 88.85 | 88.8167
5.6 | 84.3 | 85.1 | 86 | 86.5286 | 86.9 | 87.825 | 88.48 | 88.55
5.7 | 83.82 | 84.5857 | 85.5 | 86.0429 | 86.55 | 87.46 | 88.1667 | 88.15
5.8 | 83.4 | 84.14 | 85.1 | 85.7 | 86.15 | 87.1 | 87.8 | 87.84
5.9 | 82.92 | 83.65 | 84.6 | 85.24 | 85.675 | 86.7333 | 87.375 | 87.4857
6 | 82.5 | 83.12 | 84.2 | 84.8 | 85.28 | 86.4 | 87.1 | 87.18
6.1 | 81.95 | 82.8 | 83.7667 | 84.35 | 84.9 | 86.05 | 86.7667 | 86.925
6.2 | 81.44 | 82.35 | 83.35 | 84.12 | 84.5 | 85.7 | 86.45 | 86.54
6.3 | 81.05 | 81.85 | 83 | 83.6857 | 84.2 | 85.35 | 86.125 | 86.1667
6.4 | 80.475 | 81.46 | 82.5 | 83.34 | 83.8 | 85 | 85.8429 | 85.925
6.5 | 80.075 | 81.0333 | 82.04 | 82.9333 | 83.42 | 84.6333 | 85.5 | 85.5857
6.6 | 79.75 | 80.56 | 81.65 | 82.4667 | 83.025 | 84.225 | 85.2286 | 85.24
6.7 | 79.18 | 80.1143 | 81.225 | 82.1 | 82.575 | 83.9667 | 84.8 | 84.94
6.8 | 78.9 | 79.66 | 80.85 | 81.725 | 82.25 | 83.5 | 84.475 | 84.64
6.9 | 78.4 | 79.2 | 80.45 | 81.18 | 81.75 | 83.2333 | 84.14 | 84.2857
7 | 78 | 78.85 | 80.04 | 80.8667 | 81.34 | 82.9 | 83.82 | 84
7.1 | 77.6 | 78.48 | 79.64 | 80.45 | 80.98 | 82.55 | 83.5 | 83.58
7.2 | 77.1 | 77.98 | 79.12 | 80.06 | 80.7 | 82.2 | 83.1571 | 83.4
7.3 | 76.6857 | 77.55 | 78.825 | 79.6 | 80.375 | 81.82 | 82.74 | 83
7.4 | 76.3 | 77.1 | 78.375 | 79.3 | 79.9667 | 81.5 | 82.3889 | 82.7
7.5 | 75.85 | 76.8 | 78.08 | 78.9429 | 79.56 | 81.24 | 82.16 | 82.4
7.6 | 75.4 | 76.4 | 77.5875 | 78.5 | 79.2 | 80.8571 | 81.8625 | 82.06
7.7 | 74.85 | 75.925 | 77.2 | 78.0875 | 78.86 | 80.52 | 81.45 | 81.8143
7.8 | 74.5 | 75.45 | 76.75 | 77.6 | 78.4 | 80.1571 | 81.1667 | 81.4429
7.9 | 74.05 | 75.15 | 76.4 | 77.3333 | 78.1 | 79.84 | 80.86 | 81.1
8 | 73.675 | 74.64 | 76 | 76.9333 | 77.72 | 79.5 | 80.5 | 80.88
8.1 | 73.1571 | 74.25 | 75.6 | 76.5333 | 77.35 | 79.05 | 80.18 | 80.52
8.2 | 72.8 | 73.8333 | 75.28 | 76.175 | 77 | 78.7667 | 79.8667 | 80.18
8.3 | 72.4 | 73.5 | 74.9 | 75.85 | 76.625 | 78.48 | 79.6 | 79.92
8.4 | 72.025 | 73.0333 | 74.4 | 75.54 | 76.175 | 78.1111 | 79.2286 | 79.66
8.5 | 71.7 | 72.6 | 74.1 | 75.1 | 75.9 | 77.76 | 78.95 | 79.36
8.6 | 71.26 | 72.25 | 73.675 | 74.75 | 75.5333 | 77.45 | 78.6333 | 79
8.7 | 70.8 | 71.9 | 73.35 | 74.3 | 75.16 | 77.14 | 78.3 | 78.68
8.8 | 70.34 | 71.5 | 72.925 | 73.94 | 74.76 | 76.84 | 78.0571 | 78.4
8.9 | 70.0875 | 71.1667 | 72.55 | 73.6 | 74.4333 | 76.5 | 77.7 | 78.16
9 | 69.7 | 70.7 | 72.15 | 73.26 | 74.1 | 76.18 | 77.4 | 77.8
9.1 | 69.3286 | 70.25 | 71.825 | 72.95 | 73.75 | 75.86 | 77.04 | 77.38
9.2 | 68.96 | 70 | 71.4 | 72.5 | 73.2857 | 75.58 | 76.7 | 77.2
9.3 | 68.6375 | 69.5 | 71.025 | 72.1667 | 73.025 | 75.16 | 76.4857 | 76.9
9.4 | 68.12 | 69.15 | 70.6667 | 71.8 | 72.7 | 74.94 | 76.16 | 76.6143
9.5 | 67.7375 | 68.75 | 70.26 | 71.45 | 72.3 | 74.62 | 75.85 | 76.3
9.6 | 67.36 | 68.4 | 69.8 | 71 | 72.02 | 74.2 | 75.52 | 75.98
9.7 | 66.95 | 68.0333 | 69.6 | 70.7 | 71.64 | 73.92 | 75.18 | 75.66
9.8 | 66.64 | 67.6429 | 69.2429 | 70.35 | 71.35 | 73.6 | 74.92 | 75.425
9.9 | 66.2571 | 67.36 | 68.9 | 70.0571 | 70.9 | 73.25 | 74.625 | 75.0667
10 | 65.95 | 66.98 | 68.48 | 69.7429 | 70.55 | 72.95 | 74.32 | 74.8
10.1 | 65.5667 | 66.55 | 68.14 | 69.4 | 70.3 | 72.625 | 73.94 | 74.46
10.2 | 65.12 | 66.15 | 67.8 | 68.92 | 69.925 | 72.4 | 73.675 | 74.25
10.3 | 64.8143 | 65.725 | 67.5 | 68.58 | 69.6333 | 72 | 73.425 | 73.8667
10.4 | 64.44 | 65.45 | 67.1286 | 68.26 | 69.275 | 71.75 | 73.125 | 73.7
10.5 | 64 | 65.08 | 66.6667 | 67.92 | 68.9 | 71.4429 | 72.8 | 73.2333
10.6 | 63.76 | 64.7 | 66.3714 | 67.52 | 68.55 | 71.1 | 72.525 | 72.95
10.7 | 63.32 | 64.38 | 66 | 67.26 | 68.2 | 70.7714 | 72.1429 | 72.7
10.8 | 62.94 | 64 | 65.65 | 66.925 | 67.825 | 70.48 | 71.9 | 72.4667
10.9 | 62.62 | 63.7 | 65.3333 | 66.56 | 67.6667 | 70.2667 | 71.55 | 72.18
11 | 62.28 | 63.35 | 65 | 66.28 | 67.1 | 69.88 | 71.24 | 71.875
11.1 | 61.88 | 62.9875 | 64.6 | 65.95 | 66.9 | 69.575 | 70.9714 | 71.5
11.2 | 61.65 | 62.6 | 64.3 | 65.625 | 66.62 | 69.28 | 70.7 | 71.2667
11.3 | 61.325 | 62.2429 | 63.875 | 65.25 | 66.275 | 68.96 | 70.375 | 71.1
11.4 | 60.9 | 62 | 63.5333 | 64.925 | 65.95 | 68.64 | 70.15 | 70.7625
11.5 | 60.575 | 61.575 | 63.25 | 64.65 | 65.575 | 68.35 | 69.86 | 70.44
11.6 | 60.22 | 61.3333 | 63 | 64.225 | 65.325 | 68.1 | 69.56 | 70.24
11.7 | 59.8571 | 60.975 | 62.64 | 63.875 | 64.975 | 67.75 | 69.2667 | 69.9
11.8 | 59.5 | 60.6 | 62.3286 | 63.66 | 64.625 | 67.5 | 69.025 | 69.7
11.9 | 59.25 | 60.275 | 61.94 | 63.35 | 64.4 | 67.15 | 68.7 | 69.36
12 | 58.8667 | 60 | 61.6 | 63 | 64.05 | 66.9 | 68.475 | 69.0143
12.1 | 58.58 | 59.575 | 61.2714 | 62.65 | 63.7 | 66.575 | 68.15 | 68.8143
12.2 | 58.25 | 59.26 | 61 | 62.3 | 63.35 | 66.325 | 67.875 | 68.5
12.3 | 57.9 | 58.95 | 60.66 | 62 | 63.0333 | 65.975 | 67.55 | 68.22
12.4 | 57.55 | 58.675 | 60.3 | 61.65 | 62.78 | 65.575 | 67.3 | 67.9333
12.5 | 57.3 | 58.25 | 59.9714 | 61.3 | 62.4 | 65.425 | 66.925 | 67.6286
12.6 | 57 | 58.025 | 59.7 | 61.02 | 62.1 | 65.1 | 66.68 | 67.42
12.7 | 56.55 | 57.7 | 59.4 | 60.7 | 61.8 | 64.8 | 66.45 | 67.08
12.8 | 56.225 | 57.35 | 59.1 | 60.42 | 61.5 | 64.5 | 66.2 | 66.8429
12.9 | 55.9667 | 57.05 | 58.7 | 60.1 | 61.16 | 64.2 | 65.9 | 66.5571
13 | 55.68 | 56.68 | 58.4 | 59.775 | 60.95 | 63.925 | 65.6 | 66.32
13.1 | 55.4 | 56.4333 | 58.05 | 59.5 | 60.62 | 63.6 | 65.28 | 66.0571
13.2 | 55.0714 | 56.08 | 57.8 | 59.225 | 60.3 | 63.34 | 65.05 | 65.74
13.3 | 54.7857 | 55.825 | 57.55 | 58.8667 | 60.06 | 63.075 | 64.76 | 65.5333
13.4 | 54.48 | 55.48 | 57.2 | 58.6 | 59.72 | 62.8 | 64.52 | 65.225
13.5 | 54.15 | 55.25 | 56.975 | 58.25 | 59.425 | 62.5714 | 64.2 | 65
13.6 | 53.8333 | 54.9 | 56.625 | 58 | 59.2 | 62.2 | 64.025 | 64.76
13.7 | 53.575 | 54.58 | 56.25 | 57.7 | 58.9333 | 61.975 | 63.68 | 64.5
13.8 | 53.325 | 54.3 | 56 | 57.5 | 58.52 | 61.725 | 63.475 | 64.18
13.9 | 52.9 | 54 | 55.675 | 57.15 | 58.275 | 61.45 | 63.2 | 63.9
14 | 52.7 | 53.6 | 55.4 | 56.8667 | 57.95 | 61.15 | 62.9 | 63.68
14.1 | 52.45 | 53.35 | 55.1333 | 56.5 | 57.66 | 60.8 | 62.66 | 63.3875
14.2 | 52.1 | 53.1 | 54.8 | 56.25 | 57.3333 | 60.6 | 62.3375 | 63.1
14.3 | 51.7333 | 52.8 | 54.55 | 56 | 57.1 | 60.3 | 62.1 | 62.8429
14.4 | 51.5667 | 52.5 | 54.2667 | 55.6333 | 56.76 | 60.05 | 61.8143 | 62.66
14.5 | 51.2 | 52.26 | 53.875 | 55.45 | 56.48 | 59.8 | 61.56 | 62.3875
14.6 | 51 | 51.9 | 53.7 | 55.0667 | 56.2667 | 59.5 | 61.3333 | 62
14.7 | 50.6143 | 51.7 | 53.4 | 54.8 | 55.9571 | 59.3 | 61.06 | 61.8333
14.8 | 50.3333 | 51.45 | 53.1667 | 54.58 | 55.65 | 59 | 60.825 | 61.56
14.9 | 50.14 | 51.16 | 52.8 | 54.2 | 55.4 | 58.82 | 60.54 | 61.35
15 | 49.84 | 50.9 | 52.575 | 53.95 | 55.1667 | 58.475 | 60.2333 | 61.12
15.1 | 49.5143 | 50.6 | 52.325 | 53.75 | 54.875 | 58.3 | 60.02 | 60.8286
15.2 | 49.375 | 50.36 | 52.02 | 53.425 | 54.64 | 57.975 | 59.8167 | 60.675
15.3 | 49.1 | 50 | 51.72 | 53.22 | 54.36 | 57.7 | 59.54 | 60.35
15.4 | 48.78 | 49.78 | 51.5 | 52.9 | 54.1 | 57.425 | 59.2571 | 60.04
15.5 | 48.475 | 49.5 | 51.24 | 52.675 | 53.82 | 57.15 | 59.05 | 59.8667
15.6 | 48.225 | 49.2667 | 50.975 | 52.425 | 53.6 | 56.875 | 58.7667 | 59.7
15.7 | 47.96 | 48.9857 | 50.7 | 52.1167 | 53.26 | 56.6 | 58.58 | 59.3667
15.8 | 47.74 | 48.68 | 50.45 | 51.8 | 53 | 56.44 | 58.275 | 59.1
15.9 | 47.42 | 48.45 | 50.15 | 51.6167 | 52.72 | 56.2 | 58.025 | 58.875
16 | 47.2 | 48.1333 | 49.825 | 51.3 | 52.5 | 55.94 | 57.7667 | 58.58
16.1 | 46.86 | 47.9 | 49.58 | 51.025 | 52.275 | 55.7 | 57.5 | 58.4
16.2 | 46.6571 | 47.66 | 49.325 | 50.72 | 51.94 | 55.4 | 57.26 | 58.1
16.3 | 46.4286 | 47.4143 | 49.075 | 50.475 | 51.675 | 55.1333 | 57.05 | 57.88
16.4 | 46.2 | 47.125 | 48.825 | 50.26 | 51.52 | 54.9 | 56.8 | 57.6286
16.5 | 45.9 | 46.84 | 48.6 | 50.1 | 51.175 | 54.6 | 56.6 | 57.4
16.6 | 45.7333 | 46.675 | 48.35 | 49.75 | 51.02 | 54.3667 | 56.35 | 57.2286
16.7 | 45.4286 | 46.38 | 48.1 | 49.5 | 50.675 | 54.1333 | 56.1 | 56.9
16.8 | 45.1 | 46.125 | 47.8 | 49.22 | 50.42 | 53.95 | 55.86 | 56.66
16.9 | 44.925 | 45.85 | 47.5429 | 48.95 | 50.26 | 53.64 | 55.6 | 56.5
17 | 44.6333 | 45.6333 | 47.3 | 48.76 | 50 | 53.45 | 55.375 | 56.275
17.1 | 44.4 | 45.36 | 47.1 | 48.4 | 49.68 | 53.14 | 55.08 | 56
17.2 | 44.2 | 45.125 | 46.8 | 48.2 | 49.44 | 52.9 | 54.85 | 55.7857
17.3 | 43.98 | 44.88 | 46.5 | 47.9 | 49.1 | 52.72 | 54.6 | 55.62
17.4 | 43.7 | 44.58 | 46.3 | 47.74 | 48.9 | 52.4667 | 54.375 | 55.3571
17.5 | 43.5 | 44.35 | 46.0667 | 47.4875 | 48.64 | 52.25 | 54.125 | 55.12
17.6 | 43.18 | 44.08 | 45.8 | 47.2 | 48.4 | 51.9667 | 53.9333 | 54.875
17.7 | 42.9 | 43.95 | 45.6 | 47.0143 | 48.2 | 51.775 | 53.675 | 54.62
17.8 | 42.7 | 43.6333 | 45.35 | 46.75 | 47.9 | 51.5429 | 53.4 | 54.3875
17.9 | 42.5 | 43.4 | 45.1 | 46.5333 | 47.75 | 51.28 | 53.25 | 54.225
18 | 42.25 | 43.2 | 44.85 | 46.3 | 47.4 | 51.02 | 52.975 | 54
18.1 | 42 | 42.95 | 44.6 | 46.0714 | 47.2 | 50.725 | 52.75 | 53.775
18.2 | 41.76 | 42.7 | 44.35 | 45.8 | 46.95 | 50.575 | 52.56 | 53.5
18.3 | 41.575 | 42.5 | 44.1 | 45.56 | 46.76 | 50.3 | 52.325 | 53.225
18.4 | 41.35 | 42.28 | 43.95 | 45.375 | 46.62 | 50.075 | 52.025 | 53.02
18.5 | 41.06 | 42.0286 | 43.64 | 45.15 | 46.275 | 49.925 | 51.8667 | 52.84
18.6 | 40.95 | 41.875 | 43.4333 | 44.92 | 46.02 | 49.675 | 51.6 | 52.62
18.7 | 40.7 | 41.625 | 43.25 | 44.65 | 45.875 | 49.425 | 51.42 | 52.38
18.8 | 40.44 | 41.375 | 43 | 44.45 | 45.625 | 49.25 | 51.2 | 52.1667
18.9 | 40.2 | 41.125 | 42.8 | 44.2 | 45.375 | 48.925 | 50.9833 | 51.9
19 | 40.0333 | 40.9 | 42.48 | 43.975 | 45.125 | 48.7667 | 50.7 | 51.7667
19.1 | 39.8667 | 40.725 | 42.325 | 43.76 | 44.95 | 48.6 | 50.5143 | 51.5
19.2 | 39.58 | 40.475 | 42.14 | 43.525 | 44.68 | 48.375 | 50.25 | 51.2857
19.3 | 39.4 | 40.24 | 41.86 | 43.32 | 44.44 | 48.125 | 50.125 | 51.0857
19.4 | 39.1667 | 40.075 | 41.64 | 43.1 | 44.3 | 47.875 | 49.8 | 50.825
19.5 | 38.9333 | 39.825 | 41.46 | 42.82 | 44.075 | 47.7 | 49.6143 | 50.7
19.6 | 38.75 | 39.675 | 41.25 | 42.64 | 43.8 | 47.425 | 49.46 | 50.4714
19.7 | 38.58 | 39.44 | 41.125 | 42.4333 | 43.5833 | 47.2 | 49.2333 | 50.2429
19.8 | 38.4 | 39.28 | 40.8 | 42.3 | 43.4 | 47.05 | 48.975 | 49.94
19.9 | 38.2 | 39.04 | 40.6167 | 42 | 43.1857 | 46.78 | 48.7667 | 49.8
20 | 38.05 | 38.8 | 40.4 | 41.75 | 43 | 46.55 | 48.575 | 49.7
20.1 | 37.8 | 38.6 | 40.2167 | 41.54 | 42.7 | 46.4 | 48.35 | 49.4429
20.2 | 37.55 | 38.45 | 39.9833 | 41.3 | 42.52 | 46.2 | 48.12 | 49.25
20.3 | 37.3 | 38.25 | 39.7571 | 41.2 | 42.3 | 45.9 | 48.0333 | 49.0222
20.4 | 37.15 | 38 | 39.5 | 40.95 | 42.1 | 45.75 | 47.78 | 48.7
20.5 | 37 | 37.8 | 39.34 | 40.7 | 41.96 | 45.55 | 47.5714 | 48.5571
20.6 | 36.75 | 37.65 | 39.12 | 40.55 | 41.7 | 45.36 | 47.38 | 48.36
20.7 | 36.5 | 37.4 | 39 | 40.3 | 41.52 | 45.2 | 47.15 | 48.125
20.8 | 36.3 | 37.175 | 38.725 | 40.15 | 41.28 | 44.96 | 46.95 | 47.9
20.9 | 36.125 | 37 | 38.6 | 39.92 | 41.16 | 44.7 | 46.7 | 47.7143
21 | 35.9 | 36.775 | 38.325 | 39.66 | 40.8375 | 44.52 | 46.55 | 47.575
21.1 | 35.72 | 36.6 | 38.2 | 39.52 | 40.62 | 44.2667 | 46.3 | 47.3444
21.2 | 35.5667 | 36.38 | 37.9 | 39.3 | 40.48 | 44.0667 | 46.1 | 47.12
21.3 | 35.42 | 36.225 | 37.72 | 39.125 | 40.32 | 43.9 | 45.84 | 46.9429
21.4 | 35.175 | 36 | 37.54 | 38.9 | 40.0571 | 43.6333 | 45.68 | 46.68
21.5 | 35 | 35.8 | 37.36 | 38.7 | 39.875 | 43.45 | 45.5833 | 46.525
21.6 | 34.825 | 35.625 | 37.125 | 38.575 | 39.7 | 43.3 | 45.3 | 46.325
21.7 | 34.575 | 35.375 | 36.9333 | 38.3 | 39.48 | 43.15 | 45.0857 | 46.1714
21.8 | 34.475 | 35.225 | 36.775 | 38.175 | 39.3 | 42.9 | 44.9 | 46
21.9 | 34.325 | 35.075 | 36.6 | 37.925 | 39.1 | 42.75 | 44.6857 | 45.7667
22 | 34.125 | 34.825 | 36.375 | 37.7 | 38.88 | 42.52 | 44.5 | 45.54
22.1 | 33.9 | 34.7 | 36.3 | 37.5 | 38.68 | 42.26 | 44.3375 | 45.3429
22.2 | 33.75 | 34.5 | 36.05 | 37.38 | 38.45 | 42.1333 | 44.16 | 45.2
22.3 | 33.58 | 34.4 | 35.85 | 37.15 | 38.28 | 41.925 | 43.9667 | 44.975
22.4 | 33.3667 | 34.15 | 35.65 | 36.94 | 38.15 | 41.74 | 43.775 | 44.74
22.5 | 33.125 | 34 | 35.42 | 36.7667 | 37.9333 | 41.5667 | 43.6 | 44.65
22.6 | 32.98 | 33.84 | 35.25 | 36.6 | 37.7375 | 41.28 | 43.3333 | 44.4
22.7 | 32.8 | 33.6 | 35.075 | 36.4 | 37.56 | 41.15 | 43.22 | 44.2
22.8 | 32.625 | 33.4167 | 34.9 | 36.25 | 37.3714 | 41 | 42.98 | 43.96
22.9 | 32.425 | 33.3 | 34.7333 | 36.1 | 37.2 | 40.8 | 42.7714 | 43.8571
23 | 32.28 | 33.04 | 34.5 | 35.86 | 37 | 40.575 | 42.58 | 43.7143
23.1 | 32.1 | 32.9 | 34.34 | 35.7 | 36.8 | 40.4 | 42.4143 | 43.5
23.2 | 31.8833 | 32.75 | 34.18 | 35.5 | 36.62 | 40.275 | 42.28 | 43.275
23.3 | 31.8 | 32.6 | 34.05 | 35.325 | 36.4667 | 40.1 | 42.0286 | 43.15
23.4 | 31.6 | 32.4 | 33.8 | 35.15 | 36.32 | 39.88 | 41.88 | 42.8857
23.5 | 31.45 | 32.22 | 33.7 | 35.0167 | 36.1 | 39.64 | 41.6833 | 42.7
23.6 | 31.3 | 32.1 | 33.55 | 34.82 | 35.9 | 39.5 | 41.5 | 42.575
23.7 | 31.2 | 31.9 | 33.3 | 34.62 | 35.725 | 39.38 | 41.325 | 42.36
23.8 | 30.975 | 31.7 | 33.1857 | 34.4 | 35.525 | 39.15 | 41.175 | 42.1857
23.9 | 30.8 | 31.55 | 33 | 34.3 | 35.375 | 38.9333 | 41 | 42.02
24 | 30.66 | 31.4 | 32.8375 | 34.08 | 35.24 | 38.8 | 40.725 | 41.85
24.1 | 30.4667 | 31.2 | 32.66 | 33.9 | 35.05 | 38.55 | 40.6 | 41.675
24.2 | 30.3 | 31.02 | 32.4857 | 33.72 | 34.875 | 38.45 | 40.4167 | 41.5
24.3 | 30.12 | 30.9 | 32.32 | 33.6 | 34.64 | 38.28 | 40.3 | 41.25
24.4 | 30 | 30.7 | 32.1286 | 33.4 | 34.5 | 38.1143 | 40.0333 | 41.1
24.5 | 29.78 | 30.6 | 31.9 | 33.25 | 34.36 | 37.9 | 39.92 | 40.94
24.6 | 29.65 | 30.4 | 31.8 | 33.125 | 34.2 | 37.7143 | 39.7286 | 40.72
24.7 | 29.475 | 30.275 | 31.6 | 32.8714 | 34 | 37.55 | 39.52 | 40.58
24.8 | 29.375 | 30.1 | 31.4333 | 32.7 | 33.84 | 37.36 | 39.375 | 40.3714
24.9 | 29.22 | 29.9857 | 31.275 | 32.525 | 33.7 | 37.1667 | 39.225 | 40.28
25 | 29.0333 | 29.7667 | 31.125 | 32.4 | 33.55 | 37.0143 | 38.975 | 40.0833
25.1 | 28.9 | 29.625 | 31 | 32.2857 | 33.325 | 36.9 | 38.9 | 39.8571
25.2 | 28.7333 | 29.5 | 30.82 | 32.1 | 33.1667 | 36.7167 | 38.675 | 39.7
25.3 | 28.6 | 29.2857 | 30.64 | 31.9167 | 33.0571 | 36.52 | 38.54 | 39.55
25.4 | 28.4714 | 29.1333 | 30.52 | 31.7 | 32.88 | 36.3286 | 38.3333 | 39.44
25.5 | 28.3 | 29 | 30.4 | 31.6 | 32.7 | 36.22 | 38.2 | 39.2
25.6 | 28.1714 | 28.9 | 30.25 | 31.4 | 32.5 | 35.98 | 38 | 39.0857
25.7 | 28 | 28.7 | 30.08 | 31.26 | 32.35 | 35.8714 | 37.8667 | 38.9429
25.8 | 27.85 | 28.6 | 29.9 | 31.12 | 32.25 | 35.7286 | 37.725 | 38.75
25.9 | 27.74 | 28.4 | 29.74 | 31 | 32.1 | 35.5 | 37.55 | 38.5
26 | 27.5714 | 28.2 | 29.6 | 30.82 | 31.9 | 35.38 | 37.3 | 38.4
26.1 | 27.4 | 28.15 | 29.4 | 30.675 | 31.8 | 35.1833 | 37.2 | 38.2429
26.2 | 27.3333 | 28 | 29.25 | 30.475 | 31.6 | 35.0286 | 37.075 | 38.05
26.3 | 27.15 | 27.85 | 29.2 | 30.325 | 31.425 | 34.9 | 36.825 | 37.9111
26.4 | 27 | 27.7 | 29 | 30.175 | 31.275 | 34.725 | 36.7 | 37.7
26.5 | 26.9 | 27.55 | 28.9 | 30.025 | 31.1 | 34.58 | 36.5 | 37.5571
26.6 | 26.7 | 27.4 | 28.7 | 29.9286 | 31 | 34.4 | 36.4 | 37.46
26.7 | 26.55 | 27.25 | 28.55 | 29.75 | 30.8333 | 34.275 | 36.2167 | 37.3125
26.8 | 26.4 | 27.15 | 28.4 | 29.6 | 30.675 | 34.0833 | 36 | 37.12
26.9 | 26.3 | 27 | 28.3 | 29.4714 | 30.5 | 33.9 | 35.9 | 36.9571
27 | 26.1375 | 26.8 | 28.15 | 29.28 | 30.4 | 33.8 | 35.75 | 36.8
27.1 | 26.06 | 26.7 | 28.02 | 29.1833 | 30.2333 | 33.6 | 35.6333 | 36.6444
27.2 | 25.9 | 26.575 | 27.86 | 29.0286 | 30.1 | 33.44 | 35.42 | 36.5
27.3 | 25.7333 | 26.4 | 27.72 | 28.84 | 29.9 | 33.36 | 35.3 | 36.3333
27.4 | 25.6 | 26.3 | 27.575 | 28.76 | 29.8 | 33.14 | 35.175 | 36.12
27.5 | 25.5 | 26.175 | 27.42 | 28.6 | 29.6 | 33.075 | 34.925 | 36.0333
27.6 | 25.375 | 26.025 | 27.225 | 28.4833 | 29.5 | 32.9 | 34.825 | 35.9
27.7 | 25.225 | 25.875 | 27.1 | 28.2857 | 29.3 | 32.7 | 34.68 | 35.7333
27.8 | 25.1 | 25.725 | 27 | 28.2 | 29.16 | 32.575 | 34.4833 | 35.525
27.9 | 24.9833 | 25.6 | 26.8667 | 28.04 | 29.1 | 32.4 | 34.3 | 35.4429
28 | 24.9 | 25.5 | 26.7 | 27.875 | 28.9 | 32.3 | 34.1857 | 35.18
28.1 | 24.7 | 25.32 | 26.6 | 27.7 | 28.775 | 32.175 | 34.0429 | 35.0833
28.2 | 24.6 | 25.2714 | 26.4714 | 27.6 | 28.6 | 32 | 33.94 | 34.98
28.3 | 24.5 | 25.1 | 26.28 | 27.46 | 28.5 | 31.8 | 33.7429 | 34.775
28.4 | 24.3 | 24.95 | 26.125 | 27.3 | 28.3 | 31.66 | 33.54 | 34.6
28.5 | 24.2 | 24.825 | 26.04 | 27.2 | 28.2 | 31.5125 | 33.45 | 34.48
28.6 | 24.1 | 24.75 | 25.9 | 27.075 | 28 | 31.42 | 33.3 | 34.3
28.7 | 24 | 24.6 | 25.8 | 26.96 | 27.925 | 31.2571 | 33.1 | 34.1889
28.8 | 23.9 | 24.45 | 25.66 | 26.8 | 27.8 | 31.16 | 33 | 34.06
28.9 | 23.7 | 24.3 | 25.5667 | 26.65 | 27.7 | 30.9444 | 32.9 | 33.9
29 | 23.6 | 24.2 | 25.4143 | 26.5 | 27.52 | 30.82 | 32.76 | 33.8143
29.1 | 23.5 | 24.1 | 25.3 | 26.36 | 27.4143 | 30.6571 | 32.6 | 33.6
29.2 | 23.3 | 23.95 | 25.175 | 26.28 | 27.275 | 30.5 | 32.4 | 33.5
29.3 | 23.2 | 23.8 | 24.9714 | 26.16 | 27.1125 | 30.4 | 32.2667 | 33.34
29.4 | 23.1 | 23.7 | 24.9 | 26 | 27 | 30.275 | 32.175 | 33.16
29.5 | 23 | 23.6 | 24.8 | 25.875 | 26.9 | 30.1714 | 32.0125 | 33.0667
29.6 | 22.9 | 23.45 | 24.64 | 25.7125 | 26.72 | 29.98 | 31.9 | 32.925
29.7 | 22.8 | 23.3 | 24.5714 | 25.62 | 26.6 | 29.825 | 31.7143 | 32.8
29.8 | 22.65 | 23.2 | 24.4 | 25.5 | 26.5 | 29.7 | 31.6 | 32.6
29.9 | 22.5 | 23.1 | 24.2833 | 25.4 | 26.35 | 29.5333 | 31.4714 | 32.475
30 | 22.4 | 23 | 24.2 | 25.275 | 26.2 | 29.425 | 31.3 | 32.3

## Open Field Profiles at 2cm
Lateral Profiles
Open field
SSD=100cm
Depth=2cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.3 | [7 empty] | 2.8
-28.2 | [7 empty] | 2.8
-28.1 | [7 empty] | 2.8429
-28 | [7 empty] | 2.9
-27.9 | [7 empty] | 2.9
-27.8 | [7 empty] | 2.9
-27.7 | [7 empty] | 3
-27.6 | [7 empty] | 3
-27.5 | [7 empty] | 3.0714
-27.4 | [7 empty] | 3.1
-27.3 | [7 empty] | 3.1083
-27.2 | [7 empty] | 3.1917
-27.1 | [7 empty] | 3.2
-27 | [7 empty] | 3.2286
-26.9 | [7 empty] | 3.3
-26.8 | [7 empty] | 3.3222
-26.7 | [7 empty] | 3.4
-26.6 | [7 empty] | 3.4
-26.5 | [7 empty] | 3.5
-26.4 | [7 empty] | 3.5125
-26.3 | [7 empty] | 3.6
-26.2 | [7 empty] | 3.6
-26.1 | [7 empty] | 3.7
-26 | [7 empty] | 3.7333
-25.9 | [7 empty] | 3.8
-25.8 | [7 empty] | 3.8308
-25.7 | [7 empty] | 3.9
-25.6 | [7 empty] | 3.95
-25.5 | [7 empty] | 4.0167
-25.4 | [7 empty] | 4.1
-25.3 | [7 empty] | 4.2
-25.2 | [7 empty] | 4.2833
-25.1 | [7 empty] | 4.3
-25 | [7 empty] | 4.4333
-24.9 | [7 empty] | 4.5143
-24.8 | [7 empty] | 4.6
-24.7 | [7 empty] | 4.7
-24.6 | [7 empty] | 4.8
-24.5 | [7 empty] | 4.9125
-24.4 | [7 empty] | 5
-24.3 | [7 empty] | 5.075
-24.2 | [7 empty] | 5.2167
-24.1 | [7 empty] | 5.3
-24 | [7 empty] | 5.4
-23.9 | [7 empty] | 5.6
-23.8 | [7 empty] | 5.7
-23.7 | [7 empty] | 5.8
-23.6 | [7 empty] | 6
-23.5 | [7 empty] | 6.1429
-23.4 | [7 empty] | 6.3
-23.3 | [7 empty] | 6.4
-23.2 | [7 empty] | 6.6
-23.1 | [7 empty] | 6.7
-23 | [7 empty] | 6.9
-22.9 | [7 empty] | 7.1
-22.8 | [7 empty] | 7.2
-22.7 | [7 empty] | 7.4
-22.6 | [7 empty] | 7.6
-22.5 | [7 empty] | 7.7833
-22.4 | [7 empty] | 8
-22.3 | [7 empty] | 8.1
-22.2 | [7 empty] | 8.2833
-22.1 | [7 empty] | 8.58
-22 | [7 empty] | 8.78
-21.9 | [7 empty] | 8.98
-21.8 | [7 empty] | 9.26
-21.7 | [7 empty] | 9.48
-21.6 | [7 empty] | 9.78
-21.5 | [7 empty] | 10.16
-21.4 | [7 empty] | 10.56
-21.3 | [7 empty] | 11.04
-21.2 | [7 empty] | 11.72
-21.1 | [7 empty] | 12.52
-21 | [7 empty] | 13.76
-20.9 | [7 empty] | 15.68
-20.8 | [7 empty] | 19
-20.7 | [7 empty] | 24.46
-20.6 | [7 empty] | 32.78
-20.5 | [7 empty] | 43.0857
-20.4 | [7 empty] | 54.9714
-20.3 | [7 empty] | 66.4625
-20.2 | [6 empty] | 4.0667 | 77.3714
-20.1 | [6 empty] | 4.2 | 87.1
-20 | [6 empty] | 4.3 | 94.4833
-19.9 | [6 empty] | 4.35 | 98.8857
-19.8 | [6 empty] | 4.4 | 101.2667
-19.7 | [6 empty] | 4.5 | 102.8143
-19.6 | [6 empty] | 4.6 | 103.8625
-19.5 | [6 empty] | 4.7 | 104.4714
-19.4 | [6 empty] | 4.8 | 104.925
-19.3 | [6 empty] | 4.85 | 105.2667
-19.2 | [6 empty] | 4.94 | 105.5714
-19.1 | [6 empty] | 5.075 | 105.76
-19 | [6 empty] | 5.16 | 105.96
-18.9 | [6 empty] | 5.3 | 106.16
-18.8 | [6 empty] | 5.4 | 106.25
-18.7 | [6 empty] | 5.5 | 106.4
-18.6 | [6 empty] | 5.6 | 106.5
-18.5 | [6 empty] | 5.7333 | 106.5375
-18.4 | [6 empty] | 5.82 | 106.6
-18.3 | [6 empty] | 6 | 106.7
-18.2 | [6 empty] | 6.12 | 106.7333
-18.1 | [6 empty] | 6.22 | 106.8
-18 | [6 empty] | 6.4 | 106.8
-17.9 | [6 empty] | 6.5 | 106.825
-17.8 | [6 empty] | 6.68 | 106.9
-17.7 | [6 empty] | 6.8 | 106.9
-17.6 | [6 empty] | 7 | 106.9
-17.5 | [6 empty] | 7.125 | 106.9
-17.4 | [6 empty] | 7.275 | 106.9
-17.3 | [6 empty] | 7.48 | 106.9
-17.2 | [6 empty] | 7.6 | 106.9
-17.1 | [6 empty] | 7.825 | 106.9
-17 | [6 empty] | 7.98 | 106.8091
-16.9 | [6 empty] | 8.1667 | 106.8
-16.8 | [6 empty] | 8.375 | 106.8
-16.7 | [6 empty] | 8.5833 | 106.8
-16.6 | [6 empty] | 8.84 | 106.8
-16.5 | [6 empty] | 9.1333 | 106.8
-16.4 | [6 empty] | 9.5 | 106.7375
-16.3 | [6 empty] | 9.8667 | 106.7
-16.2 | [6 empty] | 10.38 | 106.6714
-16.1 | [6 empty] | 11.125 | 106.6
-16 | [6 empty] | 12.02 | 106.6
-15.9 | [6 empty] | 13.25 | 106.6
-15.8 | [6 empty] | 15.35 | 106.5429
-15.7 | [6 empty] | 19.02 | 106.5
-15.6 | [6 empty] | 24.78 | 106.4857
-15.5 | [6 empty] | 33.4286 | 106.4
-15.4 | [6 empty] | 43.76 | 106.4
-15.3 | [6 empty] | 56.0333 | 106.3857
-15.2 | [6 empty] | 66.78 | 106.3
-15.1 | [5 empty] | 2.9 | 77.5667 | 106.3
-15 | [5 empty] | 3 | 86.7 | 106.2875
-14.9 | [5 empty] | 3.04 | 93.1714 | 106.2
-14.8 | [5 empty] | 3.1 | 97 | 106.2
-14.7 | [5 empty] | 3.18 | 99.34 | 106.2
-14.6 | [5 empty] | 3.2 | 100.68 | 106.1
-14.5 | [5 empty] | 3.3 | 101.52 | 106.1
-14.4 | [5 empty] | 3.4 | 102.1 | 106.1
-14.3 | [5 empty] | 3.425 | 102.5 | 106.0875
-14.2 | [5 empty] | 3.5 | 102.8667 | 106
-14.1 | [5 empty] | 3.6 | 103.075 | 105.925
-14 | [5 empty] | 3.7 | 103.2 | 105.9
-13.9 | [5 empty] | 3.775 | 103.3 | 105.9
-13.8 | [5 empty] | 3.8667 | 103.44 | 105.8
-13.7 | [5 empty] | 3.9 | 103.5 | 105.7625
-13.6 | [5 empty] | 4 | 103.6 | 105.7
-13.5 | [5 empty] | 4.1 | 103.625 | 105.7
-13.4 | [5 empty] | 4.2 | 103.7 | 105.6
-13.3 | [5 empty] | 4.3 | 103.7 | 105.6
-13.2 | [5 empty] | 4.4 | 103.7 | 105.6
-13.1 | [5 empty] | 4.575 | 103.8 | 105.5714
-13 | [5 empty] | 4.6667 | 103.8 | 105.5
-12.9 | [5 empty] | 4.8 | 103.8 | 105.5
-12.8 | [5 empty] | 4.9 | 103.8 | 105.4
-12.7 | [5 empty] | 5 | 103.8 | 105.3429
-12.6 | [5 empty] | 5.2 | 103.8 | 105.3
-12.5 | [5 empty] | 5.3 | 103.8 | 105.275
-12.4 | [5 empty] | 5.475 | 103.8 | 105.2
-12.3 | [5 empty] | 5.6 | 103.8 | 105.1
-12.2 | [5 empty] | 5.7667 | 103.8 | 105.1
-12.1 | [5 empty] | 5.9 | 103.7 | 105.0571
-12 | [5 empty] | 6.1 | 103.7 | 105
-11.9 | [5 empty] | 6.3 | 103.7 | 104.9
-11.8 | [5 empty] | 6.5 | 103.7 | 104.8333
-11.7 | [5 empty] | 6.65 | 103.7 | 104.8
-11.6 | [5 empty] | 6.9 | 103.7 | 104.8
-11.5 | [5 empty] | 7.14 | 103.66 | 104.7
-11.4 | [5 empty] | 7.44 | 103.6 | 104.7
-11.3 | [5 empty] | 7.8 | 103.6 | 104.6429
-11.2 | [5 empty] | 8.2333 | 103.6 | 104.6
-11.1 | [5 empty] | 8.7333 | 103.6 | 104.6
-11 | [5 empty] | 9.4 | 103.6 | 104.5125
-10.9 | [5 empty] | 10.3333 | 103.5833 | 104.5
-10.8 | [5 empty] | 11.7333 | 103.5 | 104.5
-10.7 | [5 empty] | 13.7 | 103.5 | 104.4125
-10.6 | [5 empty] | 17.025 | 103.5 | 104.3
-10.5 | [5 empty] | 22.9 | 103.5 | 104.3
-10.4 | [5 empty] | 31.1 | 103.5 | 104.2125
-10.3 | [5 empty] | 40.8 | 103.4 | 104.2
-10.2 | [5 empty] | 52.3333 | 103.4 | 104.15
-10.1 | [5 empty] | 63.9 | 103.4 | 104.1
-10 |  |  |  |  | 1.5 | 75.0667 | 103.4 | 104
-9.9 |  |  |  |  | 1.6 | 83.65 | 103.3 | 104
-9.8 |  |  |  |  | 1.6 | 90.4 | 103.3 | 103.925
-9.7 |  |  |  |  | 1.6 | 94.3 | 103.3 | 103.9
-9.6 |  |  |  |  | 1.7 | 96.6 | 103.2 | 103.87
-9.5 |  |  |  |  | 1.7 | 97.925 | 103.2 | 103.7769
-9.4 |  |  |  |  | 1.8 | 98.825 | 103.2 | 103.7
-9.3 |  |  |  |  | 1.8 | 99.4667 | 103.2 | 103.7
-9.2 |  |  |  |  | 1.8667 | 99.9 | 103.1 | 103.6615
-9.1 |  |  |  |  | 1.9 | 100.2333 | 103.1 | 103.6
-9 |  |  |  | 1.2 | 1.9 | 100.48 | 103.08 | 103.6
-8.9 |  |  |  | 1.3 | 2 | 100.65 | 103 | 103.5375
-8.8 |  |  |  | 1.3 | 2 | 100.9 | 103 | 103.5
-8.7 |  |  |  | 1.3 | 2.1 | 101 | 102.925 | 103.47
-8.6 |  |  |  | 1.3 | 2.1 | 101.0333 | 102.9 | 103.3667
-8.5 |  |  |  | 1.4 | 2.2 | 101.1 | 102.9 | 103.3
-8.4 |  |  |  | 1.4 | 2.2333 | 101.2 | 102.9 | 103.2786
-8.3 |  |  |  | 1.4 | 2.3 | 101.2 | 102.8 | 103.2071
-8.2 |  |  |  | 1.4 | 2.4 | 101.3 | 102.8 | 103.1182
-8.1 |  |  |  | 1.5 | 2.425 | 101.3 | 102.74 | 103.1
-8 |  |  | 0.9 | 1.5 | 2.5 | 101.3 | 102.7 | 103
-7.9 |  |  | 0.9 | 1.525 | 2.6 | 101.3 | 102.7 | 103
-7.8 |  |  | 0.95 | 1.6 | 2.7 | 101.38 | 102.6 | 103
-7.7 |  |  | 1 | 1.6 | 2.8 | 101.4 | 102.6 | 102.9
-7.6 |  |  | 1 | 1.7 | 2.9 | 101.4 | 102.58 | 102.9
-7.5 |  |  | 1 | 1.7 | 3 | 101.4 | 102.5 | 102.8286
-7.4 |  |  | 1 | 1.8 | 3.1 | 101.4 | 102.5 | 102.8
-7.3 |  |  | 1.1 | 1.8 | 3.2 | 101.4 | 102.5 | 102.7714
-7.2 |  |  | 1.1 | 1.9 | 3.3 | 101.4 | 102.4 | 102.7
-7.1 |  |  | 1.1 | 1.9 | 3.4 | 101.4 | 102.4 | 102.7
-7 |  | 0.6 | 1.1 | 2 | 3.5333 | 101.4 | 102.35 | 102.6
-6.9 |  | 0.7 | 1.2 | 2.1 | 3.7 | 101.4 | 102.3 | 102.6
-6.8 |  | 0.7 | 1.2 | 2.2 | 3.825 | 101.4 | 102.3 | 102.5
-6.7 |  | 0.7 | 1.2 | 2.2 | 4 | 101.4 | 102.2 | 102.4714
-6.6 |  | 0.7 | 1.3 | 2.3 | 4.14 | 101.4 | 102.2 | 102.4
-6.5 | 0.5 | 0.7 | 1.3 | 2.4 | 4.35 | 101.36 | 102.125 | 102.4
-6.4 | 0.5 | 0.7 | 1.3 | 2.5 | 4.6 | 101.3 | 102.1 | 102.4
-6.3 | 0.5 | 0.7 | 1.4 | 2.575 | 4.8333 | 101.3 | 102.1 | 102.3
-6.2 | 0.5 | 0.7 | 1.4 | 2.675 | 5.175 | 101.3 | 102.075 | 102.2667
-6.1 | 0.5 | 0.7 | 1.5 | 2.8 | 5.65 | 101.3 | 102 | 102.2
-6 | 0.5 | 0.8 | 1.5 | 2.9 | 6.15 | 101.3 | 102 | 102.15
-5.9 | 0.6 | 0.8 | 1.6 | 3 | 6.775 | 101.3 | 102 | 102.1
-5.8 | 0.6 | 0.8 | 1.6 | 3.1333 | 7.725 | 101.3 | 101.9 | 102.0286
-5.7 | 0.6 | 0.8 | 1.7 | 3.3 | 9.025 | 101.3 | 101.86 | 102
-5.6 | 0.6 | 0.8 | 1.7 | 3.5 | 11 | 101.3 | 101.8 | 102
-5.5 | 0.6 | 0.9 | 1.8 | 3.6 | 14.44 | 101.3 | 101.8 | 101.9
-5.4 | 0.6 | 0.9 | 1.9 | 3.86 | 19.925 | 101.3 | 101.8 | 101.9
-5.3 | 0.6 | 0.9 | 1.9 | 4.04 | 27.9 | 101.2 | 101.7 | 101.8875
-5.2 | 0.6 | 0.9 | 2 | 4.3333 | 36.8 | 101.2 | 101.7 | 101.8
-5.1 | 0.66667 | 0.95 | 2.1 | 4.7 | 48.56 | 101.2 | 101.6375 | 101.8
-5 | 0.7 | 1 | 2.2 | 5.2333 | 59.9 | 101.14 | 101.6 | 101.8
-4.9 | 0.7 | 1 | 2.3 | 5.8 | 71.06 | 101.1 | 101.6 | 101.7
-4.8 | 0.7 | 1.1 | 2.4 | 6.6 | 80.4 | 101.1 | 101.5 | 101.6
-4.7 | 0.7 | 1.1 | 2.5 | 7.7 | 87.1 | 101.05 | 101.5 | 101.6
-4.6 | 0.7 | 1.2 | 2.7 | 9.4 | 91.5 | 101 | 101.5 | 101.525
-4.5 | 0.8 | 1.2 | 2.8 | 12.2 | 94.4 | 101 | 101.4571 | 101.5
-4.4 | 0.8 | 1.2 | 3 | 16.9 | 96.0333 | 101 | 101.4 | 101.5
-4.3 | 0.8 | 1.3 | 3.2 | 23.62 | 97 | 101 | 101.4 | 101.4
-4.2 | 0.8 | 1.3 | 3.5 | 33.56 | 97.75 | 101 | 101.3 | 101.3714
-4.1 | 0.85 | 1.4 | 3.7 | 44.24 | 98.325 | 101 | 101.3 | 101.3
-4 | 0.9 | 1.5 | 4.1667 | 55.45 | 98.6333 | 100.9 | 101.2 | 101.3
-3.9 | 0.9 | 1.5 | 4.65 | 67.35 | 98.94 | 100.9 | 101.2 | 101.2
-3.8 | 1 | 1.6 | 5.5 | 77.2667 | 99.22 | 100.9 | 101.1333 | 101.1714
-3.7 | 1 | 1.7 | 6.5 | 85.44 | 99.3857 | 100.8 | 101.1 | 101.1
-3.6 | 1.1 | 1.8 | 8 | 90.6667 | 99.5 | 100.8 | 101.0667 | 101.1
-3.5 | 1.1 | 1.9333 | 10.75 | 93.5667 | 99.6 | 100.8 | 101 | 101
-3.4 | 1.2 | 2.1 | 14.5 | 95.5667 | 99.7 | 100.8 | 101 | 101
-3.3 | 1.2 | 2.2 | 21.8 | 96.8 | 99.8 | 100.7 | 101 | 100.9125
-3.2 | 1.3 | 2.5 | 30.8 | 97.6 | 99.8 | 100.7 | 100.96 | 100.9
-3.1 | 1.4 | 2.7 | 41.4 | 98.1 | 99.9 | 100.6 | 100.9 | 100.86
-3 | 1.5 | 3.1 | 53.1 | 98.5 | 99.9 | 100.6 | 100.9 | 100.8
-2.9 | 1.6 | 3.6 | 64.6 | 98.8 | 99.9 | 100.6 | 100.8 | 100.8
-2.8 | 1.7 | 4.3 | 75.4 | 99.0286 | 100 | 100.6 | 100.7 | 100.7833
-2.7 | 1.9 | 5.2 | 83.9 | 99.28 | 100 | 100.6 | 100.7 | 100.7
-2.6 | 2.1 | 6.6333 | 89.5 | 99.375 | 100.0333 | 100.5 | 100.7 | 100.7
-2.5 | 2.45 | 8.9 | 93.1333 | 99.5 | 100.1 | 100.5 | 100.6 | 100.6
-2.4 | 2.9 | 12.4 | 95.3 | 99.6 | 100.1 | 100.4 | 100.6 | 100.6
-2.3 | 3.5 | 18.2 | 96.5 | 99.6667 | 100.1 | 100.4 | 100.5429 | 100.5286
-2.2 | 4.4 | 27.2667 | 97.6 | 99.7 | 100.1 | 100.4 | 100.5 | 100.5
-2.1 | 5.85 | 37.2 | 98.2 | 99.8 | 100.1 | 100.4 | 100.5 | 100.4625
-2 | 7.7667 | 49.2 | 98.6 | 99.8 | 100.1 | 100.4 | 100.4833 | 100.4
-1.9 | 11.2 | 61.5 | 98.9667 | 99.8333 | 100.1 | 100.3 | 100.4 | 100.4
-1.8 | 16.9 | 72.7667 | 99.1 | 99.9 | 100.1 | 100.3 | 100.4 | 100.3833
-1.7 | 25.2 | 81.6 | 99.3 | 99.975 | 100.1 | 100.3 | 100.4 | 100.3
-1.6 | 35.8 | 88.7 | 99.4 | 100 | 100.1 | 100.2 | 100.32 | 100.3
-1.5 | 47.7333 | 93 | 99.5 | 100 | 100.1 | 100.2 | 100.3 | 100.3
-1.4 | 60.1 | 95.6 | 99.6 | 100 | 100.1 | 100.2 | 100.3 | 100.26
-1.3 | 72.1 | 97 | 99.7 | 100 | 100.1 | 100.2 | 100.3 | 100.2
-1.2 | 81.05 | 98 | 99.7333 | 100 | 100.1 | 100.1 | 100.2667 | 100.2
-1.1 | 88.2 | 98.8 | 99.8 | 100 | 100.1 | 100.1 | 100.2 | 100.16
-1 | 93.1 | 99.1 | 99.9 | 100 | 100.1 | 100.1 | 100.2 | 100.1
-0.9 | 95.5 | 99.4 | 99.9 | 100 | 100.1 | 100.1 | 100.2 | 100.1
-0.8 | 97.3 | 99.6 | 99.9 | 100 | 100.1 | 100.1 | 100.1 | 100.1
-0.7 | 98.2 | 99.7 | 99.9 | 100 | 100.1 | 100.1 | 100.1 | 100
-0.6 | 98.9333 | 99.9 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100
-0.5 | 99.5 | 100 | 100 | 100 | 100 | 100.025 | 100.1 | 100
-0.4 | 99.7 | 100 | 100 | 100 | 100 | 100 | 100.1 | 100
-0.3 | 99.8 | 100 | 100 | 100 | 100 | 100 | 100 | 100
-0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.8 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.4 | 99.7 | 100 | 100 | 100 | 100 | 100 | 100.1 | 100
0.5 | 99.5 | 100 | 100 | 100 | 100 | 100.025 | 100.1 | 100
0.6 | 98.95 | 99.9 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100
0.7 | 98.2 | 99.7 | 99.9 | 100 | 100.1 | 100.1 | 100.1 | 100
0.8 | 97.3 | 99.6 | 99.9 | 100 | 100.1 | 100.1 | 100.12 | 100.1
0.9 | 95.5 | 99.3667 | 99.9 | 100 | 100.1 | 100.1 | 100.2 | 100.1
1 | 93.05 | 99.1 | 99.9 | 100 | 100.1 | 100.1 | 100.2 | 100.1
1.1 | 88.4667 | 98.75 | 99.8 | 100 | 100.1 | 100.1 | 100.2 | 100.16
1.2 | 81 | 98 | 99.8 | 100 | 100.1 | 100.1 | 100.3 | 100.2
1.3 | 71.55 | 97 | 99.7 | 100 | 100.1 | 100.2 | 100.3 | 100.2
1.4 | 60.2 | 95.5 | 99.6 | 100 | 100.1 | 100.2 | 100.3 | 100.26
1.5 | 48.1 | 93 | 99.5 | 100 | 100.1 | 100.2 | 100.3 | 100.3
1.6 | 35.8 | 88.6 | 99.4 | 100 | 100.1 | 100.2 | 100.3 | 100.3
1.7 | 25.2 | 82 | 99.3 | 99.975 | 100.1 | 100.3 | 100.4 | 100.3
1.8 | 17.4 | 72.8 | 99.1 | 99.9 | 100.1 | 100.3 | 100.4 | 100.3833
1.9 | 11.2 | 61.6 | 98.9333 | 99.8333 | 100.1 | 100.3 | 100.4 | 100.4
2 | 7.75 | 49.2 | 98.6 | 99.8 | 100.1 | 100.4 | 100.5 | 100.4
2.1 | 5.8333 | 38.1 | 98.2 | 99.8 | 100.1 | 100.4 | 100.5 | 100.4625
2.2 | 4.4 | 27 | 97.6333 | 99.7 | 100.1 | 100.4 | 100.5 | 100.5
2.3 | 3.5 | 18.8 | 96.6 | 99.65 | 100.1 | 100.4 | 100.525 | 100.5286
2.4 | 2.9 | 12.5 | 95.25 | 99.6 | 100.1 | 100.4 | 100.6 | 100.6
2.5 | 2.4 | 8.9 | 93.1333 | 99.5 | 100.1 | 100.5 | 100.6 | 100.6
2.6 | 2.15 | 6.6 | 89.6 | 99.4 | 100 | 100.5 | 100.7 | 100.7
2.7 | 1.9 | 5.2 | 83.65 | 99.28 | 100 | 100.6 | 100.7 | 100.7
2.8 | 1.7 | 4.2667 | 75.7 | 99 | 100 | 100.6 | 100.7167 | 100.7833
2.9 | 1.6 | 3.6 | 65.1 | 98.8 | 99.9 | 100.6 | 100.8 | 100.8
3 | 1.5 | 3.1 | 53.7 | 98.5 | 99.9 | 100.6 | 100.8833 | 100.8
3.1 | 1.4 | 2.7 | 41.6 | 98.1 | 99.9 | 100.625 | 100.9 | 100.86
3.2 | 1.3 | 2.5 | 30.9 | 97.6 | 99.8 | 100.7 | 100.9333 | 100.9
3.3 | 1.2 | 2.2 | 21.5 | 96.8 | 99.8 | 100.7 | 101 | 100.9125
3.4 | 1.2 | 2.1 | 14.6 | 95.6 | 99.7 | 100.8 | 101 | 101
3.5 | 1.1 | 1.9333 | 10.7 | 93.6 | 99.6 | 100.8 | 101 | 101
3.6 | 1.1 | 1.8 | 8 | 90.4 | 99.5 | 100.8 | 101.1 | 101.1
3.7 | 1 | 1.7 | 6.5 | 85.24 | 99.3833 | 100.8 | 101.1 | 101.1
3.8 | 1 | 1.6 | 5.5 | 77.3333 | 99.2167 | 100.9 | 101.1 | 101.1714
3.9 | 0.9 | 1.5 | 4.65 | 67.1667 | 98.94 | 100.9 | 101.2 | 101.2
4 | 0.9 | 1.5 | 4.15 | 56.1 | 98.6333 | 100.9 | 101.24 | 101.3
4.1 | 0.85 | 1.4 | 3.7 | 44.425 | 98.325 | 101 | 101.3 | 101.3
4.2 | 0.8 | 1.3 | 3.5 | 33.375 | 97.7333 | 101 | 101.34 | 101.3714
4.3 | 0.8 | 1.3 | 3.2 | 23.775 | 97.025 | 101 | 101.4 | 101.4
4.4 | 0.8 | 1.2 | 3 | 16.625 | 96.025 | 101 | 101.4 | 101.5
4.5 | 0.8 | 1.2 | 2.8 | 12.3 | 94.4 | 101 | 101.4333 | 101.5
4.6 | 0.7 | 1.2 | 2.7 | 9.4 | 91.7 | 101 | 101.5 | 101.525
4.7 | 0.7 | 1.1 | 2.5 | 7.8 | 87.04 | 101.05 | 101.5 | 101.6
4.8 | 0.7 | 1.1 | 2.4 | 6.6 | 80.6 | 101.1 | 101.5 | 101.6
4.9 | 0.7 | 1 | 2.3 | 5.7333 | 70.7 | 101.1 | 101.6 | 101.7
5 | 0.7 | 1 | 2.2 | 5.2 | 60 | 101.14 | 101.6 | 101.8
5.1 | 0.66667 | 1 | 2.1 | 4.7 | 48.86 | 101.2 | 101.6143 | 101.8
5.2 | 0.6 | 0.9 | 2 | 4.35 | 36.9 | 101.2 | 101.7 | 101.8
5.3 | 0.6 | 0.9 | 1.9 | 4.05 | 27.1 | 101.2 | 101.7 | 101.8875
5.4 | 0.6 | 0.9 | 1.9 | 3.86 | 19.52 | 101.3 | 101.8 | 101.9
5.5 | 0.6 | 0.9 | 1.8 | 3.6 | 14.1 | 101.3 | 101.8 | 101.9
5.6 | 0.6 | 0.8 | 1.7 | 3.5 | 11 | 101.3 | 101.8 | 102
5.7 | 0.6 | 0.8 | 1.7 | 3.3 | 9.05 | 101.3 | 101.84 | 102
5.8 | 0.6 | 0.8 | 1.6 | 3.1333 | 7.725 | 101.3 | 101.9 | 102.0286
5.9 | 0.6 | 0.8 | 1.6 | 3 | 6.775 | 101.3 | 101.96 | 102.1
6 | 0.5 | 0.8 | 1.5 | 2.9 | 6.1 | 101.3 | 102 | 102.15
6.1 | 0.5 | 0.7 | 1.5 | 2.8 | 5.64 | 101.3 | 102 | 102.2
6.2 | 0.5 | 0.7 | 1.4 | 2.675 | 5.2333 | 101.3 | 102.1 | 102.2667
6.3 | 0.5 | 0.7 | 1.4 | 2.575 | 4.9 | 101.3 | 102.1 | 102.3
6.4 | 0.5 | 0.7 | 1.3 | 2.5 | 4.6 | 101.3 | 102.1 | 102.4
6.5 | 0.5 | 0.7 | 1.3 | 2.4 | 4.35 | 101.36 | 102.1 | 102.4
6.6 |  | 0.7 | 1.3 | 2.3 | 4.1667 | 101.4 | 102.2 | 102.4
6.7 |  | 0.7 | 1.2 | 2.2 | 4 | 101.4 | 102.2 | 102.4714
6.8 |  | 0.7 | 1.2 | 2.2 | 3.82 | 101.4 | 102.275 | 102.5
6.9 |  | 0.7 | 1.2 | 2.1 | 3.7 | 101.4 | 102.3 | 102.6
7 |  | 0.6 | 1.1 | 2 | 3.525 | 101.4 | 102.3333 | 102.6
7.1 |  |  | 1.1 | 1.9 | 3.4 | 101.4 | 102.4 | 102.7
7.2 |  |  | 1.1 | 1.9 | 3.3 | 101.4 | 102.4 | 102.7
7.3 |  |  | 1.1 | 1.8 | 3.2 | 101.4 | 102.48 | 102.7714
7.4 |  |  | 1 | 1.8 | 3.1 | 101.4 | 102.5 | 102.8
7.5 |  |  | 1 | 1.7 | 3 | 101.4 | 102.5 | 102.8286
7.6 |  |  | 1 | 1.7 | 2.9 | 101.4 | 102.6 | 102.9
7.7 |  |  | 1 | 1.6 | 2.8 | 101.4 | 102.6 | 102.9
7.8 |  |  | 0.95 | 1.6 | 2.7 | 101.4 | 102.6 | 103
7.9 |  |  | 0.9 | 1.5 | 2.6 | 101.3 | 102.7 | 103
8 |  |  | 0.9 | 1.5 | 2.5 | 101.3 | 102.7 | 103
8.1 |  |  |  | 1.5 | 2.5 | 101.3 | 102.8 | 103.1
8.2 |  |  |  | 1.4 | 2.4 | 101.3 | 102.8 | 103.1182
8.3 |  |  |  | 1.4 | 2.3 | 101.2 | 102.8 | 103.2071
8.4 |  |  |  | 1.4 | 2.225 | 101.2 | 102.9 | 103.2786
8.5 |  |  |  | 1.38 | 2.2 | 101.1 | 102.9 | 103.3
8.6 |  |  |  | 1.3 | 2.1 | 101.0333 | 102.9 | 103.3667
8.7 |  |  |  | 1.3 | 2.1 | 100.9667 | 103 | 103.47
8.8 |  |  |  | 1.3 | 2 | 100.875 | 103 | 103.5
8.9 |  |  |  | 1.3 | 2 | 100.7 | 103 | 103.5375
9 |  |  |  | 1.24 | 1.9 | 100.475 | 103.05 | 103.6
9.1 |  |  |  |  | 1.9 | 100.225 | 103.1 | 103.6
9.2 |  |  |  |  | 1.8667 | 99.9 | 103.1 | 103.6615
9.3 |  |  |  |  | 1.8 | 99.4667 | 103.2 | 103.7
9.4 |  |  |  |  | 1.8 | 98.825 | 103.2 | 103.7
9.5 |  |  |  |  | 1.7 | 98 | 103.2 | 103.7769
9.6 |  |  |  |  | 1.7 | 96.6 | 103.26 | 103.87
9.7 |  |  |  |  | 1.6 | 94.2333 | 103.3 | 103.9
9.8 |  |  |  |  | 1.6 | 90.2667 | 103.3 | 103.925
9.9 |  |  |  |  | 1.6 | 83.7 | 103.3 | 104
10 |  |  |  |  | 1.5 | 75.0667 | 103.4 | 104
10.1 | [5 empty] | 63.9 | 103.4 | 104.1
10.2 | [5 empty] | 52.7 | 103.4 | 104.15
10.3 | [5 empty] | 40.8 | 103.4 | 104.2
10.4 | [5 empty] | 30.7 | 103.5 | 104.2125
10.5 | [5 empty] | 22.525 | 103.5 | 104.3
10.6 | [5 empty] | 17.25 | 103.5 | 104.3
10.7 | [5 empty] | 13.8 | 103.5 | 104.4125
10.8 | [5 empty] | 11.7333 | 103.5 | 104.5
10.9 | [5 empty] | 10.3333 | 103.6 | 104.5
11 | [5 empty] | 9.4 | 103.6 | 104.5125
11.1 | [5 empty] | 8.7333 | 103.6 | 104.6
11.2 | [5 empty] | 8.2333 | 103.6 | 104.6
11.3 | [5 empty] | 7.8 | 103.6 | 104.6429
11.4 | [5 empty] | 7.45 | 103.6 | 104.7
11.5 | [5 empty] | 7.14 | 103.7 | 104.7
11.6 | [5 empty] | 6.9 | 103.7 | 104.8
11.7 | [5 empty] | 6.6333 | 103.7 | 104.8
11.8 | [5 empty] | 6.5 | 103.7 | 104.8333
11.9 | [5 empty] | 6.2667 | 103.7 | 104.9
12 | [5 empty] | 6.1 | 103.7 | 105
12.1 | [5 empty] | 5.9 | 103.7 | 105.0571
12.2 | [5 empty] | 5.7667 | 103.8 | 105.1
12.3 | [5 empty] | 5.6 | 103.8 | 105.1
12.4 | [5 empty] | 5.46 | 103.8 | 105.2
12.5 | [5 empty] | 5.3 | 103.8 | 105.275
12.6 | [5 empty] | 5.2 | 103.8 | 105.3
12.7 | [5 empty] | 5.06 | 103.8 | 105.3429
12.8 | [5 empty] | 4.9 | 103.8 | 105.4
12.9 | [5 empty] | 4.8 | 103.8 | 105.5
13 | [5 empty] | 4.6667 | 103.8 | 105.5
13.1 | [5 empty] | 4.56 | 103.78 | 105.5714
13.2 | [5 empty] | 4.4 | 103.7 | 105.6
13.3 | [5 empty] | 4.3 | 103.7 | 105.6
13.4 | [5 empty] | 4.2 | 103.7 | 105.6
13.5 | [5 empty] | 4.1 | 103.66 | 105.7
13.6 | [5 empty] | 4 | 103.6 | 105.7
13.7 | [5 empty] | 3.9 | 103.5 | 105.7625
13.8 | [5 empty] | 3.85 | 103.46 | 105.8
13.9 | [5 empty] | 3.7667 | 103.325 | 105.9
14 | [5 empty] | 3.7 | 103.2 | 105.9
14.1 | [5 empty] | 3.6 | 103 | 105.925
14.2 | [5 empty] | 3.5 | 102.8 | 106
14.3 | [5 empty] | 3.4 | 102.54 | 106.0875
14.4 | [5 empty] | 3.4 | 102.1 | 106.1
14.5 | [5 empty] | 3.3 | 101.55 | 106.1
14.6 | [5 empty] | 3.2 | 100.6 | 106.1
14.7 | [5 empty] | 3.2 | 99.3 | 106.2
14.8 | [5 empty] | 3.1 | 97.15 | 106.2
14.9 | [5 empty] | 3.025 | 93.36 | 106.2
15 | [5 empty] | 3 | 86.74 | 106.2875
15.1 | [5 empty] | 2.9 | 77.8333 | 106.3
15.2 | [6 empty] | 67.34 | 106.3
15.3 | [6 empty] | 55.1 | 106.3857
15.4 | [6 empty] | 43.3 | 106.4
15.5 | [6 empty] | 33.3667 | 106.4
15.6 | [6 empty] | 25.08 | 106.4857
15.7 | [6 empty] | 18.8 | 106.5
15.8 | [6 empty] | 15.3833 | 106.5429
15.9 | [6 empty] | 13.35 | 106.6
16 | [6 empty] | 12 | 106.6
16.1 | [6 empty] | 11.1 | 106.6
16.2 | [6 empty] | 10.4 | 106.6714
16.3 | [6 empty] | 9.8667 | 106.7
16.4 | [6 empty] | 9.52 | 106.7375
16.5 | [6 empty] | 9.2 | 106.8
16.6 | [6 empty] | 8.8667 | 106.8
16.7 | [6 empty] | 8.6286 | 106.8
16.8 | [6 empty] | 8.4 | 106.8
16.9 | [6 empty] | 8.16 | 106.8
17 | [6 empty] | 7.9667 | 106.8091
17.1 | [6 empty] | 7.76 | 106.9
17.2 | [6 empty] | 7.6 | 106.9
17.3 | [6 empty] | 7.4 | 106.9
17.4 | [6 empty] | 7.3 | 106.9
17.5 | [6 empty] | 7.1 | 106.9
17.6 | [6 empty] | 6.94 | 106.9
17.7 | [6 empty] | 6.8 | 106.9
17.8 | [6 empty] | 6.7 | 106.9
17.9 | [6 empty] | 6.5333 | 106.825
18 | [6 empty] | 6.4 | 106.8
18.1 | [6 empty] | 6.24 | 106.8
18.2 | [6 empty] | 6.1 | 106.7333
18.3 | [6 empty] | 6 | 106.7
18.4 | [6 empty] | 5.86 | 106.6
18.5 | [6 empty] | 5.75 | 106.5375
18.6 | [6 empty] | 5.5833 | 106.5
18.7 | [6 empty] | 5.5 | 106.4
18.8 | [6 empty] | 5.4 | 106.25
18.9 | [6 empty] | 5.3 | 106.16
19 | [6 empty] | 5.2 | 105.96
19.1 | [6 empty] | 5.1 | 105.76
19.2 | [6 empty] | 4.9667 | 105.5714
19.3 | [6 empty] | 4.875 | 105.2667
19.4 | [6 empty] | 4.76 | 104.925
19.5 | [6 empty] | 4.675 | 104.4714
19.6 | [6 empty] | 4.6 | 103.8625
19.7 | [6 empty] | 4.5 | 102.8143
19.8 | [6 empty] | 4.4 | 101.2667
19.9 | [6 empty] | 4.3 | 98.8857
20 | [6 empty] | 4.26 | 94.4833
20.1 | [6 empty] | 4.14 | 87.1
20.2 | [6 empty] | 4.1 | 77.3714
20.3 | [7 empty] | 66.4625
20.4 | [7 empty] | 54.9714
20.5 | [7 empty] | 43.0857
20.6 | [7 empty] | 32.78
20.7 | [7 empty] | 24.46
20.8 | [7 empty] | 19
20.9 | [7 empty] | 15.68
21 | [7 empty] | 13.76
21.1 | [7 empty] | 12.52
21.2 | [7 empty] | 11.72
21.3 | [7 empty] | 11.04
21.4 | [7 empty] | 10.56
21.5 | [7 empty] | 10.16
21.6 | [7 empty] | 9.78
21.7 | [7 empty] | 9.48
21.8 | [7 empty] | 9.26
21.9 | [7 empty] | 8.98
22 | [7 empty] | 8.78
22.1 | [7 empty] | 8.58
22.2 | [7 empty] | 8.2833
22.3 | [7 empty] | 8.1
22.4 | [7 empty] | 8
22.5 | [7 empty] | 7.7833
22.6 | [7 empty] | 7.6
22.7 | [7 empty] | 7.4
22.8 | [7 empty] | 7.2
22.9 | [7 empty] | 7.1
23 | [7 empty] | 6.9
23.1 | [7 empty] | 6.7
23.2 | [7 empty] | 6.6
23.3 | [7 empty] | 6.4
23.4 | [7 empty] | 6.3
23.5 | [7 empty] | 6.1429
23.6 | [7 empty] | 6
23.7 | [7 empty] | 5.8
23.8 | [7 empty] | 5.7
23.9 | [7 empty] | 5.6
24 | [7 empty] | 5.4
24.1 | [7 empty] | 5.3
24.2 | [7 empty] | 5.2167
24.3 | [7 empty] | 5.075
24.4 | [7 empty] | 5
24.5 | [7 empty] | 4.9125
24.6 | [7 empty] | 4.8
24.7 | [7 empty] | 4.7
24.8 | [7 empty] | 4.6
24.9 | [7 empty] | 4.5143
25 | [7 empty] | 4.4333
25.1 | [7 empty] | 4.3
25.2 | [7 empty] | 4.2833
25.3 | [7 empty] | 4.2
25.4 | [7 empty] | 4.1
25.5 | [7 empty] | 4.0167
25.6 | [7 empty] | 3.95
25.7 | [7 empty] | 3.9
25.8 | [7 empty] | 3.8308
25.9 | [7 empty] | 3.8
26 | [7 empty] | 3.7333
26.1 | [7 empty] | 3.7
26.2 | [7 empty] | 3.6
26.3 | [7 empty] | 3.6
26.4 | [7 empty] | 3.5125
26.5 | [7 empty] | 3.5
26.6 | [7 empty] | 3.4
26.7 | [7 empty] | 3.4
26.8 | [7 empty] | 3.3222
26.9 | [7 empty] | 3.3
27 | [7 empty] | 3.2286
27.1 | [7 empty] | 3.2
27.2 | [7 empty] | 3.1917
27.3 | [7 empty] | 3.1083
27.4 | [7 empty] | 3.1
27.5 | [7 empty] | 3.0714
27.6 | [7 empty] | 3
27.7 | [7 empty] | 3
27.8 | [7 empty] | 2.9
27.9 | [7 empty] | 2.9
28 | [7 empty] | 2.9
28.1 | [7 empty] | 2.8429
28.2 | [7 empty] | 2.8
28.3 | [7 empty] | 2.8

## Open Field Profiles at 5cm
Lateral Profiles
Open field
SSD=100cm
Depth=5cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-28.9 | [7 empty] | 3.4
-28.8 | [7 empty] | 3.4
-28.7 | [7 empty] | 3.4857
-28.6 | [7 empty] | 3.5
-28.5 | [7 empty] | 3.5857
-28.4 | [7 empty] | 3.6
-28.3 | [7 empty] | 3.6667
-28.2 | [7 empty] | 3.7
-28.1 | [7 empty] | 3.7
-28 | [7 empty] | 3.8
-27.9 | [7 empty] | 3.8667
-27.8 | [7 empty] | 3.9
-27.7 | [7 empty] | 3.9286
-27.6 | [7 empty] | 4
-27.5 | [7 empty] | 4.0714
-27.4 | [7 empty] | 4.1
-27.3 | [7 empty] | 4.2
-27.2 | [7 empty] | 4.2
-27.1 | [7 empty] | 4.2667
-27 | [7 empty] | 4.3
-26.9 | [7 empty] | 4.4
-26.8 | [7 empty] | 4.4714
-26.7 | [7 empty] | 4.5
-26.6 | [7 empty] | 4.6
-26.5 | [7 empty] | 4.6625
-26.4 | [7 empty] | 4.8
-26.3 | [7 empty] | 4.8167
-26.2 | [7 empty] | 4.9
-26.1 | [7 empty] | 5
-26 | [7 empty] | 5.08
-25.9 | [7 empty] | 5.18
-25.8 | [7 empty] | 5.28
-25.7 | [7 empty] | 5.35
-25.6 | [7 empty] | 5.4
-25.5 | [7 empty] | 5.5667
-25.4 | [7 empty] | 5.66
-25.3 | [7 empty] | 5.72
-25.2 | [7 empty] | 5.8857
-25.1 | [7 empty] | 6.0125
-25 | [7 empty] | 6.1
-24.9 | [7 empty] | 6.2333
-24.8 | [7 empty] | 6.4125
-24.7 | [7 empty] | 6.5
-24.6 | [7 empty] | 6.6
-24.5 | [7 empty] | 6.76
-24.4 | [7 empty] | 6.9
-24.3 | [7 empty] | 7.0429
-24.2 | [7 empty] | 7.25
-24.1 | [7 empty] | 7.4125
-24 | [7 empty] | 7.5429
-23.9 | [7 empty] | 7.725
-23.8 | [7 empty] | 7.96
-23.7 | [7 empty] | 8.1429
-23.6 | [7 empty] | 8.275
-23.5 | [7 empty] | 8.5125
-23.4 | [7 empty] | 8.65
-23.3 | [7 empty] | 8.9
-23.2 | [7 empty] | 9.1286
-23.1 | [7 empty] | 9.325
-23 | [7 empty] | 9.6
-22.9 | [7 empty] | 9.7571
-22.8 | [7 empty] | 10.025
-22.7 | [7 empty] | 10.2571
-22.6 | [7 empty] | 10.5667
-22.5 | [7 empty] | 10.8143
-22.4 | [7 empty] | 11.1625
-22.3 | [7 empty] | 11.4429
-22.2 | [7 empty] | 11.8333
-22.1 | [7 empty] | 12.25
-22 | [7 empty] | 12.75
-21.9 | [7 empty] | 13.3714
-21.8 | [7 empty] | 14.1333
-21.7 | [7 empty] | 15.05
-21.6 | [7 empty] | 16.4667
-21.5 | [7 empty] | 18.5
-21.4 | [7 empty] | 21.9
-21.3 | [7 empty] | 26.95
-21.2 | [7 empty] | 34.6375
-21.1 | [7 empty] | 43.5
-21 | [7 empty] | 54.06
-20.9 | [7 empty] | 64.625
-20.8 | [7 empty] | 74.4167
-20.7 | [6 empty] | 4.7 | 82.9286
-20.6 | [6 empty] | 4.8 | 89.2
-20.5 | [6 empty] | 4.8 | 93.3625
-20.4 | [6 empty] | 4.9 | 95.8714
-20.3 | [6 empty] | 5 | 97.4
-20.2 | [6 empty] | 5.1 | 98.5625
-20.1 | [6 empty] | 5.2 | 99.3571
-20 | [6 empty] | 5.3 | 100
-19.9 | [6 empty] | 5.45 | 100.5125
-19.8 | [6 empty] | 5.52 | 100.8875
-19.7 | [6 empty] | 5.6 | 101.3
-19.6 | [6 empty] | 5.8 | 101.5571
-19.5 | [6 empty] | 5.9 | 101.7625
-19.4 | [6 empty] | 6 | 102
-19.3 | [6 empty] | 6.1 | 102.2571
-19.2 | [6 empty] | 6.2 | 102.4625
-19.1 | [6 empty] | 6.3571 | 102.5875
-19 | [6 empty] | 6.46 | 102.825
-18.9 | [6 empty] | 6.65 | 102.9
-18.8 | [6 empty] | 6.8 | 102.9857
-18.7 | [6 empty] | 6.925 | 103.1143
-18.6 | [6 empty] | 7.075 | 103.25
-18.5 | [6 empty] | 7.26 | 103.3857
-18.4 | [6 empty] | 7.4 | 103.5
-18.3 | [6 empty] | 7.5333 | 103.55
-18.2 | [6 empty] | 7.775 | 103.6
-18.1 | [6 empty] | 7.925 | 103.7
-18 | [6 empty] | 8.1 | 103.7636
-17.9 | [6 empty] | 8.325 | 103.8
-17.8 | [6 empty] | 8.4857 | 103.8067
-17.7 | [6 empty] | 8.6857 | 103.8733
-17.6 | [6 empty] | 8.9 | 103.94
-17.5 | [6 empty] | 9.225 | 104
-17.4 | [6 empty] | 9.375 | 104
-17.3 | [6 empty] | 9.625 | 104
-17.2 | [6 empty] | 9.975 | 104.0067
-17.1 | [6 empty] | 10.2333 | 104.0733
-17 | [6 empty] | 10.6 | 104.1
-16.9 | [6 empty] | 10.975 | 104.1
-16.8 | [6 empty] | 11.45 | 104.1
-16.7 | [6 empty] | 12.025 | 104.1
-16.6 | [6 empty] | 12.7 | 104.1
-16.5 | [6 empty] | 13.6 | 104.1429
-16.4 | [6 empty] | 14.7 | 104.2
-16.3 | [6 empty] | 16.5 | 104.2
-16.2 | [6 empty] | 19.3 | 104.2
-16.1 | [6 empty] | 23.7 | 104.2
-16 | [6 empty] | 30.15 | 104.2
-15.9 | [6 empty] | 39.2 | 104.2
-15.8 | [6 empty] | 49.45 | 104.2
-15.7 | [6 empty] | 60 | 104.2
-15.6 | [6 empty] | 70.1 | 104.2
-15.5 | [6 empty] | 79.4 | 104.2
-15.4 | [5 empty] | 3.32 | 86.15 | 104.1
-15.3 | [5 empty] | 3.42 | 90.8333 | 104.1
-15.2 | [5 empty] | 3.5 | 93.825 | 104.1
-15.1 | [5 empty] | 3.525 | 95.62 | 104.1
-15 | [5 empty] | 3.6 | 96.725 | 104.1
-14.9 | [5 empty] | 3.7 | 97.625 | 104.1
-14.8 | [5 empty] | 3.8 | 98.3667 | 104.1
-14.7 | [5 empty] | 3.8286 | 98.7714 | 104.1
-14.6 | [5 empty] | 3.9625 | 99.22 | 104.1
-14.5 | [5 empty] | 4.0286 | 99.6143 | 104.1
-14.4 | [5 empty] | 4.1 | 99.76 | 104.0167
-14.3 | [5 empty] | 4.2 | 100.0143 | 104
-14.2 | [5 empty] | 4.3 | 100.26 | 104
-14.1 | [5 empty] | 4.4 | 100.4 | 104
-14 | [5 empty] | 4.5 | 100.525 | 103.9375
-13.9 | [5 empty] | 4.62 | 100.6375 | 103.9
-13.8 | [5 empty] | 4.7 | 100.8 | 103.9
-13.7 | [5 empty] | 4.85 | 100.9 | 103.9
-13.6 | [5 empty] | 4.9 | 101.05 | 103.875
-13.5 | [5 empty] | 5.1 | 101.1 | 103.8
-13.4 | [5 empty] | 5.24 | 101.2 | 103.8
-13.3 | [5 empty] | 5.35 | 101.2 | 103.8
-13.2 | [5 empty] | 5.5143 | 101.3 | 103.8
-13.1 | [5 empty] | 5.6 | 101.4 | 103.8
-13 | [5 empty] | 5.7625 | 101.4 | 103.7143
-12.9 | [5 empty] | 5.95 | 101.5 | 103.7
-12.8 | [5 empty] | 6.175 | 101.5 | 103.7
-12.7 | [5 empty] | 6.3 | 101.5 | 103.7
-12.6 | [5 empty] | 6.5 | 101.54 | 103.6
-12.5 | [5 empty] | 6.64 | 101.6 | 103.6
-12.4 | [5 empty] | 6.84 | 101.6 | 103.6
-12.3 | [5 empty] | 7.04 | 101.6 | 103.5
-12.2 | [5 empty] | 7.325 | 101.7 | 103.5
-12.1 | [5 empty] | 7.54 | 101.7 | 103.46
-12 | [5 empty] | 7.82 | 101.72 | 103.4
-11.9 | [5 empty] | 8.08 | 101.8 | 103.4
-11.8 | [5 empty] | 8.425 | 101.8 | 103.3125
-11.7 | [5 empty] | 8.825 | 101.875 | 103.3
-11.6 | [5 empty] | 9.15 | 101.9 | 103.3
-11.5 | [5 empty] | 9.75 | 101.9 | 103.3
-11.4 | [5 empty] | 10.28 | 101.9 | 103.3
-11.3 | [5 empty] | 11.12 | 101.9 | 103.2
-11.2 | [5 empty] | 12.1714 | 101.9 | 103.2
-11.1 | [5 empty] | 13.64 | 101.9 | 103.1455
-11 | [5 empty] | 15.8429 | 101.9 | 103.1
-10.9 | [5 empty] | 19.06 | 101.8 | 103.1
-10.8 | [5 empty] | 24.42 | 101.8 | 103.1
-10.7 | [5 empty] | 32.5667 | 101.8 | 103.0417
-10.6 | [5 empty] | 41.8 | 101.8 | 103
-10.5 | [5 empty] | 52.1 | 101.8 | 103
-10.4 | [5 empty] | 62.4 | 101.9 | 102.9
-10.3 | [5 empty] | 72.32 | 101.9 | 102.9
-10.2 |  |  |  |  | 1.8 | 80.375 | 101.9 | 102.9
-10.1 |  |  |  |  | 1.8 | 86.625 | 101.9 | 102.87
-10 |  |  |  |  | 1.8 | 90.7333 | 101.9 | 102.8
-9.9 |  |  |  |  | 1.9 | 93.075 | 101.9 | 102.8
-9.8 |  |  |  |  | 1.9 | 94.65 | 101.8 | 102.8
-9.7 |  |  |  |  | 1.9 | 95.775 | 101.8 | 102.77
-9.6 |  |  |  |  | 2 | 96.525 | 101.8 | 102.7
-9.5 |  |  |  |  | 2 | 97.15 | 101.8 | 102.7
-9.4 |  |  |  |  | 2.1 | 97.575 | 101.8 | 102.7
-9.3 |  |  |  |  | 2.12 | 97.9 | 101.8 | 102.6125
-9.2 |  |  |  |  | 2.2 | 98.325 | 101.8 | 102.6
-9.1 |  |  |  | 1.4 | 2.2 | 98.575 | 101.8 | 102.5429
-9 |  |  |  | 1.4 | 2.3 | 98.82 | 101.8 | 102.5
-8.9 |  |  |  | 1.5 | 2.4 | 98.975 | 101.75 | 102.5
-8.8 |  |  |  | 1.5 | 2.4 | 99.125 | 101.7 | 102.5
-8.7 |  |  |  | 1.5 | 2.5 | 99.25 | 101.7 | 102.5
-8.6 |  |  |  | 1.5 | 2.6 | 99.4 | 101.7 | 102.4125
-8.5 |  |  |  | 1.6 | 2.6 | 99.5 | 101.7 | 102.375
-8.4 |  |  |  | 1.6 | 2.7 | 99.6 | 101.6857 | 102.3
-8.3 |  |  |  | 1.7 | 2.8 | 99.78 | 101.6 | 102.3
-8.2 |  |  |  | 1.7 | 2.9 | 99.9 | 101.6 | 102.2
-8.1 |  |  | 1 | 1.8 | 3 | 99.9 | 101.6 | 102.2
-8 |  |  | 1.1 | 1.8 | 3.1 | 100 | 101.6 | 102.2
-7.9 |  |  | 1.1 | 1.9 | 3.2 | 100 | 101.55 | 102.1875
-7.8 |  |  | 1.1 | 1.9 | 3.325 | 100 | 101.5 | 102.1
-7.7 |  |  | 1.1 | 2 | 3.475 | 100.1 | 101.5 | 102.1
-7.6 |  |  | 1.2 | 2 | 3.5333 | 100.1 | 101.5 | 102.1
-7.5 |  |  | 1.2 | 2.1 | 3.675 | 100.2 | 101.5 | 102.1
-7.4 |  |  | 1.2 | 2.2 | 3.8 | 100.2 | 101.5 | 102.025
-7.3 |  |  | 1.2 | 2.2 | 3.9667 | 100.3 | 101.45 | 102
-7.2 |  |  | 1.3 | 2.3 | 4.2 | 100.3 | 101.4 | 101.9571
-7.1 |  |  | 1.3 | 2.4 | 4.3 | 100.3 | 101.4 | 101.9
-7 |  | 0.7 | 1.4 | 2.5 | 4.5 | 100.3 | 101.4 | 101.9
-6.9 |  | 0.73333 | 1.4 | 2.6 | 4.75 | 100.3 | 101.4 | 101.84
-6.8 |  | 0.8 | 1.4 | 2.7 | 4.9 | 100.35 | 101.4 | 101.8
-6.7 |  | 0.8 | 1.5 | 2.75 | 5.14 | 100.4 | 101.3667 | 101.8
-6.6 |  | 0.8 | 1.5 | 2.8667 | 5.4625 | 100.4 | 101.3 | 101.8
-6.5 | 0.6 | 0.8 | 1.6 | 3 | 5.76 | 100.4 | 101.3 | 101.7
-6.4 | 0.6 | 0.8 | 1.6 | 3.1 | 6.16 | 100.4 | 101.3 | 101.7
-6.3 | 0.6 | 0.8 | 1.7 | 3.2 | 6.6333 | 100.4 | 101.2 | 101.7
-6.2 | 0.6 | 0.9 | 1.7 | 3.4 | 7.2 | 100.45 | 101.2 | 101.6714
-6.1 | 0.6 | 0.9 | 1.8 | 3.5 | 7.9 | 100.5 | 101.2 | 101.6
-6 | 0.6 | 0.9 | 1.9 | 3.7 | 8.8 | 100.5 | 101.1714 | 101.6
-5.9 | 0.63333 | 0.9 | 1.9 | 3.9 | 9.9333 | 100.5 | 101.1 | 101.54
-5.8 | 0.7 | 0.93333 | 2 | 4.0667 | 11.8 | 100.5 | 101.1 | 101.5
-5.7 | 0.7 | 1 | 2.1 | 4.25 | 14.2 | 100.5 | 101 | 101.5
-5.6 | 0.7 | 1 | 2.2 | 4.5 | 18.2 | 100.5 | 101 | 101.5
-5.5 | 0.7 | 1 | 2.3 | 4.8 | 24.25 | 100.5 | 101.1 | 101.425
-5.4 | 0.7 | 1.1 | 2.4 | 5.1 | 32.575 | 100.5 | 101.1 | 101.4
-5.3 | 0.7 | 1.1 | 2.5 | 5.5333 | 43.15 | 100.5 | 101 | 101.37
-5.2 | 0.8 | 1.1 | 2.6 | 6 | 53.8 | 100.5 | 101.075 | 101.3
-5.1 | 0.8 | 1.2 | 2.7 | 6.5667 | 64.3 | 100.5 | 101 | 101.3
-5 | 0.8 | 1.2 | 2.8 | 7.3 | 73.8 | 100.5 | 101 | 101.3
-4.9 | 0.8 | 1.3 | 3 | 8.3 | 82 | 100.5 | 100.9 | 101.2687
-4.8 | 0.83333 | 1.3 | 3.1 | 9.8 | 87.4 | 100.5 | 100.9 | 101.2062
-4.7 | 0.9 | 1.4 | 3.3 | 11.8 | 90.775 | 100.5 | 100.8 | 101.2
-4.6 | 0.9 | 1.4 | 3.55 | 14.7 | 93.1 | 100.5 | 100.8 | 101.2
-4.5 | 0.9 | 1.5 | 3.7667 | 19.5 | 94.525 | 100.5 | 100.8 | 101.2
-4.4 | 1 | 1.6 | 4.05 | 26.7 | 95.6 | 100.5 | 100.8 | 101.14
-4.3 | 1 | 1.7 | 4.3667 | 36.6 | 96.425 | 100.5 | 100.7 | 101.0733
-4.2 | 1.1 | 1.7 | 4.8 | 46.5 | 96.925 | 100.4 | 100.7 | 101.0067
-4.1 | 1.1 | 1.8 | 5.3333 | 57.8667 | 97.4 | 100.4 | 100.7 | 101
-4 | 1.1 | 1.9 | 6 | 68.85 | 97.7 | 100.4 | 100.7 | 100.9636
-3.9 | 1.2 | 2 | 6.8 | 77.7 | 98.0333 | 100.4 | 100.7 | 100.9
-3.8 | 1.2 | 2.1333 | 8 | 84.65 | 98.34 | 100.4 | 100.7 | 100.9
-3.7 | 1.3 | 2.2667 | 9.6 | 89.3 | 98.5 | 100.4 | 100.7 | 100.8
-3.6 | 1.4 | 2.4 | 12.1 | 92 | 98.7 | 100.3 | 100.625 | 100.8
-3.5 | 1.5 | 2.7 | 15.7 | 93.9 | 98.85 | 100.3 | 100.6 | 100.8
-3.4 | 1.5 | 2.9 | 21.8333 | 95.1 | 99 | 100.3 | 100.6 | 100.7
-3.3 | 1.6 | 3.1 | 30.3 | 96 | 99.1 | 100.3 | 100.5 | 100.7
-3.2 | 1.75 | 3.5 | 40.05 | 96.8 | 99.2 | 100.3 | 100.5 | 100.7
-3.1 | 1.9 | 3.95 | 51.05 | 97.3 | 99.3 | 100.275 | 100.5 | 100.7
-3 | 2 | 4.45 | 62.6 | 97.6667 | 99.4 | 100.275 | 100.46 | 100.6714
-2.9 | 2.2 | 5.2 | 72.3 | 98 | 99.5 | 100.2 | 100.4 | 100.6
-2.8 | 2.4 | 6.1 | 80.9 | 98.3 | 99.6 | 100.2 | 100.4 | 100.6
-2.7 | 2.7 | 7.55 | 86.9 | 98.5 | 99.6 | 100.2 | 100.375 | 100.5714
-2.6 | 3.0667 | 9.5 | 91.1667 | 98.6 | 99.7 | 100.125 | 100.3 | 100.5
-2.5 | 3.5333 | 12.5 | 93.3 | 98.8 | 99.7 | 100.2 | 100.3 | 100.5
-2.4 | 4.15 | 17.5 | 95 | 98.95 | 99.7 | 100.2 | 100.3 | 100.5
-2.3 | 5 | 24.5 | 96.1 | 99.1 | 99.8 | 100.1 | 100.3 | 100.45
-2.2 | 6.2333 | 33.8 | 97 | 99.1333 | 99.8 | 100.1 | 100.2 | 100.4
-2.1 | 8 | 44.8 | 97.5667 | 99.2 | 99.9 | 100.1 | 100.2 | 100.4
-2 | 10.7 | 56.55 | 98 | 99.3 | 99.9 | 100.1 | 100.2 | 100.4
-1.9 | 15.3 | 67.9 | 98.4 | 99.4 | 99.9 | 100.1 | 100.2 | 100.325
-1.8 | 22.3 | 77.2 | 98.6333 | 99.5 | 100 | 100.1 | 100.2 | 100.3
-1.7 | 31.1 | 84.9 | 98.9 | 99.5 | 100 | 100.1 | 100.1 | 100.3
-1.6 | 41.625 | 89.95 | 99 | 99.6 | 100 | 100.1 | 100.1 | 100.3
-1.5 | 53.2 | 93.25 | 99.2 | 99.6 | 100 | 100.1 | 100.1 | 100.3
-1.4 | 65.2 | 95.4667 | 99.4 | 99.7 | 100 | 100.1 | 100.1 | 100.24
-1.3 | 75.0667 | 96.9 | 99.5 | 99.7 | 100 | 100 | 100.1 | 100.2
-1.2 | 83.9333 | 97.6 | 99.6 | 99.7 | 100 | 100 | 100.1 | 100.2
-1.1 | 90.2 | 98.1667 | 99.6 | 99.8 | 100 | 100 | 100.1 | 100.2
-1 | 93.7 | 98.8 | 99.6 | 99.8 | 100 | 100 | 100 | 100.12
-0.9 | 95.7 | 99.1 | 99.7 | 99.8 | 100 | 100 | 100 | 100.1
-0.8 | 97.1 | 99.3 | 99.8 | 99.9 | 100 | 100 | 100 | 100.1
-0.7 | 98 | 99.55 | 99.8 | 99.9 | 100 | 100 | 100 | 100
-0.6 | 98.85 | 99.7 | 99.9 | 99.9 | 100 | 100 | 100 | 100
-0.5 | 99.2 | 99.8 | 99.9 | 99.9 | 100 | 99.975 | 100 | 100
-0.4 | 99.5333 | 99.8 | 99.9 | 100 | 100 | 99.9 | 100 | 100
-0.3 | 99.8 | 99.9 | 100 | 100 | 100 | 99.9 | 100 | 100
-0.2 | 99.9 | 99.9 | 100 | 100 | 100 | 99.9 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100 | 100 | 100
0.4 | 99.5 | 99.9 | 100 | 99.9 | 100.1 | 100 | 100 | 100
0.5 | 99.3 | 99.8 | 100 | 99.9 | 100.1 | 100 | 100 | 100
0.6 | 98.875 | 99.7 | 100 | 99.9 | 100.1 | 100 | 100.0667 | 100
0.7 | 98.2 | 99.5 | 99.9 | 99.9 | 100.1 | 100 | 100.1 | 100
0.8 | 97.1 | 99.3 | 99.9 | 99.9 | 100.1 | 100.1 | 100.1 | 100.075
0.9 | 95.65 | 99.1 | 99.9 | 99.9 | 100.1 | 100.1 | 100.1 | 100.1
1 | 93.3 | 98.8 | 99.8 | 99.9 | 100.1 | 100.1 | 100.175 | 100.1
1.1 | 89.7 | 98.35 | 99.7 | 99.9 | 100.1 | 100.1 | 100.2 | 100.2
1.2 | 83.5667 | 97.6 | 99.6 | 99.8 | 100.1 | 100.1 | 100.2 | 100.2
1.3 | 74.8 | 96.7 | 99.6 | 99.8 | 100.1 | 100.1 | 100.2 | 100.2
1.4 | 64.2 | 95.3 | 99.5 | 99.8 | 100.1 | 100.1 | 100.2 | 100.2
1.5 | 53.6 | 93.1 | 99.3 | 99.7 | 100.1 | 100.2 | 100.2 | 100.3
1.6 | 42.275 | 89.9 | 99.2 | 99.7 | 100.1 | 100.2 | 100.2333 | 100.3
1.7 | 31 | 84.6 | 99 | 99.6 | 100.075 | 100.2 | 100.3 | 100.3
1.8 | 22.075 | 77 | 98.75 | 99.5333 | 100 | 100.2 | 100.3714 | 100.3
1.9 | 15 | 67.8 | 98.5 | 99.5 | 100 | 100.3 | 100.4 | 100.3
2 | 10.5 | 56.2 | 98.15 | 99.4 | 100 | 100.3 | 100.4 | 100.4
2.1 | 8 | 44.5 | 97.4 | 99.4 | 100 | 100.3 | 100.4 | 100.4
2.2 | 6.1 | 33.9 | 96.9333 | 99.3 | 100 | 100.3 | 100.4 | 100.4
2.3 | 5 | 24.1 | 96.2 | 99.2 | 100 | 100.32 | 100.5 | 100.4143
2.4 | 4.15 | 17.3 | 95.1 | 99 | 99.95 | 100.4 | 100.5 | 100.5
2.5 | 3.5 | 12.2 | 93.3 | 98.9 | 99.9 | 100.4 | 100.525 | 100.5
2.6 | 3.05 | 9.4 | 90.8 | 98.8 | 99.8625 | 100.4 | 100.6 | 100.5
2.7 | 2.7 | 7.4 | 86.8 | 98.6 | 99.8 | 100.5 | 100.6 | 100.5375
2.8 | 2.4333 | 5.9 | 80.9667 | 98.3 | 99.7 | 100.5 | 100.6 | 100.6
2.9 | 2.1667 | 5 | 72.25 | 98.0667 | 99.6 | 100.5 | 100.6 | 100.6
3 | 2 | 4.3667 | 62.45 | 97.7333 | 99.5 | 100.5 | 100.6 | 100.6375
3.1 | 1.9 | 3.8 | 51.3 | 97.4 | 99.4571 | 100.5 | 100.7 | 100.7
3.2 | 1.7 | 3.45 | 40.05 | 96.8 | 99.3 | 100.5 | 100.7 | 100.7
3.3 | 1.6 | 3.1 | 30.4 | 96.2 | 99.3 | 100.5 | 100.7 | 100.7
3.4 | 1.55 | 2.8 | 21.9 | 95.35 | 99.2 | 100.5 | 100.7857 | 100.7625
3.5 | 1.4 | 2.65 | 15.6 | 94 | 99 | 100.5 | 100.8 | 100.8
3.6 | 1.3667 | 2.4 | 11.9 | 92.1 | 98.9 | 100.6 | 100.8 | 100.8
3.7 | 1.3 | 2.3 | 9.45 | 89.4 | 98.725 | 100.5 | 100.9 | 100.8
3.8 | 1.2 | 2.2 | 8 | 84.6 | 98.4 | 100.6 | 100.9 | 100.9
3.9 | 1.2 | 2.1 | 6.8 | 77.7 | 98.18 | 100.6 | 100.9 | 100.9
4 | 1.1 | 1.9 | 5.9 | 68.4 | 97.9 | 100.6 | 100.9 | 100.9308
4.1 | 1.1 | 1.8 | 5.3 | 58.1333 | 97.6 | 100.6 | 100.98 | 101.0067
4.2 | 1.1 | 1.7 | 4.8 | 46.6 | 97.2 | 100.7 | 101 | 101.0733
4.3 | 1 | 1.6 | 4.275 | 36.5667 | 96.5667 | 100.7 | 101.1 | 101.1
4.4 | 1 | 1.6 | 4.05 | 26.9 | 95.875 | 100.7 | 101.1 | 101.1067
4.5 | 0.9 | 1.5 | 3.775 | 20 | 94.6333 | 100.7 | 101.1 | 101.1733
4.6 | 0.9 | 1.4 | 3.5 | 14.8 | 93.2 | 100.8 | 101.1 | 101.2
4.7 | 0.9 | 1.4 | 3.3 | 11.6667 | 90.975 | 100.8 | 101.1 | 101.2
4.8 | 0.9 | 1.3 | 3.175 | 9.8 | 87.2 | 100.8 | 101.1 | 101.2
4.9 | 0.8 | 1.3 | 3 | 8.4 | 81.8 | 100.8 | 101.1 | 101.2357
5 | 0.8 | 1.2 | 2.8 | 7.4 | 73.8667 | 100.8 | 101.2 | 101.3
5.1 | 0.8 | 1.2 | 2.7 | 6.6 | 64.2 | 100.8 | 101.2 | 101.3
5.2 | 0.8 | 1.1 | 2.6 | 6 | 53.8 | 100.8 | 101.225 | 101.3
5.3 | 0.7 | 1.1 | 2.5 | 5.5333 | 42.8 | 100.8 | 101.3 | 101.34
5.4 | 0.7 | 1.1 | 2.4 | 5.2 | 32.75 | 100.7 | 101.3 | 101.4
5.5 | 0.7 | 1.0333 | 2.3 | 4.8 | 24.6667 | 100.7 | 101.3 | 101.4
5.6 | 0.7 | 1 | 2.2 | 4.5 | 18.4 | 100.8 | 101.3 | 101.5
5.7 | 0.7 | 1 | 2.1 | 4.3 | 14.3 | 100.8 | 101.3 | 101.5
5.8 | 0.7 | 0.96667 | 2 | 4.1 | 11.625 | 100.8 | 101.3 | 101.5
5.9 | 0.6 | 0.9 | 1.9 | 3.9 | 9.95 | 100.8 | 101.4 | 101.5
6 | 0.6 | 0.9 | 1.9 | 3.7 | 8.7 | 100.7 | 101.4 | 101.6
6.1 | 0.6 | 0.9 | 1.8 | 3.5333 | 7.8 | 100.7 | 101.4 | 101.6
6.2 | 0.6 | 0.9 | 1.7 | 3.4 | 7.075 | 100.7 | 101.4 | 101.6375
6.3 | 0.6 | 0.8 | 1.7 | 3.3 | 6.6 | 100.6 | 101.4 | 101.7
6.4 | 0.6 | 0.8 | 1.6 | 3.1 | 6.1 | 100.6 | 101.5 | 101.7
6.5 | 0.6 | 0.8 | 1.6 | 3 | 5.725 | 100.6 | 101.5 | 101.7571
6.6 |  | 0.8 | 1.5 | 2.9 | 5.45 | 100.6 | 101.5 | 101.8
6.7 |  | 0.8 | 1.5 | 2.8 | 5.2 | 100.6 | 101.6 | 101.8
6.8 |  | 0.7 | 1.4 | 2.7 | 4.94 | 100.6 | 101.6 | 101.8
6.9 |  | 0.7 | 1.4 | 2.6 | 4.7 | 100.6 | 101.6 | 101.8
7 |  | 0.7 | 1.3 | 2.5 | 4.5 | 100.6 | 101.6 | 101.9
7.1 |  |  | 1.3 | 2.4 | 4.3 | 100.54 | 101.6 | 101.9
7.2 |  |  | 1.3 | 2.3333 | 4.15 | 100.5 | 101.6 | 101.925
7.3 |  |  | 1.2 | 2.3 | 3.95 | 100.5 | 101.6 | 102
7.4 |  |  | 1.2 | 2.2 | 3.8 | 100.4 | 101.7 | 102
7.5 |  |  | 1.2 | 2.1 | 3.65 | 100.4 | 101.7 | 102.1
7.6 |  |  | 1.2 | 2.05 | 3.52 | 100.3 | 101.725 | 102.1
7.7 |  |  | 1.1 | 2 | 3.4 | 100.2667 | 101.8 | 102.1
7.8 |  |  | 1.1 | 1.9 | 3.3 | 100.2 | 101.8 | 102.125
7.9 |  |  | 1.1 | 1.9 | 3.2 | 100.2 | 101.8 | 102.2
8 |  |  | 1 | 1.8 | 3.05 | 100.1 | 101.8 | 102.2
8.1 |  |  | 1 | 1.8 | 3 | 100 | 101.8 | 102.2
8.2 |  |  |  | 1.7 | 2.9 | 99.925 | 101.8 | 102.2
8.3 |  |  |  | 1.7 | 2.8 | 99.9 | 101.8 | 102.275
8.4 |  |  |  | 1.6 | 2.7 | 99.78 | 101.8 | 102.3125
8.5 |  |  |  | 1.6 | 2.625 | 99.7 | 101.8 | 102.4
8.6 |  |  |  | 1.6 | 2.6 | 99.5 | 101.9 | 102.475
8.7 |  |  |  | 1.5 | 2.5 | 99.4 | 101.9 | 102.5
8.8 |  |  |  | 1.5 | 2.425 | 99.3 | 101.9 | 102.5
8.9 |  |  |  | 1.5 | 2.4 | 99.14 | 101.9 | 102.5
9 |  |  |  | 1.4 | 2.3 | 98.9 | 101.9 | 102.5
9.1 |  |  |  | 1.4 | 2.2667 | 98.76 | 101.94 | 102.5125
9.2 |  |  |  |  | 2.2 | 98.4667 | 102 | 102.6
9.3 |  |  |  |  | 2.1 | 98.0429 | 102 | 102.6
9.4 |  |  |  |  | 2.1 | 97.8 | 102 | 102.7
9.5 |  |  |  |  | 2.05 | 97.2333 | 102 | 102.7
9.6 |  |  |  |  | 2 | 96.6 | 102 | 102.725
9.7 |  |  |  |  | 1.9 | 95.8 | 102 | 102.8
9.8 |  |  |  |  | 1.9 | 94.64 | 102 | 102.8
9.9 |  |  |  |  | 1.8714 | 93.2 | 102 | 102.8
10 |  |  |  |  | 1.8 | 90.7 | 102 | 102.8
10.1 |  |  |  |  | 1.8 | 86.66 | 102 | 102.8375
10.2 |  |  |  |  | 1.75 | 80.5 | 102 | 102.9
10.3 | [5 empty] | 72.3 | 102 | 102.9
10.4 | [5 empty] | 63.05 | 102 | 102.9286
10.5 | [5 empty] | 52.35 | 102 | 103
10.6 | [5 empty] | 41.65 | 102 | 103
10.7 | [5 empty] | 32.125 | 102 | 103.1
10.8 | [5 empty] | 24.475 | 102 | 103.1
10.9 | [5 empty] | 18.9 | 102 | 103.1
11 | [5 empty] | 15.6714 | 102 | 103.1
11.1 | [5 empty] | 13.6 | 102 | 103.1182
11.2 | [5 empty] | 12.175 | 101.9 | 103.2
11.3 | [5 empty] | 11.14 | 101.9 | 103.2
11.4 | [5 empty] | 10.35 | 101.9 | 103.2857
11.5 | [5 empty] | 9.7 | 101.9 | 103.3
11.6 | [5 empty] | 9.16 | 101.8 | 103.3
11.7 | [5 empty] | 8.8 | 101.8 | 103.3
11.8 | [5 empty] | 8.3714 | 101.8 | 103.3
11.9 | [5 empty] | 8.1 | 101.8 | 103.4
12 | [5 empty] | 7.775 | 101.8 | 103.4
12.1 | [5 empty] | 7.5333 | 101.7 | 103.4286
12.2 | [5 empty] | 7.225 | 101.7 | 103.5
12.3 | [5 empty] | 7.1 | 101.7 | 103.5
12.4 | [5 empty] | 6.85 | 101.7 | 103.6
12.5 | [5 empty] | 6.65 | 101.7 | 103.6
12.6 | [5 empty] | 6.5 | 101.7 | 103.6
12.7 | [5 empty] | 6.32 | 101.6 | 103.7
12.8 | [5 empty] | 6.125 | 101.5 | 103.7
12.9 | [5 empty] | 5.9 | 101.5 | 103.7
13 | [5 empty] | 5.75 | 101.44 | 103.7
13.1 | [5 empty] | 5.6 | 101.4 | 103.8
13.2 | [5 empty] | 5.4714 | 101.4 | 103.8
13.3 | [5 empty] | 5.375 | 101.3 | 103.8
13.4 | [5 empty] | 5.2 | 101.2 | 103.8
13.5 | [5 empty] | 5.0667 | 101.2 | 103.8
13.6 | [5 empty] | 4.9667 | 101.1 | 103.8429
13.7 | [5 empty] | 4.8 | 100.9375 | 103.9
13.8 | [5 empty] | 4.78 | 100.86 | 103.9
13.9 | [5 empty] | 4.6 | 100.6857 | 103.9
14 | [5 empty] | 4.5 | 100.6 | 103.9
14.1 | [5 empty] | 4.4 | 100.4 | 104
14.2 | [5 empty] | 4.3 | 100.26 | 104
14.3 | [5 empty] | 4.2 | 100.0667 | 104
14.4 | [5 empty] | 4.1 | 99.85 | 104
14.5 | [5 empty] | 4 | 99.5 | 104.1
14.6 | [5 empty] | 3.95 | 99.325 | 104.1
14.7 | [5 empty] | 3.875 | 98.85 | 104.1
14.8 | [5 empty] | 3.8 | 98.35 | 104.1
14.9 | [5 empty] | 3.7 | 97.74 | 104.1
15 | [5 empty] | 3.6 | 96.825 | 104.1
15.1 | [5 empty] | 3.525 | 95.64 | 104.1
15.2 | [5 empty] | 3.5 | 93.8 | 104.1
15.3 | [5 empty] | 3.4 | 90.84 | 104.1
15.4 | [5 empty] | 3.32 | 85.975 | 104.1625
15.5 | [6 empty] | 78.825 | 104.2
15.6 | [6 empty] | 69.84 | 104.2
15.7 | [6 empty] | 60 | 104.2
15.8 | [6 empty] | 48.84 | 104.2
15.9 | [6 empty] | 39.08 | 104.2
16 | [6 empty] | 30.5333 | 104.2
16.1 | [6 empty] | 23.5 | 104.2
16.2 | [6 empty] | 19.1333 | 104.2
16.3 | [6 empty] | 16.4 | 104.2
16.4 | [6 empty] | 14.7 | 104.1545
16.5 | [6 empty] | 13.5 | 104.1
16.6 | [6 empty] | 12.65 | 104.1
16.7 | [6 empty] | 11.9 | 104.1
16.8 | [6 empty] | 11.375 | 104.1
16.9 | [6 empty] | 10.95 | 104.1
17 | [6 empty] | 10.5 | 104.1
17.1 | [6 empty] | 10.2286 | 104.1
17.2 | [6 empty] | 9.9 | 104.04
17.3 | [6 empty] | 9.6167 | 104
17.4 | [6 empty] | 9.36 | 104
17.5 | [6 empty] | 9.14 | 103.94
17.6 | [6 empty] | 8.9 | 103.9
17.7 | [6 empty] | 8.65 | 103.9
17.8 | [6 empty] | 8.46 | 103.84
17.9 | [6 empty] | 8.2714 | 103.7714
18 | [6 empty] | 8.075 | 103.7
18.1 | [6 empty] | 7.875 | 103.7
18.2 | [6 empty] | 7.72 | 103.6143
18.3 | [6 empty] | 7.5333 | 103.5
18.4 | [6 empty] | 7.38 | 103.45
18.5 | [6 empty] | 7.175 | 103.3143
18.6 | [6 empty] | 7.02 | 103.2857
18.7 | [6 empty] | 6.8875 | 103.15
18.8 | [6 empty] | 6.725 | 103.0143
18.9 | [6 empty] | 6.56 | 102.8875
19 | [6 empty] | 6.44 | 102.7571
19.1 | [6 empty] | 6.3 | 102.625
19.2 | [6 empty] | 6.1571 | 102.4
19.3 | [6 empty] | 6.06 | 102.225
19.4 | [6 empty] | 5.95 | 101.9571
19.5 | [6 empty] | 5.8 | 101.8
19.6 | [6 empty] | 5.7 | 101.5625
19.7 | [6 empty] | 5.5429 | 101.3125
19.8 | [6 empty] | 5.44 | 100.9
19.9 | [6 empty] | 5.3714 | 100.45
20 | [6 empty] | 5.2667 | 99.95
20.1 | [6 empty] | 5.15 | 99.4
20.2 | [6 empty] | 5.04 | 98.5
20.3 | [6 empty] | 4.95 | 97.3571
20.4 | [6 empty] | 4.9 | 95.8
20.5 | [6 empty] | 4.7571 | 93.3857
20.6 | [6 empty] | 4.7 | 89.2875
20.7 | [6 empty] | 4.6 | 82.7
20.8 | [7 empty] | 74.2
20.9 | [7 empty] | 64.8625
21 | [7 empty] | 53.85
21.1 | [7 empty] | 43.45
21.2 | [7 empty] | 34.6
21.3 | [7 empty] | 27
21.4 | [7 empty] | 21.675
21.5 | [7 empty] | 18.4143
21.6 | [7 empty] | 16.4625
21.7 | [7 empty] | 15.0429
21.8 | [7 empty] | 14.2
21.9 | [7 empty] | 13.35
22 | [7 empty] | 12.7571
22.1 | [7 empty] | 12.2714
22.2 | [7 empty] | 11.9
22.3 | [7 empty] | 11.45
22.4 | [7 empty] | 11.1286
22.5 | [7 empty] | 10.7833
22.6 | [7 empty] | 10.5
22.7 | [7 empty] | 10.3143
22.8 | [7 empty] | 10
22.9 | [7 empty] | 9.825
23 | [7 empty] | 9.5571
23.1 | [7 empty] | 9.3
23.2 | [7 empty] | 9.125
23.3 | [7 empty] | 8.875
23.4 | [7 empty] | 8.6833
23.5 | [7 empty] | 8.5375
23.6 | [7 empty] | 8.325
23.7 | [7 empty] | 8.0714
23.8 | [7 empty] | 7.9
23.9 | [7 empty] | 7.7429
24 | [7 empty] | 7.575
24.1 | [7 empty] | 7.4
24.2 | [7 empty] | 7.2
24.3 | [7 empty] | 7.0714
24.4 | [7 empty] | 6.9375
24.5 | [7 empty] | 6.8
24.6 | [7 empty] | 6.6375
24.7 | [7 empty] | 6.5
24.8 | [7 empty] | 6.3375
24.9 | [7 empty] | 6.26
25 | [7 empty] | 6.0857
25.1 | [7 empty] | 6
25.2 | [7 empty] | 5.9143
25.3 | [7 empty] | 5.7667
25.4 | [7 empty] | 5.7
25.5 | [7 empty] | 5.5
25.6 | [7 empty] | 5.4333
25.7 | [7 empty] | 5.3
25.8 | [7 empty] | 5.3
25.9 | [7 empty] | 5.1
26 | [7 empty] | 5.1
26.1 | [7 empty] | 4.9571
26.2 | [7 empty] | 4.8875
26.3 | [7 empty] | 4.8
26.4 | [7 empty] | 4.75
26.5 | [7 empty] | 4.6875
26.6 | [7 empty] | 4.6
26.7 | [7 empty] | 4.525
26.8 | [7 empty] | 4.5
26.9 | [7 empty] | 4.4
27 | [7 empty] | 4.3375
27.1 | [7 empty] | 4.3
27.2 | [7 empty] | 4.2
27.3 | [7 empty] | 4.2
27.4 | [7 empty] | 4.1
27.5 | [7 empty] | 4
27.6 | [7 empty] | 4
27.7 | [7 empty] | 3.9625
27.8 | [7 empty] | 3.9
27.9 | [7 empty] | 3.9
28 | [7 empty] | 3.8
28.1 | [7 empty] | 3.7
28.2 | [7 empty] | 3.7
28.3 | [7 empty] | 3.7
28.4 | [7 empty] | 3.6
28.5 | [7 empty] | 3.6
28.6 | [7 empty] | 3.5
28.7 | [7 empty] | 3.5
28.8 | [7 empty] | 3.4

## Open Field Profiles at 10cm
Lateral Profiles
Open field
SSD=100cm
Depth=10cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 4.725
-29.7 | [7 empty] | 4.8
-29.6 | [7 empty] | 4.8583
-29.5 | [7 empty] | 4.9556
-29.4 | [7 empty] | 5
-29.3 | [7 empty] | 5.0083
-29.2 | [7 empty] | 5.0917
-29.1 | [7 empty] | 5.1818
-29 | [7 empty] | 5.3
-28.9 | [7 empty] | 5.3
-28.8 | [7 empty] | 5.375
-28.7 | [7 empty] | 5.4286
-28.6 | [7 empty] | 5.5
-28.5 | [7 empty] | 5.6
-28.4 | [7 empty] | 5.7
-28.3 | [7 empty] | 5.7833
-28.2 | [7 empty] | 5.8667
-28.1 | [7 empty] | 5.9286
-28 | [7 empty] | 6
-27.9 | [7 empty] | 6.1286
-27.8 | [7 empty] | 6.2
-27.7 | [7 empty] | 6.3
-27.6 | [7 empty] | 6.4
-27.5 | [7 empty] | 6.5
-27.4 | [7 empty] | 6.6
-27.3 | [7 empty] | 6.7
-27.2 | [7 empty] | 6.8
-27.1 | [7 empty] | 6.9
-27 | [7 empty] | 7.0333
-26.9 | [7 empty] | 7.12
-26.8 | [7 empty] | 7.3143
-26.7 | [7 empty] | 7.4
-26.6 | [7 empty] | 7.5429
-26.5 | [7 empty] | 7.6667
-26.4 | [7 empty] | 7.7583
-26.3 | [7 empty] | 7.9
-26.2 | [7 empty] | 8.125
-26.1 | [7 empty] | 8.27
-26 | [7 empty] | 8.3778
-25.9 | [7 empty] | 8.56
-25.8 | [7 empty] | 8.7455
-25.7 | [7 empty] | 8.8778
-25.6 | [7 empty] | 9.0455
-25.5 | [7 empty] | 9.2556
-25.4 | [7 empty] | 9.4067
-25.3 | [7 empty] | 9.5667
-25.2 | [7 empty] | 9.8
-25.1 | [7 empty] | 9.9857
-25 | [7 empty] | 10.225
-24.9 | [7 empty] | 10.45
-24.8 | [7 empty] | 10.6667
-24.7 | [7 empty] | 10.85
-24.6 | [7 empty] | 11.0714
-24.5 | [7 empty] | 11.325
-24.4 | [7 empty] | 11.54
-24.3 | [7 empty] | 11.7533
-24.2 | [7 empty] | 12.02
-24.1 | [7 empty] | 12.2867
-24 | [7 empty] | 12.5533
-23.9 | [7 empty] | 12.82
-23.8 | [7 empty] | 13.0867
-23.7 | [7 empty] | 13.3727
-23.6 | [7 empty] | 13.7286
-23.5 | [7 empty] | 14.0143
-23.4 | [7 empty] | 14.38
-23.3 | [7 empty] | 14.78
-23.2 | [7 empty] | 15.18
-23.1 | [7 empty] | 15.6267
-23 | [7 empty] | 16.1333
-22.9 | [7 empty] | 16.8
-22.8 | [7 empty] | 17.56
-22.7 | [7 empty] | 18.4667
-22.6 | [7 empty] | 19.8
-22.5 | [7 empty] | 21.8333
-22.4 | [7 empty] | 24.5143
-22.3 | [7 empty] | 28.5857
-22.2 | [7 empty] | 35.2182
-22.1 | [7 empty] | 42.9562
-22 | [7 empty] | 50.9933
-21.9 | [7 empty] | 59.9267
-21.8 | [7 empty] | 68.18
-21.7 | [7 empty] | 75.8091
-21.6 | [7 empty] | 81.9
-21.5 | [7 empty] | 85.4
-21.4 | [6 empty] | 6.4 | 88.0545
-21.3 | [6 empty] | 6.575 | 89.7429
-21.2 | [6 empty] | 6.6333 | 90.8889
-21.1 | [6 empty] | 6.8 | 91.8182
-21 | [6 empty] | 6.9 | 92.5333
-20.9 | [6 empty] | 7 | 93.1143
-20.8 | [6 empty] | 7.1667 | 93.65
-20.7 | [6 empty] | 7.2333 | 94.025
-20.6 | [6 empty] | 7.44 | 94.4
-20.5 | [6 empty] | 7.6 | 94.875
-20.4 | [6 empty] | 7.7 | 95.1875
-20.3 | [6 empty] | 7.8556 | 95.5
-20.2 | [6 empty] | 7.9857 | 95.75
-20.1 | [6 empty] | 8.12 | 96
-20 | [6 empty] | 8.35 | 96.2333
-19.9 | [6 empty] | 8.5 | 96.51
-19.8 | [6 empty] | 8.625 | 96.775
-19.7 | [6 empty] | 8.8 | 96.98
-19.6 | [6 empty] | 9.0429 | 97.1929
-19.5 | [6 empty] | 9.2 | 97.4
-19.4 | [6 empty] | 9.4 | 97.5833
-19.3 | [6 empty] | 9.6 | 97.7333
-19.2 | [6 empty] | 9.8 | 97.8571
-19.1 | [6 empty] | 9.96 | 98.075
-19 | [6 empty] | 10.2 | 98.2222
-18.9 | [6 empty] | 10.4 | 98.3429
-18.8 | [6 empty] | 10.72 | 98.4857
-18.7 | [6 empty] | 10.92 | 98.6333
-18.6 | [6 empty] | 11.15 | 98.775
-18.5 | [6 empty] | 11.375 | 98.8
-18.4 | [6 empty] | 11.66 | 98.9375
-18.3 | [6 empty] | 11.975 | 99.0556
-18.2 | [6 empty] | 12.2333 | 99.1667
-18.1 | [6 empty] | 12.525 | 99.2538
-18 | [6 empty] | 12.775 | 99.3364
-17.9 | [6 empty] | 13.15 | 99.4273
-17.8 | [6 empty] | 13.52 | 99.5222
-17.7 | [6 empty] | 13.9 | 99.6
-17.6 | [6 empty] | 14.4429 | 99.6875
-17.5 | [6 empty] | 14.94 | 99.8
-17.4 | [6 empty] | 15.66 | 99.8167
-17.3 | [6 empty] | 16.35 | 99.9
-17.2 | [6 empty] | 17.4 | 99.98
-17.1 | [6 empty] | 18.85 | 100
-17 | [6 empty] | 21 | 100.06
-16.9 | [6 empty] | 23.85 | 100.1143
-16.8 | [6 empty] | 28.7 | 100.2
-16.7 | [6 empty] | 35.7 | 100.3
-16.6 | [6 empty] | 43.5 | 100.3
-16.5 | [6 empty] | 52.8 | 100.4
-16.4 | [6 empty] | 61.9 | 100.4
-16.3 | [6 empty] | 70.5 | 100.4286
-16.2 | [6 empty] | 77.2667 | 100.5
-16.1 | [6 empty] | 82.72 | 100.6
-16 | [6 empty] | 86.5333 | 100.6
-15.9 | [5 empty] | 4.775 | 88.66 | 100.62
-15.8 | [5 empty] | 4.82 | 90.2 | 100.6867
-15.7 | [5 empty] | 4.9667 | 91.225 | 100.7
-15.6 | [5 empty] | 5.05 | 92.0857 | 100.72
-15.5 | [5 empty] | 5.1143 | 92.75 | 100.7867
-15.4 | [5 empty] | 5.28 | 93.3667 | 100.8889
-15.3 | [5 empty] | 5.3556 | 93.7857 | 100.9
-15.2 | [5 empty] | 5.45 | 94.22 | 100.9
-15.1 | [5 empty] | 5.6 | 94.4 | 100.9
-15 | [5 empty] | 5.7 | 94.85 | 100.9857
-14.9 | [5 empty] | 5.8 | 95.1 | 101
-14.8 | [5 empty] | 5.975 | 95.35 | 101
-14.7 | [5 empty] | 6.0333 | 95.65 | 101
-14.6 | [5 empty] | 6.2 | 95.88 | 101.0111
-14.5 | [5 empty] | 6.4 | 96.06 | 101.1
-14.4 | [5 empty] | 6.5 | 96.2 | 101.1
-14.3 | [5 empty] | 6.6571 | 96.45 | 101.1
-14.2 | [5 empty] | 6.8 | 96.6 | 101.1
-14.1 | [5 empty] | 6.9429 | 96.7 | 101.1
-14 | [5 empty] | 7.1714 | 96.8429 | 101.1375
-13.9 | [5 empty] | 7.3143 | 97.04 | 101.2
-13.8 | [5 empty] | 7.5 | 97.2 | 101.2
-13.7 | [5 empty] | 7.6625 | 97.34 | 101.2
-13.6 | [5 empty] | 7.875 | 97.5 | 101.2
-13.5 | [5 empty] | 8.025 | 97.6 | 101.2
-13.4 | [5 empty] | 8.2 | 97.7 | 101.3
-13.3 | [5 empty] | 8.4833 | 97.85 | 101.3
-13.2 | [5 empty] | 8.7 | 98 | 101.3
-13.1 | [5 empty] | 8.9667 | 98 | 101.3
-13 | [5 empty] | 9.2 | 98.1 | 101.3
-12.9 | [5 empty] | 9.4667 | 98.18 | 101.3
-12.8 | [5 empty] | 9.7333 | 98.2 | 101.3
-12.7 | [5 empty] | 10 | 98.3 | 101.3
-12.6 | [5 empty] | 10.28 | 98.4667 | 101.3
-12.5 | [5 empty] | 10.6 | 98.6 | 101.3
-12.4 | [5 empty] | 10.94 | 98.6286 | 101.3
-12.3 | [5 empty] | 11.28 | 98.8 | 101.3
-12.2 | [5 empty] | 11.7333 | 98.8 | 101.3
-12.1 | [5 empty] | 12.175 | 98.9 | 101.3
-12 | [5 empty] | 12.75 | 98.9 | 101.3
-11.9 | [5 empty] | 13.45 | 99 | 101.3
-11.8 | [5 empty] | 14.2 | 99 | 101.3
-11.7 | [5 empty] | 15.3333 | 99.125 | 101.3
-11.6 | [5 empty] | 16.7 | 99.2 | 101.3
-11.5 | [5 empty] | 18.82 | 99.26 | 101.3
-11.4 | [5 empty] | 22.1667 | 99.3 | 101.3
-11.3 | [5 empty] | 27.06 | 99.4 | 101.3
-11.2 | [5 empty] | 34 | 99.4 | 101.3
-11.1 | [5 empty] | 41.9 | 99.48 | 101.3
-11 | [5 empty] | 50.8 | 99.5 | 101.3
-10.9 | [5 empty] | 60.88 | 99.5 | 101.3
-10.8 | [5 empty] | 69.675 | 99.54 | 101.3
-10.7 | [5 empty] | 76.65 | 99.6 | 101.3
-10.6 | [5 empty] | 82.46 | 99.6 | 101.3
-10.5 | [5 empty] | 86.5 | 99.675 | 101.3
-10.4 |  |  |  |  | 2.6 | 88.64 | 99.7 | 101.3
-10.3 |  |  |  |  | 2.675 | 90.3 | 99.8 | 101.3
-10.2 |  |  |  |  | 2.7 | 91.4 | 99.8 | 101.3
-10.1 |  |  |  |  | 2.8 | 92.36 | 99.9 | 101.3
-10 |  |  |  |  | 2.8 | 92.98 | 99.9 | 101.3
-9.9 |  |  |  |  | 2.9 | 93.62 | 99.9 | 101.3
-9.8 |  |  |  |  | 3 | 94.12 | 99.9 | 101.2
-9.7 |  |  |  |  | 3.1 | 94.5 | 99.94 | 101.2
-9.6 |  |  |  |  | 3.1 | 94.92 | 100 | 101.2
-9.5 |  |  |  |  | 3.2 | 95.14 | 100 | 101.2
-9.4 |  |  |  |  | 3.3 | 95.56 | 100.0429 | 101.2
-9.3 |  |  |  | 2.1 | 3.4 | 95.74 | 100.1 | 101.2
-9.2 |  |  |  | 2.15 | 3.5 | 96.0286 | 100.1 | 101.2
-9.1 |  |  |  | 2.2 | 3.6 | 96.2667 | 100.1 | 101.2
-9 |  |  |  | 2.2 | 3.7 | 96.325 | 100.1833 | 101.1556
-8.9 |  |  |  | 2.3 | 3.8 | 96.5 | 100.2 | 101.1
-8.8 |  |  |  | 2.3667 | 3.9 | 96.8 | 100.2 | 101.1
-8.7 |  |  |  | 2.4 | 4 | 96.8833 | 100.2 | 101.1
-8.6 |  |  |  | 2.5 | 4.1 | 97 | 100.2 | 101.1
-8.5 |  |  |  | 2.5 | 4.25 | 97.2 | 100.2 | 101.1
-8.4 |  |  |  | 2.6 | 4.4 | 97.3 | 100.2 | 101.1
-8.3 |  |  |  | 2.7 | 4.5 | 97.55 | 100.2 | 101.1
-8.2 |  |  | 1.6 | 2.8 | 4.7 | 97.7 | 100.25 | 101.1
-8.1 |  |  | 1.6 | 2.8 | 4.8 | 97.8 | 100.3 | 101.1
-8 |  |  | 1.68 | 2.9 | 5 | 97.9 | 100.3 | 101.0571
-7.9 |  |  | 1.7 | 3 | 5.1 | 97.975 | 100.3 | 101
-7.8 |  |  | 1.7 | 3.1 | 5.3667 | 98.1 | 100.3 | 101
-7.7 |  |  | 1.8 | 3.2 | 5.5 | 98.1444 | 100.3 | 101
-7.6 |  |  | 1.9 | 3.3 | 5.7 | 98.3 | 100.3 | 101
-7.5 |  |  | 1.9 | 3.4 | 5.9 | 98.4167 | 100.3 | 101
-7.4 |  |  | 1.9 | 3.56 | 6.2 | 98.5 | 100.3 | 101
-7.3 |  |  | 2 | 3.6 | 6.375 | 98.6 | 100.3 | 101
-7.2 |  |  | 2.025 | 3.75 | 6.65 | 98.6 | 100.3 | 101
-7.1 |  | 1 | 2.1 | 3.8667 | 6.9 | 98.7 | 100.3 | 101
-7 |  | 1.1 | 2.2 | 4.025 | 7.25 | 98.725 | 100.3 | 101
-6.9 |  | 1.1 | 2.25 | 4.2 | 7.5 | 98.9 | 100.3 | 101
-6.8 |  | 1.1 | 2.3 | 4.3 | 7.8333 | 98.9 | 100.3 | 101
-6.7 |  | 1.2 | 2.4 | 4.5 | 8.24 | 99 | 100.3 | 101
-6.6 | 0.8 | 1.2 | 2.5 | 4.7 | 8.75 | 99 | 100.3 | 101
-6.5 | 0.8 | 1.2 | 2.525 | 4.85 | 9.3 | 99.1 | 100.3 | 100.9333
-6.4 | 0.8 | 1.3 | 2.6 | 5 | 9.9667 | 99.1 | 100.3 | 100.9
-6.3 | 0.9 | 1.3 | 2.725 | 5.225 | 10.8 | 99.2 | 100.3 | 100.9
-6.2 | 0.9 | 1.3 | 2.8 | 5.4 | 12 | 99.3 | 100.3 | 100.9
-6.1 | 0.9 | 1.4 | 2.9 | 5.7 | 13.4 | 99.3 | 100.3 | 100.9
-6 | 0.9 | 1.4 | 3 | 5.9 | 15.6667 | 99.3 | 100.3 | 100.9
-5.9 | 0.95 | 1.4 | 3.1 | 6.1667 | 18.8 | 99.4 | 100.3 | 100.9
-5.8 | 1 | 1.5 | 3.24 | 6.55 | 23.5 | 99.4 | 100.3 | 100.9
-5.7 | 1 | 1.5 | 3.4 | 6.8 | 30.2 | 99.5 | 100.3 | 100.9
-5.6 | 1 | 1.6 | 3.525 | 7.2 | 39 | 99.575 | 100.3 | 100.88
-5.5 | 1.1 | 1.6 | 3.7 | 7.65 | 48.3 | 99.6 | 100.3 | 100.8
-5.4 | 1.1 | 1.7 | 3.8 | 8.2 | 58.9 | 99.6 | 100.3 | 100.8
-5.3 | 1.1 | 1.7667 | 3.9333 | 8.875 | 68.1 | 99.65 | 100.3 | 100.8
-5.2 | 1.1 | 1.8 | 4.1 | 9.65 | 76.4 | 99.7 | 100.3 | 100.7
-5.1 | 1.2 | 1.9 | 4.3 | 10.8 | 82.6667 | 99.8 | 100.32 | 100.7
-5 | 1.2 | 2 | 4.5 | 12.225 | 87.1 | 99.8 | 100.4 | 100.7
-4.9 | 1.275 | 2 | 4.7 | 14.2 | 89.8333 | 99.8 | 100.4 | 100.7
-4.8 | 1.3 | 2.1 | 5 | 17.3 | 91.55 | 99.8286 | 100.4 | 100.7
-4.7 | 1.375 | 2.2333 | 5.2 | 21.8 | 92.9 | 99.9 | 100.4 | 100.7
-4.6 | 1.4 | 2.3 | 5.5 | 28.68 | 93.8667 | 99.9 | 100.35 | 100.7
-4.5 | 1.4667 | 2.4 | 6 | 37.4 | 94.625 | 99.9 | 100.3 | 100.6857
-4.4 | 1.5 | 2.5 | 6.2667 | 47.3 | 95.1333 | 99.9 | 100.3 | 100.6
-4.3 | 1.6 | 2.6 | 6.8 | 57.875 | 95.8 | 99.9 | 100.3 | 100.6
-4.2 | 1.6 | 2.7667 | 7.4667 | 68.05 | 96.0667 | 99.9 | 100.3 | 100.6
-4.1 | 1.7 | 2.9 | 8.2667 | 75.8667 | 96.45 | 99.9 | 100.3 | 100.6
-4 | 1.7667 | 3 | 9.4333 | 82.6 | 96.6333 | 99.9 | 100.3 | 100.6
-3.9 | 1.8 | 3.2 | 10.9 | 87.15 | 97.05 | 99.9 | 100.3 | 100.6
-3.8 | 1.95 | 3.4 | 12.825 | 90.3 | 97.3 | 99.9 | 100.3 | 100.6
-3.7 | 2.1 | 3.6 | 15.8 | 92.125 | 97.45 | 99.9 | 100.3 | 100.6
-3.6 | 2.15 | 3.8 | 20.35 | 93.55 | 97.7 | 99.925 | 100.3 | 100.5667
-3.5 | 2.3 | 4.1 | 27.3 | 94.575 | 97.85 | 100 | 100.275 | 100.5
-3.4 | 2.4 | 4.35 | 36.4 | 95.3333 | 98.1 | 100 | 100.2 | 100.5
-3.3 | 2.5 | 4.75 | 46.45 | 95.9333 | 98.2 | 100 | 100.2 | 100.5
-3.2 | 2.6 | 5.15 | 57 | 96.35 | 98.3 | 100 | 100.2 | 100.4538
-3.1 | 2.8333 | 5.85 | 66.85 | 96.85 | 98.4 | 100 | 100.2 | 100.4
-3 | 3 | 6.6 | 75.75 | 97.15 | 98.5 | 100 | 100.2 | 100.4
-2.9 | 3.3 | 7.6 | 83.025 | 97.45 | 98.7 | 100 | 100.2 | 100.4
-2.8 | 3.5 | 9 | 87.7 | 97.6 | 98.8 | 100 | 100.1 | 100.3625
-2.7 | 3.9 | 11.2 | 90.9 | 97.9 | 98.925 | 100 | 100.1 | 100.3
-2.6 | 4.4 | 14.3 | 93 | 98.1 | 99 | 100 | 100.1 | 100.3
-2.5 | 5.1 | 19 | 94.5 | 98.3667 | 99.1 | 100 | 100.2 | 100.3
-2.4 | 5.9 | 25.6 | 95.3667 | 98.4 | 99.25 | 100 | 100.2 | 100.3
-2.3 | 7.1 | 35.02 | 96.15 | 98.7 | 99.35 | 100 | 100.1 | 100.3
-2.2 | 8.65 | 45.2 | 96.85 | 98.8 | 99.4 | 100 | 100.1 | 100.3
-2.1 | 11.1 | 56.48 | 97.2333 | 98.9 | 99.5 | 100 | 100.1 | 100.3
-2 | 15.1 | 66.86 | 97.7 | 99.0667 | 99.5 | 100.1 | 100.1 | 100.3
-1.9 | 20.8 | 76.4 | 97.9333 | 99.1 | 99.5 | 100.02 | 100.1 | 100.2
-1.8 | 29.45 | 83.6 | 98.275 | 99.3 | 99.6 | 100.1 | 100.1 | 100.2
-1.7 | 39.8 | 89.3 | 98.5 | 99.3 | 99.6 | 100.1 | 100.1 | 100.2
-1.6 | 50.1 | 92.36 | 98.675 | 99.4 | 99.6333 | 100.0429 | 100.1 | 100.2
-1.5 | 61.6 | 94.3 | 98.9 | 99.5 | 99.7 | 100 | 100.1 | 100.2
-1.4 | 71.4 | 95.95 | 99.05 | 99.6 | 99.8 | 100 | 100.1 | 100.15
-1.3 | 80.85 | 97 | 99.2 | 99.7 | 99.8 | 100 | 100.1 | 100.1
-1.2 | 87.4 | 97.8 | 99.3 | 99.7 | 99.9 | 100 | 100.1 | 100.1
-1.1 | 91.8 | 98.3 | 99.4 | 99.8 | 99.9 | 100 | 100 | 100.1
-1 | 94.6 | 98.8 | 99.6 | 99.8 | 99.9 | 100 | 100 | 100.1
-0.9 | 96.45 | 99 | 99.7 | 99.9 | 99.9 | 100 | 100 | 100.1
-0.8 | 97.6 | 99.25 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100.1
-0.7 | 98.25 | 99.45 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100.0143
-0.6 | 98.85 | 99.6 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100
-0.5 | 99.4 | 99.8 | 99.8 | 99.9 | 100 | 100 | 100 | 100
-0.4 | 99.5 | 99.9 | 99.9 | 99.9 | 100 | 100 | 100 | 100
-0.3 | 99.7 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100
-0.2 | 99.9 | 100 | 99.9 | 100 | 100 | 99.975 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 99.9 | 100 | 100 | 100 | 100.05 | 100 | 100
0.3 | 99.9 | 99.9 | 100 | 100 | 100 | 100.1 | 100 | 100
0.4 | 99.6 | 99.8 | 99.9 | 100 | 100 | 100.1 | 100 | 100
0.5 | 99.35 | 99.75 | 99.9 | 99.9 | 100 | 100 | 100.075 | 100
0.6 | 99 | 99.6 | 99.8 | 99.9 | 100 | 100 | 100 | 100
0.7 | 98.3 | 99.4667 | 99.8 | 99.9 | 100 | 100.1 | 100 | 100.0833
0.8 | 97.75 | 99.3 | 99.8 | 99.9 | 100 | 100.1 | 100 | 100.1
0.9 | 96.2667 | 99.1 | 99.75 | 99.9 | 100 | 100.1 | 100.0667 | 100.1
1 | 94.7333 | 98.8333 | 99.6 | 99.9 | 100 | 100.1 | 100.0167 | 100.1
1.1 | 91.875 | 98.3333 | 99.5 | 99.9 | 99.9 | 100.1 | 100.1 | 100.1
1.2 | 87.7667 | 97.7 | 99.5 | 99.825 | 99.9 | 100.1 | 100.1 | 100.1
1.3 | 81.5 | 96.9667 | 99.3 | 99.7 | 99.9 | 100.1 | 100.1 | 100.1
1.4 | 72.3 | 96.1 | 99.1 | 99.7 | 99.8 | 100.1 | 100.1 | 100.1143
1.5 | 61.7 | 94.75 | 99.0333 | 99.6 | 99.8 | 100.1889 | 100.1 | 100.2
1.6 | 50.6 | 92.4 | 98.85 | 99.5 | 99.8 | 100.2 | 100.1 | 100.2
1.7 | 39.6 | 88.9 | 98.7 | 99.4 | 99.7 | 100.2 | 100.125 | 100.2
1.8 | 29.7 | 83.6 | 98.5 | 99.4 | 99.7 | 100.1 | 100.2 | 100.2
1.9 | 21 | 76.45 | 98.2 | 99.2 | 99.7 | 100.1 | 100.2 | 100.2
2 | 15.4 | 66.7333 | 97.95 | 99.1 | 99.7 | 100.1 | 100.2 | 100.3
2.1 | 11.4 | 56.5 | 97.5 | 99 | 99.6 | 100.1 | 100.24 | 100.3
2.2 | 8.8 | 45.825 | 97.12 | 98.9 | 99.5 | 100.12 | 100.3 | 100.3
2.3 | 7.1 | 34.9 | 96.6667 | 98.7667 | 99.4333 | 100.2 | 100.3 | 100.3
2.4 | 6 | 26.1333 | 95.7 | 98.5 | 99.4 | 100.2 | 100.3 | 100.3
2.5 | 5.1 | 19.15 | 94.7 | 98.4 | 99.3 | 100.2 | 100.3 | 100.3
2.6 | 4.4667 | 14.275 | 93.2 | 98.3 | 99.2 | 100.2 | 100.3 | 100.3
2.7 | 4 | 11 | 91 | 98.1 | 99.1 | 100.2 | 100.3 | 100.3
2.8 | 3.5333 | 9.1 | 87.9 | 97.9333 | 99 | 100.2 | 100.3 | 100.4
2.9 | 3.3 | 7.6 | 82.8 | 97.5 | 99 | 100.2 | 100.3 | 100.4
3 | 3.05 | 6.6 | 76.1 | 97.3 | 98.9 | 100.2 | 100.35 | 100.4
3.1 | 2.85 | 5.7667 | 67.3667 | 96.875 | 98.7 | 100.1571 | 100.4 | 100.4
3.2 | 2.6 | 5.2 | 57.1 | 96.45 | 98.6 | 100.1 | 100.4 | 100.4231
3.3 | 2.5 | 4.8 | 45.8 | 96 | 98.4 | 100.1 | 100.4 | 100.5
3.4 | 2.4 | 4.35 | 36.5 | 95.44 | 98.3 | 100.1 | 100.4 | 100.5
3.5 | 2.225 | 4 | 27.5667 | 94.75 | 98.1 | 100.1 | 100.4 | 100.5
3.6 | 2.15 | 3.8 | 20.6 | 93.6 | 97.9 | 100.2 | 100.4 | 100.53
3.7 | 2 | 3.5667 | 15.9 | 92.2 | 97.7 | 100.2 | 100.4 | 100.6
3.8 | 1.9 | 3.4 | 12.8 | 90.2 | 97.42 | 100.1 | 100.42 | 100.6
3.9 | 1.85 | 3.2 | 10.85 | 87.1 | 97.1 | 100.1 | 100.45 | 100.6
4 | 1.8 | 3 | 9.4 | 82.2 | 96.9 | 100.1 | 100.5 | 100.6
4.1 | 1.7 | 2.9 | 8.3333 | 75.4 | 96.5 | 100.1 | 100.5 | 100.6
4.2 | 1.6 | 2.725 | 7.5 | 67.2 | 96.3 | 100.1 | 100.5 | 100.6
4.3 | 1.6 | 2.6 | 6.9 | 57.8667 | 95.95 | 100.1 | 100.5 | 100.6
4.4 | 1.5 | 2.5 | 6.4 | 47.5 | 95.3 | 100.025 | 100.5 | 100.6
4.5 | 1.5 | 2.4 | 5.9 | 37.45 | 94.7 | 100 | 100.5 | 100.65
4.6 | 1.4 | 2.3 | 5.5667 | 28.6 | 94.0667 | 100 | 100.5 | 100.7
4.7 | 1.4 | 2.2 | 5.3 | 22.1 | 92.9667 | 100 | 100.5 | 100.7
4.8 | 1.3 | 2.1 | 5.05 | 17.325 | 91.6667 | 100 | 100.5 | 100.7
4.9 | 1.3 | 2.025 | 4.8 | 14.3333 | 89.6 | 100 | 100.5 | 100.7
5 | 1.2 | 2 | 4.5 | 12.175 | 86.95 | 100 | 100.6 | 100.7
5.1 | 1.2 | 1.9 | 4.3 | 10.7 | 82.5 | 99.9625 | 100.6 | 100.7
5.2 | 1.1 | 1.8 | 4.175 | 9.7 | 76.05 | 99.9 | 100.6 | 100.7
5.3 | 1.1 | 1.8 | 4 | 8.8333 | 67.8 | 99.9 | 100.6 | 100.78
5.4 | 1.1 | 1.7 | 3.8 | 8.25 | 58.7 | 99.9 | 100.6 | 100.8
5.5 | 1.1 | 1.6 | 3.7 | 7.675 | 48.2 | 99.9 | 100.6 | 100.8
5.6 | 1 | 1.6 | 3.5 | 7.2333 | 39.45 | 99.8 | 100.6 | 100.84
5.7 | 1 | 1.5 | 3.4 | 6.8 | 30.1333 | 99.8 | 100.6 | 100.9
5.8 | 1 | 1.5 | 3.3 | 6.55 | 23.8 | 99.7 | 100.6 | 100.9
5.9 | 1 | 1.4 | 3.2 | 6.2 | 18.7 | 99.6 | 100.6 | 100.9
6 | 0.9 | 1.4 | 3.025 | 6 | 15.5 | 99.6 | 100.6 | 100.9
6.1 | 0.9 | 1.4 | 2.9 | 5.7 | 13.4 | 99.5 | 100.6 | 100.9
6.2 | 0.9 | 1.3 | 2.8 | 5.5 | 11.8333 | 99.425 | 100.6 | 100.9
6.3 | 0.9 | 1.3 | 2.7 | 5.2 | 10.8 | 99.4 | 100.6 | 100.9
6.4 | 0.9 | 1.275 | 2.6 | 5.02 | 9.95 | 99.4 | 100.6 | 100.9
6.5 | 0.8 | 1.2 | 2.55 | 4.8 | 9.35 | 99.3 | 100.6 | 100.9
6.6 | 0.8 | 1.2 | 2.5 | 4.6333 | 8.75 | 99.3 | 100.575 | 101
6.7 |  | 1.2 | 2.4 | 4.5 | 8.2667 | 99.2 | 100.5 | 101
6.8 |  | 1.1 | 2.3 | 4.3 | 7.9 | 99.1833 | 100.58 | 101
6.9 |  | 1.1 | 2.2 | 4.2 | 7.5 | 99.1 | 100.5 | 101
7 |  | 1.1 | 2.2 | 4 | 7.2 | 99.075 | 100.5 | 101
7.1 |  | 1 | 2.1 | 3.9 | 6.9 | 98.9 | 100.5 | 101
7.2 |  |  | 2.1 | 3.8 | 6.6333 | 98.8857 | 100.5 | 101
7.3 |  |  | 2 | 3.625 | 6.4 | 98.7 | 100.5 | 101
7.4 |  |  | 1.9 | 3.5667 | 6.2 | 98.6 | 100.5 | 101
7.5 |  |  | 1.9 | 3.44 | 5.95 | 98.5 | 100.5 | 101
7.6 |  |  | 1.8 | 3.3 | 5.7333 | 98.5 | 100.5 | 101
7.7 |  |  | 1.8 | 3.2 | 5.6 | 98.4 | 100.5 | 101
7.8 |  |  | 1.8 | 3.1 | 5.3333 | 98.26 | 100.4 | 101
7.9 |  |  | 1.7 | 3 | 5.2 | 98.15 | 100.4 | 101
8 |  |  | 1.65 | 2.9 | 5.05 | 98.025 | 100.4 | 101.025
8.1 |  |  | 1.6 | 2.8667 | 4.9 | 97.9 | 100.4 | 101.1
8.2 |  |  | 1.6 | 2.725 | 4.7 | 97.8 | 100.4 | 101.1
8.3 |  |  |  | 2.7 | 4.525 | 97.7 | 100.4 | 101.1
8.4 |  |  |  | 2.6 | 4.4 | 97.6 | 100.4 | 101.1
8.5 |  |  |  | 2.6 | 4.3 | 97.4 | 100.3 | 101.1
8.6 |  |  |  | 2.5 | 4.1667 | 97.275 | 100.3 | 101.1
8.7 |  |  |  | 2.4 | 4.025 | 97.125 | 100.3 | 101.1
8.8 |  |  |  | 2.3 | 3.9 | 96.875 | 100.3 | 101.1
8.9 |  |  |  | 2.3 | 3.8 | 96.625 | 100.25 | 101.1
9 |  |  |  | 2.2 | 3.7 | 96.46 | 100.2 | 101.1222
9.1 |  |  |  | 2.2 | 3.6 | 96.28 | 100.2 | 101.2
9.2 |  |  |  | 2.1 | 3.5 | 96.125 | 100.1 | 101.2
9.3 |  |  |  | 2.1 | 3.4 | 95.8667 | 100.1 | 101.2
9.4 |  |  |  |  | 3.3 | 95.65 | 100.1 | 101.2
9.5 |  |  |  |  | 3.2 | 95.3 | 100.2 | 101.2
9.6 |  |  |  |  | 3.1 | 95.04 | 100.2 | 101.2
9.7 |  |  |  |  | 3.1 | 94.6 | 100.1 | 101.2
9.8 |  |  |  |  | 3.0333 | 94.2125 | 100.1 | 101.2
9.9 |  |  |  |  | 2.9 | 93.75 | 100.1 | 101.275
10 |  |  |  |  | 2.9 | 93.15 | 100.06 | 101.3
10.1 |  |  |  |  | 2.8 | 92.5 | 100 | 101.3
10.2 |  |  |  |  | 2.7 | 91.5 | 99.92 | 101.3
10.3 |  |  |  |  | 2.7 | 90.4 | 99.9 | 101.3
10.4 |  |  |  |  | 2.6 | 88.7857 | 99.825 | 101.3
10.5 | [5 empty] | 86.46 | 99.8 | 101.3
10.6 | [5 empty] | 82.8 | 99.8 | 101.3
10.7 | [5 empty] | 76.8 | 99.7 | 101.3
10.8 | [5 empty] | 69.6667 | 99.64 | 101.3
10.9 | [5 empty] | 60.46 | 99.6 | 101.3
11 | [5 empty] | 51.54 | 99.6 | 101.3
11.1 | [5 empty] | 42.175 | 99.5 | 101.3
11.2 | [5 empty] | 33.825 | 99.4833 | 101.3
11.3 | [5 empty] | 26.9 | 99.4 | 101.3
11.4 | [5 empty] | 21.925 | 99.3375 | 101.3
11.5 | [5 empty] | 18.675 | 99.3 | 101.3
11.6 | [5 empty] | 16.725 | 99.2 | 101.3
11.7 | [5 empty] | 15.3 | 99.2 | 101.3
11.8 | [5 empty] | 14.2 | 99.2 | 101.3
11.9 | [5 empty] | 13.42 | 99.1 | 101.3
12 | [5 empty] | 12.66 | 99 | 101.3
12.1 | [5 empty] | 12.2 | 98.9 | 101.3
12.2 | [5 empty] | 11.66 | 98.8 | 101.3
12.3 | [5 empty] | 11.3 | 98.7 | 101.3
12.4 | [5 empty] | 10.94 | 98.625 | 101.3
12.5 | [5 empty] | 10.6 | 98.6 | 101.3
12.6 | [5 empty] | 10.25 | 98.58 | 101.3
12.7 | [5 empty] | 9.975 | 98.5 | 101.3
12.8 | [5 empty] | 9.6857 | 98.4 | 101.3
12.9 | [5 empty] | 9.4333 | 98.3 | 101.3
13 | [5 empty] | 9.25 | 98.1 | 101.3
13.1 | [5 empty] | 9 | 98.0143 | 101.3
13.2 | [5 empty] | 8.75 | 97.975 | 101.3
13.3 | [5 empty] | 8.525 | 97.8 | 101.3
13.4 | [5 empty] | 8.3 | 97.7 | 101.3
13.5 | [5 empty] | 8.0857 | 97.6 | 101.2375
13.6 | [5 empty] | 7.9 | 97.5 | 101.2
13.7 | [5 empty] | 7.7 | 97.34 | 101.2
13.8 | [5 empty] | 7.5 | 97.3 | 101.2
13.9 | [5 empty] | 7.3 | 97.05 | 101.2
14 | [5 empty] | 7.15 | 96.9 | 101.1
14.1 | [5 empty] | 7 | 96.85 | 101.1
14.2 | [5 empty] | 6.8 | 96.6 | 101.1
14.3 | [5 empty] | 6.625 | 96.4 | 101.1
14.4 | [5 empty] | 6.475 | 96.25 | 101.1
14.5 | [5 empty] | 6.4 | 96 | 101.1
14.6 | [5 empty] | 6.25 | 95.8 | 101.0444
14.7 | [5 empty] | 6.1 | 95.55 | 101
14.8 | [5 empty] | 6 | 95.4 | 101
14.9 | [5 empty] | 5.86 | 95.15 | 101
15 | [5 empty] | 5.72 | 94.8 | 101
15.1 | [5 empty] | 5.6 | 94.5 | 100.9
15.2 | [5 empty] | 5.5 | 94.1 | 100.9
15.3 | [5 empty] | 5.325 | 93.8667 | 100.9
15.4 | [5 empty] | 5.2 | 93.35 | 100.9
15.5 | [5 empty] | 5.1556 | 92.7333 | 100.825
15.6 | [5 empty] | 5.0375 | 92.0429 | 100.7533
15.7 | [5 empty] | 4.95 | 91.3 | 100.7
15.8 | [5 empty] | 4.88 | 90.2 | 100.7
15.9 | [5 empty] | 4.7143 | 88.5833 | 100.65
16 | [6 empty] | 86.25 | 100.6
16.1 | [6 empty] | 82.7 | 100.5286
16.2 | [6 empty] | 77.6 | 100.5
16.3 | [6 empty] | 70.16 | 100.4625
16.4 | [6 empty] | 62.2 | 100.4
16.5 | [6 empty] | 53.1 | 100.4
16.6 | [6 empty] | 43.8 | 100.3
16.7 | [6 empty] | 35.7 | 100.3
16.8 | [6 empty] | 28.8 | 100.2143
16.9 | [6 empty] | 24 | 100.1
17 | [6 empty] | 20.8 | 100.08
17.1 | [6 empty] | 18.8 | 100
17.2 | [6 empty] | 17.3667 | 100
17.3 | [6 empty] | 16.4 | 99.9
17.4 | [6 empty] | 15.58 | 99.8
17.5 | [6 empty] | 14.9 | 99.7833
17.6 | [6 empty] | 14.3857 | 99.7
17.7 | [6 empty] | 13.9 | 99.6
17.8 | [6 empty] | 13.45 | 99.5545
17.9 | [6 empty] | 13.12 | 99.4636
18 | [6 empty] | 12.8 | 99.375
18.1 | [6 empty] | 12.4 | 99.2917
18.2 | [6 empty] | 12.1667 | 99.2083
18.3 | [6 empty] | 11.8 | 99
18.4 | [6 empty] | 11.55 | 98.9625
18.5 | [6 empty] | 11.32 | 98.8286
18.6 | [6 empty] | 11.0571 | 98.8
18.7 | [6 empty] | 10.875 | 98.6714
18.8 | [6 empty] | 10.65 | 98.4571
18.9 | [6 empty] | 10.42 | 98.275
19 | [6 empty] | 10.1667 | 98.15
19.1 | [6 empty] | 9.98 | 98.025
19.2 | [6 empty] | 9.7556 | 97.8875
19.3 | [6 empty] | 9.525 | 97.7667
19.4 | [6 empty] | 9.32 | 97.6
19.5 | [6 empty] | 9.14 | 97.3667
19.6 | [6 empty] | 9 | 97.2
19.7 | [6 empty] | 8.78 | 96.95
19.8 | [6 empty] | 8.56 | 96.74
19.9 | [6 empty] | 8.42 | 96.5
20 | [6 empty] | 8.26 | 96.27
20.1 | [6 empty] | 8.125 | 95.975
20.2 | [6 empty] | 7.94 | 95.7571
20.3 | [6 empty] | 7.825 | 95.55
20.4 | [6 empty] | 7.64 | 95.2
20.5 | [6 empty] | 7.54 | 94.8429
20.6 | [6 empty] | 7.3714 | 94.4143
20.7 | [6 empty] | 7.2 | 94.1
20.8 | [6 empty] | 7.1 | 93.6
20.9 | [6 empty] | 6.975 | 93.1
21 | [6 empty] | 6.84 | 92.5615
21.1 | [6 empty] | 6.7 | 91.7875
21.2 | [6 empty] | 6.5571 | 90.95
21.3 | [6 empty] | 6.475 | 89.8
21.4 | [6 empty] | 6.325 | 88.0923
21.5 | [7 empty] | 85.5538
21.6 | [7 empty] | 81.4
21.7 | [7 empty] | 75.5231
21.8 | [7 empty] | 68.2667
21.9 | [7 empty] | 59.9333
22 | [7 empty] | 51.06
22.1 | [7 empty] | 42.5267
22.2 | [7 empty] | 34.5933
22.3 | [7 empty] | 28.8091
22.4 | [7 empty] | 24.82
22.5 | [7 empty] | 21.74
22.6 | [7 empty] | 19.8733
22.7 | [7 empty] | 18.5933
22.8 | [7 empty] | 17.58
22.9 | [7 empty] | 16.8467
23 | [7 empty] | 16.22
23.1 | [7 empty] | 15.64
23.2 | [7 empty] | 15.1067
23.3 | [7 empty] | 14.7333
23.4 | [7 empty] | 14.4
23.5 | [7 empty] | 14.0667
23.6 | [7 empty] | 13.7333
23.7 | [7 empty] | 13.4154
23.8 | [7 empty] | 13.12
23.9 | [7 empty] | 12.8533
24 | [7 empty] | 12.5867
24.1 | [7 empty] | 12.32
24.2 | [7 empty] | 12.0533
24.3 | [7 empty] | 11.7867
24.4 | [7 empty] | 11.52
24.5 | [7 empty] | 11.2533
24.6 | [7 empty] | 11.0222
24.7 | [7 empty] | 10.8714
24.8 | [7 empty] | 10.6273
24.9 | [7 empty] | 10.42
25 | [7 empty] | 10.1533
25.1 | [7 empty] | 10.0111
25.2 | [7 empty] | 9.7875
25.3 | [7 empty] | 9.6455
25.4 | [7 empty] | 9.4571
25.5 | [7 empty] | 9.2429
25.6 | [7 empty] | 9.1
25.7 | [7 empty] | 8.9333
25.8 | [7 empty] | 8.7
25.9 | [7 empty] | 8.5333
26 | [7 empty] | 8.42
26.1 | [7 empty] | 8.2
26.2 | [7 empty] | 8.1
26.3 | [7 empty] | 7.9636
26.4 | [7 empty] | 7.7917
26.5 | [7 empty] | 7.7083
26.6 | [7 empty] | 7.5833
26.7 | [7 empty] | 7.4
26.8 | [7 empty] | 7.3
26.9 | [7 empty] | 7.16
27 | [7 empty] | 7.0667
27.1 | [7 empty] | 6.9
27.2 | [7 empty] | 6.8
27.3 | [7 empty] | 6.7
27.4 | [7 empty] | 6.6
27.5 | [7 empty] | 6.5
27.6 | [7 empty] | 6.4
27.7 | [7 empty] | 6.3
27.8 | [7 empty] | 6.1875
27.9 | [7 empty] | 6.1
28 | [7 empty] | 6.0167
28.1 | [7 empty] | 5.9571
28.2 | [7 empty] | 5.9
28.3 | [7 empty] | 5.8
28.4 | [7 empty] | 5.6625
28.5 | [7 empty] | 5.6
28.6 | [7 empty] | 5.5
28.7 | [7 empty] | 5.4667
28.8 | [7 empty] | 5.4
28.9 | [7 empty] | 5.3
29 | [7 empty] | 5.2375
29.1 | [7 empty] | 5.2
29.2 | [7 empty] | 5.1308
29.3 | [7 empty] | 5.0333
29.4 | [7 empty] | 5
29.5 | [7 empty] | 4.9
29.6 | [7 empty] | 4.89
29.7 | [7 empty] | 4.8

## Open Field Profiles at 20cm
Lateral Profiles
Open field
SSD=100cm
Depth=20cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 9.31
-29.7 | [7 empty] | 9.4091
-29.6 | [7 empty] | 9.5
-29.5 | [7 empty] | 9.66
-29.4 | [7 empty] | 9.8
-29.3 | [7 empty] | 9.925
-29.2 | [7 empty] | 10.04
-29.1 | [7 empty] | 10.1875
-29 | [7 empty] | 10.35
-28.9 | [7 empty] | 10.5571
-28.8 | [7 empty] | 10.7286
-28.7 | [7 empty] | 10.8
-28.6 | [7 empty] | 11.0143
-28.5 | [7 empty] | 11.1
-28.4 | [7 empty] | 11.3143
-28.3 | [7 empty] | 11.5
-28.2 | [7 empty] | 11.675
-28.1 | [7 empty] | 11.825
-28 | [7 empty] | 12.0429
-27.9 | [7 empty] | 12.25
-27.8 | [7 empty] | 12.4
-27.7 | [7 empty] | 12.575
-27.6 | [7 empty] | 12.8
-27.5 | [7 empty] | 12.96
-27.4 | [7 empty] | 13.2
-27.3 | [7 empty] | 13.4375
-27.2 | [7 empty] | 13.625
-27.1 | [7 empty] | 13.9
-27 | [7 empty] | 14.06
-26.9 | [7 empty] | 14.3286
-26.8 | [7 empty] | 14.525
-26.7 | [7 empty] | 14.775
-26.6 | [7 empty] | 15.05
-26.5 | [7 empty] | 15.2571
-26.4 | [7 empty] | 15.475
-26.3 | [7 empty] | 15.76
-26.2 | [7 empty] | 16.05
-26.1 | [7 empty] | 16.3143
-26 | [7 empty] | 16.5556
-25.9 | [7 empty] | 16.8909
-25.8 | [7 empty] | 17.175
-25.7 | [7 empty] | 17.4375
-25.6 | [7 empty] | 17.8
-25.5 | [7 empty] | 18.075
-25.4 | [7 empty] | 18.4429
-25.3 | [7 empty] | 18.8714
-25.2 | [7 empty] | 19.2
-25.1 | [7 empty] | 19.6714
-25 | [7 empty] | 20.2
-24.9 | [7 empty] | 20.7833
-24.8 | [7 empty] | 21.5571
-24.7 | [7 empty] | 22.4125
-24.6 | [7 empty] | 23.6286
-24.5 | [7 empty] | 25.325
-24.4 | [7 empty] | 27.9429
-24.3 | [7 empty] | 31.375
-24.2 | [7 empty] | 36.3875
-24.1 | [7 empty] | 42.225
-24 | [7 empty] | 48.475
-23.9 | [7 empty] | 55.3
-23.8 | [7 empty] | 61.8857
-23.7 | [7 empty] | 67.3143
-23.6 | [7 empty] | 72.1
-23.5 | [7 empty] | 75.1
-23.4 | [7 empty] | 77.2467
-23.3 | [7 empty] | 78.94
-23.2 | [7 empty] | 80.0733
-23.1 | [7 empty] | 80.9222
-23 | [7 empty] | 81.6545
-22.9 | [7 empty] | 82.2444
-22.8 | [7 empty] | 82.8375
-22.7 | [7 empty] | 83.2286
-22.6 | [6 empty] | 9.9 | 83.7
-22.5 | [6 empty] | 10.0167 | 84.1
-22.4 | [6 empty] | 10.225 | 84.5571
-22.3 | [6 empty] | 10.36 | 84.8
-22.2 | [6 empty] | 10.56 | 85.2
-22.1 | [6 empty] | 10.7667 | 85.4714
-22 | [6 empty] | 10.96 | 85.7375
-21.9 | [6 empty] | 11.1286 | 86.075
-21.8 | [6 empty] | 11.3 | 86.4143
-21.7 | [6 empty] | 11.5111 | 86.7
-21.6 | [6 empty] | 11.64 | 86.9571
-21.5 | [6 empty] | 11.86 | 87.2429
-21.4 | [6 empty] | 12.05 | 87.5
-21.3 | [6 empty] | 12.3 | 87.775
-21.2 | [6 empty] | 12.5 | 88.025
-21.1 | [6 empty] | 12.7 | 88.26
-21 | [6 empty] | 12.9333 | 88.475
-20.9 | [6 empty] | 13.14 | 88.6714
-20.8 | [6 empty] | 13.4 | 88.9
-20.7 | [6 empty] | 13.625 | 89.15
-20.6 | [6 empty] | 13.8 | 89.3571
-20.5 | [6 empty] | 14.125 | 89.575
-20.4 | [6 empty] | 14.35 | 89.7143
-20.3 | [6 empty] | 14.58 | 89.95
-20.2 | [6 empty] | 14.8667 | 90.1857
-20.1 | [6 empty] | 15.1714 | 90.3143
-20 | [6 empty] | 15.4 | 90.4571
-19.9 | [6 empty] | 15.7 | 90.5875
-19.8 | [6 empty] | 15.95 | 90.8286
-19.7 | [6 empty] | 16.2 | 91.05
-19.6 | [6 empty] | 16.55 | 91.175
-19.5 | [6 empty] | 16.9 | 91.325
-19.4 | [6 empty] | 17.2667 | 91.5429
-19.3 | [6 empty] | 17.7 | 91.675
-19.2 | [6 empty] | 18.1 | 91.8125
-19.1 | [6 empty] | 18.5 | 91.9375
-19 | [6 empty] | 19.0667 | 92.1429
-18.9 | [6 empty] | 19.7 | 92.2429
-18.8 | [6 empty] | 20.48 | 92.375
-18.7 | [6 empty] | 21.5 | 92.52
-18.6 | [6 empty] | 22.8 | 92.7
-18.5 | [6 empty] | 24.6 | 92.7571
-18.4 | [6 empty] | 27.3 | 92.8875
-18.3 | [6 empty] | 31.2 | 93.125
-18.2 | [6 empty] | 36.5 | 93.2
-18.1 | [6 empty] | 43.05 | 93.375
-18 | [6 empty] | 50.26 | 93.525
-17.9 | [6 empty] | 57.5 | 93.6
-17.8 | [6 empty] | 64.34 | 93.6875
-17.7 | [6 empty] | 70.2 | 93.75
-17.6 | [6 empty] | 74.6 | 93.9571
-17.5 | [6 empty] | 78.1 | 94.0875
-17.4 | [6 empty] | 80.1 | 94.2125
-17.3 | [6 empty] | 81.5143 | 94.33
-17.2 | [6 empty] | 82.65 | 94.4214
-17.1 | [6 empty] | 83.55 | 94.4929
-17 | [6 empty] | 84.28 | 94.5818
-16.9 | [5 empty] | 7.5 | 84.85 | 94.7
-16.8 | [5 empty] | 7.6667 | 85.4 | 94.7
-16.7 | [5 empty] | 7.82 | 85.8333 | 94.85
-16.6 | [5 empty] | 7.9 | 86.28 | 94.9583
-16.5 | [5 empty] | 8.1 | 86.6 | 95.0417
-16.4 | [5 empty] | 8.2 | 86.98 | 95.1273
-16.3 | [5 empty] | 8.3286 | 87.275 | 95.2143
-16.2 | [5 empty] | 8.5 | 87.64 | 95.2857
-16.1 | [5 empty] | 8.65 | 87.975 | 95.4231
-16 | [5 empty] | 8.9 | 88.225 | 95.5
-15.9 | [5 empty] | 9.075 | 88.575 | 95.5625
-15.8 | [5 empty] | 9.2 | 88.725 | 95.6875
-15.7 | [5 empty] | 9.3 | 89.05 | 95.725
-15.6 | [5 empty] | 9.4714 | 89.2333 | 95.8571
-15.5 | [5 empty] | 9.7 | 89.5 | 95.9
-15.4 | [5 empty] | 9.86 | 89.7667 | 96
-15.3 | [5 empty] | 10.075 | 90 | 96.05
-15.2 | [5 empty] | 10.3 | 90.2667 | 96.1
-15.1 | [5 empty] | 10.4857 | 90.35 | 96.2
-15 | [5 empty] | 10.7 | 90.6 | 96.25
-14.9 | [5 empty] | 10.9 | 90.82 | 96.375
-14.8 | [5 empty] | 11.175 | 91.02 | 96.4
-14.7 | [5 empty] | 11.325 | 91.15 | 96.5375
-14.6 | [5 empty] | 11.55 | 91.4 | 96.6
-14.5 | [5 empty] | 11.85 | 91.5125 | 96.7
-14.4 | [5 empty] | 12.1 | 91.72 | 96.7714
-14.3 | [5 empty] | 12.35 | 91.96 | 96.8
-14.2 | [5 empty] | 12.5286 | 92.1667 | 96.8
-14.1 | [5 empty] | 12.8143 | 92.325 | 96.9
-14 | [5 empty] | 13.2 | 92.525 | 96.9
-13.9 | [5 empty] | 13.38 | 92.6 | 96.9625
-13.8 | [5 empty] | 13.7 | 92.7 | 97
-13.7 | [5 empty] | 13.98 | 92.825 | 97.1
-13.6 | [5 empty] | 14.35 | 93.025 | 97.1571
-13.5 | [5 empty] | 14.7 | 93.18 | 97.2
-13.4 | [5 empty] | 15.05 | 93.3714 | 97.225
-13.3 | [5 empty] | 15.45 | 93.52 | 97.3
-13.2 | [5 empty] | 15.95 | 93.7125 | 97.3857
-13.1 | [5 empty] | 16.45 | 93.86 | 97.4143
-13 | [5 empty] | 17 | 94 | 97.5
-12.9 | [5 empty] | 17.72 | 94.1 | 97.5
-12.8 | [5 empty] | 18.5 | 94.2 | 97.6
-12.7 | [5 empty] | 19.5 | 94.3 | 97.6133
-12.6 | [5 empty] | 20.98 | 94.5 | 97.68
-12.5 | [5 empty] | 23.04 | 94.5125 | 97.7875
-12.4 | [5 empty] | 25.85 | 94.7 | 97.8
-12.3 | [5 empty] | 29.925 | 94.8 | 97.85
-12.2 | [5 empty] | 35.66 | 94.9 | 97.9
-12.1 | [5 empty] | 43.025 | 95 | 97.9438
-12 | [5 empty] | 50.64 | 95.2 | 98
-11.9 | [5 empty] | 58 | 95.3 | 98
-11.8 | [5 empty] | 65.85 | 95.4 | 98.0462
-11.7 | [5 empty] | 71.84 | 95.5 | 98.125
-11.6 | [5 empty] | 76.72 | 95.6 | 98.2067
-11.5 | [5 empty] | 80.4667 | 95.7 | 98.2733
-11.4 | [5 empty] | 82.8 | 95.75 | 98.3
-11.3 | [5 empty] | 84.425 | 95.9 | 98.3
-11.2 | [5 empty] | 85.6 | 96 | 98.3
-11.1 | [5 empty] | 86.5429 | 96.1143 | 98.3667
-11 | [5 empty] | 87.4 | 96.28 | 98.4
-10.9 |  |  |  |  | 4.5 | 87.95 | 96.35 | 98.4
-10.8 |  |  |  |  | 4.6 | 88.48 | 96.425 | 98.44
-10.7 |  |  |  |  | 4.7 | 88.95 | 96.5 | 98.5
-10.6 |  |  |  |  | 4.8 | 89.35 | 96.575 | 98.5778
-10.5 |  |  |  |  | 4.9 | 89.75 | 96.6 | 98.6
-10.4 |  |  |  |  | 5 | 90.14 | 96.7 | 98.6429
-10.3 |  |  |  |  | 5.1 | 90.4857 | 96.8 | 98.7
-10.2 |  |  |  |  | 5.2 | 90.9 | 96.875 | 98.7
-10.1 |  |  |  |  | 5.34 | 91.1 | 96.9 | 98.7
-10 |  |  |  |  | 5.475 | 91.38 | 96.975 | 98.7909
-9.9 |  |  |  |  | 5.525 | 91.725 | 97.025 | 98.8
-9.8 |  |  |  |  | 5.7 | 91.9 | 97.2 | 98.8889
-9.7 |  |  |  | 3.7 | 5.825 | 92.14 | 97.3 | 98.9
-9.6 |  |  |  | 3.8 | 6 | 92.4333 | 97.4 | 98.9
-9.5 |  |  |  | 3.8 | 6.1 | 92.64 | 97.45 | 98.9
-9.4 |  |  |  | 3.9 | 6.275 | 92.8571 | 97.5 | 98.9429
-9.3 |  |  |  | 4 | 6.475 | 93.12 | 97.6 | 99
-9.2 |  |  |  | 4.1 | 6.6 | 93.34 | 97.6 | 99
-9.1 |  |  |  | 4.2 | 6.7 | 93.5 | 97.7 | 99.0182
-9 |  |  |  | 4.3 | 6.84 | 93.6375 | 97.7 | 99.1
-8.9 |  |  |  | 4.4 | 7.06 | 93.975 | 97.7 | 99.1
-8.8 |  |  |  | 4.5333 | 7.3286 | 94.1333 | 97.8667 | 99.1375
-8.7 |  |  |  | 4.6 | 7.5 | 94.3 | 97.9 | 99.2
-8.6 |  |  |  | 4.7 | 7.6714 | 94.54 | 98 | 99.2
-8.5 |  |  | 2.8 | 4.8667 | 7.86 | 94.7 | 98.0667 | 99.2
-8.4 |  |  | 2.9 | 5 | 8.1 | 94.85 | 98.1 | 99.2625
-8.3 |  |  | 2.9 | 5.15 | 8.3 | 95 | 98.2 | 99.3
-8.2 |  |  | 3 | 5.25 | 8.5 | 95.2125 | 98.2 | 99.3
-8.1 |  |  | 3.1 | 5.4 | 8.8 | 95.3 | 98.2571 | 99.3
-8 |  |  | 3.1333 | 5.5 | 9 | 95.475 | 98.3 | 99.3875
-7.9 |  |  | 3.2 | 5.6 | 9.24 | 95.58 | 98.4 | 99.4
-7.8 |  |  | 3.3 | 5.8 | 9.5857 | 95.7 | 98.4 | 99.4
-7.7 |  |  | 3.4 | 6 | 9.8714 | 95.8857 | 98.5 | 99.4
-7.6 |  |  | 3.5 | 6.125 | 10.1 | 96.06 | 98.5 | 99.4143
-7.5 |  |  | 3.6 | 6.3 | 10.5 | 96.2 | 98.55 | 99.5
-7.4 |  |  | 3.65 | 6.5 | 10.86 | 96.3 | 98.6 | 99.5
-7.3 |  | 1.8 | 3.7 | 6.675 | 11.2714 | 96.5 | 98.7 | 99.5
-7.2 |  | 1.9 | 3.8 | 6.9 | 11.6667 | 96.5 | 98.7 | 99.5429
-7.1 |  | 1.9 | 4 | 7.12 | 12.2 | 96.6 | 98.775 | 99.6
-7 |  | 2 | 4.1 | 7.3 | 12.8333 | 96.7 | 98.8 | 99.6
-6.9 |  | 2 | 4.1 | 7.5 | 13.44 | 96.9 | 98.8 | 99.6
-6.8 |  | 2.1 | 4.3 | 7.7 | 14.325 | 97.0333 | 98.8 | 99.6
-6.7 | 1.375 | 2.1 | 4.4 | 7.9333 | 15.6 | 97.175 | 98.9 | 99.6125
-6.6 | 1.4 | 2.1 | 4.5 | 8.2 | 17.2 | 97.225 | 99 | 99.7
-6.5 | 1.45 | 2.2 | 4.65 | 8.5 | 19.175 | 97.375 | 99.05 | 99.7
-6.4 | 1.5 | 2.2 | 4.8 | 8.8 | 22.3 | 97.5 | 99.1 | 99.7125
-6.3 | 1.5 | 2.3 | 4.9 | 9.1 | 26.84 | 97.6 | 99.1 | 99.8
-6.2 | 1.5 | 2.4 | 5.0667 | 9.4333 | 33.125 | 97.7 | 99.1333 | 99.8
-6.1 | 1.5333 | 2.4 | 5.2 | 9.8 | 40.775 | 97.8 | 99.2 | 99.8
-6 | 1.6 | 2.5 | 5.4 | 10.25 | 49.625 | 97.94 | 99.2 | 99.9
-5.9 | 1.6 | 2.6 | 5.6 | 10.75 | 58.5 | 98 | 99.2 | 99.9
-5.8 | 1.7 | 2.7 | 5.7 | 11.3667 | 67.025 | 98.1 | 99.3 | 99.9
-5.7 | 1.7 | 2.7 | 5.9 | 12 | 73.65 | 98.2 | 99.3 | 99.9
-5.6 | 1.7333 | 2.8 | 6.1 | 12.8667 | 79.8 | 98.24 | 99.4 | 99.9
-5.5 | 1.8 | 2.9 | 6.3333 | 14 | 84.1667 | 98.3 | 99.4 | 99.98
-5.4 | 1.9 | 3 | 6.6 | 15.5 | 86.8 | 98.4 | 99.4857 | 100
-5.3 | 1.9 | 3.1 | 6.8333 | 17.6 | 88.8 | 98.5 | 99.5 | 100
-5.2 | 2 | 3.2 | 7 | 20.6 | 90.3 | 98.5125 | 99.5 | 100
-5.1 | 2 | 3.3 | 7.3 | 25.15 | 91.35 | 98.6 | 99.5 | 100
-5 | 2.1 | 3.375 | 7.6 | 31.25 | 92.14 | 98.625 | 99.5375 | 100
-4.9 | 2.1 | 3.5 | 7.9667 | 39.05 | 92.8429 | 98.7 | 99.6 | 100
-4.8 | 2.2 | 3.6 | 8.3667 | 47.7333 | 93.4 | 98.8 | 99.6 | 100
-4.7 | 2.3 | 3.725 | 8.85 | 57.35 | 93.86 | 98.8 | 99.7 | 100
-4.6 | 2.375 | 3.875 | 9.4 | 66.1667 | 94.4 | 98.9 | 99.7 | 100
-4.5 | 2.4 | 4 | 10.0667 | 73.9 | 94.725 | 98.9 | 99.7 | 100
-4.4 | 2.575 | 4.15 | 11 | 80.55 | 95.05 | 99 | 99.7 | 100
-4.3 | 2.6 | 4.3 | 12.1333 | 85.1 | 95.375 | 99.0833 | 99.7 | 100
-4.2 | 2.7 | 4.5 | 13.5 | 88.2333 | 95.65 | 99.1 | 99.7 | 100
-4.1 | 2.8 | 4.7 | 15.75 | 90.2 | 96 | 99.1 | 99.7 | 100.1
-4 | 2.825 | 4.9 | 18.6 | 91.7 | 96.24 | 99.2 | 99.7 | 100.1
-3.9 | 3 | 5.1667 | 23.2 | 92.8 | 96.425 | 99.2 | 99.7 | 100.1
-3.8 | 3.1 | 5.35 | 29.4 | 93.6 | 96.6333 | 99.2 | 99.7 | 100.1
-3.7 | 3.2 | 5.6333 | 37.9 | 94.3 | 96.7667 | 99.3 | 99.8 | 100.1
-3.6 | 3.32 | 6 | 46.9 | 94.9 | 96.95 | 99.3 | 99.8 | 100.1
-3.5 | 3.5 | 6.4667 | 56.9 | 95.5 | 97.1 | 99.4 | 99.8 | 100.1
-3.4 | 3.62 | 7 | 66.5 | 95.85 | 97.34 | 99.42 | 99.9 | 100.1
-3.3 | 3.9 | 7.6 | 74.5 | 96.25 | 97.575 | 99.5 | 99.9 | 100.1
-3.2 | 4.1 | 8.5 | 81.2333 | 96.4 | 97.8 | 99.5 | 99.9 | 100.1
-3.1 | 4.3 | 9.6 | 86.3 | 96.8 | 97.9 | 99.5143 | 99.9 | 100.1
-3 | 4.7 | 11.18 | 89.4333 | 96.9667 | 98.025 | 99.6 | 99.9 | 100.1
-2.9 | 5 | 13.4 | 91.8 | 97.18 | 98.175 | 99.6 | 99.8 | 100.1
-2.8 | 5.55 | 16.6 | 93.25 | 97.3667 | 98.3 | 99.7 | 99.9 | 100.1
-2.7 | 6.1 | 21.6 | 94.4 | 97.675 | 98.525 | 99.7 | 99.9 | 100.1
-2.6 | 6.9 | 27.525 | 95.15 | 97.8 | 98.6 | 99.7857 | 99.9 | 100.2
-2.5 | 7.9667 | 35.8 | 96.1 | 98.125 | 98.65 | 99.8 | 100 | 100.15
-2.4 | 9.4 | 45.5 | 96.575 | 98.3 | 98.8 | 99.8 | 100 | 100.1
-2.3 | 11.8 | 56.3333 | 96.9 | 98.46 | 98.9 | 99.8 | 100 | 100.1
-2.2 | 15 | 65.8 | 97.2667 | 98.5 | 99 | 99.8 | 100 | 100.1
-2.1 | 19.8333 | 75.05 | 97.7 | 98.65 | 99.1 | 99.8571 | 100 | 100.1
-2 | 26.9 | 82.725 | 98.05 | 98.7 | 99.2 | 99.9 | 100 | 100.1
-1.9 | 35 | 87.8 | 98.2 | 98.94 | 99.3 | 99.9 | 100 | 100.1
-1.8 | 45.35 | 91.2 | 98.3 | 99.1 | 99.38 | 99.9 | 100 | 100.1
-1.7 | 55.1 | 93.6 | 98.6 | 99.225 | 99.425 | 99.9 | 100 | 100.1
-1.6 | 66.2 | 95.15 | 98.8 | 99.375 | 99.5 | 99.9 | 100 | 100.0714
-1.5 | 75.45 | 96.4 | 98.9 | 99.4 | 99.6 | 99.9286 | 100 | 100
-1.4 | 83.35 | 97.3333 | 99 | 99.5 | 99.6 | 99.98 | 100 | 100.1
-1.3 | 88.7 | 97.9 | 99.1333 | 99.6 | 99.7 | 100 | 100 | 100.1
-1.2 | 92.375 | 98.2 | 99.3 | 99.6 | 99.7 | 100 | 100 | 100.1
-1.1 | 94.8 | 98.6333 | 99.4 | 99.7 | 99.7 | 99.9 | 100 | 100.1
-1 | 96.4 | 98.925 | 99.5 | 99.7 | 99.7 | 100 | 100 | 100.1
-0.9 | 97.4 | 99.3 | 99.6 | 99.8 | 99.8 | 100 | 100 | 100.1
-0.8 | 98.2 | 99.4 | 99.7 | 99.8333 | 99.825 | 100 | 100 | 100.1
-0.7 | 98.7 | 99.5333 | 99.8 | 99.9 | 99.9 | 100 | 100 | 100
-0.6 | 99.2 | 99.7 | 99.8 | 99.9 | 99.975 | 100 | 100.1 | 100
-0.5 | 99.4 | 99.8 | 99.8 | 100 | 99.9 | 100 | 100.1 | 100
-0.4 | 99.6 | 99.8 | 99.9 | 100 | 99.9 | 100 | 100.1 | 100
-0.3 | 99.9 | 99.9 | 99.9 | 100 | 99.95 | 100.1 | 100 | 100
-0.2 | 100 | 100 | 99.9 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.3 | 99.8 | 99.9 | 100 | 100 | 100 | 100 | 100.075 | 100
0.4 | 99.8 | 99.8 | 100 | 100 | 99.9 | 100.1 | 100.1 | 100
0.5 | 99.6333 | 99.6667 | 99.9 | 100 | 100 | 100.1 | 100.1 | 100
0.6 | 98.9 | 99.6 | 99.9 | 100 | 99.9 | 100.1 | 100.1 | 100
0.7 | 98.6 | 99.5 | 99.8 | 99.9 | 99.9 | 100.1 | 100.1 | 100.0125
0.8 | 98.25 | 99.3 | 99.7 | 99.8333 | 99.9 | 100.1 | 100.0571 | 100.1
0.9 | 97.4 | 99.1 | 99.6667 | 99.8 | 99.88 | 100 | 100 | 100.1
1 | 96.2667 | 98.9 | 99.6 | 99.8 | 99.8 | 100 | 100 | 100.1
1.1 | 94.55 | 98.7 | 99.5 | 99.7 | 99.8 | 99.975 | 100 | 100.1
1.2 | 91.9667 | 98.3333 | 99.4 | 99.7 | 99.8 | 99.9 | 100 | 100.1
1.3 | 87.9 | 98 | 99.4 | 99.7 | 99.7 | 99.9 | 100 | 100.1
1.4 | 82.8 | 97.35 | 99.2667 | 99.6 | 99.7 | 99.9 | 100 | 100.0625
1.5 | 75.2 | 96.4 | 99 | 99.5 | 99.6 | 100 | 100 | 100
1.6 | 65.8333 | 95.275 | 98.8 | 99.3 | 99.54 | 100 | 100 | 100.1
1.7 | 55.4 | 93.675 | 98.7 | 99.3 | 99.5 | 100 | 100 | 100.1
1.8 | 45.2 | 91.4 | 98.55 | 99.2 | 99.4 | 100 | 100.1 | 100.1
1.9 | 34.525 | 87.58 | 98.4 | 99 | 99.36 | 100 | 100.1 | 100.1
2 | 26.1333 | 82.2 | 98.0333 | 98.9 | 99.3 | 99.9 | 100 | 100.1
2.1 | 19.2 | 74.7667 | 97.75 | 98.8 | 99.225 | 99.9 | 100.0111 | 100.1
2.2 | 14.7 | 66.5 | 97.45 | 98.7 | 99.1 | 99.9 | 100.05 | 100.1
2.3 | 11.5333 | 56.5 | 97.2 | 98.55 | 98.9667 | 99.9 | 100.0143 | 100.1
2.4 | 9.2 | 46 | 96.6 | 98.3 | 98.9 | 99.85 | 100.02 | 100.1
2.5 | 7.825 | 36.4333 | 96.1333 | 98.1 | 98.8 | 99.8 | 100.05 | 100.1
2.6 | 6.8 | 27.5 | 95.4 | 98.0667 | 98.6 | 99.8 | 100.1 | 100.2
2.7 | 6.1 | 21.0667 | 94.6 | 97.8 | 98.5 | 99.8 | 100.0714 | 100.15
2.8 | 5.45 | 16.2 | 93.3 | 97.6 | 98.4 | 99.7 | 100 | 100.1
2.9 | 5 | 13.05 | 91.65 | 97.3 | 98.4 | 99.7 | 100 | 100.1
3 | 4.55 | 11.05 | 89.4 | 97.1 | 98.2 | 99.6667 | 100 | 100.1
3.1 | 4.3 | 9.5 | 86 | 96.8 | 98.1 | 99.6 | 100 | 100.1
3.2 | 4.025 | 8.3 | 81.25 | 96.6 | 97.96 | 99.6 | 100 | 100.1
3.3 | 3.8 | 7.6 | 74.6 | 96.1667 | 97.7429 | 99.6 | 100 | 100.1
3.4 | 3.65 | 6.9667 | 66.12 | 95.8 | 97.5 | 99.6 | 100 | 100.1
3.5 | 3.5 | 6.5 | 56.9 | 95.5 | 97.4 | 99.5 | 99.94 | 100.1
3.6 | 3.35 | 6 | 46.7 | 95 | 97.2 | 99.5 | 99.9 | 100.1
3.7 | 3.2 | 5.6 | 37.8 | 94.4 | 96.8 | 99.4 | 99.9 | 100.1
3.8 | 3.05 | 5.3667 | 29.6 | 93.8 | 96.6571 | 99.4 | 99.8 | 100.1
3.9 | 2.9 | 5.14 | 23.3 | 92.8 | 96.5 | 99.3 | 99.8 | 100.1
4 | 2.85 | 4.9 | 18.6 | 91.75 | 96.1 | 99.3 | 99.8 | 100.1
4.1 | 2.7 | 4.64 | 15.5 | 90.05 | 96 | 99.275 | 99.8 | 100.0875
4.2 | 2.7 | 4.5 | 13.4 | 87.925 | 95.68 | 99.2 | 99.8 | 100
4.3 | 2.6 | 4.2667 | 12 | 84.9333 | 95.425 | 99.2 | 99.8 | 100
4.4 | 2.5 | 4.2 | 10.9 | 79.9 | 95.2 | 99.1 | 99.7 | 100
4.5 | 2.4 | 4 | 10 | 73.7 | 94.64 | 99.1 | 99.7 | 100
4.6 | 2.35 | 3.85 | 9.3 | 65.6 | 94.2 | 99 | 99.7 | 100
4.7 | 2.3 | 3.7 | 8.9 | 56.9667 | 94.0667 | 98.925 | 99.7 | 100
4.8 | 2.2 | 3.6 | 8.3 | 47.5 | 93.625 | 98.9 | 99.7 | 100
4.9 | 2.2 | 3.5 | 8 | 38.9333 | 93.0286 | 98.8 | 99.7 | 100
5 | 2.1 | 3.375 | 7.64 | 30.8 | 92.3 | 98.725 | 99.7 | 100
5.1 | 2 | 3.3 | 7.3 | 25.1 | 91.4167 | 98.7 | 99.7 | 100
5.2 | 2 | 3.175 | 7 | 20.35 | 90.5 | 98.6 | 99.6 | 100
5.3 | 1.9 | 3.1 | 6.8 | 17.35 | 89.1 | 98.5 | 99.6 | 100
5.4 | 1.875 | 3 | 6.55 | 15.4 | 87.025 | 98.5 | 99.58 | 100
5.5 | 1.8 | 2.9 | 6.3 | 13.9 | 84.275 | 98.4 | 99.5 | 100
5.6 | 1.8 | 2.9 | 6.1 | 12.8 | 79.975 | 98.3 | 99.48 | 99.9583
5.7 | 1.7 | 2.8 | 5.9 | 12 | 74.275 | 98.2286 | 99.4 | 99.9
5.8 | 1.7 | 2.7 | 5.75 | 11.3 | 66.7667 | 98.125 | 99.4 | 99.9
5.9 | 1.7 | 2.6 | 5.5 | 10.7 | 58.1 | 98 | 99.3 | 99.9
6 | 1.6 | 2.5 | 5.4 | 10.1 | 49.75 | 98 | 99.3 | 99.87
6.1 | 1.6 | 2.475 | 5.2 | 9.7667 | 41 | 97.82 | 99.24 | 99.8
6.2 | 1.5 | 2.4 | 5.0667 | 9.4 | 33 | 97.74 | 99.2 | 99.8
6.3 | 1.5 | 2.3 | 4.9 | 9.04 | 26.9 | 97.62 | 99.2 | 99.8
6.4 | 1.4667 | 2.2 | 4.8 | 8.7 | 22.1 | 97.54 | 99.125 | 99.7
6.5 | 1.4 | 2.2 | 4.7 | 8.475 | 19.175 | 97.5 | 99.1 | 99.7
6.6 | 1.4 | 2.1 | 4.5 | 8.1333 | 17.0667 | 97.3429 | 99 | 99.7
6.7 | 1.4 | 2.1 | 4.4 | 7.9 | 15.52 | 97.2 | 99 | 99.675
6.8 |  | 2 | 4.25 | 7.7333 | 14.4 | 97.1 | 99 | 99.6
6.9 |  | 2 | 4.1 | 7.4 | 13.4 | 97.075 | 98.95 | 99.6
7 |  | 1.9 | 4 | 7.25 | 12.76 | 96.825 | 98.9 | 99.6
7.1 |  | 1.9 | 3.9 | 7 | 12.15 | 96.7 | 98.8667 | 99.6
7.2 |  | 1.8 | 3.8 | 6.8 | 11.66 | 96.6857 | 98.8 | 99.5143
7.3 |  | 1.8 | 3.7 | 6.7 | 11.2 | 96.6 | 98.7 | 99.5
7.4 |  |  | 3.6 | 6.4 | 10.75 | 96.5 | 98.7 | 99.5
7.5 |  |  | 3.5 | 6.2667 | 10.42 | 96.3429 | 98.6333 | 99.5
7.6 |  |  | 3.4 | 6.1 | 10.12 | 96.18 | 98.6 | 99.4857
7.7 |  |  | 3.3 | 5.9667 | 9.8429 | 95.9667 | 98.5 | 99.4
7.8 |  |  | 3.24 | 5.78 | 9.48 | 95.86 | 98.5 | 99.4
7.9 |  |  | 3.2 | 5.6 | 9.2 | 95.7571 | 98.4 | 99.4
8 |  |  | 3.1 | 5.5 | 8.9667 | 95.575 | 98.4 | 99.35
8.1 |  |  | 3 | 5.3 | 8.7 | 95.45 | 98.4 | 99.3
8.2 |  |  | 2.975 | 5.2 | 8.56 | 95.3 | 98.4 | 99.3
8.3 |  |  | 2.9 | 5.1 | 8.28 | 95.1333 | 98.28 | 99.3
8.4 |  |  | 2.875 | 5 | 8.025 | 94.98 | 98.15 | 99.2286
8.5 |  |  | 2.725 | 4.9 | 7.825 | 94.8 | 98.05 | 99.2
8.6 |  |  |  | 4.8 | 7.6 | 94.7 | 98 | 99.2
8.7 |  |  |  | 4.6 | 7.4 | 94.38 | 98 | 99.2
8.8 |  |  |  | 4.5 | 7.22 | 94.2 | 97.9 | 99.1
8.9 |  |  |  | 4.4 | 7 | 94 | 97.875 | 99.1
9 |  |  |  | 4.3 | 6.9 | 93.825 | 97.8 | 99.075
9.1 |  |  |  | 4.2 | 6.775 | 93.54 | 97.8 | 99
9.2 |  |  |  | 4.1 | 6.525 | 93.3857 | 97.7 | 99
9.3 |  |  |  | 4 | 6.4 | 93.2 | 97.6 | 99
9.4 |  |  |  | 3.9 | 6.3 | 93 | 97.56 | 98.9125
9.5 |  |  |  | 3.8 | 6.1 | 92.75 | 97.5 | 98.9
9.6 |  |  |  | 3.7 | 6 | 92.5 | 97.475 | 98.9
9.7 |  |  |  | 3.7 | 5.85 | 92.18 | 97.325 | 98.9
9.8 |  |  |  |  | 5.6875 | 92.06 | 97.2 | 98.85
9.9 |  |  |  |  | 5.54 | 91.7 | 97.15 | 98.8
10 |  |  |  |  | 5.44 | 91.475 | 97.1 | 98.8
10.1 |  |  |  |  | 5.3 | 91.2 | 97 | 98.7692
10.2 |  |  |  |  | 5.2 | 90.85 | 96.95 | 98.7
10.3 |  |  |  |  | 5.1375 | 90.6 | 96.9 | 98.7
10.4 |  |  |  |  | 5 | 90.25 | 96.8 | 98.6125
10.5 |  |  |  |  | 4.85 | 89.85 | 96.7 | 98.6
10.6 |  |  |  |  | 4.76 | 89.55 | 96.625 | 98.55
10.7 |  |  |  |  | 4.7 | 89.125 | 96.5 | 98.5
10.8 |  |  |  |  | 4.6 | 88.625 | 96.4 | 98.5
10.9 |  |  |  |  | 4.5 | 88.05 | 96.3 | 98.4
11 | [5 empty] | 87.45 | 96.2 | 98.4
11.1 | [5 empty] | 86.625 | 96.15 | 98.325
11.2 | [5 empty] | 85.675 | 96.1 | 98.3
11.3 | [5 empty] | 84.4 | 96.075 | 98.3
11.4 | [5 empty] | 82.9 | 95.925 | 98.3
11.5 | [5 empty] | 80.3333 | 95.7667 | 98.24
11.6 | [5 empty] | 77.1 | 95.7 | 98.1714
11.7 | [5 empty] | 72.05 | 95.55 | 98.1
11.8 | [5 empty] | 65.6 | 95.4 | 98.1
11.9 | [5 empty] | 58.3 | 95.3 | 98.04
12 | [5 empty] | 50.6 | 95.2 | 97.975
12.1 | [5 empty] | 42.36 | 95.1 | 97.9125
12.2 | [5 empty] | 35.4 | 95 | 97.9
12.3 | [5 empty] | 29.9 | 94.9 | 97.9
12.4 | [5 empty] | 25.7 | 94.7 | 97.8
12.5 | [5 empty] | 22.9667 | 94.6 | 97.75
12.6 | [5 empty] | 20.92 | 94.5 | 97.7
12.7 | [5 empty] | 19.54 | 94.3 | 97.675
12.8 | [5 empty] | 18.5 | 94.2333 | 97.5857
12.9 | [5 empty] | 17.62 | 94.1 | 97.5
13 | [5 empty] | 16.95 | 94 | 97.5
13.1 | [5 empty] | 16.425 | 93.9 | 97.4
13.2 | [5 empty] | 15.95 | 93.78 | 97.35
13.3 | [5 empty] | 15.375 | 93.5714 | 97.3
13.4 | [5 empty] | 15.025 | 93.48 | 97.2875
13.5 | [5 empty] | 14.675 | 93.2875 | 97.2
13.6 | [5 empty] | 14.325 | 93.14 | 97.2
13.7 | [5 empty] | 13.975 | 93 | 97.1
13.8 | [5 empty] | 13.7333 | 92.7714 | 97.0625
13.9 | [5 empty] | 13.5 | 92.58 | 96.9286
14 | [5 empty] | 13.1333 | 92.425 | 96.9
14.1 | [5 empty] | 12.8333 | 92.24 | 96.8625
14.2 | [5 empty] | 12.6 | 92.15 | 96.8
14.3 | [5 empty] | 12.3333 | 91.95 | 96.8
14.4 | [5 empty] | 12.1 | 91.8 | 96.7333
14.5 | [5 empty] | 11.7714 | 91.575 | 96.7
14.6 | [5 empty] | 11.6 | 91.44 | 96.6
14.7 | [5 empty] | 11.325 | 91.2429 | 96.5
14.8 | [5 empty] | 11.1 | 91.05 | 96.475
14.9 | [5 empty] | 10.86 | 90.8 | 96.3429
15 | [5 empty] | 10.625 | 90.7 | 96.3
15.1 | [5 empty] | 10.475 | 90.375 | 96.175
15.2 | [5 empty] | 10.3 | 90.2 | 96.1
15.3 | [5 empty] | 10.075 | 90 | 96.0143
15.4 | [5 empty] | 9.9 | 89.8 | 96
15.5 | [5 empty] | 9.64 | 89.55 | 95.9571
15.6 | [5 empty] | 9.5 | 89.3 | 95.825
15.7 | [5 empty] | 9.3 | 89.0333 | 95.7
15.8 | [5 empty] | 9.1 | 88.8 | 95.6625
15.9 | [5 empty] | 9 | 88.45 | 95.6
16 | [5 empty] | 8.8 | 88.2 | 95.52
16.1 | [5 empty] | 8.55 | 87.9667 | 95.4273
16.2 | [5 empty] | 8.4 | 87.6667 | 95.3067
16.3 | [5 empty] | 8.3 | 87.375 | 95.1818
16.4 | [5 empty] | 8.2 | 87.05 | 95.0923
16.5 | [5 empty] | 8.08 | 86.54 | 95.0154
16.6 | [5 empty] | 7.86 | 86.32 | 94.9273
16.7 | [5 empty] | 7.7 | 85.775 | 94.8364
16.8 | [5 empty] | 7.6 | 85.4 | 94.7538
16.9 | [5 empty] | 7.475 | 84.8 | 94.6667
17 | [6 empty] | 84.3 | 94.55
17.1 | [6 empty] | 83.6 | 94.46
17.2 | [6 empty] | 82.64 | 94.3909
17.3 | [6 empty] | 81.6 | 94.3
17.4 | [6 empty] | 80.08 | 94.1778
17.5 | [6 empty] | 77.9 | 94.1
17.6 | [6 empty] | 74.4143 | 93.95
17.7 | [6 empty] | 70.0333 | 93.7875
17.8 | [6 empty] | 64.0889 | 93.6571
17.9 | [6 empty] | 57.7429 | 93.6
18 | [6 empty] | 50.1111 | 93.4875
18.1 | [6 empty] | 43.05 | 93.3571
18.2 | [6 empty] | 36.6 | 93.225
18.3 | [6 empty] | 31.3333 | 93.0875
18.4 | [6 empty] | 27.34 | 92.9571
18.5 | [6 empty] | 24.5714 | 92.825
18.6 | [6 empty] | 22.8 | 92.6714
18.7 | [6 empty] | 21.425 | 92.6
18.8 | [6 empty] | 20.46 | 92.4
18.9 | [6 empty] | 19.7 | 92.3125
19 | [6 empty] | 19.1 | 92.15
19.1 | [6 empty] | 18.525 | 91.925
19.2 | [6 empty] | 18 | 91.775
19.3 | [6 empty] | 17.6 | 91.7
19.4 | [6 empty] | 17.25 | 91.525
19.5 | [6 empty] | 16.9 | 91.35
19.6 | [6 empty] | 16.55 | 91.1429
19.7 | [6 empty] | 16.2 | 91.0222
19.8 | [6 empty] | 15.9 | 90.8714
19.9 | [6 empty] | 15.66 | 90.6
20 | [6 empty] | 15.35 | 90.4143
20.1 | [6 empty] | 15.075 | 90.2857
20.2 | [6 empty] | 14.82 | 90.15
20.3 | [6 empty] | 14.5714 | 89.9286
20.4 | [6 empty] | 14.32 | 89.6875
20.5 | [6 empty] | 14.0333 | 89.5571
20.6 | [6 empty] | 13.76 | 89.35
20.7 | [6 empty] | 13.575 | 89.1778
20.8 | [6 empty] | 13.325 | 88.9429
20.9 | [6 empty] | 13.16 | 88.675
21 | [6 empty] | 12.9 | 88.4
21.1 | [6 empty] | 12.7 | 88.3
21.2 | [6 empty] | 12.475 | 88.0273
21.3 | [6 empty] | 12.225 | 87.8
21.4 | [6 empty] | 12.1 | 87.5429
21.5 | [6 empty] | 11.84 | 87.275
21.6 | [6 empty] | 11.68 | 87
21.7 | [6 empty] | 11.48 | 86.6714
21.8 | [7 empty] | 86.4125
21.9 | [7 empty] | 86.1
22 | [7 empty] | 85.75
22.1 | [7 empty] | 85.5
22.2 | [7 empty] | 85.2286
22.3 | [7 empty] | 84.85
22.4 | [7 empty] | 84.55
22.5 | [7 empty] | 84.1
22.6 | [7 empty] | 83.7
22.7 | [7 empty] | 83.3
22.8 | [7 empty] | 82.775
22.9 | [7 empty] | 82.2
23 | [7 empty] | 81.65
23.1 | [7 empty] | 80.9889
23.2 | [7 empty] | 80.0615
23.3 | [7 empty] | 78.88
23.4 | [7 empty] | 77.28
23.5 | [7 empty] | 75.05
23.6 | [7 empty] | 72.1
23.7 | [7 empty] | 67.4077
23.8 | [7 empty] | 61.94
23.9 | [7 empty] | 55.7571
24 | [7 empty] | 48.9
24.1 | [7 empty] | 42.4625
24.2 | [7 empty] | 36.375
24.3 | [7 empty] | 31.4714
24.4 | [7 empty] | 27.9125
24.5 | [7 empty] | 25.3429
24.6 | [7 empty] | 23.65
24.7 | [7 empty] | 22.425
24.8 | [7 empty] | 21.5571
24.9 | [7 empty] | 20.8286
25 | [7 empty] | 20.1714
25.1 | [7 empty] | 19.65
25.2 | [7 empty] | 19.2333
25.3 | [7 empty] | 18.875
25.4 | [7 empty] | 18.425
25.5 | [7 empty] | 18.1
25.6 | [7 empty] | 17.7429
25.7 | [7 empty] | 17.4714
25.8 | [7 empty] | 17.15
25.9 | [7 empty] | 16.8364
26 | [7 empty] | 16.5333
26.1 | [7 empty] | 16.25
26.2 | [7 empty] | 16.07
26.3 | [7 empty] | 15.775
26.4 | [7 empty] | 15.5143
26.5 | [7 empty] | 15.25
26.6 | [7 empty] | 15.0625
26.7 | [7 empty] | 14.7571
26.8 | [7 empty] | 14.55
26.9 | [7 empty] | 14.3
27 | [7 empty] | 14.075
27.1 | [7 empty] | 13.8625
27.2 | [7 empty] | 13.675
27.3 | [7 empty] | 13.4
27.4 | [7 empty] | 13.2429
27.5 | [7 empty] | 13.0167
27.6 | [7 empty] | 12.76
27.7 | [7 empty] | 12.6
27.8 | [7 empty] | 12.3714
27.9 | [7 empty] | 12.2375
28 | [7 empty] | 12.025
28.1 | [7 empty] | 11.875
28.2 | [7 empty] | 11.6857
28.3 | [7 empty] | 11.5125
28.4 | [7 empty] | 11.2857
28.5 | [7 empty] | 11.15
28.6 | [7 empty] | 11.025
28.7 | [7 empty] | 10.8
28.8 | [7 empty] | 10.6833
28.9 | [7 empty] | 10.5286
29 | [7 empty] | 10.375
29.1 | [7 empty] | 10.1571
29.2 | [7 empty] | 10.1
29.3 | [7 empty] | 9.9333
29.4 | [7 empty] | 9.8
29.5 | [7 empty] | 9.6778
29.6 | [7 empty] | 9.475
29.7 | [7 empty] | 9.3667

## Open Field Profiles at 30cm
Lateral Profiles
Open field
SSD=100cm
Depth=30cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Field Size: 3x3 cm2 | Field Size: 4x4 cm2 | Field Size: 6x6 cm2 | Field Size: 8x8 cm2 | Field Size: 10x10 cm2 | Field Size: 20x20 cm2 | Field Size: 30x30 cm2 | Field Size: 40x40 cm2
-29.8 | [7 empty] | 15.04
-29.7 | [7 empty] | 15.3
-29.6 | [7 empty] | 15.4667
-29.5 | [7 empty] | 15.66
-29.4 | [7 empty] | 15.94
-29.3 | [7 empty] | 16.1
-29.2 | [7 empty] | 16.3286
-29.1 | [7 empty] | 16.5429
-29 | [7 empty] | 16.7
-28.9 | [7 empty] | 16.9571
-28.8 | [7 empty] | 17.1625
-28.7 | [7 empty] | 17.4
-28.6 | [7 empty] | 17.6571
-28.5 | [7 empty] | 17.925
-28.4 | [7 empty] | 18.1
-28.3 | [7 empty] | 18.35
-28.2 | [7 empty] | 18.6143
-28.1 | [7 empty] | 18.9
-28 | [7 empty] | 19.2
-27.9 | [7 empty] | 19.3875
-27.8 | [7 empty] | 19.75
-27.7 | [7 empty] | 20.0143
-27.6 | [7 empty] | 20.22
-27.5 | [7 empty] | 20.5571
-27.4 | [6 empty] | 8.47 | 20.875
-27.3 | [6 empty] | 8.5778 | 21.2714
-27.2 | [6 empty] | 8.7 | 21.6625
-27.1 | [6 empty] | 8.7143 | 22.0571
-27 | [6 empty] | 8.8667 | 22.55
-26.9 | [6 empty] | 9.0167 | 23.1286
-26.8 | [6 empty] | 9.1 | 23.7857
-26.7 | [6 empty] | 9.2 | 24.65
-26.6 | [6 empty] | 9.3429 | 25.7714
-26.5 | [6 empty] | 9.48 | 27.2
-26.4 | [6 empty] | 9.6 | 29.35
-26.3 | [6 empty] | 9.7375 | 32.175
-26.2 | [6 empty] | 9.8714 | 36.15
-26.1 | [6 empty] | 10.0125 | 40.9714
-26 | [6 empty] | 10.1333 | 46.275
-25.9 | [6 empty] | 10.2571 | 52.0125
-25.8 | [6 empty] | 10.4 | 57.1857
-25.7 | [6 empty] | 10.4909 | 61.85
-25.6 | [6 empty] | 10.62 | 65.7875
-25.5 | [6 empty] | 10.74 | 68.475
-25.4 | [6 empty] | 10.9222 | 70.5286
-25.3 | [6 empty] | 11.0429 | 72
-25.2 | [6 empty] | 11.1857 | 73
-25.1 | [6 empty] | 11.35 | 73.875
-25 | [6 empty] | 11.5 | 74.7
-24.9 | [6 empty] | 11.5667 | 75.3
-24.8 | [6 empty] | 11.7 | 75.7
-24.7 | [6 empty] | 11.8571 | 76.1857
-24.6 | [6 empty] | 12.0625 | 76.7
-24.5 | [6 empty] | 12.1875 | 77.1143
-24.4 | [6 empty] | 12.325 | 77.4571
-24.3 | [6 empty] | 12.4571 | 77.7429
-24.2 | [6 empty] | 12.675 | 78.08
-24.1 | [6 empty] | 12.825 | 78.4
-24 | [6 empty] | 13.0143 | 78.7875
-23.9 | [6 empty] | 13.2 | 79.075
-23.8 | [6 empty] | 13.38 | 79.3571
-23.7 | [6 empty] | 13.48 | 79.7
-23.6 | [6 empty] | 13.7143 | 80.025
-23.5 | [6 empty] | 13.8667 | 80.275
-23.4 | [6 empty] | 14.0286 | 80.55
-23.3 | [6 empty] | 14.2444 | 80.8714
-23.2 | [6 empty] | 14.45 | 81.175
-23.1 | [6 empty] | 14.6625 | 81.425
-23 | [6 empty] | 14.875 | 81.66
-22.9 | [6 empty] | 15.0545 | 81.9143
-22.8 | [6 empty] | 15.25 | 82.15
-22.7 | [6 empty] | 15.4571 | 82.3857
-22.6 | [6 empty] | 15.675 | 82.7286
-22.5 | [6 empty] | 15.9 | 82.95
-22.4 | [6 empty] | 16.15 | 83.075
-22.3 | [6 empty] | 16.3889 | 83.325
-22.2 | [6 empty] | 16.6364 | 83.5429
-22.1 | [6 empty] | 16.8889 | 83.9
-22 | [6 empty] | 17.1333 | 84
-21.9 | [6 empty] | 17.4 | 84.2429
-21.8 | [6 empty] | 17.6 | 84.52
-21.7 | [6 empty] | 17.8429 | 84.7
-21.6 | [6 empty] | 18.125 | 84.8
-21.5 | [6 empty] | 18.375 | 85
-21.4 | [6 empty] | 18.625 | 85.2333
-21.3 | [6 empty] | 18.9 | 85.4714
-21.2 | [6 empty] | 19.2333 | 85.5571
-21.1 | [6 empty] | 19.525 | 85.825
-21 | [6 empty] | 19.9 | 86
-20.9 | [6 empty] | 20.2571 | 86.1286
-20.8 | [6 empty] | 20.5875 | 86.2625
-20.7 | [6 empty] | 21 | 86.475
-20.6 | [6 empty] | 21.5143 | 86.77
-20.5 | [6 empty] | 22.1125 | 86.925
-20.4 | [6 empty] | 22.65 | 87.05
-20.3 | [6 empty] | 23.4 | 87.175
-20.2 | [6 empty] | 24.4 | 87.3778
-20.1 | [6 empty] | 25.5714 | 87.58
-20 | [6 empty] | 27.3286 | 87.69
-19.9 | [6 empty] | 30.0143 | 87.85
-19.8 | [6 empty] | 33.45 | 88.05
-19.7 | [6 empty] | 38.0571 | 88.22
-19.6 | [6 empty] | 43.7 | 88.3429
-19.5 | [6 empty] | 49.6857 | 88.44
-19.4 | [6 empty] | 55.8143 | 88.64
-19.3 | [6 empty] | 61.55 | 88.8077
-19.2 | [6 empty] | 66.4286 | 88.9333
-19.1 | [6 empty] | 70.1 | 89.0267
-19 | [6 empty] | 72.85 | 89.16
-18.9 | [6 empty] | 74.8286 | 89.3273
-18.8 | [6 empty] | 76.2125 | 89.4923
-18.7 | [6 empty] | 77.3 | 89.6375
-18.6 | [6 empty] | 78.15 | 89.7714
-18.5 | [6 empty] | 78.8625 | 89.9
-18.4 | [6 empty] | 79.4714 | 90.05
-18.3 | [6 empty] | 79.925 | 90.2
-18.2 | [6 empty] | 80.35 | 90.4333
-18.1 | [6 empty] | 80.85 | 90.5
-18 | [6 empty] | 81.3143 | 90.5875
-17.9 | [5 empty] | 9.92 | 81.625 | 90.725
-17.8 | [5 empty] | 10.1 | 81.95 | 90.8571
-17.7 | [5 empty] | 10.2125 | 82.45 | 90.9875
-17.6 | [5 empty] | 10.36 | 82.675 | 91.025
-17.5 | [5 empty] | 10.5 | 83.0143 | 91.1571
-17.4 | [5 empty] | 10.6286 | 83.2 | 91.2778
-17.3 | [5 empty] | 10.8 | 83.5125 | 91.4125
-17.2 | [5 empty] | 11.075 | 83.825 | 91.5
-17.1 | [5 empty] | 11.25 | 84.1 | 91.65
-17 | [5 empty] | 11.4 | 84.35 | 91.8
-16.9 | [5 empty] | 11.5333 | 84.6714 | 91.8375
-16.8 | [5 empty] | 11.68 | 84.8875 | 91.9
-16.7 | [5 empty] | 11.9 | 85.05 | 92.1
-16.6 | [5 empty] | 12.1429 | 85.3143 | 92.2375
-16.5 | [5 empty] | 12.34 | 85.7 | 92.3714
-16.4 | [5 empty] | 12.4875 | 85.925 | 92.5
-16.3 | [5 empty] | 12.68 | 86.1143 | 92.55
-16.2 | [5 empty] | 12.925 | 86.375 | 92.78
-16.1 | [5 empty] | 13.1 | 86.55 | 92.8
-16 | [5 empty] | 13.35 | 86.8143 | 92.9
-15.9 | [5 empty] | 13.58 | 86.9875 | 92.975
-15.8 | [5 empty] | 13.8333 | 87.25 | 93.1
-15.7 | [5 empty] | 14.02 | 87.45 | 93.1429
-15.6 | [5 empty] | 14.35 | 87.5857 | 93.275
-15.5 | [5 empty] | 14.6 | 87.85 | 93.4
-15.4 | [5 empty] | 14.78 | 88.05 | 93.4429
-15.3 | [5 empty] | 15.05 | 88.2714 | 93.6
-15.2 | [5 empty] | 15.36 | 88.4286 | 93.7364
-15.1 | [5 empty] | 15.7 | 88.6 | 93.8
-15 | [5 empty] | 15.925 | 88.7714 | 93.8667
-14.9 | [5 empty] | 16.175 | 88.9286 | 93.9
-14.8 | [5 empty] | 16.525 | 89.15 | 94.02
-14.7 | [5 empty] | 16.8 | 89.2857 | 94.1333
-14.6 | [5 empty] | 17.1714 | 89.4286 | 94.2167
-14.5 | [5 empty] | 17.475 | 89.65 | 94.3
-14.4 | [5 empty] | 17.9 | 89.7857 | 94.3364
-14.3 | [5 empty] | 18.2667 | 90.0125 | 94.5125
-14.2 | [5 empty] | 18.725 | 90.1857 | 94.6
-14.1 | [5 empty] | 19.1 | 90.375 | 94.6
-14 | [5 empty] | 19.8 | 90.5125 | 94.7
-13.9 | [5 empty] | 20.5 | 90.66 | 94.7375
-13.8 | [5 empty] | 21.2 | 90.8125 | 94.8
-13.7 | [5 empty] | 22.35 | 90.9429 | 94.9286
-13.6 | [5 empty] | 23.8 | 91.15 | 95.0556
-13.5 | [5 empty] | 25.7 | 91.4125 | 95.1857
-13.4 | [5 empty] | 28.4 | 91.5429 | 95.2
-13.3 | [5 empty] | 32.04 | 91.675 | 95.25
-13.2 | [5 empty] | 37.4 | 91.8125 | 95.3857
-13.1 | [5 empty] | 43.18 | 91.9375 | 95.4125
-13 | [5 empty] | 49.6 | 92.0714 | 95.5429
-12.9 | [5 empty] | 56.725 | 92.2 | 95.6
-12.8 | [5 empty] | 63.22 | 92.325 | 95.6125
-12.7 | [5 empty] | 68.825 | 92.4 | 95.7
-12.6 | [5 empty] | 73.55 | 92.6571 | 95.7714
-12.5 | [5 empty] | 76.75 | 92.7875 | 95.9
-12.4 | [5 empty] | 79 | 92.9286 | 96
-12.3 | [5 empty] | 80.675 | 93.1 | 96
-12.2 | [5 empty] | 81.9667 | 93.2 | 96.1
-12.1 | [5 empty] | 82.9 | 93.3 | 96.1286
-12 | [5 empty] | 83.525 | 93.4 | 96.2625
-11.9 | [5 empty] | 84.275 | 93.5 | 96.3
-11.8 | [5 empty] | 84.9 | 93.6714 | 96.4
-11.7 | [5 empty] | 85.25 | 93.7 | 96.4625
-11.6 | [5 empty] | 85.6667 | 93.86 | 96.5
-11.5 | [5 empty] | 86.1 | 93.96 | 96.5222
-11.4 |  |  |  |  | 6.2 | 86.6 | 94.15 | 96.6333
-11.3 |  |  |  |  | 6.3 | 86.9 | 94.2286 | 96.7
-11.2 |  |  |  |  | 6.4 | 87.3 | 94.3833 | 96.7
-11.1 |  |  |  |  | 6.6143 | 87.6 | 94.5 | 96.79
-11 |  |  |  |  | 6.7 | 88 | 94.58 | 96.89
-10.9 |  |  |  |  | 6.8 | 88.34 | 94.6 | 96.9111
-10.8 |  |  |  |  | 6.94 | 88.5 | 94.7 | 97
-10.7 |  |  |  |  | 7.1 | 88.78 | 94.8125 | 97
-10.6 |  |  |  |  | 7.2 | 88.95 | 94.9429 | 97.0933
-10.5 |  |  |  |  | 7.34 | 89.26 | 95.075 | 97.2125
-10.4 |  |  |  |  | 7.575 | 89.58 | 95.2125 | 97.275
-10.3 |  |  |  |  | 7.6 | 89.92 | 95.3 | 97.34
-10.2 |  |  |  |  | 7.825 | 90.15 | 95.375 | 97.4077
-10.1 |  |  |  | 5.2 | 7.975 | 90.375 | 95.5125 | 97.4846
-10 |  |  |  | 5.3 | 8.125 | 90.6 | 95.6 | 97.5
-9.9 |  |  |  | 5.4 | 8.35 | 90.85 | 95.7 | 97.5714
-9.8 |  |  |  | 5.4667 | 8.5 | 91.1 | 95.7429 | 97.64
-9.7 |  |  |  | 5.6 | 8.7 | 91.375 | 95.875 | 97.7
-9.6 |  |  |  | 5.7 | 8.9 | 91.525 | 96 | 97.7
-9.5 |  |  |  | 5.8 | 9 | 91.725 | 96.0429 | 97.7545
-9.4 |  |  |  | 5.9 | 9.3 | 91.95 | 96.1667 | 97.8
-9.3 |  |  |  | 6.1 | 9.475 | 92.2 | 96.2 | 97.8667
-9.2 |  |  |  | 6.2 | 9.6667 | 92.4 | 96.3444 | 97.9
-9.1 |  |  |  | 6.3 | 9.86 | 92.64 | 96.52 | 97.9
-9 |  |  |  | 6.4 | 10.04 | 92.8 | 96.5867 | 97.9
-8.9 |  |  |  | 6.6 | 10.26 | 92.9 | 96.6533 | 97.9583
-8.8 |  |  | 3.9 | 6.7 | 10.5 | 93.125 | 96.72 | 98.0333
-8.7 |  |  | 4 | 6.8333 | 10.7667 | 93.275 | 96.7867 | 98.1
-8.6 |  |  | 4.1 | 7 | 11.02 | 93.5 | 96.8533 | 98.1769
-8.5 |  |  | 4.2 | 7.125 | 11.25 | 93.56 | 96.92 | 98.2
-8.4 |  |  | 4.3 | 7.3 | 11.5333 | 93.82 | 96.9867 | 98.2583
-8.3 |  |  | 4.3 | 7.5 | 11.8 | 93.95 | 97.0533 | 98.3
-8.2 |  |  | 4.4 | 7.65 | 12.14 | 94.0333 | 97.1375 | 98.375
-8.1 |  |  | 4.55 | 7.825 | 12.38 | 94.2 | 97.2625 | 98.4083
-8 |  |  | 4.6 | 7.9 | 12.68 | 94.4333 | 97.3583 | 98.4917
-7.9 |  |  | 4.7 | 8.225 | 13.08 | 94.5571 | 97.4 | 98.5
-7.8 |  |  | 4.8 | 8.4 | 13.525 | 94.76 | 97.4154 | 98.5889
-7.7 |  |  | 4.95 | 8.6 | 14.05 | 95 | 97.4923 | 98.6125
-7.6 |  |  | 5.075 | 8.8 | 14.56 | 95.1 | 97.6 | 98.7
-7.5 |  | 2.5 | 5.2 | 9.1 | 15.1 | 95.28 | 97.6375 | 98.7
-7.4 |  | 2.6 | 5.375 | 9.2 | 15.9 | 95.4 | 97.7 | 98.7333
-7.3 |  | 2.6 | 5.5 | 9.5 | 16.8 | 95.5 | 97.78 | 98.8
-7.2 |  | 2.7 | 5.6 | 9.7 | 18 | 95.6222 | 97.9 | 98.8
-7.1 |  | 2.7 | 5.7 | 9.9667 | 19.7 | 95.7429 | 97.9 | 98.8
-7 |  | 2.8 | 5.8 | 10.2 | 21.85 | 95.9 | 97.975 | 98.8286
-6.9 | 1.7 | 2.8 | 6 | 10.45 | 25 | 96 | 98 | 98.9625
-6.8 | 1.8 | 2.9 | 6.1 | 10.8 | 29.4 | 96.1667 | 98.1 | 99
-6.7 | 1.9 | 3 | 6.3 | 11 | 35.3 | 96.2 | 98.225 | 99.0286
-6.6 | 1.95 | 3.1 | 6.4 | 11.55 | 41.9143 | 96.38 | 98.3 | 99.1
-6.5 | 2 | 3.1 | 6.55 | 11.8 | 50.06 | 96.4571 | 98.3875 | 99.1
-6.4 | 2 | 3.2 | 6.7 | 12.3333 | 58.2 | 96.525 | 98.5 | 99.2
-6.3 | 2.0667 | 3.3 | 6.925 | 12.9 | 65.75 | 96.68 | 98.5571 | 99.2
-6.2 | 2.1 | 3.4 | 7.1 | 13.4 | 72.6 | 96.86 | 98.6 | 99.2333
-6.1 | 2.2 | 3.5 | 7.325 | 14.1333 | 78.1 | 97.075 | 98.625 | 99.3
-6 | 2.2 | 3.6 | 7.55 | 15.1 | 82.3857 | 97.2 | 98.7 | 99.3
-5.9 | 2.3 | 3.6 | 7.62 | 16.2333 | 85.22 | 97.2 | 98.7 | 99.4
-5.8 | 2.3 | 3.7 | 7.9333 | 17.9 | 87.3 | 97.3167 | 98.7 | 99.4
-5.7 | 2.3 | 3.8 | 8.14 | 19.9667 | 88.84 | 97.4 | 98.7429 | 99.4
-5.6 | 2.4 | 3.9 | 8.4 | 22.95 | 89.86 | 97.4625 | 98.8 | 99.4
-5.5 | 2.5 | 4 | 8.6 | 27.35 | 90.8 | 97.64 | 98.9 | 99.5
-5.4 | 2.5 | 4.1 | 8.8 | 33.2 | 91.4 | 97.7 | 98.9429 | 99.5
-5.3 | 2.6 | 4.225 | 9.275 | 40.5 | 91.9 | 97.7 | 99 | 99.5
-5.2 | 2.6 | 4.35 | 9.6333 | 48.1 | 92.4 | 97.8 | 99 | 99.5
-5.1 | 2.7 | 4.5 | 10.075 | 56.8 | 93 | 97.9 | 99.075 | 99.5
-5 | 2.8 | 4.5 | 10.5667 | 64.9 | 93.5 | 98 | 99.1 | 99.6
-4.9 | 2.9 | 4.8 | 11.1333 | 72.9333 | 93.7 | 98.075 | 99.1 | 99.6
-4.8 | 3 | 4.9333 | 11.86 | 78.6 | 94.05 | 98.1667 | 99.1 | 99.6
-4.7 | 3.0333 | 5 | 12.7 | 83.375 | 94.4333 | 98.2 | 99.2 | 99.7
-4.6 | 3.1 | 5.1333 | 13.825 | 86.65 | 94.7 | 98.3 | 99.2 | 99.7667
-4.5 | 3.2 | 5.3 | 15.425 | 88.75 | 95.06 | 98.4 | 99.3 | 99.8
-4.4 | 3.3 | 5.5 | 17.675 | 90 | 95.4 | 98.5333 | 99.3 | 99.8
-4.3 | 3.4 | 5.7 | 20.825 | 91.5 | 95.5833 | 98.6 | 99.3 | 99.8
-4.2 | 3.525 | 6 | 25.25 | 92.6 | 95.78 | 98.6 | 99.3571 | 99.7571
-4.1 | 3.7 | 6.2 | 31.2 | 93.3 | 96.12 | 98.7 | 99.4167 | 99.7
-4 | 3.8 | 6.4 | 38.2 | 93.9 | 96.325 | 98.76 | 99.5 | 99.8
-3.9 | 3.9 | 6.7 | 46.9 | 94.38 | 96.55 | 98.8 | 99.5 | 99.8375
-3.8 | 4 | 7 | 56.3 | 94.9 | 96.675 | 98.9 | 99.575 | 99.9
-3.7 | 4.1667 | 7.6333 | 65.7 | 95.1 | 96.8143 | 98.9 | 99.5 | 99.9
-3.6 | 4.3 | 8.1 | 73.8 | 95.55 | 97.06 | 98.9625 | 99.5 | 99.9
-3.5 | 4.5667 | 8.7333 | 80.25 | 95.7 | 97.2 | 99 | 99.6 | 99.9
-3.4 | 4.8 | 9.7 | 85 | 96.1333 | 97.4 | 99.1 | 99.6875 | 99.9
-3.3 | 5 | 10.82 | 88.4 | 96.4 | 97.5 | 99.1 | 99.7 | 99.9
-3.2 | 5.3 | 12.6 | 90.65 | 96.6333 | 97.6 | 99.175 | 99.7 | 99.9714
-3.1 | 5.6 | 14.8667 | 92.1 | 97 | 97.825 | 99.2 | 99.7 | 100
-3 | 6.1 | 18 | 93.7 | 97.2 | 98 | 99.35 | 99.7 | 100
-2.9 | 6.6667 | 22.4667 | 94.5 | 97.25 | 98.3 | 99.4 | 99.775 | 100
-2.8 | 7.4 | 29.1 | 95.3667 | 97.5 | 98.4 | 99.5 | 99.8 | 100
-2.7 | 8.45 | 36.8 | 96.05 | 97.7 | 98.4 | 99.5 | 99.8 | 100
-2.6 | 9.7667 | 46 | 96.4 | 97.9 | 98.55 | 99.5 | 99.8625 | 100.0571
-2.5 | 11.6 | 55.525 | 96.8 | 98.1 | 98.68 | 99.5 | 99.9 | 100.1
-2.4 | 14.1 | 65.6 | 97.2 | 98.2 | 98.7 | 99.55 | 99.9 | 100.0286
-2.3 | 18.1 | 73.7 | 97.4667 | 98.4 | 98.88 | 99.6 | 99.9 | 100.1
-2.2 | 23.58 | 80.6 | 97.7 | 98.5 | 98.9667 | 99.6 | 99.9 | 100.1
-2.1 | 31 | 86.3 | 97.9 | 98.7 | 99.025 | 99.6 | 99.9286 | 100.1
-2 | 40.05 | 90.3 | 98.2 | 98.8 | 99.2333 | 99.7 | 99.9375 | 100.1
-1.9 | 49.6667 | 92.55 | 98.46 | 98.85 | 99.3889 | 99.7 | 99.9 | 100.1
-1.8 | 59.7 | 94.5 | 98.6 | 99 | 99.5 | 99.8 | 99.9 | 100.1
-1.7 | 69.6 | 95.85 | 98.7 | 99 | 99.6 | 99.8 | 99.9 | 100.1
-1.6 | 77.9 | 96.7 | 98.8 | 99.2 | 99.6 | 99.8 | 99.9 | 100.1
-1.5 | 85 | 97.3 | 98.9333 | 99.375 | 99.6 | 99.8 | 99.9 | 100.1
-1.4 | 89.3 | 98.0333 | 99.1 | 99.4 | 99.6 | 99.8 | 99.9286 | 100.1
-1.3 | 92.7667 | 98.3 | 99.2 | 99.5 | 99.7 | 99.7 | 100 | 100.1
-1.2 | 94.78 | 98.6 | 99.3 | 99.6 | 99.7 | 99.75 | 100 | 100.1
-1.1 | 96.15 | 98.95 | 99.4 | 99.65 | 99.8 | 99.8 | 100 | 100.1
-1 | 97.275 | 99.3 | 99.5 | 99.7 | 99.8 | 99.8 | 100 | 100.1
-0.9 | 97.95 | 99.5 | 99.6 | 99.8 | 99.9143 | 99.8 | 100 | 100.1
-0.8 | 98.6 | 99.6 | 99.8 | 99.8 | 100 | 99.9 | 100 | 100.1
-0.7 | 98.9333 | 99.7 | 99.8 | 99.9 | 100.1 | 99.9 | 100 | 100.1
-0.6 | 99.3 | 99.8 | 99.9 | 99.9667 | 100.1 | 99.9 | 99.9143 | 100.1
-0.5 | 99.5333 | 99.8 | 99.9 | 100 | 100.1 | 99.9 | 99.9 | 100
-0.4 | 99.6 | 99.8 | 100 | 100 | 100.1 | 99.9 | 99.95 | 100
-0.3 | 99.8 | 99.9 | 100 | 100 | 100 | 99.9 | 100 | 100
-0.2 | 99.9 | 99.9333 | 100 | 100 | 100 | 99.9 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.2 | 99.9667 | 100 | 100 | 100 | 100.06 | 99.9833 | 100 | 100
0.3 | 99.8 | 100 | 100 | 100 | 100 | 99.9 | 100 | 100
0.4 | 99.6 | 99.9 | 100 | 99.975 | 100 | 99.9 | 99.95 | 100
0.5 | 99.5 | 99.9 | 100 | 99.9 | 100 | 99.9 | 99.9 | 100
0.6 | 99.35 | 99.8 | 99.9 | 99.9 | 100 | 99.9 | 99.9143 | 100.0571
0.7 | 99.1 | 99.7 | 99.8 | 99.8 | 99.9 | 99.9 | 100 | 100.1
0.8 | 98.6 | 99.58 | 99.8 | 99.8 | 99.9 | 99.9 | 100 | 100.1
0.9 | 97.7 | 99.5 | 99.7 | 99.7 | 99.9 | 99.9 | 100 | 100.1
1 | 97.1 | 99.2 | 99.7 | 99.7 | 99.9 | 99.9 | 100 | 100.1
1.1 | 96.1 | 99 | 99.6 | 99.7 | 99.8 | 99.8 | 100 | 100.1
1.2 | 94.5 | 98.7 | 99.5 | 99.625 | 99.8 | 99.8 | 100 | 100.1
1.3 | 92.5 | 98.45 | 99.35 | 99.5 | 99.7 | 99.8 | 100 | 100.1
1.4 | 89.2667 | 97.94 | 99.3 | 99.45 | 99.7 | 99.8 | 99.9286 | 100.1
1.5 | 84.2 | 97.7333 | 99.2 | 99.3 | 99.7 | 99.8 | 99.9 | 100.1
1.6 | 78.1 | 96.98 | 99.0333 | 99.25 | 99.65 | 99.8 | 99.9 | 100.1
1.7 | 69.9 | 95.8 | 98.9 | 99.2 | 99.6 | 99.8 | 99.9 | 100.1
1.8 | 60.1 | 94.7 | 98.7 | 99.0667 | 99.5 | 99.8 | 99.9 | 100.1
1.9 | 49.75 | 92.9 | 98.5 | 98.9 | 99.4 | 99.8 | 99.9 | 100.1
2 | 39.5 | 89.9667 | 98.35 | 98.8 | 99.3 | 99.75 | 99.9375 | 100.1
2.1 | 30.76 | 86.35 | 98.125 | 98.7 | 99.22 | 99.7 | 99.9286 | 100.1
2.2 | 23.9 | 81.3 | 97.9 | 98.6 | 99.1667 | 99.7 | 99.9 | 100.1
2.3 | 18.1 | 74.14 | 97.625 | 98.5 | 99 | 99.675 | 99.9 | 100.1
2.4 | 14.0667 | 65.4 | 97.35 | 98.3 | 98.9 | 99.6 | 99.9 | 100.0125
2.5 | 11.6 | 56.1 | 97 | 98.1667 | 98.875 | 99.5333 | 99.9 | 100.1
2.6 | 9.6 | 46.1333 | 96.6333 | 98 | 98.8 | 99.5 | 99.8625 | 100.0286
2.7 | 8.25 | 36.875 | 96.1333 | 97.7333 | 98.7 | 99.5 | 99.8 | 100
2.8 | 7.325 | 28.4 | 95.675 | 97.4 | 98.525 | 99.5 | 99.8 | 100
2.9 | 6.75 | 22.3333 | 94.7667 | 97.3 | 98.28 | 99.5 | 99.775 | 100
3 | 6.1 | 17.95 | 93.625 | 97.1333 | 98.2 | 99.5 | 99.7 | 100
3.1 | 5.6667 | 14.5667 | 92.4333 | 97 | 98.1 | 99.475 | 99.7 | 100
3.2 | 5.3 | 12.4667 | 90.8 | 96.74 | 97.95 | 99.325 | 99.7 | 100
3.3 | 5 | 10.75 | 88.1 | 96.5 | 97.8 | 99.3 | 99.7 | 99.9
3.4 | 4.7 | 9.7 | 84.9 | 96.2333 | 97.575 | 99.2 | 99.6875 | 99.9
3.5 | 4.5333 | 8.8 | 79.9 | 95.9 | 97.3333 | 99.1143 | 99.6 | 99.9
3.6 | 4.34 | 8.1 | 73.4 | 95.5667 | 97.2 | 99.1 | 99.5 | 99.9
3.7 | 4.2 | 7.5667 | 65.2333 | 95.25 | 97.1 | 99 | 99.5 | 99.9
3.8 | 4.0667 | 7.2 | 56.5 | 94.9667 | 96.85 | 98.9 | 99.575 | 99.9
3.9 | 3.9 | 6.8 | 47.6 | 94.4 | 96.6 | 98.9 | 99.5 | 99.8
4 | 3.8 | 6.5 | 38.8 | 93.8 | 96.4 | 98.8 | 99.5 | 99.8
4.1 | 3.7 | 6.3 | 31.2 | 93.2333 | 96.2 | 98.8 | 99.4167 | 99.7429
4.2 | 3.5 | 6 | 25.2 | 92.4333 | 95.9 | 98.7 | 99.3571 | 99.7875
4.3 | 3.4 | 5.8 | 20.5 | 91.6 | 95.675 | 98.6667 | 99.3 | 99.8
4.4 | 3.3 | 5.5 | 17.6 | 90.45 | 95.425 | 98.58 | 99.3 | 99.8
4.5 | 3.2 | 5.35 | 15.5 | 88.8 | 95.24 | 98.4 | 99.3 | 99.8
4.6 | 3.175 | 5.3 | 13.825 | 86.5333 | 94.9 | 98.3286 | 99.2 | 99.8
4.7 | 3.1 | 5.1 | 12.8 | 83.38 | 94.76 | 98.3 | 99.2 | 99.7
4.8 | 3 | 4.9 | 11.775 | 78.65 | 94.3333 | 98.2 | 99.1 | 99.66
4.9 | 3 | 4.8 | 11.1 | 72.4 | 93.8667 | 98.1 | 99.1 | 99.6
5 | 2.9 | 4.65 | 10.4 | 65.4 | 93.54 | 98.1 | 99.1 | 99.6
5.1 | 2.8 | 4.5 | 10.1 | 57.15 | 93.2286 | 98 | 99.075 | 99.56
5.2 | 2.7 | 4.4 | 9.6 | 48.375 | 92.75 | 97.9571 | 99 | 99.5
5.3 | 2.6 | 4.3 | 9.3 | 40.125 | 92.2 | 97.86 | 99 | 99.5
5.4 | 2.6 | 4.2 | 8.9 | 33.2 | 91.66 | 97.74 | 98.9429 | 99.5
5.5 | 2.5 | 4.1 | 8.7 | 26.95 | 91 | 97.62 | 98.9 | 99.48
5.6 | 2.5 | 4 | 8.35 | 23 | 90.08 | 97.6 | 98.8 | 99.4
5.7 | 2.4 | 3.9 | 8.1 | 19.85 | 88.9 | 97.5 | 98.7429 | 99.4
5.8 | 2.4 | 3.8 | 7.9 | 17.8 | 87.54 | 97.4 | 98.7 | 99.4
5.9 | 2.3 | 3.7 | 7.7 | 16.3 | 85.25 | 97.3 | 98.7 | 99.4
6 | 2.2 | 3.6 | 7.4333 | 15 | 82.475 | 97.2 | 98.7 | 99.3
6.1 | 2.15 | 3.5 | 7.3 | 14.125 | 78.3 | 97.1 | 98.625 | 99.3
6.2 | 2.1 | 3.4 | 7.05 | 13.4667 | 72.825 | 97 | 98.6 | 99.2857
6.3 | 2.1 | 3.4 | 6.9 | 12.8 | 66.075 | 96.9 | 98.5571 | 99.2
6.4 | 2 | 3.3 | 6.75 | 12.3333 | 58.1 | 96.8 | 98.5 | 99.1875
6.5 | 2 | 3.2 | 6.525 | 11.85 | 50.0333 | 96.6 | 98.3875 | 99.1
6.6 | 1.9 | 3.1 | 6.4 | 11.5333 | 41.85 | 96.5 | 98.3 | 99.1
6.7 | 1.9 | 3 | 6.15 | 11.1 | 34.98 | 96.4 | 98.225 | 99.1
6.8 | 1.8 | 3 | 6 | 10.7667 | 29 | 96.2 | 98.1 | 99
6.9 |  | 2.9 | 5.925 | 10.55 | 24.62 | 96.1 | 98 | 98.9286
7 |  | 2.9 | 5.8 | 10.15 | 21.7 | 96 | 97.975 | 98.9
7.1 |  | 2.8 | 5.6 | 9.85 | 19.5 | 95.825 | 97.9 | 98.8
7.2 |  | 2.7 | 5.5 | 9.6 | 18 | 95.775 | 97.9 | 98.8
7.3 |  | 2.6 | 5.4 | 9.4 | 16.8 | 95.6 | 97.78 | 98.8
7.4 |  | 2.6 | 5.3 | 9.2 | 15.85 | 95.4 | 97.7 | 98.7
7.5 |  | 2.5 | 5.2 | 8.925 | 15.16 | 95.2667 | 97.6375 | 98.7
7.6 |  |  | 5 | 8.7 | 14.56 | 95.1 | 97.6 | 98.7
7.7 |  |  | 4.9 | 8.6 | 14.06 | 95.025 | 97.4923 | 98.6
7.8 |  |  | 4.8 | 8.4333 | 13.6 | 94.9 | 97.4154 | 98.56
7.9 |  |  | 4.7 | 8.175 | 13.2 | 94.75 | 97.4 | 98.5
8 |  |  | 4.6 | 8 | 12.825 | 94.6 | 97.3583 | 98.4571
8.1 |  |  | 4.5 | 7.8 | 12.475 | 94.48 | 97.2625 | 98.4
8.2 |  |  | 4.4 | 7.6 | 12.2 | 94.28 | 97.1375 | 98.4
8.3 |  |  | 4.3 | 7.5 | 11.825 | 94.08 | 97.0533 | 98.3125
8.4 |  |  | 4.2 | 7.35 | 11.5667 | 93.925 | 96.9867 | 98.3
8.5 |  |  | 4.1 | 7.15 | 11.34 | 93.7 | 96.92 | 98.23
8.6 |  |  | 4.025 | 7 | 11.05 | 93.48 | 96.8533 | 98.1417
8.7 |  |  | 4 | 6.8 | 10.8 | 93.25 | 96.7867 | 98.0667
8.8 |  |  | 3.9 | 6.7 | 10.5 | 93.1 | 96.72 | 98
8.9 |  |  |  | 6.6 | 10.3 | 92.96 | 96.6533 | 98
9 |  |  |  | 6.45 | 10.1333 | 92.8 | 96.5867 | 97.925
9.1 |  |  |  | 6.3 | 9.875 | 92.625 | 96.52 | 97.9
9.2 |  |  |  | 6.1667 | 9.6714 | 92.3 | 96.3444 | 97.9
9.3 |  |  |  | 6 | 9.5286 | 92.1 | 96.2 | 97.8417
9.4 |  |  |  | 5.8667 | 9.32 | 92 | 96.1667 | 97.8
9.5 |  |  |  | 5.775 | 9.0571 | 91.8 | 96.0429 | 97.7125
9.6 |  |  |  | 5.7 | 8.78 | 91.5 | 96 | 97.7
9.7 |  |  |  | 5.6 | 8.7 | 91.32 | 95.875 | 97.6733
9.8 |  |  |  | 5.4 | 8.54 | 91.1667 | 95.7429 | 97.6067
9.9 |  |  |  | 5.275 | 8.325 | 90.92 | 95.7 | 97.6
10 |  |  |  | 5.2 | 8.1 | 90.675 | 95.6 | 97.5556
10.1 |  |  |  | 5.1 | 8.0167 | 90.4429 | 95.5125 | 97.4545
10.2 |  |  |  |  | 7.9 | 90.08 | 95.375 | 97.3733
10.3 |  |  |  |  | 7.7 | 89.8833 | 95.3 | 97.3067
10.4 |  |  |  |  | 7.55 | 89.68 | 95.2125 | 97.2438
10.5 |  |  |  |  | 7.35 | 89.4 | 95.075 | 97.18
10.6 |  |  |  |  | 7.2 | 89.06 | 94.9429 | 97.1133
10.7 |  |  |  |  | 7.1 | 88.8 | 94.8125 | 97.0467
10.8 |  |  |  |  | 6.96 | 88.5714 | 94.7 | 97
10.9 |  |  |  |  | 6.7778 | 88.2 | 94.6 | 96.9857
11 |  |  |  |  | 6.6667 | 87.88 | 94.58 | 96.9
11.1 |  |  |  |  | 6.6 | 87.6 | 94.5 | 96.86
11.2 |  |  |  |  | 6.5 | 87.35 | 94.3833 | 96.75
11.3 |  |  |  |  | 6.4 | 86.9 | 94.2286 | 96.7
11.4 | [5 empty] | 86.55 | 94.15 | 96.59
11.5 | [5 empty] | 86.1444 | 93.96 | 96.5
11.6 | [5 empty] | 85.6 | 93.86 | 96.5
11.7 | [5 empty] | 85.25 | 93.7 | 96.5
11.8 | [5 empty] | 84.7 | 93.6714 | 96.4
11.9 | [5 empty] | 84.2 | 93.5 | 96.3
12 | [5 empty] | 83.5667 | 93.4 | 96.2286
12.1 | [5 empty] | 82.8 | 93.3 | 96.1
12.2 | [5 empty] | 81.8 | 93.2 | 96.1
12.3 | [5 empty] | 80.6333 | 93.1 | 96.025
12.4 | [5 empty] | 79.025 | 92.9286 | 96
12.5 | [5 empty] | 76.4 | 92.7875 | 95.9
12.6 | [5 empty] | 72.95 | 92.6571 | 95.8375
12.7 | [5 empty] | 68.88 | 92.4 | 95.7
12.8 | [5 empty] | 63.125 | 92.325 | 95.675
12.9 | [5 empty] | 56.775 | 92.2 | 95.6
13 | [5 empty] | 50.05 | 92.0714 | 95.5125
13.1 | [5 empty] | 43.125 | 91.9375 | 95.4
13.2 | [5 empty] | 37.025 | 91.8125 | 95.35
13.3 | [5 empty] | 32.06 | 91.675 | 95.2143
13.4 | [5 empty] | 28.16 | 91.5429 | 95.2
13.5 | [5 empty] | 25.7 | 91.4125 | 95.15
13.6 | [5 empty] | 23.75 | 91.15 | 95.025
13.7 | [5 empty] | 22.3 | 90.9429 | 94.9
13.8 | [5 empty] | 21.3 | 90.8125 | 94.8571
13.9 | [5 empty] | 20.4 | 90.66 | 94.8
14 | [5 empty] | 19.76 | 90.5125 | 94.7
14.1 | [5 empty] | 19.2 | 90.375 | 94.6375
14.2 | [5 empty] | 18.7 | 90.1857 | 94.6
14.3 | [5 empty] | 18.25 | 90.0125 | 94.48
14.4 | [5 empty] | 17.84 | 89.7857 | 94.3778
14.5 | [5 empty] | 17.45 | 89.65 | 94.2727
14.6 | [5 empty] | 17.1 | 89.4286 | 94.1818
14.7 | [5 empty] | 16.82 | 89.2857 | 94.0933
14.8 | [5 empty] | 16.48 | 89.15 | 94.0267
14.9 | [5 empty] | 16.2 | 88.9286 | 93.9455
15 | [5 empty] | 15.85 | 88.7714 | 93.8375
15.1 | [5 empty] | 15.7 | 88.6 | 93.8
15.2 | [5 empty] | 15.35 | 88.4286 | 93.6929
15.3 | [5 empty] | 15.025 | 88.2714 | 93.6214
15.4 | [5 empty] | 14.8 | 88.05 | 93.5125
15.5 | [5 empty] | 14.6 | 87.85 | 93.375
15.6 | [5 empty] | 14.3 | 87.5857 | 93.2429
15.7 | [5 empty] | 14.06 | 87.45 | 93.1125
15.8 | [5 empty] | 13.86 | 87.25 | 93.0857
15.9 | [5 empty] | 13.625 | 86.9875 | 93
16 | [5 empty] | 13.375 | 86.8143 | 92.9143
16.1 | [5 empty] | 13.2 | 86.55 | 92.8
16.2 | [5 empty] | 13.0167 | 86.375 | 92.74
16.3 | [5 empty] | 12.75 | 86.1143 | 92.5875
16.4 | [5 empty] | 12.5571 | 85.925 | 92.4625
16.5 | [5 empty] | 12.3 | 85.7 | 92.3375
16.6 | [5 empty] | 12.175 | 85.3143 | 92.2
16.7 | [5 empty] | 11.9 | 85.05 | 92.0714
16.8 | [5 empty] | 11.7 | 84.8875 | 91.9375
16.9 | [5 empty] | 11.5333 | 84.6714 | 91.8
17 | [5 empty] | 11.32 | 84.35 | 91.775
17.1 | [5 empty] | 11.14 | 84.1 | 91.6429
17.2 | [5 empty] | 11 | 83.825 | 91.5125
17.3 | [5 empty] | 10.8 | 83.5125 | 91.475
17.4 | [5 empty] | 10.68 | 83.2 | 91.35
17.5 | [5 empty] | 10.5 | 83.0143 | 91.225
17.6 | [5 empty] | 10.28 | 82.675 | 91.0875
17.7 | [5 empty] | 10.2 | 82.45 | 90.9571
17.8 | [5 empty] | 10.0714 | 81.95 | 90.825
17.9 | [5 empty] | 9.82 | 81.625 | 90.6875
18 | [6 empty] | 81.3143 | 90.6
18.1 | [6 empty] | 80.85 | 90.525
18.2 | [6 empty] | 80.35 | 90.4
18.3 | [6 empty] | 79.925 | 90.2429
18.4 | [6 empty] | 79.4714 | 90.1
18.5 | [6 empty] | 78.8625 | 89.9714
18.6 | [6 empty] | 78.15 | 89.775
18.7 | [6 empty] | 77.3 | 89.6
18.8 | [6 empty] | 76.2125 | 89.5091
18.9 | [6 empty] | 74.8286 | 89.35
19 | [6 empty] | 72.85 | 89.2
19.1 | [6 empty] | 70.1 | 89.0769
19.2 | [6 empty] | 66.4286 | 89
19.3 | [6 empty] | 61.55 | 88.8333
19.4 | [6 empty] | 55.8143 | 88.64
19.5 | [6 empty] | 49.6857 | 88.46
19.6 | [6 empty] | 43.7 | 88.3267
19.7 | [6 empty] | 38.0571 | 88.22
19.8 | [6 empty] | 33.45 | 88.0222
19.9 | [6 empty] | 30.0143 | 87.9
20 | [6 empty] | 27.3286 | 87.66
20.1 | [6 empty] | 25.5714 | 87.52
20.2 | [6 empty] | 24.4 | 87.3
20.3 | [6 empty] | 23.4 | 87.14
20.4 | [6 empty] | 22.65 | 87.0143
20.5 | [6 empty] | 22.1125 | 86.8875
20.6 | [6 empty] | 21.5143 | 86.7333
20.7 | [6 empty] | 21 | 86.5111
20.8 | [6 empty] | 20.5875 | 86.2571
20.9 | [6 empty] | 20.2571 | 86.1
21 | [6 empty] | 19.9 | 85.9625
21.1 | [6 empty] | 19.525 | 85.775
21.2 | [6 empty] | 19.2333 | 85.6
21.3 | [6 empty] | 18.9 | 85.4333
21.4 | [6 empty] | 18.625 | 85.2667
21.5 | [6 empty] | 18.375 | 85.0167
21.6 | [6 empty] | 18.125 | 84.8333
21.7 | [6 empty] | 17.8429 | 84.7
21.8 | [6 empty] | 17.6 | 84.5
21.9 | [6 empty] | 17.4 | 84.2375
22 | [6 empty] | 17.1333 | 84
22.1 | [6 empty] | 16.8889 | 83.82
22.2 | [6 empty] | 16.6364 | 83.525
22.3 | [6 empty] | 16.3889 | 83.35
22.4 | [6 empty] | 16.15 | 83.1
22.5 | [6 empty] | 15.9 | 82.9143
22.6 | [6 empty] | 15.675 | 82.6714
22.7 | [6 empty] | 15.4571 | 82.4
22.8 | [6 empty] | 15.25 | 82.1286
22.9 | [6 empty] | 15.0545 | 81.8875
23 | [6 empty] | 14.875 | 81.7
23.1 | [6 empty] | 14.6625 | 81.475
23.2 | [6 empty] | 14.45 | 81.2143
23.3 | [6 empty] | 14.2444 | 80.95
23.4 | [6 empty] | 14.0286 | 80.5625
23.5 | [6 empty] | 13.8667 | 80.2143
23.6 | [6 empty] | 13.7143 | 80.025
23.7 | [6 empty] | 13.48 | 79.72
23.8 | [6 empty] | 13.38 | 79.375
23.9 | [6 empty] | 13.2 | 79.125
24 | [6 empty] | 13.0143 | 78.7857
24.1 | [6 empty] | 12.825 | 78.4429
24.2 | [6 empty] | 12.675 | 78.1
24.3 | [6 empty] | 12.4571 | 77.7
24.4 | [6 empty] | 12.325 | 77.4
24.5 | [6 empty] | 12.1875 | 77.1286
24.6 | [6 empty] | 12.0625 | 76.6714
24.7 | [6 empty] | 11.8571 | 76.2
24.8 | [6 empty] | 11.7 | 75.6875
24.9 | [6 empty] | 11.5667 | 75.2857
25 | [6 empty] | 11.5 | 74.6571
25.1 | [6 empty] | 11.35 | 73.8714
25.2 | [6 empty] | 11.1857 | 73
25.3 | [6 empty] | 11.0429 | 71.925
25.4 | [6 empty] | 10.9222 | 70.55
25.5 | [6 empty] | 10.74 | 68.4
25.6 | [6 empty] | 10.62 | 65.625
25.7 | [6 empty] | 10.4909 | 61.8286
25.8 | [6 empty] | 10.4 | 57.2625
25.9 | [6 empty] | 10.2571 | 51.85
26 | [6 empty] | 10.1333 | 46.3
26.1 | [6 empty] | 10.0125 | 41
26.2 | [6 empty] | 9.8714 | 36.025
26.3 | [6 empty] | 9.7375 | 32.2143
26.4 | [6 empty] | 9.6 | 29.2625
26.5 | [6 empty] | 9.48 | 27.2286
26.6 | [6 empty] | 9.3429 | 25.75
26.7 | [6 empty] | 9.2 | 24.6143
26.8 | [6 empty] | 9.1 | 23.8143
26.9 | [6 empty] | 9.0167 | 23.1
27 | [6 empty] | 8.8667 | 22.5571
27.1 | [6 empty] | 8.7143 | 22.0571
27.2 | [6 empty] | 8.7 | 21.65
27.3 | [6 empty] | 8.5778 | 21.275
27.4 | [6 empty] | 8.47 | 20.875
27.5 | [7 empty] | 20.5714
27.6 | [7 empty] | 20.3
27.7 | [7 empty] | 20.0571
27.8 | [7 empty] | 19.675
27.9 | [7 empty] | 19.425
28 | [7 empty] | 19.1333
28.1 | [7 empty] | 18.925
28.2 | [7 empty] | 18.6571
28.3 | [7 empty] | 18.4
28.4 | [7 empty] | 18.125
28.5 | [7 empty] | 17.8571
28.6 | [7 empty] | 17.7
28.7 | [7 empty] | 17.425
28.8 | [7 empty] | 17.1571
28.9 | [7 empty] | 16.9
29 | [7 empty] | 16.7625
29.1 | [7 empty] | 16.575
29.2 | [7 empty] | 16.3
29.3 | [7 empty] | 16.1429
29.4 | [7 empty] | 15.9
29.5 | [7 empty] | 15.6571
29.6 | [7 empty] | 15.44
29.7 | [7 empty] | 15.2714
29.8 | [7 empty] | 15.1

## Diagonal Profiles
Diagonal Profiles
Open field
SSD=100cm
Field Size=40x40cm2
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-42 |  |  | 2.2 | 3.8451 | 6.4644
-41.9 |  |  | 2.2 | 3.9 | 6.5
-41.8 |  |  | 2.2 | 3.9 | 6.5899
-41.7 |  |  | 2.2 | 3.9273 | 6.6061
-41.6 |  |  | 2.2 | 4 | 6.7
-41.5 |  |  | 2.2 | 4 | 6.7627
-41.4 |  |  | 2.2465 | 4.1 | 6.8
-41.3 |  |  | 2.3 | 4.1 | 6.8807
-41.2 |  |  | 2.3 | 4.1144 | 6.9
-41.1 |  |  | 2.3 | 4.2 | 6.9827
-41 |  |  | 2.3 | 4.2 | 7.0837
-40.9 |  | 1.8 | 2.3 | 4.2386 | 7.1847
-40.8 |  | 1.8 | 2.3572 | 4.3 | 7.2001
-40.7 |  | 1.8 | 2.4 | 4.3509 | 7.3215
-40.6 |  | 1.8449 | 2.4 | 4.4 | 7.4
-40.5 |  | 1.9 | 2.4 | 4.4277 | 7.4603
-40.4 |  | 1.9 | 2.4756 | 4.5 | 7.5536
-40.3 |  | 1.9 | 2.5 | 4.5052 | 7.6
-40.2 |  | 1.9 | 2.5 | 4.6 | 7.649
-40.1 | 1.9 | 1.9 | 2.5 | 4.6 | 7.75
-40 | 1.9 | 1.9 | 2.5 | 4.6082 | 7.8
-39.9 | 1.9 | 1.9 | 2.5 | 4.7 | 7.852
-39.8 | 1.9 | 1.9191 | 2.596 | 4.7102 | 7.953
-39.7 | 1.9 | 1.9976 | 2.6 | 4.8 | 8.054
-39.6 | 1.9 | 2 | 2.6 | 4.8061 | 8.1771
-39.5 | 1.9 | 2 | 2.6 | 4.8566 | 8.2154
-39.4 | 1.9 | 2 | 2.7 | 4.9071 | 8.3286
-39.3 | 2 | 2 | 2.7 | 4.9576 | 8.4296
-39.2 | 2 | 2 | 2.7 | 5.0104 | 8.5357
-39.1 | 2 | 2.003 | 2.7 | 5.0747 | 8.6
-39 | 2 | 2.1 | 2.7184 | 5.1306 | 8.6755
-38.9 | 2 | 2.1 | 2.8 | 5.1811 | 8.7765
-38.8 | 2 | 2.1 | 2.8 | 5.2316 | 8.8775
-38.7 | 2 | 2.1 | 2.8187 | 5.2821 | 8.9687
-38.6 | 2 | 2.1 | 2.9 | 5.3653 | 9.0653
-38.5 | 2.033 | 2.1 | 2.9 | 5.4663 | 9.2325
-38.4 | 2.1 | 2.1959 | 2.9 | 5.5336 | 9.3
-38.3 | 2.1 | 2.2 | 2.9 | 5.5841 | 9.3956
-38.2 | 2.1 | 2.2 | 2.9371 | 5.6441 | 9.4979
-38.1 | 2.1 | 2.2 | 3 | 5.7132 | 9.5989
-38 | 2.1 | 2.2 | 3 | 5.8142 | 9.6874
-37.9 | 2.1 | 2.2009 | 3.0813 | 5.9071 | 9.7867
-37.8 | 2.1 | 2.3 | 3.1 | 5.9542 | 9.9754
-37.7 | 2.1459 | 2.3 | 3.1 | 6.0015 | 10.0887
-37.6 | 2.2 | 2.3 | 3.1469 | 6.052 | 10.1897
-37.5 | 2.2 | 2.3 | 3.2 | 6.105 | 10.2907
-37.4 | 2.2 | 2.306 | 3.2 | 6.206 | 10.3917
-37.3 | 2.2 | 2.4 | 3.2785 | 6.3071 | 10.5855
-37.2 | 2.2 | 2.4 | 3.3 | 6.4081 | 10.6937
-37.1 | 2.2 | 2.4 | 3.3 | 6.5045 | 10.7947
-37 | 2.2 | 2.4 | 3.3815 | 6.5551 | 10.9915
-36.9 | 2.2968 | 2.4111 | 3.4 | 6.6078 | 11.0847
-36.8 | 2.3 | 2.5 | 3.4121 | 6.6785 | 11.1835
-36.7 | 2.3 | 2.5 | 3.5 | 6.8093 | 11.3691
-36.6 | 2.3 | 2.5 | 3.5141 | 6.9214 | 11.4856
-36.5 | 2.3 | 2.5294 | 3.6 | 6.9719 | 11.6731
-36.4 | 2.3767 | 2.6 | 3.6 | 7.0522 | 11.7876
-36.3 | 2.4 | 2.6 | 3.6367 | 7.1701 | 11.9772
-36.2 | 2.4 | 2.6 | 3.7 | 7.2754 | 12.0896
-36.1 | 2.4 | 2.6 | 3.7764 | 7.3764 | 12.2812
-36 | 2.4 | 2.6202 | 3.8 | 7.4774 | 12.4604
-35.9 | 2.4 | 2.6791 | 3.8212 | 7.561 | 12.5784
-35.8 | 2.4093 | 2.7351 | 3.9 | 7.6593 | 12.7588
-35.7 | 2.5 | 2.7895 | 3.9232 | 7.7662 | 12.9608
-35.6 | 2.5 | 2.8 | 4 | 7.8783 | 13.1628
-35.5 | 2.5 | 2.8 | 4.0253 | 7.9824 | 13.2824
-35.4 | 2.5 | 2.8263 | 4.1 | 8.0834 | 13.514
-35.3 | 2.6 | 2.9 | 4.1274 | 8.1986 | 13.6383
-35.2 | 2.6 | 2.9 | 4.2 | 8.3197 | 13.8139
-35.1 | 2.6 | 2.9 | 4.2412 | 8.4509 | 14.016
-35 | 2.6 | 2.9446 | 4.3 | 8.5 | 14.218
-34.9 | 2.7 | 3 | 4.4 | 8.7 | 14.42
-34.8 | 2.7 | 3.0467 | 4.4854 | 8.7378 | 14.6709
-34.7 | 2.7 | 3.1 | 4.5 | 8.8557 | 14.8535
-34.6 | 2.7368 | 3.1 | 4.5853 | 9.0261 | 15.1364
-34.5 | 2.7957 | 3.1348 | 4.608 | 9.164 | 15.3828
-34.4 | 2.8 | 3.2 | 4.7222 | 9.2651 | 15.6518
-34.3 | 2.8 | 3.2 | 4.8 | 9.3925 | 15.9124
-34.2 | 2.8837 | 3.2957 | 4.8616 | 9.5673 | 16.2155
-34.1 | 2.9 | 3.3 | 4.9 | 9.6823 | 16.5185
-34 | 2.9 | 3.3119 | 4.9834 | 9.8668 | 16.9025
-33.9 | 2.9701 | 3.4 | 5.1 | 9.9844 | 17.4076
-33.8 | 3 | 3.414 | 5.1796 | 10.1709 | 17.9126
-33.7 | 3 | 3.5 | 5.2 | 10.3176 | 18.4612
-33.6 | 3.0731 | 3.5225 | 5.2875 | 10.4606 | 19.164
-33.5 | 3.1 | 3.6 | 5.3885 | 10.6626 | 20.0588
-33.4 | 3.1283 | 3.6 | 5.4895 | 10.8323 | 20.9578
-33.3 | 3.2 | 3.6619 | 5.5905 | 10.9667 | 21.9811
-33.2 | 3.2 | 3.7629 | 5.7 | 11.1687 | 23.281
-33.1 | 3.2211 | 3.8 | 5.7354 | 11.3707 | 24.823
-33 | 3.3 | 3.8649 | 5.8364 | 11.5727 | 26.5691
-32.9 | 3.3 | 3.9 | 5.9436 | 11.7747 | 28.5854
-32.8 | 3.3241 | 3.9782 | 6.0527 | 11.9336 | 31.1673
-32.7 | 3.4 | 4.0153 | 6.1537 | 12.0755 | 33.9816
-32.6 | 3.4262 | 4.1354 | 6.2547 | 12.3523 | 37.2099
-32.5 | 3.5272 | 4.2 | 6.3558 | 12.5543 | 40.3146
-32.4 | 3.6 | 4.2 | 6.4568 | 12.7564 | 43.3606
-32.3 | 3.6 | 4.3099 | 6.5578 | 12.9584 | 46.3045
-32.2 | 3.6779 | 4.4163 | 6.6686 | 13.1906 | 48.9717
-32.1 | 3.717 | 4.4707 | 6.8026 | 13.4624 | 51.2444
-32 | 3.8 | 4.5362 | 6.9037 | 13.6967 | 52.9085
-31.9 | 3.819 | 4.619 | 7.0055 | 13.9873 | 54.3663
-31.8 | 3.9 | 4.7 | 7.12 | 14.26 | 55.6261
-31.7 | 3.921 | 4.7245 | 7.221 | 14.584 | 56.6818
-31.6 | 4.022 | 4.8364 | 7.344 | 14.9881 | 57.4541
-31.5 | 4.1 | 4.9374 | 7.523 | 15.3921 | 58.3615
-31.4 | 4.1336 | 5 | 7.6481 | 15.8202 | 59.0686
-31.3 | 4.2125 | 5.0394 | 7.8 | 16.3503 | 59.7865
-31.2 | 4.263 | 5.1404 | 7.953 | 16.9824 | 60.5689
-31.1 | 4.319 | 5.2483 | 8.1379 | 17.6896 | 61.1624
-31 | 4.3897 | 5.3662 | 8.2793 | 18.4958 | 61.8295
-30.9 | 4.5008 | 5.4 | 8.4011 | 19.6068 | 62.6041
-30.8 | 4.6 | 5.4873 | 8.5777 | 20.7801 | 63.2952
-30.7 | 4.6437 | 5.6026 | 8.6883 | 22.3441 | 63.9181
-30.6 | 4.7 | 5.7032 | 8.8786 | 24.0866 | 64.6255
-30.5 | 4.7761 | 5.7915 | 9.0806 | 26.0102 | 65.2458
-30.4 | 4.8899 | 5.8913 | 9.1913 | 28.8218 | 65.8393
-30.3 | 4.9078 | 6.0848 | 9.4366 | 31.5633 | 66.497
-30.2 | 5.0219 | 6.1934 | 9.6154 | 35.0055 | 67.103
-30.1 | 5.1201 | 6.2944 | 9.8174 | 38.7536 | 67.7091
-30 | 5.2104 | 6.3955 | 10.0194 | 42.693 | 68.2626
-29.9 | 5.3193 | 6.5228 | 10.225 | 46.5931 | 68.7343
-29.8 | 5.4261 | 6.6353 | 10.4608 | 50.2703 | 69.2418
-29.7 | 5.5237 | 6.7631 | 10.6826 | 53.4981 | 69.7653
-29.6 | 5.616 | 6.9423 | 10.8988 | 56.2581 | 70.2116
-29.5 | 5.7 | 7.0433 | 11.1154 | 58.5233 | 70.7167
-29.4 | 5.7602 | 7.1518 | 11.3761 | 60.273 | 71.1553
-29.3 | 5.9 | 7.3392 | 11.6791 | 61.6523 | 71.521
-29.2 | 5.9541 | 7.5035 | 11.97 | 62.8189 | 71.9124
-29.1 | 6.072 | 7.6091 | 12.3191 | 63.905 | 72.2792
-29 | 6.177 | 7.8065 | 12.6956 | 64.795 | 72.6355
-28.9 | 6.278 | 7.9366 | 13.157 | 65.7153 | 73.0345
-28.8 | 6.4088 | 8.1088 | 13.7455 | 66.5636 | 73.4412
-28.7 | 6.5229 | 8.2459 | 14.3687 | 67.4644 | 73.715
-28.6 | 6.6239 | 8.4479 | 15.1864 | 68.2675 | 73.9471
-28.5 | 6.7499 | 8.6499 | 16.1114 | 69.2203 | 74.3297
-28.4 | 6.926 | 8.8519 | 17.1724 | 70.0364 | 74.7113
-28.3 | 7.027 | 9.0539 | 18.479 | 70.8444 | 74.9445
-28.2 | 7.156 | 9.249 | 20.2891 | 71.6525 | 75.2784
-28.1 | 7.3254 | 9.4375 | 22.4109 | 72.4606 | 75.5612
-28 | 7.4315 | 9.6946 | 25.1467 | 73.2687 | 75.884
-27.9 | 7.6168 | 10.0086 | 28.4945 | 74.0172 | 76.0835
-27.8 | 7.7356 | 10.3622 | 32.458 | 74.7242 | 76.3747
-27.7 | 7.9188 | 10.7948 | 37.1267 | 75.4315 | 76.7105
-27.6 | 8.0554 | 11.2848 | 42.2578 | 76.0903 | 76.9462
-27.5 | 8.2 | 11.8547 | 47.0324 | 76.6964 | 77.1819
-27.4 | 8.3865 | 12.5745 | 52.3413 | 77.252 | 77.4509
-27.3 | 8.5885 | 13.4511 | 57.4847 | 77.7398 | 77.6866
-27.2 | 8.7905 | 14.5861 | 61.5982 | 78.2734 | 77.9223
-27.1 | 9.0888 | 15.8827 | 65.0253 | 78.8776 | 78.203
-27 | 9.3918 | 17.7436 | 67.8712 | 79.2891 | 78.4375
-26.9 | 9.7931 | 20.158 | 69.9262 | 79.7631 | 78.6396
-26.8 | 10.1972 | 23.0795 | 71.5621 | 80.2545 | 78.9124
-26.7 | 10.8032 | 26.8178 | 72.9336 | 80.6586 | 79.1436
-26.6 | 11.6127 | 31.2721 | 74.1739 | 81.0313 | 79.4185
-26.5 | 12.5315 | 36.4285 | 75.4471 | 81.2833 | 79.5738
-26.4 | 13.8549 | 41.8643 | 76.4909 | 81.7354 | 79.8096
-26.3 | 15.4844 | 48.9216 | 77.557 | 81.9654 | 80.0462
-26.2 | 17.4604 | 54.9927 | 78.7897 | 82.3164 | 80.2751
-26.1 | 20.5216 | 60.8909 | 79.7711 | 82.6194 | 80.4982
-26 | 24.0822 | 66.103 | 80.8822 | 82.9299 | 80.7808
-25.9 | 28.5976 | 70.2949 | 81.9848 | 83.317 | 80.9796
-25.8 | 34.3647 | 73.5435 | 82.9949 | 83.538 | 81.1934
-25.7 | 40.4998 | 75.8891 | 84.0177 | 83.9276 | 81.2963
-25.6 | 47.2175 | 77.6727 | 85.0273 | 84.1928 | 81.4945
-25.5 | 54.1788 | 79.1754 | 85.9819 | 84.3719 | 81.6966
-25.4 | 61.0374 | 80.5895 | 86.7943 | 84.6548 | 81.9317
-25.3 | 67.5995 | 82.0035 | 87.6022 | 84.8505 | 82.1539
-25.2 | 72.5533 | 83.3711 | 88.2794 | 85.1182 | 82.3424
-25.1 | 76.3955 | 84.7271 | 88.9164 | 85.3936 | 82.5331
-25 | 79.2126 | 85.9054 | 89.5915 | 85.6587 | 82.7352
-24.9 | 81.4194 | 87.1091 | 90.1976 | 85.8943 | 82.9372
-24.8 | 83.0843 | 88.398 | 90.7697 | 86.0964 | 83.1393
-24.7 | 84.5888 | 89.6333 | 91.2398 | 86.2984 | 83.3395
-24.6 | 86.0733 | 90.7667 | 91.6511 | 86.5406 | 83.5323
-24.5 | 87.3454 | 91.8778 | 92.1058 | 86.8117 | 83.6834
-24.4 | 88.8264 | 92.9081 | 92.5232 | 87.0064 | 83.8188
-24.3 | 89.9933 | 93.8364 | 92.8455 | 87.2893 | 84.0208
-24.2 | 91.3608 | 94.8291 | 93.2316 | 87.4943 | 84.2184
-24.1 | 92.6912 | 95.5876 | 93.5202 | 87.6535 | 84.3817
-24 | 93.9194 | 96.2946 | 93.8118 | 87.8556 | 84.5486
-23.9 | 95.2745 | 97.0014 | 94.1148 | 88.0576 | 84.7145
-23.8 | 96.4496 | 97.5927 | 94.4179 | 88.2596 | 84.8467
-23.7 | 97.422 | 98.0917 | 94.7209 | 88.4616 | 85.1166
-23.6 | 98.4238 | 98.7113 | 94.9743 | 88.6636 | 85.2246
-23.5 | 99.3626 | 99.065 | 95.2372 | 88.8985 | 85.3825
-23.4 | 100.1666 | 99.464 | 95.4392 | 89.1338 | 85.5549
-23.3 | 100.9442 | 99.7761 | 95.6618 | 89.2697 | 85.7574
-23.2 | 101.6436 | 100.1721 | 95.9432 | 89.4314 | 85.9102
-23.1 | 102.2497 | 100.4821 | 96.1396 | 89.5475 | 86
-23 | 102.8263 | 100.7852 | 96.3282 | 89.7597 | 86.0931
-22.9 | 103.2178 | 101.0882 | 96.6208 | 89.9321 | 86.3431
-22.8 | 103.5918 | 101.2941 | 96.8114 | 90.0953 | 86.4959
-22.7 | 103.9483 | 101.5576 | 96.9248 | 90.306 | 86.6311
-22.6 | 104.3411 | 101.7376 | 97.1403 | 90.4563 | 86.7588
-22.5 | 104.6255 | 101.9861 | 97.4144 | 90.6145 | 86.9001
-22.4 | 104.8453 | 102.1881 | 97.5309 | 90.7583 | 87.0416
-22.3 | 105.0473 | 102.3901 | 97.7165 | 90.9384 | 87.183
-22.2 | 105.3239 | 102.5921 | 97.8349 | 91.087 | 87.3444
-22.1 | 105.5513 | 102.7471 | 98.0185 | 91.2049 | 87.4859
-22 | 105.7533 | 102.8962 | 98.139 | 91.3341 | 87.6
-21.9 | 105.8777 | 103.0491 | 98.3205 | 91.5063 | 87.7774
-21.8 | 106.0835 | 103.1877 | 98.443 | 91.6145 | 87.9301
-21.7 | 106.2879 | 103.3322 | 98.6225 | 91.8083 | 88.0894
-21.6 | 106.4163 | 103.4235 | 98.7235 | 91.913 | 88.133
-21.5 | 106.5531 | 103.543 | 98.8313 | 92.0389 | 88.2744
-21.4 | 106.6541 | 103.7113 | 98.9599 | 92.1798 | 88.3966
-21.3 | 106.7483 | 103.8144 | 99.0695 | 92.3409 | 88.5123
-21.2 | 106.8733 | 103.9386 | 99.1705 | 92.4419 | 88.6133
-21.1 | 107.025 | 104.0572 | 99.3073 | 92.5429 | 88.7251
-21 | 107.1119 | 104.2164 | 99.4588 | 92.6769 | 88.8582
-20.9 | 107.1905 | 104.3 | 99.5644 | 92.8307 | 89.0185
-20.8 | 107.2 | 104.3603 | 99.7055 | 92.9317 | 89.1603
-20.7 | 107.2898 | 104.4 | 99.847 | 93.0327 | 89.2613
-20.6 | 107.3909 | 104.4623 | 99.948 | 93.1337 | 89.3623
-20.5 | 107.4919 | 104.5633 | 100.049 | 93.2347 | 89.5267
-20.4 | 107.5929 | 104.6563 | 100.1 | 93.3313 | 89.6643
-20.3 | 107.6 | 104.751 | 100.1596 | 93.4449 | 89.7572
-20.2 | 107.7 | 104.8607 | 100.2664 | 93.6 | 89.852
-20.1 | 107.7245 | 104.9 | 100.3857 | 93.6245 | 89.953
-20 | 107.8199 | 104.9797 | 100.5079 | 93.7358 | 90.0541
-19.9 | 107.8984 | 105 | 100.5786 | 93.8643 | 90.1552
-19.8 | 107.9 | 105.1 | 100.6704 | 93.9704 | 90.2562
-19.7 | 107.9857 | 105.16 | 100.7 | 94.0714 | 90.4143
-19.6 | 108 | 105.2 | 100.7724 | 94.1 | 90.5582
-19.5 | 108 | 105.2734 | 100.8 | 94.2469 | 90.6592
-19.4 | 108.0888 | 105.3 | 100.8869 | 94.3 | 90.7602
-19.3 | 108.1 | 105.3881 | 101 | 94.4 | 90.8612
-19.2 | 108.1471 | 105.4 | 101.0051 | 94.4336 | 90.9
-19.1 | 108.2 | 105.4 | 101.1 | 94.5346 | 90.9632
-19 | 108.2 | 105.4 | 101.1 | 94.6357 | 91.0999
-18.9 | 108.2 | 105.4081 | 101.1 | 94.7367 | 91.2257
-18.8 | 108.2 | 105.5 | 101.1948 | 94.8377 | 91.2964
-18.7 | 108.2 | 105.5 | 101.2 | 94.9339 | 91.3959
-18.6 | 108.2 | 105.5 | 101.2969 | 95.0223 | 91.4969
-18.5 | 108.2 | 105.5 | 101.3 | 95.1107 | 91.6246
-18.4 | 108.2 | 105.5 | 101.3 | 95.199 | 91.7417
-18.3 | 108.2 | 105.5 | 101.3713 | 95.2 | 91.8
-18.2 | 108.2 | 105.5 | 101.4 | 95.2642 | 91.8438
-18.1 | 108.2 | 105.5 | 101.4 | 95.3224 | 91.9448
-18 | 108.2 | 105.4479 | 101.4242 | 95.3729 | 92.0802
-17.9 | 108.1413 | 105.4 | 101.5 | 95.4469 | 92.2325
-17.8 | 108.0628 | 105.4 | 101.5 | 95.5479 | 92.3
-17.7 | 108 | 105.4 | 101.5 | 95.6285 | 92.3884
-17.6 | 108 | 105.4 | 101.5 | 95.6874 | 92.4249
-17.5 | 108 | 105.4 | 101.5652 | 95.7795 | 92.532
-17.4 | 108 | 105.4 | 101.6 | 95.8805 | 92.6233
-17.3 | 107.9383 | 105.4 | 101.6 | 95.9815 | 92.7243
-17.2 | 107.9 | 105.4 | 101.6 | 96.0156 | 92.8254
-17.1 | 107.8269 | 105.3796 | 101.6 | 96.1 | 92.9264
-17 | 107.8 | 105.3291 | 101.6 | 96.1784 | 93.0274
-16.9 | 107.7668 | 105.3 | 101.6 | 96.2 | 93.1284
-16.8 | 107.7078 | 105.3 | 101.6 | 96.2866 | 93.2294
-16.7 | 107.6489 | 105.2739 | 101.6027 | 96.3 | 93.3304
-16.6 | 107.6 | 105.215 | 101.7 | 96.3886 | 93.4
-16.5 | 107.5818 | 105.2 | 101.7 | 96.4896 | 93.4455
-16.4 | 107.5 | 105.2 | 101.7 | 96.5 | 93.5724
-16.3 | 107.4798 | 105.2 | 101.7 | 96.5802 | 93.6677
-16.2 | 107.4 | 105.1103 | 101.7297 | 96.6 | 93.7
-16.1 | 107.4 | 105.1 | 101.8 | 96.7 | 93.7652
-16 | 107.3675 | 105.0895 | 101.8 | 96.7 | 93.8662
-15.9 | 107.3 | 105 | 101.8 | 96.71 | 93.9672
-15.8 | 107.2318 | 105 | 101.8 | 96.8 | 94
-15.7 | 107.1308 | 104.9736 | 101.8 | 96.8 | 94.0692
-15.6 | 107.1 | 104.9 | 101.8 | 96.8988 | 94.1702
-15.5 | 107.0801 | 104.9 | 101.8 | 96.9998 | 94.2623
-15.4 | 107.0095 | 104.8743 | 101.8 | 97.0211 | 94.358
-15.3 | 107 | 104.8 | 101.8 | 97.1 | 94.4
-15.2 | 106.94 | 104.8 | 101.8 | 97.1033 | 94.46
-15.1 | 106.839 | 104.8 | 101.8 | 97.2 | 94.561
-15 | 106.8 | 104.7094 | 101.8 | 97.2 | 94.662
-14.9 | 106.7918 | 104.7 | 101.8 | 97.2202 | 94.7
-14.8 | 106.7 | 104.7 | 101.8 | 97.3 | 94.764
-14.7 | 106.689 | 104.7 | 101.8 | 97.3079 | 94.8569
-14.6 | 106.6 | 104.6339 | 101.8 | 97.4 | 94.9
-14.5 | 106.5187 | 104.6 | 101.8 | 97.4099 | 94.9528
-14.4 | 106.5 | 104.5319 | 101.8 | 97.5109 | 95.0419
-14.3 | 106.4166 | 104.5 | 101.8 | 97.6 | 95.1306
-14.2 | 106.4 | 104.5 | 101.8 | 97.6129 | 95.2485
-14.1 | 106.3145 | 104.5 | 101.8 | 97.7 | 95.3
-14 | 106.3 | 104.4421 | 101.8 | 97.7 | 95.3579
-13.9 | 106.2125 | 104.4 | 101.8 | 97.7 | 95.4
-13.8 | 106.1226 | 104.3134 | 101.8 | 97.7038 | 95.4457
-13.7 | 106.1 | 104.3 | 101.8 | 97.8 | 95.5
-13.6 | 106.0238 | 104.3 | 101.8 | 97.8066 | 95.5477
-13.5 | 106 | 104.3 | 101.8 | 97.9 | 95.6487
-13.4 | 106 | 104.2075 | 101.8 | 98 | 95.7497
-13.3 | 105.9493 | 104.2 | 101.7922 | 98.0425 | 95.8507
-13.2 | 105.8483 | 104.2 | 101.7 | 98.1 | 95.9
-13.1 | 105.8 | 104.1718 | 101.7 | 98.1 | 95.9527
-13 | 105.7463 | 104.1 | 101.7 | 98.1 | 96
-12.9 | 105.7 | 104.1 | 101.7 | 98.1 | 96.0547
-12.8 | 105.6442 | 104.0157 | 101.7 | 98.1545 | 96.1
-12.7 | 105.6 | 104 | 101.7 | 98.2 | 96.1568
-12.6 | 105.5422 | 103.9136 | 101.7 | 98.2506 | 96.2578
-12.5 | 105.4485 | 103.9 | 101.7 | 98.3 | 96.3588
-12.4 | 105.4 | 103.8116 | 101.7 | 98.3 | 96.4
-12.3 | 105.4 | 103.8 | 101.6948 | 98.3466 | 96.4608
-12.2 | 105.3238 | 103.8 | 101.6 | 98.4 | 96.5
-12.1 | 105.2 | 103.7508 | 101.6 | 98.4 | 96.5486
-12 | 105.179 | 103.7 | 101.6 | 98.4 | 96.6
-11.9 | 105.1 | 103.7 | 101.6 | 98.4506 | 96.6506
-11.8 | 105.0678 | 103.6821 | 101.6 | 98.5 | 96.7516
-11.7 | 105 | 103.6035 | 101.6 | 98.5 | 96.8461
-11.6 | 104.9 | 103.5325 | 101.6 | 98.5536 | 96.9
-11.5 | 104.9 | 103.5 | 101.6 | 98.6 | 96.9
-11.4 | 104.9 | 103.4763 | 101.6 | 98.6557 | 96.9986
-11.3 | 104.8147 | 103.4 | 101.5147 | 98.7 | 97
-11.2 | 104.7 | 103.4 | 101.5 | 98.7404 | 97.0782
-11.1 | 104.7 | 103.4 | 101.5 | 98.8 | 97.1
-11 | 104.7 | 103.4 | 101.5 | 98.8 | 97.1523
-10.9 | 104.6821 | 103.3821 | 101.5 | 98.8032 | 97.2
-10.8 | 104.6 | 103.2811 | 101.5 | 98.8915 | 97.2475
-10.7 | 104.5601 | 103.2 | 101.5 | 98.9 | 97.3
-10.6 | 104.4505 | 103.2 | 101.5 | 98.9247 | 97.3433
-10.5 | 104.3495 | 103.2 | 101.5 | 98.9753 | 97.4
-10.4 | 104.3 | 103.177 | 101.4199 | 99 | 97.4373
-10.3 | 104.2475 | 103.1 | 101.4947 | 99 | 97.5383
-10.2 | 104.2 | 103.1 | 101.4035 | 99.0417 | 97.6
-10.1 | 104.1454 | 103.0739 | 101.4 | 99.1 | 97.6403
-10 | 104.1 | 103 | 101.4 | 99.1 | 97.7
-9.9 | 104.0434 | 102.9782 | 101.4 | 99.1 | 97.7281
-9.8 | 104 | 102.9 | 101.4 | 99.1 | 97.8291
-9.7 | 104 | 102.9 | 101.4 | 99.1729 | 97.9
-9.6 | 103.9304 | 102.9 | 101.4 | 99.2 | 97.9
-9.5 | 103.9 | 102.8459 | 101.4 | 99.2 | 97.9
-9.4 | 103.8097 | 102.8 | 101.4 | 99.2 | 97.9331
-9.3 | 103.8 | 102.7826 | 101.3087 | 99.2627 | 98.0299
-9.2 | 103.759 | 102.7 | 101.3 | 99.3 | 98.1
-9.1 | 103.6558 | 102.6924 | 101.3 | 99.3 | 98.1307
-9 | 103.6 | 102.6 | 101.3 | 99.3 | 98.2515
-8.9 | 103.6 | 102.6 | 101.3 | 99.3 | 98.3
-8.8 | 103.5465 | 102.5876 | 101.3 | 99.3343 | 98.3
-8.7 | 103.5 | 102.5 | 101.3 | 99.4 | 98.3402
-8.6 | 103.4445 | 102.5 | 101.3 | 99.4 | 98.4
-8.5 | 103.3505 | 102.5 | 101.2149 | 99.4327 | 98.4423
-8.4 | 103.3 | 102.4495 | 101.2 | 99.5 | 98.5
-8.3 | 103.2487 | 102.3414 | 101.2 | 99.5 | 98.5
-8.2 | 103.2 | 102.3 | 101.2 | 99.5 | 98.5453
-8.1 | 103.1459 | 102.3 | 101.2 | 99.5463 | 98.6
-8 | 103.1 | 102.2835 | 101.2 | 99.5526 | 98.6386
-7.9 | 103.1 | 102.233 | 101.1102 | 99.5484 | 98.7
-7.8 | 103.0221 | 102.2 | 101.1 | 99.6 | 98.7
-7.7 | 102.9386 | 102.2 | 101.1 | 99.6 | 98.7
-7.6 | 102.9 | 102.174 | 101.1 | 99.6 | 98.7371
-7.5 | 102.8619 | 102.1033 | 101.1 | 99.6381 | 98.8
-7.4 | 102.7608 | 102.1 | 101.1 | 99.7 | 98.8685
-7.3 | 102.7 | 102.0027 | 101 | 99.7 | 98.9
-7.2 | 102.7 | 102 | 101 | 99.736 | 98.9
-7.1 | 102.6174 | 101.9292 | 101 | 99.8 | 98.9708
-7 | 102.6 | 101.9 | 101 | 99.8 | 99
-6.9 | 102.6 | 101.9 | 101 | 99.8 | 99
-6.8 | 102.5 | 101.8262 | 101 | 99.8 | 99.0881
-6.7 | 102.4627 | 101.8 | 101 | 99.8 | 99.1
-6.6 | 102.4 | 101.8 | 101 | 99.8164 | 99.1
-6.5 | 102.4 | 101.7946 | 101 | 99.9 | 99.1912
-6.4 | 102.3793 | 101.7 | 100.9507 | 99.9 | 99.2
-6.3 | 102.3 | 101.6925 | 100.9 | 99.9 | 99.2932
-6.2 | 102.2934 | 101.6 | 100.9 | 99.9 | 99.3
-6.1 | 102.2148 | 101.6 | 100.9 | 99.9 | 99.3
-6 | 102.1181 | 101.5876 | 100.9 | 99.9478 | 99.3962
-5.9 | 102.1 | 101.5371 | 100.9 | 100 | 99.4
-5.8 | 102.016 | 101.5 | 100.9 | 100 | 99.4
-5.7 | 102 | 101.5 | 100.8823 | 100 | 99.485
-5.6 | 101.914 | 101.4747 | 100.818 | 100 | 99.5
-5.5 | 101.9 | 101.4 | 100.8 | 100 | 99.5
-5.4 | 101.812 | 101.4 | 100.8 | 100.0469 | 99.5
-5.3 | 101.8 | 101.4 | 100.8 | 100.1 | 99.5
-5.2 | 101.8 | 101.3956 | 100.8 | 100.1 | 99.5901
-5.1 | 101.7089 | 101.3 | 100.8 | 100.1 | 99.6
-5 | 101.6463 | 101.3 | 100.7508 | 100.1 | 99.6806
-4.9 | 101.6 | 101.2926 | 100.7 | 100.1 | 99.7
-4.8 | 101.6 | 101.2 | 100.7 | 100.1 | 99.7
-4.7 | 101.5576 | 101.2 | 100.7 | 100.1 | 99.7
-4.6 | 101.4944 | 101.1896 | 100.7 | 100.1 | 99.7
-4.5 | 101.44 | 101.1 | 100.66 | 100.1 | 99.7
-4.4 | 101.4 | 101.1 | 100.6 | 100.1 | 99.7
-4.3 | 101.4 | 101.0865 | 100.6 | 100.1 | 99.7
-4.2 | 101.3197 | 101 | 100.6 | 100.1 | 99.7336
-4.1 | 101.2131 | 101 | 100.6 | 100.1 | 99.8
-4 | 101.2 | 100.923 | 100.6 | 100.1 | 99.8
-3.9 | 101.1253 | 100.8627 | 100.6 | 100.1 | 99.8461
-3.8 | 101.1 | 100.8121 | 100.6 | 100.1 | 99.9
-3.7 | 101.1 | 100.8 | 100.5126 | 100.1053 | 99.9
-3.6 | 101.0223 | 100.7937 | 100.5 | 100.2 | 99.9
-3.5 | 101 | 100.7 | 100.5 | 100.2 | 99.9
-3.4 | 101 | 100.7 | 100.4604 | 100.2 | 99.9
-3.3 | 100.9 | 100.7 | 100.4 | 100.2 | 99.9
-3.2 | 100.9 | 100.7 | 100.4 | 100.2 | 99.9
-3.1 | 100.86 | 100.6029 | 100.4 | 100.2 | 99.9
-3 | 100.7681 | 100.6 | 100.4 | 100.2 | 99.9
-2.9 | 100.7 | 100.6 | 100.4 | 100.2 | 99.9
-2.8 | 100.7 | 100.5713 | 100.4 | 100.2 | 99.943
-2.7 | 100.6703 | 100.5 | 100.3274 | 100.2 | 100
-2.6 | 100.6 | 100.4693 | 100.3 | 100.1835 | 100
-2.5 | 100.6 | 100.4 | 100.3 | 100.1 | 100
-2.4 | 100.6 | 100.4 | 100.3 | 100.1 | 100
-2.3 | 100.5662 | 100.4 | 100.2376 | 100.1 | 100
-2.2 | 100.5 | 100.3713 | 100.2 | 100.1 | 100
-2.1 | 100.5 | 100.3 | 100.2 | 100.1 | 100
-2 | 100.453 | 100.3 | 100.2 | 100.1 | 100
-1.9 | 100.4025 | 100.3 | 100.2 | 100.1 | 100
-1.8 | 100.344 | 100.2468 | 100.2 | 100.1 | 100
-1.7 | 100.3 | 100.2 | 100.1042 | 100.1 | 100
-1.6 | 100.3 | 100.2 | 100.1 | 100.1 | 100
-1.5 | 100.2738 | 100.2 | 100.1 | 100.1 | 100
-1.4 | 100.2266 | 100.2 | 100.1 | 100.1 | 100
-1.3 | 100.2 | 100.2 | 100.1 | 100.1 | 100
-1.2 | 100.2 | 100.1732 | 100.1 | 100.1 | 100
-1.1 | 100.1578 | 100.1 | 100.0969 | 100.1 | 100
-1 | 100.1 | 100.1 | 100 | 100.1 | 100
-0.9 | 100.1 | 100.1 | 100 | 100.1 | 100
-0.8 | 100.0946 | 100.0755 | 100 | 100.1 | 100
-0.7 | 100.0062 | 100.025 | 100 | 100.1 | 100
-0.6 | 100 | 100 | 100 | 100.1 | 100
-0.5 | 100 | 100 | 100 | 100.1 | 100
-0.4 | 100 | 100 | 100 | 100.1 | 100
-0.3 | 100 | 100 | 100 | 100.0459 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100.0173 | 100
0.4 | 100 | 100 | 100 | 100.1 | 100
0.5 | 100 | 100 | 100 | 100.1 | 100
0.6 | 100 | 100.0028 | 100 | 100.1 | 100
0.7 | 100.0214 | 100.05 | 100 | 100.1 | 100
0.8 | 100.1 | 100.0971 | 100 | 100.1 | 100
0.9 | 100.1 | 100.1 | 100 | 100.1 | 100
1 | 100.1156 | 100.1 | 100 | 100.1 | 100
1.1 | 100.1798 | 100.1 | 100.0683 | 100.1 | 100
1.2 | 100.2 | 100.1 | 100.1 | 100.1 | 100
1.3 | 100.2 | 100.1769 | 100.1 | 100.1 | 100
1.4 | 100.25 | 100.2 | 100.1 | 100.1 | 100
1.5 | 100.3 | 100.2 | 100.1 | 100.1 | 100
1.6 | 100.3 | 100.2 | 100.1 | 100.1 | 100
1.7 | 100.3019 | 100.2 | 100.1701 | 100.1 | 100
1.8 | 100.3662 | 100.2 | 100.2 | 100.1 | 100
1.9 | 100.4239 | 100.2764 | 100.2 | 100.1 | 100
2 | 100.4744 | 100.3 | 100.2 | 100.1 | 100
2.1 | 100.5 | 100.3 | 100.2 | 100.1 | 100
2.2 | 100.5 | 100.3927 | 100.2081 | 100.1 | 100
2.3 | 100.5829 | 100.4 | 100.3 | 100.1 | 100
2.4 | 100.6 | 100.4 | 100.3 | 100.1 | 100
2.5 | 100.6 | 100.4 | 100.3 | 100.1 | 100
2.6 | 100.6 | 100.4975 | 100.3106 | 100.155 | 100
2.7 | 100.6988 | 100.5 | 100.399 | 100.2 | 99.9583
2.8 | 100.7 | 100.5874 | 100.4 | 100.2 | 99.9
2.9 | 100.7009 | 100.6 | 100.4 | 100.2 | 99.9
3 | 100.8 | 100.6 | 100.4 | 100.2 | 99.9
3.1 | 100.8 | 100.6315 | 100.4 | 100.1986 | 99.9
3.2 | 100.8896 | 100.7 | 100.4 | 100.1515 | 99.9
3.3 | 100.9069 | 100.7 | 100.4 | 100.1044 | 99.9
3.4 | 101 | 100.7 | 100.4354 | 100.1713 | 99.9
3.5 | 101 | 100.7212 | 100.5 | 100.2 | 99.9
3.6 | 101.0365 | 100.8 | 100.5 | 100.2 | 99.9
3.7 | 101.1 | 100.8 | 100.5 | 100.1339 | 99.9
3.8 | 101.1 | 100.8407 | 100.6 | 100.1 | 99.9
3.9 | 101.1396 | 100.8912 | 100.6 | 100.1 | 99.8747
4 | 101.2 | 100.9531 | 100.6 | 100.1 | 99.8
4.1 | 101.2416 | 101 | 100.6 | 100.1 | 99.8
4.2 | 101.3 | 101.0141 | 100.6 | 100.1 | 99.7603
4.3 | 101.4 | 101.1 | 100.6 | 100.1 | 99.7
4.4 | 101.4011 | 101.1 | 100.6188 | 100.1 | 99.7
4.5 | 101.46 | 101.115 | 100.7 | 100.1 | 99.7
4.6 | 101.5 | 101.2 | 100.7 | 100.1 | 99.7
4.7 | 101.5 | 101.2 | 100.7 | 100.1 | 99.7
4.8 | 101.5492 | 101.2 | 100.7 | 100.1 | 99.7
4.9 | 101.6124 | 101.2069 | 100.7 | 100.1 | 99.7
5 | 101.6713 | 101.3 | 100.7222 | 100.1 | 99.7
5.1 | 101.7 | 101.3 | 100.8 | 100.1 | 99.6054
5.2 | 101.7385 | 101.3099 | 100.8 | 100.1 | 99.6
5.3 | 101.8 | 101.4 | 100.8 | 100.1 | 99.5154
5.4 | 101.8405 | 101.4 | 100.8 | 100.0705 | 99.5
5.5 | 101.9 | 101.4 | 100.8 | 100.0161 | 99.5
5.6 | 101.9 | 101.4 | 100.8 | 100 | 99.4146
5.7 | 101.9436 | 101.4542 | 100.8605 | 100 | 99.4864
5.8 | 102 | 101.508 | 100.9 | 99.9988 | 99.4125
5.9 | 102.0399 | 101.5585 | 100.9 | 99.9281 | 99.4
6 | 102.1323 | 101.6 | 100.9 | 99.9 | 99.3105
6.1 | 102.2292 | 101.6 | 100.9 | 99.9 | 99.3
6.2 | 102.3 | 101.6301 | 100.9 | 99.9 | 99.3
6.3 | 102.3 | 101.7 | 100.9211 | 99.9 | 99.219
6.4 | 102.3061 | 101.7221 | 101 | 99.9 | 99.2
6.5 | 102.3847 | 101.8 | 101 | 99.8265 | 99.1197
6.6 | 102.4 | 101.8 | 101 | 99.8 | 99.1
6.7 | 102.4 | 101.8 | 101 | 99.8 | 99.1
6.8 | 102.5 | 101.8548 | 101 | 99.8 | 99.1
6.9 | 102.5272 | 101.9 | 101 | 99.8 | 99.0157
7 | 102.6 | 101.9 | 101 | 99.8 | 99
7.1 | 102.6 | 101.9578 | 101 | 99.755 | 98.9992
7.2 | 102.6686 | 102 | 101 | 99.7 | 98.9
7.3 | 102.7 | 102 | 101 | 99.7 | 98.9
7.4 | 102.7894 | 102.0323 | 101.0876 | 99.7 | 98.8948
7.5 | 102.8904 | 102.1 | 101.1 | 99.6667 | 98.8
7.6 | 102.9 | 102.1029 | 101.1 | 99.6 | 98.7657
7.7 | 102.9647 | 102.1534 | 101.1 | 99.6 | 98.7
7.8 | 103 | 102.2039 | 101.1 | 99.5682 | 98.7
7.9 | 103.0516 | 102.2544 | 101.1923 | 99.5231 | 98.7
8 | 103.1 | 102.3 | 101.2 | 99.5759 | 98.6616
8.1 | 103.1793 | 102.3 | 101.2 | 99.5 | 98.6
8.2 | 103.2 | 102.3 | 101.2 | 99.5 | 98.5596
8.3 | 103.27 | 102.37 | 101.2 | 99.5 | 98.5
8.4 | 103.3 | 102.4 | 101.2 | 99.5 | 98.4719
8.5 | 103.372 | 102.5 | 101.2006 | 99.4566 | 98.4
8.6 | 103.4 | 102.5 | 101.3 | 99.4 | 98.4
8.7 | 103.474 | 102.5026 | 101.3 | 99.4 | 98.3688
8.8 | 103.5751 | 102.6 | 101.3 | 99.3535 | 98.3
8.9 | 103.6 | 102.6 | 101.3 | 99.3 | 98.3
9 | 103.6 | 102.6 | 101.3 | 99.3 | 98.2658
9.1 | 103.6781 | 102.621 | 101.3 | 99.3 | 98.1589
9.2 | 103.7 | 102.7 | 101.3 | 99.3 | 98.0495
9.3 | 103.8 | 102.7146 | 101.3 | 99.2913 | 98
9.4 | 103.8383 | 102.7789 | 101.3835 | 99.2 | 97.9474
9.5 | 103.9 | 102.8679 | 101.4 | 99.2 | 97.9
9.6 | 103.9 | 102.9 | 101.4 | 99.2 | 97.9
9.7 | 103.9699 | 102.9 | 101.4 | 99.1872 | 97.9
9.8 | 104 | 102.936 | 101.4 | 99.1 | 97.8576
9.9 | 104.0719 | 103 | 101.4 | 99.1 | 97.7566
10 | 104.1 | 103.0015 | 101.4 | 99.1 | 97.7
10.1 | 104.1739 | 103.1 | 101.4 | 99.1 | 97.6546
10.2 | 104.2 | 103.1 | 101.4 | 99.075 | 97.6
10.3 | 104.276 | 103.1045 | 101.4 | 99 | 97.5585
10.4 | 104.3 | 103.2 | 101.4056 | 99 | 97.4659
10.5 | 104.3781 | 103.2 | 101.5 | 98.9967 | 97.3648
10.6 | 104.4 | 103.2 | 101.5 | 98.9462 | 97.3
10.7 | 104.4801 | 103.2086 | 101.5 | 98.9 | 97.3
10.8 | 104.6 | 103.3 | 101.5 | 98.9 | 97.2618
10.9 | 104.6 | 103.3107 | 101.5 | 98.8179 | 97.2
11 | 104.6974 | 103.4 | 101.5 | 98.8 | 97.1773
11.1 | 104.7 | 103.4 | 101.5 | 98.8 | 97.1
11.2 | 104.716 | 103.4 | 101.5 | 98.756 | 97.1
11.3 | 104.8338 | 103.4004 | 101.5004 | 98.7 | 97.0219
11.4 | 104.9 | 103.5 | 101.6 | 98.6737 | 97
11.5 | 104.9 | 103.5 | 101.6 | 98.6 | 96.9261
11.6 | 104.9207 | 103.5624 | 101.6 | 98.5822 | 96.8476
11.7 | 105 | 103.6368 | 101.6 | 98.5 | 96.7711
11.8 | 105.0877 | 103.7 | 101.6 | 98.5 | 96.7
11.9 | 105.1066 | 103.7 | 101.6 | 98.4792 | 96.6792
12 | 105.2 | 103.7195 | 101.6 | 98.4 | 96.6
12.1 | 105.212 | 103.7738 | 101.6 | 98.4 | 96.5772
12.2 | 105.3445 | 103.8 | 101.6 | 98.4 | 96.5
12.3 | 105.4 | 103.8 | 101.6624 | 98.3752 | 96.4752
12.4 | 105.4 | 103.8402 | 101.7 | 98.3 | 96.4
12.5 | 105.4698 | 103.9 | 101.7 | 98.3 | 96.3731
12.6 | 105.5 | 103.9 | 101.7 | 98.2721 | 96.2721
12.7 | 105.5718 | 103.9432 | 101.7 | 98.2 | 96.1747
12.8 | 105.6728 | 104.0442 | 101.7 | 98.1863 | 96.1
12.9 | 105.7 | 104.1 | 101.7 | 98.1 | 96.069
13 | 105.7748 | 104.1 | 101.7 | 98.1 | 95.972
13.1 | 105.8 | 104.2 | 101.7 | 98.1 | 95.9
13.2 | 105.8769 | 104.2 | 101.7 | 98.0936 | 95.8803
13.3 | 105.9779 | 104.2 | 101.7636 | 98 | 95.8
13.4 | 106 | 104.2361 | 101.8 | 98 | 95.7783
13.5 | 106 | 104.3 | 101.8 | 97.9 | 95.6773
13.6 | 106.0523 | 104.3 | 101.8 | 97.8334 | 95.5762
13.7 | 106.1 | 104.3 | 101.8 | 97.8 | 95.5
13.8 | 106.1543 | 104.3401 | 101.8 | 97.7438 | 95.4741
13.9 | 106.2 | 104.4 | 101.8 | 97.7 | 95.3731
14 | 106.2421 | 104.4565 | 101.8 | 97.7 | 95.3
14.1 | 106.3431 | 104.5 | 101.8 | 97.6283 | 95.2747
14.2 | 106.4 | 104.5 | 101.8 | 97.6 | 95.1832
14.3 | 106.4452 | 104.5 | 101.8 | 97.6 | 95.0744
14.4 | 106.5 | 104.5605 | 101.8 | 97.5395 | 95
14.5 | 106.5472 | 104.6 | 101.8 | 97.4385 | 94.9813
14.6 | 106.6 | 104.6625 | 101.8 | 97.4 | 94.8803
14.7 | 106.7 | 104.7 | 101.8 | 97.3365 | 94.8
14.8 | 106.7 | 104.7 | 101.8 | 97.3 | 94.7783
14.9 | 106.8 | 104.7 | 101.8 | 97.2344 | 94.7
15 | 106.8 | 104.738 | 101.8 | 97.2 | 94.6792
15.1 | 106.8676 | 104.8 | 101.8 | 97.2 | 94.5896
15.2 | 106.9686 | 104.8 | 101.8 | 97.1239 | 94.4886
15.3 | 107 | 104.811 | 101.8 | 97.1 | 94.3875
15.4 | 107 | 104.8993 | 101.8 | 97.0412 | 94.3
15.5 | 107.0001 | 104.9 | 101.8 | 96.9284 | 94.2855
15.6 | 107.0644 | 104.9012 | 101.8 | 96.9 | 94.1844
15.7 | 107.1527 | 105 | 101.8 | 96.8264 | 94.0855
15.8 | 107.2 | 105 | 101.8 | 96.8 | 94
15.9 | 107.2717 | 105.0042 | 101.8 | 96.7242 | 93.9958
16 | 107.3895 | 105.1 | 101.8 | 96.7 | 93.8947
16.1 | 107.4 | 105.1 | 101.8 | 96.7 | 93.7937
16.2 | 107.4074 | 105.1 | 101.7697 | 96.6098 | 93.7
16.3 | 107.5 | 105.1764 | 101.7 | 96.506 | 93.6916
16.4 | 107.5094 | 105.2 | 101.7 | 96.5 | 93.5906
16.5 | 107.6 | 105.2 | 101.7 | 96.5 | 93.4854
16.6 | 107.608 | 105.2436 | 101.7 | 96.415 | 93.4
16.7 | 107.6787 | 105.3 | 101.6355 | 96.3022 | 93.3447
16.8 | 107.738 | 105.3 | 101.6 | 96.3 | 93.2508
16.9 | 107.7924 | 105.3 | 101.6 | 96.2141 | 93.157
17 | 107.8 | 105.3 | 101.6 | 96.1986 | 93.056
17.1 | 107.8451 | 105.3082 | 101.6 | 96.1 | 92.9099
17.2 | 107.9 | 105.3587 | 101.6 | 96.0463 | 92.8
17.3 | 107.9549 | 105.4 | 101.6 | 95.9958 | 92.7529
17.4 | 108 | 105.4 | 101.5829 | 95.8954 | 92.6519
17.5 | 108 | 105.4 | 101.5 | 95.8071 | 92.5509
17.6 | 108 | 105.4 | 101.5 | 95.7071 | 92.4499
17.7 | 108.0064 | 105.4 | 101.5 | 95.653 | 92.4
17.8 | 108.085 | 105.4 | 101.5 | 95.6025 | 92.3
17.9 | 108.1715 | 105.4 | 101.5 | 95.5039 | 92.2611
18 | 108.2 | 105.4679 | 101.4441 | 95.4029 | 92.1202
18.1 | 108.2 | 105.5 | 101.4 | 95.351 | 91.9591
18.2 | 108.2 | 105.5 | 101.4 | 95.3005 | 91.8581
18.3 | 108.2 | 105.5 | 101.3999 | 95.2222 | 91.7624
18.4 | 108.2 | 105.5 | 101.3 | 95.1275 | 91.7
18.5 | 108.2 | 105.5 | 101.3 | 95.0265 | 91.6642
18.6 | 108.2 | 105.5 | 101.3 | 94.942 | 91.5463
18.7 | 108.2 | 105.5 | 101.2244 | 94.9 | 91.408
18.8 | 108.2 | 105.483 | 101.2 | 94.858 | 91.2565
18.9 | 108.2 | 105.4 | 101.1224 | 94.7653 | 91.1507
19 | 108.2 | 105.4 | 101.1 | 94.6642 | 91.0722
19.1 | 108.2 | 105.4 | 101.0303 | 94.5632 | 90.9918
19.2 | 108.1071 | 105.4 | 101 | 94.4622 | 90.8908
19.3 | 108.1 | 105.4 | 100.9214 | 94.4 | 90.7898
19.4 | 108.0602 | 105.3459 | 100.9 | 94.3203 | 90.6888
19.5 | 108 | 105.2448 | 100.8024 | 94.2024 | 90.6
19.6 | 108 | 105.2 | 100.8 | 94.101 | 90.5735
19.7 | 107.9572 | 105.1333 | 100.7 | 94 | 90.3857
19.8 | 107.9 | 105.0988 | 100.699 | 93.9988 | 90.2846
19.9 | 107.8686 | 105 | 100.5974 | 93.8809 | 90.1836
20 | 107.797 | 104.9631 | 100.4688 | 93.7631 | 90.0826
20.1 | 107.7 | 104.9 | 100.3635 | 93.653 | 89.9816
20.2 | 107.7 | 104.8378 | 100.2806 | 93.552 | 89.8806
20.3 | 107.6654 | 104.7368 | 100.1796 | 93.451 | 89.7592
20.4 | 107.5643 | 104.6358 | 100.0786 | 93.3 | 89.5813
20.5 | 107.5 | 104.5347 | 100 | 93.1554 | 89.4919
20.6 | 107.4623 | 104.5 | 99.9766 | 93.1 | 89.3909
20.7 | 107.3613 | 104.4327 | 99.8511 | 93.0613 | 89.2898
20.8 | 107.2603 | 104.3317 | 99.6745 | 92.9603 | 89.1776
20.9 | 107.1644 | 104.3 | 99.5629 | 92.8185 | 88.9878
21 | 107.0786 | 104.1594 | 99.4215 | 92.6582 | 88.8846
21.1 | 107.0001 | 104.0286 | 99.32 | 92.5625 | 88.7601
21.2 | 106.8233 | 103.9155 | 99.2257 | 92.4705 | 88.5838
21.3 | 106.7233 | 103.7981 | 99.0898 | 92.3695 | 88.5
21.4 | 106.6256 | 103.6971 | 98.9383 | 92.237 | 88.4132
21.5 | 106.5245 | 103.596 | 98.8413 | 92.0349 | 88.3103
21.6 | 106.3941 | 103.4912 | 98.7442 | 91.9 | 88.1
21.7 | 106.2594 | 103.3145 | 98.5737 | 91.7737 | 88.0116
21.8 | 106.0502 | 103.2188 | 98.4439 | 91.5717 | 87.9627
21.9 | 105.8572 | 103.041 | 98.3491 | 91.4348 | 87.7175
22 | 105.7481 | 102.9195 | 98.1962 | 91.3421 | 87.5591
22.1 | 105.5412 | 102.737 | 97.9941 | 91.2382 | 87.5
22.2 | 105.2921 | 102.5349 | 97.8461 | 91.1045 | 87.3645
22.3 | 105.0901 | 102.4165 | 97.7451 | 90.9042 | 87.2231
22.4 | 104.8321 | 102.2309 | 97.5881 | 90.7726 | 87.0617
22.5 | 104.5291 | 102.0289 | 97.443 | 90.6503 | 86.9335
22.6 | 104.2376 | 101.7975 | 97.2261 | 90.4735 | 86.7984
22.7 | 103.9515 | 101.539 | 96.982 | 90.296 | 86.6478
22.8 | 103.6519 | 101.337 | 96.7799 | 90.0838 | 86.516
22.9 | 103.238 | 101.1024 | 96.6389 | 89.9013 | 86.3491
23 | 102.8711 | 100.7994 | 96.4136 | 89.7382 | 86.1164
23.1 | 102.2788 | 100.4964 | 96.1737 | 89.5797 | 85.9717
23.2 | 101.6515 | 100.1578 | 95.9717 | 89.4004 | 85.9
23.3 | 100.9371 | 99.7903 | 95.7291 | 89.2492 | 85.6888
23.4 | 100.1414 | 99.4497 | 95.4964 | 89.1093 | 85.5474
23.5 | 99.3339 | 99.0842 | 95.2767 | 88.8842 | 85.4059
23.6 | 98.4532 | 98.6782 | 95.041 | 88.6208 | 85.2645
23.7 | 97.5233 | 98.1124 | 94.7078 | 88.4594 | 85.0902
23.8 | 96.4436 | 97.6061 | 94.4036 | 88.3168 | 84.8881
23.9 | 95.2361 | 96.965 | 94.1006 | 88.1147 | 84.743
24 | 94.0473 | 96.2945 | 93.7984 | 87.8691 | 84.6098
24.1 | 92.8047 | 95.5875 | 93.5927 | 87.6479 | 84.4392
24.2 | 91.4916 | 94.7891 | 93.1901 | 87.4601 | 84.2372
24.3 | 90.1461 | 93.8333 | 92.8364 | 87.2243 | 84.0514
24.4 | 88.6932 | 92.8232 | 92.4806 | 86.9732 | 83.8998
24.5 | 87.3884 | 91.8126 | 92.0795 | 86.7636 | 83.7196
24.6 | 85.9134 | 90.7522 | 91.6812 | 86.5172 | 83.484
24.7 | 84.6405 | 89.5903 | 91.1778 | 86.3556 | 83.2483
24.8 | 83.0755 | 88.3782 | 90.725 | 86.1303 | 83.0438
24.9 | 81.5261 | 87.1132 | 90.1766 | 85.8515 | 82.8923
25 | 79.4617 | 85.8879 | 89.563 | 85.6242 | 82.7263
25.1 | 76.4799 | 84.6764 | 88.957 | 85.3387 | 82.5377
25.2 | 72.8074 | 83.3706 | 88.3509 | 85.1031 | 82.3455
25.3 | 67.9151 | 82.0763 | 87.6525 | 84.9006 | 82.1434
25.4 | 61.3417 | 80.7901 | 86.7601 | 84.5981 | 81.9441
25.5 | 53.9778 | 79.3742 | 85.9024 | 84.3329 | 81.7513
25.6 | 46.9062 | 77.8536 | 85.0849 | 84.1451 | 81.5516
25.7 | 40.0332 | 76.1058 | 84.0191 | 83.8742 | 81.3495
25.8 | 34.0736 | 73.4665 | 83.0808 | 83.5332 | 81.2111
25.9 | 28.3933 | 70.1975 | 82.0078 | 83.2741 | 80.9648
26 | 23.8372 | 65.8948 | 80.8967 | 82.9442 | 80.7209
26.1 | 20.3198 | 60.6278 | 79.8505 | 82.6052 | 80.538
26.2 | 17.5134 | 54.8084 | 78.7085 | 82.3021 | 80.3095
26.3 | 15.4896 | 48.3435 | 77.6303 | 81.9991 | 80.0324
26.4 | 13.873 | 42.1385 | 76.5064 | 81.6961 | 79.8248
26.5 | 12.6122 | 36.2396 | 75.3283 | 81.3584 | 79.5762
26.6 | 11.6097 | 31.1773 | 74.0448 | 81.0327 | 79.3774
26.7 | 10.9164 | 26.7727 | 72.8292 | 80.6729 | 79.2007
26.8 | 10.3253 | 22.9951 | 71.5894 | 80.2111 | 78.8986
26.9 | 9.8256 | 20.1356 | 70.0657 | 79.7255 | 78.6948
27 | 9.4063 | 17.9007 | 68.0198 | 79.2542 | 78.3945
27.1 | 9.103 | 15.9782 | 64.998 | 78.887 | 78.1925
27.2 | 8.8333 | 14.543 | 61.5396 | 78.343 | 77.9801
27.3 | 8.6313 | 13.3539 | 57.2229 | 77.7731 | 77.6313
27.4 | 8.4293 | 12.5527 | 52.4473 | 77.2739 | 77.4505
27.5 | 8.2491 | 11.8458 | 47.1236 | 76.6899 | 77.2485
27.6 | 8.1076 | 11.2419 | 41.7493 | 76.1388 | 76.8898
27.7 | 7.8804 | 10.7697 | 36.6394 | 75.4317 | 76.7279
27.8 | 7.6892 | 10.3747 | 32.4613 | 74.7246 | 76.3683
27.9 | 7.5882 | 10.0533 | 28.6474 | 74.0175 | 76.1034
28 | 7.4775 | 9.7543 | 25.33 | 73.2974 | 75.862
28.1 | 7.3007 | 9.4637 | 22.6247 | 72.4893 | 75.6413
28.2 | 7.1989 | 9.1988 | 20.3876 | 71.6812 | 75.2979
28.3 | 7.0223 | 8.9968 | 18.6315 | 70.8731 | 74.9755
28.4 | 6.8974 | 8.7947 | 17.1634 | 70.0651 | 74.689
28.5 | 6.7936 | 8.5927 | 16.0495 | 69.2248 | 74.2873
28.6 | 6.6168 | 8.3954 | 15.2008 | 68.3676 | 73.987
28.7 | 6.4943 | 8.2887 | 14.3139 | 67.4328 | 73.7041
28.8 | 6.3907 | 8.0867 | 13.6899 | 66.5842 | 73.3813
28.9 | 6.2577 | 7.8923 | 13.1975 | 65.6355 | 73.0477
29 | 6.1484 | 7.7797 | 12.7543 | 64.7347 | 72.6757
29.1 | 6.0474 | 7.576 | 12.3159 | 63.8382 | 72.3659
29.2 | 5.9375 | 7.4708 | 12.0099 | 62.7933 | 71.9498
29.3 | 5.9 | 7.3529 | 11.6648 | 61.6283 | 71.5321
29.4 | 5.8018 | 7.2221 | 11.3745 | 60.1581 | 71.1553
29.5 | 5.7005 | 7.0295 | 11.1519 | 58.3535 | 70.7315
29.6 | 5.5995 | 6.9137 | 10.8418 | 56.1759 | 70.1836
29.7 | 5.4985 | 6.8127 | 10.6398 | 53.597 | 69.7796
29.8 | 5.3971 | 6.5971 | 10.4274 | 50.4111 | 69.3133
29.9 | 5.2792 | 6.4792 | 10.1916 | 46.7794 | 68.7394
30 | 5.1768 | 6.371 | 9.9766 | 42.8747 | 68.2343
30.1 | 5.1061 | 6.2801 | 9.7745 | 38.9026 | 67.661
30.2 | 5 | 6.1791 | 9.5725 | 35.1345 | 66.9891
30.3 | 4.9744 | 6.0781 | 9.4353 | 31.679 | 66.4192
30.4 | 4.8566 | 5.9541 | 9.2465 | 28.7453 | 65.883
30.5 | 4.8 | 5.7761 | 9.0378 | 26.3232 | 65.2325
30.6 | 4.7465 | 5.7 | 8.9179 | 24.0458 | 64.6254
30.7 | 4.6236 | 5.674 | 8.7337 | 22.3598 | 63.9184
30.8 | 4.6 | 5.573 | 8.5317 | 20.7601 | 63.2761
30.9 | 4.5 | 5.4673 | 8.4148 | 19.5663 | 62.5846
31 | 4.4621 | 5.3495 | 8.1994 | 18.5391 | 61.84
31.1 | 4.37 | 5.218 | 8.058 | 17.6899 | 61.1898
31.2 | 4.269 | 5.1228 | 7.9165 | 17.0279 | 60.4825
31.3 | 4.1776 | 5.0251 | 7.7626 | 16.4218 | 59.7004
31.4 | 4.1069 | 5 | 7.6526 | 15.8631 | 59.1014
31.5 | 4.1 | 4.923 | 7.5516 | 15.4065 | 58.3151
31.6 | 3.9934 | 4.8091 | 7.4012 | 15.0024 | 57.4635
31.7 | 3.9 | 4.7924 | 7.2496 | 14.5984 | 56.6541
31.8 | 3.8914 | 4.6914 | 7.1486 | 14.2458 | 55.6576
31.9 | 3.8 | 4.5916 | 7.0388 | 13.9427 | 54.457
32 | 3.7894 | 4.5033 | 6.9322 | 13.6931 | 52.9561
32.1 | 3.7 | 4.4476 | 6.8198 | 13.457 | 51.0499
32.2 | 3.7 | 4.3926 | 6.6874 | 13.2176 | 49.021
32.3 | 3.6174 | 4.3337 | 6.5864 | 12.9733 | 46.4036
32.4 | 3.5996 | 4.2664 | 6.4854 | 12.7135 | 43.4182
32.5 | 3.5 | 4.1878 | 6.3843 | 12.5115 | 40.3008
32.6 | 3.4976 | 4.1093 | 6.2667 | 12.3095 | 37.0321
32.7 | 3.4 | 4.0952 | 6.1 | 12.1075 | 34.1136
32.8 | 3.3961 | 3.9537 | 6.0782 | 11.9055 | 31.2364
32.9 | 3.3077 | 3.9 | 5.966 | 11.7034 | 28.6215
33 | 3.3 | 3.8364 | 5.8649 | 11.5014 | 26.6836
33.1 | 3.2068 | 3.8 | 5.758 | 11.3119 | 24.8804
33.2 | 3.2 | 3.7343 | 5.6281 | 11.1259 | 23.2897
33.3 | 3.2 | 3.7 | 5.6 | 10.9619 | 21.9956
33.4 | 3.1037 | 3.6323 | 5.5181 | 10.8218 | 20.9055
33.5 | 3.1 | 3.6 | 5.4171 | 10.6599 | 19.9925
33.6 | 3.052 | 3.5 | 5.3161 | 10.5178 | 19.1388
33.7 | 3 | 3.4865 | 5.2151 | 10.341 | 18.5327
33.8 | 3 | 3.4 | 5.1996 | 10.2163 | 17.9267
33.9 | 2.9489 | 3.3987 | 5.1 | 10.0259 | 17.3804
34 | 2.9 | 3.3 | 5.0139 | 9.9119 | 16.9455
34.1 | 2.9 | 3.3 | 4.8967 | 9.7218 | 16.5042
34.2 | 2.8587 | 3.2671 | 4.8 | 9.5198 | 16.2012
34.3 | 2.8 | 3.2 | 4.7724 | 9.4089 | 15.9321
34.4 | 2.8 | 3.1778 | 4.7 | 9.2926 | 15.7021
34.5 | 2.7652 | 3.1135 | 4.6355 | 9.1566 | 15.4192
34.6 | 2.7108 | 3.1 | 4.6 | 8.9801 | 15.1706
34.7 | 2.7 | 3.1 | 4.5 | 8.889 | 14.8668
34.8 | 2.7 | 3.0181 | 4.5 | 8.7653 | 14.6134
34.9 | 2.7 | 3 | 4.4 | 8.6366 | 14.4485
35 | 2.6 | 3 | 4.3025 | 8.5025 | 14.2465
35.1 | 2.6 | 2.9176 | 4.2611 | 8.4 | 14.0167
35.2 | 2.6 | 2.9 | 4.2 | 8.3498 | 13.7712
35.3 | 2.5983 | 2.8988 | 4.1489 | 8.1638 | 13.6583
35.4 | 2.5 | 2.8 | 4.1 | 8.1 | 13.4613
35.5 | 2.5 | 2.8 | 4.0538 | 8.011 | 13.3128
35.6 | 2.5 | 2.8 | 4 | 7.895 | 13.12
35.7 | 2.5 | 2.7719 | 3.9518 | 7.7804 | 12.918
35.8 | 2.476 | 2.713 | 3.9 | 7.6794 | 12.808
35.9 | 2.4 | 2.7 | 3.8415 | 7.5832 | 12.6139
36 | 2.4 | 2.6951 | 3.8 | 7.5046 | 12.4119
36.1 | 2.4 | 2.6362 | 3.8 | 7.4049 | 12.3049
36.2 | 2.4 | 2.6 | 3.7 | 7.3039 | 12.1079
36.3 | 2.4 | 2.6 | 3.66 | 7.1301 | 11.9059
36.4 | 2.3447 | 2.6 | 3.6 | 7.0549 | 11.8141
36.5 | 2.3 | 2.5152 | 3.5412 | 7.0005 | 11.7152
36.6 | 2.3 | 2.5 | 3.5 | 6.8832 | 11.5283
36.7 | 2.3 | 2.5 | 3.5 | 6.7792 | 11.3263
36.8 | 2.3 | 2.4835 | 3.4407 | 6.7085 | 11.2121
36.9 | 2.2682 | 2.4 | 3.4 | 6.6378 | 11.0222
37 | 2.2 | 2.4 | 3.4 | 6.5529 | 10.9101
37.1 | 2.2 | 2.4 | 3.3091 | 6.4519 | 10.8091
37.2 | 2.2 | 2.4 | 3.3 | 6.3755 | 10.7195
37.3 | 2.2 | 2.4 | 3.2071 | 6.3249 | 10.5426
37.4 | 2.2 | 2.3775 | 3.2 | 6.2523 | 10.4203
37.5 | 2.2 | 2.3 | 3.2 | 6.158 | 10.3193
37.6 | 2.2 | 2.3 | 3.1713 | 6.0806 | 10.2183
37.7 | 2.2 | 2.3 | 3.1 | 6.0301 | 10.1173
37.8 | 2.1163 | 2.3 | 3.1 | 5.9795 | 10.0163
37.9 | 2.1 | 2.2723 | 3.1 | 5.929 | 9.8305
38 | 2.1 | 2.2 | 3 | 5.8571 | 9.7142
38.1 | 2.1 | 2.2 | 3 | 5.7561 | 9.6132
38.2 | 2.1 | 2.2 | 2.957 | 5.6714 | 9.5122
38.3 | 2.1 | 2.2 | 2.9 | 5.6071 | 9.4111
38.4 | 2.1 | 2.1673 | 2.9 | 5.5551 | 9.2959
38.5 | 2.0091 | 2.1 | 2.9 | 5.5045 | 9.1948
38.6 | 2 | 2.1 | 2.8367 | 5.4142 | 9.0938
38.7 | 2 | 2.1 | 2.8 | 5.32 | 8.9928
38.8 | 2 | 2.1 | 2.8 | 5.2602 | 8.8918
38.9 | 2 | 2.1 | 2.8 | 5.2097 | 8.792
39 | 2 | 2.1 | 2.7469 | 5.156 | 8.7035
39.1 | 2 | 2.0744 | 2.7 | 5.1016 | 8.6
39.2 | 2 | 2 | 2.7 | 5.0428 | 8.5524
39.3 | 1.9815 | 2 | 2.7 | 4.9862 | 8.4509
39.4 | 1.9 | 2 | 2.7 | 4.9357 | 8.3572
39.5 | 1.9 | 2 | 2.6322 | 4.8827 | 8.2386
39.6 | 1.9 | 2 | 2.6 | 4.8238 | 8.1976
39.7 | 1.9 | 1.9723 | 2.6 | 4.8 | 8.0826
39.8 | 1.9 | 1.9 | 2.6 | 4.7388 | 7.9816
39.9 | 1.9 | 1.9 | 2.5235 | 4.7 | 7.8806
40 | 1.9 | 1.9 | 2.5 | 4.6368 | 7.8
40.1 | 1.9 | 1.9 | 2.5 | 4.6 | 7.7786
40.2 |  | 1.9 | 2.5 | 4.5347 | 7.6776
40.3 |  | 1.9 | 2.5 | 4.5 | 7.6
40.4 |  | 1.9 | 2.4041 | 4.4477 | 7.5756
40.5 |  | 1.9 | 2.4 | 4.4 | 7.4745
40.6 |  | 1.8269 | 2.4 | 4.3643 | 7.3691
40.7 |  | 1.8 | 2.4 | 4.3 | 7.3
40.8 |  | 1.8 | 2.3858 | 4.2834 | 7.2201
40.9 |  | 1.8 | 2.3 | 4.2 | 7.2
41 |  |  | 2.3 | 4.2 | 7.1123
41.1 |  |  | 2.3 | 4.1947 | 7.0113
41.2 |  |  | 2.3 | 4.1 | 6.9103
41.3 |  |  | 2.3 | 4.1 | 6.9
41.4 |  |  | 2.266 | 4.1 | 6.8083
41.5 |  |  | 2.2116 | 4.0102 | 6.7
41.6 |  |  | 2.2 | 4 | 6.7
41.7 |  |  | 2.2 | 3.9472 | 6.6337
41.8 |  |  | 2.2 | 3.9 | 6.6
41.9 |  |  | 2.2 | 3.9 | 6.5037
42 |  |  | 2.1573 | 3.8736 | 6.4879

## Open field Output Factors
Output factors
Open field
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30 | 40
| 3 | 0.894 | 0.906 | 0.913 | 0.921 | 0.926 | 0.931 | 0.933 | 0.936 | 0.937
| 4 | 0.907 | 0.921 | 0.929 | 0.94 | 0.946 | 0.952 | 0.956 | 0.959 | 0.96
| 5 | 0.914 | 0.93 | 0.94 | 0.952 | 0.961 | 0.968 | 0.971 | 0.976 | 0.977
| 7 | 0.924 | 0.944 | 0.956 | 0.971 | 0.981 | 0.99 | 0.996 | 1.001 | 1.003
| 10 | 0.934 | 0.955 | 0.969 | 0.987 | 1 | 1.012 | 1.019 | 1.027 | 1.029
| 15 | 0.94 | 0.963 | 0.979 | 0.999 | 1.016 | 1.031 | 1.042 | 1.052 | 1.055
| 20 | 0.944 | 0.968 | 0.984 | 1.006 | 1.024 | 1.042 | 1.055 | 1.067 | 1.071
| 30 | 0.948 | 0.973 | 0.99 | 1.014 | 1.035 | 1.056 | 1.071 | 1.087 | 1.093
| 40 | 0.951 | 0.976 | 0.994 | 1.018 | 1.04 | 1.063 | 1.079 | 1.096 | 1.102

## Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 33.4 | 39.3 | 48.8 | 57.4 | 60.8
0.1 | 36.75 | 42.4 | 52.1 | 60.4 | 63.9
0.2 | 41.95 | 47.3 | 56.6 | 64.7 | 67.9
0.3 | 51.7 | 56 | 64.15 | 71.65 | 74.6
0.4 | 60.06 | 64.8 | 71.7 | 77.9 | 80.4
0.5 | 67.975 | 72 | 78.1 | 83.2 | 85.4
0.6 | 74.1 | 77.6 | 82.925 | 87.1667 | 88.775
0.7 | 79.425 | 82.14 | 86.9 | 90.3 | 91.8333
0.8 | 83.775 | 85.975 | 89.825 | 92.9333 | 94
0.9 | 87.225 | 89.2 | 92.45 | 94.9 | 95.7333
1 | 89.9 | 91.7333 | 94.375 | 96.3 | 97.075
1.1 | 92.5 | 93.775 | 95.875 | 97.55 | 98.0333
1.2 | 94.3 | 95.45 | 97.02 | 98.5 | 98.75
1.3 | 95.9 | 96.775 | 98.1 | 98.95 | 99.25
1.4 | 97.05 | 97.75 | 98.9 | 99.4667 | 99.6333
1.5 | 98.1 | 98.6 | 99.3 | 99.725 | 99.9
1.6 | 98.8 | 99.1 | 99.6 | 99.975 | 100
1.7 | 99.3333 | 99.6 | 99.9 | 100 | 100
1.8 | 99.6667 | 99.84 | 100 | 100 | 100
1.9 | 99.9 | 100 | 100 | 99.9667 | 99.95
2 | 100 | 100 | 99.975 | 99.8 | 99.8
2.1 | 100 | 100 | 99.8 | 99.6 | 99.6
2.2 | 99.9 | 99.9 | 99.725 | 99.45 | 99.325
2.3 | 99.8 | 99.8 | 99.525 | 99.1 | 99.075
2.4 | 99.625 | 99.65 | 99.375 | 98.8333 | 98.875
2.5 | 99.25 | 99.3667 | 98.975 | 98.575 | 98.48
2.6 | 98.95 | 99 | 98.6333 | 98.25 | 98.15
2.7 | 98.6 | 98.725 | 98.35 | 97.975 | 97.8
2.8 | 98.3 | 98.4 | 98.05 | 97.66 | 97.6333
2.9 | 97.8667 | 97.95 | 97.7667 | 97.35 | 97.2333
3 | 97.4 | 97.65 | 97.3667 | 97 | 96.975
3.1 | 96.95 | 97.25 | 97.025 | 96.65 | 96.55
3.2 | 96.55 | 96.8667 | 96.675 | 96.3 | 96.25
3.3 | 96.15 | 96.55 | 96.3 | 95.96 | 95.9
3.4 | 95.725 | 96.12 | 95.94 | 95.55 | 95.4
3.5 | 95.3 | 95.75 | 95.66 | 95.275 | 95.125
3.6 | 94.7 | 95.325 | 95.2 | 95 | 94.85
3.7 | 94.3667 | 95.0333 | 94.9 | 94.6 | 94.5
3.8 | 93.9333 | 94.6 | 94.5 | 94.25 | 94.2
3.9 | 93.525 | 94.1333 | 94.15 | 93.9333 | 93.75
4 | 93 | 93.7667 | 93.75 | 93.5 | 93.4667
4.1 | 92.5 | 93.35 | 93.35 | 93.175 | 93.1
4.2 | 92 | 92.9333 | 92.9 | 92.86 | 92.7333
4.3 | 91.6 | 92.6 | 92.6 | 92.45 | 92.375
4.4 | 91.16 | 92.075 | 92.3 | 92.15 | 92.15
4.5 | 90.74 | 91.6 | 91.95 | 91.8 | 91.75
4.6 | 90.15 | 91.4 | 91.55 | 91.475 | 91.375
4.7 | 89.7 | 90.8333 | 91.25 | 91.05 | 91.08
4.8 | 89.3 | 90.55 | 90.8 | 90.675 | 90.7
4.9 | 88.8 | 90.15 | 90.5 | 90.35 | 90.4
5 | 88.4 | 89.8 | 90.1333 | 90.1 | 90.075
5.1 | 88 | 89.4 | 89.76 | 89.7 | 89.675
5.2 | 87.6 | 89 | 89.45 | 89.525 | 89.425
5.3 | 87.1 | 88.625 | 89.05 | 89 | 89.05
5.4 | 86.6 | 88.2667 | 88.625 | 88.7 | 88.75
5.5 | 86.15 | 87.85 | 88.4 | 88.35 | 88.4
5.6 | 85.65 | 87.5 | 87.9 | 87.975 | 88
5.7 | 85.25 | 87.1 | 87.62 | 87.5667 | 87.65
5.8 | 84.9 | 86.65 | 87.325 | 87.3 | 87.32
5.9 | 84.3 | 86.3 | 86.95 | 87 | 87.025
6 | 83.9 | 86 | 86.56 | 86.7 | 86.75
6.1 | 83.525 | 85.5 | 86.2 | 86.35 | 86.3667
6.2 | 83.025 | 85.1 | 85.9 | 85.925 | 85.95
6.3 | 82.65 | 84.7 | 85.5 | 85.65 | 85.65
6.4 | 82.3333 | 84.3 | 85.3 | 85.35 | 85.3
6.5 | 81.8 | 84.0333 | 84.8 | 85 | 85
6.6 | 81.4333 | 83.6333 | 84.525 | 84.6 | 84.75
6.7 | 80.9 | 83.3 | 84.125 | 84.3333 | 84.3667
6.8 | 80.4 | 82.85 | 83.825 | 83.9 | 83.95
6.9 | 80.05 | 82.5 | 83.45 | 83.65 | 83.625
7 | 79.55 | 82.0333 | 83.0667 | 83.3 | 83.34
7.1 | 79.2 | 81.775 | 82.64 | 82.925 | 82.95
7.2 | 78.7 | 81.375 | 82.375 | 82.5 | 82.6
7.3 | 78.45 | 81.025 | 82.025 | 82.3 | 82.35
7.4 | 78 | 80.6 | 81.68 | 82 | 82.05
7.5 | 77.65 | 80.25 | 81.45 | 81.65 | 81.65
7.6 | 77.15 | 79.9 | 81 | 81.3333 | 81.3667
7.7 | 76.8 | 79.5 | 80.7 | 81 | 81
7.8 | 76.25 | 79 | 80.32 | 80.7 | 80.7
7.9 | 75.8 | 78.675 | 79.975 | 80.42 | 80.4
8 | 75.4 | 78.275 | 79.55 | 79.9667 | 80.1
8.1 | 75.1 | 77.925 | 79.35 | 79.675 | 79.74
8.2 | 74.825 | 77.65 | 79.0667 | 79.275 | 79.525
8.3 | 74.35 | 77.3 | 78.6667 | 79 | 79.15
8.4 | 73.85 | 76.9667 | 78.35 | 78.8 | 78.85
8.5 | 73.48 | 76.7 | 78.025 | 78.4667 | 78.475
8.6 | 73.1 | 76.325 | 77.7 | 78.1667 | 78.275
8.7 | 72.75 | 75.8 | 77.275 | 77.8333 | 77.875
8.8 | 72.4 | 75.525 | 77 | 77.5667 | 77.6
8.9 | 71.86 | 75.225 | 76.7667 | 77.2 | 77.3
9 | 71.525 | 74.8667 | 76.4 | 76.8333 | 76.95
9.1 | 71.18 | 74.5 | 76.1 | 76.525 | 76.65
9.2 | 70.84 | 74.15 | 75.8 | 76.15 | 76.325
9.3 | 70.45 | 73.75 | 75.375 | 75.84 | 76.1
9.4 | 70.1 | 73.45 | 75.05 | 75.55 | 75.74
9.5 | 69.7 | 73.125 | 74.8 | 75.3 | 75.4333
9.6 | 69.3667 | 72.9 | 74.5 | 74.95 | 75.1
9.7 | 68.95 | 72.45 | 74.15 | 74.7 | 74.825
9.8 | 68.5667 | 72.1 | 73.9 | 74.35 | 74.55
9.9 | 68.1667 | 71.75 | 73.46 | 74.1 | 74.175
10 | 67.8 | 71.34 | 73.15 | 73.8667 | 73.925
10.1 | 67.4333 | 70.95 | 72.8667 | 73.5 | 73.6
10.2 | 67.15 | 70.725 | 72.7 | 73.25 | 73.34
10.3 | 66.7 | 70.425 | 72.2333 | 72.9 | 73
10.4 | 66.4 | 70.1 | 71.95 | 72.6 | 72.7
10.5 | 66.1 | 69.8 | 71.7 | 72.2 | 72.425
10.6 | 65.7 | 69.3667 | 71.3 | 71.875 | 72.1
10.7 | 65.4 | 68.9667 | 71.075 | 71.675 | 71.775
10.8 | 65 | 68.7 | 70.75 | 71.375 | 71.6
10.9 | 64.65 | 68.4 | 70.5 | 71.15 | 71.3333
11 | 64.2667 | 68.15 | 70.1333 | 70.825 | 71
11.1 | 63.975 | 67.7 | 69.8 | 70.55 | 70.7
11.2 | 63.625 | 67.375 | 69.6 | 70.25 | 70.375
11.3 | 63.26 | 67.075 | 69.2 | 69.9667 | 70.05
11.4 | 62.95 | 66.775 | 68.95 | 69.62 | 69.84
11.5 | 62.6 | 66.55 | 68.54 | 69.375 | 69.6
11.6 | 62.25 | 66.15 | 68.275 | 69.075 | 69.3
11.7 | 61.9 | 65.825 | 68.025 | 68.8333 | 68.9
11.8 | 61.575 | 65.4667 | 67.6667 | 68.5 | 68.6333
11.9 | 61.25 | 65.125 | 67.5 | 68.15 | 68.3333
12 | 60.925 | 64.9 | 67.15 | 67.9 | 68.1333
12.1 | 60.575 | 64.65 | 66.85 | 67.6 | 67.825
12.2 | 60.25 | 64.25 | 66.5667 | 67.4 | 67.55
12.3 | 59.9 | 64 | 66.225 | 67.075 | 67.2667
12.4 | 59.6 | 63.7 | 65.8667 | 66.775 | 66.95
12.5 | 59.3 | 63.3 | 65.6667 | 66.54 | 66.7
12.6 | 59 | 63 | 65.4 | 66.2 | 66.45
12.7 | 58.6667 | 62.6 | 65.1 | 66 | 66.1333
12.8 | 58.325 | 62.325 | 64.85 | 65.7 | 65.95
12.9 | 58.0333 | 62.125 | 64.55 | 65.42 | 65.6667
13 | 57.74 | 61.8 | 64.25 | 65.125 | 65.35
13.1 | 57.4667 | 61.56 | 63.9333 | 64.8 | 65
13.2 | 57 | 61.2 | 63.65 | 64.625 | 64.8
13.3 | 56.7 | 60.9 | 63.4 | 64.35 | 64.65
13.4 | 56.425 | 60.5333 | 63.075 | 64.025 | 64.225
13.5 | 56.175 | 60.3 | 62.9 | 63.8 | 63.975
13.6 | 55.9 | 60.1 | 62.5 | 63.55 | 63.78
13.7 | 55.64 | 59.6667 | 62.3 | 63.2667 | 63.55
13.8 | 55.3 | 59.4333 | 61.95 | 62.9667 | 63.3
13.9 | 55 | 59.1333 | 61.725 | 62.72 | 63
14 | 54.7 | 58.8333 | 61.45 | 62.375 | 62.7333
14.1 | 54.42 | 58.5 | 61.2 | 62.175 | 62.4333
14.2 | 54.04 | 58.2 | 60.95 | 61.9 | 62.175
14.3 | 53.775 | 58.05 | 60.65 | 61.7333 | 61.9
14.4 | 53.475 | 57.65 | 60.4 | 61.4 | 61.66
14.5 | 53.18 | 57.375 | 60.1 | 61.1 | 61.45
14.6 | 53 | 57.125 | 59.8333 | 60.9 | 61.2
14.7 | 52.7 | 56.875 | 59.6 | 60.62 | 60.85
14.8 | 52.4 | 56.6 | 59.3 | 60.425 | 60.62
14.9 | 52.125 | 56.3 | 59.1 | 60.1667 | 60.475
15 | 51.925 | 56.1 | 58.7333 | 59.9 | 60.1
15.1 | 51.55 | 55.8333 | 58.6 | 59.575 | 59.875
15.2 | 51.24 | 55.4333 | 58.3 | 59.375 | 59.7
15.3 | 51.08 | 55.175 | 58.1 | 59.1 | 59.4
15.4 | 50.8 | 54.975 | 57.7 | 58.8 | 59.0667
15.5 | 50.5 | 54.7 | 57.5 | 58.65 | 58.9
15.6 | 50.3 | 54.4667 | 57.3 | 58.4 | 58.7
15.7 | 49.9667 | 54.125 | 57.05 | 58.2 | 58.375
15.8 | 49.725 | 53.95 | 56.75 | 57.925 | 58.2
15.9 | 49.45 | 53.6333 | 56.5333 | 57.625 | 57.9
16 | 49.14 | 53.3333 | 56.2333 | 57.3333 | 57.675
16.1 | 48.9 | 53.1 | 55.975 | 57.15 | 57.4
16.2 | 48.65 | 52.775 | 55.725 | 56.9667 | 57.2333
16.3 | 48.42 | 52.5667 | 55.54 | 56.6667 | 56.9
16.4 | 48.0667 | 52.375 | 55.2 | 56.4333 | 56.7
16.5 | 47.825 | 52.025 | 55.05 | 56.2 | 56.4
16.6 | 47.64 | 51.8667 | 54.8 | 56.025 | 56.2
16.7 | 47.4 | 51.6 | 54.5 | 55.6333 | 55.9667
16.8 | 47.2 | 51.4 | 54.275 | 55.4333 | 55.8
16.9 | 46.85 | 51.1 | 54.075 | 55.3 | 55.6
17 | 46.6 | 50.825 | 53.875 | 54.9333 | 55.35
17.1 | 46.35 | 50.6 | 53.625 | 54.7667 | 55.075
17.2 | 46.125 | 50.3 | 53.3333 | 54.5667 | 54.8
17.3 | 45.85 | 50.0333 | 53.1 | 54.35 | 54.6
17.4 | 45.6 | 49.825 | 52.8 | 54.0667 | 54.3333
17.5 | 45.5 | 49.625 | 52.5 | 53.7667 | 54.1
17.6 | 45.15 | 49.35 | 52.3333 | 53.5 | 53.9
17.7 | 44.9 | 49.075 | 52.1 | 53.4 | 53.6
17.8 | 44.65 | 48.8 | 51.9 | 53.15 | 53.4
17.9 | 44.4 | 48.5 | 51.6667 | 52.875 | 53.125
18 | 44.1 | 48.4 | 51.35 | 52.675 | 52.9
18.1 | 43.975 | 48.125 | 51.1 | 52.425 | 52.725
18.2 | 43.8 | 47.9 | 50.925 | 52.15 | 52.525
18.3 | 43.525 | 47.7 | 50.6667 | 52 | 52.2
18.4 | 43.28 | 47.4667 | 50.4333 | 51.8 | 52.1
18.5 | 43.0333 | 47.3 | 50.275 | 51.6 | 51.85
18.6 | 42.85 | 47.0333 | 50 | 51.3333 | 51.6
18.7 | 42.6 | 46.85 | 49.7667 | 51.15 | 51.42
18.8 | 42.4 | 46.5 | 49.6 | 50.95 | 51.175
18.9 | 42.1333 | 46.3 | 49.45 | 50.65 | 51.025
19 | 41.9333 | 46.1 | 49.2 | 50.55 | 50.775
19.1 | 41.7 | 45.9 | 48.96 | 50.25 | 50.6
19.2 | 41.5 | 45.675 | 48.775 | 50 | 50.4
19.3 | 41.3 | 45.475 | 48.5 | 49.8 | 50.1
19.4 | 41.0333 | 45.2 | 48.275 | 49.625 | 49.95
19.5 | 40.9 | 44.9667 | 48.1 | 49.475 | 49.7
19.6 | 40.625 | 44.7 | 47.85 | 49.25 | 49.46
19.7 | 40.425 | 44.5 | 47.65 | 48.95 | 49.225
19.8 | 40.2 | 44.325 | 47.4 | 48.8 | 49.075
19.9 | 40.1 | 44.125 | 47.25 | 48.5 | 48.825
20 | 39.75 | 43.9 | 47 | 48.375 | 48.68
20.1 | 39.6 | 43.7 | 46.825 | 48.125 | 48.5
20.2 | 39.4 | 43.475 | 46.575 | 47.975 | 48.2
20.3 | 39.175 | 43.2 | 46.4 | 47.8 | 48.0333
20.4 | 38.925 | 43 | 46.2 | 47.55 | 47.875
20.5 | 38.775 | 42.86 | 46 | 47.35 | 47.625
20.6 | 38.58 | 42.7 | 45.8 | 47.1 | 47.475
20.7 | 38.35 | 42.45 | 45.55 | 46.9 | 47.225
20.8 | 38.1 | 42.2 | 45.325 | 46.76 | 47.025
20.9 | 37.95 | 41.95 | 45.125 | 46.525 | 46.825
21 | 37.775 | 41.825 | 44.88 | 46.3 | 46.6
21.1 | 37.5 | 41.6333 | 44.75 | 46.125 | 46.4
21.2 | 37.3 | 41.45 | 44.5667 | 46 | 46.2
21.3 | 37.1 | 41.15 | 44.3 | 45.7 | 46
21.4 | 36.875 | 41 | 44.1667 | 45.5 | 45.85
21.5 | 36.675 | 40.8 | 43.9 | 45.35 | 45.66
21.6 | 36.5 | 40.6 | 43.7667 | 45.12 | 45.5
21.7 | 36.4 | 40.4 | 43.575 | 44.925 | 45.3
21.8 | 36.25 | 40.175 | 43.375 | 44.75 | 45.0333
21.9 | 36.02 | 39.975 | 43.175 | 44.525 | 44.9
22 | 35.875 | 39.8 | 42.94 | 44.38 | 44.7
22.1 | 35.625 | 39.6 | 42.8 | 44.15 | 44.5333
22.2 | 35.48 | 39.35 | 42.55 | 43.9 | 44.325
22.3 | 35.25 | 39.2 | 42.3667 | 43.75 | 44.15
22.4 | 35.05 | 39 | 42.2 | 43.52 | 43.875
22.5 | 34.92 | 38.875 | 41.95 | 43.3 | 43.7
22.6 | 34.775 | 38.6 | 41.8 | 43.2 | 43.6
22.7 | 34.525 | 38.475 | 41.675 | 43.04 | 43.35
22.8 | 34.4 | 38.3 | 41.4 | 42.8 | 43.2
22.9 | 34.225 | 38.14 | 41.3 | 42.6333 | 42.9667
23 | 34 | 37.9 | 41.0333 | 42.5 | 42.8
23.1 | 33.78 | 37.8 | 40.85 | 42.3 | 42.6667
23.2 | 33.65 | 37.5 | 40.7 | 42.075 | 42.4667
23.3 | 33.42 | 37.35 | 40.5667 | 41.925 | 42.3333
23.4 | 33.3 | 37.225 | 40.35 | 41.7 | 42.125
23.5 | 33.125 | 37.075 | 40.125 | 41.575 | 41.95
23.6 | 33.025 | 36.85 | 39.9667 | 41.425 | 41.7
23.7 | 32.8 | 36.6333 | 39.8 | 41.2 | 41.55
23.8 | 32.7 | 36.45 | 39.6 | 41 | 41.4
23.9 | 32.5 | 36.25 | 39.4 | 40.85 | 41.225
24 | 32.3 | 36.1 | 39.3 | 40.6 | 41.025
24.1 | 32.1 | 35.9667 | 39.1 | 40.475 | 40.8
24.2 | 31.9 | 35.8 | 38.925 | 40.325 | 40.64
24.3 | 31.775 | 35.56 | 38.725 | 40.125 | 40.5
24.4 | 31.6 | 35.425 | 38.575 | 39.94 | 40.3
24.5 | 31.4 | 35.275 | 38.4 | 39.8 | 40.15
24.6 | 31.3 | 35.08 | 38.2 | 39.6 | 40
24.7 | 31.18 | 34.94 | 38.0667 | 39.5 | 39.8
24.8 | 31 | 34.8 | 37.9 | 39.25 | 39.7
24.9 | 30.8 | 34.6 | 37.7 | 39.1 | 39.475
25 | 30.7 | 34.4667 | 37.5333 | 38.9 | 39.3
25.1 | 30.575 | 34.22 | 37.3667 | 38.825 | 39.1
25.2 | 30.375 | 34.1 | 37.1667 | 38.58 | 39
25.3 | 30.2 | 33.9 | 37 | 38.4 | 38.85
25.4 | 30.08 | 33.725 | 36.9 | 38.25 | 38.6
25.5 | 29.85 | 33.5 | 36.7 | 38.1 | 38.5
25.6 | 29.7 | 33.45 | 36.55 | 37.95 | 38.3
25.7 | 29.575 | 33.25 | 36.3 | 37.75 | 38.175
25.8 | 29.4 | 33.1 | 36.2 | 37.6 | 38.04
25.9 | 29.3 | 32.9 | 36.075 | 37.45 | 37.8
26 | 29.125 | 32.8 | 35.9 | 37.3 | 37.675
26.1 | 29 | 32.625 | 35.725 | 37.1 | 37.5
26.2 | 28.9 | 32.475 | 35.6 | 37 | 37.35
26.3 | 28.7 | 32.3 | 35.4 | 36.85 | 37.15
26.4 | 28.6 | 32.2 | 35.25 | 36.6 | 37.02
26.5 | 28.45 | 31.95 | 35.1 | 36.5 | 36.85
26.6 | 28.2667 | 31.85 | 34.9 | 36.3333 | 36.7
26.7 | 28.125 | 31.7 | 34.7 | 36.2 | 36.6
26.8 | 28 | 31.5 | 34.6 | 36.0333 | 36.4
26.9 | 27.825 | 31.3667 | 34.4 | 35.875 | 36.2
27 | 27.7 | 31.225 | 34.3 | 35.7 | 36.1
27.1 | 27.6 | 31.1 | 34.1 | 35.5 | 35.9
27.2 | 27.45 | 31 | 34 | 35.325 | 35.8
27.3 | 27.3 | 30.7 | 33.8 | 35.2 | 35.6333
27.4 | 27.15 | 30.6 | 33.675 | 35.1 | 35.5
27.5 | 27 | 30.5 | 33.575 | 34.9 | 35.3
27.6 | 26.9 | 30.35 | 33.4 | 34.8 | 35.2
27.7 | 26.8 | 30.2 | 33.2 | 34.6 | 35.1
27.8 | 26.575 | 30.075 | 33.1 | 34.475 | 34.825
27.9 | 26.5 | 29.92 | 32.9 | 34.3 | 34.7
28 | 26.4 | 29.775 | 32.8 | 34.2 | 34.6
28.1 | 26.2 | 29.625 | 32.65 | 34.05 | 34.4
28.2 | 26 | 29.4667 | 32.5 | 33.9 | 34.3
28.3 | 25.9 | 29.3 | 32.3 | 33.7 | 34.1
28.4 | 25.8 | 29.2 | 32.175 | 33.6 | 34
28.5 | 25.625 | 29 | 32.025 | 33.425 | 33.8
28.6 | 25.5 | 28.9 | 31.9 | 33.275 | 33.6667
28.7 | 25.4 | 28.8 | 31.8 | 33.125 | 33.525
28.8 | 25.3 | 28.625 | 31.6 | 33 | 33.4
28.9 | 25.15 | 28.475 | 31.45 | 32.825 | 33.3
29 | 25 | 28.375 | 31.35 | 32.7 | 33.1
29.1 | 24.9 | 28.2 | 31.2 | 32.5333 | 33
29.2 | 24.7 | 28.125 | 31.06 | 32.5 | 32.8
29.3 | 24.6333 | 27.975 | 30.9 | 32.3 | 32.7
29.4 | 24.5 | 27.775 | 30.7 | 32.1 | 32.6
29.5 | 24.4 | 27.675 | 30.6 | 32.0333 | 32.4
29.6 | 24.22 | 27.6 | 30.45 | 31.875 | 32.3
29.7 | 24.125 | 27.45 | 30.4 | 31.725 | 32.1
29.8 | 24.025 | 27.3 | 30.2 | 31.6 | 32
29.9 | 23.9 | 27.2 | 30.1 | 31.5 | 31.85
30 | 23.8 | 27.05 | 29.925 | 31.3 | 31.75

## SSD80 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 35.2 | 43.3 | 56.8 | 68.3 | 72.2
0.1 | 38.8667 | 47.1 | 60.5 | 71.6 | 75.38
0.2 | 45.1333 | 52.96 | 65.9 | 76.26 | 79.7
0.3 | 54.5 | 61.56 | 73.32 | 82.14 | 84.98
0.4 | 63.68 | 70.1 | 79.6667 | 87.2 | 89.34
0.5 | 71.08 | 76.12 | 84.8 | 90.84 | 92.4
0.6 | 77.2 | 81.56 | 88.7 | 93.5571 | 94.88
0.7 | 81.9 | 85.76 | 91.7333 | 95.5333 | 96.62
0.8 | 86 | 89.0333 | 94.1 | 97.08 | 97.82
0.9 | 88.92 | 91.72 | 95.8667 | 98.2 | 98.76
1 | 91.58 | 93.84 | 97.12 | 98.96 | 99.36
1.1 | 93.8 | 95.475 | 98.18 | 99.46 | 99.74
1.2 | 95.58 | 96.9 | 98.96 | 99.82 | 99.95
1.3 | 96.9 | 97.9 | 99.44 | 100 | 100
1.4 | 97.975 | 98.7 | 99.82 | 100 | 100
1.5 | 98.74 | 99.3 | 100 | 100 | 100
1.6 | 99.34 | 99.7 | 100 | 99.9 | 99.8
1.7 | 99.74 | 99.9667 | 100 | 99.7 | 99.6
1.8 | 100 | 100 | 99.9 | 99.5 | 99.4
1.9 | 100 | 100 | 99.7 | 99.3 | 99.0667
2 | 100 | 100 | 99.5 | 99 | 98.78
2.1 | 99.9 | 99.8167 | 99.2 | 98.6 | 98.425
2.2 | 99.7 | 99.62 | 98.9 | 98.3 | 98.1
2.3 | 99.475 | 99.32 | 98.6 | 97.9 | 97.7
2.4 | 99.2 | 98.94 | 98.2 | 97.5 | 97.3
2.5 | 98.8 | 98.54 | 97.8 | 97.14 | 97.0167
2.6 | 98.42 | 98.24 | 97.35 | 96.74 | 96.5
2.7 | 98 | 97.84 | 97.1 | 96.34 | 96.16
2.8 | 97.5143 | 97.44 | 96.6333 | 95.94 | 95.82
2.9 | 97.05 | 96.96 | 96.2 | 95.54 | 95.34
3 | 96.52 | 96.54 | 95.84 | 95.14 | 95.04
3.1 | 96.08 | 96.1667 | 95.42 | 94.7 | 94.56
3.2 | 95.62 | 95.62 | 94.94 | 94.28 | 94.14
3.3 | 95.075 | 95.26 | 94.54 | 93.88 | 93.82
3.4 | 94.54 | 94.82 | 94.1 | 93.48 | 93.3667
3.5 | 94.1 | 94.36 | 93.7 | 93.11 | 92.98
3.6 | 93.56 | 93.86 | 93.26 | 92.75 | 92.58
3.7 | 93 | 93.3667 | 92.8667 | 92.28 | 92.22
3.8 | 92.5 | 93 | 92.42 | 91.9 | 91.78
3.9 | 92 | 92.4667 | 92 | 91.5111 | 91.4667
4 | 91.5 | 92.1 | 91.58 | 91.08 | 90.98
4.1 | 91.0444 | 91.6 | 91.15 | 90.7 | 90.68
4.2 | 90.46 | 91.2 | 90.78 | 90.3333 | 90.22
4.3 | 89.98 | 90.8 | 90.35 | 89.96 | 89.78
4.4 | 89.48 | 90.3 | 89.86 | 89.46 | 89.44
4.5 | 88.88 | 89.9 | 89.475 | 89.08 | 89.1
4.6 | 88.45 | 89.4 | 89.1 | 88.76 | 88.68
4.7 | 87.98 | 88.9 | 88.7 | 88.4 | 88.32
4.8 | 87.5 | 88.5 | 88.3 | 87.8667 | 87.9
4.9 | 87 | 88.12 | 87.9 | 87.5 | 87.55
5 | 86.46 | 87.7 | 87.5 | 87.2 | 87.1
5.1 | 86.05 | 87.2333 | 87.1 | 86.8 | 86.675
5.2 | 85.55 | 86.74 | 86.72 | 86.4 | 86.3333
5.3 | 85.02 | 86.34 | 86.32 | 86.02 | 85.9333
5.4 | 84.58 | 85.94 | 85.9 | 85.6 | 85.58
5.5 | 84.08 | 85.46 | 85.52 | 85.22 | 85.26
5.6 | 83.62 | 85.04 | 85.14 | 84.86 | 84.8667
5.7 | 83.02 | 84.64 | 84.66 | 84.48 | 84.46
5.8 | 82.62 | 84.14 | 84.275 | 84.02 | 84.1
5.9 | 82.08 | 83.74 | 83.96 | 83.62 | 83.76
6 | 81.64 | 83.24 | 83.5 | 83.32 | 83.2667
6.1 | 81.18 | 82.9 | 83.08 | 82.92 | 83
6.2 | 80.75 | 82.58 | 82.68 | 82.6 | 82.6
6.3 | 80.28 | 82.1 | 82.3091 | 82.22 | 82.2
6.4 | 79.82 | 81.575 | 81.99 | 81.74 | 81.8333
6.5 | 79.38 | 81.2 | 81.62 | 81.4333 | 81.44
6.6 | 78.88 | 80.76 | 81.22 | 81.05 | 81.04
6.7 | 78.32 | 80.35 | 80.82 | 80.6667 | 80.74
6.8 | 77.96 | 79.9 | 80.425 | 80.4 | 80.38
6.9 | 77.48 | 79.6 | 80.0286 | 79.98 | 80.04
7 | 77 | 79.1 | 79.6375 | 79.62 | 79.74
7.1 | 76.56 | 78.7 | 79.24 | 79.1667 | 79.28
7.2 | 76.16 | 78.325 | 78.8286 | 78.8 | 78.88
7.3 | 75.66 | 77.9 | 78.46 | 78.52 | 78.54
7.4 | 75.26 | 77.45 | 78.05 | 78.12 | 78.18
7.5 | 74.86 | 77.1 | 77.68 | 77.76 | 77.8
7.6 | 74.3 | 76.7 | 77.3 | 77.5 | 77.5
7.7 | 73.9 | 76.3 | 77 | 77.1 | 77.12
7.8 | 73.5 | 75.88 | 76.6 | 76.7143 | 76.8
7.9 | 73.1 | 75.4 | 76.2 | 76.35 | 76.45
8 | 72.7 | 75.02 | 75.8667 | 76 | 76.075
8.1 | 72.2333 | 74.62 | 75.5 | 75.6333 | 75.75
8.2 | 71.76 | 74.32 | 75.12 | 75.3286 | 75.3333
8.3 | 71.34 | 73.92 | 74.76 | 74.98 | 75.06
8.4 | 70.94 | 73.52 | 74.4 | 74.58 | 74.65
8.5 | 70.56 | 73.12 | 74.02 | 74.26 | 74.32
8.6 | 70.18 | 72.75 | 73.72 | 73.96 | 74.02
8.7 | 69.84 | 72.36 | 73.36 | 73.56 | 73.6333
8.8 | 69.38 | 71.98 | 73.06 | 73.26 | 73.3333
8.9 | 68.98 | 71.66 | 72.66 | 72.88 | 73.0333
9 | 68.58 | 71.2333 | 72.36 | 72.5333 | 72.72
9.1 | 68.15 | 70.86 | 71.96 | 72.26 | 72.36
9.2 | 67.72 | 70.56 | 71.66 | 71.88 | 71.98
9.3 | 67.32 | 70.16 | 71.26 | 71.55 | 71.7667
9.4 | 66.92 | 69.74 | 70.96 | 71.3 | 71.34
9.5 | 66.56 | 69.36 | 70.56 | 70.88 | 71.02
9.6 | 66.12 | 69.06 | 70.26 | 70.6 | 70.66
9.7 | 65.72 | 68.7 | 69.88 | 70.18 | 70.36
9.8 | 65.36 | 68.3 | 69.56 | 69.9 | 70.06
9.9 | 65.05 | 67.9667 | 69.18 | 69.6 | 69.68
10 | 64.62 | 67.6 | 68.9 | 69.2667 | 69.35
10.1 | 64.28 | 67.3 | 68.56 | 68.98 | 69.08
10.2 | 63.9571 | 66.8 | 68.2 | 68.6 | 68.7667
10.3 | 63.5667 | 66.5 | 67.9 | 68.32 | 68.425
10.4 | 63.2 | 66.1 | 67.5 | 68.02 | 68.075
10.5 | 62.8667 | 65.8 | 67.26 | 67.72 | 67.84
10.6 | 62.38 | 65.4 | 66.9 | 67.32 | 67.5333
10.7 | 61.9875 | 65.1333 | 66.6 | 67.02 | 67.175
10.8 | 61.6571 | 64.7333 | 66.2 | 66.72 | 66.95
10.9 | 61.38 | 64.36 | 65.9 | 66.42 | 66.6
11 | 60.98 | 64.02 | 65.6 | 66.12 | 66.3
11.1 | 60.6 | 63.72 | 65.3 | 65.82 | 65.94
11.2 | 60.2 | 63.3 | 64.9 | 65.52 | 65.7
11.3 | 59.9 | 63.02 | 64.6 | 65.16 | 65.36
11.4 | 59.5 | 62.72 | 64.3 | 64.8333 | 65.05
11.5 | 59.3 | 62.36 | 64.0333 | 64.62 | 64.75
11.6 | 58.9 | 62.02 | 63.7 | 64.2 | 64.45
11.7 | 58.54 | 61.66 | 63.35 | 63.96 | 64.16
11.8 | 58.24 | 61.36 | 63.05 | 63.66 | 63.86
11.9 | 57.84 | 61.06 | 62.8 | 63.38 | 63.56
12 | 57.54 | 60.68 | 62.5 | 63.08 | 63.26
12.1 | 57.14 | 60.36 | 62.1 | 62.76 | 62.94
12.2 | 56.92 | 60.06 | 61.82 | 62.46 | 62.66
12.3 | 56.52 | 59.68 | 61.52 | 62.125 | 62.3429
12.4 | 56.16 | 59.3667 | 61.16 | 61.86 | 62.0333
12.5 | 55.86 | 59.1 | 60.925 | 61.625 | 61.8
12.6 | 55.575 | 58.76 | 60.5667 | 61.25 | 61.54
12.7 | 55.28 | 58.38 | 60.2667 | 60.9667 | 61.22
12.8 | 54.86 | 58.1 | 60.02 | 60.675 | 60.9571
12.9 | 54.56 | 57.8 | 59.66 | 60.4 | 60.58
13 | 54.26 | 57.5 | 59.45 | 60.075 | 60.28
13.1 | 53.88 | 57.2 | 59.12 | 59.8 | 59.98
13.2 | 53.56 | 56.9 | 58.82 | 59.55 | 59.72
13.3 | 53.26 | 56.6 | 58.52 | 59.3 | 59.5273
13.4 | 52.98 | 56.3 | 58.22 | 59 | 59.2714
13.5 | 52.68 | 56.02 | 57.94 | 58.7 | 58.95
13.6 | 52.38 | 55.7 | 57.71 | 58.44 | 58.66
13.7 | 52.1 | 55.4 | 57.41 | 58.2 | 58.36
13.8 | 51.7667 | 55.1 | 57.11 | 57.9 | 58.06
13.9 | 51.48 | 54.8 | 56.825 | 57.64 | 57.84
14 | 51.2667 | 54.5 | 56.575 | 57.4 | 57.56
14.1 | 51 | 54.2 | 56.2667 | 57.1 | 57.31
14.2 | 50.6 | 53.9 | 55.9818 | 56.825 | 57.01
14.3 | 50.3 | 53.625 | 55.69 | 56.525 | 56.74
14.4 | 50 | 53.38 | 55.39 | 56.3 | 56.51
14.5 | 49.8 | 53.0667 | 55.16 | 56.025 | 56.21
14.6 | 49.5 | 52.78 | 54.89 | 55.725 | 55.9455
14.7 | 49.2 | 52.54 | 54.66 | 55.525 | 55.72
14.8 | 48.9 | 52.24 | 54.39 | 55.225 | 55.5125
14.9 | 48.62 | 52 | 54.09 | 55 | 55.22
15 | 48.34 | 51.6 | 53.84 | 54.7333 | 54.96
15.1 | 48.12 | 51.4 | 53.55 | 54.4333 | 54.7333
15.2 | 47.82 | 51.1 | 53.2857 | 54.2333 | 54.48
15.3 | 47.54 | 50.9 | 52.99 | 53.9333 | 54.1778
15.4 | 47.24 | 50.5 | 52.76 | 53.6333 | 53.9333
15.5 | 46.94 | 50.3 | 52.49 | 53.4333 | 53.6286
15.6 | 46.72 | 50.1 | 52.225 | 53.1 | 53.48
15.7 | 46.42 | 49.82 | 51.98 | 52.9 | 53.14
15.8 | 46.22 | 49.52 | 51.77 | 52.65 | 52.94
15.9 | 45.92 | 49.22 | 51.48 | 52.42 | 52.74
16 | 45.54 | 48.92 | 51.27 | 52.14 | 52.44
16.1 | 45.3333 | 48.72 | 50.98 | 51.9333 | 52.14
16.2 | 45.125 | 48.4667 | 50.77 | 51.72 | 51.94
16.3 | 44.92 | 48.225 | 50.47 | 51.44 | 51.72
16.4 | 44.64 | 47.94 | 50.18 | 51.24 | 51.44
16.5 | 44.3667 | 47.725 | 49.97 | 50.98 | 51.24
16.6 | 44.1667 | 47.4667 | 49.68 | 50.6833 | 51.04
16.7 | 43.8667 | 47.1667 | 49.47 | 50.5 | 50.75
16.8 | 43.55 | 46.94 | 49.18 | 50.2143 | 50.55
16.9 | 43.375 | 46.68 | 48.97 | 50 | 50.24
17 | 43.15 | 46.38 | 48.6778 | 49.75 | 50.1167
17.1 | 42.88 | 46.16 | 48.45 | 49.55 | 49.8
17.2 | 42.625 | 45.88 | 48.2 | 49.24 | 49.575
17.3 | 42.38 | 45.65 | 48 | 49.04 | 49.36
17.4 | 42.0833 | 45.375 | 47.7 | 48.8 | 49.1
17.5 | 41.9 | 45.2 | 47.5 | 48.575 | 48.84
17.6 | 41.6 | 44.86 | 47.2 | 48.32 | 48.6333
17.7 | 41.44 | 44.68 | 47.0167 | 48.06 | 48.3333
17.8 | 41.14 | 44.48 | 46.74 | 47.9 | 48.1
17.9 | 40.88 | 44.18 | 46.52 | 47.64 | 47.86
18 | 40.65 | 43.96 | 46.26 | 47.42 | 47.7333
18.1 | 40.44 | 43.68 | 46.06 | 47.18 | 47.4333
18.2 | 40.24 | 43.48 | 45.8667 | 46.92 | 47.2333
18.3 | 40.05 | 43.28 | 45.65 | 46.7429 | 47.05
18.4 | 39.8 | 43.0667 | 45.4 | 46.56 | 46.86
18.5 | 39.5 | 42.7833 | 45.18 | 46.28 | 46.66
18.6 | 39.4 | 42.6 | 44.98 | 46.1 | 46.36
18.7 | 39.1 | 42.4 | 44.78 | 45.88 | 46.16
18.8 | 38.9 | 42.2 | 44.5 | 45.68 | 45.96
18.9 | 38.7167 | 41.8833 | 44.38 | 45.4286 | 45.76
19 | 38.5167 | 41.7 | 44.08 | 45.2143 | 45.46
19.1 | 38.2 | 41.5 | 43.88 | 44.9857 | 45.32
19.2 | 38.02 | 41.2 | 43.66 | 44.8 | 45.06
19.3 | 37.82 | 41 | 43.38 | 44.58 | 44.8667
19.4 | 37.62 | 40.8 | 43.18 | 44.4 | 44.6667
19.5 | 37.42 | 40.6 | 42.98 | 44.1833 | 44.4667
19.6 | 37.2333 | 40.4 | 42.78 | 43.9 | 44.3143
19.7 | 37.02 | 40.2 | 42.58 | 43.7 | 44.02
19.8 | 36.7571 | 40 | 42.4 | 43.5 | 43.82
19.9 | 36.6143 | 39.7 | 42.14 | 43.3 | 43.62
20 | 36.42 | 39.5 | 41.94 | 43.1 | 43.42
20.1 | 36.2 | 39.3 | 41.74 | 42.9 | 43.22
20.2 | 36.04 | 39.1 | 41.55 | 42.7 | 43.04
20.3 | 35.84 | 38.9167 | 41.36 | 42.56 | 42.8333
20.4 | 35.56 | 38.7 | 41.16 | 42.36 | 42.64
20.5 | 35.42 | 38.5 | 40.96 | 42.12 | 42.44
20.6 | 35.16 | 38.3 | 40.76 | 41.86 | 42.24
20.7 | 34.94 | 38.1 | 40.56 | 41.66 | 42.04
20.8 | 34.8273 | 37.92 | 40.36 | 41.46 | 41.85
20.9 | 34.6444 | 37.72 | 40.16 | 41.3 | 41.65
21 | 34.3818 | 37.52 | 39.96 | 41.16 | 41.45
21.1 | 34.23 | 37.32 | 39.76 | 40.92 | 41.26
21.2 | 34.025 | 37.12 | 39.56 | 40.76 | 41.06
21.3 | 33.82 | 36.92 | 39.36 | 40.56 | 40.86
21.4 | 33.64 | 36.72 | 39.16 | 40.36 | 40.66
21.5 | 33.42 | 36.52 | 38.98 | 40.16 | 40.46
21.6 | 33.24 | 36.32 | 38.78 | 39.98 | 40.26
21.7 | 33.1 | 36.12 | 38.58 | 39.78 | 40.06
21.8 | 32.9333 | 35.92 | 38.38 | 39.58 | 39.9
21.9 | 32.72 | 35.75 | 38.18 | 39.38 | 39.76
22 | 32.5 | 35.54 | 38.0333 | 39.18 | 39.5667
22.1 | 32.34 | 35.34 | 37.8 | 38.98 | 39.36
22.2 | 32.14 | 35.2 | 37.66 | 38.78 | 39.125
22.3 | 32 | 34.98 | 37.48 | 38.58 | 39
22.4 | 31.825 | 34.8333 | 37.1833 | 38.3833 | 38.8
22.5 | 31.64 | 34.64 | 37 | 38.28 | 38.6
22.6 | 31.44 | 34.44 | 36.9 | 38.1 | 38.375
22.7 | 31.3 | 34.24 | 36.7 | 37.9 | 38.225
22.8 | 31.14 | 34.04 | 36.5286 | 37.7 | 38.0571
22.9 | 31 | 33.9 | 36.34 | 37.6 | 37.88
23 | 30.74 | 33.7 | 36.15 | 37.4 | 37.7
23.1 | 30.64 | 33.5 | 35.9444 | 37.2 | 37.4833
23.2 | 30.35 | 33.34 | 35.76 | 37 | 37.3833
23.3 | 30.25 | 33.14 | 35.6222 | 36.85 | 37.1833
23.4 | 30.06 | 33 | 35.44 | 36.7 | 37
23.5 | 29.96 | 32.8667 | 35.32 | 36.5 | 36.84
23.6 | 29.76 | 32.7 | 35.14 | 36.3 | 36.64
23.7 | 29.66 | 32.5 | 34.96 | 36.1667 | 36.4625
23.8 | 29.46 | 32.3667 | 34.8 | 36 | 36.3286
23.9 | 29.36 | 32.1667 | 34.6 | 35.8167 | 36.14
24 | 29.16 | 32 | 34.4 | 35.62 | 35.96
24.1 | 29 | 31.9 | 34.2833 | 35.5 | 35.7714
24.2 | 28.82 | 31.7 | 34.1 | 35.32 | 35.7
24.3 | 28.6667 | 31.5 | 33.9 | 35.12 | 35.5
24.4 | 28.52 | 31.325 | 33.8 | 35 | 35.3167
24.5 | 28.32 | 31.2 | 33.6 | 34.825 | 35.1
24.6 | 28.22 | 31.04 | 33.4 | 34.62 | 35
24.7 | 28.02 | 30.8 | 33.2 | 34.4 | 34.8429
24.8 | 27.88 | 30.7 | 33.1 | 34.3 | 34.66
24.9 | 27.78 | 30.525 | 32.9 | 34.18 | 34.46
25 | 27.58 | 30.4 | 32.8 | 33.98 | 34.2667
25.1 | 27.4 | 30.1333 | 32.6 | 33.78 | 34.1667
25.2 | 27.28 | 30.05 | 32.4 | 33.68 | 33.9667
25.3 | 27.18 | 29.9 | 32.3 | 33.48 | 33.86
25.4 | 26.98 | 29.75 | 32.1 | 33.38 | 33.6667
25.5 | 26.8 | 29.6 | 32.0167 | 33.18 | 33.5
25.6 | 26.68 | 29.44 | 31.8 | 32.98 | 33.375
25.7 | 26.6 | 29.3 | 31.64 | 32.9 | 33.2
25.8 | 26.4 | 29.1 | 31.52 | 32.7 | 33.0667
25.9 | 26.3 | 28.94 | 31.32 | 32.58 | 32.98
26 | 26.1 | 28.84 | 31.22 | 32.4 | 32.7667
26.1 | 26 | 28.65 | 31.02 | 32.3 | 32.575
26.2 | 25.8 | 28.56 | 30.92 | 32.15 | 32.48
26.3 | 25.7 | 28.4 | 30.72 | 31.96 | 32.3
26.4 | 25.6 | 28.26 | 30.62 | 31.8 | 32.18
26.5 | 25.5 | 28.06 | 30.42 | 31.72 | 31.98
26.6 | 25.3 | 27.96 | 30.3 | 31.52 | 31.88
26.7 | 25.1 | 27.8 | 30.12 | 31.4 | 31.66
26.8 | 25 | 27.66 | 30.02 | 31.22 | 31.58
26.9 | 24.9 | 27.5 | 29.84 | 31.02 | 31.4
27 | 24.8 | 27.36 | 29.64 | 30.9 | 31.28
27.1 | 24.6333 | 27.26 | 29.54 | 30.8 | 31.1
27.2 | 24.5 | 27.08 | 29.3625 | 30.6 | 30.98
27.3 | 24.34 | 26.9714 | 29.25 | 30.4714 | 30.8
27.4 | 24.24 | 26.8167 | 29.1167 | 30.3286 | 30.7
27.5 | 24.0833 | 26.625 | 29 | 30.2 | 30.55
27.6 | 24 | 26.52 | 28.9 | 30 | 30.36
27.7 | 23.8 | 26.4 | 28.7111 | 29.9 | 30.25
27.8 | 23.7 | 26.24 | 28.5167 | 29.7545 | 30.16
27.9 | 23.6 | 26.14 | 28.42 | 29.6556 | 29.95
28 | 23.4 | 25.94 | 28.3 | 29.5167 | 29.8
28.1 | 23.3 | 25.84 | 28.12 | 29.32 | 29.6714
28.2 | 23.2 | 25.7 | 28.02 | 29.21 | 29.5375
28.3 | 23.1 | 25.56 | 27.82 | 29.11 | 29.36
28.4 | 23 | 25.46 | 27.75 | 28.92 | 29.26
28.5 | 22.8167 | 25.3 | 27.6 | 28.72 | 29.16
28.6 | 22.7 | 25.2 | 27.48 | 28.61 | 29
28.7 | 22.6 | 25.06 | 27.38 | 28.51 | 28.86
28.8 | 22.5 | 24.9 | 27.2 | 28.41 | 28.7
28.9 | 22.32 | 24.8 | 27.08 | 28.22 | 28.56
29 | 22.22 | 24.7 | 26.9 | 28.12 | 28.46
29.1 | 22.1 | 24.56 | 26.8 | 27.92 | 28.26
29.2 | 22 | 24.4 | 26.6667 | 27.82 | 28.18
29.3 | 21.82 | 24.26 | 26.5 | 27.74 | 28.08
29.4 | 21.72 | 24.2 | 26.5 | 27.6 | 27.9
29.5 | 21.6 | 24.06 | 26.3 | 27.44 | 27.74
29.6 | 21.5 | 23.9 | 26.1 | 27.34 | 27.64
29.7 | 21.4 | 23.8 | 26 | 27.2 | 27.5
29.8 | 21.24 | 23.7 | 25.9 | 27.1 | 27.4
29.9 | 21.14 | 23.56 | 25.8 | 26.95 | 27.3
30 | 21.04 | 23.4 | 25.65 | 26.8 | 27.1667

## SSD120 Wedge 15 Depth Doses
Percentage Depth Doses
Wedge 15
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 35.7 | 41.3 | 49.4 | 56.6 | 59.7
0.1 | 40.0333 | 45.425 | 53.7 | 60.7 | 63.625
0.2 | 48.15 | 53.2 | 60.425 | 66.92 | 69.4
0.3 | 57.15 | 61.575 | 68.15 | 74.1 | 75.65
0.4 | 65.25 | 69.2 | 74.8 | 79.7 | 81.3
0.5 | 72.3333 | 75.45 | 80.2 | 84.05 | 85.7
0.6 | 77.675 | 80.45 | 84.525 | 87.625 | 89
0.7 | 82.1 | 84.275 | 87.875 | 90.55 | 91.8333
0.8 | 85.6571 | 87.8 | 90.56 | 92.925 | 93.9
0.9 | 88.7778 | 90.4 | 92.7 | 94.75 | 95.5
1 | 91.25 | 92.68 | 94.6 | 96.2 | 96.8
1.1 | 93.4 | 94.4 | 96 | 97.2667 | 97.64
1.2 | 95.025 | 95.95 | 97.15 | 98.2 | 98.475
1.3 | 96.425 | 96.98 | 98 | 98.8333 | 99.05
1.4 | 97.5 | 98 | 98.7 | 99.4 | 99.525
1.5 | 98.3 | 98.6 | 99.2 | 99.6 | 99.775
1.6 | 99 | 99.2 | 99.65 | 99.8667 | 99.9
1.7 | 99.44 | 99.5 | 99.85 | 100 | 100
1.8 | 99.7 | 99.85 | 100 | 100 | 100
1.9 | 99.94 | 99.9 | 100 | 100 | 100
2 | 100 | 100 | 100 | 99.9 | 99.9
2.1 | 100 | 100 | 99.95 | 99.7 | 99.7
2.2 | 99.925 | 99.975 | 99.7667 | 99.6 | 99.55
2.3 | 99.8 | 99.825 | 99.6 | 99.4 | 99.35
2.4 | 99.55 | 99.575 | 99.4333 | 99.14 | 99.1
2.5 | 99.2 | 99.26 | 99.075 | 98.95 | 98.8
2.6 | 99 | 99.125 | 98.875 | 98.725 | 98.6
2.7 | 98.7333 | 98.85 | 98.625 | 98.36 | 98.2333
2.8 | 98.3333 | 98.525 | 98.2 | 98.025 | 98
2.9 | 98.05 | 98.1 | 98.05 | 97.82 | 97.8
3 | 97.6 | 97.7667 | 97.8 | 97.525 | 97.4
3.1 | 97.1857 | 97.5 | 97.3667 | 97.175 | 97
3.2 | 96.825 | 97.05 | 97 | 96.925 | 96.7667
3.3 | 96.4571 | 96.625 | 96.7 | 96.575 | 96.4
3.4 | 95.96 | 96.45 | 96.35 | 96.225 | 96.1
3.5 | 95.4625 | 96.05 | 96 | 96 | 95.875
3.6 | 95.0875 | 95.625 | 95.65 | 95.65 | 95.6
3.7 | 94.7125 | 95.2857 | 95.3 | 95.3 | 95.16
3.8 | 94.2889 | 94.9 | 94.975 | 94.95 | 94.875
3.9 | 93.9 | 94.48 | 94.68 | 94.7 | 94.625
4 | 93.5 | 94.2 | 94.34 | 94.2 | 94.25
4.1 | 93.1 | 93.85 | 94.075 | 94 | 93.95
4.2 | 92.7 | 93.475 | 93.725 | 93.75 | 93.725
4.3 | 92.24 | 93.1 | 93.275 | 93.4 | 93.3667
4.4 | 91.8143 | 92.8 | 93.025 | 93.14 | 93
4.5 | 91.1857 | 92.35 | 92.65 | 92.775 | 92.6
4.6 | 90.925 | 91.975 | 92.325 | 92.45 | 92.3333
4.7 | 90.46 | 91.5 | 91.975 | 92.05 | 92.0667
4.8 | 90.12 | 91.25 | 91.725 | 91.75 | 91.7667
4.9 | 89.65 | 90.95 | 91.36 | 91.3 | 91.5
5 | 89.15 | 90.55 | 91.0333 | 91.1333 | 91.2
5.1 | 88.7 | 90.175 | 90.75 | 90.825 | 90.8333
5.2 | 88.38 | 89.825 | 90.35 | 90.56 | 90.45
5.3 | 87.9 | 89.5 | 90.1 | 90.325 | 90.225
5.4 | 87.6 | 89.15 | 89.78 | 89.9 | 90
5.5 | 87.2 | 88.84 | 89.4 | 89.525 | 89.65
5.6 | 86.7 | 88.375 | 89.1 | 89.2 | 89.3333
5.7 | 86.3 | 88.08 | 88.75 | 89 | 88.975
5.8 | 85.92 | 87.8 | 88.5 | 88.6 | 88.65
5.9 | 85.375 | 87.35 | 88.12 | 88.3 | 88.325
6 | 84.975 | 87 | 87.8 | 88.05 | 87.9667
6.1 | 84.725 | 86.7 | 87.4857 | 87.7667 | 87.7
6.2 | 84.175 | 86.3 | 87.2286 | 87.4 | 87.4333
6.3 | 83.825 | 86 | 86.925 | 87.05 | 87.125
6.4 | 83.425 | 85.65 | 86.5333 | 86.825 | 86.8
6.5 | 83.04 | 85.4 | 86.25 | 86.475 | 86.5333
6.6 | 82.65 | 84.95 | 85.8667 | 86.225 | 86.2667
6.7 | 82.325 | 84.6 | 85.575 | 85.875 | 86
6.8 | 81.9 | 84.3 | 85.325 | 85.625 | 85.7
6.9 | 81.3667 | 83.8 | 84.95 | 85.3167 | 85.35
7 | 81.1 | 83.55 | 84.575 | 84.9625 | 85
7.1 | 80.7 | 83.325 | 84.35 | 84.625 | 84.68
7.2 | 80.3 | 82.875 | 84 | 84.3125 | 84.45
7.3 | 79.9 | 82.525 | 83.65 | 83.9375 | 84.2
7.4 | 79.5 | 82.1667 | 83.4 | 83.8 | 83.86
7.5 | 79 | 81.78 | 83.1 | 83.4857 | 83.5
7.6 | 78.7 | 81.4 | 82.7 | 83.0889 | 83.275
7.7 | 78.4 | 81.15 | 82.45 | 82.85 | 82.85
7.8 | 77.9 | 80.75 | 82.1 | 82.6 | 82.575
7.9 | 77.65 | 80.45 | 81.8333 | 82.3 | 82.325
8 | 77.1 | 80.125 | 81.5 | 81.8 | 82.075
8.1 | 76.85 | 79.8 | 81.1667 | 81.62 | 81.725
8.2 | 76.45 | 79.525 | 80.8 | 81.34 | 81.475
8.3 | 76.075 | 79.1 | 80.55 | 81.1 | 81.15
8.4 | 75.725 | 78.75 | 80.3 | 80.7333 | 80.8
8.5 | 75.375 | 78.3 | 79.9667 | 80.475 | 80.5167
8.6 | 74.95 | 78.12 | 79.6 | 80.1625 | 80.3
8.7 | 74.5 | 77.75 | 79.32 | 79.9333 | 80.0286
8.8 | 74.2 | 77.5 | 79.05 | 79.5333 | 79.6
8.9 | 73.8333 | 77.1 | 78.725 | 79.2714 | 79.35
9 | 73.4 | 76.7333 | 78.4 | 79 | 79.1
9.1 | 73.1 | 76.4 | 78.1 | 78.7 | 78.85
9.2 | 72.75 | 76.1 | 77.8 | 78.4 | 78.5
9.3 | 72.2 | 75.76 | 77.5 | 78.0778 | 78.2
9.4 | 71.8667 | 75.425 | 77.2 | 77.8143 | 77.925
9.5 | 71.62 | 75.15 | 76.95 | 77.5 | 77.575
9.6 | 71.32 | 74.825 | 76.5 | 77.25 | 77.325
9.7 | 70.85 | 74.475 | 76.3 | 76.9 | 77.075
9.8 | 70.4625 | 74.15 | 76 | 76.7 | 76.65
9.9 | 70.0875 | 73.875 | 75.75 | 76.35 | 76.375
10 | 69.7125 | 73.525 | 75.4 | 76.0333 | 76.225
10.1 | 69.5125 | 73.175 | 75.05 | 75.82 | 75.9
10.2 | 69.12 | 72.925 | 74.8 | 75.525 | 75.7
10.3 | 68.825 | 72.575 | 74.55 | 75.275 | 75.3
10.4 | 68.4 | 72.225 | 74.1 | 75.025 | 75.1333
10.5 | 68.15 | 71.95 | 73.96 | 74.68 | 74.78
10.6 | 67.8 | 71.6333 | 73.625 | 74.375 | 74.55
10.7 | 67.45 | 71.3 | 73.375 | 74.08 | 74.18
10.8 | 67.1 | 71.05 | 73.05 | 73.8 | 74
10.9 | 66.75 | 70.7 | 72.775 | 73.5 | 73.6
11 | 66.4 | 70.35 | 72.625 | 73.2 | 73.35
11.1 | 66.05 | 70.0667 | 72.175 | 72.95 | 73.1
11.2 | 65.7333 | 69.7333 | 71.925 | 72.72 | 72.85
11.3 | 65.425 | 69.375 | 71.65 | 72.4 | 72.6
11.4 | 65.15 | 69.15 | 71.325 | 72.1 | 72.3667
11.5 | 64.825 | 68.78 | 71.075 | 71.85 | 72
11.6 | 64.4333 | 68.5 | 70.7333 | 71.5667 | 71.775
11.7 | 64.15 | 68.2333 | 70.5 | 71.42 | 71.4667
11.8 | 63.7 | 67.85 | 70.1667 | 71.05 | 71.2
11.9 | 63.5 | 67.6 | 69.9 | 70.8 | 70.9333
12 | 63.2 | 67.35 | 69.65 | 70.55 | 70.7
12.1 | 62.8 | 67 | 69.5 | 70.3 | 70.45
12.2 | 62.5 | 66.66 | 69.15 | 70 | 70.3
12.3 | 62.3333 | 66.5 | 68.86 | 69.7 | 69.9
12.4 | 61.9667 | 66.25 | 68.55 | 69.45 | 69.625
12.5 | 61.6 | 65.925 | 68.3 | 69.2 | 69.375
12.6 | 61.275 | 65.55 | 68 | 68.85 | 69.125
12.7 | 61.025 | 65.3 | 67.7 | 68.7 | 68.8
12.8 | 60.575 | 64.95 | 67.45 | 68.5 | 68.625
12.9 | 60.4 | 64.625 | 67.2 | 68.15 | 68.3
13 | 60.0667 | 64.375 | 66.9333 | 67.92 | 68.075
13.1 | 59.775 | 64.125 | 66.7 | 67.65 | 67.8
13.2 | 59.4667 | 63.875 | 66.45 | 67.4 | 67.575
13.3 | 59.1667 | 63.55 | 66.2 | 67.175 | 67.275
13.4 | 58.8 | 63.275 | 65.975 | 66.8 | 67.0333
13.5 | 58.525 | 62.975 | 65.725 | 66.56 | 66.76
13.6 | 58.3 | 62.6667 | 65.35 | 66.325 | 66.4
13.7 | 57.9 | 62.36 | 65.025 | 66.075 | 66.2
13.8 | 57.7 | 62.05 | 64.85 | 65.84 | 65.95
13.9 | 57.3 | 61.9 | 64.525 | 65.6 | 65.74
14 | 57.05 | 61.6 | 64.25 | 65.3 | 65.4
14.1 | 56.8 | 61.2 | 64.025 | 65 | 65.2
14.2 | 56.4 | 60.95 | 63.675 | 64.8 | 65
14.3 | 56.2667 | 60.6 | 63.425 | 64.58 | 64.7
14.4 | 55.9 | 60.45 | 63.175 | 64.3 | 64.45
14.5 | 55.6333 | 60.1 | 63 | 64.0667 | 64.1
14.6 | 55.45 | 59.85 | 62.7 | 63.8 | 63.95
14.7 | 55.075 | 59.5 | 62.4 | 63.4667 | 63.7
14.8 | 54.725 | 59.25 | 62.25 | 63.2 | 63.45
14.9 | 54.475 | 59 | 62 | 63 | 63.2
15 | 54.325 | 58.75 | 61.75 | 62.8 | 62.925
15.1 | 54 | 58.5 | 61.5 | 62.625 | 62.775
15.2 | 53.68 | 58.2667 | 61.25 | 62.35 | 62.475
15.3 | 53.4333 | 58 | 61 | 62.1 | 62.3
15.4 | 53.175 | 57.75 | 60.65 | 61.92 | 62.1
15.5 | 52.9 | 57.525 | 60.5 | 61.62 | 61.8
15.6 | 52.6857 | 57.175 | 60.25 | 61.325 | 61.5667
15.7 | 52.45 | 57.025 | 59.9 | 61.125 | 61.325
15.8 | 52.175 | 56.675 | 59.65 | 60.9 | 61.075
15.9 | 51.9143 | 56.525 | 59.4 | 60.675 | 60.825
16 | 51.6625 | 56.25 | 59.2667 | 60.425 | 60.575
16.1 | 51.325 | 55.975 | 59 | 60.1 | 60.425
16.2 | 51.075 | 55.625 | 58.7 | 59.925 | 60.175
16.3 | 50.825 | 55.4375 | 58.525 | 59.775 | 59.9
16.4 | 50.575 | 55.18 | 58.3 | 59.425 | 59.7
16.5 | 50.34 | 54.95 | 58.02 | 59.175 | 59.425
16.6 | 50.1 | 54.65 | 57.725 | 59.025 | 59.2
16.7 | 49.8 | 54.4 | 57.575 | 58.775 | 59
16.8 | 49.5 | 54.15 | 57.325 | 58.525 | 58.8
16.9 | 49.35 | 53.9 | 57.075 | 58.275 | 58.5
17 | 49.1 | 53.6667 | 56.825 | 58.025 | 58.3
17.1 | 48.86 | 53.45 | 56.575 | 57.775 | 58.05
17.2 | 48.625 | 53.2 | 56.325 | 57.54 | 57.8
17.3 | 48.375 | 53 | 56.175 | 57.375 | 57.55
17.4 | 48.125 | 52.75 | 55.85 | 57.14 | 57.3
17.5 | 47.875 | 52.5 | 55.6 | 56.9 | 57.15
17.6 | 47.625 | 52.25 | 55.4 | 56.75 | 56.9
17.7 | 47.375 | 52 | 55.2 | 56.4 | 56.6333
17.8 | 47.15 | 51.775 | 55 | 56.24 | 56.5
17.9 | 47 | 51.62 | 54.8 | 56 | 56.25
18 | 46.75 | 51.275 | 54.55 | 55.75 | 56
18.1 | 46.5 | 51.05 | 54.3 | 55.5 | 55.85
18.2 | 46.2 | 50.9 | 54 | 55.3 | 55.625
18.3 | 45.95 | 50.65 | 53.8 | 55.1 | 55.36
18.4 | 45.7 | 50.45 | 53.6 | 54.95 | 55.1
18.5 | 45.46 | 50.125 | 53.35 | 54.72 | 54.975
18.6 | 45.3 | 49.8333 | 53.2 | 54.5 | 54.8
18.7 | 45.075 | 49.6 | 52.975 | 54.325 | 54.5
18.8 | 44.8625 | 49.4333 | 52.7 | 54.06 | 54.325
18.9 | 44.675 | 49.2 | 52.475 | 53.825 | 54.075
19 | 44.4429 | 49.05 | 52.3 | 53.575 | 53.9
19.1 | 44.175 | 48.8 | 52.075 | 53.375 | 53.675
19.2 | 43.9 | 48.6 | 51.825 | 53.2 | 53.525
19.3 | 43.675 | 48.3 | 51.575 | 52.975 | 53.25
19.4 | 43.4 | 48.2 | 51.325 | 52.725 | 53
19.5 | 43.375 | 47.95 | 51.2 | 52.475 | 52.825
19.6 | 43.125 | 47.7 | 51.075 | 52.325 | 52.575
19.7 | 42.8889 | 47.55 | 50.8 | 52.125 | 52.325
19.8 | 42.65 | 47.3 | 50.575 | 51.9 | 52.18
19.9 | 42.5 | 47.0667 | 50.34 | 51.775 | 52.04
20 | 42.25 | 46.9 | 50.1 | 51.54 | 51.7667
20.1 | 42 | 46.5667 | 49.9333 | 51.3333 | 51.6
20.2 | 41.8 | 46.4 | 49.8 | 51.075 | 51.35
20.3 | 41.625 | 46.2333 | 49.525 | 50.9 | 51.2
20.4 | 41.4286 | 46 | 49.3 | 50.6 | 50.95
20.5 | 41.2 | 45.8333 | 49.1 | 50.44 | 50.7
20.6 | 40.975 | 45.575 | 48.96 | 50.3 | 50.5
20.7 | 40.8 | 45.425 | 48.65 | 50.1 | 50.4
20.8 | 40.625 | 45.2 | 48.5 | 49.8 | 50.2
20.9 | 40.4 | 44.95 | 48.3 | 49.7 | 50.02
21 | 40.16 | 44.7 | 48.1 | 49.4667 | 49.76
21.1 | 40 | 44.5 | 47.85 | 49.25 | 49.6
21.2 | 39.775 | 44.3 | 47.7 | 49.1 | 49.35
21.3 | 39.6111 | 44.1 | 47.55 | 48.9 | 49.12
21.4 | 39.375 | 43.9 | 47.3 | 48.65 | 48.975
21.5 | 39.2125 | 43.75 | 47.05 | 48.5 | 48.8
21.6 | 38.975 | 43.5333 | 46.9 | 48.35 | 48.6
21.7 | 38.7625 | 43.3 | 46.65 | 48.125 | 48.4
21.8 | 38.575 | 43.1 | 46.6 | 47.975 | 48.225
21.9 | 38.4125 | 42.9 | 46.3 | 47.725 | 47.975
22 | 38.275 | 42.75 | 46.1 | 47.575 | 47.725
22.1 | 38.025 | 42.5 | 45.975 | 47.325 | 47.575
22.2 | 37.8375 | 42.3 | 45.725 | 47.1667 | 47.425
22.3 | 37.7125 | 42.1 | 45.5 | 46.9 | 47.25
22.4 | 37.525 | 42 | 45.4 | 46.8 | 47.025
22.5 | 37.3 | 41.8 | 45.175 | 46.575 | 46.88
22.6 | 37.1 | 41.6 | 45 | 46.425 | 46.625
22.7 | 36.95 | 41.4 | 44.8 | 46.25 | 46.475
22.8 | 36.8 | 41.1667 | 44.6 | 46.0333 | 46.325
22.9 | 36.575 | 41.025 | 44.425 | 45.875 | 46.1333
23 | 36.4 | 40.8 | 44.275 | 45.7 | 46
23.1 | 36.2714 | 40.625 | 44.025 | 45.48 | 45.7333
23.2 | 36 | 40.475 | 43.875 | 45.275 | 45.5
23.3 | 35.825 | 40.225 | 43.7 | 45.075 | 45.45
23.4 | 35.7 | 40.1 | 43.475 | 44.925 | 45.3
23.5 | 35.4778 | 39.925 | 43.3 | 44.775 | 45.05
23.6 | 35.325 | 39.7 | 43.08 | 44.525 | 44.9
23.7 | 35.175 | 39.525 | 42.95 | 44.4 | 44.7
23.8 | 35.0125 | 39.375 | 42.8 | 44.18 | 44.5
23.9 | 34.7875 | 39.225 | 42.5 | 44.04 | 44.35
24 | 34.6571 | 38.975 | 42.35 | 43.8 | 44.12
24.1 | 34.4667 | 38.8 | 42.15 | 43.65 | 43.95
24.2 | 34.3125 | 38.7 | 42 | 43.5 | 43.8
24.3 | 34.0875 | 38.5 | 41.85 | 43.35 | 43.6667
24.4 | 33.9625 | 38.3 | 41.6333 | 43.1 | 43.5
24.5 | 33.8375 | 38.1 | 41.5 | 42.94 | 43.275
24.6 | 33.625 | 37.95 | 41.35 | 42.8 | 43.1
24.7 | 33.45 | 37.8 | 41.1667 | 42.625 | 42.925
24.8 | 33.3125 | 37.65 | 40.9333 | 42.4 | 42.8
24.9 | 33.0875 | 37.4 | 40.8 | 42.24 | 42.575
25 | 32.9625 | 37.25 | 40.6 | 42.1 | 42.425
25.1 | 32.8444 | 37.1 | 40.475 | 41.95 | 42.275
25.2 | 32.65 | 36.95 | 40.3 | 41.8 | 42.04
25.3 | 32.5 | 36.8 | 40.1 | 41.6 | 41.875
25.4 | 32.35 | 36.55 | 40.025 | 41.4 | 41.75
25.5 | 32.15 | 36.42 | 39.8 | 41.2 | 41.6
25.6 | 32.025 | 36.275 | 39.625 | 41.1 | 41.4
25.7 | 31.9 | 36.12 | 39.475 | 40.875 | 41.24
25.8 | 31.675 | 35.9 | 39.325 | 40.7 | 41.1
25.9 | 31.55 | 35.725 | 39.1 | 40.55 | 40.9
26 | 31.425 | 35.575 | 38.925 | 40.4 | 40.7
26.1 | 31.2 | 35.4 | 38.8 | 40.25 | 40.55
26.2 | 31.075 | 35.275 | 38.6 | 40.0333 | 40.4
26.3 | 30.9333 | 35.125 | 38.4 | 39.9 | 40.25
26.4 | 30.8 | 34.975 | 38.35 | 39.8 | 40.1
26.5 | 30.65 | 34.8 | 38.1 | 39.62 | 39.9
26.6 | 30.525 | 34.6 | 38 | 39.5 | 39.7
26.7 | 30.3 | 34.5 | 37.8 | 39.3 | 39.55
26.8 | 30.2 | 34.3 | 37.7 | 39.1 | 39.4
26.9 | 30.0556 | 34.125 | 37.5 | 38.9 | 39.25
27 | 29.8571 | 33.975 | 37.35 | 38.8 | 39.1
27.1 | 29.7 | 33.875 | 37.2 | 38.6 | 38.95
27.2 | 29.5875 | 33.625 | 37.05 | 38.5 | 38.8
27.3 | 29.4625 | 33.5 | 36.9 | 38.3 | 38.6
27.4 | 29.3375 | 33.4 | 36.7 | 38.1667 | 38.5
27.5 | 29.2125 | 33.2 | 36.5 | 37.95 | 38.3333
27.6 | 29.075 | 33.05 | 36.4667 | 37.825 | 38.2
27.7 | 28.8625 | 32.9 | 36.225 | 37.7 | 38.05
27.8 | 28.7375 | 32.8 | 36.075 | 37.525 | 37.9
27.9 | 28.6125 | 32.6 | 35.925 | 37.3667 | 37.7
28 | 28.4875 | 32.5 | 35.775 | 37.2 | 37.6
28.1 | 28.3625 | 32.3 | 35.625 | 37.1 | 37.4333
28.2 | 28.175 | 32.15 | 35.475 | 36.975 | 37.275
28.3 | 28.1 | 32.1 | 35.325 | 36.8 | 37.125
28.4 | 27.9 | 31.85 | 35.175 | 36.675 | 36.975
28.5 | 27.775 | 31.7 | 35.025 | 36.5 | 36.825
28.6 | 27.65 | 31.6 | 34.9167 | 36.3 | 36.7
28.7 | 27.55 | 31.5 | 34.7 | 36.2 | 36.525
28.8 | 27.4 | 31.325 | 34.6 | 36 | 36.4
28.9 | 27.275 | 31.2 | 34.4 | 35.9 | 36.2
29 | 27.15 | 31.025 | 34.25 | 35.775 | 36.075
29.1 | 27.0143 | 30.875 | 34.1 | 35.6 | 35.9333
29.2 | 26.9 | 30.725 | 34 | 35.375 | 35.8
29.3 | 26.725 | 30.675 | 33.875 | 35.3 | 35.625
29.4 | 26.575 | 30.525 | 33.75 | 35.1 | 35.475
29.5 | 26.425 | 30.3667 | 33.6 | 35 | 35.35
29.6 | 26.375 | 30.175 | 33.4333 | 34.8667 | 35.25
29.7 | 26.225 | 30.1 | 33.3 | 34.7 | 35.075
29.8 | 26.1 | 29.9 | 33.2 | 34.6 | 34.9
29.9 | 26.025 | 29.8 | 33.0333 | 34.475 | 34.8
30 | 25.9 | 29.675 | 32.875 | 34.325 | 34.6375

## Wedge 15 Lateral Profiles
Lateral Profiles
Wedge 15
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:2 cm FieldSize:4x4 cm2 | Depth:2 cm FieldSize:10x10 cm2 | Depth:2 cm FieldSize:20x20 cm2 | Depth:2 cm FieldSize:30x30 cm2 | Depth:2 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-21.7 |  |  |  | 5.82 | 6.8 |  |  |  | 6.4 | 7.44 |  |  |  | 8.9 | 9.95 |  |  |  | 14.6333 | 15.8833 |  |  | 7.3 | 21.26 | 22.6
-21.6 |  |  |  | 5.9 | 6.9 |  |  |  | 6.5 | 7.55 |  |  |  | 9 | 10.1 |  |  |  | 14.8833 | 16.15 |  |  | 7.3333 | 21.5 | 22.9167
-21.5 |  |  |  | 5.96 | 6.9667 |  |  |  | 6.6 | 7.6 |  |  |  | 9.18 | 10.2 |  |  |  | 15.1 | 16.3167 |  |  | 7.5 | 21.88 | 23.2333
-21.4 |  |  |  | 6.075 | 7.0286 |  |  |  | 6.7 | 7.78 |  |  |  | 9.26 | 10.3667 |  |  |  | 15.3167 | 16.5667 |  |  | 7.5 | 22.1333 | 23.55
-21.3 |  |  |  | 6.1 | 7.1 |  |  |  | 6.85 | 7.8167 |  |  |  | 9.4333 | 10.475 |  |  |  | 15.5667 | 16.85 |  |  | 7.6667 | 22.5 | 23.8333
-21.2 |  |  |  | 6.2 | 7.2 |  |  |  | 6.95 | 7.9 |  |  |  | 9.6 | 10.6429 |  |  |  | 15.7667 | 17 |  |  | 7.74 | 22.7429 | 24.12
-21.1 |  |  |  | 6.3 | 7.3 |  |  |  | 7.025 | 8.0833 |  |  |  | 9.7 | 10.8 |  |  |  | 16 | 17.2667 |  |  | 7.8 | 23.16 | 24.5667
-21 |  |  |  | 6.375 | 7.4 |  |  |  | 7.125 | 8.2429 |  |  |  | 9.8833 | 10.9667 |  |  |  | 16.3 | 17.5667 |  |  | 7.9 | 23.4667 | 24.9
-20.9 |  |  |  | 6.5 | 7.5 |  |  |  | 7.3 | 8.34 |  |  |  | 10 | 11.1 |  |  |  | 16.5333 | 17.8 |  |  | 7.9 | 23.9 | 25.3
-20.8 |  |  |  | 6.6 | 7.58 |  |  |  | 7.4 | 8.4286 |  |  |  | 10.1 | 11.25 |  |  |  | 16.82 | 18.0667 |  |  | 8 | 24.4 | 25.8
-20.7 |  |  |  | 6.7 | 7.6667 |  |  |  | 7.55 | 8.5833 |  |  |  | 10.28 | 11.3667 |  |  |  | 17.0286 | 18.34 |  |  | 8.1 | 24.86 | 26.16
-20.6 |  |  |  | 6.78 | 7.75 |  |  |  | 7.6167 | 8.7 |  |  |  | 10.46 | 11.6 |  |  |  | 17.35 | 18.6571 |  |  | 8.2 | 25.3714 | 26.64
-20.5 |  |  |  | 6.85 | 7.9 |  |  |  | 7.7 | 8.8 |  |  |  | 10.66 | 11.75 |  |  |  | 17.6 | 18.8833 |  |  | 8.3 | 25.92 | 27.3
-20.4 |  |  |  | 7 | 8 |  |  |  | 7.8714 | 8.9667 |  |  |  | 10.8 | 11.9 |  |  |  | 17.9333 | 19.2 |  |  | 8.3667 | 26.66 | 27.96
-20.3 |  |  |  | 7.0833 | 8.0667 |  |  |  | 8.0333 | 9.04 |  |  |  | 10.975 | 12.08 |  |  | 5.9 | 18.1833 | 19.4333 |  |  | 8.5 | 27.5 | 28.8
-20.2 |  |  |  | 7.2429 | 8.2667 |  |  |  | 8.1 | 9.1167 |  |  |  | 11.1714 | 12.2833 |  |  | 5.9 | 18.5 | 19.7667 |  |  | 8.6 | 28.7 | 29.85
-20.1 |  |  |  | 7.3 | 8.3333 |  |  |  | 8.2 | 9.2833 |  |  |  | 11.4333 | 12.5 |  |  | 6 | 18.8333 | 20.05 |  |  | 8.74 | 29.9 | 31.22
-20 |  |  |  | 7.4 | 8.4857 |  |  |  | 8.34 | 9.4 |  |  |  | 11.55 | 12.6167 |  |  | 6 | 19.0833 | 20.4286 |  |  | 8.825 | 31.95 | 33.2
-19.9 |  |  |  | 7.55 | 8.5 |  |  |  | 8.5 | 9.56 |  |  |  | 11.7667 | 12.8667 |  |  | 6.1 | 19.4333 | 20.76 |  |  | 9 | 34.64 | 35.7
-19.8 |  |  |  | 7.6167 | 8.6667 |  |  |  | 8.6 | 9.7 |  |  |  | 11.9667 | 13.05 |  |  | 6.2 | 19.7286 | 21.0333 |  |  | 9.0667 | 38.46 | 39.4
-19.7 |  |  |  | 7.7167 | 8.8 |  |  |  | 8.7571 | 9.8 |  |  |  | 12.2 | 13.26 |  |  | 6.2833 | 20.0333 | 21.4 |  |  | 9.1333 | 43.54 | 44.5
-19.6 |  |  |  | 7.92 | 8.92 |  |  |  | 8.92 | 9.9833 |  |  |  | 12.36 | 13.45 |  |  | 6.4 | 20.4667 | 21.7333 |  |  | 9.3 | 50.3 | 50.44
-19.5 |  |  |  | 8 | 9.0167 |  |  |  | 9 | 10.15 |  |  |  | 12.55 | 13.675 |  |  | 6.4 | 20.9 | 22.1 |  |  | 9.3833 | 57.1 | 57.5
-19.4 |  |  |  | 8.1 | 9.1 |  |  |  | 9.15 | 10.22 |  |  |  | 12.8167 | 13.9 |  |  | 6.5 | 21.3 | 22.5 |  |  | 9.5 | 64.75 | 64.9
-19.3 |  |  |  | 8.25 | 9.2833 |  |  |  | 9.3167 | 10.4167 |  |  |  | 13.0667 | 14.1429 |  |  | 6.6 | 21.7 | 22.92 |  |  | 9.6333 | 71.46 | 71.44
-19.2 |  |  |  | 8.32 | 9.4429 |  |  |  | 9.5167 | 10.5714 |  |  |  | 13.3 | 14.3667 |  |  | 6.6333 | 22.1714 | 23.35 |  |  | 9.7 | 77.84 | 77.58
-19.1 |  |  |  | 8.5 | 9.5 |  |  |  | 9.6833 | 10.7667 |  |  |  | 13.52 | 14.6 |  |  | 6.8 | 22.7143 | 23.92 |  |  | 9.88 | 82.88 | 82.55
-19 |  |  |  | 8.6 | 9.6667 |  |  |  | 9.775 | 10.96 |  |  |  | 13.7333 | 14.9 |  |  | 6.8 | 23.2833 | 24.5667 |  |  | 10 | 86.975 | 86.5
-18.9 |  |  |  | 8.7 | 9.8 |  |  |  | 9.95 | 11.0286 |  |  |  | 13.9833 | 15.15 |  |  | 6.9 | 24.0333 | 25.2333 |  |  | 10.0833 | 89.5 | 88.9
-18.8 |  |  |  | 8.9 | 9.96 |  |  |  | 10.1333 | 11.2667 |  |  |  | 14.2857 | 15.4 |  |  | 6.96 | 24.9 | 26.1 |  |  | 10.15 | 91.32 | 90.8333
-18.7 |  |  |  | 9 | 10.1333 |  |  |  | 10.4 | 11.4 |  |  |  | 14.5 | 15.68 |  |  | 7.0333 | 26.075 | 27.3333 |  |  | 10.3167 | 92.5333 | 92
-18.6 |  |  |  | 9.1667 | 10.2 |  |  |  | 10.5 | 11.6167 |  |  | 4.5 | 14.8333 | 15.9 |  |  | 7.2 | 27.575 | 28.6667 |  |  | 10.5 | 93.48 | 92.96
-18.5 |  |  |  | 9.32 | 10.3667 |  |  |  | 10.7333 | 11.7833 |  |  | 4.6 | 15.12 | 16.24 |  |  | 7.2833 | 29.75 | 30.9 |  |  | 10.7 | 94.2 | 93.62
-18.4 |  |  |  | 9.4857 | 10.55 |  |  |  | 10.9 | 12 |  |  | 4.6 | 15.4333 | 16.54 |  |  | 7.35 | 32.8 | 33.94 |  |  | 10.8 | 94.85 | 94.22
-18.3 |  |  |  | 9.6 | 10.7 |  |  |  | 11.1333 | 12.2 |  |  | 4.7 | 15.7333 | 16.84 |  |  | 7.5 | 37.3 | 38.525 |  |  | 10.925 | 95.14 | 94.62
-18.2 |  |  |  | 9.76 | 10.82 |  |  |  | 11.3333 | 12.4167 |  |  | 4.74 | 16.02 | 17.1333 |  |  | 7.52 | 43.1 | 44.12 |  |  | 11.1167 | 95.65 | 95
-18.1 |  |  |  | 9.96 | 11.02 |  |  |  | 11.5167 | 12.5714 |  |  | 4.8 | 16.46 | 17.5143 |  |  | 7.7 | 51.1333 | 51.8167 |  |  | 11.2333 | 95.9833 | 95.32
-18 |  |  |  | 10.1 | 11.2 |  |  |  | 11.7667 | 12.8 |  |  | 4.84 | 16.8 | 17.8 |  |  | 7.8 | 59.34 | 60.4 |  |  | 11.4 | 96.3 | 95.6
-17.9 |  |  |  | 10.3 | 11.2833 |  |  |  | 11.95 | 13.05 |  |  | 4.9 | 17.18 | 18.2667 |  |  | 7.9 | 68.18 | 68.88 |  |  | 11.5333 | 96.45 | 95.94
-17.8 |  |  |  | 10.48 | 11.54 |  |  | 3.9 | 12.225 | 13.2667 |  |  | 4.9 | 17.65 | 18.7 |  |  | 8 | 77 | 77 |  |  | 11.7 | 96.72 | 96.2
-17.7 |  |  |  | 10.6 | 11.74 |  |  | 3.9 | 12.4333 | 13.5 |  |  | 5 | 18.15 | 19.2 |  |  | 8.1 | 84.34 | 84.35 |  |  | 11.8571 | 96.98 | 96.4167
-17.6 |  |  |  | 10.82 | 11.9 |  |  | 4 | 12.6833 | 13.7667 |  |  | 5.1 | 18.65 | 19.7333 |  |  | 8.2 | 89.8667 | 89.9 |  |  | 12.0167 | 97.18 | 96.6
-17.5 |  |  |  | 11 | 12.1 |  |  | 4 | 12.9333 | 14.0667 |  |  | 5.2 | 19.26 | 20.3429 |  |  | 8.3714 | 94.2 | 93.94 |  |  | 12.1714 | 97.4 | 96.7
-17.4 |  |  |  | 11.2 | 12.3 |  |  | 4.1 | 13.2 | 14.3571 |  |  | 5.2 | 20.05 | 21.12 |  |  | 8.5 | 96.9833 | 96.5 |  |  | 12.35 | 97.6 | 96.86
-17.3 |  |  |  | 11.4 | 12.5143 |  |  | 4.1 | 13.5333 | 14.6 |  |  | 5.3 | 21 | 22.05 |  |  | 8.6167 | 98.6 | 98.2 |  |  | 12.6 | 97.7667 | 96.96
-17.2 |  |  |  | 11.6 | 12.7286 |  |  | 4.1 | 13.9 | 14.9667 |  |  | 5.4 | 22.2 | 23.25 |  |  | 8.7 | 99.9 | 99.5 |  |  | 12.7667 | 97.86 | 97.1667
-17.1 |  |  | 4 | 11.825 | 12.9167 |  |  | 4.2 | 14.24 | 15.35 |  |  | 5.4143 | 24 | 24.9167 |  |  | 8.82 | 100.8167 | 100.3 |  |  | 12.9286 | 97.95 | 97.38
-17 |  |  | 4.0667 | 12.05 | 13.0833 |  |  | 4.2 | 14.6 | 15.68 |  |  | 5.5 | 26.5 | 27.5 |  |  | 9.025 | 101.34 | 100.9 |  |  | 13.1 | 98.0167 | 97.5
-16.9 |  |  | 4.1 | 12.275 | 13.2833 |  |  | 4.25 | 15.1 | 16.15 |  |  | 5.6 | 30.5 | 31.1571 |  |  | 9.2 | 101.9 | 101.35 |  |  | 13.3 | 98.2 | 97.6
-16.8 |  |  | 4.1 | 12.5 | 13.6 |  |  | 4.3 | 15.6333 | 16.7333 |  |  | 5.62 | 36.7333 | 36.5667 |  |  | 9.26 | 102.28 | 101.68 |  |  | 13.4833 | 98.3 | 97.7667
-16.7 |  |  | 4.2 | 12.75 | 13.88 |  |  | 4.375 | 16.28 | 17.36 |  |  | 5.7 | 45.0143 | 44.75 |  |  | 9.4 | 102.5667 | 102.0333 |  |  | 13.66 | 98.3167 | 97.8
-16.6 |  |  | 4.2 | 13.1 | 14.1667 |  |  | 4.4 | 17.1 | 18.16 |  |  | 5.8 | 54.9 | 54.65 |  |  | 9.6 | 102.8 | 102.3333 |  |  | 13.96 | 98.4 | 97.9
-16.5 |  |  | 4.275 | 13.4 | 14.4667 |  |  | 4.5 | 18.14 | 19.18 |  |  | 5.8333 | 65.65 | 65.1 |  |  | 9.7167 | 103.08 | 102.58 |  |  | 14.1667 | 98.5 | 98
-16.4 |  |  | 4.3 | 13.78 | 14.9 |  |  | 4.5 | 19.66 | 20.6667 |  |  | 5.9333 | 76.86 | 76.2 |  |  | 9.8 | 103.2667 | 102.7667 |  |  | 14.4 | 98.6 | 98.1
-16.3 |  |  | 4.3 | 14.32 | 15.4 |  |  | 4.6 | 21.76 | 22.9 |  |  | 6 | 87.14 | 86.5667 |  |  | 10.0667 | 103.42 | 102.84 |  |  | 14.52 | 98.6 | 98.1
-16.2 |  |  | 4.4 | 14.98 | 16.0333 |  |  | 4.6 | 25 | 26.2 |  |  | 6.1 | 95.75 | 94.78 |  |  | 10.2429 | 103.5 | 103.04 |  |  | 14.8 | 98.7667 | 98.2
-16.1 |  |  | 4.5 | 15.8 | 16.8 |  |  | 4.7 | 30.1333 | 31.36 |  |  | 6.26 | 102.1333 | 101.1667 |  |  | 10.42 | 103.7 | 103.14 |  |  | 14.98 | 98.8 | 98.2
-16 |  |  | 4.5 | 16.96 | 18 |  |  | 4.7 | 37.9 | 39.28 |  |  | 6.3 | 105.8714 | 105.3333 |  |  | 10.5167 | 103.8 | 103.2 |  |  | 15.2 | 98.9 | 98.3
-15.9 |  |  | 4.5 | 18.6 | 19.5857 |  |  | 4.8 | 49.12 | 49.7 |  |  | 6.4 | 108.54 | 107.65 |  |  | 10.6714 | 103.8 | 103.2833 |  |  | 15.5 | 98.9833 | 98.4
-15.8 |  |  | 4.6 | 21.45 | 22.25 |  |  | 4.9 | 61.325 | 62.3 |  |  | 6.5 | 110.0429 | 109.32 |  |  | 10.8667 | 103.9 | 103.35 |  |  | 15.7143 | 99.05 | 98.5
-15.7 |  |  | 4.6 | 25.96 | 26.5167 |  |  | 4.9333 | 74.04 | 75.12 |  |  | 6.68 | 111.1 | 110.3833 |  |  | 11.12 | 103.9 | 103.4167 |  |  | 15.9333 | 99.1 | 98.56
-15.6 |  |  | 4.7 | 33.5 | 33.85 |  |  | 5 | 86.28 | 86.3 |  |  | 6.7 | 111.9 | 111.05 |  |  | 11.2 | 104 | 103.5 |  |  | 16.2667 | 99.1833 | 98.6667
-15.5 |  |  | 4.7 | 43.65 | 44.4 |  |  | 5.1 | 97.4 | 97 |  |  | 6.8 | 112.3 | 111.4667 |  |  | 11.4 | 104 | 103.5 |  | 3.9 | 16.5 | 99.25 | 98.7
-15.4 |  |  | 4.8 | 56.64 | 56.95 |  |  | 5.175 | 106 | 105.3 |  |  | 6.9 | 112.6333 | 111.78 |  |  | 11.6 | 104 | 103.6 |  | 4 | 16.8167 | 99.3 | 98.8
-15.3 |  |  | 4.9 | 70.55 | 70.26 |  |  | 5.2 | 111.8 | 111.04 |  |  | 7.0667 | 112.84 | 112.0333 |  |  | 11.8333 | 104 | 103.6 |  | 4 | 16.9833 | 99.3833 | 98.96
-15.2 |  |  | 4.9 | 84.45 | 83.5 |  |  | 5.25 | 114.7 | 114.1333 |  |  | 7.1333 | 113 | 112.2 |  |  | 12.0571 | 104.1 | 103.66 |  | 4.0667 | 17.275 | 99.4429 | 99
-15.1 |  |  | 5 | 97.24 | 96.2 |  |  | 5.4 | 116.65 | 116.2 |  |  | 7.3 | 113.025 | 112.2667 |  |  | 12.2 | 104.1 | 103.7 |  | 4.1 | 17.56 | 99.6 | 99.08
-15 |  |  | 5.0333 | 107.62 | 106.4667 |  |  | 5.4333 | 117.85 | 117.2333 |  |  | 7.4 | 113.1 | 112.3 |  |  | 12.4 | 104.1 | 103.7 |  | 4.125 | 17.8667 | 99.6667 | 99.18
-14.9 |  |  | 5.1 | 115 | 114.1333 |  |  | 5.5 | 118.575 | 117.8857 |  |  | 7.525 | 113.1 | 112.3 |  |  | 12.6667 | 104.1 | 103.7 |  | 4.2 | 18.1667 | 99.8 | 99.2
-14.8 |  |  | 5.2 | 119.25 | 118.5 |  |  | 5.6 | 118.9333 | 118.3167 |  |  | 7.7 | 113.1 | 112.3 |  |  | 12.9 | 104.1 | 103.7 |  | 4.2 | 18.5 | 99.9 | 99.3
-14.7 |  |  | 5.2 | 121.35 | 120.5667 |  |  | 5.7 | 119.2 | 118.4833 |  |  | 7.85 | 113.0167 | 112.3 |  |  | 13.1667 | 104.1 | 103.7 |  | 4.3 | 18.9 | 100 | 99.475
-14.6 |  |  | 5.3 | 122.5 | 121.7143 |  |  | 5.8 | 119.3 | 118.56 |  |  | 8 | 113 | 112.25 |  |  | 13.3857 | 104.1 | 103.7 |  | 4.3 | 19.18 | 100.0571 | 99.6
-14.5 |  |  | 5.4 | 123.1 | 122.3 |  |  | 5.9 | 119.3 | 118.6 |  |  | 8.2 | 112.9 | 112.1833 |  |  | 13.6 | 104.1 | 103.7 |  | 4.4 | 19.56 | 100.2 | 99.75
-14.4 |  |  | 5.5 | 123.3667 | 122.5833 |  |  | 6 | 119.2333 | 118.5833 |  |  | 8.28 | 112.8333 | 112.1 |  |  | 13.8667 | 104.1 | 103.6 |  | 4.5 | 19.94 | 100.3 | 99.8
-14.3 |  |  | 5.575 | 123.4 | 122.6 |  |  | 6.1 | 119.1667 | 118.5 |  |  | 8.4 | 112.725 | 112 |  | 3 | 14.22 | 104.1 | 103.62 |  | 4.5 | 20.4 | 100.4 | 99.8667
-14.2 |  |  | 5.7 | 123.4 | 122.5833 |  |  | 6.2 | 118.9667 | 118.4 |  |  | 8.5833 | 112.5333 | 111.9 |  | 3.025 | 14.44 | 104.1 | 103.7 |  | 4.6 | 20.8 | 100.5 | 100.0667
-14.1 |  |  | 5.75 | 123.28 | 122.4167 |  |  | 6.3 | 118.9 | 118.2 |  |  | 8.7667 | 112.44 | 111.8 |  | 3.125 | 14.72 | 104.1 | 103.7 |  | 4.6 | 21.3 | 100.56 | 100.1
-14 |  |  | 5.8 | 123.1 | 122.22 |  |  | 6.4 | 118.72 | 118.02 |  |  | 8.95 | 112.32 | 111.6333 |  | 3.2 | 15.04 | 104.1 | 103.7 |  | 4.7 | 21.96 | 100.7 | 100.2
-13.9 |  |  | 5.9 | 122.7667 | 122.0143 |  |  | 6.5 | 118.425 | 117.8333 |  |  | 9.16 | 112.2 | 111.5 |  | 3.2 | 15.3571 | 104.2 | 103.8 |  | 4.7 | 22.66 | 100.8 | 100.2833
-13.8 |  |  | 6.075 | 122.4667 | 121.7 |  |  | 6.6 | 118.22 | 117.5667 |  |  | 9.375 | 112.05 | 111.4 |  | 3.25 | 15.6 | 104.2 | 103.85 |  | 4.8 | 23.6 | 100.8833 | 100.4
-13.7 |  |  | 6.1 | 122.1333 | 121.34 |  |  | 6.8 | 118 | 117.3 |  |  | 9.4833 | 111.95 | 111.3 |  | 3.3 | 15.9667 | 104.2 | 103.9 |  | 4.8667 | 24.7167 | 101 | 100.5
-13.6 |  |  | 6.2 | 121.7 | 121.05 |  |  | 6.9 | 117.68 | 117.0333 |  |  | 9.75 | 111.78 | 111.16 |  | 3.3 | 16.32 | 104.3 | 104 |  | 4.9 | 26.0667 | 101.1 | 100.6
-13.5 |  |  | 6.3 | 121.3833 | 120.6333 |  |  | 7.05 | 117.45 | 116.8333 |  |  | 9.94 | 111.6 | 111 |  | 3.4 | 16.66 | 104.4 | 104.08 |  | 5 | 28.0667 | 101.26 | 100.7
-13.4 |  |  | 6.45 | 121.05 | 120.3 |  |  | 7.2 | 117.2 | 116.5286 |  |  | 10.2 | 111.425 | 110.8 |  | 3.4 | 17.08 | 104.5 | 104.1 |  | 5.05 | 30.85 | 101.3 | 100.76
-13.3 |  |  | 6.5333 | 120.75 | 119.9333 |  |  | 7.325 | 116.875 | 116.22 |  |  | 10.3833 | 111.2333 | 110.6333 |  | 3.5 | 17.54 | 104.5 | 104.15 |  | 5.2 | 34.82 | 101.3 | 100.85
-13.2 |  |  | 6.675 | 120.28 | 119.6429 |  |  | 7.5 | 116.62 | 115.95 |  |  | 10.6 | 111.1 | 110.4833 |  | 3.5 | 17.96 | 104.6 | 104.2 |  | 5.2 | 39.62 | 101.4 | 101
-13.1 |  |  | 6.8 | 119.96 | 119.1667 |  |  | 7.6333 | 116.32 | 115.6667 |  | 2.2 | 10.85 | 110.94 | 110.3833 |  | 3.6 | 18.4 | 104.7 | 104.26 |  | 5.3 | 46.5 | 101.5 | 101.0667
-13 |  |  | 6.975 | 119.5333 | 118.8857 |  |  | 7.825 | 116.1 | 115.4167 |  | 2.3 | 11.1333 | 110.8571 | 110.28 |  | 3.6 | 19.1 | 104.8 | 104.3 |  | 5.3 | 53.625 | 101.5833 | 101.15
-12.9 |  |  | 7.05 | 119.2 | 118.55 |  |  | 8 | 115.75 | 115.15 |  | 2.3 | 11.4 | 110.7857 | 110.2 |  | 3.7 | 19.9 | 104.8 | 104.3 |  | 5.4 | 60.9333 | 101.6143 | 101.2
-12.8 |  |  | 7.2 | 118.8 | 118.05 |  |  | 8.15 | 115.44 | 114.86 |  | 2.3 | 11.7 | 110.7 | 110.1333 |  | 3.7 | 20.7667 | 104.8 | 104.4 |  | 5.4 | 68.12 | 101.7 | 101.3
-12.7 |  |  | 7.3 | 118.4 | 117.725 |  |  | 8.4 | 115.18 | 114.6333 |  | 2.325 | 12 | 110.6 | 110.0714 |  | 3.775 | 22 | 104.9 | 104.4 |  | 5.5 | 74.6 | 101.85 | 101.32
-12.6 |  |  | 7.5 | 117.9667 | 117.4 |  |  | 8.55 | 114.82 | 114.3 |  | 2.4 | 12.3 | 110.5 | 110 |  | 3.8 | 23.4 | 104.9 | 104.5 |  | 5.6 | 79.7 | 102 | 101.4
-12.5 |  |  | 7.625 | 117.7 | 116.975 |  | 1.9 | 8.8 | 114.55 | 114.02 |  | 2.4 | 12.6714 | 110.5 | 109.9 |  | 3.9 | 25.55 | 104.9 | 104.6 |  | 5.7 | 83.8 | 102.06 | 101.4714
-12.4 |  |  | 7.8 | 117.2667 | 116.7 |  | 1.9 | 8.98 | 114.3833 | 113.8 |  | 2.4 | 13 | 110.48 | 109.9 |  | 3.9333 | 28.5833 | 104.9 | 104.625 |  | 5.825 | 86.3 | 102.1 | 101.575
-12.3 |  |  | 8 | 116.9333 | 116.2 |  | 1.9 | 9.175 | 114.2 | 113.6167 |  | 2.4 | 13.35 | 110.4 | 109.9 |  | 4 | 33.2 | 104.98 | 104.7 |  | 5.85 | 88.18 | 102.1 | 101.6
-12.2 |  |  | 8.1333 | 116.6 | 115.9667 |  | 2 | 9.375 | 114.0167 | 113.5 |  | 2.5 | 13.85 | 110.4 | 109.86 |  | 4.1 | 39.2 | 105.05 | 104.7333 |  | 6 | 89.6 | 102.2 | 101.7
-12.1 |  |  | 8.3 | 116.25 | 115.54 |  | 2 | 9.7 | 113.8333 | 113.4 |  | 2.5 | 14.35 | 110.3 | 109.8 |  | 4.1 | 46.72 | 105.1 | 104.8 |  | 6.1 | 90.64 | 102.2 | 101.8
-12 |  | 2 | 8.5 | 115.8833 | 115.3333 |  | 2 | 9.95 | 113.7 | 113.22 |  | 2.55 | 14.9 | 110.3 | 109.8 |  | 4.2 | 54.8714 | 105.2 | 104.8 |  | 6.2 | 91.48 | 102.3 | 101.9
-11.9 |  | 2 | 8.7 | 115.7167 | 115.1333 |  | 2 | 10.275 | 113.6333 | 113.1333 |  | 2.6 | 15.7 | 110.2 | 109.8 |  | 4.2 | 62.9 | 105.28 | 104.8 |  | 6.4 | 92.05 | 102.4 | 101.9833
-11.8 |  | 2 | 8.9 | 115.5571 | 114.9 |  | 2 | 10.6333 | 113.4 | 113 |  | 2.6 | 16.52 | 110.2 | 109.7 |  | 4.3 | 71.3 | 105.3 | 104.875 |  | 6.4 | 92.72 | 102.4 | 102.05
-11.7 |  | 2 | 9.15 | 115.3667 | 114.7333 |  | 2 | 11 | 113.4 | 112.9 |  | 2.7 | 17.72 | 110.1 | 109.7 |  | 4.4 | 78.5667 | 105.3 | 104.9 |  | 6.5 | 93.22 | 102.4 | 102.1
-11.6 |  | 2 | 9.4 | 115.1167 | 114.5667 |  | 2 | 11.5 | 113.3 | 112.8167 |  | 2.7 | 19.3 | 110.1 | 109.6333 |  | 4.5 | 84.2 | 105.3 | 104.9 | 1.7 | 6.6 | 93.55 | 102.5 | 102.2
-11.5 |  | 2 | 9.6667 | 115 | 114.5 |  | 2 | 12 | 113.2 | 112.74 |  | 2.7 | 21.7667 | 110.1 | 109.6 |  | 4.6 | 88.2333 | 105.3 | 104.9 | 1.7 | 6.6 | 93.9 | 102.5 | 102.2
-11.4 |  | 2.1 | 9.9667 | 114.875 | 114.3 |  | 2.1 | 12.65 | 113.1 | 112.6 |  | 2.8 | 25.2667 | 110.1 | 109.5 |  | 4.7 | 90.7857 | 105.3 | 105 | 1.65 | 6.75 | 94.28 | 102.5 | 102.3
-11.3 |  | 2.1 | 10.38 | 114.725 | 114.1833 |  | 2.1 | 13.55 | 113 | 112.45 |  | 2.8 | 30.3 | 110 | 109.5 |  | 4.7 | 92.55 | 105.3 | 105 | 1.7 | 6.85 | 94.65 | 102.5667 | 102.4
-11.2 |  | 2.1 | 10.825 | 114.5667 | 114.0286 |  | 2.1 | 14.8 | 112.88 | 112.4 |  | 2.9 | 37.0833 | 109.98 | 109.5 |  | 4.8 | 94 | 105.3 | 105.1 | 1.7 | 7.04 | 94.9333 | 102.64 | 102.4
-11.1 |  | 2.1 | 11.375 | 114.4 | 113.925 |  | 2.2 | 16.42 | 112.78 | 112.3 |  | 2.9 | 46.7 | 109.9 | 109.5 |  | 4.9 | 95 | 105.4 | 105.1 | 1.7 | 7.2 | 95.26 | 102.8 | 102.4
-11 |  | 2.175 | 12.1 | 114.3 | 113.84 |  | 2.2 | 18.625 | 112.6857 | 112.22 |  | 2.95 | 56.14 | 109.9 | 109.4 |  | 4.95 | 95.65 | 105.4 | 105.1 | 1.7 | 7.3 | 95.5 | 102.8 | 102.4
-10.9 |  | 2.2 | 13.275 | 114.2333 | 113.72 |  | 2.2 | 22.35 | 112.6 | 112.1 |  | 3 | 66.6 | 109.8 | 109.4 |  | 5.1 | 96.35 | 105.4 | 105.1 | 1.7 | 7.4 | 95.8 | 102.9 | 102.4
-10.8 |  | 2.2 | 14.825 | 114.0667 | 113.54 |  | 2.275 | 28.45 | 112.4667 | 112 |  | 3 | 76.5667 | 109.8 | 109.3 |  | 5.1 | 96.9333 | 105.4 | 105.1 | 1.7 | 7.5 | 95.9833 | 102.9 | 102.4
-10.7 |  | 2.2 | 17.16 | 114 | 113.3667 |  | 2.3 | 36.2 | 112.3667 | 111.9 |  | 3.1 | 84.7333 | 109.7 | 109.3 | 1.2 | 5.2 | 97.36 | 105.4 | 105.1 | 1.7 | 7.6 | 96.26 | 102.9 | 102.48
-10.6 |  | 2.2 | 20.7 | 113.85 | 113.3 |  | 2.3 | 47.15 | 112.1833 | 111.8333 |  | 3.16 | 90.44 | 109.64 | 109.275 | 1.2 | 5.325 | 97.7286 | 105.4 | 105.1 | 1.8 | 7.8 | 96.46 | 102.98 | 102.58
-10.5 |  | 2.3 | 27.05 | 113.7 | 113.2 |  | 2.3 | 57.85 | 112.1 | 111.7 |  | 3.2 | 94.84 | 109.56 | 109.1833 | 1.2 | 5.4 | 98 | 105.44 | 105.1 | 1.8 | 7.9667 | 96.6 | 103 | 102.6
-10.4 |  | 2.3 | 36.25 | 113.6167 | 113.1 |  | 2.4 | 69.3 | 112 | 111.6 |  | 3.2 | 97.5 | 109.5 | 109.1 | 1.2 | 5.55 | 98.4333 | 105.4 | 105.1 | 1.8 | 8.1 | 96.9 | 103 | 102.64
-10.3 |  | 2.3 | 47.6 | 113.52 | 112.9 |  | 2.4 | 79.8 | 111.96 | 111.5 |  | 3.3 | 99.3 | 109.5 | 109.0667 | 1.3 | 5.7 | 98.7667 | 105.45 | 105.15 | 1.8 | 8.275 | 97.0667 | 103 | 102.7
-10.2 |  | 2.3 | 59.6 | 113.3 | 112.72 |  | 2.4 | 89 | 111.88 | 111.4167 |  | 3.3 | 100.6833 | 109.4 | 109 | 1.3 | 5.8 | 99 | 105.4 | 105.2 | 1.8 | 8.4 | 97.3333 | 103 | 102.78
-10.1 |  | 2.4 | 71.35 | 113.2 | 112.65 |  | 2.5 | 95.8 | 111.7 | 111.3 |  | 3.4 | 101.4 | 109.4 | 109 | 1.3 | 5.9 | 99.3 | 105.4 | 105.1286 | 1.9 | 8.525 | 97.5 | 103.08 | 102.8
-10 |  | 2.4 | 83.24 | 113.0571 | 112.55 |  | 2.5 | 100.1 | 111.6143 | 111.14 |  | 3.5 | 102.1667 | 109.3167 | 108.9 | 1.3 | 6 | 99.5286 | 105.4 | 105.1 | 1.9 | 8.7 | 97.7 | 103.1 | 102.8
-9.9 |  | 2.4 | 92.6 | 112.94 | 112.42 |  | 2.5667 | 102.6 | 111.5 | 111.1 |  | 3.6 | 102.55 | 109.22 | 108.8857 | 1.3 | 6.2 | 99.6833 | 105.4 | 105.1 | 1.95 | 8.85 | 97.86 | 103.18 | 102.8
-9.8 |  | 2.5 | 99.825 | 112.74 | 112.25 |  | 2.6 | 104.1 | 111.3833 | 111.02 | 0.9 | 3.7 | 103.12 | 109.1333 | 108.7333 | 1.325 | 6.3 | 99.875 | 105.4 | 105.1 | 2 | 9.1 | 98.1571 | 103.2 | 102.8
-9.7 |  | 2.5 | 103.9 | 112.625 | 112.1857 |  | 2.65 | 105.2 | 111.2286 | 110.92 | 0.9 | 3.7 | 103.475 | 109.05 | 108.65 | 1.4 | 6.4 | 100 | 105.4 | 105.1 | 2 | 9.2 | 98.34 | 103.2 | 102.8
-9.6 |  | 2.5 | 106.25 | 112.5 | 112.0333 |  | 2.7 | 106.125 | 111.15 | 110.8571 | 0.9 | 3.8 | 103.8 | 109 | 108.6 | 1.4 | 6.56 | 100.25 | 105.4 | 105.1 | 2 | 9.425 | 98.4667 | 103.2 | 102.8
-9.5 |  | 2.6 | 107.675 | 112.4 | 111.9 |  | 2.7 | 106.575 | 111.1 | 110.7 | 0.9 | 3.9 | 104.05 | 108.95 | 108.5167 | 1.4 | 6.7 | 100.4167 | 105.4 | 105.1 | 2 | 9.675 | 98.6 | 103.2 | 102.9
-9.4 |  | 2.6 | 108.5 | 112.25 | 111.76 | 0.75 | 2.8 | 107.025 | 110.9667 | 110.6 | 0.9 | 3.96 | 104.2167 | 108.88 | 108.5 | 1.4 | 6.8 | 100.5833 | 105.4 | 105.1 | 2 | 9.8 | 98.7 | 103.3 | 102.9
-9.3 |  | 2.6 | 109.1 | 112.2 | 111.6667 | 0.7 | 2.825 | 107.275 | 110.875 | 110.5333 | 0.9 | 4 | 104.3714 | 108.8 | 108.4 | 1.4 | 7 | 100.75 | 105.4 | 105.1 | 2.025 | 9.9 | 98.9167 | 103.3 | 102.95
-9.2 |  | 2.7 | 109.6 | 112.0833 | 111.6 | 0.7 | 2.9 | 107.625 | 110.8 | 110.4 | 0.9 | 4.1 | 104.5 | 108.74 | 108.3857 | 1.4 | 7.1 | 100.92 | 105.3857 | 105.1 | 2.1 | 10.2 | 99 | 103.3 | 103
-9.1 |  | 2.78 | 109.8 | 111.9 | 111.4333 | 0.7 | 2.9 | 107.8 | 110.7 | 110.2833 | 0.9 | 4.2 | 104.7 | 108.7 | 108.3 | 1.4667 | 7.3 | 101.1 | 105.3 | 105.1 | 2.1 | 10.4 | 99.1714 | 103.3 | 103.1
-9 | 0.7 | 2.8 | 109.9 | 111.7 | 111.36 | 0.7 | 3 | 107.95 | 110.6333 | 110.2 | 0.9 | 4.3 | 104.9 | 108.6 | 108.2 | 1.5 | 7.45 | 101.1667 | 105.3 | 105.02 | 2.2 | 10.725 | 99.3333 | 103.35 | 103.1
-8.9 | 0.7 | 2.8 | 110 | 111.5429 | 111.15 | 0.7 | 3.1 | 108.1 | 110.4833 | 110.06 | 0.95 | 4.5 | 105 | 108.5 | 108.1167 | 1.5 | 7.6 | 101.3667 | 105.3 | 105 | 2.2 | 10.875 | 99.5 | 103.4 | 103.1
-8.8 | 0.7 | 2.9 | 110.1 | 111.45 | 111.075 | 0.7 | 3.2 | 108.1 | 110.3 | 109.9667 | 1 | 4.5 | 105.08 | 108.48 | 108.1 | 1.5 | 7.775 | 101.4333 | 105.3 | 104.9 | 2.2 | 11.175 | 99.6 | 103.4 | 103.1
-8.7 | 0.7 | 2.9 | 110.1 | 111.375 | 110.92 | 0.8 | 3.2 | 108.2 | 110.2167 | 109.8833 | 1 | 4.65 | 105.2 | 108.4 | 108.04 | 1.5 | 8 | 101.5333 | 105.3 | 104.95 | 2.2 | 11.4 | 99.7 | 103.3 | 103.1
-8.6 | 0.7 | 3 | 110.1 | 111.1857 | 110.8 | 0.8 | 3.3 | 108.2 | 110.05 | 109.7 | 1 | 4.75 | 105.3 | 108.3 | 107.9 | 1.6 | 8.175 | 101.6 | 105.3 | 104.9 | 2.3 | 11.7 | 99.7714 | 103.3 | 103
-8.5 | 0.7 | 3.075 | 110.1 | 111.0333 | 110.7 | 0.8 | 3.4 | 108.2 | 110 | 109.6 | 1 | 4.9 | 105.3 | 108.2 | 107.8 | 1.6 | 8.4 | 101.7 | 105.2 | 104.9 | 2.4 | 11.925 | 99.9 | 103.4 | 103
-8.4 | 0.7 | 3.1 | 110 | 110.8667 | 110.4833 | 0.8 | 3.4 | 108.2 | 109.9 | 109.5 | 1 | 5.1 | 105.3 | 108.1 | 107.75 | 1.6 | 8.7 | 101.8167 | 105.2 | 104.9 | 2.4 | 12.175 | 100 | 103.4 | 103
-8.3 | 0.7 | 3.2 | 110 | 110.8 | 110.3286 | 0.8 | 3.5 | 108.2 | 109.8 | 109.45 | 1 | 5.2 | 105.4 | 108.1 | 107.7 | 1.7 | 8.8667 | 101.9 | 105.2 | 104.9 | 2.5 | 12.4 | 100.1 | 103.4 | 103.1
-8.2 | 0.8 | 3.3 | 109.9 | 110.6 | 110.26 | 0.8 | 3.65 | 108.2 | 109.6 | 109.3 | 1.1 | 5.3 | 105.4 | 107.98 | 107.6 | 1.7 | 9.0333 | 102.06 | 105.16 | 104.9 | 2.5 | 12.675 | 100.1 | 103.4 | 103.1
-8.1 | 0.8 | 3.35 | 109.9 | 110.5 | 110.1714 | 0.8 | 3.7 | 108.2 | 109.5 | 109.2 | 1.1 | 5.5 | 105.5 | 107.8833 | 107.5167 | 1.7 | 9.325 | 102.15 | 105.1 | 104.9 | 2.5 | 13.1 | 100.2667 | 103.4 | 103.15
-8 | 0.8 | 3.4 | 109.8 | 110.4 | 110 | 0.8 | 3.8333 | 108.1 | 109.4 | 109.1 | 1.1 | 5.7 | 105.5 | 107.8 | 107.5 | 1.7667 | 9.625 | 102.2 | 105.1 | 104.8 | 2.5667 | 13.35 | 100.35 | 103.4 | 103.1
-7.9 | 0.8 | 3.5 | 109.75 | 110.3 | 109.85 | 0.8 | 3.9 | 108.1 | 109.32 | 108.9333 | 1.1 | 5.85 | 105.5 | 107.75 | 107.4 | 1.8 | 9.8667 | 102.2714 | 105 | 104.8 | 2.6 | 13.65 | 100.5 | 103.4 | 103.1
-7.8 | 0.8 | 3.6 | 109.7 | 110.1 | 109.725 | 0.8 | 4 | 108.1 | 109.2 | 108.8 | 1.1 | 6 | 105.5667 | 107.625 | 107.3 | 1.8 | 10.12 | 102.35 | 105 | 104.7667 | 2.6 | 14.075 | 100.5833 | 103.4 | 103.1
-7.7 | 0.8 | 3.7 | 109.5667 | 109.9167 | 109.62 | 0.8 | 4.1667 | 108 | 109.1 | 108.75 | 1.1 | 6.175 | 105.5 | 107.6 | 107.2 | 1.8 | 10.4 | 102.4 | 105 | 104.7 | 2.7 | 14.5667 | 100.7 | 103.4 | 103.1
-7.6 | 0.8 | 3.7667 | 109.5 | 109.85 | 109.45 | 0.8 | 4.25 | 108 | 108.9833 | 108.65 | 1.12 | 6.35 | 105.5 | 107.5167 | 107.1167 | 1.8 | 10.6667 | 102.5 | 104.95 | 104.7 | 2.7 | 15.16 | 100.7 | 103.4 | 103.1
-7.5 | 0.8 | 3.9 | 109.375 | 109.7 | 109.3 | 0.8 | 4.4 | 107.9 | 108.9 | 108.56 | 1.2 | 6.55 | 105.5 | 107.3667 | 107.1 | 1.9 | 11.0333 | 102.54 | 104.95 | 104.7 | 2.7 | 15.78 | 100.8 | 103.4 | 103.1
-7.4 | 0.8 | 4 | 109.3 | 109.66 | 109.2 | 0.8 | 4.5667 | 107.8 | 108.8 | 108.4667 | 1.2 | 6.8 | 105.5 | 107.3 | 107 | 1.9 | 11.4 | 102.6 | 104.9 | 104.66 | 2.8 | 16.5667 | 100.9 | 103.4 | 103.1
-7.3 | 0.8 | 4.08 | 109.1 | 109.46 | 109.06 | 0.9 | 4.7 | 107.8 | 108.6167 | 108.3 | 1.2 | 7 | 105.5 | 107.2 | 106.9 | 2 | 11.8 | 102.6 | 104.9 | 104.6 | 2.9 | 17.4 | 101 | 103.4 | 103
-7.2 | 0.8 | 4.2 | 109 | 109.2833 | 108.98 | 0.9 | 4.8667 | 107.7 | 108.55 | 108.2 | 1.2 | 7.3 | 105.5 | 107.1 | 106.8333 | 2 | 12.3333 | 102.7 | 104.8 | 104.5833 | 2.9 | 18.7 | 101.1 | 103.4 | 103
-7.1 | 0.8 | 4.3 | 108.9 | 109.2 | 108.9 | 0.9 | 5.05 | 107.6 | 108.4 | 108.1 | 1.2 | 7.5667 | 105.5 | 107.05 | 106.74 | 2.1 | 12.775 | 102.7 | 104.775 | 104.5 | 2.9 | 20.3 | 101.1 | 103.4 | 103
-7 | 0.8 | 4.475 | 108.875 | 109.0667 | 108.7333 | 0.9 | 5.24 | 107.6 | 108.3 | 107.98 | 1.3 | 7.8 | 105.5 | 107 | 106.66 | 2.1 | 13.4 | 102.7167 | 104.7 | 104.5 | 3 | 22.4 | 101.2 | 103.4 | 103
-6.9 | 0.8 | 4.6 | 108.75 | 108.9 | 108.56 | 0.9 | 5.475 | 107.5 | 108.1 | 107.8833 | 1.3 | 8.1333 | 105.475 | 106.9 | 106.5667 | 2.14 | 14.125 | 102.8 | 104.7 | 104.5 | 3 | 25.56 | 101.2 | 103.4 | 103
-6.8 | 0.9 | 4.725 | 108.6 | 108.8 | 108.45 | 0.9 | 5.7 | 107.4 | 108 | 107.7167 | 1.3 | 8.5 | 105.4 | 106.8 | 106.5 | 2.2 | 15.0333 | 102.8 | 104.7 | 104.4833 | 3.1 | 30 | 101.3 | 103.32 | 103
-6.7 | 0.9 | 4.9 | 108.5 | 108.6571 | 108.3 | 0.9 | 5.9 | 107.325 | 107.9333 | 107.6 | 1.3 | 9.05 | 105.4 | 106.7 | 106.5 | 2.2 | 16.35 | 102.8 | 104.6 | 104.4 | 3.2 | 35.7333 | 101.3 | 103.4 | 103
-6.6 | 0.9 | 5.15 | 108.3333 | 108.5 | 108.2 | 1 | 6.1667 | 107.2 | 107.8 | 107.5 | 1.3333 | 9.5 | 105.38 | 106.6 | 106.4 | 2.3 | 17.9 | 102.8 | 104.6 | 104.4 | 3.3 | 42.55 | 101.4 | 103.3 | 103
-6.5 | 0.9 | 5.325 | 108.2 | 108.4 | 108.075 | 1 | 6.5 | 107.1 | 107.62 | 107.4 | 1.4 | 10.15 | 105.3 | 106.5833 | 106.2 | 2.325 | 19.8667 | 102.9 | 104.5 | 104.4 | 3.4 | 50.75 | 101.4 | 103.3 | 103
-6.4 | 0.9 | 5.54 | 108.15 | 108.25 | 107.9833 | 1 | 6.95 | 107.1 | 107.52 | 107.3 | 1.4 | 10.84 | 105.3 | 106.5 | 106.2 | 2.4 | 23.3333 | 102.9 | 104.5 | 104.4 | 3.4333 | 59.475 | 101.4 | 103.3 | 103
-6.3 | 0.9 | 5.8667 | 107.9667 | 108.1 | 107.8167 | 1 | 7.4 | 107 | 107.4 | 107.16 | 1.5 | 11.625 | 105.2167 | 106.42 | 106.1 | 2.5 | 27.9 | 102.9667 | 104.4 | 104.3 | 3.5 | 67.9 | 101.4 | 103.3 | 103
-6.2 | 0.9 | 6.225 | 107.82 | 107.98 | 107.7 | 1 | 8 | 106.875 | 107.325 | 107.075 | 1.5 | 12.85 | 105.2 | 106.25 | 106.0286 | 2.5 | 33.6 | 103 | 104.4 | 104.2 | 3.5 | 75.225 | 101.5 | 103.25 | 102.98
-6.1 | 0.9 | 6.65 | 107.725 | 107.8143 | 107.5857 | 1 | 8.7 | 106.8 | 107.225 | 106.96 | 1.55 | 14.5 | 105.1 | 106.2 | 105.94 | 2.6 | 41.6 | 103 | 104.3714 | 104.2 | 3.7 | 81.02 | 101.6 | 103.2 | 102.9
-6 | 0.9 | 7.175 | 107.6 | 107.7 | 107.42 | 1.1 | 9.7 | 106.7 | 107.15 | 106.9 | 1.6 | 16.92 | 105.1 | 106.175 | 105.86 | 2.7 | 51 | 103 | 104.3 | 104.125 | 3.7667 | 85.5 | 101.7 | 103.2 | 102.9
-5.9 | 1 | 7.9333 | 107.6 | 107.5833 | 107.2833 | 1.1 | 11 | 106.66 | 107 | 106.7333 | 1.6 | 20.3 | 105.1 | 106.06 | 105.7 | 2.725 | 60.3 | 103 | 104.25 | 104.1 | 3.8 | 88.6667 | 101.7667 | 103.2 | 102.9
-5.8 | 1 | 9 | 107.475 | 107.4286 | 107.1286 | 1.1 | 12.8667 | 106.5 | 106.9 | 106.7 | 1.6 | 25.8333 | 105.1 | 105.9667 | 105.6167 | 2.8 | 68.425 | 103 | 104.2 | 104.075 | 3.9 | 90.9 | 101.8 | 103.2 | 102.9
-5.7 | 1 | 10.3 | 107.375 | 107.36 | 107 | 1.1 | 15.4 | 106.4667 | 106.8 | 106.5429 | 1.7 | 32.6 | 105 | 105.8 | 105.6 | 2.9 | 76.7 | 103 | 104.2 | 103.9714 | 4 | 92.3667 | 101.8 | 103.1833 | 102.9
-5.6 | 1 | 12.6 | 107.24 | 107.2 | 106.9 | 1.2 | 19.6667 | 106.4 | 106.68 | 106.4714 | 1.7 | 42.3667 | 105 | 105.7333 | 105.5 | 2.9 | 82.875 | 103 | 104.1 | 103.9 | 4.1 | 93.65 | 101.8 | 103.1 | 102.9
-5.5 | 1 | 16.3 | 107.1 | 107.1 | 106.8 | 1.2 | 26.2 | 106.3 | 106.5 | 106.4 | 1.8 | 52.575 | 104.9 | 105.66 | 105.3667 | 3 | 87.6 | 103 | 104 | 103.9 | 4.2 | 94.5 | 101.8 | 103.1 | 102.8
-5.4 | 1.1 | 22.1 | 107 | 106.94 | 106.7 | 1.2 | 35.3667 | 106.125 | 106.4 | 106.22 | 1.8 | 62.925 | 104.8833 | 105.5667 | 105.2 | 3.1 | 90.775 | 103 | 104 | 103.9 | 4.3 | 95.18 | 101.8 | 103.1 | 102.8714
-5.3 | 1.1 | 30 | 106.8 | 106.8667 | 106.6 | 1.3 | 45.6 | 106 | 106.3 | 106.15 | 1.9 | 72.85 | 104.8 | 105.4833 | 105.2 | 3.2 | 93 | 103 | 103.9 | 103.8 | 4.4 | 95.75 | 101.8 | 103.0333 | 102.84
-5.2 | 1.1 | 40.525 | 106.75 | 106.7 | 106.5 | 1.3 | 56.3 | 105.9 | 106.2 | 105.9667 | 1.94 | 80.35 | 104.7 | 105.3167 | 105.1 | 3.3 | 94.3333 | 103 | 103.88 | 103.8 | 4.5 | 96.1333 | 101.8 | 103 | 102.8
-5.1 | 1.1 | 52.4 | 106.6 | 106.6 | 106.4 | 1.35 | 67.8 | 105.8 | 106.1 | 105.88 | 2 | 87.1 | 104.65 | 105.2 | 105.05 | 3.4 | 95.5 | 103 | 103.8 | 103.78 | 4.6 | 96.75 | 101.8 | 103 | 102.8
-5 | 1.1 | 64.45 | 106.5 | 106.4333 | 106.2 | 1.4 | 77.8 | 105.8 | 106 | 105.7 | 2.1 | 91.64 | 104.6 | 105.125 | 104.9667 | 3.5 | 96.4 | 102.9 | 103.8 | 103.7 | 4.8 | 97.15 | 101.8833 | 103 | 102.8
-4.9 | 1.2 | 76 | 106.35 | 106.35 | 106.0833 | 1.4 | 86.4 | 105.7 | 105.82 | 105.6167 | 2.2 | 94.32 | 104.6 | 105.04 | 104.9 | 3.6 | 96.8333 | 102.9 | 103.7333 | 103.6 | 4.9 | 97.3667 | 101.9 | 102.9167 | 102.7
-4.8 | 1.2 | 85.4667 | 106.225 | 106.15 | 105.9 | 1.5 | 91.775 | 105.6 | 105.75 | 105.5 | 2.3 | 96.175 | 104.5 | 104.95 | 104.8 | 3.7 | 97.4 | 102.9 | 103.625 | 103.5167 | 5 | 97.5667 | 101.9 | 102.9 | 102.7
-4.7 | 1.3 | 92.5333 | 106.1 | 106.075 | 105.8 | 1.525 | 95.5333 | 105.45 | 105.6 | 105.4286 | 2.3 | 97.475 | 104.4 | 104.9 | 104.7 | 3.8667 | 98 | 102.9 | 103.6 | 103.5 | 5.18 | 97.9667 | 101.9833 | 102.8 | 102.625
-4.6 | 1.3 | 96.7 | 106 | 105.96 | 105.6833 | 1.6 | 97.55 | 105.4 | 105.5 | 105.3 | 2.46 | 98.3 | 104.3 | 104.8 | 104.6 | 4 | 98.2 | 102.9 | 103.5333 | 103.4 | 5.3 | 98.1 | 101.9 | 102.8 | 102.6833
-4.5 | 1.325 | 99.35 | 105.8 | 105.7667 | 105.58 | 1.6 | 99.025 | 105.175 | 105.3833 | 105.2 | 2.5 | 99.08 | 104.2 | 104.7167 | 104.5 | 4.1 | 98.5 | 102.8 | 103.4 | 103.4 | 5.52 | 98.3 | 101.9 | 102.7 | 102.6
-4.4 | 1.4 | 100.9 | 105.7 | 105.6 | 105.4 | 1.7 | 100.2 | 105.1 | 105.2167 | 105.1 | 2.65 | 99.55 | 104.1 | 104.6 | 104.4 | 4.225 | 98.9 | 102.8 | 103.4 | 103.3 | 5.6 | 98.525 | 101.9 | 102.7 | 102.6
-4.3 | 1.5 | 101.8 | 105.6 | 105.4667 | 105.3286 | 1.8 | 100.85 | 104.9 | 105.2 | 105 | 2.725 | 100.025 | 104.0833 | 104.5333 | 104.3 | 4.475 | 99.075 | 102.8 | 103.35 | 103.2 | 5.9 | 98.7 | 101.9 | 102.7 | 102.6
-4.2 | 1.5 | 102.5 | 105.5 | 105.4 | 105.1667 | 1.9 | 101.3667 | 104.84 | 105.0667 | 104.8571 | 2.825 | 100.3 | 104 | 104.3429 | 104.2 | 4.6 | 99.4 | 102.75 | 103.26 | 103.175 | 6.05 | 98.9 | 101.9 | 102.7 | 102.5167
-4.1 | 1.58 | 102.85 | 105.3 | 105.2 | 105 | 1.95 | 101.7667 | 104.7 | 104.9 | 104.7857 | 3 | 100.6 | 103.9 | 104.3 | 104.1 | 4.75 | 99.5 | 102.7 | 103.2 | 103.1 | 6.3 | 99.075 | 101.9 | 102.7 | 102.5
-4 | 1.6 | 103.1 | 105.2 | 105.12 | 104.9333 | 2.1 | 102.0667 | 104.7 | 104.8 | 104.6333 | 3.1 | 100.8 | 103.8167 | 104.2 | 104 | 4.94 | 99.625 | 102.7 | 103.2 | 103 | 6.625 | 99.2 | 101.8833 | 102.7 | 102.4
-3.9 | 1.7 | 103.3667 | 105.0667 | 104.9833 | 104.7333 | 2.2 | 102.2 | 104.54 | 104.6 | 104.5333 | 3.25 | 101.1 | 103.75 | 104.1 | 103.86 | 5.175 | 99.775 | 102.64 | 103.1 | 102.9 | 6.925 | 99.3 | 101.8 | 102.6 | 102.3
-3.8 | 1.8 | 103.4333 | 104.9 | 104.8167 | 104.64 | 2.3 | 102.35 | 104.4 | 104.5 | 104.325 | 3.5 | 101.2 | 103.6333 | 104 | 103.7667 | 5.425 | 99.9 | 102.6 | 103 | 102.9 | 7.2 | 99.525 | 101.8 | 102.5 | 102.3
-3.7 | 1.82 | 103.5 | 104.775 | 104.675 | 104.5 | 2.4 | 102.5 | 104.3 | 104.4 | 104.24 | 3.7 | 101.3 | 103.6 | 103.9 | 103.7 | 5.7 | 100.1 | 102.6 | 102.9 | 102.9 | 7.62 | 99.675 | 101.9 | 102.5 | 102.3
-3.6 | 1.9667 | 103.5 | 104.62 | 104.5 | 104.4167 | 2.6 | 102.6 | 104.1 | 104.3 | 104.1 | 3.925 | 101.4 | 103.4167 | 103.8 | 103.6 | 6.05 | 100.1667 | 102.5 | 102.8333 | 102.88 | 8.1 | 99.7 | 101.9 | 102.4 | 102.2
-3.5 | 2.1 | 103.5 | 104.525 | 104.38 | 104.25 | 2.7 | 102.7 | 104.05 | 104.15 | 103.9833 | 4.175 | 101.5 | 103.3667 | 103.68 | 103.5 | 6.55 | 100.3 | 102.5 | 102.8 | 102.8 | 8.7 | 99.8 | 101.8167 | 102.36 | 102.2
-3.4 | 2.2 | 103.475 | 104.4 | 104.3 | 104.15 | 3.05 | 102.7 | 103.9667 | 104.05 | 103.9 | 4.5 | 101.6 | 103.2 | 103.5833 | 103.4 | 7.15 | 100.4 | 102.4 | 102.7 | 102.8 | 9.725 | 99.9 | 101.8 | 102.3 | 102.1
-3.3 | 2.4 | 103.4 | 104.18 | 104.2167 | 104 | 3.25 | 102.7 | 103.875 | 103.9 | 103.8 | 4.9 | 101.7 | 103.2 | 103.5 | 103.3 | 7.7 | 100.5 | 102.4 | 102.7 | 102.7 | 10.825 | 100 | 101.8 | 102.24 | 102
-3.2 | 2.58 | 103.4 | 104.1 | 104 | 103.9 | 3.65 | 102.7 | 103.7 | 103.8 | 103.7 | 5.3 | 101.7 | 103.1 | 103.4 | 103.2 | 8.6 | 100.6 | 102.3 | 102.64 | 102.64 | 12.475 | 100 | 101.7 | 102.2 | 102
-3.1 | 2.85 | 103.3 | 104 | 103.9 | 103.7167 | 4.05 | 102.7 | 103.6 | 103.62 | 103.6 | 5.95 | 101.8 | 103.0333 | 103.28 | 103.1333 | 9.72 | 100.7 | 102.3 | 102.56 | 102.5 | 14.6 | 100.2 | 101.7 | 102.1 | 101.9
-3 | 3.3 | 103.3 | 103.9 | 103.8 | 103.6286 | 4.65 | 102.625 | 103.5 | 103.55 | 103.4167 | 6.7 | 101.8 | 103 | 103.2 | 103.1 | 11.2 | 100.7 | 102.2 | 102.4667 | 102.4667 | 17.7 | 100.3667 | 101.6 | 102.1 | 101.9
-2.9 | 3.75 | 103.275 | 103.7 | 103.5667 | 103.4667 | 5.35 | 102.6 | 103.4 | 103.5 | 103.3 | 7.825 | 101.8 | 102.9 | 103.1 | 103 | 13.5 | 100.7 | 102.2 | 102.3833 | 102.3 | 22.3 | 100.45 | 101.6 | 102 | 101.9
-2.8 | 4.5 | 103.1 | 103.6 | 103.5 | 103.38 | 6.4 | 102.6 | 103.275 | 103.3333 | 103.1857 | 9.325 | 101.8 | 102.72 | 103 | 102.8667 | 16.5 | 100.8 | 102.1 | 102.3 | 102.22 | 28.5 | 100.425 | 101.5 | 101.9333 | 101.8
-2.7 | 5.5 | 103 | 103.4 | 103.3333 | 103.2667 | 7.94 | 102.6 | 103.2 | 103.2 | 103 | 11.4333 | 101.8 | 102.6 | 102.9143 | 102.7143 | 21.2 | 100.9 | 102.1 | 102.25 | 102.2 | 36.15 | 100.5 | 101.5 | 101.9 | 101.8
-2.6 | 6.975 | 102.9333 | 103.3 | 103.225 | 103.18 | 10 | 102.6 | 103 | 103.1 | 102.9 | 14.7 | 101.8 | 102.5 | 102.8 | 102.6429 | 27.5 | 100.9 | 102 | 102.2 | 102.1 | 44.95 | 100.5 | 101.5 | 101.8667 | 101.8
-2.5 | 9.5333 | 102.875 | 103.2 | 103.02 | 103.08 | 13.325 | 102.4667 | 102.9 | 102.94 | 102.8 | 19.675 | 101.725 | 102.4 | 102.7 | 102.5 | 36.05 | 101 | 101.9167 | 102.1 | 102.0667 | 55.25 | 100.5 | 101.4167 | 101.8 | 101.7
-2.4 | 13.375 | 102.8 | 103.075 | 102.92 | 102.8667 | 18.5 | 102.4 | 102.8 | 102.84 | 102.7 | 26.925 | 101.7 | 102.3 | 102.5714 | 102.4333 | 45.4333 | 101 | 101.9 | 102 | 102 | 64.75 | 100.5 | 101.4 | 101.6714 | 101.6667
-2.3 | 19.7 | 102.7 | 102.9 | 102.9 | 102.7143 | 25.5 | 102.3 | 102.6333 | 102.7833 | 102.6 | 35.95 | 101.7 | 102.3 | 102.5 | 102.35 | 56.125 | 101 | 101.8 | 101.9333 | 101.9 | 73.425 | 100.5333 | 101.4 | 101.6 | 101.6
-2.2 | 28.625 | 102.575 | 102.8 | 102.75 | 102.65 | 35.7333 | 102.2 | 102.6 | 102.6167 | 102.5 | 45.9 | 101.7 | 102.2 | 102.4 | 102.2 | 66 | 101 | 101.8 | 101.9 | 101.8167 | 80.7 | 100.6 | 101.3667 | 101.6 | 101.6
-2.1 | 39.2 | 102.475 | 102.625 | 102.6 | 102.56 | 46.1 | 102.2 | 102.46 | 102.48 | 102.4167 | 56.85 | 101.6 | 102.1 | 102.3 | 102.1 | 75.175 | 101 | 101.7 | 101.8 | 101.8 | 86.7 | 100.5333 | 101.3 | 101.6 | 101.5
-2 | 51.7667 | 102.325 | 102.5667 | 102.3571 | 102.36 | 58.2 | 102.1 | 102.275 | 102.3 | 102.3 | 67.9 | 101.6 | 102 | 102.16 | 102.0167 | 83.2 | 100.9 | 101.7 | 101.7667 | 101.7167 | 90.7 | 100.5 | 101.25 | 101.5 | 101.5
-1.9 | 63.54 | 102.275 | 102.4 | 102.3 | 102.2 | 69.325 | 102 | 102.125 | 102.2167 | 102.2 | 77.45 | 101.55 | 101.9 | 102 | 101.9571 | 88.6 | 100.9 | 101.6 | 101.6143 | 101.6571 | 93.7 | 100.5 | 101.2 | 101.5 | 101.4
-1.8 | 74.1667 | 102.1333 | 102.3 | 102.2 | 102.1 | 78.7 | 102 | 102 | 102.14 | 102.1 | 84.9 | 101.5 | 101.8 | 101.88 | 101.7857 | 92.525 | 100.9 | 101.6 | 101.5 | 101.5667 | 95.4 | 100.5 | 101.1 | 101.4167 | 101.3167
-1.7 | 83.25 | 102 | 102.2 | 102.1 | 102.1 | 86.8667 | 101.8 | 101.9 | 102.02 | 102 | 90.1 | 101.5 | 101.7 | 101.8 | 101.7 | 95.025 | 100.9 | 101.45 | 101.38 | 101.5 | 96.7 | 100.5667 | 101.1 | 101.4 | 101.2
-1.6 | 90.3 | 101.85 | 102.075 | 101.9667 | 101.9 | 91.65 | 101.7 | 101.825 | 101.92 | 101.84 | 93.4 | 101.4 | 101.6429 | 101.7 | 101.5667 | 96.5 | 100.9 | 101.3667 | 101.3 | 101.35 | 97.8667 | 100.5333 | 101 | 101.3 | 101.2
-1.5 | 94.4 | 101.7 | 101.9667 | 101.8 | 101.8 | 94.98 | 101.65 | 101.8 | 101.8167 | 101.7 | 95.78 | 101.4 | 101.58 | 101.6 | 101.4 | 97.5 | 100.9 | 101.2667 | 101.2167 | 101.2833 | 98.4 | 100.575 | 101 | 101.2333 | 101.15
-1.4 | 96.7 | 101.6 | 101.8 | 101.7 | 101.66 | 96.9333 | 101.55 | 101.7 | 101.6667 | 101.6 | 97 | 101.3 | 101.4143 | 101.5 | 101.34 | 98.2 | 100.8 | 101.2 | 101.2 | 101.2 | 99 | 100.5 | 100.9 | 101.14 | 101.08
-1.3 | 98.3333 | 101.5 | 101.75 | 101.6 | 101.5667 | 98.05 | 101.4 | 101.6 | 101.6 | 101.4 | 98.2 | 101.225 | 101.4 | 101.3667 | 101.3 | 98.95 | 100.8 | 101.1 | 101.1 | 101.1 | 99.3333 | 100.575 | 100.9 | 101.1 | 100.98
-1.2 | 99.2 | 101.4 | 101.5 | 101.48 | 101.4 | 98.8 | 101.3333 | 101.5 | 101.4429 | 101.3167 | 98.9 | 101.2 | 101.2833 | 101.2667 | 101.2 | 99.3 | 100.8 | 101 | 101.04 | 101 | 99.6 | 100.5 | 100.9 | 101 | 100.9
-1.1 | 99.9 | 101.3 | 101.375 | 101.3667 | 101.2 | 99.44 | 101.2 | 101.2667 | 101.2714 | 101.2333 | 99.5 | 101.1 | 101.1167 | 101.2 | 101.1 | 99.6 | 100.75 | 100.9 | 101 | 100.98 | 99.84 | 100.5 | 100.8167 | 100.98 | 100.8
-1 | 100.125 | 101.1 | 101.225 | 101.25 | 101.15 | 99.75 | 101.1 | 101.2 | 101.1167 | 101.05 | 99.7 | 101 | 101 | 101.1 | 100.9833 | 99.9 | 100.7 | 100.8167 | 100.9 | 100.9 | 100.1 | 100.4 | 100.74 | 100.9 | 100.7
-0.9 | 100.3 | 101.075 | 101.1 | 101.1 | 101.08 | 100.02 | 101 | 101 | 101.1 | 100.9 | 99.9 | 100.9 | 100.9 | 101 | 100.8167 | 100 | 100.7 | 100.8 | 100.8 | 100.8 | 100.1 | 100.4 | 100.625 | 100.8167 | 100.7
-0.8 | 100.4 | 100.9 | 101 | 100.94 | 100.9 | 100.2 | 100.875 | 100.9 | 100.975 | 100.8333 | 100.1 | 100.8 | 100.8 | 100.9 | 100.7667 | 100.1 | 100.6 | 100.7 | 100.7333 | 100.72 | 100.1 | 100.4 | 100.6 | 100.8 | 100.5143
-0.7 | 100.4333 | 100.8 | 100.875 | 100.8 | 100.8 | 100.3 | 100.7 | 100.75 | 100.825 | 100.7667 | 100.22 | 100.8 | 100.7167 | 100.8 | 100.6 | 100.175 | 100.6 | 100.7 | 100.6333 | 100.65 | 100.1 | 100.35 | 100.5 | 100.7429 | 100.5
-0.6 | 100.5 | 100.7 | 100.7 | 100.7 | 100.6857 | 100.3 | 100.7 | 100.625 | 100.7833 | 100.7 | 100.3 | 100.625 | 100.54 | 100.6833 | 100.5 | 100.2 | 100.5 | 100.55 | 100.5 | 100.5 | 100.1 | 100.4 | 100.4 | 100.6 | 100.4
-0.5 | 100.4 | 100.5 | 100.6 | 100.6 | 100.6 | 100.3 | 100.6 | 100.5 | 100.7 | 100.5429 | 100.3 | 100.6 | 100.4667 | 100.5167 | 100.4167 | 100.2 | 100.4 | 100.48 | 100.4 | 100.4333 | 100.1 | 100.4 | 100.3 | 100.5167 | 100.3
-0.4 | 100.4 | 100.4 | 100.5 | 100.4333 | 100.45 | 100.3 | 100.5 | 100.4 | 100.5167 | 100.475 | 100.2 | 100.4 | 100.36 | 100.4 | 100.3333 | 100.1 | 100.3 | 100.3 | 100.4 | 100.3667 | 100.15 | 100.3 | 100.2 | 100.34 | 100.28
-0.3 | 100.3 | 100.3 | 100.4 | 100.34 | 100.325 | 100.2 | 100.4 | 100.4 | 100.4 | 100.25 | 100.2 | 100.3 | 100.2667 | 100.3 | 100.3 | 100.1 | 100.2 | 100.3 | 100.3 | 100.3 | 100.2 | 100.3 | 100.2 | 100.2 | 100.2
-0.2 | 100.2 | 100.2 | 100.2 | 100.26 | 100.25 | 100.2 | 100.2 | 100.275 | 100.2667 | 100.175 | 100.1 | 100.15 | 100.1667 | 100.2 | 100.24 | 100.1 | 100.18 | 100.2 | 100.2 | 100.2 | 100.15 | 100.225 | 100.1 | 100.2 | 100.1
-0.1 | 100.1 | 100.1 | 100.1 | 100.16 | 100.1667 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1 | 100.1667 | 100.1 | 100.05 | 100.1 | 100.1 | 100.1 | 100.14 | 100.1 | 100.075 | 100.1 | 100.1 | 100.1
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.9 | 99.9 | 100 | 99.8833 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.9 | 99.8167 | 99.9 | 99.9 | 100 | 99.9 | 99.9 | 99.9167 | 99.9 | 99.95 | 99.875 | 99.9 | 99.9 | 99.9
0.2 | 99.7667 | 99.7 | 99.825 | 99.8 | 99.8 | 99.8 | 99.8 | 99.8 | 99.8 | 99.8 | 99.775 | 99.8 | 99.68 | 99.8 | 99.75 | 99.74 | 99.8 | 99.8167 | 99.82 | 99.86 | 99.85 | 99.775 | 99.8833 | 99.8 | 99.8667
0.3 | 99.65 | 99.55 | 99.7 | 99.7 | 99.7 | 99.6 | 99.7 | 99.7 | 99.7 | 99.7167 | 99.625 | 99.7 | 99.6 | 99.72 | 99.6833 | 99.6 | 99.7 | 99.8 | 99.75 | 99.7667 | 99.7 | 99.6 | 99.78 | 99.8 | 99.8
0.4 | 99.575 | 99.5 | 99.6 | 99.6143 | 99.6 | 99.4 | 99.6 | 99.5 | 99.65 | 99.6 | 99.4 | 99.5 | 99.5 | 99.65 | 99.5 | 99.5 | 99.6 | 99.7 | 99.66 | 99.7 | 99.575 | 99.5 | 99.6 | 99.8 | 99.7
0.5 | 99.375 | 99.3 | 99.5 | 99.5 | 99.5 | 99.2 | 99.475 | 99.5 | 99.5333 | 99.48 | 99.3 | 99.4 | 99.4 | 99.5833 | 99.4 | 99.3 | 99.5 | 99.64 | 99.56 | 99.6 | 99.475 | 99.4 | 99.5143 | 99.7167 | 99.62
0.6 | 99.075 | 99.2 | 99.4 | 99.4 | 99.4167 | 99 | 99.3 | 99.4 | 99.44 | 99.4 | 99.1 | 99.3 | 99.3571 | 99.4857 | 99.3 | 99.1 | 99.5 | 99.5 | 99.45 | 99.5 | 99.4 | 99.3 | 99.4 | 99.7 | 99.5167
0.7 | 98.925 | 99.1 | 99.3 | 99.4 | 99.4 | 98.7 | 99.2 | 99.3 | 99.35 | 99.3 | 98.825 | 99.3 | 99.3 | 99.32 | 99.275 | 99 | 99.4 | 99.4143 | 99.4 | 99.4 | 99.275 | 99.2 | 99.3 | 99.6 | 99.4333
0.8 | 98.625 | 99 | 99.2 | 99.26 | 99.3 | 98.5 | 99.125 | 99.125 | 99.2571 | 99.2 | 98.5333 | 99.1 | 99.2 | 99.2286 | 99.15 | 98.74 | 99.325 | 99.3 | 99.3 | 99.3 | 99.1 | 99.1667 | 99.3 | 99.55 | 99.4
0.9 | 98.3333 | 98.9 | 99.1 | 99.18 | 99.2 | 98.1 | 99 | 99 | 99.18 | 99.1333 | 98.25 | 99 | 99.1 | 99.14 | 99.0333 | 98.4667 | 99.2 | 99.1 | 99.25 | 99.1833 | 98.7333 | 99.1 | 99.2333 | 99.3833 | 99.3
1 | 97.8 | 98.7 | 99 | 99.0833 | 99.04 | 97.5 | 98.9 | 98.9 | 99.06 | 99.1 | 97.8 | 98.825 | 98.9333 | 99.0667 | 98.96 | 98.075 | 99.075 | 99.06 | 99.1667 | 99.1 | 98.4333 | 99 | 99.1 | 99.3 | 99.2
1.1 | 97.3 | 98.6 | 98.9 | 99 | 98.94 | 96.9 | 98.8 | 98.8 | 98.94 | 98.9833 | 97.275 | 98.775 | 98.86 | 99 | 98.9 | 97.85 | 98.9 | 99 | 99.1 | 99 | 98.14 | 98.9 | 99.0667 | 99.2333 | 99.15
1.2 | 96.4 | 98.5 | 98.8 | 98.9 | 98.8667 | 96.14 | 98.6 | 98.74 | 98.85 | 98.9 | 96.65 | 98.7 | 98.7714 | 98.9 | 98.8 | 97.275 | 98.8 | 98.9333 | 98.95 | 98.9 | 97.85 | 98.7 | 99 | 99.1 | 99
1.3 | 95.35 | 98.3333 | 98.6 | 98.8 | 98.7143 | 95.1 | 98.5 | 98.64 | 98.8 | 98.8 | 95.8 | 98.5 | 98.6 | 98.8 | 98.7 | 96.82 | 98.6 | 98.84 | 98.86 | 98.85 | 97.5667 | 98.6 | 98.92 | 99 | 98.9
1.4 | 93.5 | 98.2 | 98.575 | 98.7 | 98.6 | 93.6 | 98.4 | 98.5 | 98.7 | 98.7 | 94.6 | 98.5 | 98.54 | 98.7 | 98.65 | 96.2 | 98.5 | 98.75 | 98.7 | 98.6857 | 97 | 98.5 | 98.8 | 98.9667 | 98.8333
1.5 | 91 | 98.2 | 98.4667 | 98.6 | 98.6 | 91.3 | 98.3 | 98.4333 | 98.6 | 98.6 | 93.125 | 98.325 | 98.45 | 98.6 | 98.5833 | 95.1667 | 98.375 | 98.6667 | 98.6833 | 98.6 | 96.3667 | 98.3 | 98.7 | 98.9 | 98.8
1.6 | 86.6 | 98.1 | 98.4 | 98.48 | 98.5 | 87.95 | 98.2 | 98.35 | 98.5 | 98.5 | 90.9 | 98.2 | 98.3833 | 98.5 | 98.4167 | 93.8 | 98.2333 | 98.6 | 98.6 | 98.5 | 95.875 | 98.175 | 98.55 | 98.8 | 98.7833
1.7 | 80.0333 | 97.975 | 98.275 | 98.4 | 98.4 | 82.66 | 98.0667 | 98.2 | 98.4 | 98.4 | 87.625 | 98.05 | 98.2167 | 98.4 | 98.4 | 92.15 | 98.08 | 98.4429 | 98.5 | 98.5 | 94.775 | 97.94 | 98.4833 | 98.7 | 98.7
1.8 | 70.4 | 97.8 | 98.125 | 98.3 | 98.3333 | 75.3 | 98 | 98.1 | 98.32 | 98.3 | 82.275 | 97.9 | 98.15 | 98.3167 | 98.3 | 89.82 | 97.95 | 98.3833 | 98.4 | 98.4143 | 93.475 | 97.8 | 98.4 | 98.6 | 98.5167
1.9 | 59.975 | 97.7 | 97.98 | 98.2 | 98.2 | 65.7333 | 97.75 | 98 | 98.25 | 98.2167 | 75.3 | 97.72 | 98 | 98.3 | 98.2 | 86.3 | 97.7 | 98.2167 | 98.2333 | 98.3 | 91.8 | 97.7 | 98.3 | 98.5167 | 98.45
2 | 48.1333 | 97.6 | 97.9 | 98.15 | 98.1286 | 55.5667 | 97.6 | 97.9 | 98.08 | 98.1 | 65.85 | 97.5333 | 97.9333 | 98.12 | 98.0833 | 81.0667 | 97.525 | 98.125 | 98.0667 | 98.2 | 89 | 97.575 | 98.1667 | 98.3 | 98.3667
2.1 | 36.75 | 97.5 | 97.8 | 98 | 98.0667 | 43.9 | 97.45 | 97.9 | 97.9833 | 98 | 55.7667 | 97.4 | 97.8 | 98.0167 | 98 | 74.6667 | 97.36 | 98 | 98 | 98.1 | 85.7667 | 97.4667 | 98.1 | 98.3 | 98.3
2.2 | 26.5 | 97.3 | 97.7 | 98 | 97.88 | 33.3 | 97.3 | 97.6333 | 97.9 | 97.9 | 45.45 | 97.2 | 97.7 | 97.9667 | 97.8333 | 66.525 | 97.12 | 97.8833 | 97.875 | 97.9833 | 81 | 97.3333 | 97.9333 | 98.2 | 98.12
2.3 | 18.44 | 97.2 | 97.6 | 97.8571 | 97.7833 | 24.42 | 97.12 | 97.5 | 97.8 | 97.8 | 34.9667 | 96.975 | 97.6 | 97.8 | 97.7571 | 56.9 | 97 | 97.8 | 97.8 | 97.9 | 74.4667 | 97.15 | 97.8 | 98.1857 | 98
2.4 | 12.6 | 97.1 | 97.5 | 97.7 | 97.7 | 17 | 96.975 | 97.5 | 97.6857 | 97.7 | 26.1 | 96.825 | 97.4667 | 97.7333 | 97.7 | 46.35 | 96.875 | 97.7 | 97.7 | 97.8 | 66.6 | 97 | 97.6714 | 98.1 | 97.9
2.5 | 8.7667 | 96.9333 | 97.4 | 97.6 | 97.6 | 12.5 | 96.8 | 97.375 | 97.6 | 97.7 | 19.2667 | 96.64 | 97.35 | 97.66 | 97.5 | 37.0667 | 96.55 | 97.52 | 97.6 | 97.6667 | 57.2 | 96.775 | 97.5167 | 97.95 | 97.8167
2.6 | 6.7 | 96.8 | 97.32 | 97.5 | 97.5 | 9.3333 | 96.6333 | 97.25 | 97.55 | 97.52 | 14.48 | 96.4 | 97.2667 | 97.46 | 97.4 | 28.675 | 96.3333 | 97.4 | 97.48 | 97.56 | 47.75 | 96.575 | 97.45 | 97.82 | 97.7333
2.7 | 5.25 | 96.6 | 97.2 | 97.4 | 97.4167 | 7.3667 | 96.4 | 97.1 | 97.46 | 97.4 | 11.4333 | 96.25 | 97.1 | 97.4 | 97.35 | 21.7667 | 96.1 | 97.3 | 97.38 | 97.4833 | 38.4 | 96.4 | 97.375 | 97.7 | 97.6
2.8 | 4.1 | 96.4 | 97.1 | 97.3 | 97.34 | 6.05 | 96.175 | 97 | 97.35 | 97.375 | 9.25 | 95.9333 | 96.95 | 97.3 | 97.3 | 16.7 | 96 | 97.18 | 97.28 | 97.3167 | 30.1 | 96.1 | 97.3 | 97.7 | 97.5
2.9 | 3.6333 | 96.3 | 96.9 | 97.2333 | 97.3 | 5.0667 | 96 | 96.9 | 97.2 | 97.2 | 7.7 | 95.7667 | 96.8 | 97.22 | 97.2 | 13.625 | 95.9 | 97.1 | 97.18 | 97.3 | 23.6333 | 95.92 | 97.2 | 97.5667 | 97.4
3 | 3.1 | 96.1 | 96.8 | 97.1 | 97.2 | 4.35 | 95.84 | 96.9 | 97.1 | 97.1 | 6.7 | 95.6 | 96.7 | 97.15 | 97.1167 | 11.3 | 95.6 | 96.9429 | 97.1 | 97.2 | 19.025 | 95.675 | 97.1 | 97.4 | 97.3833
3.1 | 2.8 | 96 | 96.7 | 97 | 97.1 | 3.8667 | 95.8 | 96.7333 | 97 | 97 | 5.9 | 95.3 | 96.6167 | 97 | 97 | 9.65 | 95.3 | 96.78 | 96.9333 | 97.1 | 15.475 | 95.5 | 96.8 | 97.3 | 97.3
3.2 | 2.5 | 95.7667 | 96.7 | 96.9 | 96.9667 | 3.5 | 95.5333 | 96.6 | 96.9167 | 96.98 | 5.25 | 95.1 | 96.5 | 96.9 | 96.9 | 8.6 | 95.05 | 96.6 | 96.9 | 96.9714 | 13.2 | 95.35 | 96.7 | 97.2 | 97.2286
3.3 | 2.3 | 95.52 | 96.6 | 96.84 | 96.9 | 3.1 | 95.3 | 96.5 | 96.84 | 96.9 | 4.75 | 94.95 | 96.3714 | 96.8 | 96.8 | 7.6667 | 94.9 | 96.4833 | 96.72 | 96.8 | 11.3667 | 95.1333 | 96.6 | 97.18 | 97.1667
3.4 | 2.1 | 95.4 | 96.5 | 96.7 | 96.8 | 2.8 | 95.05 | 96.3 | 96.725 | 96.7571 | 4.42 | 94.5667 | 96.2 | 96.6833 | 96.6333 | 6.9333 | 94.62 | 96.4 | 96.6333 | 96.65 | 10.1333 | 94.9333 | 96.5 | 96.9833 | 96.9667
3.5 | 2 | 95.15 | 96.4 | 96.6 | 96.65 | 2.625 | 94.7 | 96.2 | 96.62 | 96.6 | 4.075 | 94.4 | 96.1 | 96.6 | 96.6 | 6.5 | 94.3667 | 96.2167 | 96.5 | 96.5667 | 9.175 | 94.6667 | 96.4 | 96.9 | 96.8
3.6 | 1.9 | 94.9 | 96.2 | 96.56 | 96.5667 | 2.5 | 94.475 | 96.1 | 96.55 | 96.5 | 3.85 | 94 | 96 | 96.5 | 96.5 | 6 | 94 | 96.1 | 96.5 | 96.4 | 8.4333 | 94.4 | 96.2 | 96.75 | 96.7833
3.7 | 1.8 | 94.7 | 96.1 | 96.46 | 96.4 | 2.3 | 94.225 | 95.975 | 96.4 | 96.3667 | 3.575 | 93.7333 | 95.8333 | 96.3667 | 96.4 | 5.6333 | 93.725 | 95.9 | 96.35 | 96.35 | 7.8 | 94.1333 | 96.1333 | 96.7 | 96.7
3.8 | 1.7 | 94.35 | 96 | 96.4 | 96.3 | 2.15 | 93.775 | 95.825 | 96.3 | 96.3 | 3.375 | 93.425 | 95.6833 | 96.28 | 96.2 | 5.4 | 93.475 | 95.8 | 96.15 | 96.2667 | 7.4 | 93.8333 | 95.95 | 96.575 | 96.6
3.9 | 1.6 | 94 | 95.8 | 96.2 | 96.1833 | 2.05 | 93.45 | 95.7 | 96.2 | 96.2 | 3.2 | 93.06 | 95.5167 | 96.15 | 96.1 | 5.1 | 93.24 | 95.6333 | 96.06 | 96.18 | 6.85 | 93.575 | 95.8833 | 96.42 | 96.5
4 | 1.525 | 93.5667 | 95.7667 | 96.1 | 96.1 | 2 | 93.15 | 95.6 | 96.0167 | 96.1 | 3.1 | 92.7333 | 95.45 | 96.1 | 95.9667 | 4.96 | 92.875 | 95.5667 | 95.9 | 96.1 | 6.7 | 93.3 | 95.8 | 96.3167 | 96.34
4.1 | 1.5 | 93 | 95.7 | 96 | 96.0167 | 1.9 | 92.75 | 95.5 | 95.9333 | 96 | 2.975 | 92.35 | 95.3833 | 95.9667 | 95.8 | 4.675 | 92.525 | 95.4143 | 95.8857 | 95.9167 | 6.35 | 92.925 | 95.6 | 96.1667 | 96.1667
4.2 | 1.4 | 92.4 | 95.55 | 95.9 | 95.92 | 1.8 | 92.075 | 95.4 | 95.9 | 95.9 | 2.8 | 91.85 | 95.2 | 95.8 | 95.7 | 4.5 | 92.2 | 95.3 | 95.8 | 95.75 | 6.1 | 92.675 | 95.5 | 96.1 | 96
4.3 | 1.4 | 91.425 | 95.4 | 95.8 | 95.8571 | 1.7 | 91.6 | 95.3 | 95.8 | 95.84 | 2.7 | 91.3 | 95.1 | 95.6333 | 95.5667 | 4.3 | 91.8 | 95.15 | 95.6667 | 95.5857 | 5.975 | 92.375 | 95.3 | 96 | 95.9
4.4 | 1.3 | 90.325 | 95.325 | 95.725 | 95.7 | 1.68 | 90.7 | 95.15 | 95.7167 | 95.68 | 2.6 | 90.8 | 94.9286 | 95.5 | 95.5 | 4.24 | 91.4 | 95 | 95.5 | 95.5 | 5.8 | 91.975 | 95.1571 | 95.8333 | 95.8
4.5 | 1.3 | 88.5667 | 95.2 | 95.6 | 95.66 | 1.6 | 89.4 | 95 | 95.5 | 95.56 | 2.5 | 90.16 | 94.7667 | 95.4 | 95.42 | 4.1 | 91.1 | 94.8667 | 95.4 | 95.4 | 5.5333 | 91.7 | 95 | 95.7 | 95.62
4.6 | 1.2 | 85.7 | 95.1 | 95.5 | 95.5 | 1.5 | 87.85 | 94.925 | 95.46 | 95.48 | 2.375 | 89.35 | 94.5667 | 95.3 | 95.3 | 3.95 | 90.6 | 94.6167 | 95.24 | 95.4 | 5.275 | 91.38 | 94.9 | 95.6 | 95.5
4.7 | 1.2 | 80.75 | 95 | 95.4 | 95.4333 | 1.5 | 85.4 | 94.8 | 95.3 | 95.3 | 2.3 | 88.2 | 94.46 | 95.2 | 95.14 | 3.8 | 90 | 94.54 | 95.2 | 95.225 | 5.15 | 91.1 | 94.7333 | 95.4167 | 95.4
4.8 | 1.2 | 73.5667 | 94.8 | 95.2167 | 95.3333 | 1.4 | 82.1 | 94.66 | 95.15 | 95.2 | 2.2 | 86.9 | 94.3667 | 95.1429 | 95 | 3.6333 | 89.55 | 94.35 | 95.0333 | 95.14 | 5 | 90.7333 | 94.5 | 95.32 | 95.3
4.9 | 1.1 | 64.7333 | 94.7 | 95.2 | 95.2 | 1.4 | 77.3 | 94.5 | 95.05 | 95.1 | 2.1 | 84.9 | 94.2429 | 95 | 94.9167 | 3.5 | 88.95 | 94.2 | 94.86 | 94.96 | 4.9 | 90.325 | 94.3333 | 95.2 | 95.15
5 | 1.1 | 54.6 | 94.5 | 95.1 | 95.1 | 1.3 | 69.5 | 94.325 | 94.9 | 95 | 2.075 | 82 | 94.0167 | 94.9286 | 94.85 | 3.4 | 88.125 | 94.1 | 94.7 | 94.8 | 4.775 | 89.75 | 94.16 | 95.0714 | 95
5.1 | 1.1 | 43.9 | 94.4 | 95.0167 | 95.0167 | 1.3 | 60.94 | 94.2 | 94.8 | 94.8 | 2 | 77.625 | 93.94 | 94.7 | 94.6667 | 3.3 | 87.375 | 93.8571 | 94.6 | 94.7 | 4.625 | 89.4 | 94 | 95 | 94.9167
5.2 | 1.1 | 33.5 | 94.3667 | 94.85 | 94.82 | 1.3 | 51.05 | 94.1 | 94.7 | 94.7167 | 1.9 | 71.65 | 93.8 | 94.6 | 94.575 | 3.2 | 86.2 | 93.7 | 94.55 | 94.6 | 4.5 | 88.9 | 93.8667 | 94.85 | 94.8
5.3 | 1 | 24.6 | 94.2 | 94.8 | 94.75 | 1.2 | 40.6 | 93.9 | 94.5667 | 94.65 | 1.9 | 64.2 | 93.7 | 94.5333 | 94.4833 | 3.1 | 84.7667 | 93.6 | 94.32 | 94.5 | 4.38 | 88.275 | 93.7 | 94.7857 | 94.7
5.4 | 1 | 17.9 | 94.025 | 94.7 | 94.64 | 1.2 | 31.6 | 93.8 | 94.4667 | 94.5 | 1.8 | 55.08 | 93.56 | 94.4 | 94.4 | 3 | 82.9 | 93.35 | 94.2 | 94.4 | 4.2 | 87.5667 | 93.6 | 94.6333 | 94.6
5.5 | 1 | 13.4333 | 93.875 | 94.5333 | 94.5 | 1.12 | 24.25 | 93.7 | 94.36 | 94.4 | 1.725 | 45.8 | 93.4 | 94.3 | 94.3 | 3 | 80.2 | 93.25 | 94.1 | 94.25 | 4.1667 | 86.95 | 93.45 | 94.5 | 94.55
5.6 | 1 | 10.7 | 93.76 | 94.4 | 94.4 | 1.1 | 18.18 | 93.6 | 94.2667 | 94.3429 | 1.7 | 36.9 | 93.2 | 94.2 | 94.14 | 2.9 | 76.5 | 93.0833 | 94 | 94.08 | 4.1 | 86.1 | 93.24 | 94.3667 | 94.3833
5.7 | 0.9 | 8.95 | 93.6 | 94.24 | 94.2571 | 1.1 | 14.1 | 93.4333 | 94.1 | 94.16 | 1.6 | 28.7 | 93.1 | 94.06 | 94 | 2.8 | 71.4 | 92.9167 | 93.8833 | 93.98 | 3.9333 | 84.9 | 93.04 | 94.2 | 94.3
5.8 | 0.9 | 7.725 | 93.6 | 94.1 | 94.2 | 1.1 | 11.7333 | 93.2 | 94 | 94 | 1.6 | 22.5 | 92.92 | 93.8 | 93.9 | 2.7 | 64.3667 | 92.7 | 93.7 | 93.9 | 3.9 | 83.4 | 92.8667 | 94.0429 | 94.06
5.9 | 0.9 | 6.925 | 93.4 | 94.0143 | 94.0333 | 1.025 | 10.0667 | 93.05 | 93.9 | 93.9 | 1.5 | 17.8 | 92.7429 | 93.7 | 93.6833 | 2.7 | 56.3333 | 92.4833 | 93.6 | 93.8143 | 3.8 | 81.6333 | 92.8 | 93.8833 | 93.94
6 | 0.9 | 6.325 | 93.2 | 93.9 | 93.9 | 1 | 8.9 | 92.92 | 93.85 | 93.8333 | 1.5 | 15 | 92.5833 | 93.7 | 93.5833 | 2.6 | 47.98 | 92.3167 | 93.475 | 93.6 | 3.7 | 78.7 | 92.5714 | 93.8 | 93.84
6.1 | 0.9 | 5.9 | 93.04 | 93.72 | 93.78 | 1 | 7.9667 | 92.775 | 93.75 | 93.64 | 1.5 | 13.175 | 92.4167 | 93.5714 | 93.4286 | 2.5 | 40.18 | 92.15 | 93.24 | 93.4 | 3.6 | 74.95 | 92.3333 | 93.65 | 93.75
6.2 | 0.9 | 5.48 | 92.95 | 93.6429 | 93.7 | 1 | 7.25 | 92.6 | 93.58 | 93.6 | 1.4 | 11.775 | 92.2167 | 93.4 | 93.3 | 2.5 | 32.9 | 91.8857 | 93.1667 | 93.3 | 3.5 | 69.8 | 92.1667 | 93.56 | 93.6667
6.3 | 0.9 | 5.2 | 92.75 | 93.5833 | 93.6 | 1 | 6.7 | 92.5 | 93.3833 | 93.5 | 1.4 | 10.6 | 92.02 | 93.2286 | 93.2 | 2.4 | 26.575 | 91.72 | 93.0833 | 93.1333 | 3.5 | 63.7 | 92.0143 | 93.44 | 93.5
6.4 | 0.875 | 4.9667 | 92.7 | 93.5 | 93.45 | 0.9 | 6.3 | 92.35 | 93.275 | 93.4 | 1.4 | 9.875 | 91.9 | 93.14 | 93.14 | 2.36 | 22.1667 | 91.6 | 92.9 | 93.025 | 3.375 | 56.2 | 91.7429 | 93.25 | 93.3286
6.5 | 0.8 | 4.7667 | 92.5 | 93.36 | 93.3 | 0.9 | 5.975 | 92.2 | 93.15 | 93.3 | 1.3 | 9.15 | 91.75 | 92.95 | 92.94 | 2.3 | 18.975 | 91.3333 | 92.8 | 92.9 | 3.3 | 48.8 | 91.575 | 93.1 | 93.1667
6.6 | 0.8 | 4.6 | 92.4 | 93.25 | 93.2167 | 0.9 | 5.75 | 92.1333 | 93.05 | 93.12 | 1.3 | 8.725 | 91.6 | 92.85 | 92.8571 | 2.2 | 16.875 | 91.1 | 92.68 | 92.75 | 3.2 | 41.45 | 91.3333 | 93 | 93
6.7 | 0.8 | 4.4 | 92.25 | 93.12 | 93.05 | 0.9 | 5.425 | 91.875 | 92.9 | 93 | 1.3 | 8.3 | 91.4333 | 92.7 | 92.7 | 2.2 | 15.4 | 90.92 | 92.575 | 92.66 | 3.1 | 34.54 | 91.1 | 92.9 | 92.8167
6.8 | 0.8 | 4.3 | 92.1 | 92.96 | 92.96 | 0.9 | 5.225 | 91.75 | 92.7667 | 92.8667 | 1.3 | 7.9 | 91.16 | 92.7 | 92.58 | 2.1 | 14.2 | 90.74 | 92.3857 | 92.46 | 3.0333 | 29 | 90.9167 | 92.7 | 92.75
6.9 | 0.8 | 4.175 | 92 | 92.86 | 92.9 | 0.9 | 5.05 | 91.575 | 92.68 | 92.7143 | 1.2 | 7.55 | 90.95 | 92.52 | 92.4 | 2.1 | 13.35 | 90.54 | 92.2333 | 92.35 | 3 | 24.6 | 90.75 | 92.5667 | 92.5833
7 | 0.8 | 4 | 91.8 | 92.7 | 92.74 | 0.9 | 4.8 | 91.3 | 92.5 | 92.55 | 1.2 | 7.2 | 90.8 | 92.3333 | 92.3167 | 2.075 | 12.7 | 90.3833 | 92.1 | 92.24 | 3 | 21.625 | 90.6 | 92.4 | 92.4167
7.1 | 0.8 | 3.9 | 91.7 | 92.6 | 92.66 | 0.8 | 4.65 | 91.25 | 92.4 | 92.4 | 1.2 | 6.95 | 90.6333 | 92.1714 | 92.2 | 2 | 12.0667 | 90.2167 | 91.96 | 92.1 | 2.9 | 19.5 | 90.3167 | 92.25 | 92.225
7.2 | 0.8 | 3.775 | 91.575 | 92.4667 | 92.4667 | 0.8 | 4.5 | 91.05 | 92.3 | 92.3 | 1.2 | 6.7 | 90.4 | 92 | 92.1 | 2 | 11.6 | 90 | 91.8 | 91.9833 | 2.8 | 17.9 | 90.15 | 92.16 | 92.1
7.3 | 0.8 | 3.7 | 91.34 | 92.3143 | 92.4 | 0.8 | 4.375 | 90.85 | 92.1833 | 92.2 | 1.1 | 6.5 | 90.2 | 91.98 | 91.85 | 1.9 | 11.2 | 89.7667 | 91.64 | 91.8 | 2.7 | 16.72 | 89.8833 | 92.08 | 92.0167
7.4 | 0.8 | 3.6 | 91.2 | 92.125 | 92.2333 | 0.8 | 4.175 | 90.675 | 92.1 | 92.1 | 1.1 | 6.3 | 90.0167 | 91.8 | 91.7 | 1.9 | 10.8667 | 89.5167 | 91.5 | 91.62 | 2.7 | 15.8 | 89.7 | 91.7833 | 91.8333
7.5 | 0.7 | 3.5 | 91.025 | 92.1 | 92.16 | 0.8 | 4.08 | 90.5 | 91.9 | 91.9 | 1.1 | 6.1 | 89.8429 | 91.7167 | 91.6333 | 1.8 | 10.4 | 89.35 | 91.4 | 91.5 | 2.6 | 15.15 | 89.4571 | 91.6 | 91.6667
7.6 | 0.7 | 3.4 | 90.9 | 92 | 92 | 0.8 | 4 | 90.3 | 91.7833 | 91.7833 | 1.1 | 5.85 | 89.575 | 91.55 | 91.5 | 1.8 | 10.15 | 89.1714 | 91.28 | 91.3 | 2.6 | 14.4667 | 89.3 | 91.4833 | 91.5833
7.7 | 0.7 | 3.3 | 90.65 | 91.9 | 91.9143 | 0.8 | 3.85 | 90.2 | 91.6286 | 91.6 | 1.1 | 5.7 | 89.3857 | 91.3857 | 91.4 | 1.8 | 9.925 | 88.9333 | 91.16 | 91.225 | 2.6 | 13.9 | 89.06 | 91.3167 | 91.4167
7.8 | 0.7 | 3.2 | 90.5 | 91.75 | 91.8167 | 0.8 | 3.7 | 90 | 91.4571 | 91.55 | 1.1 | 5.56 | 89.22 | 91.2333 | 91.24 | 1.7 | 9.5 | 88.6333 | 90.96 | 91.08 | 2.6 | 13.5 | 88.85 | 91.15 | 91.24
7.9 | 0.7 | 3.175 | 90.4 | 91.6 | 91.625 | 0.8 | 3.6 | 89.88 | 91.4 | 91.4 | 1 | 5.375 | 88.9571 | 91.1333 | 91 | 1.7 | 9.3 | 88.4333 | 90.8 | 90.9 | 2.5 | 13.2 | 88.7 | 91.1 | 91.1
8 | 0.7 | 3.1 | 90.2 | 91.4833 | 91.5 | 0.8 | 3.6 | 89.65 | 91.2 | 91.275 | 1 | 5.2333 | 88.8 | 90.94 | 90.9 | 1.7 | 9.1 | 88.2333 | 90.7167 | 90.8 | 2.5 | 12.7 | 88.4333 | 90.96 | 91
8.1 | 0.7 | 3.08 | 89.975 | 91.4 | 91.325 | 0.8 | 3.5 | 89.4 | 91.1 | 91.16 | 1 | 5.15 | 88.54 | 90.8 | 90.8 | 1.7 | 8.775 | 87.98 | 90.5 | 90.6333 | 2.4 | 12.375 | 88.2833 | 90.7667 | 90.9
8.2 | 0.7 | 3 | 89.8 | 91.26 | 91.2 | 0.7 | 3.4 | 89.3 | 91 | 91 | 1 | 4.92 | 88.3 | 90.7333 | 90.7 | 1.7 | 8.6 | 87.85 | 90.3 | 90.4333 | 2.3 | 12.125 | 88.0333 | 90.6 | 90.72
8.3 | 0.7 | 2.9 | 89.625 | 91.1667 | 91.1167 | 0.7 | 3.3 | 89 | 90.8 | 90.95 | 1 | 4.8 | 88.1 | 90.54 | 90.575 | 1.6 | 8.4 | 87.5667 | 90.22 | 90.24 | 2.3 | 11.775 | 87.75 | 90.4 | 90.54
8.4 | 0.7 | 2.825 | 89.425 | 91 | 91.0333 | 0.7 | 3.2 | 88.8 | 90.7 | 90.8 | 1 | 4.7 | 87.9 | 90.45 | 90.3667 | 1.6 | 8.26 | 87.26 | 89.98 | 90.0833 | 2.275 | 11.625 | 87.54 | 90.3 | 90.34
8.5 | 0.7 | 2.8 | 89.225 | 90.8333 | 90.8667 | 0.7 | 3.1 | 88.55 | 90.6333 | 90.7 | 0.95 | 4.55 | 87.62 | 90.36 | 90.3 | 1.6 | 8 | 87 | 89.88 | 90 | 2.2 | 11.35 | 87.1 | 90.12 | 90.125
8.6 | 0.7 | 2.7 | 89 | 90.7 | 90.7 | 0.7 | 3.025 | 88.32 | 90.5 | 90.6 | 0.9 | 4.45 | 87.42 | 90.1833 | 90.1333 | 1.5 | 7.825 | 86.8667 | 89.7 | 89.9 | 2.175 | 11.025 | 87 | 90 | 90.04
8.7 | 0.7 | 2.7 | 88.7667 | 90.62 | 90.6 | 0.7 | 3 | 88.1 | 90.4 | 90.4333 | 0.9 | 4.4 | 87.1167 | 90.0167 | 89.9833 | 1.5 | 7.6 | 86.55 | 89.5333 | 89.8 | 2.1 | 10.925 | 86.775 | 89.8 | 89.85
8.8 | 0.7 | 2.675 | 88.5 | 90.5429 | 90.5 | 0.7 | 2.9 | 87.825 | 90.2 | 90.2333 | 0.9 | 4.2667 | 86.8667 | 89.85 | 89.8167 | 1.5 | 7.4667 | 86.2667 | 89.38 | 89.5 | 2.1 | 10.6 | 86.55 | 89.625 | 89.7
8.9 | 0.7 | 2.6 | 88.3 | 90.46 | 90.475 | 0.7 | 2.9 | 87.575 | 90.1 | 90.15 | 0.9 | 4.1 | 86.5667 | 89.7 | 89.7 | 1.5 | 7.2667 | 85.9333 | 89.2667 | 89.3857 | 2.1 | 10.4 | 86.24 | 89.46 | 89.575
9 | 0.7 | 2.6 | 87.9 | 90.325 | 90.4 | 0.7 | 2.8 | 87.25 | 89.98 | 90 | 0.9 | 4 | 86.3 | 89.6 | 89.5667 | 1.5 | 7.1 | 85.66 | 89.1143 | 89.22 | 2.1 | 10.15 | 86.1 | 89.3 | 89.4
9.1 |  | 2.5 | 87.6333 | 90.3 | 90.3 | 0.7 | 2.7667 | 86.9333 | 89.9 | 89.9 | 0.9 | 3.9 | 86 | 89.4857 | 89.5 | 1.4333 | 7 | 85.3667 | 88.95 | 89.12 | 2 | 9.94 | 85.75 | 89.2 | 89.3
9.2 |  | 2.5 | 87.2333 | 90.2 | 90.2 | 0.7 | 2.7 | 86.5667 | 89.8 | 89.8 | 0.9 | 3.8333 | 85.675 | 89.32 | 89.3333 | 1.4 | 6.8 | 85.16 | 88.78 | 88.925 | 2 | 9.8 | 85.56 | 89.1 | 89.15
9.3 |  | 2.4 | 86.74 | 90.15 | 90.1 | 0.7 | 2.6 | 86.2 | 89.7 | 89.68 | 0.9 | 3.75 | 85.45 | 89.2 | 89.1667 | 1.4 | 6.7 | 84.86 | 88.6833 | 88.75 | 2 | 9.6 | 85.3667 | 88.9 | 88.9714
9.4 |  | 2.4 | 86.075 | 90.1 | 90.04 | 0.7 | 2.6 | 85.8 | 89.6 | 89.6 | 0.9 | 3.7 | 85.08 | 89.1 | 89.0143 | 1.4 | 6.4333 | 84.56 | 88.5 | 88.66 | 2 | 9.4 | 85.0333 | 88.72 | 88.8167
9.5 |  | 2.4 | 85.2 | 90.0167 | 90 |  | 2.525 | 85.3 | 89.5 | 89.5 | 0.8 | 3.6 | 84.68 | 88.9 | 88.9 | 1.4 | 6.3 | 84.2286 | 88.36 | 88.375 | 1.9 | 9.225 | 84.8 | 88.54 | 88.6667
9.6 |  | 2.3 | 83.7 | 90 | 89.925 |  | 2.5 | 84.6 | 89.44 | 89.5 | 0.8 | 3.5667 | 84.28 | 88.8 | 88.8 | 1.32 | 6.2 | 84 | 88.1833 | 88.14 | 1.9 | 9.075 | 84.5 | 88.35 | 88.5
9.7 |  | 2.3 | 81.5 | 89.9 | 89.9 |  | 2.45 | 83.8 | 89.35 | 89.4 | 0.8 | 3.5 | 83.88 | 88.7429 | 88.6667 | 1.3 | 6.075 | 83.8 | 88.0286 | 88.1 | 1.9 | 8.825 | 84.1833 | 88.175 | 88.3
9.8 |  | 2.3 | 77.8 | 89.9 | 89.9 |  | 2.4 | 82.8333 | 89.3 | 89.3 | 0.8 | 3.4 | 83.4667 | 88.5833 | 88.5 | 1.3 | 5.925 | 83.45 | 87.8667 | 88 | 1.9 | 8.7 | 83.8833 | 88 | 88.1167
9.9 |  | 2.2 | 72.4 | 89.9 | 89.82 |  | 2.4 | 81.2 | 89.2 | 89.2 |  | 3.375 | 82.9 | 88.5 | 88.4 | 1.3 | 5.8 | 83.075 | 87.6667 | 87.8833 | 1.8 | 8.55 | 83.6333 | 87.85 | 88
10 |  | 2.2 | 64.9 | 89.8 | 89.8 |  | 2.3 | 79.25 | 89.2 | 89.1833 |  | 3.3 | 82.4 | 88.4 | 88.3714 | 1.3 | 5.7 | 82.6714 | 87.6 | 87.7167 | 1.8 | 8.4 | 83.325 | 87.6833 | 87.78
10.1 |  | 2.2 | 55.3 | 89.8 | 89.7 |  | 2.3 | 75.7 | 89.1 | 89.1 |  | 3.2 | 81.7333 | 88.35 | 88.3 | 1.2 | 5.6 | 82.34 | 87.45 | 87.54 | 1.8 | 8.2667 | 83.1 | 87.5286 | 87.6
10.2 |  | 2.2 | 45.8 | 89.75 | 89.6667 |  | 2.3 | 70.4 | 89 | 89 |  | 3.1 | 80.8667 | 88.2 | 88.2 | 1.2 | 5.5 | 82.1 | 87.2333 | 87.3667 | 1.7 | 8 | 82.9333 | 87.36 | 87.42
10.3 |  | 2.1 | 36.5 | 89.66 | 89.6 |  | 2.2 | 62.875 | 89 | 89 |  | 3.1 | 79.8 | 88.14 | 88.1 | 1.2 | 5.4 | 81.7333 | 87 | 87.2 | 1.7 | 7.8333 | 82.5 | 87.26 | 87.3333
10.4 |  | 2.1 | 28.3 | 89.6 | 89.6 |  | 2.2 | 54.65 | 89 | 89 |  | 3.075 | 78.2 | 88.04 | 88 | 1.2 | 5.3 | 81.4 | 86.9 | 87.1 | 1.7 | 7.775 | 82.22 | 87.0667 | 87.1
10.5 |  | 2.1 | 21.5 | 89.54 | 89.6 |  | 2.2 | 45.6 | 88.96 | 88.9 |  | 3 | 76 | 87.94 | 87.92 | 1.2 | 5.2 | 80.9 | 86.8 | 86.92 | 1.6 | 7.6 | 81.8667 | 86.8333 | 86.84
10.6 |  | 2.1 | 16.7333 | 89.5 | 89.5 |  | 2.1 | 37.3 | 88.8667 | 88.82 |  | 2.9 | 72.6667 | 87.85 | 87.8 | 1.2 | 5.0667 | 80.48 | 86.65 | 86.7 | 1.6 | 7.48 | 81.525 | 86.6 | 86.8
10.7 |  | 2 | 13.9 | 89.5 | 89.5 |  | 2.1 | 29.3 | 88.8 | 88.8 |  | 2.9 | 67.775 | 87.72 | 87.7 | 1.2 | 5 | 80.0333 | 86.5833 | 86.5286 | 1.6 | 7.35 | 81.28 | 86.5 | 86.7
10.8 |  | 2 | 12.2667 | 89.5 | 89.4 |  | 2.1 | 22.95 | 88.7 | 88.75 |  | 2.8 | 61.1 | 87.6333 | 87.6833 |  | 4.9 | 79.6 | 86.4167 | 86.36 | 1.6 | 7.2 | 80.925 | 86.32 | 86.5333
10.9 |  | 2 | 11.2 | 89.5 | 89.4 |  | 2.1 | 18.25 | 88.6667 | 88.7 |  | 2.8 | 54.1 | 87.6 | 87.5 |  | 4.8 | 79.12 | 86.3167 | 86.3 | 1.5 | 7.1 | 80.5 | 86.14 | 86.34
11 |  | 2 | 10.4 | 89.5 | 89.35 |  | 2 | 15.5 | 88.6 | 88.6167 |  | 2.75 | 45.8 | 87.5 | 87.5 |  | 4.7 | 78.46 | 86.1167 | 86.2 | 1.5 | 6.94 | 80.2667 | 86.05 | 86.1429
11.1 |  | 1.9333 | 9.8 | 89.4 | 89.2857 |  | 2 | 13.7 | 88.5 | 88.5 |  | 2.7 | 37.9667 | 87.44 | 87.3333 |  | 4.7 | 77.82 | 85.92 | 86.05 | 1.5 | 6.85 | 79.85 | 85.85 | 85.9167
11.2 |  | 1.9 | 9.35 | 89.4 | 89.2 |  | 2 | 12.4 | 88.5 | 88.48 |  | 2.7 | 30.66 | 87.4 | 87.3 |  | 4.6 | 76.95 | 85.8667 | 85.95 | 1.5 | 6.8 | 79.52 | 85.6833 | 85.8571
11.3 |  | 1.9 | 8.9667 | 89.3 | 89.2 |  | 2 | 11.45 | 88.4 | 88.3833 |  | 2.6 | 24.92 | 87.2 | 87.2143 |  | 4.5 | 75.85 | 85.7 | 85.7833 | 1.5 | 6.6 | 79.1333 | 85.5 | 85.6
11.4 |  | 1.9 | 8.6333 | 89.2 | 89.1 |  | 1.925 | 10.8333 | 88.4 | 88.3 |  | 2.6 | 20.7 | 87.1167 | 87.1 |  | 4.4 | 74.2667 | 85.6 | 85.6286 | 1.5 | 6.5 | 78.8 | 85.3167 | 85.4429
11.5 |  | 1.9 | 8.3667 | 89.1333 | 89.04 |  | 1.9 | 10.25 | 88.3 | 88.2 |  | 2.55 | 18.2 | 87.0571 | 87 |  | 4.3 | 72.05 | 85.48 | 85.55 | 1.5 | 6.4 | 78.46 | 85.1667 | 85.36
11.6 |  | 1.825 | 8.2 | 89.1 | 89 |  | 1.9 | 9.85 | 88.22 | 88.2 |  | 2.5 | 16.3667 | 87 | 86.9 |  | 4.2 | 69.15 | 85.4 | 85.3667 | 1.5 | 6.3 | 77.9 | 85.1 | 85.16
11.7 |  | 1.8 | 8 | 89 | 89 |  | 1.9 | 9.475 | 88.2 | 88.12 |  | 2.5 | 15.02 | 86.9 | 86.8 |  | 4.2 | 64.72 | 85.2 | 85.3 |  | 6.2 | 77.4857 | 84.9333 | 84.98
11.8 |  | 1.8 | 7.8 | 89 | 88.9 |  | 1.9 | 9.15 | 88.1 | 88.1 |  | 2.5 | 14.05 | 86.8 | 86.7286 |  | 4.1 | 59.7 | 85.1143 | 85.1 |  | 6.1 | 77.02 | 84.7667 | 84.82
11.9 |  | 1.8 | 7.65 | 88.95 | 88.9 |  | 1.8 | 8.875 | 88 | 88 |  | 2.4 | 13.28 | 86.7667 | 86.7 |  | 4 | 53.2 | 85 | 85.0667 |  | 6 | 76.55 | 84.6714 | 84.6333
12 |  | 1.8 | 7.5 | 88.9 | 88.85 |  | 1.8 | 8.6 | 88 | 87.9333 |  | 2.4 | 12.75 | 86.6 | 86.55 |  | 3.92 | 46.4333 | 84.8167 | 84.8286 |  | 5.9 | 75.9667 | 84.5 | 84.5
12.1 |  |  | 7.4 | 88.9 | 88.8 |  | 1.8 | 8.4 | 87.9 | 87.8667 |  | 2.4 | 12.25 | 86.5 | 86.4833 |  | 3.9 | 39.42 | 84.7167 | 84.7 |  | 5.8 | 75.175 | 84.45 | 84.4
12.2 |  |  | 7.2 | 88.8 | 88.7 |  | 1.8 | 8.2 | 87.82 | 87.7667 |  | 2.3 | 11.9 | 86.4 | 86.3 |  | 3.8 | 33.38 | 84.6 | 84.6 |  | 5.7 | 74.35 | 84.2333 | 84.24
12.3 |  |  | 7.1 | 88.8 | 88.7 |  | 1.8 | 8 | 87.8 | 87.6714 |  | 2.3 | 11.5 | 86.3 | 86.2 |  | 3.8 | 28.3 | 84.4833 | 84.5 |  | 5.6 | 73.1667 | 84.05 | 84.1
12.4 |  |  | 6.9 | 88.7 | 88.6 |  | 1.8 | 7.8 | 87.7 | 87.6 |  | 2.3 | 11.2 | 86.15 | 86.1 |  | 3.7 | 24.6667 | 84.3 | 84.3833 |  | 5.5 | 71.7 | 83.85 | 84
12.5 |  |  | 6.8 | 88.6 | 88.5 |  | 1.7 | 7.7 | 87.6 | 87.5 |  | 2.2 | 10.9 | 86 | 86 |  | 3.7 | 22.04 | 84.2 | 84.2167 |  | 5.5 | 69.66 | 83.6833 | 83.8286
12.6 |  |  | 6.7 | 88.6 | 88.5 |  |  | 7.5 | 87.6 | 87.4833 |  | 2.2 | 10.6 | 85.9167 | 85.9 |  | 3.6 | 20.35 | 84 | 84.025 |  | 5.4 | 66.8 | 83.5857 | 83.66
12.7 |  |  | 6.58 | 88.5 | 88.4667 |  |  | 7.325 | 87.5 | 87.4 |  | 2.2 | 10.3 | 85.9 | 85.7833 |  | 3.6 | 19.0167 | 83.9 | 83.84 |  | 5.3 | 63.5 | 83.4333 | 83.5
12.8 |  |  | 6.4333 | 88.4 | 88.3 |  |  | 7.2 | 87.3833 | 87.3 |  | 2.2 | 10.1 | 85.7 | 85.7 |  | 3.5 | 18.06 | 83.7333 | 83.7 |  | 5.2333 | 58.5 | 83.26 | 83.3833
12.9 |  |  | 6.3 | 88.4 | 88.3 |  |  | 7 | 87.2 | 87.2 |  | 2.1 | 9.92 | 85.68 | 85.5667 |  | 3.5 | 17.22 | 83.5667 | 83.6 |  | 5.175 | 52.62 | 83.1667 | 83.2167
13 |  |  | 6.25 | 88.3 | 88.2 |  |  | 6.9 | 87.14 | 87.0857 |  | 2.1 | 9.68 | 85.58 | 85.4143 |  | 3.4667 | 16.65 | 83.5 | 83.48 |  | 5.1 | 46.7333 | 82.9 | 83.0667
13.1 |  |  | 6.1 | 88.2 | 88.1 |  |  | 6.8 | 87.1 | 87 |  | 2.1 | 9.475 | 85.4 | 85.25 |  | 3.4 | 16.05 | 83.2333 | 83.3 |  | 4.98 | 40.68 | 82.8 | 83
13.2 |  |  | 6.05 | 88.2 | 88.0167 |  |  | 6.6 | 86.95 | 86.92 |  |  | 9.26 | 85.3333 | 85.16 |  | 3.3 | 15.56 | 83.1833 | 83.22 |  | 4.9 | 35.3 | 82.62 | 82.8
13.3 |  |  | 5.9 | 88.15 | 88 |  |  | 6.5 | 86.8 | 86.82 |  |  | 9.08 | 85.1667 | 85 |  | 3.3 | 15.2333 | 82.98 | 83.05 |  | 4.9 | 30.9857 | 82.5 | 82.6167
13.4 |  |  | 5.825 | 88.0833 | 87.9 |  |  | 6.4 | 86.8 | 86.7 |  |  | 8.8833 | 85 | 84.8333 |  | 3.275 | 14.8 | 82.9 | 82.86 |  | 4.8 | 27.5 | 82.4 | 82.4167
13.5 |  |  | 5.7333 | 87.9 | 87.8 |  |  | 6.3 | 86.6667 | 86.6333 |  |  | 8.7167 | 84.8667 | 84.7333 |  | 3.2 | 14.54 | 82.7167 | 82.7667 |  | 4.7 | 24.9 | 82.1833 | 82.2833
13.6 |  |  | 5.7 | 87.8286 | 87.7 |  |  | 6.2 | 86.5 | 86.54 |  |  | 8.5571 | 84.7 | 84.6 |  | 3.2 | 14.14 | 82.55 | 82.5833 |  | 4.7 | 23.1 | 82.0167 | 82.1167
13.7 |  |  | 5.6 | 87.74 | 87.5667 |  |  | 6.0333 | 86.4 | 86.3667 |  |  | 8.4 | 84.58 | 84.4 |  | 3.1 | 13.8667 | 82.3833 | 82.38 |  | 4.6 | 21.8 | 81.85 | 81.95
13.8 |  |  | 5.5 | 87.5667 | 87.5 |  |  | 6 | 86.32 | 86.2 |  |  | 8.2429 | 84.475 | 84.3 |  | 3.1 | 13.66 | 82.2 | 82.275 |  | 4.525 | 20.75 | 81.7 | 81.7667
13.9 |  |  | 5.4 | 87.4 | 87.3 |  |  | 5.9 | 86.2 | 86.1167 |  |  | 8.0667 | 84.34 | 84.1167 |  | 3.1 | 13.4333 | 82.0571 | 82.025 |  | 4.5 | 19.98 | 81.5667 | 81.6
14 |  |  | 5.3 | 87.3 | 87.1667 |  |  | 5.8 | 86 | 85.95 |  |  | 7.9 | 84.15 | 84 |  | 3 | 13.1 | 81.9 | 81.95 |  | 4.4 | 19.42 | 81.425 | 81.5
14.1 |  |  | 5.3 | 87.02 | 87 |  |  | 5.7 | 85.8333 | 85.775 |  |  | 7.8 | 83.925 | 83.7833 |  | 3 | 12.8 | 81.74 | 81.7 |  | 4.4 | 18.86 | 81.22 | 81.4
14.2 |  |  | 5.2 | 86.7667 | 86.6667 |  |  | 5.6 | 85.64 | 85.475 |  |  | 7.6 | 83.7 | 83.7 |  | 2.975 | 12.6167 | 81.6667 | 81.5714 |  | 4.3 | 18.42 | 81.1 | 81.1
14.3 |  |  | 5.1 | 86.4667 | 86.36 |  |  | 5.5 | 85.475 | 85.3 |  |  | 7.4429 | 83.5167 | 83.5 |  | 2.9 | 12.44 | 81.4 | 81.32 |  | 4.25 | 17.95 | 80.9 | 80.8333
14.4 |  |  | 5 | 85.9667 | 85.9143 |  |  | 5.4 | 85.1833 | 85.0667 |  |  | 7.325 | 83.3333 | 83.2833 |  |  | 12.2 | 81.1667 | 81.1333 |  | 4.2 | 17.55 | 80.6857 | 80.7333
14.5 |  |  | 5 | 85.4 | 85.3 |  |  | 5.3 | 85 | 84.8 |  |  | 7.2 | 83.14 | 83.1167 |  |  | 11.925 | 80.95 | 80.9833 |  | 4.2 | 17.18 | 80.5667 | 80.5
14.6 |  |  | 4.9 | 84.6 | 84.4833 |  |  | 5.2 | 84.58 | 84.4667 |  |  | 7.08 | 82.9333 | 82.9333 |  |  | 11.7333 | 80.75 | 80.8833 |  | 4.1 | 16.9286 | 80.3 | 80.38
14.7 |  |  | 4.8 | 83.2333 | 83.3167 |  |  | 5.2 | 84.18 | 84.04 |  |  | 6.9833 | 82.7667 | 82.7 |  |  | 11.5333 | 80.56 | 80.68 |  | 4.1 | 16.58 | 80.2 | 80.3
14.8 |  |  | 4.8 | 81.15 | 81.36 |  |  | 5.1 | 83.78 | 83.56 |  |  | 6.8167 | 82.6 | 82.42 |  |  | 11.3 | 80.46 | 80.5143 |  | 4 | 16.2429 | 80.0333 | 80.1
14.9 |  |  | 4.7 | 77.5 | 77.9667 |  |  | 5 | 82.95 | 82.9333 |  |  | 6.7333 | 82.32 | 82.1333 |  |  | 11.1833 | 80.26 | 80.3 |  | 3.9 | 16.0167 | 79.9 | 79.9286
15 |  |  | 4.7 | 72.275 | 72.5833 |  |  | 5 | 82.2143 | 82.1 |  |  | 6.6667 | 81.94 | 81.9333 |  |  | 10.9833 | 80.0667 | 80.0667 |  | 3.9 | 15.85 | 79.68 | 79.74
15.1 |  |  | 4.6 | 64.8 | 65.2 |  |  | 4.9 | 81.2333 | 81.04 |  |  | 6.5 | 81.7333 | 81.6 |  |  | 10.8167 | 79.9 | 79.8167 |  | 3.9 | 15.66 | 79.5 | 79.5333
15.2 |  |  | 4.525 | 56.1 | 57.0667 |  |  | 4.8 | 79.6 | 79.3714 |  |  | 6.4167 | 81.425 | 81.2857 |  |  | 10.55 | 79.6167 | 79.62 |  | 3.8 | 15.3667 | 79.35 | 79.3
15.3 |  |  | 4.5 | 46.5 | 48.06 |  |  | 4.775 | 77.2 | 76.85 |  |  | 6.35 | 81 | 80.8667 |  |  | 10.4 | 79.4167 | 79.45 |  | 3.7 | 15.0833 | 79.16 | 79.15
15.4 |  |  | 4.45 | 37.5333 | 39.05 |  |  | 4.7 | 73.2857 | 72.45 |  |  | 6.2667 | 80.5 | 80.45 |  |  | 10.28 | 79.2333 | 79.1857 |  | 3.7 | 14.8333 | 79.05 | 78.96
15.5 |  |  | 4.4 | 29.44 | 31.0333 |  |  | 4.6 | 67.175 | 66.7 |  |  | 6.16 | 80.05 | 79.98 |  |  | 10.075 | 79 | 79.0333 |  | 3.7 | 14.6 | 78.8 | 78.72
15.6 |  |  | 4.4 | 23.25 | 24.4 |  |  | 4.6 | 60.0429 | 59.2 |  |  | 6.1 | 79.4 | 79.35 |  |  | 9.94 | 78.7571 | 78.78 |  |  | 14.4 | 78.62 | 78.4667
15.7 |  |  | 4.3 | 18.65 | 20.02 |  |  | 4.5 | 51.9333 | 51.08 |  |  | 6 | 78.6 | 78.5333 |  |  | 9.7667 | 78.5667 | 78.46 |  |  | 14.2 | 78.32 | 78.2833
15.8 |  |  | 4.3 | 16.2 | 17.5167 |  |  | 4.5 | 43.0167 | 43.3333 |  |  | 5.86 | 77.6 | 77.6286 |  |  | 9.6 | 78.3 | 78.24 |  |  | 14 | 78.08 | 78.1
15.9 |  |  | 4.225 | 14.54 | 15.85 |  |  | 4.4 | 34.84 | 34.9667 |  |  | 5.8 | 76.28 | 76.3 |  |  | 9.4833 | 78.04 | 78.04 |  |  | 13.8 | 77.88 | 77.875
16 |  |  | 4.2 | 13.5 | 14.75 |  |  | 4.4 | 28.1 | 28.4 |  |  | 5.7 | 74.2667 | 74.5 |  |  | 9.3 | 77.8333 | 77.825 |  |  | 13.5333 | 77.66 | 77.68
16.1 |  |  | 4.2 | 12.82 | 14.02 |  |  | 4.3 | 22.6167 | 23.2 |  |  | 5.6143 | 71.1 | 71.3 |  |  | 9.2 | 77.52 | 77.5333 |  |  | 13.3143 | 77.54 | 77.4833
16.2 |  |  | 4.1 | 12.25 | 13.45 |  |  | 4.3 | 19.2 | 19.8 |  |  | 5.55 | 66.38 | 67 |  |  | 9.0167 | 77.225 | 77.3 |  |  | 13.1 | 77.3333 | 77.2333
16.3 |  |  | 4.1 | 11.8333 | 13.05 |  |  | 4.2 | 16.9667 | 17.7 |  |  | 5.4857 | 60.25 | 60.9333 |  |  | 8.8333 | 76.98 | 76.94 |  |  | 12.96 | 77.12 | 77.1
16.4 |  |  | 4.1 | 11.45 | 12.6 |  |  | 4.2 | 15.4667 | 16.3333 |  |  | 5.4 | 53.66 | 53.7 |  |  | 8.7333 | 76.68 | 76.62 |  |  | 12.75 | 76.8333 | 76.88
16.5 |  |  | 4 | 11.24 | 12.36 |  |  | 4.2 | 14.4 | 15.3571 |  |  | 5.3667 | 45.5 | 46.6 |  |  | 8.6 | 76.3667 | 76.2667 |  |  | 12.6 | 76.5333 | 76.7
16.6 |  |  | 4 | 11.04 | 12.175 |  |  | 4.1 | 13.7 | 14.75 |  |  | 5.275 | 38.2 | 39.4 |  |  | 8.5 | 76.04 | 75.98 |  |  | 12.4 | 76.32 | 76.3667
16.7 |  |  | 4 | 10.825 | 11.98 |  |  | 4.1 | 13.1 | 14.15 |  |  | 5.2 | 31.925 | 33.12 |  |  | 8.34 | 75.65 | 75.5714 |  |  | 12.2167 | 76.08 | 76.15
16.8 |  |  | 3.9 | 10.6 | 11.78 |  |  | 4.1 | 12.6286 | 13.6667 |  |  | 5.1 | 26.425 | 28.02 |  |  | 8.25 | 75.15 | 75.2333 |  |  | 12.0333 | 75.85 | 75.8667
16.9 |  |  | 3.9 | 10.5 | 11.6 |  |  | 4 | 12.32 | 13.3333 |  |  | 5.1 | 22.46 | 24 |  |  | 8.2 | 74.74 | 74.74 |  |  | 11.9333 | 75.62 | 75.675
17 |  |  | 3.9 | 10.3 | 11.5 |  |  | 4 | 11.96 | 13.04 |  |  | 5.0167 | 20.2 | 21.3 |  |  | 8.0333 | 74.16 | 74.3333 |  |  | 11.6333 | 75.42 | 75.3429
17.1 |  |  | 3.8 | 10.1667 | 11.3333 |  |  | 4 | 11.6833 | 12.7333 |  |  | 5 | 18.46 | 19.55 |  |  | 7.9 | 73.56 | 73.65 |  |  | 11.5 | 75.2 | 75.0833
17.2 |  |  |  | 10.08 | 11.1714 |  |  | 3.9 | 11.4333 | 12.5 |  |  | 4.9 | 17.3 | 18.4 |  |  | 7.7714 | 72.8 | 72.8333 |  |  | 11.32 | 74.96 | 74.8333
17.3 |  |  |  | 9.98 | 11.0167 |  |  | 3.9 | 11.2167 | 12.3333 |  |  | 4.9 | 16.5667 | 17.5667 |  |  | 7.6167 | 71.9 | 71.9 |  |  | 11.22 | 74.6 | 74.6833
17.4 |  |  |  | 9.8 | 10.95 |  |  | 3.825 | 11.02 | 12.05 |  |  | 4.8 | 15.82 | 16.95 |  |  | 7.6 | 70.5 | 70.7 |  |  | 11.0667 | 74.25 | 74.4333
17.5 |  |  |  | 9.7 | 10.8 |  |  | 3.8 | 10.84 | 11.9667 |  |  | 4.8 | 15.28 | 16.4 |  |  | 7.5 | 68.7 | 68.9 |  |  | 10.9 | 74.1 | 74.1
17.6 |  |  |  | 9.58 | 10.7167 |  |  | 3.8 | 10.65 | 11.7667 |  |  | 4.7 | 14.78 | 16 |  |  | 7.3667 | 66.1 | 66.1 |  |  | 10.7333 | 73.8 | 73.8
17.7 |  |  |  | 9.4667 | 10.6 |  |  | 3.725 | 10.5 | 11.575 |  |  | 4.7 | 14.4 | 15.6 |  |  | 7.3 | 62.38 | 62.4571 |  |  | 10.64 | 73.5 | 73.54
17.8 |  |  |  | 9.375 | 10.5 |  |  | 3.7 | 10.3 | 11.4833 |  |  | 4.6 | 14.0667 | 15.2667 |  |  | 7.2 | 57.45 | 57.78 |  |  | 10.48 | 73.1333 | 73.28
17.9 |  |  |  | 9.26 | 10.38 |  |  |  | 10.2 | 11.3 |  |  | 4.6 | 13.7333 | 14.9333 |  |  | 7.1 | 51.6833 | 52.1833 |  |  | 10.3 | 72.8667 | 72.82
18 |  |  |  | 9.1833 | 10.28 |  |  |  | 10.08 | 11.2 |  |  | 4.5 | 13.4714 | 14.6 |  |  | 7 | 45.5167 | 46.0167 |  |  | 10.15 | 72.55 | 72.5
18.1 |  |  |  | 9.075 | 10.175 |  |  |  | 9.9 | 11 |  |  | 4.5 | 13.2333 | 14.3857 |  |  | 6.9167 | 39.64 | 40.18 |  |  | 10.06 | 72.1667 | 72.14
18.2 |  |  |  | 8.96 | 10.025 |  |  |  | 9.8 | 10.8 |  |  | 4.4 | 13.025 | 14.1667 |  |  | 6.8 | 34.125 | 34.9 |  |  | 9.96 | 71.8 | 71.825
18.3 |  |  |  | 8.86 | 9.9333 |  |  |  | 9.6333 | 10.7167 |  |  | 4.4 | 12.7667 | 13.9167 |  |  | 6.8 | 29.7 | 30.6 |  |  | 9.8 | 71.35 | 71.4
18.4 |  |  |  | 8.76 | 9.8 |  |  |  | 9.4667 | 10.55 |  |  | 4.3167 | 12.5667 | 13.65 |  |  | 6.7 | 26.1 | 27.1429 |  |  | 9.7 | 70.85 | 70.95
18.5 |  |  |  | 8.66 | 9.75 |  |  |  | 9.4 | 10.4667 |  |  | 4.3 | 12.4 | 13.4833 |  |  | 6.6 | 23.72 | 24.7 |  |  | 9.5286 | 70.35 | 70.4
18.6 |  |  |  | 8.5 | 9.6 |  |  |  | 9.3 | 10.38 |  |  | 4.25 | 12.1333 | 13.3 |  |  | 6.5 | 22 | 23.2 |  |  | 9.4 | 69.7 | 69.7333
18.7 |  |  |  | 8.4143 | 9.5167 |  |  |  | 9.1333 | 10.25 |  |  |  | 11.96 | 13.1286 |  |  | 6.46 | 20.8 | 22.02 |  |  | 9.3 | 68.85 | 69.0667
18.8 |  |  |  | 8.3167 | 9.4167 |  |  |  | 9.06 | 10.1 |  |  |  | 11.8 | 12.94 |  |  | 6.3667 | 19.98 | 21.1 |  |  | 9.2 | 67.85 | 68
18.9 |  |  |  | 8.2167 | 9.32 |  |  |  | 8.9 | 9.9667 |  |  |  | 11.65 | 12.7667 |  |  | 6.3 | 19.28 | 20.4333 |  |  | 9.14 | 66.64 | 66.7333
19 |  |  |  | 8.15 | 9.24 |  |  |  | 8.8 | 9.9 |  |  |  | 11.46 | 12.6 |  |  | 6.2167 | 18.78 | 19.95 |  |  | 8.95 | 64.8 | 65.0429
19.1 |  |  |  | 8 | 9.15 |  |  |  | 8.7167 | 9.7429 |  |  |  | 11.2833 | 12.4333 |  |  | 6.125 | 18.2667 | 19.42 |  |  | 8.8 | 62.3 | 62.6429
19.2 |  |  |  | 7.9 | 9.02 |  |  |  | 8.6 | 9.66 |  |  |  | 11.1167 | 12.2667 |  |  | 6.1 | 17.8667 | 19.0429 |  |  | 8.7833 | 58.98 | 59.44
19.3 |  |  |  | 7.8 | 8.9 |  |  |  | 8.5 | 9.56 |  |  |  | 11 | 12.1 |  |  | 6 | 17.5333 | 18.7333 |  |  | 8.7 | 54.75 | 55.3667
19.4 |  |  |  | 7.8 | 8.8167 |  |  |  | 8.4 | 9.4 |  |  |  | 10.82 | 11.9167 |  |  | 5.95 | 17.2 | 18.3333 |  |  | 8.5857 | 49.78 | 50.6
19.5 |  |  |  | 7.7 | 8.7333 |  |  |  | 8.275 | 9.38 |  |  |  | 10.7 | 11.8 |  |  | 5.9 | 16.85 | 18.1 |  |  | 8.5 | 44.56 | 45.5333
19.6 |  |  |  | 7.5667 | 8.6 |  |  |  | 8.2 | 9.2 |  |  |  | 10.5333 | 11.65 |  |  | 5.8 | 16.6 | 17.775 |  |  | 8.4 | 39.36 | 40.5
19.7 |  |  |  | 7.5 | 8.5833 |  |  |  | 8.1 | 9.1 |  |  |  | 10.3667 | 11.4833 |  |  | 5.7857 | 16.32 | 17.5667 |  |  | 8.2333 | 34.75 | 36.2143
19.8 |  |  |  | 7.4 | 8.4167 |  |  |  | 8 | 9.02 |  |  |  | 10.2 | 11.3167 |  |  | 5.7 | 16.0333 | 17.3 |  |  | 8.2 | 30.9 | 32.24
19.9 |  |  |  | 7.3 | 8.34 |  |  |  | 7.9 | 8.9333 |  |  |  | 10.12 | 11.25 |  |  | 5.6 | 15.8 | 16.96 |  |  | 8.1 | 28 | 29.4
20 |  |  |  | 7.2 | 8.24 |  |  |  | 7.8 | 8.8 |  |  |  | 9.92 | 11.08 |  |  | 5.58 | 15.58 | 16.7857 |  |  | 8.08 | 25.74 | 27.2
20.1 |  |  |  | 7.1 | 8.1667 |  |  |  | 7.7 | 8.7 |  |  |  | 9.8 | 10.8833 |  |  | 5.5 | 15.2833 | 16.5667 |  |  | 8 | 24.2667 | 25.6
20.2 |  |  |  | 7 | 8.1 |  |  |  | 7.6 | 8.6167 |  |  |  | 9.6667 | 10.8 |  |  | 5.4833 | 15.1 | 16.3143 |  |  | 7.8333 | 23.2 | 24.6
20.3 |  |  |  | 6.9 | 7.94 |  |  |  | 7.5 | 8.525 |  |  |  | 9.5 | 10.6333 |  |  | 5.4 | 14.9167 | 16.0667 |  |  | 7.7 | 22.3 | 23.7
20.4 |  |  |  | 6.8 | 7.8833 |  |  |  | 7.4 | 8.4 |  |  |  | 9.4429 | 10.55 |  |  |  | 14.6333 | 15.85 |  |  | 7.7 | 21.6333 | 23.04
20.5 |  |  |  | 6.7833 | 7.7167 |  |  |  | 7.3 | 8.3 |  |  |  | 9.3 | 10.4 |  |  |  | 14.45 | 15.7333 |  |  | 7.6 | 21.15 | 22.6
20.6 |  |  |  | 6.6167 | 7.6667 |  |  |  | 7.2333 | 8.2167 |  |  |  | 9.2 | 10.3 |  |  |  | 14.25 | 15.44 |  |  | 7.5 | 20.65 | 22.1
20.7 |  |  |  | 6.5667 | 7.5667 |  |  |  | 7.1 | 8.15 |  |  |  | 9.06 | 10.175 |  |  |  | 14.1 | 15.3 |  |  | 7.48 | 20.2333 | 21.7
20.8 |  |  |  | 6.4 | 7.48 |  |  |  | 7.0167 | 8 |  |  |  | 8.9833 | 10.04 |  |  |  | 13.88 | 15.18 |  |  | 7.3833 | 19.84 | 21.3
20.9 |  |  |  | 6.3333 | 7.375 |  |  |  | 6.94 | 7.9 |  |  |  | 8.9 | 9.9 |  |  |  | 13.6833 | 14.9667 |  |  | 7.2 | 19.5 | 20.9
21 |  |  |  | 6.25 | 7.3 |  |  |  | 6.85 | 7.8 |  |  |  | 8.78 | 9.8143 |  |  |  | 13.5 | 14.7167 |  |  | 7.2 | 19.3 | 20.65
21.1 |  |  |  | 6.2 | 7.2 |  |  |  | 6.7 | 7.7 |  |  |  | 8.6 | 9.7 |  |  |  | 13.34 | 14.55 |  |  | 7.1 | 19.0167 | 20.38
21.2 |  |  |  | 6.1 | 7.1 |  |  |  | 6.6 | 7.6 |  |  |  | 8.5 | 9.5333 |  |  |  | 13.2 | 14.36 |  |  | 7.025 | 18.7 | 20.1286
21.3 |  |  |  | 6.05 | 7 |  |  |  | 6.55 | 7.5167 |  |  |  | 8.4286 | 9.45 |  |  |  | 13 | 14.2 |  |  | 6.92 | 18.4 | 19.84
21.4 |  |  |  | 6 | 7 |  |  |  | 6.46 | 7.45 |  |  |  | 8.3 | 9.375 |  |  |  | 12.7667 | 14.1 |  |  | 6.9 | 18.1 | 19.6
21.5 |  |  |  | 5.9 | 6.9 |  |  |  | 6.3833 | 7.3333 |  |  |  | 8.2 | 9.275 |  |  |  | 12.6667 | 13.9 |  |  | 6.84 | 17.9333 | 19.3833
21.6 |  |  |  | 5.9 | 6.9 |  |  |  | 6.28 | 7.2429 |  |  |  | 8.1 | 9.15 |  |  |  | 12.5 | 13.7 |  |  | 6.725 | 17.6333 | 19.1333
21.7 |  |  |  | 5.8 | 6.8 |  |  |  | 6.2 | 7.2 |  |  |  | 8 | 9.06 |  |  |  | 12.3 | 13.5667 |  |  | 6.62 | 17.4333 | 18.925

## Wedge 15 Longitudinal Profiles
Longitudinal Profiles
Wedge 15
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 9.9
-34.9 |  |  |  |  | 9.9625
-34.8 |  |  |  |  | 10.0875
-34.7 |  |  |  |  | 10.1286
-34.6 |  |  |  |  | 10.2
-34.5 |  |  |  |  | 10.2875
-34.4 |  |  |  |  | 10.4143
-34.3 |  |  |  |  | 10.5
-34.2 |  |  |  |  | 10.5857
-34.1 |  |  |  |  | 10.65
-34 |  |  |  |  | 10.7
-33.9 |  |  |  |  | 10.79
-33.8 |  |  |  |  | 10.9
-33.7 |  |  |  |  | 11.0333
-33.6 |  |  |  |  | 11.1571
-33.5 |  |  |  |  | 11.2875
-33.4 |  |  |  |  | 11.325
-33.3 |  |  |  |  | 11.4
-33.2 |  |  |  |  | 11.4875
-33.1 |  |  |  |  | 11.625
-33 |  |  |  |  | 11.7571
-32.9 |  |  |  |  | 11.8778
-32.8 |  |  |  |  | 11.9286
-32.7 |  |  |  |  | 12.1571
-32.6 |  |  |  |  | 12.2
-32.5 |  |  |  |  | 12.3429
-32.4 |  |  |  |  | 12.4
-32.3 |  |  |  | 8.7 | 12.6125
-32.2 |  |  |  | 8.8 | 12.7429
-32.1 |  |  |  | 8.8625 | 12.875
-32 |  |  |  | 9 | 13
-31.9 |  |  |  | 9.06 | 13.075
-31.8 |  |  |  | 9.15 | 13.2
-31.7 |  |  |  | 9.2 | 13.3
-31.6 |  |  |  | 9.2875 | 13.425
-31.5 |  |  |  | 9.4 | 13.5
-31.4 |  |  |  | 9.5 | 13.675
-31.3 |  |  |  | 9.58 | 13.825
-31.2 |  |  |  | 9.7 | 13.95
-31.1 |  |  |  | 9.75 | 14.1
-31 |  |  |  | 9.8 | 14.25
-30.9 |  |  |  | 9.95 | 14.3857
-30.8 |  |  |  | 10 | 14.5125
-30.7 |  |  |  | 10.125 | 14.6429
-30.6 |  |  |  | 10.25 | 14.775
-30.5 |  |  |  | 10.3 | 14.9125
-30.4 |  |  |  | 10.4143 | 15.0857
-30.3 |  |  |  | 10.55 | 15.275
-30.2 |  |  |  | 10.6857 | 15.425
-30.1 |  |  |  | 10.74 | 15.6375
-30 |  |  |  | 10.9 | 15.7714
-29.9 |  |  |  | 10.9333 | 15.9
-29.8 |  |  |  | 11.1 | 16.0857
-29.7 |  |  |  | 11.2 | 16.3
-29.6 |  |  | 7.4 | 11.3 | 16.4667
-29.5 |  |  | 7.525 | 11.4833 | 16.64
-29.4 |  |  | 7.6 | 11.6 | 16.8333
-29.3 |  |  | 7.65 | 11.7 | 16.975
-29.2 |  |  | 7.72 | 11.84 | 17.2
-29.1 |  |  | 7.7867 | 11.9 | 17.375
-29 |  |  | 7.8615 | 12.06 | 17.5714
-28.9 |  |  | 7.9625 | 12.2375 | 17.8
-28.8 |  |  | 8 | 12.3625 | 18.0375
-28.7 |  |  | 8.08 | 12.5 | 18.1625
-28.6 |  |  | 8.1 | 12.6286 | 18.4
-28.5 |  |  | 8.2286 | 12.7625 | 18.6286
-28.4 |  |  | 8.3 | 12.9 | 18.825
-28.3 |  | 6.8 | 8.4 | 13.0286 | 19.1
-28.2 |  | 6.8 | 8.5 | 13.225 | 19.3286
-28.1 |  | 6.9 | 8.5714 | 13.4 | 19.525
-28 |  | 6.9 | 8.6 | 13.5286 | 19.6875
-27.9 |  | 7 | 8.7 | 13.6625 | 19.95
-27.8 |  | 7 | 8.7571 | 13.8 | 20.2714
-27.7 |  | 7.0875 | 8.8875 | 14.025 | 20.4875
-27.6 |  | 7.125 | 9.025 | 14.2143 | 20.775
-27.5 |  | 7.2 | 9.1 | 14.3875 | 21.1143
-27.4 |  | 7.2875 | 9.1875 | 14.525 | 21.4625
-27.3 | 6.8286 | 7.36 | 9.3125 | 14.7 | 21.7429
-27.2 | 6.9 | 7.4 | 9.4 | 14.8857 | 22.1
-27.1 | 6.9 | 7.5 | 9.475 | 15.1167 | 22.5429
-27 | 7 | 7.5 | 9.6125 | 15.3111 | 23.08
-26.9 | 7 | 7.6 | 9.7429 | 15.4857 | 23.56
-26.8 | 7.0778 | 7.625 | 9.8 | 15.7143 | 24.3333
-26.7 | 7.1 | 7.75 | 9.9333 | 15.9 | 25.3
-26.6 | 7.175 | 7.8 | 10.075 | 16.0857 | 26.4714
-26.5 | 7.2 | 7.9 | 10.1125 | 16.3286 | 27.9571
-26.4 | 7.3 | 7.95 | 10.2429 | 16.6 | 30
-26.3 | 7.3714 | 8 | 10.3857 | 16.7857 | 32.56
-26.2 | 7.4 | 8.1 | 10.5143 | 17.0143 | 35.82
-26.1 | 7.5 | 8.14 | 10.65 | 17.25 | 39.54
-26 | 7.52 | 8.2733 | 10.7857 | 17.4857 | 43.8125
-25.9 | 7.6 | 8.3533 | 10.9143 | 17.7286 | 48.1
-25.8 | 7.6167 | 8.42 | 11.05 | 18 | 52.7
-25.7 | 7.7 | 8.4867 | 11.1857 | 18.2714 | 57.175
-25.6 | 7.8 | 8.6067 | 11.3143 | 18.6286 | 61.125
-25.5 | 7.9 | 8.7333 | 11.45 | 18.9 | 64.225
-25.4 | 7.9 | 8.8 | 11.5857 | 19.2571 | 66.9571
-25.3 | 8 | 8.86 | 11.8125 | 19.5375 | 69
-25.2 | 8.1 | 8.975 | 11.9429 | 19.9125 | 70.675
-25.1 | 8.2 | 9.1 | 12.075 | 20.3857 | 71.8714
-25 | 8.2 | 9.2 | 12.225 | 20.8 | 72.8
-24.9 | 8.4 | 9.3429 | 12.4429 | 21.4875 | 73.5143
-24.8 | 8.4 | 9.4 | 12.65 | 22.2286 | 74.1125
-24.7 | 8.5 | 9.5167 | 12.8143 | 23.2 | 74.7
-24.6 | 8.575 | 9.6625 | 12.9889 | 24.5714 | 75.2
-24.5 | 8.7 | 9.8 | 13.1714 | 26.4 | 75.6333
-24.4 | 8.75 | 9.8286 | 13.4 | 28.8625 | 76.05
-24.3 | 8.875 | 9.9625 | 13.575 | 32.1571 | 76.4
-24.2 | 8.9125 | 10.1 | 13.8429 | 36 | 76.74
-24.1 | 9.0429 | 10.225 | 14 | 40.7 | 77.08
-24 | 9.175 | 10.3571 | 14.2375 | 45.9857 | 77.4667
-23.9 | 9.2143 | 10.4875 | 14.4429 | 51.5 | 77.92
-23.8 | 9.35 | 10.6333 | 14.7 | 56.9125 | 78.1
-23.7 | 9.4 | 10.7857 | 14.975 | 61.7875 | 78.3667
-23.6 | 9.5 | 10.95 | 15.2429 | 66.1 | 78.675
-23.5 | 9.5909 | 11.1571 | 15.5 | 70.2 | 78.9571
-23.4 | 9.7 | 11.2875 | 15.8 | 72.94 | 79.2625
-23.3 | 9.8375 | 11.4143 | 16.1125 | 74.84 | 79.525
-23.2 | 9.9714 | 11.6 | 16.5143 | 76.42 | 79.7143
-23.1 | 10.1 | 11.7857 | 16.95 | 77.7 | 79.975
-23 | 10.2375 | 11.9286 | 17.5571 | 78.65 | 80.3143
-22.9 | 10.3714 | 12.15 | 18.1444 | 79.5 | 80.55
-22.8 | 10.5 | 12.3714 | 19 | 80.18 | 80.8714
-22.7 | 10.6375 | 12.5286 | 20.125 | 80.5857 | 81.1286
-22.6 | 10.7714 | 12.8 | 21.8 | 81.05 | 81.35
-22.5 | 10.9 | 13.0714 | 23.925 | 81.4625 | 81.55
-22.4 | 11.075 | 13.3286 | 26.9286 | 81.95 | 81.7125
-22.3 | 11.2625 | 13.6 | 30.8 | 82.2143 | 81.9286
-22.2 | 11.4 | 13.925 | 35.875 | 82.65 | 82.25
-22.1 | 11.525 | 14.3375 | 41.7286 | 82.975 | 82.4
-22 | 11.6571 | 14.8143 | 48.5125 | 83.3143 | 82.6333
-21.9 | 11.8909 | 15.475 | 55.525 | 83.575 | 82.9
-21.8 | 12.1385 | 16.3167 | 62.6571 | 83.875 | 83.0714
-21.7 | 12.3667 | 17.4 | 68.9125 | 84.2 | 83.2667
-21.6 | 12.6231 | 19.1 | 74.0857 | 84.4714 | 83.5
-21.5 | 12.9444 | 21.48 | 78.35 | 84.7333 | 83.7
-21.4 | 13.4 | 24.85 | 81.6714 | 84.95 | 83.9
-21.3 | 13.925 | 29.1 | 83.675 | 85.225 | 84.075
-21.2 | 14.7308 | 35.05 | 85.2857 | 85.4857 | 84.2714
-21.1 | 15.8733 | 42.26 | 86.475 | 85.75 | 84.5
-21 | 17.6067 | 49.6 | 87.275 | 86.0125 | 84.675
-20.9 | 20.1 | 57.7714 | 87.9714 | 86.16 | 84.8714
-20.8 | 23.4 | 65.82 | 88.575 | 86.3857 | 85.1
-20.7 | 28.275 | 73.5667 | 89.05 | 86.65 | 85.2333
-20.6 | 35.0846 | 79.2 | 89.5286 | 86.825 | 85.4
-20.5 | 42.6867 | 83.8 | 89.85 | 87.0429 | 85.6375
-20.4 | 51.5533 | 87.2 | 90.2375 | 87.25 | 85.7714
-20.3 | 60.1625 | 89.4 | 90.5857 | 87.5125 | 85.9
-20.2 | 68.6 | 90.9 | 90.85 | 87.66 | 86.0571
-20.1 | 76.2 | 91.94 | 91.125 | 87.8429 | 86.2625
-20 | 82.6333 | 92.7714 | 91.4286 | 88.05 | 86.4
-19.9 | 87.9 | 93.4857 | 91.75 | 88.3 | 86.55
-19.8 | 90.9 | 93.92 | 91.8222 | 88.4364 | 86.7571
-19.7 | 93.2 | 94.4 | 92.0857 | 88.6111 | 86.975
-19.6 | 94.8 | 94.7 | 92.4429 | 88.725 | 87.1143
-19.5 | 95.6 | 95.0333 | 92.5 | 88.8444 | 87.25
-19.4 | 96.2667 | 95.2 | 92.7286 | 88.9909 | 87.3857
-19.3 | 96.8 | 95.5143 | 92.925 | 89.1889 | 87.6
-19.2 | 97.1333 | 95.8 | 93.2 | 89.3714 | 87.7375
-19.1 | 97.4333 | 96 | 93.3286 | 89.525 | 87.8556
-19 | 97.7 | 96.2 | 93.525 | 89.7429 | 88.02
-18.9 | 97.9 | 96.3333 | 93.7 | 89.9 | 88.2167
-18.8 | 98.0938 | 96.4333 | 93.9286 | 90.1 | 88.3833
-18.7 | 98.2812 | 96.6 | 94.0625 | 90.175 | 88.4818
-18.6 | 98.42 | 96.7 | 94.2 | 90.3 | 88.6333
-18.5 | 98.5533 | 96.88 | 94.325 | 90.475 | 88.7429
-18.4 | 98.6867 | 97.025 | 94.4 | 90.7 | 88.825
-18.3 | 98.76 | 97.1571 | 94.4875 | 90.7667 | 88.9571
-18.2 | 98.8533 | 97.3143 | 94.725 | 90.9 | 89.0875
-18.1 | 98.9867 | 97.45 | 94.8571 | 91 | 89.25
-18 | 99.1 | 97.5 | 94.9875 | 91.16 | 89.4571
-17.9 | 99.125 | 97.5143 | 95.0143 | 91.35 | 89.52
-17.8 | 99.2571 | 97.68 | 95.15 | 91.5 | 89.72
-17.7 | 99.3 | 97.75 | 95.2857 | 91.6 | 89.82
-17.6 | 99.4 | 97.8 | 95.4143 | 91.7 | 90
-17.5 | 99.4 | 97.9 | 95.5 | 91.8 | 90.02
-17.4 | 99.4875 | 98 | 95.575 | 92 | 90.1857
-17.3 | 99.525 | 98.0833 | 95.7 | 92.0833 | 90.28
-17.2 | 99.6 | 98.2 | 95.7375 | 92.2 | 90.38
-17.1 | 99.6 | 98.2 | 95.8714 | 92.2625 | 90.5143
-17 | 99.7 | 98.3 | 95.9 | 92.3875 | 90.7
-16.9 | 99.7 | 98.3667 | 96.0286 | 92.525 | 90.875
-16.8 | 99.7 | 98.4 | 96.1 | 92.6727 | 91.0125
-16.7 | 99.8 | 98.5 | 96.2 | 92.83 | 91.1429
-16.6 | 99.85 | 98.5556 | 96.3333 | 92.92 | 91.2
-16.5 | 99.9 | 98.6 | 96.4 | 92.9867 | 91.3125
-16.4 | 100 | 98.6187 | 96.42 | 93.1067 | 91.4429
-16.3 | 100 | 98.6813 | 96.52 | 93.2375 | 91.575
-16.2 | 100 | 98.7636 | 96.62 | 93.3625 | 91.6125
-16.1 | 100.1 | 98.8 | 96.72 | 93.4538 | 91.7
-16 | 100.1 | 98.8222 | 96.8 | 93.5444 | 91.85
-15.9 | 100.1 | 98.9 | 96.88 | 93.6417 | 91.9
-15.8 | 100.2 | 98.9 | 97 | 93.7 | 92.0286
-15.7 | 100.2 | 99 | 97.04 | 93.7714 | 92.1625
-15.6 | 100.2 | 99 | 97.1 | 93.9 | 92.3
-15.5 | 100.3 | 99.0625 | 97.14 | 93.925 | 92.3625
-15.4 | 100.3 | 99.1 | 97.24 | 94 | 92.5
-15.3 | 100.3 | 99.2 | 97.3 | 94.0875 | 92.6
-15.2 | 100.3 | 99.2 | 97.4 | 94.2 | 92.7
-15.1 | 100.3333 | 99.2 | 97.5 | 94.2571 | 92.82
-15 | 100.4 | 99.2125 | 97.52 | 94.3 | 92.92
-14.9 | 100.4 | 99.3 | 97.6 | 94.5143 | 93
-14.8 | 100.5 | 99.3 | 97.7 | 94.65 | 93.1
-14.7 | 100.5 | 99.4 | 97.725 | 94.775 | 93.1875
-14.6 | 100.5 | 99.4 | 97.8 | 94.8 | 93.325
-14.5 | 100.5 | 99.4714 | 97.8 | 94.9067 | 93.4
-14.4 | 100.5 | 99.5 | 97.8 | 94.9733 | 93.4875
-14.3 | 100.5 | 99.5375 | 97.8467 | 95.04 | 93.6143
-14.2 | 100.6 | 99.6 | 97.9133 | 95.1133 | 93.7
-14.1 | 100.6625 | 99.6 | 97.98 | 95.2467 | 93.72
-14 | 100.7 | 99.7 | 98 | 95.34 | 93.8
-13.9 | 100.7 | 99.7 | 98.0133 | 95.4067 | 93.9
-13.8 | 100.7 | 99.7 | 98.08 | 95.4733 | 94.0143
-13.7 | 100.7 | 99.8 | 98.1 | 95.54 | 94.1
-13.6 | 100.7857 | 99.8 | 98.1385 | 95.6062 | 94.1857
-13.5 | 100.75 | 99.8 | 98.2125 | 95.6688 | 94.3
-13.4 | 100.8 | 99.8 | 98.275 | 95.7556 | 94.3429
-13.3 | 100.8 | 99.8 | 98.36 | 95.8 | 94.475
-13.2 | 100.8 | 99.8 | 98.45 | 95.8667 | 94.6
-13.1 | 100.8 | 99.9 | 98.5 | 95.96 | 94.6375
-13 | 100.8 | 99.9 | 98.5 | 96.075 | 94.7
-12.9 | 100.8 | 99.9 | 98.5 | 96.1125 | 94.7143
-12.8 | 100.8 | 99.9 | 98.525 | 96.2 | 94.8667
-12.7 | 100.8 | 99.9 | 98.6 | 96.2625 | 94.9125
-12.6 | 100.8 | 99.975 | 98.6 | 96.3 | 95.0375
-12.5 | 100.8714 | 100 | 98.7 | 96.4 | 95.1
-12.4 | 100.9 | 100 | 98.7 | 96.45 | 95.2
-12.3 | 100.9 | 100 | 98.7 | 96.5 | 95.2286
-12.2 | 100.9 | 100.1 | 98.8 | 96.525 | 95.425
-12.1 | 100.9 | 100.0571 | 98.8 | 96.6667 | 95.5
-12 | 100.9 | 100.075 | 98.8 | 96.76 | 95.5333
-11.9 | 100.9 | 100.1 | 98.9 | 96.8 | 95.6167
-11.8 | 100.9 | 100.1 | 98.9 | 96.8 | 95.7
-11.7 | 100.9 | 100.1 | 98.9 | 96.9 | 95.8
-11.6 | 100.9 | 100.1 | 98.94 | 96.9 | 95.9
-11.5 | 100.9 | 100.2 | 99 | 96.9714 | 95.98
-11.4 | 100.9 | 100.2 | 99 | 97 | 96.08
-11.3 | 100.9 | 100.2 | 99 | 97.0286 | 96.1
-11.2 | 100.9 | 100.2 | 99 | 97.1625 | 96.16
-11.1 | 100.9 | 100.2 | 99.0125 | 97.2 | 96.225
-11 | 100.975 | 100.2 | 99.1 | 97.3 | 96.35
-10.9 | 101 | 100.2 | 99.1625 | 97.3 | 96.4
-10.8 | 100.9 | 100.3 | 99.2 | 97.3 | 96.4125
-10.7 | 100.9 | 100.3 | 99.3 | 97.4286 | 96.5
-10.6 | 100.9667 | 100.3 | 99.3 | 97.5 | 96.52
-10.5 | 101 | 100.3 | 99.3 | 97.5875 | 96.5867
-10.4 | 101 | 100.3 | 99.3 | 97.6 | 96.6533
-10.3 | 101 | 100.3 | 99.3625 | 97.6 | 96.7231
-10.2 | 101 | 100.3 | 99.4 | 97.6875 | 96.8
-10.1 | 101 | 100.3 | 99.4 | 97.7143 | 96.9
-10 | 101 | 100.3 | 99.4 | 97.8 | 97
-9.9 | 101 | 100.3125 | 99.4 | 97.88 | 97
-9.8 | 101 | 100.4 | 99.4125 | 97.9 | 97.1
-9.7 | 101 | 100.4 | 99.5 | 98 | 97.2
-9.6 | 101 | 100.4 | 99.5 | 98 | 97.2571
-9.5 | 101 | 100.4 | 99.5 | 98.1 | 97.3
-9.4 | 101 | 100.4 | 99.5375 | 98.15 | 97.4
-9.3 | 101 | 100.4 | 99.6 | 98.2 | 97.4
-9.2 | 101 | 100.4 | 99.6 | 98.2 | 97.4
-9.1 | 101 | 100.475 | 99.6 | 98.2 | 97.4727
-9 | 100.9 | 100.5 | 99.6 | 98.3 | 97.5438
-8.9 | 100.9 | 100.5 | 99.6 | 98.3 | 97.6
-8.8 | 100.9 | 100.5 | 99.6857 | 98.4 | 97.6
-8.7 | 100.9 | 100.5 | 99.7 | 98.4 | 97.6875
-8.6 | 100.9 | 100.5 | 99.7 | 98.4 | 97.725
-8.5 | 100.9 | 100.5 | 99.7 | 98.4 | 97.85
-8.4 | 100.9 | 100.5 | 99.7143 | 98.4625 | 97.975
-8.3 | 100.9 | 100.5 | 99.8 | 98.5 | 98.0125
-8.2 | 100.9 | 100.5 | 99.8 | 98.6 | 98.1429
-8.1 | 100.9714 | 100.5 | 99.8889 | 98.6 | 98.2
-8 | 100.9667 | 100.5 | 99.8286 | 98.7 | 98.2
-7.9 | 100.9 | 100.5 | 99.9 | 98.7 | 98.2
-7.8 | 100.9 | 100.5 | 99.9 | 98.7 | 98.2714
-7.7 | 100.9 | 100.5 | 99.9 | 98.7286 | 98.3
-7.6 | 100.9 | 100.5 | 99.9 | 98.8 | 98.325
-7.5 | 100.9 | 100.5 | 99.9 | 98.8875 | 98.4
-7.4 | 100.9 | 100.5 | 99.9857 | 98.9 | 98.4
-7.3 | 100.9 | 100.5 | 100 | 98.9 | 98.475
-7.2 | 100.9 | 100.5 | 100 | 99 | 98.5875
-7.1 | 100.9 | 100.5 | 99.9714 | 99 | 98.6
-7 | 100.9 | 100.5 | 99.9 | 99 | 98.6
-6.9 | 100.9 | 100.5 | 99.9 | 99 | 98.6
-6.8 | 100.9 | 100.5 | 100 | 99.1 | 98.7091
-6.7 | 100.9 | 100.5 | 100 | 99.1286 | 98.8
-6.6 | 100.9 | 100.5 | 100 | 99.2 | 98.8111
-6.5 | 100.8143 | 100.5 | 100.0167 | 99.2 | 98.9
-6.4 | 100.8 | 100.5 | 100.1 | 99.2 | 98.9
-6.3 | 100.8 | 100.5 | 100.1 | 99.3 | 99
-6.2 | 100.8 | 100.5 | 100.1 | 99.3 | 99
-6.1 | 100.8 | 100.5 | 100.1 | 99.3 | 99.0571
-6 | 100.8 | 100.5 | 100.1 | 99.4 | 99.1
-5.9 | 100.8 | 100.5 | 100.1 | 99.4 | 99.1833
-5.8 | 100.8 | 100.5 | 100.1 | 99.4714 | 99.2
-5.7 | 100.8 | 100.5 | 100.2 | 99.5 | 99.2
-5.6 | 100.8 | 100.5 | 100.2 | 99.5 | 99.2
-5.5 | 100.8 | 100.45 | 100.2 | 99.5375 | 99.2286
-5.4 | 100.7625 | 100.4 | 100.2 | 99.6 | 99.3
-5.3 | 100.7 | 100.4125 | 100.2 | 99.6 | 99.3
-5.2 | 100.7 | 100.5 | 100.2 | 99.65 | 99.4
-5.1 | 100.7 | 100.5 | 100.1 | 99.7 | 99.4
-5 | 100.7 | 100.5 | 100.175 | 99.7 | 99.44
-4.9 | 100.7 | 100.5 | 100.2 | 99.7 | 99.5
-4.8 | 100.7 | 100.5 | 100.2 | 99.7 | 99.5
-4.7 | 100.7 | 100.4 | 100.2 | 99.7 | 99.5
-4.6 | 100.7 | 100.4 | 100.2 | 99.7 | 99.6
-4.5 | 100.7 | 100.4 | 100.2 | 99.7 | 99.6
-4.4 | 100.6 | 100.4 | 100.2875 | 99.7 | 99.6
-4.3 | 100.6 | 100.4 | 100.2429 | 99.7 | 99.6
-4.2 | 100.6 | 100.4 | 100.2 | 99.7 | 99.6
-4.1 | 100.6 | 100.4 | 100.2875 | 99.7 | 99.6
-4 | 100.5 | 100.4 | 100.2429 | 99.7 | 99.6
-3.9 | 100.5 | 100.4 | 100.2 | 99.7833 | 99.6429
-3.8 | 100.5 | 100.4 | 100.2857 | 99.8125 | 99.7
-3.7 | 100.5 | 100.4 | 100.2 | 99.9 | 99.8
-3.6 | 100.5 | 100.4 | 100.2 | 99.9 | 99.8
-3.5 | 100.5 | 100.3333 | 100.2857 | 99.9 | 99.8
-3.4 | 100.4875 | 100.3 | 100.2 | 99.9 | 99.8875
-3.3 | 100.4 | 100.3 | 100.2 | 99.9 | 99.8
-3.2 | 100.4 | 100.3 | 100.2 | 99.9 | 99.8714
-3.1 | 100.4 | 100.3 | 100.2 | 99.9222 | 99.9
-3 | 100.4 | 100.3 | 100.2 | 100 | 99.9375
-2.9 | 100.4 | 100.3 | 100.2 | 100 | 100
-2.8 | 100.3 | 100.3 | 100.2 | 100 | 100
-2.7 | 100.3 | 100.2375 | 100.2 | 100 | 100.0375
-2.6 | 100.3 | 100.2 | 100.2 | 100 | 100.1
-2.5 | 100.3 | 100.2 | 100.2 | 100 | 100.1
-2.4 | 100.3 | 100.2 | 100.2 | 100 | 100.1
-2.3 | 100.3 | 100.2 | 100.2 | 100 | 100.1
-2.2 | 100.2 | 100.2 | 100.2 | 100 | 100.1
-2.1 | 100.2 | 100.2 | 100.2 | 100 | 100.1
-2 | 100.2 | 100.2 | 100.2 | 100 | 100.08
-1.9 | 100.2 | 100.2 | 100.2 | 100 | 100
-1.8 | 100.1 | 100.2 | 100.2 | 100 | 100
-1.7 | 100.1 | 100.18 | 100.1 | 100 | 100
-1.6 | 100.1 | 100.1 | 100.1 | 100 | 100
-1.5 | 100.1 | 100.1 | 100.1 | 100 | 100
-1.4 | 100.1 | 100.1 | 100.1 | 100 | 100
-1.3 | 100.1 | 100.1 | 100.1 | 99.9 | 100
-1.2 | 100.1 | 100.1 | 100.1 | 99.9571 | 100
-1.1 | 100.1 | 100.06 | 100.1 | 100 | 100
-1 | 100.1 | 100 | 100.1 | 100 | 100
-0.9 | 100.1 | 100 | 100.1 | 100 | 100
-0.8 | 100 | 100 | 100.1 | 100 | 100
-0.7 | 100 | 100 | 100.1 | 100 | 100
-0.6 | 100 | 100 | 100.1 | 100 | 100
-0.5 | 99.95 | 100 | 100.1 | 100 | 100
-0.4 | 99.94 | 100 | 100.1 | 100 | 100
-0.3 | 100 | 100 | 100.03 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100.1 | 100 | 100
0.4 | 99.94 | 100 | 100.1 | 100 | 100
0.5 | 99.95 | 100 | 100.1 | 100 | 100
0.6 | 100 | 100 | 100.1 | 100 | 100
0.7 | 100 | 100 | 100.1 | 100 | 100
0.8 | 100 | 100 | 100.1 | 100 | 100
0.9 | 100.1 | 100 | 100.1 | 100 | 100
1 | 100.1 | 100 | 100.1 | 100 | 100
1.1 | 100.1 | 100 | 100.1 | 100 | 100
1.2 | 100.1 | 100.1 | 100.1 | 99.9571 | 100
1.3 | 100.1 | 100.1 | 100.1 | 99.9 | 100
1.4 | 100.1 | 100.1 | 100.1 | 100 | 100
1.5 | 100.1 | 100.1 | 100.1 | 100 | 100
1.6 | 100.1 | 100.1 | 100.1 | 100 | 100
1.7 | 100.1 | 100.1 | 100.1 | 100 | 100
1.8 | 100.1 | 100.2 | 100.2 | 100 | 100
1.9 | 100.2 | 100.2 | 100.2 | 100 | 100
2 | 100.2 | 100.2 | 100.2 | 100 | 100.1
2.1 | 100.2 | 100.2 | 100.2 | 100 | 100.1
2.2 | 100.2 | 100.2 | 100.2 | 100 | 100.1
2.3 | 100.3 | 100.2 | 100.2 | 100 | 100.1
2.4 | 100.3 | 100.2 | 100.2 | 100 | 100.1
2.5 | 100.3 | 100.2 | 100.2 | 100 | 100.1
2.6 | 100.3 | 100.2 | 100.2 | 100 | 100.1
2.7 | 100.3 | 100.2 | 100.2 | 100 | 100
2.8 | 100.3 | 100.3 | 100.2 | 100 | 100
2.9 | 100.4 | 100.3 | 100.2 | 100 | 100
3 | 100.4 | 100.3 | 100.2 | 100 | 100
3.1 | 100.4 | 100.3 | 100.2 | 99.9222 | 99.9
3.2 | 100.4 | 100.3 | 100.2 | 99.9 | 99.9
3.3 | 100.4 | 100.3 | 100.2 | 99.9 | 99.8
3.4 | 100.4875 | 100.3 | 100.2429 | 99.9 | 99.875
3.5 | 100.5 | 100.3667 | 100.3 | 99.9 | 99.8
3.6 | 100.5 | 100.4 | 100.2 | 99.9 | 99.8
3.7 | 100.5 | 100.4 | 100.2 | 99.9 | 99.8
3.8 | 100.5 | 100.4 | 100.275 | 99.8125 | 99.75
3.9 | 100.5 | 100.4 | 100.2 | 99.7833 | 99.6143
4 | 100.5 | 100.4 | 100.2167 | 99.7 | 99.6
4.1 | 100.6 | 100.4 | 100.3 | 99.7 | 99.6
4.2 | 100.6 | 100.4 | 100.2 | 99.7 | 99.6
4.3 | 100.6 | 100.4 | 100.2167 | 99.7 | 99.6
4.4 | 100.6 | 100.4 | 100.3 | 99.7 | 99.6
4.5 | 100.7 | 100.4 | 100.2 | 99.7 | 99.6
4.6 | 100.7 | 100.4 | 100.2 | 99.7 | 99.58
4.7 | 100.7 | 100.4 | 100.2 | 99.7 | 99.5
4.8 | 100.7 | 100.4625 | 100.2 | 99.7 | 99.5
4.9 | 100.7 | 100.5 | 100.2 | 99.7 | 99.5
5 | 100.7 | 100.5 | 100.2 | 99.7 | 99.42
5.1 | 100.7 | 100.5 | 100.1 | 99.7 | 99.4
5.2 | 100.7 | 100.5 | 100.2 | 99.65 | 99.4
5.3 | 100.7 | 100.475 | 100.2 | 99.6 | 99.3
5.4 | 100.7625 | 100.4 | 100.2 | 99.6 | 99.3
5.5 | 100.8 | 100.4 | 100.2 | 99.5375 | 99.3
5.6 | 100.8 | 100.5 | 100.2 | 99.5 | 99.2
5.7 | 100.8 | 100.5 | 100.2 | 99.5 | 99.2
5.8 | 100.8 | 100.5 | 100.1 | 99.4714 | 99.2
5.9 | 100.8 | 100.5 | 100.1 | 99.4 | 99.2
6 | 100.8 | 100.5 | 100.1 | 99.4 | 99.1
6.1 | 100.8 | 100.5 | 100.1 | 99.3 | 99.1
6.2 | 100.8 | 100.5 | 100.1 | 99.3 | 99
6.3 | 100.8 | 100.5 | 100.1 | 99.3 | 99
6.4 | 100.8 | 100.5 | 100.1 | 99.2 | 98.9
6.5 | 100.8143 | 100.5 | 100 | 99.2 | 98.9
6.6 | 100.9 | 100.5 | 100 | 99.2 | 98.875
6.7 | 100.9 | 100.5 | 100 | 99.1286 | 98.8
6.8 | 100.9 | 100.5 | 100 | 99.1 | 98.775
6.9 | 100.9 | 100.5 | 99.92 | 99 | 98.6429
7 | 100.9 | 100.5 | 99.9 | 99 | 98.6
7.1 | 100.9 | 100.5 | 99.94 | 99 | 98.6
7.2 | 100.9 | 100.5 | 100 | 99 | 98.5571
7.3 | 100.9 | 100.5 | 100 | 98.9 | 98.4455
7.4 | 100.9 | 100.5 | 100 | 98.9 | 98.4
7.5 | 100.9 | 100.5 | 99.9 | 98.8875 | 98.4
7.6 | 100.9 | 100.5 | 99.9 | 98.8 | 98.3
7.7 | 100.9 | 100.5 | 99.9 | 98.7286 | 98.3
7.8 | 100.9 | 100.5 | 99.9 | 98.7 | 98.2375
7.9 | 100.9 | 100.5 | 99.8778 | 98.7 | 98.2
8 | 100.9667 | 100.5 | 99.8 | 98.7 | 98.2
8.1 | 100.9714 | 100.5 | 99.8625 | 98.6 | 98.2
8.2 | 100.9 | 100.5 | 99.8125 | 98.6 | 98.1125
8.3 | 100.9 | 100.5 | 99.8 | 98.5 | 98.075
8.4 | 100.9 | 100.5 | 99.75 | 98.4625 | 97.95
8.5 | 100.9 | 100.5 | 99.7 | 98.4 | 97.825
8.6 | 100.9 | 100.5 | 99.7 | 98.4 | 97.7875
8.7 | 100.9 | 100.5 | 99.7 | 98.4 | 97.7
8.8 | 100.9 | 100.5 | 99.7 | 98.4 | 97.65
8.9 | 100.9 | 100.5 | 99.6 | 98.3 | 97.5733
9 | 100.9 | 100.5 | 99.6 | 98.3 | 97.5067
9.1 | 101 | 100.4429 | 99.6 | 98.2 | 97.4357
9.2 | 101 | 100.4 | 99.6 | 98.2 | 97.4
9.3 | 101 | 100.4 | 99.6 | 98.2 | 97.4
9.4 | 101 | 100.4 | 99.5714 | 98.15 | 97.3857
9.5 | 101 | 100.4 | 99.5 | 98.1 | 97.3
9.6 | 101 | 100.4 | 99.5 | 98 | 97.225
9.7 | 101 | 100.4 | 99.4667 | 98 | 97.1875
9.8 | 101 | 100.4 | 99.4 | 97.9 | 97.1
9.9 | 101 | 100.3 | 99.4 | 97.88 | 97.0375
10 | 101 | 100.3 | 99.4 | 97.8 | 97
10.1 | 101 | 100.3 | 99.4 | 97.7143 | 96.8778
10.2 | 101 | 100.3 | 99.4 | 97.6875 | 96.7727
10.3 | 101 | 100.3 | 99.4 | 97.6 | 96.7
10.4 | 101 | 100.3 | 99.3 | 97.6 | 96.7
10.5 | 101 | 100.3 | 99.3 | 97.5875 | 96.6533
10.6 | 100.9667 | 100.3 | 99.3 | 97.5 | 96.5818
10.7 | 100.9 | 100.3 | 99.2625 | 97.4286 | 96.49
10.8 | 100.9 | 100.3 | 99.2 | 97.3 | 96.4
10.9 | 101 | 100.2 | 99.1 | 97.3 | 96.4
11 | 100.975 | 100.2 | 99.1 | 97.3 | 96.3143
11.1 | 100.9 | 100.2 | 99.0429 | 97.2 | 96.3
11.2 | 100.9 | 100.2 | 99 | 97.1625 | 96.2
11.3 | 100.9 | 100.2 | 99 | 97.0286 | 96.1
11.4 | 100.9 | 100.2 | 99 | 97 | 96.04
11.5 | 100.9 | 100.2 | 99 | 96.9714 | 95.94
11.6 | 100.9 | 100.1714 | 98.97 | 96.9 | 95.9
11.7 | 100.9 | 100.1 | 98.9 | 96.9 | 95.8
11.8 | 100.9 | 100.1 | 98.9 | 96.8 | 95.76
11.9 | 100.9 | 100.1 | 98.8333 | 96.8 | 95.6
12 | 100.9 | 100.0429 | 98.8 | 96.76 | 95.6
12.1 | 100.9 | 100.0875 | 98.8 | 96.6667 | 95.5
12.2 | 100.9 | 100.1 | 98.8 | 96.525 | 95.4286
12.3 | 100.9 | 100.0429 | 98.7125 | 96.5 | 95.3
12.4 | 100.9 | 100 | 98.7 | 96.45 | 95.1714
12.5 | 100.8714 | 100 | 98.65 | 96.4 | 95.1
12.6 | 100.8 | 99.9429 | 98.6 | 96.3 | 95
12.7 | 100.8 | 99.9 | 98.6 | 96.2625 | 94.9714
12.8 | 100.8 | 99.9 | 98.5571 | 96.2 | 94.8286
12.9 | 100.8 | 99.9 | 98.5 | 96.1125 | 94.775
13 | 100.8 | 99.9 | 98.5 | 96.075 | 94.7
13.1 | 100.8 | 99.8857 | 98.4889 | 95.96 | 94.7
13.2 | 100.8 | 99.8 | 98.3818 | 95.8667 | 94.575
13.3 | 100.8 | 99.8 | 98.3 | 95.8 | 94.4429
13.4 | 100.8 | 99.8 | 98.3 | 95.7556 | 94.4
13.5 | 100.75 | 99.8 | 98.2438 | 95.6688 | 94.275
13.6 | 100.7857 | 99.8 | 98.1727 | 95.6062 | 94.15
13.7 | 100.7 | 99.7875 | 98.1 | 95.54 | 94.1
13.8 | 100.7 | 99.7 | 98.1 | 95.4733 | 93.9857
13.9 | 100.7 | 99.7 | 98.0467 | 95.4067 | 93.9
14 | 100.7 | 99.6875 | 97.98 | 95.34 | 93.8
14.1 | 100.6625 | 99.6 | 97.9133 | 95.2467 | 93.7
14.2 | 100.6 | 99.6 | 97.9 | 95.1133 | 93.7
14.3 | 100.5 | 99.5 | 97.8813 | 95.04 | 93.5875
14.4 | 100.5 | 99.5 | 97.8187 | 94.9733 | 93.4571
14.5 | 100.5 | 99.5 | 97.8 | 94.9067 | 93.4
14.6 | 100.5 | 99.4125 | 97.75 | 94.8 | 93.2875
14.7 | 100.5 | 99.3857 | 97.7 | 94.775 | 93.1571
14.8 | 100.5 | 99.3 | 97.7 | 94.65 | 93.1
14.9 | 100.4 | 99.3 | 97.625 | 94.5143 | 93
15 | 100.4 | 99.2 | 97.56 | 94.3 | 92.9
15.1 | 100.3333 | 99.2 | 97.4667 | 94.2571 | 92.8
15.2 | 100.3 | 99.2 | 97.4 | 94.2 | 92.68
15.3 | 100.3 | 99.1875 | 97.3286 | 94.0875 | 92.58
15.4 | 100.3 | 99.1 | 97.2 | 94 | 92.5
15.5 | 100.3 | 99.1 | 97.1667 | 93.925 | 92.4
15.6 | 100.2 | 99 | 97.1 | 93.9 | 92.2714
15.7 | 100.2 | 99 | 97.0833 | 93.7714 | 92.1286
15.8 | 100.2 | 98.9375 | 97 | 93.7 | 92.1
15.9 | 100.1 | 98.9 | 96.9 | 93.6417 | 91.9667
16 | 100.1 | 98.8923 | 96.8 | 93.5444 | 91.8429
16.1 | 100.1 | 98.8154 | 96.7 | 93.4538 | 91.7125
16.2 | 100 | 98.8 | 96.64 | 93.3625 | 91.6
16.3 | 100 | 98.7533 | 96.5 | 93.2375 | 91.5429
16.4 | 100 | 98.6867 | 96.46 | 93.1067 | 91.4125
16.5 | 99.9 | 98.62 | 96.4 | 92.9867 | 91.275
16.6 | 99.85 | 98.6 | 96.3714 | 92.92 | 91.2
16.7 | 99.8 | 98.5 | 96.2286 | 92.83 | 91.1125
16.8 | 99.7 | 98.4625 | 96.1 | 92.6727 | 90.975
16.9 | 99.7 | 98.3286 | 96.0625 | 92.525 | 90.8429
17 | 99.7 | 98.3 | 95.9375 | 92.3875 | 90.7125
17.1 | 99.6 | 98.2125 | 95.9 | 92.2625 | 90.5714
17.2 | 99.6 | 98.175 | 95.7714 | 92.2 | 90.3333
17.3 | 99.525 | 98.05 | 95.6286 | 92.0833 | 90.24
17.4 | 99.4875 | 97.98 | 95.6 | 92 | 90.1571
17.5 | 99.4 | 97.9 | 95.475 | 91.8 | 90.08
17.6 | 99.4 | 97.8 | 95.35 | 91.7 | 89.98
17.7 | 99.3 | 97.8 | 95.3 | 91.6 | 89.88
17.8 | 99.2571 | 97.64 | 95.1857 | 91.5 | 89.68
17.9 | 99.125 | 97.5857 | 95.05 | 91.35 | 89.6
18 | 99.1 | 97.5 | 94.9143 | 91.16 | 89.4286
18.1 | 98.9867 | 97.4143 | 94.7875 | 91 | 89.2875
18.2 | 98.8533 | 97.3 | 94.6571 | 90.9 | 89.1571
18.3 | 98.76 | 97.2286 | 94.525 | 90.7667 | 89.025
18.4 | 98.6867 | 97 | 94.3875 | 90.7 | 88.8833
18.5 | 98.5533 | 96.9429 | 94.2625 | 90.475 | 88.7167
18.6 | 98.42 | 96.7 | 94.1286 | 90.3 | 88.6308
18.7 | 98.2812 | 96.625 | 94 | 90.175 | 88.5
18.8 | 98.0938 | 96.4833 | 93.8625 | 90.1 | 88.36
18.9 | 97.9 | 96.3 | 93.6571 | 89.9 | 88.225
19 | 97.7 | 96.25 | 93.5 | 89.7429 | 88.0375
19.1 | 97.4333 | 95.96 | 93.325 | 89.525 | 87.8545
19.2 | 97.1333 | 95.7667 | 93.1286 | 89.3714 | 87.7125
19.3 | 96.8 | 95.55 | 93 | 89.1889 | 87.575
19.4 | 96.2667 | 95.2333 | 92.7625 | 88.9909 | 87.45
19.5 | 95.6 | 95 | 92.5571 | 88.8444 | 87.3143
19.6 | 94.8 | 94.6875 | 92.4111 | 88.725 | 87.0875
19.7 | 93.2 | 94.3333 | 92.1429 | 88.6111 | 86.9571
19.8 | 90.9 | 93.98 | 91.875 | 88.4364 | 86.75
19.9 | 87.9 | 93.42 | 91.7125 | 88.3 | 86.6
20 | 82.6333 | 92.8 | 91.425 | 88.05 | 86.3625
20.1 | 76.2 | 91.96 | 91.1429 | 87.8429 | 86.2286
20.2 | 68.6 | 90.8571 | 90.8375 | 87.66 | 86.1
20.3 | 60.1625 | 89.34 | 90.55 | 87.5125 | 85.9625
20.4 | 51.5533 | 87.32 | 90.2286 | 87.25 | 85.8375
20.5 | 42.6867 | 84.08 | 89.8375 | 87.0429 | 85.7
20.6 | 35.0846 | 79.46 | 89.525 | 86.825 | 85.4429
20.7 | 28.275 | 73.45 | 89.0714 | 86.65 | 85.2286
20.8 | 23.4 | 66.2833 | 88.55 | 86.3857 | 85.0714
20.9 | 20.1 | 57.7714 | 87.975 | 86.16 | 84.875
21 | 17.6067 | 49.9667 | 87.3143 | 86.0125 | 84.6
21.1 | 15.8733 | 42.0333 | 86.4875 | 85.75 | 84.4714
21.2 | 14.7308 | 35.1167 | 85.225 | 85.4857 | 84.275
21.3 | 13.925 | 29.4143 | 83.7 | 85.225 | 84.1125
21.4 | 13.4 | 24.7 | 81.4571 | 84.95 | 83.96
21.5 | 12.9444 | 21.34 | 78.5125 | 84.7333 | 83.725
21.6 | 12.6231 | 19.1 | 74.3429 | 84.4714 | 83.4667
21.7 | 12.3667 | 17.4429 | 69.025 | 84.2 | 83.3
21.8 | 12.1385 | 16.3 | 62.5875 | 83.875 | 83.1
21.9 | 11.8909 | 15.4143 | 55.5 | 83.575 | 82.9667
22 | 11.6571 | 14.7625 | 48.575 | 83.3143 | 82.6667
22.1 | 11.525 | 14.325 | 41.7125 | 82.975 | 82.4
22.2 | 11.4 | 13.95 | 35.8857 | 82.65 | 82.1857
22.3 | 11.2625 | 13.6286 | 30.9 | 82.2143 | 81.925
22.4 | 11.075 | 13.3714 | 26.8875 | 81.95 | 81.75
22.5 | 10.9 | 13.1 | 23.9286 | 81.4625 | 81.4857
22.6 | 10.7714 | 12.8286 | 21.7 | 81.05 | 81.3125
22.7 | 10.6375 | 12.5714 | 20.2 | 80.5857 | 81.1571
22.8 | 10.5 | 12.35 | 18.9429 | 80.18 | 80.85
22.9 | 10.3714 | 12.1286 | 18.15 | 79.5 | 80.5429
23 | 10.2375 | 11.9714 | 17.5 | 78.65 | 80.275
23.1 | 10.1 | 11.75 | 16.95 | 77.7 | 80.0143
23.2 | 9.9714 | 11.625 | 16.5 | 76.42 | 79.75
23.3 | 9.8375 | 11.3857 | 16.1143 | 74.84 | 79.4875
23.4 | 9.7 | 11.25 | 15.8167 | 72.94 | 79.2714
23.5 | 9.5909 | 11.125 | 15.5429 | 70.2 | 78.95
23.6 | 9.5 | 10.975 | 15.2 | 66.1 | 78.675
23.7 | 9.4 | 10.7571 | 14.9429 | 61.7875 | 78.4143
23.8 | 9.35 | 10.6875 | 14.675 | 56.9125 | 78.08
23.9 | 9.2143 | 10.5571 | 14.5 | 51.5 | 77.86
24 | 9.175 | 10.35 | 14.2429 | 45.9857 | 77.5
24.1 | 9.0429 | 10.2 | 13.975 | 40.7 | 77.14
24.2 | 8.9125 | 10.0625 | 13.8 | 36 | 76.82
24.3 | 8.875 | 9.9286 | 13.5714 | 32.1571 | 76.425
24.4 | 8.75 | 9.9 | 13.375 | 28.8625 | 76.0857
24.5 | 8.7 | 9.7625 | 13.2125 | 26.4 | 75.66
24.6 | 8.575 | 9.6286 | 12.975 | 24.5714 | 75.2
24.7 | 8.5 | 9.5 | 12.7857 | 23.2 | 74.6875
24.8 | 8.4 | 9.4 | 12.6125 | 22.2286 | 74.1429
24.9 | 8.4 | 9.3 | 12.375 | 21.4875 | 73.5
25 | 8.2 | 9.2 | 12.2429 | 20.8 | 72.775
25.1 | 8.2 | 9.1 | 12.1125 | 20.3857 | 71.9
25.2 | 8.1 | 9 | 11.95 | 19.9125 | 70.7
25.3 | 8 | 8.9222 | 11.75 | 19.5375 | 69.0714
25.4 | 7.9 | 8.8111 | 11.6143 | 19.2571 | 66.9375
25.5 | 7.9 | 8.6929 | 11.4857 | 18.9 | 64.375
25.6 | 7.8 | 8.6214 | 11.35 | 18.6286 | 61.075
25.7 | 7.7 | 8.5533 | 11.2143 | 18.2714 | 57.1286
25.8 | 7.6167 | 8.4733 | 10.9857 | 18 | 53
25.9 | 7.6 | 8.34 | 10.85 | 17.7286 | 48.2
26 | 7.52 | 8.2533 | 10.8 | 17.4857 | 43.5857
26.1 | 7.5 | 8.1818 | 10.6857 | 17.25 | 39.3833
26.2 | 7.4 | 8.0857 | 10.55 | 17.0143 | 35.74
26.3 | 7.3714 | 8 | 10.4143 | 16.7857 | 32.5
26.4 | 7.3 | 7.9143 | 10.275 | 16.6 | 29.92
26.5 | 7.2 | 7.8857 | 10.1429 | 16.3286 | 27.8875
26.6 | 7.175 | 7.8 | 10.0125 | 16.0857 | 26.45
26.7 | 7.1 | 7.7143 | 9.9 | 15.9 | 25.3
26.8 | 7.0778 | 7.6857 | 9.8143 | 15.7143 | 24.3
26.9 | 7 | 7.6 | 9.7 | 15.4857 | 23.5
27 | 7 | 7.52 | 9.6429 | 15.3111 | 23.1
27.1 | 6.9 | 7.5 | 9.5125 | 15.1167 | 22.5
27.2 | 6.9 | 7.4 | 9.3857 | 14.8857 | 22.1286
27.3 | 6.8286 | 7.32 | 9.3 | 14.7 | 21.7571
27.4 |  | 7.3 | 9.225 | 14.525 | 21.4
27.5 |  | 7.225 | 9.0875 | 14.3875 | 21.1286
27.6 |  | 7.1875 | 9 | 14.2143 | 20.775
27.7 |  | 7.1 | 8.925 | 14.025 | 20.5143
27.8 |  | 7.025 | 8.7875 | 13.8 | 20.25
27.9 |  | 7 | 8.7 | 13.6625 | 19.975
28 |  | 6.9 | 8.625 | 13.5286 | 19.725
28.1 |  | 6.9 | 8.6 | 13.4 | 19.5286
28.2 |  | 6.8333 | 8.4625 | 13.225 | 19.3
28.3 |  | 6.8 | 8.4 | 13.0286 | 19.0625
28.4 |  |  | 8.3 | 12.9 | 18.7857
28.5 |  |  | 8.2 | 12.7625 | 18.6
28.6 |  |  | 8.1375 | 12.6286 | 18.425
28.7 |  |  | 8.1 | 12.5 | 18.1571
28.8 |  |  | 8.01 | 12.3625 | 18
28.9 |  |  | 7.9 | 12.2375 | 17.7625
29 |  |  | 7.8867 | 12.06 | 17.575
29.1 |  |  | 7.82 | 11.9 | 17.4
29.2 |  |  | 7.7067 | 11.84 | 17.15
29.3 |  |  | 7.6 | 11.7 | 16.9429
29.4 |  |  | 7.5857 | 11.6 | 16.8
29.5 |  |  | 7.5 | 11.4833 | 16.62
29.6 |  |  | 7.4375 | 11.3 | 16.4667
29.7 |  |  |  | 11.2 | 16.2667
29.8 |  |  |  | 11.1 | 16.1
29.9 |  |  |  | 10.9333 | 15.8714
30 |  |  |  | 10.9 | 15.7375
30.1 |  |  |  | 10.74 | 15.6
30.2 |  |  |  | 10.6857 | 15.375
30.3 |  |  |  | 10.55 | 15.2429
30.4 |  |  |  | 10.4143 | 15.1125
30.5 |  |  |  | 10.3 | 14.95
30.6 |  |  |  | 10.25 | 14.7429
30.7 |  |  |  | 10.125 | 14.6125
30.8 |  |  |  | 10 | 14.475
30.9 |  |  |  | 9.95 | 14.3429
31 |  |  |  | 9.8 | 14.2125
31.1 |  |  |  | 9.75 | 14.0667
31.2 |  |  |  | 9.7 | 13.9143
31.3 |  |  |  | 9.58 | 13.7875
31.4 |  |  |  | 9.5 | 13.6571
31.5 |  |  |  | 9.4 | 13.525
31.6 |  |  |  | 9.2875 | 13.4
31.7 |  |  |  | 9.2 | 13.3
31.8 |  |  |  | 9.15 | 13.2286
31.9 |  |  |  | 9.06 | 13.1
32 |  |  |  | 9 | 12.9714
32.1 |  |  |  | 8.8625 | 12.8375
32.2 |  |  |  | 8.8 | 12.7125
32.3 |  |  |  | 8.7 | 12.575
32.4 |  |  |  |  | 12.4429
32.5 |  |  |  |  | 12.3125
32.6 |  |  |  |  | 12.2667
32.7 |  |  |  |  | 12.125
32.8 |  |  |  |  | 11.9857
32.9 |  |  |  |  | 11.85
33 |  |  |  |  | 11.725
33.1 |  |  |  |  | 11.6875
33.2 |  |  |  |  | 11.5571
33.3 |  |  |  |  | 11.425
33.4 |  |  |  |  | 11.3
33.5 |  |  |  |  | 11.2571
33.6 |  |  |  |  | 11.125
33.7 |  |  |  |  | 11
33.8 |  |  |  |  | 10.975
33.9 |  |  |  |  | 10.86
34 |  |  |  |  | 10.7556
34.1 |  |  |  |  | 10.7
34.2 |  |  |  |  | 10.6
34.3 |  |  |  |  | 10.525
34.4 |  |  |  |  | 10.3875
34.5 |  |  |  |  | 10.2625
34.6 |  |  |  |  | 10.2
34.7 |  |  |  |  | 10.1
34.8 |  |  |  |  | 10.0571
34.9 |  |  |  |  | 10
35 |  |  |  |  | 9.9111

## Wedge 15 Output Factors
Output factors
Wedge 15 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.891 | 0.903 | 0.91 | 0.917 | 0.923 | 0.928 | 0.932 | 0.936
| 4 | 0.904 | 0.918 | 0.926 | 0.936 | 0.943 | 0.95 | 0.955 | 0.961
| 5 | 0.91 | 0.927 | 0.937 | 0.949 | 0.958 | 0.966 | 0.972 | 0.978
| 7 | 0.921 | 0.941 | 0.953 | 0.967 | 0.979 | 0.991 | 0.998 | 1.007
| 10 | 0.931 | 0.952 | 0.966 | 0.985 | 0.999 | 1.015 | 1.025 | 1.038
| 15 | 0.939 | 0.962 | 0.978 | 1 | 1.019 | 1.039 | 1.054 | 1.072
| 20 | 0.943 | 0.967 | 0.984 | 1.009 | 1.03 | 1.054 | 1.073 | 1.096
| 30 | 0.949 | 0.975 | 0.994 | 1.02 | 1.047 | 1.076 | 1.1 | 1.13
| 40 | 0.953 | 0.98 | 0.999 | 1.028 | 1.057 | 1.089 | 1.116 | 1.15

## Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 32.9 | 39.3 | 49.7 | 58.9 | 62.5
0.1 | 36.9 | 43.2143 | 53.5 | 62.65 | 65.9
0.2 | 43.8 | 49.275 | 59.7 | 68.275 | 71.075
0.3 | 53.55 | 58.68 | 67.5 | 74.95 | 77.4
0.4 | 62.15 | 66.4 | 74.3 | 80.9667 | 82.975
0.5 | 69.2 | 73.2625 | 80.1 | 85.1 | 87.175
0.6 | 75.25 | 78.8 | 84.5 | 89 | 90.55
0.7 | 80.0333 | 83.2333 | 88.06 | 91.6 | 93.025
0.8 | 84.1 | 86.8 | 90.65 | 93.9 | 94.95
0.9 | 87.6 | 89.8 | 93.05 | 95.7 | 96.425
1 | 90.4 | 92.04 | 94.95 | 97 | 97.6
1.1 | 92.6 | 94.1 | 96.375 | 98 | 98.4333
1.2 | 94.45 | 95.65 | 97.4 | 98.7667 | 99.12
1.3 | 96.1 | 97 | 98.4 | 99.3 | 99.55
1.4 | 97.1 | 98 | 98.9333 | 99.7 | 99.86
1.5 | 98.2 | 98.7143 | 99.625 | 99.9 | 100
1.6 | 98.78 | 99.3 | 99.875 | 100 | 100
1.7 | 99.26 | 99.6667 | 100 | 100 | 100
1.8 | 99.8 | 99.9 | 100 | 100 | 99.975
1.9 | 99.9 | 100 | 100 | 99.9 | 99.8
2 | 100 | 100 | 99.9 | 99.8 | 99.64
2.1 | 100 | 100 | 99.76 | 99.6 | 99.5
2.2 | 99.9 | 99.9 | 99.525 | 99.3 | 99.15
2.3 | 99.7 | 99.825 | 99.375 | 99.05 | 98.96
2.4 | 99.52 | 99.52 | 99.125 | 98.7 | 98.58
2.5 | 99.3 | 99.2667 | 98.84 | 98.52 | 98.2667
2.6 | 98.95 | 98.96 | 98.5 | 98.1667 | 98
2.7 | 98.6 | 98.6333 | 98.2 | 97.775 | 97.7
2.8 | 98.225 | 98.16 | 97.8333 | 97.425 | 97.3
2.9 | 97.8857 | 97.7857 | 97.5 | 97.175 | 96.9
3 | 97.4 | 97.48 | 97.12 | 96.825 | 96.56
3.1 | 96.9625 | 97.1286 | 96.8 | 96.45 | 96.325
3.2 | 96.58 | 96.8 | 96.4571 | 96.075 | 95.975
3.3 | 96.2 | 96.46 | 96.05 | 95.75 | 95.65
3.4 | 95.7 | 96 | 95.7571 | 95.35 | 95.225
3.5 | 95.16 | 95.56 | 95.425 | 94.98 | 95
3.6 | 94.58 | 95.1 | 95.05 | 94.58 | 94.55
3.7 | 94.22 | 94.86 | 94.56 | 94.25 | 94.2
3.8 | 93.9 | 94.5286 | 94.3 | 94 | 93.8333
3.9 | 93.3333 | 94.1 | 94 | 93.65 | 93.5
4 | 92.9 | 93.6429 | 93.6 | 93.2 | 93.2
4.1 | 92.5 | 93.2 | 93.2143 | 93 | 92.8
4.2 | 92.125 | 92.92 | 92.9 | 92.6 | 92.45
4.3 | 91.525 | 92.52 | 92.42 | 92.22 | 92.1
4.4 | 91.225 | 92.0333 | 92.1 | 91.875 | 91.7
4.5 | 90.725 | 91.72 | 91.75 | 91.55 | 91.36
4.6 | 90.1 | 91.22 | 91.4 | 91.15 | 91.12
4.7 | 89.7667 | 90.8667 | 91 | 90.775 | 90.6667
4.8 | 89.2 | 90.475 | 90.68 | 90.48 | 90.4
4.9 | 88.925 | 89.975 | 90.2333 | 90.1667 | 90.18
5 | 88.4 | 89.7333 | 89.95 | 89.7 | 89.75
5.1 | 87.925 | 89.35 | 89.6 | 89.5143 | 89.3
5.2 | 87.64 | 88.85 | 89.28 | 89.1571 | 89.075
5.3 | 87.06 | 88.6 | 88.9 | 88.84 | 88.7
5.4 | 86.5667 | 88.28 | 88.5 | 88.46 | 88.325
5.5 | 86.22 | 87.7286 | 88.2167 | 88.12 | 87.96
5.6 | 85.75 | 87.4 | 87.8 | 87.68 | 87.64
5.7 | 85.25 | 87.0556 | 87.4286 | 87.45 | 87.28
5.8 | 84.8667 | 86.66 | 87.06 | 87.1 | 87.0429
5.9 | 84.46 | 86.3143 | 86.7 | 86.6714 | 86.68
6 | 84.06 | 85.8571 | 86.3571 | 86.3286 | 86.3429
6.1 | 83.4857 | 85.6 | 86 | 85.98 | 86
6.2 | 83.08 | 85.12 | 85.7556 | 85.65 | 85.575
6.3 | 82.6222 | 84.8333 | 85.32 | 85.35 | 85.25
6.4 | 82.28 | 84.475 | 85.0143 | 85.05 | 84.9714
6.5 | 81.8571 | 84.02 | 84.7286 | 84.72 | 84.7
6.6 | 81.4 | 83.62 | 84.2 | 84.3375 | 84.375
6.7 | 80.875 | 83.24 | 83.95 | 84.02 | 84.02
6.8 | 80.6 | 82.96 | 83.55 | 83.64 | 83.66
6.9 | 80.25 | 82.55 | 83.2571 | 83.3 | 83.32
7 | 79.9 | 82.1125 | 82.8667 | 83 | 83.0429
7.1 | 79.3714 | 81.78 | 82.5571 | 82.6667 | 82.62
7.2 | 79.0286 | 81.5 | 82.225 | 82.32 | 82.3286
7.3 | 78.6 | 81.0571 | 81.85 | 81.98 | 81.98
7.4 | 78.14 | 80.675 | 81.44 | 81.66 | 81.6
7.5 | 77.6857 | 80.35 | 81.22 | 81.3333 | 81.2
7.6 | 77.35 | 79.9714 | 80.88 | 80.9429 | 80.9429
7.7 | 76.9333 | 79.5 | 80.5 | 80.72 | 80.66
7.8 | 76.58 | 79.18 | 80.1 | 80.45 | 80.28
7.9 | 76.0625 | 78.88 | 79.78 | 79.9429 | 80.05
8 | 75.74 | 78.4375 | 79.52 | 79.7286 | 79.7
8.1 | 75.2625 | 78.14 | 79.2 | 79.425 | 79.275
8.2 | 74.94 | 77.78 | 78.825 | 79.08 | 79.1
8.3 | 74.6 | 77.3857 | 78.4667 | 78.8571 | 78.7889
8.4 | 74.15 | 77.2 | 78.15 | 78.4286 | 78.4
8.5 | 73.8 | 76.76 | 77.85 | 78.08 | 78.1
8.6 | 73.375 | 76.375 | 77.55 | 77.8 | 77.7714
8.7 | 73 | 76.125 | 77.1 | 77.5571 | 77.5
8.8 | 72.58 | 75.6429 | 76.8 | 77.1857 | 77.2
8.9 | 72.2111 | 75.4 | 76.5667 | 76.95 | 76.85
9 | 71.86 | 75.06 | 76.3 | 76.5 | 76.65
9.1 | 71.4429 | 74.6 | 75.9 | 76.3 | 76.32
9.2 | 71.1 | 74.26 | 75.6 | 75.9143 | 75.98
9.3 | 70.6875 | 73.7333 | 75.26 | 75.7 | 75.75
9.4 | 70.3 | 73.48 | 74.925 | 75.38 | 75.3571
9.5 | 70.05 | 73.2143 | 74.62 | 75.0857 | 75.125
9.6 | 69.72 | 72.86 | 74.3 | 74.7 | 74.76
9.7 | 69.2667 | 72.42 | 74.02 | 74.46 | 74.4
9.8 | 68.8 | 72.1714 | 73.6143 | 74.1 | 74.2
9.9 | 68.54 | 71.875 | 73.375 | 73.86 | 73.85
10 | 68.14 | 71.6 | 73.1 | 73.5 | 73.6
10.1 | 67.7429 | 71.2 | 72.64 | 73.2 | 73.2571
10.2 | 67.38 | 70.85 | 72.44 | 72.96 | 72.9429
10.3 | 67.075 | 70.5 | 72.125 | 72.6286 | 72.6833
10.4 | 66.62 | 70.12 | 71.8 | 72.2286 | 72.35
10.5 | 66.32 | 69.8333 | 71.4571 | 72 | 72
10.6 | 65.98 | 69.45 | 71.3143 | 71.66 | 71.775
10.7 | 65.575 | 69.1667 | 70.9 | 71.4333 | 71.45
10.8 | 65.275 | 68.875 | 70.625 | 71.15 | 71.125
10.9 | 65.025 | 68.55 | 70.2 | 70.8429 | 70.85
11 | 64.6333 | 68.2 | 69.94 | 70.56 | 70.575
11.1 | 64.35 | 67.88 | 69.6111 | 70.2286 | 70.375
11.2 | 63.92 | 67.56 | 69.38 | 70.05 | 70.04
11.3 | 63.625 | 67.15 | 69.1286 | 69.725 | 69.68
11.4 | 63.24 | 66.9 | 68.8 | 69.46 | 69.5
11.5 | 62.96 | 66.5 | 68.54 | 69.125 | 69.1
11.6 | 62.525 | 66.3 | 68.22 | 68.8 | 68.9
11.7 | 62.2667 | 65.9 | 67.9143 | 68.525 | 68.55
11.8 | 61.94 | 65.6 | 67.65 | 68.3 | 68.4
11.9 | 61.6 | 65.35 | 67.3429 | 67.9714 | 68.15
12 | 61.3571 | 65.025 | 66.96 | 67.6857 | 67.7
12.1 | 60.9 | 64.75 | 66.65 | 67.45 | 67.5
12.2 | 60.6667 | 64.475 | 66.38 | 67.14 | 67.2333
12.3 | 60.225 | 64.125 | 66.1429 | 66.7571 | 66.84
12.4 | 60 | 63.85 | 65.9 | 66.5 | 66.6
12.5 | 59.62 | 63.425 | 65.6375 | 66.3 | 66.375
12.6 | 59.32 | 63.175 | 65.34 | 65.98 | 66.05
12.7 | 59.04 | 62.925 | 64.96 | 65.6429 | 65.8
12.8 | 58.725 | 62.6 | 64.7 | 65.5 | 65.56
12.9 | 58.38 | 62.28 | 64.48 | 65.175 | 65.24
13 | 58.025 | 62 | 64.16 | 64.8375 | 65
13.1 | 57.8 | 61.7 | 63.8571 | 64.62 | 64.7
13.2 | 57.4 | 61.4 | 63.6 | 64.2143 | 64.5
13.3 | 57.15 | 61.2 | 63.2667 | 63.98 | 64.1
13.4 | 56.925 | 61 | 63 | 63.75 | 63.96
13.5 | 56.54 | 60.55 | 62.76 | 63.525 | 63.66
13.6 | 56.25 | 60.2 | 62.44 | 63.36 | 63.35
13.7 | 56.02 | 59.85 | 62.2 | 63.0143 | 63.075
13.8 | 55.6 | 59.66 | 61.9571 | 62.675 | 62.85
13.9 | 55.46 | 59.425 | 61.68 | 62.5143 | 62.5333
14 | 55.1143 | 59.1 | 61.4333 | 62.2286 | 62.28
14.1 | 54.75 | 58.7667 | 61.125 | 61.9 | 62.075
14.2 | 54.5 | 58.475 | 60.9 | 61.675 | 61.825
14.3 | 54.25 | 58.225 | 60.64 | 61.4429 | 61.5667
14.4 | 53.9 | 57.975 | 60.3286 | 61.175 | 61.3
14.5 | 53.6333 | 57.7333 | 60.1 | 60.86 | 61.08
14.6 | 53.4 | 57.5 | 59.7857 | 60.66 | 60.8
14.7 | 53.05 | 57.05 | 59.5 | 60.375 | 60.55
14.8 | 52.8333 | 56.9 | 59.2714 | 60.175 | 60.4
14.9 | 52.6 | 56.6 | 59.1 | 59.85 | 60.05
15 | 52.3 | 56.3 | 58.8 | 59.6 | 59.8
15.1 | 52 | 56.1 | 58.52 | 59.375 | 59.575
15.2 | 51.8333 | 55.75 | 58.2667 | 59.1 | 59.4
15.3 | 51.55 | 55.5 | 57.95 | 58.9 | 59.075
15.4 | 51.3 | 55.25 | 57.7 | 58.72 | 58.925
15.5 | 50.95 | 55 | 57.48 | 58.44 | 58.6
15.6 | 50.7 | 54.7 | 57.2 | 58.2 | 58.35
15.7 | 50.42 | 54.525 | 57 | 58 | 58.1
15.8 | 50.26 | 54.2 | 56.7333 | 57.725 | 57.88
15.9 | 49.94 | 54.025 | 56.45 | 57.45 | 57.625
16 | 49.65 | 53.675 | 56.25 | 57.125 | 57.4
16.1 | 49.325 | 53.44 | 56 | 56.9 | 57.15
16.2 | 49.15 | 53.175 | 55.725 | 56.7333 | 56.8
16.3 | 48.875 | 52.925 | 55.5 | 56.52 | 56.5667
16.4 | 48.5333 | 52.7 | 55.3 | 56.24 | 56.4
16.5 | 48.3 | 52.44 | 55.06 | 56 | 56.15
16.6 | 48.05 | 52.2 | 54.84 | 55.7714 | 55.9
16.7 | 47.7 | 51.9 | 54.58 | 55.5429 | 55.7
16.8 | 47.55 | 51.6 | 54.275 | 55.35 | 55.5
16.9 | 47.4 | 51.35 | 54.12 | 55.1222 | 55.22
17 | 47.06 | 51.125 | 53.8667 | 54.8 | 55.05
17.1 | 46.825 | 50.875 | 53.58 | 54.5857 | 54.825
17.2 | 46.56 | 50.625 | 53.3556 | 54.36 | 54.5
17.3 | 46.325 | 50.375 | 53.08 | 54.125 | 54.2333
17.4 | 46.06 | 50.125 | 52.8 | 53.8857 | 54.075
17.5 | 45.825 | 49.8 | 52.6 | 53.6 | 53.825
17.6 | 45.6 | 49.625 | 52.375 | 53.3875 | 53.575
17.7 | 45.325 | 49.375 | 52.15 | 53.18 | 53.325
17.8 | 45.1333 | 49.24 | 51.95 | 52.9 | 53.1
17.9 | 44.9 | 49 | 51.7 | 52.6444 | 53
18 | 44.625 | 48.75 | 51.5 | 52.525 | 52.65
18.1 | 44.4 | 48.4 | 51.22 | 52.3 | 52.5
18.2 | 44.15 | 48.15 | 50.9857 | 52.0571 | 52.3
18.3 | 43.9 | 47.9 | 50.74 | 51.8286 | 52.05
18.4 | 43.65 | 47.76 | 50.5429 | 51.7 | 51.8
18.5 | 43.5 | 47.575 | 50.325 | 51.3667 | 51.6
18.6 | 43.3 | 47.3 | 50.0429 | 51.2 | 51.4
18.7 | 43.05 | 47.125 | 49.95 | 50.96 | 51.25
18.8 | 42.82 | 46.975 | 49.6857 | 50.6667 | 51.02
18.9 | 42.5667 | 46.675 | 49.5 | 50.5 | 50.75
19 | 42.4 | 46.45 | 49.28 | 50.325 | 50.52
19.1 | 42.2 | 46.2 | 49.06 | 50.2 | 50.325
19.2 | 42.025 | 46 | 48.825 | 49.9833 | 50.175
19.3 | 41.675 | 45.75 | 48.6 | 49.7 | 49.925
19.4 | 41.525 | 45.5 | 48.3571 | 49.5 | 49.78
19.5 | 41.275 | 45.35 | 48.175 | 49.3 | 49.525
19.6 | 41.025 | 45.1 | 47.9444 | 49.1 | 49.28
19.7 | 40.9 | 44.8 | 47.72 | 48.8 | 49.05
19.8 | 40.725 | 44.7 | 47.54 | 48.58 | 48.9
19.9 | 40.5 | 44.45 | 47.325 | 48.4 | 48.6
20 | 40.3 | 44.225 | 47.1429 | 48.2 | 48.45
20.1 | 40.05 | 43.975 | 46.88 | 47.9857 | 48.3
20.2 | 39.9 | 43.725 | 46.7333 | 47.86 | 48.075
20.3 | 39.65 | 43.575 | 46.44 | 47.625 | 47.9
20.4 | 39.5 | 43.3333 | 46.2429 | 47.4 | 47.6667
20.5 | 39.25 | 43.2 | 46 | 47.1333 | 47.4
20.6 | 39.1 | 43 | 45.8714 | 46.96 | 47.225
20.7 | 38.85 | 42.8 | 45.7 | 46.825 | 47.0333
20.8 | 38.62 | 42.5667 | 45.4714 | 46.6167 | 46.78
20.9 | 38.4 | 42.32 | 45.2 | 46.4286 | 46.6
21 | 38.225 | 42.1667 | 44.9889 | 46.275 | 46.5
21.1 | 38 | 41.975 | 44.86 | 46.0333 | 46.25
21.2 | 37.825 | 41.8 | 44.6571 | 45.78 | 46.1
21.3 | 37.68 | 41.525 | 44.4286 | 45.6429 | 45.85
21.4 | 37.425 | 41.375 | 44.3 | 45.38 | 45.7
21.5 | 37.28 | 41.125 | 44.02 | 45.225 | 45.45
21.6 | 37.15 | 40.875 | 43.8571 | 45.025 | 45.2
21.7 | 36.88 | 40.7 | 43.6667 | 44.8429 | 45.1
21.8 | 36.7 | 40.55 | 43.46 | 44.7 | 44.875
21.9 | 36.5 | 40.4 | 43.2714 | 44.475 | 44.675
22 | 36.3 | 40.15 | 43.02 | 44.3 | 44.5
22.1 | 36.06 | 40 | 42.8857 | 44.0857 | 44.3
22.2 | 35.9 | 39.8 | 42.72 | 43.8444 | 44.1333
22.3 | 35.7 | 39.6 | 42.5 | 43.725 | 43.9667
22.4 | 35.6 | 39.375 | 42.2333 | 43.4714 | 43.7167
22.5 | 35.375 | 39.125 | 42.1 | 43.3 | 43.5333
22.6 | 35.225 | 39 | 41.9286 | 43.1 | 43.4
22.7 | 34.975 | 38.875 | 41.76 | 43 | 43.2
22.8 | 34.825 | 38.625 | 41.475 | 42.76 | 43
22.9 | 34.675 | 38.4 | 41.32 | 42.6 | 42.85
23 | 34.45 | 38.225 | 41.1714 | 42.4 | 42.7
23.1 | 34.3 | 38.075 | 40.98 | 42.2667 | 42.42
23.2 | 34.15 | 37.9 | 40.8143 | 42.075 | 42.3
23.3 | 33.9 | 37.65 | 40.6 | 41.9 | 42.2
23.4 | 33.8 | 37.5 | 40.5286 | 41.68 | 41.98
23.5 | 33.6 | 37.25 | 40.3111 | 41.45 | 41.7667
23.6 | 33.5 | 37.1 | 40.0571 | 41.375 | 41.5286
23.7 | 33.3 | 37 | 39.9 | 41.2429 | 41.42
23.8 | 33.05 | 36.86 | 39.78 | 41.0286 | 41.2286
23.9 | 32.9 | 36.62 | 39.5833 | 40.82 | 41.025
24 | 32.775 | 36.475 | 39.3571 | 40.68 | 40.85
24.1 | 32.6 | 36.225 | 39.2 | 40.46 | 40.7
24.2 | 32.475 | 36.175 | 39.06 | 40.28 | 40.6
24.3 | 32.3 | 35.975 | 38.94 | 40.1714 | 40.4
24.4 | 32.075 | 35.725 | 38.66 | 39.9571 | 40.22
24.5 | 31.9 | 35.6 | 38.55 | 39.8 | 40.0286
24.6 | 31.775 | 35.425 | 38.375 | 39.6 | 39.875
24.7 | 31.6 | 35.275 | 38.175 | 39.425 | 39.7
24.8 | 31.5 | 35.1 | 37.94 | 39.2429 | 39.48
24.9 | 31.3 | 34.9 | 37.8 | 39.1 | 39.36
25 | 31.15 | 34.8 | 37.68 | 38.9 | 39.1
25.1 | 31 | 34.6 | 37.4857 | 38.7 | 38.9875
25.2 | 30.84 | 34.35 | 37.4 | 38.62 | 38.84
25.3 | 30.66 | 34.2 | 37.2 | 38.4286 | 38.6429
25.4 | 30.5 | 34.0667 | 37.04 | 38.28 | 38.5
25.5 | 30.3833 | 33.9 | 36.8429 | 38.0833 | 38.3
25.6 | 30.2 | 33.775 | 36.75 | 37.9 | 38.2222
25.7 | 30.075 | 33.6 | 36.5 | 37.775 | 38.025
25.8 | 29.9 | 33.475 | 36.35 | 37.6 | 37.8714
25.9 | 29.7 | 33.3 | 36.15 | 37.45 | 37.68
26 | 29.6 | 33.1 | 35.96 | 37.3 | 37.525
26.1 | 29.44 | 32.925 | 35.8571 | 37.15 | 37.4
26.2 | 29.3 | 32.8 | 35.7 | 37 | 37.2714
26.3 | 29.14 | 32.65 | 35.5667 | 36.8 | 37.1
26.4 | 29 | 32.5 | 35.375 | 36.7 | 36.925
26.5 | 28.9 | 32.35 | 35.2 | 36.4889 | 36.75
26.6 | 28.7 | 32.2 | 35.025 | 36.36 | 36.5857
26.7 | 28.6 | 32 | 34.9 | 36.1571 | 36.5
26.8 | 28.5 | 31.9 | 34.8 | 36.06 | 36.275
26.9 | 28.25 | 31.75 | 34.6 | 35.8667 | 36.15
27 | 28.2 | 31.6 | 34.5 | 35.7143 | 35.9857
27.1 | 27.975 | 31.4 | 34.3 | 35.6 | 35.82
27.2 | 27.825 | 31.3 | 34.175 | 35.4 | 35.7
27.3 | 27.7 | 31.16 | 33.975 | 35.2667 | 35.52
27.4 | 27.5667 | 31 | 33.9 | 35.0286 | 35.4
27.5 | 27.4333 | 30.875 | 33.6571 | 34.9571 | 35.275
27.6 | 27.3 | 30.65 | 33.5333 | 34.7857 | 35
27.7 | 27.175 | 30.5 | 33.42 | 34.62 | 34.925
27.8 | 27 | 30.4 | 33.3 | 34.4667 | 34.8
27.9 | 26.94 | 30.28 | 33.15 | 34.325 | 34.6
28 | 26.7 | 30.14 | 32.9 | 34.2 | 34.4714
28.1 | 26.6125 | 30 | 32.75 | 34.0857 | 34.3
28.2 | 26.5 | 29.8 | 32.62 | 33.9 | 34.1714
28.3 | 26.3571 | 29.7 | 32.52 | 33.8 | 34.0286
28.4 | 26.2 | 29.56 | 32.375 | 33.6 | 33.9
28.5 | 26.1 | 29.4 | 32.2 | 33.5 | 33.7286
28.6 | 25.9667 | 29.3 | 32.0667 | 33.3429 | 33.62
28.7 | 25.8 | 29.1 | 31.9 | 33.2 | 33.45
28.8 | 25.675 | 28.975 | 31.8 | 33.025 | 33.3
28.9 | 25.56 | 28.825 | 31.7 | 32.8 | 33.2
29 | 25.5 | 28.7 | 31.5 | 32.7 | 33.025
29.1 | 25.3 | 28.525 | 31.4 | 32.6 | 32.95
29.2 | 25.2 | 28.4333 | 31.25 | 32.5 | 32.775
29.3 | 25.0333 | 28.275 | 31.1 | 32.375 | 32.6
29.4 | 24.975 | 28.2 | 31 | 32.14 | 32.5
29.5 | 24.8 | 28.025 | 30.75 | 32.05 | 32.3556
29.6 | 24.7 | 27.9 | 30.6 | 31.9 | 32.2
29.7 | 24.56 | 27.8 | 30.5 | 31.78 | 32.0714
29.8 | 24.45 | 27.65 | 30.4 | 31.7 | 31.9286
29.9 | 24.3 | 27.5 | 30.2 | 31.54 | 31.725
30 | 24.25 | 27.4 | 30.1 | 31.3714 | 31.68

## SSD80 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 33.8 | 42.5 | 57.2 | 69.6 | 73.5
0.1 | 37.4 | 46.2 | 60.82 | 72.725 | 76.65
0.2 | 43.6 | 52.3 | 66.12 | 77.48 | 80.74
0.3 | 53.22 | 60.6 | 73.1667 | 83.3333 | 85.98
0.4 | 62.38 | 68.775 | 79.9 | 87.7 | 90.04
0.5 | 69.86 | 75.2333 | 84.9667 | 91.4 | 93.16
0.6 | 75.8667 | 80.85 | 88.84 | 94.1 | 95.48
0.7 | 80.84 | 85.06 | 91.86 | 96 | 97.02
0.8 | 85 | 88.5333 | 94.2 | 97.42 | 98.1
0.9 | 88.3 | 91.2667 | 95.975 | 98.5 | 98.96
1 | 91 | 93.64 | 97.3 | 99.1 | 99.54
1.1 | 93.2 | 95.4333 | 98.28 | 99.7 | 99.92
1.2 | 95.2 | 96.78 | 98.98 | 99.9 | 100
1.3 | 96.5 | 97.9 | 99.62 | 100 | 100
1.4 | 97.75 | 98.72 | 99.82 | 100 | 100
1.5 | 98.72 | 99.26 | 100 | 100 | 99.9
1.6 | 99.28 | 99.74 | 100 | 99.825 | 99.8
1.7 | 99.74 | 99.9333 | 100 | 99.7 | 99.5
1.8 | 99.94 | 100 | 99.9 | 99.4 | 99.3
1.9 | 100 | 100 | 99.7 | 99.22 | 98.9
2 | 100 | 100 | 99.5 | 98.8333 | 98.6
2.1 | 99.98 | 99.8667 | 99.2 | 98.4333 | 98.24
2.2 | 99.88 | 99.6667 | 98.9 | 98.14 | 97.92
2.3 | 99.66 | 99.3667 | 98.5 | 97.74 | 97.5333
2.4 | 99.36 | 99 | 98.1 | 97.34 | 97.14
2.5 | 98.96 | 98.65 | 97.7 | 96.94 | 96.66
2.6 | 98.54 | 98.2667 | 97.3333 | 96.54 | 96.24
2.7 | 98.08 | 97.875 | 96.84 | 96.06 | 95.925
2.8 | 97.68 | 97.52 | 96.44 | 95.74 | 95.46
2.9 | 97.24 | 97 | 95.96 | 95.24 | 95.06
3 | 96.8 | 96.6 | 95.54 | 94.88 | 94.66
3.1 | 96.3 | 96.22 | 95.14 | 94.4333 | 94.1714
3.2 | 95.8 | 95.8 | 94.74 | 94.06 | 93.8
3.3 | 95.3 | 95.3 | 94.34 | 93.575 | 93.44
3.4 | 94.85 | 94.82 | 93.85 | 93.15 | 93.02
3.5 | 94.32 | 94.3333 | 93.35 | 92.8 | 92.58
3.6 | 93.75 | 94 | 93.0333 | 92.375 | 92.18
3.7 | 93.3 | 93.45 | 92.58 | 92 | 91.75
3.8 | 92.7 | 93 | 92.18 | 91.6 | 91.35
3.9 | 92.2 | 92.55 | 91.78 | 91.1667 | 90.925
4 | 91.7 | 92.1333 | 91.38 | 90.7 | 90.6
4.1 | 91.25 | 91.6 | 90.98 | 90.3 | 90.2
4.2 | 90.74 | 91.2333 | 90.48 | 89.9 | 89.7
4.3 | 90.24 | 90.74 | 90.08 | 89.5 | 89.4
4.4 | 89.84 | 90.34 | 89.6 | 89.1 | 89
4.5 | 89.34 | 89.84 | 89.3 | 88.7 | 88.5
4.6 | 88.8 | 89.375 | 88.78 | 88.3 | 88.2
4.7 | 88.28 | 89.04 | 88.4 | 87.92 | 87.7
4.8 | 87.82 | 88.5333 | 87.92 | 87.5333 | 87.3
4.9 | 87.28 | 88.1 | 87.55 | 87.12 | 86.95
5 | 86.85 | 87.7 | 87.2333 | 86.62 | 86.62
5.1 | 86.2857 | 87.24 | 86.75 | 86.3 | 86.2
5.2 | 85.8 | 86.8333 | 86.38 | 85.92 | 85.86
5.3 | 85.3 | 86.3333 | 85.9333 | 85.54 | 85.44
5.4 | 84.8 | 85.8667 | 85.58 | 85.14 | 85.06
5.5 | 84.4 | 85.5 | 85.18 | 84.7333 | 84.66
5.6 | 83.9 | 85.1 | 84.72 | 84.38 | 84.3
5.7 | 83.3 | 84.6 | 84.36 | 83.98 | 83.88
5.8 | 82.86 | 84.2 | 83.9 | 83.58 | 83.5143
5.9 | 82.44 | 83.7667 | 83.5 | 83.18 | 83.1
6 | 81.94 | 83.3667 | 83.1 | 82.8 | 82.7
6.1 | 81.54 | 82.9667 | 82.7 | 82.4 | 82.4
6.2 | 81.12 | 82.54 | 82.3 | 81.9667 | 82
6.3 | 80.68 | 82.08 | 81.9333 | 81.66 | 81.6
6.4 | 80.18 | 81.68 | 81.5 | 81.26 | 81.2
6.5 | 79.62 | 81.18 | 81.2333 | 80.86 | 80.8286
6.6 | 79.1714 | 80.9 | 80.84 | 80.48 | 80.42
6.7 | 78.8 | 80.45 | 80.44 | 80.16 | 80.09
6.8 | 78.3 | 80.1 | 79.94 | 79.8 | 79.67
6.9 | 77.8 | 79.58 | 79.54 | 79.4 | 79.36
7 | 77.4 | 79.22 | 79.14 | 78.97 | 78.96
7.1 | 76.9 | 78.775 | 78.82 | 78.65 | 78.57
7.2 | 76.5 | 78.34 | 78.44 | 78.2 | 78.25
7.3 | 76.04 | 77.84 | 78.12 | 77.92 | 77.7727
7.4 | 75.72 | 77.54 | 77.78 | 77.52 | 77.5
7.5 | 75.16 | 77.14 | 77.35 | 77.22 | 77.1636
7.6 | 74.74 | 76.74 | 76.8667 | 76.8 | 76.8
7.7 | 74.32 | 76.275 | 76.4833 | 76.44 | 76.4
7.8 | 73.8667 | 75.92 | 76.2 | 76.04 | 76.1
7.9 | 73.48 | 75.54 | 75.8 | 75.74 | 75.7
8 | 73.02 | 75.18 | 75.5 | 75.34 | 75.3667
8.1 | 72.58 | 74.72 | 75.1 | 75.04 | 75
8.2 | 72.12 | 74.32 | 74.72 | 74.6667 | 74.65
8.3 | 71.78 | 73.94 | 74.36 | 74.24 | 74.3333
8.4 | 71.3 | 73.6 | 74.05 | 73.9667 | 74.03
8.5 | 70.98 | 73.18 | 73.62 | 73.64 | 73.64
8.6 | 70.55 | 72.86 | 73.25 | 73.34 | 73.33
8.7 | 70.22 | 72.52 | 72.9667 | 72.88 | 72.96
8.8 | 69.76 | 72.02 | 72.5667 | 72.58 | 72.66
8.9 | 69.5 | 71.72 | 72.2333 | 72.28 | 72.36
9 | 69.05 | 71.32 | 71.8667 | 71.98 | 71.96
9.1 | 68.65 | 70.92 | 71.575 | 71.55 | 71.6286
9.2 | 68.16 | 70.5429 | 71.175 | 71.3 | 71.3
9.3 | 67.76 | 70.2333 | 70.775 | 70.96 | 71
9.4 | 67.36 | 69.84 | 70.5 | 70.62 | 70.6
9.5 | 67 | 69.5 | 70.22 | 70.22 | 70.3333
9.6 | 66.56 | 69.08 | 69.9 | 69.92 | 70.02
9.7 | 66.26 | 68.76 | 69.4 | 69.56 | 69.7
9.8 | 65.8 | 68.36 | 69.1 | 69.32 | 69.3
9.9 | 65.5 | 67.98 | 68.76 | 68.92 | 69
10 | 65.2 | 67.68 | 68.4 | 68.56 | 68.7
10.1 | 64.8 | 67.3667 | 68.1333 | 68.32 | 68.38
10.2 | 64.4 | 66.96 | 67.7667 | 68.02 | 68
10.3 | 64.1 | 66.6 | 67.42 | 67.62 | 67.8
10.4 | 63.6 | 66.1833 | 67.075 | 67.375 | 67.3667
10.5 | 63.3 | 65.9 | 66.82 | 67.02 | 67.1
10.6 | 62.9 | 65.5 | 66.475 | 66.6667 | 66.7667
10.7 | 62.4 | 65.2 | 66.18 | 66.38 | 66.425
10.8 | 62.2 | 64.86 | 65.86 | 66.06 | 66.1333
10.9 | 61.9 | 64.6 | 65.48 | 65.76 | 65.8
11 | 61.4 | 64.2333 | 65.18 | 65.46 | 65.575
11.1 | 61.16 | 63.8 | 64.86 | 65.18 | 65.175
11.2 | 60.82 | 63.4333 | 64.56 | 64.86 | 64.88
11.3 | 60.46 | 63.14 | 64.18 | 64.48 | 64.6333
11.4 | 60.12 | 62.82 | 63.9 | 64.18 | 64.36
11.5 | 59.7333 | 62.54 | 63.6 | 63.88 | 64.0333
11.6 | 59.4 | 62.14 | 63.3 | 63.58 | 63.7333
11.7 | 59.04 | 61.84 | 62.9 | 63.28 | 63.425
11.8 | 58.64 | 61.44 | 62.675 | 62.9833 | 63.08
11.9 | 58.34 | 61.14 | 62.42 | 62.7 | 62.78
12 | 57.94 | 60.82 | 62.125 | 62.3833 | 62.54
12.1 | 57.72 | 60.48 | 61.74 | 62.1 | 62.18
12.2 | 57.34 | 60.15 | 61.42 | 61.8 | 61.94
12.3 | 57.04 | 59.8667 | 61.12 | 61.5 | 61.64
12.4 | 56.64 | 59.56 | 60.825 | 61.2 | 61.34
12.5 | 56.32 | 59.18 | 60.48 | 60.9 | 61.04
12.6 | 55.94 | 58.8667 | 60.2 | 60.65 | 60.8
12.7 | 55.72 | 58.6 | 59.9571 | 60.35 | 60.4
12.8 | 55.34 | 58.26 | 59.54 | 60.0333 | 60.2
12.9 | 55.04 | 57.88 | 59.32 | 59.76 | 59.9
13 | 54.74 | 57.7 | 59.03 | 59.4667 | 59.6
13.1 | 54.44 | 57.3 | 58.7545 | 59.1667 | 59.3
13.2 | 54.18 | 57.075 | 58.46 | 58.88 | 59.1
13.3 | 53.78 | 56.75 | 58.16 | 58.625 | 58.8
13.4 | 53.54 | 56.45 | 57.86 | 58.32 | 58.5
13.5 | 53.3 | 56.15 | 57.55 | 58.14 | 58.3
13.6 | 53 | 55.85 | 57.2909 | 57.78 | 58
13.7 | 52.6 | 55.55 | 57.01 | 57.54 | 57.7167
13.8 | 52.3 | 55.26 | 56.71 | 57.25 | 57.4167
13.9 | 52 | 55 | 56.41 | 56.9833 | 57.14
14 | 51.7 | 54.7 | 56.1455 | 56.6833 | 56.92
14.1 | 51.4 | 54.4 | 55.94 | 56.4 | 56.62
14.2 | 51.22 | 54.1 | 55.64 | 56.2 | 56.32
14.3 | 50.86 | 53.8 | 55.36 | 55.9 | 56.04
14.4 | 50.56 | 53.5 | 55.0375 | 55.7 | 55.75
14.5 | 50.26 | 53.2 | 54.82 | 55.42 | 55.52
14.6 | 50.02 | 53 | 54.46 | 55.16 | 55.2667
14.7 | 49.76 | 52.7 | 54.26 | 54.82 | 55.02
14.8 | 49.42 | 52.4 | 54.08 | 54.56 | 54.78
14.9 | 49.0667 | 52.1 | 53.75 | 54.32 | 54.52
15 | 48.86 | 51.8 | 53.52 | 54.06 | 54.28
15.1 | 48.66 | 51.5 | 53.22 | 53.76 | 54.02
15.2 | 48.36 | 51.3167 | 52.92 | 53.58 | 53.75
15.3 | 48.05 | 51.04 | 52.72 | 53.28 | 53.54
15.4 | 47.775 | 50.7333 | 52.48 | 53.08 | 53.27
15.5 | 47.475 | 50.52 | 52.08 | 52.78 | 52.9818
15.6 | 47.2 | 50.22 | 51.925 | 52.575 | 52.8
15.7 | 47 | 49.94 | 51.58 | 52.3143 | 52.4818
15.8 | 46.7 | 49.72 | 51.34 | 52.0909 | 52.3
15.9 | 46.4 | 49.42 | 51.15 | 51.85 | 52
16 | 46.2 | 49.22 | 50.8 | 51.55 | 51.7
16.1 | 46 | 48.84 | 50.6 | 51.3 | 51.5
16.2 | 45.62 | 48.62 | 50.35 | 51.05 | 51.2714
16.3 | 45.4 | 48.4333 | 50.12 | 50.85 | 51.075
16.4 | 45.16 | 48.14 | 49.96 | 50.65 | 50.8
16.5 | 44.92 | 47.88 | 49.6 | 50.35 | 50.5429
16.6 | 44.66 | 47.64 | 49.36 | 50.1 | 50.26
16.7 | 44.42 | 47.34 | 49.15 | 49.9 | 50.075
16.8 | 44.0667 | 47.14 | 48.96 | 49.65 | 49.8714
16.9 | 43.88 | 46.84 | 48.6833 | 49.4 | 49.6889
17 | 43.66 | 46.64 | 48.5 | 49.15 | 49.38
17.1 | 43.38 | 46.38 | 48.2 | 48.875 | 49.18
17.2 | 43.1 | 46.08 | 48 | 48.68 | 48.86
17.3 | 42.88 | 45.84 | 47.7 | 48.4286 | 48.68
17.4 | 42.7 | 45.58 | 47.4 | 48.2 | 48.38
17.5 | 42.44 | 45.34 | 47.2 | 47.98 | 48.18
17.6 | 42.16 | 45.1 | 47 | 47.7286 | 47.98
17.7 | 41.9 | 44.88 | 46.7 | 47.52 | 47.8
17.8 | 41.6429 | 44.66 | 46.5167 | 47.2714 | 47.48
17.9 | 41.4375 | 44.42 | 46.3 | 47.05 | 47.2833
18 | 41.2 | 44.22 | 46.1 | 46.86 | 47.05
18.1 | 41.0182 | 43.96 | 45.8167 | 46.58 | 46.78
18.2 | 40.73 | 43.76 | 45.575 | 46.38 | 46.6
18.3 | 40.5 | 43.46 | 45.36 | 46.2 | 46.4
18.4 | 40.3 | 43.26 | 45.16 | 46 | 46.2
18.5 | 40.1143 | 43.06 | 44.86 | 45.74 | 46
18.6 | 39.92 | 42.8 | 44.72 | 45.5 | 45.8
18.7 | 39.6571 | 42.52 | 44.5333 | 45.3 | 45.575
18.8 | 39.5143 | 42.32 | 44.28 | 45.1 | 45.375
18.9 | 39.24 | 42.16 | 44.08 | 44.9 | 45.175
19 | 39.04 | 41.9143 | 43.88 | 44.7 | 44.88
19.1 | 38.84 | 41.7714 | 43.66 | 44.46 | 44.7
19.2 | 38.56 | 41.52 | 43.46 | 44.3 | 44.5333
19.3 | 38.3 | 41.32 | 43.18 | 44.1 | 44.3
19.4 | 38.14 | 41.08 | 42.9833 | 43.8333 | 44.1
19.5 | 37.94 | 40.82 | 42.78 | 43.6 | 43.9
19.6 | 37.74 | 40.62 | 42.6 | 43.4 | 43.65
19.7 | 37.55 | 40.4333 | 42.4 | 43.2 | 43.45
19.8 | 37.3 | 40.24 | 42.2 | 43 | 43.2
19.9 | 37.1 | 40.04 | 42 | 42.8 | 43
20 | 36.8 | 39.76 | 41.7 | 42.64 | 42.8667
20.1 | 36.6167 | 39.54 | 41.5333 | 42.4 | 42.65
20.2 | 36.52 | 39.34 | 41.3333 | 42.24 | 42.475
20.3 | 36.3167 | 39.15 | 41.2 | 42 | 42.225
20.4 | 36.02 | 38.9444 | 40.9333 | 41.78 | 42.02
20.5 | 35.92 | 38.74 | 40.6333 | 41.6 | 41.92
20.6 | 35.72 | 38.5 | 40.5 | 41.4 | 41.7
20.7 | 35.5 | 38.4 | 40.3333 | 41.1667 | 41.52
20.8 | 35.32 | 38.1286 | 40.1 | 41 | 41.32
20.9 | 35.12 | 37.96 | 39.95 | 40.8 | 41.12
21 | 34.92 | 37.7 | 39.75 | 40.6667 | 40.84
21.1 | 34.72 | 37.56 | 39.5 | 40.475 | 40.62
21.2 | 34.52 | 37.36 | 39.3 | 40.2 | 40.42
21.3 | 34.4 | 37.15 | 39.14 | 40.075 | 40.32
21.4 | 34.14 | 36.9429 | 38.96 | 39.875 | 40.12
21.5 | 34.02 | 36.8125 | 38.75 | 39.675 | 39.92
21.6 | 33.7667 | 36.5429 | 38.55 | 39.475 | 39.72
21.7 | 33.58 | 36.4125 | 38.3667 | 39.3 | 39.6
21.8 | 33.4 | 36.2 | 38.2 | 39.1 | 39.34
21.9 | 33.18 | 36.02 | 38.02 | 38.9333 | 39.12
22 | 33 | 35.82 | 37.82 | 38.75 | 39.02
22.1 | 32.8333 | 35.62 | 37.62 | 38.54 | 38.78
22.2 | 32.6333 | 35.4 | 37.42 | 38.38 | 38.6
22.3 | 32.48 | 35.2 | 37.25 | 38.14 | 38.48
22.4 | 32.3 | 35.1 | 37.0857 | 38 | 38.2
22.5 | 32.125 | 34.8333 | 36.92 | 37.78 | 38.1
22.6 | 31.88 | 34.7 | 36.7333 | 37.58 | 37.9
22.7 | 31.76 | 34.5 | 36.5333 | 37.4 | 37.7333
22.8 | 31.5 | 34.3 | 36.32 | 37.3 | 37.575
22.9 | 31.4 | 34.18 | 36.1333 | 37.1 | 37.4
23 | 31.3 | 33.98 | 35.94 | 36.9 | 37.2
23.1 | 31.14 | 33.78 | 35.8 | 36.8 | 37
23.2 | 30.95 | 33.58 | 35.66 | 36.6 | 36.84
23.3 | 30.74 | 33.48 | 35.4444 | 36.4 | 36.7
23.4 | 30.6 | 33.28 | 35.24 | 36.3 | 36.5
23.5 | 30.46 | 33.1 | 35.12 | 36.1 | 36.3667
23.6 | 30.26 | 32.9 | 34.94 | 35.875 | 36.2
23.7 | 30.1 | 32.8 | 34.74 | 35.7 | 35.98
23.8 | 29.96 | 32.5 | 34.6273 | 35.6 | 35.8333
23.9 | 29.76 | 32.4 | 34.46 | 35.4 | 35.64
24 | 29.56 | 32.2 | 34.26 | 35.3 | 35.5
24.1 | 29.46 | 32.1 | 34.06 | 35.1 | 35.3
24.2 | 29.26 | 31.9 | 33.93 | 34.9 | 35.2
24.3 | 29.16 | 31.7 | 33.76 | 34.7 | 35
24.4 | 28.96 | 31.6 | 33.63 | 34.6 | 34.8
24.5 | 28.76 | 31.4 | 33.46 | 34.4 | 34.7
24.6 | 28.6 | 31.2 | 33.26 | 34.25 | 34.5167
24.7 | 28.46 | 31.1 | 33.06 | 34.0667 | 34.32
24.8 | 28.375 | 30.92 | 32.9125 | 33.9 | 34.22
24.9 | 28.18 | 30.82 | 32.78 | 33.8 | 34.0429
25 | 28.08 | 30.62 | 32.66 | 33.6 | 33.88
25.1 | 27.9 | 30.4333 | 32.4667 | 33.475 | 33.7333
25.2 | 27.7 | 30.3 | 32.28 | 33.3 | 33.56
25.3 | 27.58 | 30.1333 | 32.14 | 33.1333 | 33.3727
25.4 | 27.4 | 30 | 31.98 | 33 | 33.24
25.5 | 27.28 | 29.8 | 31.84 | 32.85 | 33.08
25.6 | 27.18 | 29.7 | 31.68 | 32.66 | 32.94
25.7 | 27 | 29.58 | 31.54 | 32.55 | 32.78
25.8 | 26.88 | 29.38 | 31.38 | 32.4 | 32.64
25.9 | 26.68 | 29.2 | 31.3 | 32.2 | 32.48
26 | 26.5833 | 29.08 | 31.1 | 32 | 32.34
26.1 | 26.48 | 28.9 | 30.89 | 31.9 | 32.18
26.2 | 26.3 | 28.775 | 30.7875 | 31.7333 | 31.98
26.3 | 26.1 | 28.64 | 30.64 | 31.6 | 31.9
26.4 | 25.9833 | 28.5 | 30.5 | 31.5 | 31.8
26.5 | 25.9 | 28.3 | 30.375 | 31.32 | 31.6
26.6 | 25.7 | 28.1833 | 30.14 | 31.12 | 31.4
26.7 | 25.6 | 28 | 30 | 31.02 | 31.3
26.8 | 25.4 | 27.9 | 29.8 | 30.92 | 31.1
26.9 | 25.3375 | 27.8 | 29.7 | 30.72 | 31
27 | 25.16 | 27.6 | 29.6 | 30.52 | 30.8
27.1 | 25.06 | 27.5 | 29.46 | 30.44 | 30.7
27.2 | 24.9 | 27.3667 | 29.36 | 30.3333 | 30.6
27.3 | 24.7714 | 27.2 | 29.16 | 30.1571 | 30.42
27.4 | 24.6167 | 27.06 | 29.06 | 30.025 | 30.2667
27.5 | 24.52 | 27 | 28.86 | 29.84 | 30.1833
27.6 | 24.4 | 26.8 | 28.76 | 29.7 | 30
27.7 | 24.28 | 26.64 | 28.56 | 29.6 | 29.84
27.8 | 24.1 | 26.44 | 28.4333 | 29.4 | 29.7333
27.9 | 24 | 26.3714 | 28.3 | 29.3 | 29.6
28 | 23.9 | 26.2 | 28.15 | 29.1625 | 29.46
28.1 | 23.7 | 26.1 | 28.05 | 28.975 | 29.26
28.2 | 23.6 | 26 | 27.9 | 28.9 | 29.16
28.3 | 23.5 | 25.8429 | 27.8 | 28.76 | 29.0667
28.4 | 23.4 | 25.7 | 27.6 | 28.66 | 28.9
28.5 | 23.3 | 25.6 | 27.5 | 28.46 | 28.78
28.6 | 23.2 | 25.4167 | 27.4 | 28.3 | 28.6
28.7 | 23 | 25.3 | 27.2 | 28.2 | 28.48
28.8 | 22.9 | 25.2 | 27.1 | 28.1 | 28.38
28.9 | 22.8 | 25 | 27 | 27.96 | 28.2
29 | 22.6 | 24.9 | 26.9 | 27.86 | 28.1
29.1 | 22.5 | 24.8 | 26.78 | 27.7 | 27.98
29.2 | 22.4 | 24.7 | 26.58 | 27.56 | 27.88
29.3 | 22.3 | 24.52 | 26.44 | 27.46 | 27.68
29.4 | 22.2 | 24.42 | 26.34 | 27.3 | 27.58
29.5 | 22.1 | 24.3 | 26.24 | 27.2 | 27.48
29.6 | 21.92 | 24.12 | 26.1 | 27.0333 | 27.3
29.7 | 21.82 | 24.02 | 26 | 26.92 | 27.2
29.8 | 21.72 | 23.92 | 25.8 | 26.8 | 27.0833
29.9 | 21.6 | 23.8 | 25.7 | 26.68 | 26.9
30 | 21.5 | 23.62 | 25.56 | 26.5 | 26.8

## SSD120 Wedge 30 Depth Doses
Percentage Depth Doses
Wedge 30
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 30x30cm2 | 40x30cm2
0 | 33.9 | 39.8 | 47.9 | 56.2 | 58.9
0.1 | 38.65 | 44.2 | 52.6 | 60.15 | 63.475
0.2 | 46.7125 | 51.82 | 59.3625 | 66.3 | 69.5375
0.3 | 55.9 | 60.9 | 67.4667 | 73.34 | 75.66
0.4 | 63.8 | 67.7 | 74.25 | 79.32 | 81.25
0.5 | 70.5 | 74.2 | 79.46 | 84 | 85.7
0.6 | 76.3 | 79.325 | 83.9 | 87.6 | 88.825
0.7 | 80.8 | 83.5 | 87.25 | 90.6 | 91.55
0.8 | 84.775 | 87.175 | 90.3 | 92.92 | 93.8
0.9 | 87.95 | 89.95 | 92.5667 | 94.8 | 95.525
1 | 90.675 | 92.3 | 94.325 | 96.125 | 96.9
1.1 | 92.8 | 94 | 95.725 | 97.3 | 97.8
1.2 | 94.625 | 95.6 | 97.075 | 98.175 | 98.675
1.3 | 96 | 96.7 | 98.075 | 98.925 | 99.15
1.4 | 97.18 | 97.9 | 98.75 | 99.35 | 99.5
1.5 | 98.04 | 98.5 | 99.25 | 99.75 | 99.75
1.6 | 98.8 | 99.225 | 99.575 | 99.9 | 99.9
1.7 | 99.35 | 99.6 | 99.8 | 100 | 100
1.8 | 99.7 | 99.8 | 99.9 | 100 | 100
1.9 | 99.95 | 99.95 | 100 | 100 | 99.9667
2 | 100 | 100 | 100 | 99.975 | 99.8143
2.1 | 100 | 100 | 100 | 99.825 | 99.7
2.2 | 99.9 | 99.9 | 99.85 | 99.7 | 99.6375
2.3 | 99.86 | 99.75 | 99.6 | 99.6 | 99.425
2.4 | 99.725 | 99.6 | 99.42 | 99.3 | 99.0857
2.5 | 99.55 | 99.5 | 99.2 | 98.98 | 98.9
2.6 | 99.15 | 99.2333 | 98.95 | 98.6 | 98.45
2.7 | 98.875 | 98.9 | 98.625 | 98.4333 | 98.325
2.8 | 98.525 | 98.5667 | 98.375 | 98.26 | 98.075
2.9 | 98.225 | 98.325 | 98.025 | 97.8 | 97.725
3 | 97.75 | 97.9333 | 97.75 | 97.56 | 97.45
3.1 | 97.375 | 97.62 | 97.35 | 97.2667 | 97.125
3.2 | 96.95 | 97.3 | 97.08 | 97.05 | 96.675
3.3 | 96.66 | 96.85 | 96.8286 | 96.64 | 96.425
3.4 | 96.3 | 96.475 | 96.5 | 96.375 | 96.15
3.5 | 95.8 | 96.175 | 96.1 | 95.98 | 95.825
3.6 | 95.35 | 95.85 | 95.8 | 95.6 | 95.5
3.7 | 94.9 | 95.46 | 95.4 | 95.34 | 95.225
3.8 | 94.6667 | 95.15 | 95.1 | 95 | 94.9
3.9 | 94.2 | 94.8 | 94.8 | 94.7 | 94.45
4 | 93.7 | 94.36 | 94.5 | 94.4 | 94.2
4.1 | 93.3 | 94.025 | 94.1 | 94.05 | 93.95
4.2 | 92.8 | 93.55 | 93.82 | 93.7333 | 93.6
4.3 | 92.5 | 93.25 | 93.4667 | 93.35 | 93.2
4.4 | 91.9 | 92.925 | 93.225 | 93.05 | 92.9667
4.5 | 91.6 | 92.55 | 92.75 | 92.75 | 92.6125
4.6 | 91.15 | 92.25 | 92.35 | 92.4 | 92.2375
4.7 | 90.725 | 91.95 | 92.1 | 92.1571 | 92.0625
4.8 | 90.425 | 91.55 | 91.75 | 91.85 | 91.6875
4.9 | 89.925 | 91.225 | 91.4667 | 91.6 | 91.3333
5 | 89.55 | 90.85 | 91.075 | 91.2667 | 91.05
5.1 | 89.225 | 90.4 | 90.8333 | 90.92 | 90.7778
5.2 | 88.775 | 90.1 | 90.5333 | 90.6 | 90.525
5.3 | 88.3 | 89.7 | 90.2 | 90.325 | 90.15
5.4 | 87.84 | 89.4 | 89.9 | 90.02 | 89.85
5.5 | 87.45 | 89 | 89.5 | 89.6667 | 89.6
5.6 | 87.1 | 88.7 | 89.25 | 89.3333 | 89.25
5.7 | 86.75 | 88.3 | 88.9 | 89.1667 | 88.9667
5.8 | 86.4 | 88 | 88.55 | 88.8125 | 88.6571
5.9 | 85.9 | 87.6 | 88.25 | 88.4375 | 88.3
6 | 85.5 | 87.3 | 87.875 | 88.12 | 87.9556
6.1 | 85.1 | 87.0667 | 87.6 | 87.7 | 87.725
6.2 | 84.7 | 86.6 | 87.375 | 87.4571 | 87.4125
6.3 | 84.275 | 86.25 | 86.95 | 87.1333 | 87.0375
6.4 | 83.8 | 85.825 | 86.6625 | 86.9 | 86.775
6.5 | 83.475 | 85.575 | 86.325 | 86.55 | 86.4875
6.6 | 83.05 | 85.15 | 86.1 | 86.2 | 86.175
6.7 | 82.625 | 84.875 | 85.68 | 85.975 | 85.8375
6.8 | 82.15 | 84.475 | 85.38 | 85.72 | 85.5
6.9 | 81.775 | 84.15 | 85.05 | 85.35 | 85.2333
7 | 81.425 | 83.7667 | 84.8 | 85.025 | 84.95
7.1 | 81.0667 | 83.45 | 84.45 | 84.75 | 84.65
7.2 | 80.8 | 83.0333 | 84.18 | 84.425 | 84.3
7.3 | 80.225 | 82.8 | 83.85 | 84.1333 | 84.125
7.4 | 79.975 | 82.4 | 83.5 | 83.85 | 83.7857
7.5 | 79.58 | 82 | 83.2 | 83.5 | 83.5
7.6 | 79.1 | 81.675 | 82.95 | 83.2 | 83.25
7.7 | 78.8 | 81.4 | 82.5 | 82.9 | 82.825
7.8 | 78.4 | 81.1 | 82.2 | 82.6 | 82.575
7.9 | 78.1 | 80.75 | 81.9 | 82.3 | 82.26
8 | 77.7 | 80.4 | 81.6 | 81.96 | 81.925
8.1 | 77.3 | 80.1 | 81.275 | 81.675 | 81.6
8.2 | 77 | 79.7 | 81.025 | 81.425 | 81.325
8.3 | 76.5 | 79.35 | 80.7125 | 81.15 | 81.1
8.4 | 76.1 | 79.2 | 80.425 | 80.85 | 80.85
8.5 | 75.875 | 78.8333 | 80.075 | 80.6 | 80.5
8.6 | 75.5 | 78.45 | 79.7875 | 80.2333 | 80.15
8.7 | 75.05 | 78.1 | 79.475 | 79.9 | 79.9
8.8 | 74.6 | 77.7667 | 79.1375 | 79.625 | 79.65
8.9 | 74.425 | 77.475 | 78.9 | 79.3 | 79.35
9 | 73.925 | 77.15 | 78.55 | 78.975 | 79
9.1 | 73.55 | 76.775 | 78.3 | 78.7 | 78.7667
9.2 | 73.15 | 76.45 | 77.975 | 78.45 | 78.4
9.3 | 72.85 | 76.075 | 77.6 | 78.1667 | 78.1667
9.4 | 72.525 | 75.675 | 77.35 | 77.875 | 77.925
9.5 | 72.125 | 75.4 | 77.1 | 77.625 | 77.65
9.6 | 71.625 | 75.1 | 76.75 | 77.275 | 77.325
9.7 | 71.45 | 74.8 | 76.4 | 76.9375 | 76.975
9.8 | 71.125 | 74.45 | 76.15 | 76.6875 | 76.725
9.9 | 70.675 | 74.1 | 75.8 | 76.4875 | 76.4
10 | 70.3 | 73.9 | 75.55 | 76.075 | 76.2
10.1 | 70.025 | 73.4667 | 75.3 | 75.85 | 75.9
10.2 | 69.7 | 73.3 | 75.0625 | 75.6 | 75.65
10.3 | 69.325 | 73 | 74.7143 | 75.3333 | 75.4
10.4 | 69 | 72.5 | 74.375 | 75.05 | 75.05
10.5 | 68.725 | 72.3429 | 74.125 | 74.75 | 74.8
10.6 | 68.35 | 72.0333 | 73.875 | 74.5 | 74.5714
10.7 | 68.025 | 71.775 | 73.625 | 74.22 | 74.24
10.8 | 67.625 | 71.375 | 73.2625 | 73.95 | 73.975
10.9 | 67.35 | 70.94 | 73 | 73.575 | 73.7667
11 | 66.94 | 70.75 | 72.7571 | 73.3 | 73.35
11.1 | 66.68 | 70.48 | 72.4 | 73.15 | 73.1222
11.2 | 66.3333 | 70.1143 | 72.1375 | 72.85 | 72.92
11.3 | 66.04 | 69.875 | 71.925 | 72.55 | 72.5667
11.4 | 65.72 | 69.625 | 71.5625 | 72.275 | 72.375
11.5 | 65.3 | 69.25 | 71.2333 | 72.02 | 72.05
11.6 | 65.0333 | 68.9167 | 71 | 71.725 | 71.775
11.7 | 64.75 | 68.675 | 70.75 | 71.4833 | 71.525
11.8 | 64.4 | 68.35 | 70.5 | 71.1833 | 71.3
11.9 | 64.15 | 67.975 | 70.2 | 70.975 | 70.975
12 | 63.8 | 67.725 | 69.95 | 70.7 | 70.68
12.1 | 63.425 | 67.4125 | 69.75 | 70.44 | 70.4286
12.2 | 63.1429 | 67.125 | 69.3 | 70.12 | 70.2
12.3 | 62.8 | 66.9 | 69.025 | 69.8 | 69.9333
12.4 | 62.54 | 66.5375 | 68.75 | 69.6 | 69.65
12.5 | 62.175 | 66.3 | 68.55 | 69.3 | 69.4
12.6 | 61.85 | 66.025 | 68.3 | 69.05 | 69.15
12.7 | 61.65 | 65.65 | 68.025 | 68.8 | 68.9
12.8 | 61.225 | 65.35 | 67.7 | 68.575 | 68.65
12.9 | 61 | 65.1 | 67.4667 | 68.3 | 68.3
13 | 60.7333 | 64.85 | 67.225 | 68.0333 | 68.15
13.1 | 60.4167 | 64.5 | 66.9556 | 67.8 | 67.8
13.2 | 60.125 | 64.26 | 66.725 | 67.5667 | 67.56
13.3 | 59.75 | 64.0333 | 66.475 | 67.3 | 67.3429
13.4 | 59.4333 | 63.72 | 66.225 | 67.0714 | 67.0833
13.5 | 59.2 | 63.46 | 65.7875 | 66.8 | 66.8125
13.6 | 58.9 | 63.1286 | 65.5714 | 66.475 | 66.525
13.7 | 58.7 | 62.975 | 65.35 | 66.22 | 66.325
13.8 | 58.35 | 62.55 | 65.1 | 65.975 | 66.1
13.9 | 58.1 | 62.275 | 64.75 | 65.65 | 65.7875
14 | 57.7 | 62.05 | 64.5 | 65.45 | 65.58
14.1 | 57.45 | 61.7 | 64.35 | 65.125 | 65.2833
14.2 | 57.2 | 61.3833 | 63.9667 | 64.95 | 65.0333
14.3 | 56.75 | 61.2167 | 63.6667 | 64.625 | 64.775
14.4 | 56.525 | 60.8333 | 63.4 | 64.375 | 64.525
14.5 | 56.275 | 60.65 | 63.3 | 64.125 | 64.2333
14.6 | 55.95 | 60.4429 | 62.96 | 63.975 | 64
14.7 | 55.775 | 60.02 | 62.725 | 63.725 | 63.8
14.8 | 55.525 | 59.775 | 62.375 | 63.3667 | 63.5286
14.9 | 55.2 | 59.5625 | 62.2 | 63.1 | 63.3
15 | 54.9333 | 59.375 | 61.975 | 62.96 | 63
15.1 | 54.7 | 59.0375 | 61.725 | 62.75 | 62.8
15.2 | 54.4 | 58.78 | 61.48 | 62.52 | 62.55
15.3 | 54.05 | 58.5286 | 61.225 | 62.2 | 62.35
15.4 | 53.775 | 58.2778 | 61 | 61.975 | 62.15
15.5 | 53.6 | 58.025 | 60.8 | 61.7 | 61.9
15.6 | 53.275 | 57.7 | 60.52 | 61.575 | 61.6556
15.7 | 53.125 | 57.525 | 60.2 | 61.325 | 61.4143
15.8 | 52.775 | 57.2 | 59.9778 | 61.075 | 61.2333
15.9 | 52.525 | 56.975 | 59.825 | 60.75 | 61.025
16 | 52.375 | 56.8 | 59.5125 | 60.5333 | 60.5875
16.1 | 52.025 | 56.475 | 59.3125 | 60.3 | 60.425
16.2 | 51.775 | 56.2 | 59.075 | 60.15 | 60.275
16.3 | 51.5 | 55.95 | 58.78 | 59.84 | 60
16.4 | 51.2 | 55.7111 | 58.54 | 59.6 | 59.8
16.5 | 50.9333 | 55.4571 | 58.325 | 59.44 | 59.525
16.6 | 50.7 | 55.15 | 58.05 | 59.1 | 59.275
16.7 | 50.45 | 54.9 | 57.8667 | 58.9 | 59.025
16.8 | 50.2 | 54.65 | 57.575 | 58.7 | 58.8
16.9 | 50.05 | 54.5 | 57.4125 | 58.45 | 58.6333
17 | 49.7 | 54.3 | 57.175 | 58.25 | 58.4
17.1 | 49.55 | 54.05 | 56.9333 | 58 | 58.25
17.2 | 49.2 | 53.78 | 56.65 | 57.7857 | 58
17.3 | 48.95 | 53.55 | 56.425 | 57.6 | 57.675
17.4 | 48.7 | 53.3 | 56.2 | 57.3 | 57.4
17.5 | 48.55 | 53.1 | 55.975 | 57.12 | 57.275
17.6 | 48.3 | 52.8 | 55.7667 | 56.9 | 57.1
17.7 | 48.06 | 52.55 | 55.575 | 56.64 | 56.8286
17.8 | 47.82 | 52.3 | 55.35 | 56.475 | 56.6
17.9 | 47.6 | 52.08 | 55.05 | 56.2 | 56.35
18 | 47.32 | 51.85 | 54.8 | 56 | 56.1
18.1 | 47.06 | 51.6 | 54.575 | 55.8 | 55.9
18.2 | 46.925 | 51.4 | 54.3857 | 55.55 | 55.75
18.3 | 46.6 | 51.1 | 54.175 | 55.4 | 55.5
18.4 | 46.425 | 50.95 | 53.95 | 55.1 | 55.25
18.5 | 46.2 | 50.7 | 53.7 | 54.9 | 55.1
18.6 | 45.975 | 50.575 | 53.525 | 54.66 | 54.9
18.7 | 45.725 | 50.325 | 53.275 | 54.5 | 54.6429
18.8 | 45.5333 | 50.025 | 53.0571 | 54.275 | 54.4167
18.9 | 45.275 | 49.775 | 52.85 | 54.05 | 54.3
19 | 45.025 | 49.5 | 52.6 | 53.8 | 53.95
19.1 | 44.775 | 49.3 | 52.45 | 53.625 | 53.8
19.2 | 44.6 | 49.175 | 52.2 | 53.4 | 53.6
19.3 | 44.35 | 48.9 | 52 | 53.25 | 53.375
19.4 | 44.1 | 48.7 | 51.8 | 53 | 53.125
19.5 | 43.95 | 48.425 | 51.55 | 52.75 | 52.9375
19.6 | 43.7 | 48.2286 | 51.3 | 52.6333 | 52.725
19.7 | 43.5 | 48 | 51.1 | 52.3667 | 52.5
19.8 | 43.3 | 47.8 | 50.925 | 52.1333 | 52.325
19.9 | 43.05 | 47.64 | 50.675 | 51.9 | 52.075
20 | 42.9 | 47.425 | 50.5222 | 51.7 | 51.8667
20.1 | 42.65 | 47.175 | 50.3 | 51.55 | 51.6857
20.2 | 42.4 | 46.925 | 50.05 | 51.3333 | 51.5222
20.3 | 42.2 | 46.775 | 49.8 | 51.1 | 51.2889
20.4 | 42.1 | 46.55 | 49.625 | 50.925 | 51.15
20.5 | 41.85 | 46.3 | 49.5 | 50.775 | 50.95
20.6 | 41.6 | 46.1143 | 49.275 | 50.525 | 50.75
20.7 | 41.46 | 45.9 | 49.1 | 50.375 | 50.5667
20.8 | 41.3 | 45.65 | 48.85 | 50.125 | 50.35
20.9 | 41.06 | 45.5 | 48.6 | 49.975 | 50.2
21 | 40.825 | 45.25 | 48.475 | 49.725 | 49.9
21.1 | 40.675 | 45.1 | 48.3 | 49.475 | 49.75
21.2 | 40.42 | 44.95 | 48.05 | 49.325 | 49.5333
21.3 | 40.275 | 44.7 | 47.9 | 49.1 | 49.3
21.4 | 40 | 44.45 | 47.725 | 48.925 | 49.1
21.5 | 39.8 | 44.2 | 47.45 | 48.8 | 48.95
21.6 | 39.7 | 44.06 | 47.25 | 48.5 | 48.725
21.7 | 39.5 | 43.825 | 47.0222 | 48.4 | 48.55
21.8 | 39.325 | 43.76 | 46.8556 | 48.15 | 48.325
21.9 | 39.075 | 43.525 | 46.6625 | 48 | 48.1667
22 | 38.925 | 43.275 | 46.5286 | 47.75 | 48
22.1 | 38.675 | 43.0625 | 46.3222 | 47.6 | 47.85
22.2 | 38.525 | 42.9375 | 46.175 | 47.35 | 47.6
22.3 | 38.3 | 42.725 | 45.925 | 47.2 | 47.45
22.4 | 38.025 | 42.5778 | 45.7375 | 47.05 | 47.2
22.5 | 37.875 | 42.35 | 45.6125 | 46.8 | 47.05
22.6 | 37.7333 | 42.1429 | 45.3875 | 46.6667 | 46.9
22.7 | 37.6 | 41.925 | 45.225 | 46.52 | 46.675
22.8 | 37.4 | 41.8 | 44.975 | 46.25 | 46.425
22.9 | 37.2 | 41.54 | 44.8 | 46.1 | 46.38
23 | 37 | 41.3333 | 44.625 | 45.96 | 46.1714
23.1 | 36.85 | 41.1 | 44.375 | 45.74 | 46
23.2 | 36.65 | 41 | 44.2 | 45.5 | 45.76
23.3 | 36.45 | 40.8 | 44 | 45.36 | 45.55
23.4 | 36.3 | 40.7 | 43.8333 | 45.2 | 45.4
23.5 | 36.1 | 40.425 | 43.7 | 45.05 | 45.25
23.6 | 35.92 | 40.3 | 43.5 | 44.8 | 45.0571
23.7 | 35.76 | 40.125 | 43.28 | 44.65 | 44.85
23.8 | 35.625 | 39.875 | 43.075 | 44.5 | 44.725
23.9 | 35.375 | 39.725 | 42.925 | 44.2714 | 44.4857
24 | 35.3 | 39.575 | 42.775 | 44.1167 | 44.325
24.1 | 35.075 | 39.35 | 42.6 | 43.95 | 44.175
24.2 | 34.9 | 39.2 | 42.475 | 43.78 | 44
24.3 | 34.725 | 39 | 42.24 | 43.55 | 43.825
24.4 | 34.575 | 38.85 | 42.0333 | 43.4571 | 43.6
24.5 | 34.4 | 38.7429 | 41.9333 | 43.2667 | 43.4
24.6 | 34.225 | 38.5 | 41.7 | 43.12 | 43.3
24.7 | 34.075 | 38.375 | 41.5714 | 42.9 | 43.08
24.8 | 33.925 | 38.225 | 41.3333 | 42.725 | 42.9
24.9 | 33.78 | 37.94 | 41.1 | 42.575 | 42.7429
25 | 33.6 | 37.775 | 41 | 42.4 | 42.6
25.1 | 33.4 | 37.64 | 40.8 | 42.2 | 42.4889
25.2 | 33.25 | 37.5 | 40.7 | 42.025 | 42.325
25.3 | 33.1 | 37.25 | 40.5 | 41.875 | 42.1
25.4 | 32.9 | 37.1333 | 40.3 | 41.7 | 41.875
25.5 | 32.8 | 37 | 40.1 | 41.5 | 41.75
25.6 | 32.6 | 36.725 | 40 | 41.325 | 41.625
25.7 | 32.5 | 36.6 | 39.8 | 41.2333 | 41.4
25.8 | 32.3 | 36.475 | 39.6 | 41.04 | 41.275
25.9 | 32.1 | 36.3 | 39.5 | 40.8333 | 41.15
26 | 32 | 36.125 | 39.35 | 40.7 | 40.95
26.1 | 31.9 | 36 | 39.2 | 40.525 | 40.8
26.2 | 31.65 | 35.85 | 39 | 40.4 | 40.6
26.3 | 31.52 | 35.6111 | 38.82 | 40.2333 | 40.45
26.4 | 31.35 | 35.5 | 38.725 | 40.0667 | 40.325
26.5 | 31.225 | 35.3 | 38.5 | 39.9 | 40.125
26.6 | 31.075 | 35.1556 | 38.4 | 39.7 | 39.9375
26.7 | 30.925 | 34.9625 | 38.2 | 39.6 | 39.8125
26.8 | 30.775 | 34.8286 | 38.1 | 39.4 | 39.6833
26.9 | 30.625 | 34.7111 | 37.9 | 39.275 | 39.475
27 | 30.5 | 34.575 | 37.75 | 39.1 | 39.2833
27.1 | 30.4 | 34.3625 | 37.5 | 38.9 | 39.1833
27.2 | 30.2 | 34.2375 | 37.4 | 38.775 | 38.975
27.3 | 30.075 | 34.1125 | 37.3 | 38.625 | 38.825
27.4 | 29.9 | 33.9875 | 37.1 | 38.48 | 38.7
27.5 | 29.8 | 33.825 | 36.95 | 38.3333 | 38.58
27.6 | 29.7 | 33.6375 | 36.8 | 38.175 | 38.425
27.7 | 29.5 | 33.5 | 36.7 | 38 | 38.225
27.8 | 29.3 | 33.3 | 36.5 | 37.9 | 38.1
27.9 | 29.2 | 33.2 | 36.35 | 37.7667 | 37.94
28 | 29 | 33.025 | 36.1857 | 37.5 | 37.8
28.1 | 28.9 | 32.9167 | 36.05 | 37.4 | 37.64
28.2 | 28.8 | 32.775 | 35.9 | 37.2333 | 37.5
28.3 | 28.7 | 32.5625 | 35.75 | 37.1 | 37.35
28.4 | 28.5 | 32.475 | 35.6 | 37 | 37.2
28.5 | 28.4 | 32.225 | 35.475 | 36.8 | 37.025
28.6 | 28.25 | 32.1 | 35.3 | 36.6667 | 36.9333
28.7 | 28.1 | 32 | 35.175 | 36.475 | 36.775
28.8 | 28 | 31.825 | 35 | 36.3333 | 36.6
28.9 | 27.8 | 31.775 | 34.875 | 36.22 | 36.5
29 | 27.7 | 31.6 | 34.725 | 36 | 36.3
29.1 | 27.6 | 31.475 | 34.575 | 35.9 | 36.2
29.2 | 27.5 | 31.3 | 34.4 | 35.8 | 36
29.3 | 27.3 | 31.175 | 34.275 | 35.6 | 35.85
29.4 | 27.12 | 31 | 34.15 | 35.5 | 35.725
29.5 | 27.1 | 30.9 | 34.025 | 35.35 | 35.575
29.6 | 26.9 | 30.75 | 33.825 | 35.2 | 35.475
29.7 | 26.8 | 30.6 | 33.7 | 35.0333 | 35.325
29.8 | 26.675 | 30.45 | 33.525 | 34.9 | 35.1833
29.9 | 26.6 | 30.3 | 33.4167 | 34.7333 | 35.0167
30 | 26.4 | 30.175 | 33.3 | 34.6667 | 34.8833

## Wedge 30 Lateral Profiles
Lateral Profiles
Wedge 30
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%] |  |  |  |  | Dose[%]
Off axis position [cm] | Depth:2 cm FieldSize:4x4 cm2 | Depth:2 cm FieldSize:10x10 cm2 | Depth:2 cm FieldSize:20x20 cm2 | Depth:2 cm FieldSize:30x30 cm2 | Depth:2 cm FieldSize:40x30 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:30x30 cm2 | Depth:5 cm FieldSize:40x30 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:30x30 cm2 | Depth:10 cm FieldSize:40x30 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:30x30 cm2 | Depth:20 cm FieldSize:40x30 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:30x30 cm2 | Depth:30 cm FieldSize:40x30 cm2
-22.6 | [24 empty] | 23.6333
-22.5 | [18 empty] | 15.55 | 17.12 |  |  |  | 22.4249 | 23.9667
-22.4 | [18 empty] | 15.85 | 17.4143 |  |  |  | 22.6762 | 24.3286
-22.3 | [18 empty] | 16.0556 | 17.5875 |  |  |  | 22.9962 | 24.7
-22.2 | [18 empty] | 16.2545 | 17.88 |  |  |  | 23.3628 | 25.0125
-22.1 | [18 empty] | 16.5286 | 18.1 |  |  |  | 23.696 | 25.25
-22 | [18 empty] | 16.78 | 18.36 |  |  |  | 24.046 | 25.6429
-21.9 | [18 empty] | 17.0667 | 18.5 |  |  |  | 24.3759 | 25.9182
-21.8 | [18 empty] | 17.3286 | 18.8 |  |  |  | 24.7959 | 26.3
-21.7 | [18 empty] | 17.6 | 19.14 |  |  |  | 25.0558 | 26.59
-21.6 | [18 empty] | 17.8 | 19.34 |  |  |  | 25.3815 | 26.9286
-21.5 | [18 empty] | 18 | 19.62 |  |  |  | 25.7386 | 27.34
-21.4 | [13 empty] | 11.5833 | 12.95 |  |  |  | 18.36 | 19.88 |  |  |  | 26.0528 | 27.65
-21.3 | [13 empty] | 11.8 | 13.05 |  |  |  | 18.6333 | 20.16 |  |  |  | 26.4813 | 28.08
-21.2 | [13 empty] | 11.975 | 13.325 |  |  |  | 18.9667 | 20.4 |  |  |  | 26.9622 | 28.4667
-21.1 | [13 empty] | 12.1 | 13.425 |  |  |  | 19.2 | 20.7286 |  |  |  | 27.2669 | 28.8714
-21 | [13 empty] | 12.22 | 13.7 |  |  |  | 19.52 | 21.0143 |  |  |  | 27.7454 | 29.325
-20.9 | [13 empty] | 12.425 | 13.825 |  |  |  | 19.84 | 21.36 |  |  |  | 28.181 | 29.7556
-20.8 | [13 empty] | 12.65 | 13.9667 |  |  |  | 20.1 | 21.7 |  |  |  | 28.6452 | 30.2889
-20.7 | [8 empty] | 9.64 | 10.92 |  |  |  | 12.8 | 14.2 |  |  |  | 20.5 | 21.94 |  |  |  | 29.1951 | 30.75
-20.6 | [8 empty] | 9.8 | 11.1286 |  |  |  | 13.0429 | 14.375 |  |  |  | 20.775 | 22.2 |  |  |  | 29.845 | 31.35
-20.5 | [8 empty] | 9.9 | 11.2714 |  |  |  | 13.225 | 14.6 |  |  |  | 21.1143 | 22.62 |  |  |  | 30.4521 | 32.0333
-20.4 | [8 empty] | 10.1143 | 11.38 |  |  |  | 13.4 | 14.8 |  |  |  | 21.425 | 22.9 |  |  |  | 31.2948 | 32.85
-20.3 | [8 empty] | 10.2 | 11.55 |  |  |  | 13.675 | 14.98 |  |  |  | 21.7333 | 23.26 |  |  |  | 32.2613 | 33.8
-20.2 |  |  |  | 9.2667 | 10.6 |  |  |  | 10.35 | 11.6 |  |  |  | 13.925 | 15.3 |  |  |  | 22.1 | 23.66 |  |  |  | 33.7944 | 35.06
-20.1 |  |  |  | 9.43 | 10.76 |  |  |  | 10.5 | 11.8 |  |  |  | 14.06 | 15.5 |  |  |  | 22.4667 | 23.96 |  |  |  | 35.3941 | 36.7
-20 |  |  |  | 9.575 | 10.8 |  |  |  | 10.7 | 11.96 |  |  |  | 14.32 | 15.725 |  |  |  | 22.84 | 24.38 |  |  |  | 37.8537 | 39.06
-19.9 |  |  |  | 9.7 | 11.02 |  |  |  | 10.8333 | 12.08 |  |  |  | 14.62 | 15.9571 |  |  |  | 23.2333 | 24.7 |  |  |  | 41.0709 | 42.2857
-19.8 |  |  |  | 9.85 | 11.1 |  |  |  | 11 | 12.2667 |  |  |  | 14.8 | 16.1714 |  |  |  | 23.5667 | 25.075 |  |  |  | 45.8923 | 46.69
-19.7 |  |  |  | 9.9 | 11.22 |  |  |  | 11.175 | 12.425 |  |  |  | 15.05 | 16.38 |  |  |  | 24 | 25.4143 |  |  |  | 51.8114 | 52.2125
-19.6 |  |  |  | 10.1 | 11.4 |  |  |  | 11.325 | 12.675 |  |  |  | 15.2 | 16.65 |  |  |  | 24.4333 | 25.9125 |  |  |  | 59.49 | 59.9571
-19.5 |  |  |  | 10.2 | 11.5 |  |  |  | 11.475 | 12.8 |  |  |  | 15.55 | 16.9 |  |  |  | 24.95 | 26.3 |  |  |  | 68.2886 | 68.4143
-19.4 |  |  |  | 10.4 | 11.7 |  |  |  | 11.72 | 12.96 |  |  |  | 15.8 | 17.15 |  |  |  | 25.3429 | 26.7857 |  |  |  | 77.0872 | 76.8857
-19.3 |  |  |  | 10.52 | 11.8 |  |  |  | 11.825 | 13.16 |  |  |  | 16.1 | 17.5 |  |  |  | 25.85 | 27.3 |  |  |  | 85.7107 | 84.7111
-19.2 |  |  |  | 10.7286 | 12 |  |  |  | 12.025 | 13.3333 |  |  |  | 16.3333 | 17.8 |  |  |  | 26.4 | 27.9143 |  |  |  | 93.2844 | 92.1
-19.1 |  |  |  | 10.8714 | 12.1 |  |  |  | 12.1875 | 13.5429 |  |  |  | 16.66 | 18.1 |  |  |  | 27.0286 | 28.5571 |  |  |  | 99.2406 | 98.425
-19 |  |  |  | 10.98 | 12.3 |  |  |  | 12.42 | 13.7 |  |  |  | 17.02 | 18.35 |  |  |  | 27.775 | 29.225 |  |  |  | 103.563 | 102.7
-18.9 |  |  |  | 11.1667 | 12.4 |  |  |  | 12.675 | 13.98 |  |  |  | 17.28 | 18.7 |  |  |  | 28.6333 | 30.1 |  |  |  | 106.9487 | 106
-18.8 |  |  |  | 11.3 | 12.6667 |  |  |  | 12.825 | 14.15 |  |  |  | 17.62 | 18.975 |  |  |  | 29.6667 | 31.1167 |  |  |  | 109.0598 | 108
-18.7 |  |  |  | 11.5 | 12.8 |  |  |  | 13.075 | 14.4 |  |  |  | 17.9333 | 19.34 |  |  |  | 30.9714 | 32.45 |  |  |  | 110.5391 | 109.6333
-18.6 |  |  |  | 11.7 | 12.98 |  |  |  | 13.325 | 14.62 |  |  |  | 18.2 | 19.6 |  |  |  | 32.8 | 34.16 |  |  |  | 111.5588 | 110.64
-18.5 |  |  |  | 11.85 | 13.2 |  |  |  | 13.575 | 14.8667 |  |  |  | 18.5714 | 19.95 |  |  |  | 35.35 | 36.6286 |  |  |  | 112.481 | 111.45
-18.4 |  |  |  | 11.975 | 13.3667 |  |  |  | 13.725 | 15.1 |  |  |  | 19 | 20.36 |  |  |  | 38.9 | 40.04 |  |  |  | 113.2096 | 112.1286
-18.3 |  |  |  | 12.2 | 13.6 |  |  |  | 14 | 15.35 |  |  |  | 19.275 | 20.725 |  |  |  | 44 | 45.2778 |  |  |  | 113.6477 | 112.6
-18.2 |  |  |  | 12.4 | 13.7333 |  |  |  | 14.275 | 15.6 |  |  |  | 19.7 | 21.06 |  |  |  | 51.5 | 52.0429 |  |  |  | 114.231 | 113.1
-18.1 |  |  |  | 12.6 | 13.95 |  |  |  | 14.52 | 15.8 |  |  |  | 20.15 | 21.45 |  |  |  | 60.9429 | 60.9 |  |  |  | 114.481 | 113.6
-18 |  |  |  | 12.8 | 14.1 |  |  |  | 14.7667 | 16.1 |  |  |  | 20.6 | 21.95 |  |  |  | 71.2 | 70.94 |  |  |  | 114.8143 | 113.7667
-17.9 |  |  |  | 13.0222 | 14.35 |  |  |  | 15.12 | 16.3667 |  |  |  | 20.98 | 22.45 |  |  |  | 81.3857 | 82.15 |  |  | 13.0774 | 115.081 | 114
-17.8 |  |  |  | 13.2 | 14.55 |  |  |  | 15.35 | 16.64 |  |  |  | 21.6 | 22.95 |  |  |  | 91.425 | 91.75 |  |  | 13.1974 | 115.2667 | 114.3375
-17.7 |  |  |  | 13.475 | 14.8 |  |  |  | 15.6 | 16.96 |  |  |  | 22.1667 | 23.475 |  |  |  | 100.6714 | 100.75 |  |  | 13.3973 | 115.506 | 114.6143
-17.6 |  |  |  | 13.6286 | 15 |  |  |  | 15.95 | 17.3 |  |  |  | 22.775 | 24.1 |  |  |  | 107.95 | 107.3143 |  |  | 13.6723 | 115.7935 | 114.82
-17.5 |  |  |  | 13.85 | 15.2 |  |  |  | 16.3 | 17.55 |  |  |  | 23.6 | 24.9 |  |  |  | 113.3429 | 112.42 |  |  | 13.8172 | 116.0524 | 115.0125
-17.4 |  |  |  | 14.15 | 15.4333 |  |  |  | 16.5 | 17.9667 |  |  |  | 24.425 | 25.8 |  |  |  | 116.475 | 115.575 |  |  | 13.9972 | 116.2588 | 115.16
-17.3 |  |  |  | 14.38 | 15.75 |  |  |  | 16.94 | 18.2714 |  |  |  | 25.65 | 26.9 |  |  |  | 118.6429 | 117.72 |  |  | 14.0972 | 116.281 | 115.31
-17.2 |  |  |  | 14.6429 | 15.9333 |  |  |  | 17.3333 | 18.7 |  |  |  | 27.1 | 28.4 |  |  |  | 120.05 | 119.06 |  |  | 14.2971 | 116.3719 | 115.4143
-17.1 |  |  |  | 14.85 | 16.2 |  |  |  | 17.8 | 19.08 |  |  |  | 29.1 | 30.3 |  |  |  | 121.2 | 120.12 |  |  | 14.6304 | 116.5032 | 115.5364
-17 |  |  |  | 15.15 | 16.525 |  |  |  | 18.2667 | 19.52 |  |  |  | 32.1 | 33.25 |  |  |  | 121.9375 | 120.92 |  |  | 14.797 | 116.6143 | 115.6429
-16.9 |  |  |  | 15.4 | 16.775 |  |  |  | 18.8 | 20.1 |  |  |  | 36.6 | 37.7667 |  |  | 10.5833 | 122.55 | 121.36 |  |  | 15.022 | 116.731 | 115.7857
-16.8 |  |  |  | 15.7 | 17.125 |  |  |  | 19.4667 | 20.76 |  |  |  | 43.6667 | 44.5 |  |  | 10.8 | 123.05 | 121.9 |  |  | 15.2969 | 116.841 | 115.9
-16.7 |  |  |  | 16.0714 | 17.425 |  |  |  | 20.22 | 21.5 |  |  |  | 53.6571 | 54.1 |  |  | 10.925 | 123.3857 | 122.26 |  |  | 15.4302 | 116.881 | 115.95
-16.6 |  |  |  | 16.4 | 17.75 |  |  |  | 21.1333 | 22.5 |  |  |  | 64.86 | 65.6333 |  |  | 11.1 | 123.6333 | 122.6 |  |  | 15.7369 | 117.0601 | 116
-16.5 |  |  |  | 16.7714 | 18.175 |  |  |  | 22.44 | 23.65 |  |  |  | 78.9 | 79.2 |  |  | 11.3 | 123.8714 | 122.7857 |  |  | 15.9568 | 117.1 | 116.12
-16.4 |  |  |  | 17.3 | 18.675 |  |  |  | 24.025 | 25.2 |  |  |  | 92.4 | 92.3 |  |  | 11.4 | 124.0571 | 122.94 |  |  | 16.1539 | 117.18 | 116.225
-16.3 |  |  |  | 17.8444 | 19.275 |  |  |  | 26.5 | 27.65 |  |  |  | 104.92 | 104.6 |  |  | 11.6 | 124.2714 | 123.175 |  |  | 16.3967 | 117.2371 | 116.3
-16.2 |  |  |  | 18.6 | 20.025 |  |  |  | 30.35 | 31.36 |  |  |  | 115.0333 | 114.7 |  |  | 11.85 | 124.3333 | 123.3 |  |  | 16.7216 | 117.28 | 116.36
-16.1 |  |  |  | 19.6 | 20.9 |  |  |  | 36.5 | 37.36 |  |  |  | 122.8 | 122.5 |  |  | 11.98 | 124.4857 | 123.3833 |  |  | 16.8966 | 117.28 | 116.4
-16 |  |  |  | 21 | 22.2 |  |  |  | 45.92 | 46.48 |  |  |  | 128.12 | 127.025 |  |  | 12.16 | 124.5429 | 123.5 |  |  | 17.0716 | 117.38 | 116.4
-15.9 |  |  |  | 23.025 | 24.3 |  |  |  | 58.4 | 58.62 |  |  | 7.9 | 131.3 | 130.18 |  |  | 12.38 | 124.6 | 123.5 |  |  | 17.4165 | 117.4467 | 116.4571
-15.8 |  |  |  | 26.1 | 27.2 |  |  |  | 73.45 | 73.18 |  |  | 8.02 | 133.1571 | 132.05 |  |  | 12.56 | 124.7 | 123.6 |  |  | 17.6565 | 117.43 | 116.5
-15.7 |  |  |  | 31.4667 | 32.3 |  |  |  | 89.475 | 89.02 |  |  | 8.2 | 134.4429 | 133.3 |  |  | 12.775 | 124.8 | 123.6143 |  |  | 17.8964 | 117.38 | 116.5
-15.6 |  |  |  | 39.98 | 40.7 |  |  |  | 103.975 | 103.8429 |  |  | 8.26 | 135.32 | 134.175 |  |  | 13 | 124.8 | 123.7 |  |  | 18.1564 | 117.38 | 116.5
-15.5 |  |  |  | 52.2857 | 53.4 |  |  |  | 116.7 | 116.325 |  |  | 8.4 | 135.82 | 134.66 |  |  | 13.1333 | 124.8 | 123.7 |  |  | 18.4963 | 117.38 | 116.57
-15.4 |  |  |  | 68.2 | 67.98 |  |  | 6.6228 | 128.175 | 126.7 |  |  | 8.6 | 136.1857 | 134.9667 |  |  | 13.4 | 124.8 | 123.7 |  |  | 18.7962 | 117.43 | 116.6
-15.3 |  |  |  | 85.0143 | 84.2 |  |  | 6.6978 | 135.2 | 133.7 |  |  | 8.68 | 136.4714 | 135.2 |  |  | 13.56 | 124.8 | 123.7 |  |  | 19.0762 | 117.48 | 116.7
-15.2 |  |  |  | 101.7 | 101 |  |  | 6.7977 | 139.0333 | 137.6 |  |  | 8.8 | 136.58 | 135.4167 |  |  | 13.875 | 124.8 | 123.7 |  |  | 19.3836 | 117.49 | 116.7
-15.1 |  |  | 6.4833 | 116.46 | 115.25 |  |  | 6.8977 | 141.56 | 139.9 |  |  | 8.9 | 136.7 | 135.5 |  |  | 14.0667 | 124.8 | 123.7 |  |  | 19.6761 | 117.58 | 116.7
-15 |  |  | 6.5 | 129.6 | 129.1 |  |  | 7.0088 | 142.9667 | 141.3667 |  |  | 9.1 | 136.7 | 135.6 |  |  | 14.275 | 124.7444 | 123.7 |  |  | 19.9246 | 117.58 | 116.7
-14.9 |  |  | 6.6556 | 138.7 | 138 |  |  | 7.0976 | 143.7 | 142.2 |  |  | 9.2 | 136.7 | 135.6 |  |  | 14.52 | 124.7 | 123.7 |  |  | 20.2559 | 117.58 | 116.7
-14.8 |  |  | 6.725 | 144.3 | 142.7 |  |  | 7.1976 | 144.4 | 142.6143 |  |  | 9.4 | 136.7 | 135.5833 |  |  | 14.82 | 124.6 | 123.6 |  |  | 20.6833 | 117.5229 | 116.6375
-14.7 |  |  | 6.86 | 146.96 | 145.3 |  |  | 7.3198 | 144.6 | 143 |  |  | 9.54 | 136.7 | 135.5 |  |  | 15.05 | 124.6 | 123.6 |  |  | 20.9708 | 117.48 | 116.6
-14.6 |  |  | 6.9 | 148.4 | 146.675 |  |  | 7.3975 | 144.7 | 143.1 |  |  | 9.64 | 136.6 | 135.375 |  |  | 15.375 | 124.5 | 123.54 |  |  | 21.3832 | 117.43 | 116.5909
-14.5 |  |  | 7 | 149.025 | 147.4 |  |  | 7.4975 | 144.7 | 143.1 |  |  | 9.84 | 136.55 | 135.2 |  |  | 15.66 | 124.5 | 123.5 |  |  | 21.7957 | 117.38 | 116.5
-14.4 |  |  | 7.0667 | 149.3 | 147.6 |  |  | 7.5975 | 144.7 | 143.1 |  |  | 10 | 136.375 | 135.125 |  |  | 15.9143 | 124.4143 | 123.3857 |  |  | 22.2456 | 117.3657 | 116.5
-14.3 |  |  | 7.16 | 149.3 | 147.6 |  |  | 7.7724 | 144.55 | 143.0429 |  |  | 10.2 | 136.1 | 134.9 |  |  | 16.1333 | 124.35 | 123.2429 |  |  | 22.6755 | 117.28 | 116.45
-14.2 |  |  | 7.3 | 149.3 | 147.6 |  |  | 7.7974 | 144.3 | 142.8 |  |  | 10.4333 | 135.9 | 134.74 |  |  | 16.52 | 124.25 | 123.2 |  |  | 23.2398 | 117.28 | 116.4
-14.1 |  |  | 7.3875 | 149.2667 | 147.525 |  |  | 7.9973 | 144.16 | 142.5222 |  |  | 10.6 | 135.75 | 134.6 |  |  | 16.7571 | 124.1857 | 123.2 |  |  | 23.7153 | 117.28 | 116.4
-14 |  |  | 7.48 | 149.1 | 147.375 |  |  | 8.0973 | 143.9 | 142.3 |  |  | 10.8 | 135.6 | 134.5 |  |  | 17.075 | 124.1 | 123.1 |  |  | 24.4094 | 117.18 | 116.4
-13.9 |  |  | 7.6 | 148.75 | 147 |  |  | 8.1973 | 143.56 | 142.0333 |  |  | 10.95 | 135.5333 | 134.3 |  |  | 17.45 | 124 | 123.1 |  |  | 25.195 | 117.18 | 116.3625
-13.8 |  |  | 7.675 | 148.4 | 146.675 |  |  | 8.3722 | 143.3 | 141.75 |  |  | 11.2 | 135.28 | 134.2 |  |  | 17.84 | 123.9 | 123 |  |  | 26.2376 | 117.18 | 116.3
-13.7 |  |  | 7.8 | 148.0286 | 146.35 |  |  | 8.4972 | 142.8667 | 141.425 |  |  | 11.5 | 134.9714 | 133.9 |  |  | 18.2 | 123.8 | 122.8625 |  |  | 27.466 | 117.1175 | 116.3
-13.6 |  |  | 8 | 147.55 | 145.95 |  |  | 8.6971 | 142.55 | 141.1 |  |  | 11.7 | 134.7571 | 133.65 |  |  | 18.54 | 123.7 | 122.6 |  |  | 29.0942 | 117.08 | 116.3
-13.5 |  |  | 8.075 | 147.0286 | 145.55 |  |  | 8.897 | 142.16 | 140.7429 |  |  | 11.94 | 134.52 | 133.45 |  |  | 19 | 123.6 | 122.5 |  |  | 31.4537 | 117.08 | 116.2429
-13.4 |  |  | 8.2 | 146.6375 | 145.125 |  |  | 8.997 | 141.68 | 140.45 |  |  | 12.15 | 134.24 | 133.225 |  |  | 19.36 | 123.44 | 122.5 |  |  | 34.3931 | 116.981 | 116.2
-13.3 |  |  | 8.35 | 146.2 | 144.675 |  |  | 9.147 | 141.4 | 140.0875 |  |  | 12.4143 | 134.12 | 133 |  |  | 19.85 | 123.2833 | 122.45 |  |  | 38.9922 | 116.9143 | 116.1
-13.2 |  |  | 8.5 | 145.75 | 144.15 |  |  | 9.3469 | 141.2 | 139.68 |  |  | 12.7 | 133.75 | 132.7333 |  |  | 20.32 | 123.2 | 122.275 |  |  | 44.991 | 116.881 | 116.0429
-13.1 |  |  | 8.6 | 145.36 | 143.8 |  |  | 9.5368 | 140.66 | 139.38 |  |  | 12.9375 | 133.525 | 132.5 |  |  | 20.9 | 123.1 | 122.06 |  |  | 51.9496 | 116.8096 | 115.96
-13 |  |  | 8.7333 | 144.8333 | 143.28 |  |  | 9.6968 | 140.3 | 138.9333 |  |  | 13.26 | 133.24 | 132.2 |  |  | 21.64 | 122.96 | 121.975 |  |  | 60.108 | 116.7477 | 115.9
-12.9 |  |  | 8.9 | 144.3286 | 142.9 |  |  | 9.8967 | 139.98 | 138.48 |  |  | 13.55 | 133.04 | 132 |  |  | 22.5 | 122.8 | 121.9 |  |  | 69.3861 | 116.681 | 115.9
-12.8 |  |  | 9.06 | 143.88 | 142.5 |  |  | 10.0966 | 139.65 | 138.225 |  |  | 13.8 | 132.6 | 131.7 |  |  | 23.55 | 122.6571 | 121.8125 |  |  | 77.0646 | 116.581 | 115.75
-12.7 |  |  | 9.2444 | 143.4 | 142 |  |  | 10.2966 | 139.26 | 137.85 |  |  | 14.2 | 132.5 | 131.4333 |  |  | 24.85 | 122.525 | 121.7 |  |  | 84.4031 | 116.581 | 115.7
-12.6 |  |  | 9.4 | 142.875 | 141.55 |  |  | 10.5965 | 138.775 | 137.45 |  |  | 14.5 | 132.3 | 131.14 |  |  | 26.5 | 122.3857 | 121.5167 |  |  | 90.8247 | 116.4953 | 115.7
-12.5 |  |  | 9.6167 | 142.2667 | 141.08 |  |  | 10.7764 | 138.5 | 137 |  |  | 14.9333 | 131.85 | 130.9 |  |  | 28.9333 | 122.2429 | 121.42 |  |  | 95.1095 | 116.381 | 115.6
-12.4 |  |  | 9.7 | 141.9 | 140.55 |  |  | 10.9963 | 137.9 | 136.7286 |  |  | 15.3 | 131.5 | 130.6333 |  |  | 32.3 | 122.1 | 121.24 |  |  | 98.3203 | 116.231 | 115.4667
-12.3 |  |  | 9.9167 | 141.45 | 140.1 |  |  | 11.2676 | 137.56 | 136.36 |  |  | 15.75 | 131.4286 | 130.32 |  |  | 37.24 | 121.9455 | 121.1 |  |  | 100.44 | 116.0643 | 115.4
-12.2 |  |  | 10.1 | 140.9111 | 139.64 |  |  | 11.5961 | 137.26 | 135.96 |  |  | 16.25 | 131.1333 | 130.025 |  |  | 44.35 | 121.7667 | 121 |  |  | 101.88 | 115.981 | 115.4
-12.1 |  |  | 10.35 | 140.5 | 139.34 |  |  | 11.8461 | 136.95 | 135.4 |  |  | 16.775 | 130.8571 | 129.825 |  |  | 53.025 | 121.5909 | 120.9 |  |  | 102.8793 | 115.981 | 115.3
-12 |  |  | 10.6 | 139.9333 | 138.8667 |  |  | 12.1559 | 136.4 | 135.08 |  |  | 17.425 | 130.4667 | 129.48 |  |  | 62.3 | 121.4286 | 120.8 |  |  | 103.704 | 115.9477 | 115.1778
-11.9 |  |  | 10.725 | 139.4 | 138.3 |  |  | 12.5208 | 136.0857 | 134.8 |  |  | 18.3 | 130.1 | 129.35 |  |  | 72.1333 | 121.2818 | 120.7 |  |  | 104.339 | 115.881 | 115
-11.8 |  |  | 11.025 | 138.9 | 137.75 |  |  | 12.9157 | 135.5143 | 134.38 |  |  | 19.3 | 129.875 | 129.05 |  |  | 81.7 | 121.1 | 120.46 |  |  | 104.779 | 115.781 | 114.9167
-11.7 |  |  | 11.2 | 138.5 | 137.26 |  |  | 13.3456 | 135.15 | 133.94 |  |  | 20.65 | 129.6 | 128.7 |  |  | 89.6 | 121 | 120.325 |  |  | 105.279 | 115.681 | 114.8
-11.6 |  |  | 11.525 | 138.05 | 136.925 |  |  | 13.7954 | 134.8 | 133.65 |  |  | 22.425 | 129.275 | 128.45 |  |  | 96.5 | 120.925 | 120.1 |  |  | 105.529 | 115.5667 | 114.7583
-11.5 |  |  | 11.8667 | 137.6 | 136.3 |  |  | 14.5702 | 134.45 | 133.34 |  |  | 24.925 | 129 | 128.2 |  |  | 100.9 | 120.7167 | 120 |  | 7 | 105.879 | 115.481 | 114.6625
-11.4 |  |  | 12.25 | 137.0429 | 135.95 |  |  | 15.2949 | 134.1 | 132.86 |  |  | 28.95 | 128.65 | 127.95 |  |  | 103.94 | 120.5 | 119.9 |  | 7.1 | 106.229 | 115.361 | 114.5286
-11.3 |  |  | 12.6333 | 136.7 | 135.5 |  |  | 16.2946 | 133.84 | 132.525 |  |  | 34.875 | 128.5 | 127.675 |  |  | 105.9 | 120.3571 | 119.7 |  | 7.3 | 106.359 | 115.206 | 114.5
-11.2 |  |  | 13.175 | 136.0857 | 135.15 |  |  | 17.4942 | 133.4222 | 132.0333 |  |  | 43.3 | 128.2 | 127.425 |  |  | 107.38 | 120.1667 | 119.5 |  | 7.4 | 106.579 | 115.0667 | 114.38
-11.1 |  |  | 13.825 | 135.84 | 134.65 |  |  | 19.5685 | 132.9 | 131.48 |  |  | 53.6 | 127.8 | 127.075 |  |  | 108.4333 | 119.9 | 119.3 |  | 7.5 | 106.679 | 114.941 | 114.275
-11 |  |  | 14.7 | 135.18 | 134.225 |  |  | 21.9927 | 132.5286 | 131.22 |  |  | 65.375 | 127.55 | 126.75 |  | 5.4 | 109.14 | 119.75 | 119.15 |  | 7.6 | 106.879 | 114.831 | 114.1429
-10.9 |  |  | 15.8 | 134.8333 | 133.6 |  |  | 26.1913 | 131.9667 | 130.9333 |  |  | 76.6 | 127.25 | 126.5 |  | 5.5 | 109.7 | 119.5429 | 119 |  | 7.9 | 106.979 | 114.731 | 114.1
-10.8 |  |  | 17.55 | 134.44 | 133.2 |  |  | 32.949 | 131.625 | 130.52 |  |  | 88.5 | 126.975 | 126.125 |  | 5.6 | 110.1333 | 119.4857 | 118.76 |  | 8 | 107.179 | 114.6366 | 114
-10.7 |  |  | 20.2 | 133.875 | 132.9 |  |  | 42.0859 | 131.2 | 130.125 |  |  | 97.6 | 126.66 | 125.82 |  | 5.7 | 110.5 | 119.3333 | 118.56 |  | 8.1 | 107.279 | 114.5185 | 113.9667
-10.6 |  |  | 24.4 | 133.4 | 132.45 |  |  | 53.7821 | 130.8 | 129.7 |  |  | 104.76 | 126.25 | 125.6 |  | 5.8 | 110.72 | 119.0571 | 118.48 |  | 8.2 | 107.354 | 114.411 | 113.9
-10.5 |  |  | 31.6 | 133.06 | 132.1 |  |  | 66.7777 | 130.5 | 129.5 |  | 3.7 | 109.5 | 126.02 | 125.32 |  | 6 | 110.88 | 118.9 | 118.2833 |  | 8.3667 | 107.379 | 114.281 | 113.74
-10.4 |  |  | 41.8 | 132.6 | 131.62 |  |  | 80.0733 | 130.125 | 129.0333 |  | 3.8 | 112.6 | 125.6714 | 125.1 |  | 6.1 | 111.1 | 118.6 | 118.1375 |  | 8.5667 | 107.4623 | 114.181 | 113.6833
-10.3 |  |  | 54.45 | 132.025 | 131.175 |  |  | 92.4692 | 129.75 | 128.72 |  | 3.8 | 114.4 | 125.5 | 124.7333 |  | 6.1 | 111.2667 | 118.5429 | 118 |  | 8.7 | 107.578 | 114.081 | 113.5167
-10.2 |  |  | 68.25 | 131.6 | 130.65 |  | 3 | 103.066 | 129.3667 | 128.1667 |  | 3.8 | 115.6 | 125.1667 | 124.5 |  | 6.3 | 111.3 | 118.3714 | 117.8333 |  | 8.8 | 107.678 | 113.8667 | 113.35
-10.1 |  | 2.9 | 82.75 | 131.15 | 130.1333 |  | 3.1 | 110.4636 | 128.975 | 127.9 |  | 3.9 | 116.45 | 124.85 | 124.2333 |  | 6.3333 | 111.4667 | 118.2167 | 117.5286 |  | 9 | 107.678 | 113.7143 | 113.0833
-10 |  | 3 | 96.12 | 130.65 | 129.775 |  | 3.1 | 115.5414 | 128.55 | 127.45 |  | 4 | 117 | 124.54 | 123.9 |  | 6.5 | 111.5 | 117.95 | 117.4 |  | 9.1333 | 107.778 | 113.5667 | 113
-9.9 |  | 3 | 107.3333 | 130.24 | 129.35 |  | 3.1 | 118.3103 | 128.08 | 127.22 |  | 4.1 | 117.5 | 124.18 | 123.625 |  | 6.6 | 111.5 | 117.7833 | 117.26 |  | 9.3 | 107.878 | 113.481 | 112.9
-9.8 |  | 3.1 | 115.3 | 129.85 | 128.98 |  | 3.2 | 120.16 | 127.68 | 126.85 |  | 4.1 | 117.75 | 123.84 | 123.32 |  | 6.75 | 111.6 | 117.7 | 116.95 |  | 9.5 | 107.878 | 113.3524 | 112.7714
-9.7 |  | 3.1 | 120.0429 | 129.5 | 128.4 |  | 3.2 | 121.16 | 127.3 | 126.375 |  | 4.3 | 117.9 | 123.55 | 122.9556 |  | 6.9 | 111.6 | 117.4143 | 116.775 |  | 9.6 | 107.978 | 113.281 | 112.7
-9.6 |  | 3.1 | 122.56 | 128.9 | 128.1 |  | 3.3 | 121.7923 | 126.8667 | 126.025 |  | 4.35 | 118 | 123.325 | 122.725 |  | 7 | 111.6 | 117.1 | 116.6714 |  | 9.8 | 107.978 | 113.1667 | 112.6
-9.5 |  | 3.2 | 124 | 128.55 | 127.7333 |  | 3.3 | 122.284 | 126.6 | 125.75 |  | 4.4 | 118.1 | 123.0429 | 122.275 |  | 7.2 | 111.6 | 116.9333 | 116.5286 |  | 10.0667 | 107.978 | 113.0143 | 112.4625
-9.4 |  | 3.2 | 124.8 | 128.24 | 127.3 |  | 3.4 | 122.559 | 126.15 | 125.4 |  | 4.5 | 118.1 | 122.7 | 122 |  | 7.3 | 111.7 | 116.7714 | 116.42 |  | 10.3 | 107.978 | 112.9477 | 112.3444
-9.3 |  | 3.3 | 125.18 | 127.7 | 126.86 |  | 3.45 | 122.559 | 125.825 | 125 |  | 4.5667 | 118.1 | 122.275 | 121.775 |  | 7.5 | 111.7 | 116.6286 | 116.1889 |  | 10.4667 | 107.978 | 112.806 | 112.2
-9.2 |  | 3.3 | 125.4 | 127.35 | 126.6 |  | 3.5 | 122.559 | 125.5 | 124.55 |  | 4.7 | 118 | 122.025 | 121.4 |  | 7.6 | 111.6 | 116.4 | 115.95 |  | 10.7 | 107.978 | 112.7524 | 112.1571
-9.1 |  | 3.4 | 125.4 | 126.925 | 126.175 |  | 3.55 | 122.559 | 125.1429 | 124.2 |  | 4.8 | 117.9714 | 121.78 | 121.16 |  | 7.8 | 111.6 | 116.2 | 115.8 |  | 10.85 | 107.978 | 112.581 | 112.0143
-9 |  | 3.4 | 125.38 | 126.45 | 125.75 |  | 3.6 | 122.484 | 124.75 | 123.8 |  | 4.9 | 117.8667 | 121.4 | 120.8 |  | 8 | 111.6 | 116.0455 | 115.54 |  | 11.1 | 107.978 | 112.481 | 111.8667
-8.9 |  | 3.5 | 125.275 | 126.025 | 125.3333 |  | 3.7 | 122.459 | 124.28 | 123.6 |  | 4.9333 | 117.75 | 121.2 | 120.56 |  | 8.2 | 111.5333 | 115.8714 | 115.25 |  | 11.35 | 107.978 | 112.3643 | 111.6857
-8.8 |  | 3.5 | 125.1 | 125.575 | 124.88 |  | 3.8 | 122.234 | 124 | 123.125 |  | 5.1 | 117.6 | 120.8 | 120.25 |  | 8.3 | 111.5 | 115.65 | 115 |  | 11.5 | 107.978 | 112.1977 | 111.6
-8.7 |  | 3.6 | 124.85 | 125.35 | 124.5667 |  | 3.8333 | 122.019 | 123.575 | 122.6667 |  | 5.2 | 117.5 | 120.5 | 119.96 |  | 8.5 | 111.475 | 115.45 | 114.9714 |  | 11.7 | 107.978 | 112.031 | 111.4
-8.6 |  | 3.6 | 124.5 | 124.9 | 124.1667 |  | 3.9 | 121.809 | 123.1857 | 122.35 |  | 5.3 | 117.4 | 120.2 | 119.65 |  | 8.7 | 111.4 | 115.1857 | 114.8286 |  | 12.05 | 107.9669 | 111.8381 | 111.2667
-8.5 |  | 3.7 | 124.3 | 124.5 | 123.78 |  | 4 | 121.493 | 122.78 | 121.94 |  | 5.5 | 117.275 | 119.95 | 119.3 |  | 9 | 111.3 | 114.9429 | 114.62 |  | 12.4 | 107.878 | 111.581 | 111.1429
-8.4 |  | 3.8 | 124.05 | 124.05 | 123.3 |  | 4.1 | 121.28 | 122.425 | 121.6 |  | 5.6 | 117.15 | 119.5667 | 118.9333 |  | 9.2 | 111.2667 | 114.7375 | 114.4375 |  | 12.6 | 107.878 | 111.481 | 111.1
-8.3 |  | 3.9 | 123.7 | 123.55 | 122.95 |  | 4.2 | 120.9933 | 122.02 | 121.15 |  | 5.75 | 116.78 | 119.275 | 118.6 |  | 9.3 | 111.1 | 114.6125 | 114.26 |  | 12.9 | 107.778 | 111.4143 | 110.9857
-8.2 |  | 3.9333 | 123.1667 | 123.125 | 122.5429 |  | 4.3 | 120.78 | 121.6375 | 120.9 |  | 5.9 | 116.5714 | 119 | 118.45 |  | 9.6 | 111 | 114.45 | 113.9 |  | 13.2667 | 107.778 | 111.3524 | 110.8429
-8.1 |  | 4 | 122.9 | 122.725 | 122.15 |  | 4.4 | 120.46 | 121.3 | 120.65 |  | 6.1 | 116.4 | 118.575 | 118.2 |  | 9.8 | 111 | 114.1857 | 113.7 |  | 13.6 | 107.7447 | 111.181 | 110.7125
-8 |  | 4.1 | 122.65 | 122.375 | 121.72 |  | 4.5 | 120.2743 | 120.95 | 120.2 |  | 6.2 | 116.2286 | 118.34 | 117.7667 |  | 10.1 | 110.9 | 113.94 | 113.5 |  | 13.9 | 107.678 | 110.932 | 110.5778
-7.9 |  | 4.2 | 122.26 | 121.9667 | 121.36 |  | 4.6 | 119.91 | 120.66 | 119.8 |  | 6.425 | 116 | 118.14 | 117.5 |  | 10.3 | 110.82 | 113.76 | 113.3 |  | 14.2 | 107.578 | 110.682 | 110.3778
-7.8 |  | 4.2 | 121.86 | 121.5571 | 121.05 |  | 4.7 | 119.585 | 120.22 | 119.375 |  | 6.6 | 115.875 | 117.85 | 117.25 |  | 10.7 | 110.72 | 113.5 | 113.1333 |  | 14.7 | 107.578 | 110.5987 | 110.24
-7.7 |  | 4.3 | 121.375 | 121.2 | 120.5429 |  | 4.8 | 119.28 | 119.8 | 119.1 |  | 6.7 | 115.7429 | 117.42 | 116.9 |  | 10.9 | 110.6 | 113.3333 | 112.875 |  | 15.25 | 107.479 | 110.502 | 110.0667
-7.6 |  | 4.4 | 120.975 | 120.8 | 120.04 |  | 4.9 | 118.96 | 119.5 | 118.825 |  | 6.95 | 115.4571 | 117.12 | 116.68 |  | 11.2 | 110.475 | 113.1667 | 112.72 |  | 15.7667 | 107.379 | 110.3487 | 109.9
-7.5 |  | 4.6 | 120.65 | 120.3429 | 119.8 |  | 5.1 | 118.7267 | 119.1 | 118.46 |  | 7.15 | 115.22 | 116.78 | 116.4 |  | 11.6667 | 110.3222 | 112.9 | 112.6 | 2.7 | 16.4 | 107.354 | 110.2695 | 109.7778
-7.4 |  | 4.7 | 120.25 | 119.96 | 119.4 |  | 5.2 | 118.461 | 118.82 | 118.12 |  | 7.4 | 115.05 | 116.475 | 116.1778 | 2 | 11.95 | 110.225 | 112.7333 | 112.42 | 2.8 | 17.2 | 107.279 | 110.107 | 109.6667
-7.3 |  | 4.8 | 119.84 | 119.575 | 119.1 |  | 5.3667 | 118.111 | 118.35 | 117.7 |  | 7.55 | 114.875 | 116.12 | 115.92 | 2.1 | 12.4667 | 110.1 | 112.475 | 112.1 | 2.8 | 18.2 | 107.279 | 109.8987 | 109.5429
-7.2 |  | 4.9 | 119.58 | 119.14 | 118.6429 |  | 5.5 | 117.861 | 118 | 117.425 | 1.3 | 7.9 | 114.7 | 115.8714 | 115.5143 | 2.1 | 12.95 | 109.94 | 112.2667 | 111.8833 | 2.9 | 19.425 | 107.179 | 109.882 | 109.5
-7.1 |  | 5 | 119.3 | 118.7667 | 118.28 |  | 5.75 | 117.561 | 117.8 | 117.05 | 1.3 | 8.2 | 114.4 | 115.5 | 115.25 | 2.2 | 13.4 | 109.8714 | 112.0167 | 111.8 | 3 | 21.1 | 107.179 | 109.7653 | 109.4
-7 | 1 | 5.1 | 118.9 | 118.4 | 117.95 | 1.1 | 5.95 | 117.181 | 117.3667 | 116.675 | 1.3 | 8.425 | 114.2 | 115.3 | 114.94 | 2.2 | 14.0667 | 109.8 | 111.8167 | 111.6 | 3 | 23.275 | 107.179 | 109.5987 | 109.2333
-6.9 | 1 | 5.3 | 118.3 | 118.0667 | 117.6 | 1.1 | 6.12 | 116.961 | 117 | 116.3333 | 1.4 | 8.8 | 113.94 | 115.02 | 114.6857 | 2.3 | 14.8 | 109.64 | 111.66 | 111.375 | 3.1 | 26.4 | 107.079 | 109.4534 | 109.0667
-6.8 | 1 | 5.4 | 118 | 117.8 | 117.0857 | 1.1 | 6.3667 | 116.5943 | 116.6 | 116.0286 | 1.4 | 9.2 | 113.7111 | 114.7 | 114.4 | 2.3 | 15.9 | 109.575 | 111.5 | 111.125 | 3.1 | 31.1 | 107.079 | 109.3106 | 108.9
-6.7 | 1 | 5.64 | 117.725 | 117.3625 | 116.8667 | 1.1 | 6.6 | 116.236 | 116.32 | 115.75 | 1.4 | 9.65 | 113.5 | 114.45 | 114.05 | 2.4 | 17.2 | 109.45 | 111.2857 | 110.96 | 3.2 | 37.1 | 106.979 | 109.182 | 108.8333
-6.6 | 1 | 5.8 | 117.35 | 116.98 | 116.45 | 1.1 | 6.9667 | 115.961 | 115.95 | 115.275 | 1.5 | 10.15 | 113.2143 | 114.1714 | 113.8 | 2.425 | 18.76 | 109.3 | 111.0667 | 110.725 | 3.2 | 44.6 | 106.979 | 109.082 | 108.8
-6.5 | 1 | 6 | 116.95 | 116.55 | 116.15 | 1.1 | 7.3 | 115.5363 | 115.675 | 115.025 | 1.5 | 10.8 | 112.95 | 113.9571 | 113.55 | 2.5 | 21 | 109.2 | 110.9 | 110.575 | 3.2 | 52.725 | 106.879 | 108.9963 | 108.7
-6.4 | 1 | 6.275 | 116.6667 | 116.2 | 115.8 | 1.1 | 7.7 | 115.262 | 115.25 | 114.6667 | 1.5 | 11.5 | 112.7333 | 113.54 | 113.225 | 2.525 | 24.5 | 109.1 | 110.7273 | 110.34 | 3.375 | 62 | 106.879 | 108.952 | 108.6
-6.3 | 1 | 6.6 | 116.175 | 115.9 | 115.35 | 1.1 | 8.1 | 115.002 | 114.875 | 114.3333 | 1.5333 | 12.4 | 112.6 | 113.28 | 112.9333 | 2.6 | 29.25 | 109.025 | 110.6 | 110.2 | 3.5 | 70.5667 | 106.779 | 108.8487 | 108.6
-6.2 | 1.1 | 6.9333 | 115.9 | 115.6 | 115 | 1.1333 | 8.8 | 114.662 | 114.6 | 113.9667 | 1.6 | 13.6 | 112.25 | 113.02 | 112.66 | 2.6 | 35.5 | 108.975 | 110.4286 | 110.1 | 3.6 | 78.3 | 106.779 | 108.702 | 108.45
-6.1 | 1.1 | 7.3 | 115.68 | 115.35 | 114.65 | 1.2 | 9.5 | 114.412 | 114.3 | 113.7 | 1.6 | 15.2 | 112 | 112.78 | 112.4 | 2.7 | 43.8 | 108.825 | 110.31 | 110 | 3.7 | 84.6 | 106.779 | 108.532 | 108.2833
-6 | 1.1 | 7.9333 | 115.2 | 114.775 | 114.2 | 1.2 | 10.5333 | 114.062 | 113.9 | 113.35 | 1.6 | 17.6 | 111.85 | 112.575 | 112.22 | 2.8 | 53.8 | 108.8 | 110.1778 | 109.8 | 3.7667 | 89.3 | 106.679 | 108.3963 | 108.2
-5.9 | 1.1 | 8.7 | 114.94 | 114.425 | 114 | 1.2 | 11.9 | 113.7287 | 113.55 | 113.2 | 1.7 | 20.925 | 111.7 | 112.3 | 111.975 | 2.9 | 62.6333 | 108.7 | 110.0625 | 109.7 | 3.8 | 92.7333 | 106.679 | 108.382 | 108.1
-5.8 | 1.1 | 9.7667 | 114.7 | 114.1 | 113.65 | 1.2 | 13.9 | 113.482 | 113.2 | 112.85 | 1.7 | 26.6 | 111.5 | 112 | 111.8 | 2.9 | 72.25 | 108.6 | 109.9286 | 109.6 | 3.9333 | 95.1 | 106.604 | 108.222 | 108
-5.7 | 1.1 | 11.2 | 114.32 | 113.8 | 113.325 | 1.3 | 16.9 | 113.187 | 113.06 | 112.52 | 1.8 | 33.82 | 111.325 | 111.8 | 111.5 | 3 | 80.4667 | 108.5 | 109.74 | 109.4333 | 4 | 96.55 | 106.579 | 108.0445 | 107.9
-5.6 | 1.2 | 13.4667 | 114 | 113.4 | 113.125 | 1.3 | 21.4 | 112.9334 | 112.775 | 112.25 | 1.8 | 42.9333 | 111.2 | 111.7 | 111.2333 | 3.1 | 87.05 | 108.4 | 109.59 | 109.3 | 4.1 | 97.8 | 106.579 | 107.982 | 107.8
-5.5 | 1.2 | 17.1 | 113.6 | 113.1 | 112.86 | 1.3 | 28.1 | 112.762 | 112.44 | 112 | 1.9 | 53.875 | 111 | 111.4 | 111.12 | 3.2 | 91.9 | 108.3 | 109.48 | 109.1286 | 4.2667 | 98.7 | 106.479 | 107.882 | 107.6667
-5.4 | 1.2 | 23 | 113.275 | 112.875 | 112.5 | 1.4 | 36.75 | 112.443 | 112.3 | 111.84 | 1.9 | 64.7667 | 110.78 | 111.225 | 111.08 | 3.2 | 95.2 | 108.2 | 109.2714 | 108.98 | 4.4 | 99.4 | 106.379 | 107.7153 | 107.5
-5.3 | 1.2 | 30.95 | 113.1 | 112.7 | 112.35 | 1.4 | 48.5 | 112.313 | 112.04 | 111.675 | 2 | 75.22 | 110.6 | 110.9 | 110.8 | 3.3 | 97.3667 | 108 | 109.05 | 108.875 | 4.4 | 100.0333 | 106.379 | 107.6487 | 107.3
-5.2 | 1.3 | 42.3 | 112.9 | 112.45 | 112.1 | 1.4 | 59.45 | 112.0297 | 111.84 | 111.4 | 2.0333 | 84.0333 | 110.4 | 110.725 | 110.6 | 3.4333 | 98.95 | 107.9 | 108.9333 | 108.8 | 4.5 | 100.3 | 106.279 | 107.582 | 107.2
-5.1 | 1.3 | 55.1 | 112.52 | 112 | 111.85 | 1.4 | 71.3 | 111.863 | 111.62 | 111.1 | 2.1 | 91.175 | 110.25 | 110.58 | 110.4 | 3.5 | 100.0667 | 107.8 | 108.7 | 108.7 | 4.7 | 100.7 | 106.179 | 107.432 | 107.075
-5 | 1.3 | 66.8 | 112.275 | 111.875 | 111.55 | 1.5 | 81.825 | 111.563 | 111.4 | 110.875 | 2.2 | 95.8 | 110.1 | 110.4 | 110.2333 | 3.6 | 100.8 | 107.7 | 108.6571 | 108.425 | 4.8 | 101.025 | 106.079 | 107.282 | 107
-4.9 | 1.3 | 79.2 | 112 | 111.65 | 111.34 | 1.6 | 90.4 | 111.363 | 111.12 | 110.7 | 2.3 | 99.06 | 109.9 | 110.2 | 109.96 | 3.7 | 101.25 | 107.6143 | 108.48 | 108.2429 | 4.9333 | 101.2 | 105.979 | 107.182 | 106.8222
-4.8 | 1.4 | 89.1 | 111.9 | 111.4 | 111.1 | 1.6 | 96.625 | 111.163 | 110.95 | 110.5 | 2.325 | 100.75 | 109.7 | 109.95 | 109.8 | 3.8 | 101.8 | 107.5 | 108.3 | 108.1125 | 5.1 | 101.5 | 105.954 | 107.0677 | 106.7
-4.7 | 1.4 | 96.5 | 111.65 | 111.0333 | 110.8333 | 1.6 | 100.3 | 110.8963 | 110.725 | 110.375 | 2.4333 | 102.1 | 109.525 | 109.8 | 109.65 | 4 | 102.1 | 107.4 | 108.0667 | 107.98 | 5.25 | 101.65 | 105.879 | 106.9153 | 106.66
-4.6 | 1.5 | 101.05 | 111.4 | 110.8667 | 110.525 | 1.7 | 102.7 | 110.663 | 110.4667 | 110.05 | 2.5 | 103.2 | 109.3 | 109.52 | 109.325 | 4.1 | 102.45 | 107.3 | 107.8667 | 107.75 | 5.4 | 101.8 | 105.779 | 106.7391 | 106.5429
-4.5 | 1.5 | 104.3 | 111.08 | 110.65 | 110.3 | 1.8 | 104.3 | 110.3797 | 110.225 | 109.725 | 2.6 | 103.7 | 109.14 | 109.3 | 109.2 | 4.2 | 102.6 | 107.2125 | 107.7 | 107.5667 | 5.5 | 101.9 | 105.699 | 106.682 | 106.4125
-4.4 | 1.5 | 105.9 | 110.8429 | 110.3667 | 110.075 | 1.8 | 105.1 | 110.163 | 110.0333 | 109.6 | 2.8 | 104.2333 | 108.98 | 109.075 | 109.025 | 4.3 | 102.8 | 107.1 | 107.65 | 107.5 | 5.6667 | 102.1 | 105.579 | 106.5153 | 106.31
-4.3 | 1.6 | 106.9 | 110.58 | 110.2143 | 109.775 | 1.9 | 105.9 | 109.9201 | 109.9 | 109.375 | 2.8 | 104.5 | 108.8 | 108.825 | 108.8 | 4.5 | 102.9 | 107.0167 | 107.4833 | 107.4 | 5.9 | 102.25 | 105.4361 | 106.3487 | 106.16
-4.2 | 1.7 | 107.3 | 110.325 | 109.975 | 109.6 | 2 | 106.3 | 109.738 | 109.675 | 109.15 | 2.9667 | 104.72 | 108.65 | 108.7 | 108.5333 | 4.7 | 103.1 | 106.9 | 107.3 | 107.2 | 6.0667 | 102.4 | 105.279 | 106.282 | 106.1
-4.1 | 1.7 | 107.6333 | 110.15 | 109.68 | 109.375 | 2.1 | 106.52 | 109.414 | 109.34 | 108.9 | 3.1 | 105 | 108.3 | 108.5167 | 108.3 | 4.85 | 103.3 | 106.78 | 107.1167 | 106.9333 | 6.325 | 102.5 | 105.229 | 106.182 | 105.9
-4 | 1.8 | 108 | 109.82 | 109.42 | 109.2429 | 2.15 | 106.6 | 109.239 | 109.1 | 108.65 | 3.2667 | 105.1667 | 108.08 | 108.36 | 108.1667 | 5.1 | 103.3 | 106.5714 | 106.94 | 106.8 | 6.5667 | 102.6 | 105.079 | 105.9653 | 105.8333
-3.9 | 1.8 | 108.1 | 109.58 | 109.18 | 108.96 | 2.3 | 106.8 | 109.014 | 108.8 | 108.4 | 3.4 | 105.2 | 107.875 | 108.05 | 108.08 | 5.3 | 103.5 | 106.4286 | 106.76 | 106.7 | 6.9 | 102.7 | 105.079 | 105.7987 | 105.7
-3.8 | 1.9 | 108.1 | 109.375 | 108.98 | 108.6778 | 2.4 | 106.8 | 108.824 | 108.5 | 108.3 | 3.6 | 105.2 | 107.7 | 107.875 | 107.8667 | 5.58 | 103.525 | 106.225 | 106.7 | 106.525 | 7.3667 | 102.7 | 104.979 | 105.732 | 105.5667
-3.7 | 2 | 108 | 109.1143 | 108.6667 | 108.54 | 2.5 | 106.8 | 108.484 | 108.3 | 107.975 | 3.8 | 105.3 | 107.5429 | 107.65 | 107.5333 | 5.85 | 103.6 | 106.1 | 106.5 | 106.375 | 7.85 | 102.8 | 104.879 | 105.582 | 105.4143
-3.6 | 2.1 | 107.96 | 108.85 | 108.475 | 108.2429 | 2.7 | 106.8 | 108.204 | 108.22 | 107.78 | 3.9667 | 105.3 | 107.34 | 107.38 | 107.22 | 6.2 | 103.6 | 105.9857 | 106.3333 | 106.14 | 8.425 | 102.8 | 104.699 | 105.4153 | 105.3
-3.5 | 2.2 | 107.8 | 108.6 | 108.2 | 108 | 2.8 | 106.7 | 108.064 | 107.775 | 107.55 | 4.25 | 105.3 | 107.15 | 107.26 | 107.1375 | 6.6 | 103.6 | 105.9 | 106.0667 | 106 | 9.1 | 102.8 | 104.579 | 105.382 | 105.1833
-3.4 | 2.4 | 107.675 | 108.2857 | 107.925 | 107.7 | 3.1 | 106.6 | 107.914 | 107.625 | 107.25 | 4.55 | 105.3 | 107 | 107.08 | 106.96 | 7.2333 | 103.7 | 105.7714 | 105.9143 | 105.8714 | 10 | 102.8 | 104.504 | 105.2487 | 105.0167
-3.3 | 2.575 | 107.5333 | 108.1111 | 107.7 | 107.4375 | 3.4 | 106.5 | 107.639 | 107.4 | 107 | 4.95 | 105.3 | 106.78 | 106.86 | 106.6333 | 7.9 | 103.7 | 105.5 | 105.7429 | 105.66 | 11.0333 | 102.8 | 104.459 | 105.082 | 104.95
-3.2 | 2.7333 | 107.375 | 107.78 | 107.4 | 107.22 | 3.7 | 106.4 | 107.414 | 107.25 | 106.8 | 5.45 | 105.2 | 106.55 | 106.6571 | 106.375 | 8.7 | 103.7 | 105.4 | 105.6667 | 105.4714 | 12.7 | 102.7 | 104.359 | 104.8997 | 104.8
-3.1 | 3.075 | 107.2 | 107.6143 | 107.25 | 107.0143 | 4.1 | 106.4 | 107.264 | 107 | 106.66 | 6.075 | 105.2 | 106.3 | 106.475 | 106.2667 | 9.8 | 103.7 | 105.22 | 105.5143 | 105.3 | 15.0667 | 102.7 | 104.179 | 104.733 | 104.7
-3 | 3.4333 | 107.1 | 107.44 | 106.96 | 106.775 | 4.8 | 106.3 | 106.939 | 106.75 | 106.38 | 6.8 | 105.1 | 106.2143 | 106.1667 | 106.0286 | 11.325 | 103.7 | 105 | 105.3143 | 105.1833 | 18.2 | 102.6 | 104.079 | 104.5955 | 104.5333
-2.9 | 3.98 | 106.825 | 106.9833 | 106.76 | 106.65 | 5.4 | 106.0667 | 106.764 | 106.5 | 106.16 | 8.05 | 105.1 | 106 | 106.02 | 105.9 | 13.55 | 103.5 | 104.85 | 105.125 | 105 | 22.4 | 102.7 | 104.079 | 104.3687 | 104.4
-2.8 | 4.8 | 106.6 | 106.7571 | 106.475 | 106.42 | 6.5 | 105.8 | 106.5142 | 106.25 | 105.925 | 9.55 | 105 | 105.7333 | 105.8333 | 105.675 | 17 | 103.5 | 104.76 | 105.02 | 104.7857 | 28.4 | 102.7 | 103.979 | 104.233 | 104.2714
-2.7 | 5.8667 | 106.4667 | 106.54 | 106.3 | 106.12 | 7.9 | 105.8 | 106.19 | 106.1 | 105.72 | 11.8 | 104.8 | 105.6 | 105.475 | 105.52 | 21.5 | 103.4 | 104.65 | 104.7818 | 104.7 | 36.5 | 102.7 | 103.804 | 104.123 | 104.1286
-2.6 | 7.3 | 106.3 | 106.25 | 105.9889 | 105.88 | 10.1 | 105.6 | 105.965 | 105.7833 | 105.5143 | 15.1 | 104.7 | 105.3571 | 105.2714 | 105.28 | 28.05 | 103.4 | 104.5 | 104.6 | 104.5 | 45.85 | 102.7 | 103.679 | 104.0285 | 104
-2.5 | 9.9 | 106.1 | 106.02 | 105.775 | 105.65 | 13.5 | 105.4 | 105.745 | 105.58 | 105.275 | 20 | 104.6 | 105.2 | 105.1 | 105.1 | 36 | 103.3 | 104.3 | 104.4143 | 104.3 | 55.2 | 102.7 | 103.5957 | 103.858 | 103.8714
-2.4 | 13.7 | 105.8 | 105.74 | 105.5667 | 105.4 | 18.7333 | 105.3 | 105.565 | 105.4143 | 105.0444 | 26.85 | 104.45 | 104.9 | 104.95 | 104.9143 | 46 | 103.3 | 104.175 | 104.2455 | 104.2 | 65.1 | 102.6 | 103.479 | 103.7608 | 103.6889
-2.3 | 19.9 | 105.6 | 105.5 | 105.3286 | 105.175 | 26.45 | 105.1 | 105.365 | 105.15 | 104.875 | 36.3 | 104.3 | 104.7 | 104.74 | 104.7 | 57.2 | 103.2 | 104 | 104.0571 | 104.0143 | 74.2 | 102.425 | 103.279 | 103.558 | 103.5143
-2.2 | 28.9 | 105.5 | 105.375 | 105.1 | 104.9 | 36.2 | 104.85 | 105.065 | 104.925 | 104.6714 | 46.45 | 104.2 | 104.6 | 104.4714 | 104.55 | 67.05 | 103.2 | 103.8 | 103.8889 | 103.9 | 81.9 | 102.4 | 103.1647 | 103.443 | 103.41
-2.1 | 40.2 | 105.2333 | 105.125 | 104.775 | 104.675 | 47.2667 | 104.65 | 104.8983 | 104.7333 | 104.46 | 58 | 104 | 104.5 | 104.3 | 104.34 | 76.5 | 103 | 103.625 | 103.7429 | 103.6571 | 88.3 | 102.4 | 103.029 | 103.333 | 103.26
-2 | 51.3 | 104.9 | 104.8 | 104.65 | 104.4667 | 58.4 | 104.4 | 104.5983 | 104.44 | 104.16 | 68.4 | 103.9 | 104.2 | 104.1 | 104.1 | 84.2 | 102.85 | 103.575 | 103.4333 | 103.52 | 92.2 | 102.3 | 102.879 | 103.1973 | 103.1125
-1.9 | 63.95 | 104.7 | 104.5667 | 104.3714 | 104.28 | 70.35 | 104.2 | 104.365 | 104.1 | 104 | 78.25 | 103.7 | 103.975 | 103.9 | 103.875 | 89.6 | 102.6 | 103.325 | 103.3 | 103.3 | 95.1 | 102.2 | 102.779 | 103.0687 | 102.9429
-1.8 | 75 | 104.5 | 104.24 | 104.2 | 103.9714 | 79.3667 | 104.1 | 104.185 | 103.88 | 103.8 | 85.8 | 103.6 | 103.725 | 103.7667 | 103.65 | 93.7 | 102.5 | 103.175 | 103.1667 | 103.16 | 97.1 | 102.2 | 102.5592 | 102.9259 | 102.8
-1.7 | 85 | 104.15 | 104.1 | 104 | 103.68 | 87.6 | 103.9 | 103.905 | 103.7 | 103.65 | 91.1 | 103.425 | 103.5 | 103.5 | 103.4 | 96 | 102.5 | 103.025 | 103 | 103 | 98.15 | 102.0667 | 102.48 | 102.8052 | 102.64
-1.6 | 91.8 | 103.9 | 103.9 | 103.72 | 103.4714 | 92.65 | 103.6 | 103.725 | 103.54 | 103.4 | 94.9 | 103.2 | 103.4 | 103.3 | 103.175 | 97.8667 | 102.4 | 102.9 | 102.8333 | 102.9 | 99.1 | 102 | 102.28 | 102.6497 | 102.55
-1.5 | 96 | 103.65 | 103.55 | 103.4 | 103.3 | 96 | 103.3 | 103.5056 | 103.3111 | 103.3 | 97.1 | 103.0667 | 103.2 | 103.05 | 103 | 99 | 102.3 | 102.75 | 102.6 | 102.68 | 99.55 | 101.9 | 102.18 | 102.483 | 102.3857
-1.4 | 98.6 | 103.45 | 103.4 | 103.26 | 103.1 | 98.0333 | 103.2 | 103.241 | 103.1 | 103.1 | 98.6 | 102.9 | 103 | 102.9 | 102.9 | 99.5333 | 102.2 | 102.5 | 102.5167 | 102.48 | 100.2 | 101.8 | 102.08 | 102.2163 | 102.2333
-1.3 | 99.8 | 103.2 | 103.1 | 102.98 | 102.875 | 99.5 | 103 | 103.0216 | 102.8143 | 102.75 | 99.4667 | 102.6667 | 102.9 | 102.675 | 102.6 | 99.9 | 102.1 | 102.25 | 102.25 | 102.3 | 100.3667 | 101.6 | 101.88 | 102.083 | 102.0667
-1.2 | 100.9 | 102.9 | 102.85 | 102.66 | 102.74 | 100.1 | 102.7 | 102.806 | 102.6 | 102.52 | 100.05 | 102.425 | 102.55 | 102.5 | 102.45 | 100.4333 | 102 | 102.1 | 102.1833 | 102.2 | 100.5 | 101.6 | 101.73 | 101.908 | 101.9
-1.1 | 101.2 | 102.6 | 102.56 | 102.46 | 102.4 | 100.8 | 102.5 | 102.586 | 102.38 | 102.275 | 100.5 | 102.3 | 102.3 | 102.325 | 102.3 | 100.6667 | 101.9 | 101.9333 | 101.9571 | 102.06 | 100.7 | 101.5 | 101.58 | 101.783 | 101.7333
-1 | 101.4 | 102.3 | 102.38 | 102.2714 | 102.25 | 101 | 102.2 | 102.4231 | 102.16 | 102.1 | 100.7667 | 102.125 | 102.1333 | 102.1 | 102.125 | 100.8 | 101.7 | 101.825 | 101.7571 | 101.825 | 100.8 | 101.4 | 101.48 | 101.623 | 101.5833
-0.9 | 101.5 | 102.1 | 102.16 | 102 | 102 | 101.1 | 102 | 102.186 | 101.94 | 101.94 | 100.9 | 101.9 | 101.9 | 101.78 | 101.8333 | 100.9 | 101.6 | 101.6667 | 101.6 | 101.675 | 100.8667 | 101.3 | 101.34 | 101.483 | 101.5
-0.8 | 101.5 | 101.9333 | 101.9 | 101.8625 | 101.86 | 101.2 | 101.8 | 101.9327 | 101.7 | 101.675 | 100.9 | 101.725 | 101.75 | 101.6 | 101.6 | 101 | 101.4 | 101.475 | 101.4333 | 101.425 | 100.9 | 101.15 | 101.18 | 101.263 | 101.35
-0.7 | 101.4 | 101.6333 | 101.8 | 101.54 | 101.6 | 101.2 | 101.6 | 101.666 | 101.5 | 101.4667 | 100.8 | 101.6 | 101.4 | 101.45 | 101.425 | 101 | 101.2 | 101.375 | 101.1667 | 101.275 | 100.9 | 101 | 101.08 | 100.983 | 101.1833
-0.6 | 101.4 | 101.4 | 101.44 | 101.35 | 101.32 | 101.2 | 101.4 | 101.416 | 101.3 | 101.28 | 100.8 | 101.4 | 101.2 | 101.1667 | 101.275 | 100.9 | 101.1 | 101.14 | 100.9 | 101.14 | 100.8 | 100.85 | 100.855 | 100.8997 | 100.9333
-0.5 | 101.2667 | 101 | 101.2 | 101.1 | 101.025 | 101.1667 | 101.1 | 101.106 | 101.175 | 101.05 | 100.7 | 101.1 | 101.1 | 101 | 101.05 | 100.8 | 100.9 | 101 | 100.8 | 100.98 | 100.7 | 100.7 | 100.705 | 100.783 | 100.8
-0.4 | 101.0667 | 100.9 | 101.025 | 100.96 | 100.875 | 100.9667 | 100.75 | 100.966 | 100.8 | 100.82 | 100.6 | 100.95 | 100.8 | 100.75 | 100.84 | 100.7 | 100.775 | 100.74 | 100.7 | 100.8 | 100.6 | 100.5 | 100.58 | 100.683 | 100.66
-0.3 | 100.8 | 100.8 | 100.75 | 100.65 | 100.6 | 100.7 | 100.5333 | 100.741 | 100.575 | 100.6333 | 100.5 | 100.8 | 100.7 | 100.52 | 100.6333 | 100.5 | 100.6 | 100.6 | 100.55 | 100.55 | 100.45 | 100.4 | 100.48 | 100.5163 | 100.4667
-0.2 | 100.6 | 100.46 | 100.6 | 100.425 | 100.44 | 100.4 | 100.4 | 100.4917 | 100.325 | 100.38 | 100.4 | 100.4333 | 100.45 | 100.4 | 100.44 | 100.4 | 100.4 | 100.3429 | 100.35 | 100.4 | 100.325 | 100.2 | 100.3133 | 100.283 | 100.4
-0.1 | 100.3 | 100.25 | 100.3 | 100.3 | 100.175 | 100.25 | 100.1 | 100.217 | 100.2 | 100.1 | 100.3 | 100.12 | 100.22 | 100.175 | 100.2429 | 100.2 | 100.2 | 100.1 | 100.1833 | 100.2 | 100.1 | 100.1 | 100.13 | 100.1663 | 100.2429
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100.0001 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9999 | 100
0.1 | 99.9 | 99.75 | 99.8 | 99.875 | 99.8143 | 99.85 | 99.7 | 99.8467 | 99.8 | 99.82 | 99.8 | 99.8333 | 99.75 | 99.8 | 99.8556 | 99.9 | 99.7667 | 99.84 | 99.8667 | 99.9 | 99.9 | 99.8667 | 99.78 | 99.8334 | 99.7333
0.2 | 99.7 | 99.48 | 99.625 | 99.65 | 99.7 | 99.6 | 99.44 | 99.5668 | 99.6 | 99.5571 | 99.5 | 99.65 | 99.46 | 99.5 | 99.7286 | 99.7 | 99.6 | 99.6714 | 99.6 | 99.7 | 99.7667 | 99.8 | 99.6601 | 99.6834 | 99.6167
0.3 | 99.4333 | 99.3 | 99.3667 | 99.3333 | 99.4333 | 99.4 | 99.25 | 99.3669 | 99.4 | 99.4 | 99.3 | 99.4333 | 99.3125 | 99.4 | 99.52 | 99.5333 | 99.3333 | 99.5 | 99.4 | 99.525 | 99.5 | 99.6 | 99.5468 | 99.612 | 99.5
0.4 | 99.1 | 99 | 99.1 | 99.1 | 99.3 | 99.125 | 99 | 99.1669 | 99.04 | 99.2429 | 99 | 99.2 | 99.1 | 99.12 | 99.28 | 99.2 | 99.125 | 99.3857 | 99.2333 | 99.4 | 99.25 | 99.35 | 99.3801 | 99.5034 | 99.4143
0.5 | 98.6333 | 98.82 | 98.92 | 98.94 | 99.0857 | 98.8 | 98.85 | 98.987 | 98.9 | 98.98 | 98.75 | 99 | 98.9167 | 98.9556 | 99.0714 | 98.95 | 98.975 | 99.2 | 99.1 | 99.175 | 99.1 | 99.2 | 99.1802 | 99.1835 | 99.3
0.6 | 98.325 | 98.525 | 98.7 | 98.8 | 98.9 | 98.525 | 98.625 | 98.842 | 98.72 | 98.75 | 98.5 | 98.8 | 98.7667 | 98.8 | 98.9286 | 98.7 | 98.825 | 98.96 | 98.9333 | 99 | 98.975 | 99 | 99.0802 | 99.0168 | 99.1857
0.7 | 98.05 | 98.275 | 98.52 | 98.6 | 98.7375 | 98.2 | 98.4 | 98.5671 | 98.48 | 98.525 | 98.2 | 98.55 | 98.5333 | 98.6286 | 98.65 | 98.35 | 98.5 | 98.7333 | 98.6 | 98.9 | 98.7333 | 98.8 | 98.7677 | 98.8978 | 98.96
0.8 | 97.75 | 98.025 | 98.3 | 98.3 | 98.45 | 97.85 | 98.0333 | 98.3922 | 98.3 | 98.3333 | 97.6 | 98.4 | 98.375 | 98.38 | 98.5 | 98.1 | 98.35 | 98.6 | 98.5167 | 98.6333 | 98.5 | 98.7 | 98.6203 | 98.6978 | 98.725
0.9 | 97.2 | 97.8 | 98.02 | 98.1 | 98.2857 | 97.2 | 97.8 | 98.1673 | 98.25 | 98.18 | 97.3333 | 98.2 | 98.2 | 98.1667 | 98.325 | 97.7 | 98.2 | 98.45 | 98.28 | 98.4 | 98.3 | 98.5333 | 98.5303 | 98.5407 | 98.5778
1 | 96.6667 | 97.625 | 97.7714 | 97.9167 | 98.12 | 96.6667 | 97.6667 | 97.9173 | 97.95 | 98 | 96.7 | 97.9 | 98 | 97.9286 | 98.2 | 97.25 | 97.9667 | 98.26 | 98.1111 | 98.3 | 98 | 98.2333 | 98.3803 | 98.3745 | 98.4625
1.1 | 95.9 | 97.4 | 97.6167 | 97.675 | 97.7889 | 95.8667 | 97.5 | 97.7174 | 97.7429 | 97.825 | 95.95 | 97.7 | 97.7333 | 97.8 | 97.975 | 96.8 | 97.7333 | 98.1 | 97.9429 | 98.175 | 97.65 | 98.1 | 98.1604 | 98.1837 | 98.3286
1.2 | 94.9 | 97.1 | 97.42 | 97.425 | 97.5286 | 95.04 | 97.15 | 97.5075 | 97.54 | 97.575 | 95.2 | 97.5333 | 97.6 | 97.6 | 97.8 | 96.3333 | 97.475 | 97.875 | 97.7444 | 97.88 | 97 | 97.9 | 98.0054 | 98.0212 | 98.2111
1.3 | 93.75 | 96.8 | 97.2571 | 97.3333 | 97.4 | 93.8 | 96.9 | 97.3076 | 97.3 | 97.35 | 94.3 | 97.25 | 97.36 | 97.38 | 97.58 | 95.7 | 97.4 | 97.625 | 97.5286 | 97.725 | 96.3667 | 97.8 | 97.8804 | 97.8962 | 97.9714
1.4 | 91.9 | 96.6667 | 96.9 | 97.175 | 97.3 | 92.275 | 96.7 | 97.0676 | 97.1 | 97.2 | 92.9667 | 96.95 | 97.125 | 97.1 | 97.3 | 94.8 | 97.18 | 97.5 | 97.4182 | 97.55 | 95.7 | 97.5667 | 97.7304 | 97.7037 | 97.8091
1.5 | 89.1667 | 96.3 | 96.7 | 97 | 97.14 | 90.0667 | 96.5 | 96.8427 | 96.9429 | 97 | 91.4 | 96.7 | 96.84 | 97 | 97.14 | 93.8 | 96.95 | 97.35 | 97.2857 | 97.3 | 95.2 | 97.3 | 97.4805 | 97.4656 | 97.5857
1.6 | 85.1 | 96.1667 | 96.4 | 96.8 | 96.9429 | 86.7 | 96.2 | 96.6345 | 96.75 | 96.7 | 88.85 | 96.45 | 96.74 | 96.78 | 97 | 92.575 | 96.7 | 97.1 | 97.0909 | 97.1625 | 94.4 | 97.1 | 97.2555 | 97.3695 | 97.4
1.7 | 78.4333 | 95.9 | 96.2571 | 96.55 | 96.7 | 81.5667 | 95.9333 | 96.3479 | 96.6143 | 96.4 | 85.55 | 96.2667 | 96.52 | 96.56 | 96.78 | 90.85 | 96.4667 | 96.8 | 96.825 | 97.08 | 93.15 | 96.8 | 97.0556 | 97.2038 | 97.2714
1.8 | 69.5 | 95.7 | 96.0286 | 96.3 | 96.3833 | 74.2667 | 95.6 | 96.0347 | 96.32 | 96.275 | 80.6 | 96 | 96.24 | 96.34 | 96.56 | 88.56 | 96.2 | 96.55 | 96.6 | 96.9 | 92.2 | 96.5 | 96.8206 | 97.0171 | 97.02
1.9 | 59.0333 | 95.5 | 95.8 | 96 | 96.2 | 65.8 | 95.4 | 95.843 | 96.0667 | 96.125 | 73.95 | 95.6 | 96 | 96.1667 | 96.3 | 84.9 | 95.9 | 96.4 | 96.4286 | 96.68 | 90 | 96.3333 | 96.6807 | 96.7839 | 96.85
2 | 47.2 | 95.1 | 95.5833 | 95.82 | 96.0857 | 54.9 | 95.2 | 95.6281 | 95.9 | 96 | 65.2 | 95.4 | 95.7 | 95.9 | 96.1333 | 80.15 | 95.7 | 96.2 | 96.3 | 96.45 | 87.475 | 96.1 | 96.5807 | 96.6439 | 96.7333
2.1 | 36.4333 | 94.95 | 95.3667 | 95.7 | 95.8 | 44.2 | 94.9 | 95.4482 | 95.6 | 95.7 | 55.5 | 95.1667 | 95.5333 | 95.75 | 95.9 | 73.3 | 95.275 | 96 | 95.9714 | 96.34 | 83.85 | 95.95 | 96.3807 | 96.4506 | 96.6125
2.2 | 26.5 | 94.55 | 95.1 | 95.475 | 95.58 | 33.8 | 94.6 | 95.2932 | 95.475 | 95.46 | 45.225 | 94.8667 | 95.325 | 95.5 | 95.65 | 65.1 | 95.1 | 95.775 | 95.8375 | 96.075 | 79.4 | 95.7 | 96.1308 | 96.284 | 96.36
2.3 | 18.325 | 94.25 | 94.95 | 95.325 | 95.5 | 24.4 | 94.5 | 95.054 | 95.25 | 95.34 | 35.2 | 94.5 | 95.125 | 95.3 | 95.5 | 55.8 | 94.8 | 95.575 | 95.7125 | 95.9 | 73.55 | 95.3333 | 95.8808 | 96.1507 | 96.175
2.4 | 12.45 | 94.05 | 94.675 | 95.125 | 95.3 | 17.725 | 94.3 | 94.8434 | 95.075 | 95.075 | 26.3667 | 94.34 | 94.9 | 95.2 | 95.35 | 46.1 | 94.5333 | 95.38 | 95.5333 | 95.6857 | 65.4 | 95.1 | 95.6809 | 95.9673 | 96.05
2.5 | 9 | 93.7667 | 94.45 | 94.9 | 95.05 | 12.8 | 93.9 | 94.6684 | 94.875 | 94.9 | 19.5 | 94 | 94.725 | 95.05 | 95.2 | 36.4 | 94.3 | 95.0333 | 95.3143 | 95.475 | 56.4333 | 94.8667 | 95.5809 | 95.784 | 95.94
2.6 | 6.6667 | 93.6 | 94.2 | 94.7 | 94.85 | 9.5 | 93.6 | 94.4085 | 94.7 | 94.7 | 14.5 | 93.6667 | 94.475 | 94.825 | 95 | 28.2 | 93.95 | 94.8 | 95.0667 | 95.28 | 47.4 | 94.46 | 95.3059 | 95.6341 | 95.8143
2.7 | 5.4 | 93.4 | 93.95 | 94.3429 | 94.6 | 7.5 | 93.2 | 94.0544 | 94.45 | 94.5 | 11.3 | 93.4 | 94.2333 | 94.575 | 94.8 | 21.6 | 93.7 | 94.6 | 94.8667 | 95.08 | 38.2 | 94.0667 | 95.181 | 95.4441 | 95.5545
2.8 | 4.3333 | 92.9667 | 93.7 | 94.2 | 94.35 | 6.2 | 92.9 | 93.8087 | 94.2 | 94.3 | 9.1 | 93.05 | 94 | 94.325 | 94.5667 | 16.9 | 93.5 | 94.4667 | 94.6 | 94.9 | 30.1 | 93.825 | 95.081 | 95.3841 | 95.2857
2.9 | 3.66 | 92.7 | 93.4 | 93.96 | 94.1 | 5.2 | 92.7 | 93.5888 | 94.1 | 94.2 | 7.6 | 92.8 | 93.85 | 94.1 | 94.425 | 13.6 | 93.2 | 94.275 | 94.5 | 94.6 | 23.6 | 93.7 | 94.781 | 95.1508 | 95.1556
3 | 3.2 | 92.3 | 93.2667 | 93.8 | 93.9 | 4.5 | 92.4 | 93.4688 | 93.9 | 93.9 | 6.65 | 92.475 | 93.7 | 93.9 | 94.25 | 11.1667 | 92.8333 | 94.075 | 94.25 | 94.4 | 18.7 | 93.45 | 94.5144 | 94.9509 | 95.0286
3.1 | 2.8 | 92.15 | 93.025 | 93.56 | 93.7333 | 3.95 | 92.0667 | 93.1801 | 93.65 | 93.7 | 5.8 | 92.1 | 93.45 | 93.7 | 94.025 | 9.7 | 92.4 | 93.92 | 94.1 | 94.24 | 15.2 | 93.2 | 94.3811 | 94.6842 | 94.79
3.2 | 2.5 | 91.9 | 92.7667 | 93.36 | 93.55 | 3.45 | 91.8 | 92.8548 | 93.45 | 93.475 | 5.15 | 91.95 | 93.2 | 93.5 | 93.7333 | 8.5 | 92.15 | 93.6333 | 94 | 94.15 | 12.8 | 92.8 | 94.2062 | 94.5271 | 94.6714
3.3 | 2.3 | 91.6 | 92.65 | 93.15 | 93.34 | 3.2 | 91.5 | 92.5941 | 93.1 | 93.3 | 4.8 | 91.5 | 92.875 | 93.35 | 93.54 | 7.5667 | 91.9 | 93.4 | 93.7286 | 93.95 | 11.1 | 92.5333 | 94.0145 | 94.3676 | 94.4
3.4 | 2.2 | 91.25 | 92.35 | 92.96 | 93.1667 | 2.9 | 91.2 | 92.4492 | 92.96 | 93.1333 | 4.3 | 91.2 | 92.6 | 93.2 | 93.4 | 7 | 91.4 | 93.175 | 93.5571 | 93.75 | 9.95 | 92.275 | 93.7213 | 94.07 | 94.225
3.5 | 2.025 | 90.7 | 92.2 | 92.6714 | 92.975 | 2.6 | 90.6 | 92.1693 | 92.7 | 92.9 | 4 | 90.85 | 92.5167 | 92.92 | 93 | 6.4 | 91.1 | 92.94 | 93.3 | 93.56 | 8.8 | 91.95 | 93.5413 | 93.74 | 94
3.6 | 1.9 | 90.5 | 91.9333 | 92.5 | 92.7429 | 2.5 | 90.2 | 91.9693 | 92.5333 | 92.6 | 3.8 | 90.4 | 92.2 | 92.625 | 92.8333 | 5.95 | 90.7333 | 92.8 | 93.1667 | 93.4 | 8.0667 | 91.52 | 93.3213 | 93.5844 | 93.8571
3.7 | 1.8333 | 90.25 | 91.7 | 92.3 | 92.4571 | 2.3 | 89.8 | 91.7694 | 92.275 | 92.42 | 3.6 | 89.95 | 92.0333 | 92.475 | 92.6 | 5.65 | 90.3 | 92.5 | 92.9429 | 93.26 | 7.7 | 91.2333 | 93.1064 | 93.513 | 93.68
3.8 | 1.7 | 89.75 | 91.4 | 92.125 | 92.225 | 2.2 | 89.55 | 91.5945 | 92.125 | 92.1571 | 3.4 | 89.5 | 91.82 | 92.3 | 92.4667 | 5.3 | 89.975 | 92.2 | 92.7167 | 93.1 | 7.2 | 90.9 | 92.9814 | 93.2445 | 93.525
3.9 | 1.7 | 89.3333 | 91.2 | 91.9 | 92.06 | 2.1 | 89.05 | 91.3695 | 91.9 | 91.98 | 3.2 | 89.15 | 91.5667 | 92.1333 | 92.275 | 5.2 | 89.625 | 92.05 | 92.5167 | 92.75 | 6.8 | 90.65 | 92.7565 | 93.0645 | 93.3
4 | 1.6 | 88.7667 | 91 | 91.6 | 91.8333 | 2 | 88.45 | 91.1446 | 91.7 | 91.75 | 3 | 88.7 | 91.325 | 91.9 | 92.125 | 4.825 | 89.3 | 91.85 | 92.3571 | 92.52 | 6.45 | 90.42 | 92.3415 | 92.8734 | 93.1429
4.1 | 1.5 | 88.05 | 90.8 | 91.3667 | 91.6429 | 1.9 | 87.8 | 90.8947 | 91.5 | 91.5571 | 2.9 | 88.2 | 91.1333 | 91.65 | 91.975 | 4.7 | 88.85 | 91.6 | 92.2143 | 92.35 | 6.2 | 89.8667 | 92.1816 | 92.7512 | 92.9
4.2 | 1.5 | 87.3 | 90.625 | 91.2 | 91.45 | 1.8 | 87.275 | 90.6448 | 91.25 | 91.4 | 2.8 | 87.75 | 90.84 | 91.4333 | 91.725 | 4.45 | 88.5 | 91.45 | 91.9 | 92.175 | 5.9667 | 89.4 | 91.9816 | 92.6096 | 92.7375
4.3 | 1.4 | 86.3 | 90.2667 | 91.05 | 91.2 | 1.8 | 86.5 | 90.3699 | 91 | 91.2571 | 2.65 | 87.05 | 90.5857 | 91.175 | 91.5 | 4.3 | 88.02 | 91.125 | 91.6875 | 91.925 | 5.725 | 89.0333 | 91.8149 | 92.4346 | 92.5714
4.4 | 1.4 | 85.2 | 89.975 | 90.84 | 90.98 | 1.7 | 85.6 | 90.1499 | 90.85 | 90.95 | 2.6 | 86.3333 | 90.3 | 90.9 | 91.2 | 4.15 | 87.6 | 91 | 91.5571 | 91.8 | 5.6 | 88.6 | 91.6217 | 92.2046 | 92.375
4.5 | 1.4 | 83.3333 | 89.725 | 90.5667 | 90.7 | 1.6 | 84.4 | 89.87 | 90.6 | 90.7333 | 2.45 | 85.7667 | 90.0714 | 90.84 | 90.98 | 4 | 87.2 | 90.68 | 91.3286 | 91.45 | 5.325 | 88.3 | 91.2817 | 92.0276 | 92.2
4.6 | 1.3 | 80.7 | 89.375 | 90.4 | 90.55 | 1.6 | 82.7 | 89.7701 | 90.5 | 90.5 | 2.4 | 84.8 | 89.9 | 90.6167 | 90.75 | 3.9 | 86.6333 | 90.42 | 91.1143 | 91.3 | 5.175 | 88 | 91.1618 | 91.8847 | 91.9857
4.7 | 1.3 | 76.475 | 89.24 | 90.05 | 90.325 | 1.5 | 80.4 | 89.4701 | 90.275 | 90.36 | 2.3 | 83.7 | 89.6556 | 90.36 | 90.64 | 3.8 | 85.9 | 90.3 | 90.8286 | 91.15 | 5 | 87.7 | 90.9104 | 91.5348 | 91.7857
4.8 | 1.3 | 70.1333 | 89 | 89.9 | 90.0667 | 1.5 | 77 | 89.2702 | 90 | 90.1 | 2.2 | 82.14 | 89.26 | 90.15 | 90.4 | 3.66 | 85.4 | 90.0667 | 90.7111 | 90.9 | 4.9 | 87.4 | 90.6819 | 91.2848 | 91.6
4.9 | 1.2 | 62.36 | 88.8 | 89.6 | 89.88 | 1.4 | 72.2 | 88.9953 | 89.7143 | 89.88 | 2.1 | 80.5 | 89.05 | 90 | 90.1714 | 3.5 | 84.8 | 89.78 | 90.55 | 90.7 | 4.7667 | 86.75 | 90.4694 | 91.2419 | 91.4714
5 | 1.2 | 53 | 88.55 | 89.4333 | 89.65 | 1.4 | 64.9 | 88.7454 | 89.5 | 89.625 | 2.05 | 77.8 | 88.9 | 89.74 | 90 | 3.4 | 84 | 89.54 | 90.28 | 90.425 | 4.6 | 86.3 | 90.2105 | 91.0648 | 91.1667
5.1 | 1.2 | 43.02 | 88.3 | 89.2 | 89.45 | 1.3 | 56.9 | 88.4955 | 89.2571 | 89.4 | 2 | 73.7667 | 88.62 | 89.56 | 89.7714 | 3.3 | 83 | 89.3 | 90.0429 | 90.2333 | 4.4667 | 85.9 | 90.022 | 90.8135 | 90.9667
5.2 | 1.1 | 33.5 | 88 | 89 | 89.2 | 1.3 | 47.9 | 88.2456 | 89.05 | 89.2 | 1.9 | 68.15 | 88.3667 | 89.2333 | 89.5 | 3.175 | 81.9 | 88.95 | 89.8727 | 90 | 4.4 | 85.1 | 89.782 | 90.5749 | 90.8
5.3 | 1.1 | 25.1 | 87.82 | 88.825 | 89.06 | 1.3 | 38.4 | 88.0107 | 88.775 | 88.95 | 1.8 | 61 | 88.175 | 89.1 | 89.2889 | 3.1 | 80.5667 | 88.8 | 89.65 | 89.8333 | 4.3 | 84.5 | 89.5821 | 90.4563 | 90.5667
5.4 | 1.1 | 18.5 | 87.525 | 88.5 | 88.82 | 1.2 | 29.9667 | 87.8307 | 88.6429 | 88.725 | 1.8 | 53 | 87.85 | 88.9 | 89.15 | 3 | 78.8 | 88.475 | 89.44 | 89.64 | 4.2 | 83.85 | 89.2108 | 90.2406 | 90.3833
5.5 | 1.1 | 13.9 | 87.275 | 88.3167 | 88.56 | 1.2 | 22.8 | 87.4209 | 88.45 | 88.6 | 1.8 | 44.2 | 87.675 | 88.62 | 88.9 | 2.9667 | 76.0667 | 88.25 | 89.2571 | 89.5 | 4.1 | 83.0667 | 89.0679 | 90.0421 | 90.2667
5.6 | 1 | 11.2 | 87.025 | 88.075 | 88.325 | 1.2 | 17.4 | 87.2209 | 88.15 | 88.3 | 1.7 | 36 | 87.44 | 88.3 | 88.625 | 2.9 | 72.1 | 87.975 | 89 | 89.2 | 4 | 82.1 | 88.8323 | 89.685 | 90.0167
5.7 | 1 | 9.4 | 86.775 | 87.84 | 88.075 | 1.1 | 13.8 | 86.971 | 88 | 88.05 | 1.7 | 28.5 | 87.15 | 88.125 | 88.3 | 2.7667 | 67.2 | 87.85 | 88.7 | 89 | 3.9 | 80.9667 | 88.5823 | 89.5565 | 89.82
5.8 | 1 | 8.2 | 86.54 | 87.62 | 87.925 | 1.1 | 11.55 | 86.7911 | 87.76 | 87.82 | 1.6 | 22 | 86.9 | 88 | 88.16 | 2.7 | 60.9 | 87.5 | 88.5667 | 88.75 | 3.8 | 79.45 | 88.3223 | 89.3601 | 89.6125
5.9 | 1 | 7.3 | 86.4 | 87.375 | 87.775 | 1.1 | 10 | 86.3712 | 87.4333 | 87.675 | 1.5333 | 17.8 | 86.55 | 87.6714 | 87.9571 | 2.7 | 53.05 | 87.22 | 88.3167 | 88.5 | 3.7 | 77.675 | 88.0824 | 89.1851 | 89.425
6 | 1 | 6.7 | 86.1 | 87.22 | 87.5 | 1.1 | 8.9 | 86.1313 | 87.2 | 87.425 | 1.5 | 15 | 86.3 | 87.4 | 87.7286 | 2.6 | 45.4 | 86.95 | 88.1571 | 88.3667 | 3.6 | 75 | 87.7824 | 88.9519 | 89.1571
6.1 | 1 | 6.3 | 85.85 | 86.9 | 87.34 | 1 | 8.05 | 85.9838 | 87 | 87.15 | 1.5 | 13.1 | 86.1 | 87.2444 | 87.52 | 2.5 | 38.3 | 86.7 | 88.0143 | 88.1 | 3.5 | 71.5333 | 87.6325 | 88.8519 | 88.8333
6.2 | 1 | 5.9 | 85.5 | 86.74 | 86.95 | 1 | 7.45 | 85.7714 | 86.86 | 86.975 | 1.45 | 11.7 | 85.8 | 86.94 | 87.3 | 2.4 | 31.5 | 86.5 | 87.8 | 87.9333 | 3.4 | 66.6 | 87.5025 | 88.5852 | 88.7
6.3 | 1 | 5.6 | 85.32 | 86.52 | 86.8 | 1 | 6.9 | 85.5048 | 86.625 | 86.7 | 1.4 | 10.7 | 85.6 | 86.6714 | 87.12 | 2.4 | 25.5 | 86.3 | 87.5667 | 87.7143 | 3.4 | 60.5333 | 87.1826 | 88.302 | 88.5714
6.4 | 0.9 | 5.4 | 85.05 | 86.34 | 86.55 | 1 | 6.4667 | 85.2916 | 86.375 | 86.5 | 1.4 | 9.9 | 85.175 | 86.58 | 86.88 | 2.4 | 21.2 | 85.875 | 87.3286 | 87.56 | 3.2667 | 53.6333 | 86.9826 | 88.1453 | 88.4167
6.5 | 0.9 | 5.1 | 84.825 | 86.04 | 86.3 | 1 | 6.15 | 84.9288 | 86.1 | 86.25 | 1.4 | 9.3 | 85.025 | 86.35 | 86.72 | 2.3 | 18.3667 | 85.58 | 87.1333 | 87.38 | 3.2 | 46.8 | 86.6077 | 88.0853 | 88.2167
6.6 | 0.9 | 5 | 84.54 | 85.74 | 86.1 | 0.975 | 5.9 | 84.6718 | 85.925 | 86 | 1.3 | 8.7 | 84.775 | 86.08 | 86.4 | 2.3 | 16.4 | 85.3286 | 86.9143 | 87.0833 | 3.1 | 39.9 | 86.1628 | 87.7521 | 88.05
6.7 | 0.9 | 4.8 | 84.275 | 85.525 | 85.9 | 0.9 | 5.6333 | 84.4385 | 85.68 | 85.82 | 1.3 | 8.3 | 84.525 | 85.875 | 86.16 | 2.2 | 15 | 85 | 86.6857 | 86.8571 | 3 | 33.25 | 85.8828 | 87.5454 | 87.7778
6.8 | 0.9 | 4.7 | 84.025 | 85.42 | 85.7 | 0.9 | 5.4 | 84.2319 | 85.44 | 85.66 | 1.3 | 7.95 | 84.25 | 85.6 | 86 | 2.2 | 13.9 | 84.7778 | 86.4667 | 86.6 | 3 | 27.8333 | 85.7828 | 87.3568 | 87.5778
6.9 | 0.9 | 4.5333 | 83.64 | 85.15 | 85.4 | 0.9 | 5.2 | 83.932 | 85.2 | 85.4 | 1.3 | 7.6667 | 84 | 85.38 | 85.78 | 2.1 | 13.1 | 84.56 | 86.1333 | 86.45 | 2.9 | 23.8 | 85.4829 | 87.1055 | 87.3875
7 | 0.9 | 4.4 | 83.4 | 84.825 | 85.225 | 0.9 | 5.05 | 83.7121 | 85 | 85.15 | 1.2 | 7.3333 | 83.7 | 85.2 | 85.5 | 2 | 12.4667 | 84.35 | 85.95 | 86.225 | 2.8 | 21.1 | 85.183 | 86.8712 | 87.05
7.1 |  | 4.3 | 83.175 | 84.575 | 84.975 | 0.9 | 4.9 | 83.3972 | 84.75 | 84.9571 | 1.2 | 7.1 | 83.42 | 85 | 85.28 | 2 | 11.9 | 84.02 | 85.6833 | 86 | 2.8 | 19.1 | 85.033 | 86.6455 | 86.8286
7.2 |  | 4.2 | 82.9429 | 84.44 | 84.725 |  | 4.7333 | 83.0723 | 84.52 | 84.72 | 1.2 | 6.8667 | 83 | 84.82 | 85.075 | 1.975 | 11.4 | 83.7429 | 85.5167 | 85.8286 | 2.75 | 17.4 | 84.783 | 86.4523 | 86.68
7.3 |  | 4.1667 | 82.68 | 84.2 | 84.5 |  | 4.6 | 82.8057 | 84.375 | 84.4667 |  | 6.625 | 82.675 | 84.6 | 84.84 | 1.9 | 11 | 83.4571 | 85.32 | 85.58 | 2.7 | 16.35 | 84.4831 | 86.119 | 86.475
7.4 |  | 4 | 82.42 | 84 | 84.25 |  | 4.5 | 82.5296 | 84.14 | 84.28 |  | 6.45 | 82.525 | 84.3 | 84.65 | 1.9 | 10.6 | 83.14 | 85.1 | 85.36 | 2.7 | 15.55 | 84.2232 | 85.9357 | 86.225
7.5 |  | 3.9 | 82.2125 | 83.8 | 83.9 |  | 4.3 | 82.2326 | 83.8 | 84.05 |  | 6.3 | 82.24 | 84.1 | 84.4667 |  | 10.325 | 82.86 | 84.8333 | 85.16 | 2.7 | 14.78 | 83.7582 | 85.769 | 86.1
7.6 |  | 3.8 | 81.84 | 83.4 | 83.75 |  | 4.3 | 82.0948 | 83.55 | 83.825 |  | 6.1 | 81.94 | 83.88 | 84.2 |  | 10 | 82.5857 | 84.5857 | 84.8875 |  | 14.3333 | 83.5833 | 85.5429 | 85.9
7.7 |  | 3.8 | 81.5333 | 83.25 | 83.5 |  | 4.2 | 81.7927 | 83.3 | 83.54 |  | 5.8667 | 81.64 | 83.575 | 83.95 |  | 9.8 | 82.3 | 84.36 | 84.68 |  | 13.8 | 83.2833 | 85.3525 | 85.6
7.8 |  | 3.7 | 81.32 | 83.0667 | 83.325 |  | 4.0333 | 81.5328 | 83.2 | 83.325 |  | 5.8 | 81.46 | 83.325 | 83.8 |  | 9.4 | 82.0667 | 84.2125 | 84.5 |  | 13.3333 | 83.0834 | 85.1858 | 85.4571
7.9 |  | 3.6 | 81.1 | 82.72 | 83.075 |  | 3.9 | 81.173 | 83 | 83.04 |  | 5.6 | 81.18 | 83.14 | 83.6 |  | 9.2 | 81.8 | 84 | 84.325 |  | 13.05 | 82.7235 | 84.9192 | 85.3
8 |  | 3.55 | 80.76 | 82.575 | 82.825 |  | 3.9 | 80.873 | 82.7111 | 82.775 |  | 5.4 | 80.82 | 82.98 | 83.275 |  | 8.9 | 81.46 | 83.74 | 84.075 |  | 12.65 | 82.3835 | 84.7288 | 85.0625
8.1 |  | 3.5 | 80.45 | 82.225 | 82.6 |  | 3.8 | 80.5731 | 82.4 | 82.575 |  | 5.3 | 80.4 | 82.74 | 83.025 |  | 8.8 | 81.16 | 83.525 | 83.7667 |  | 12.25 | 82.155 | 84.5026 | 84.875
8.2 |  | 3.4 | 80.3 | 82.1 | 82.44 |  | 3.7 | 80.2483 | 82.2 | 82.45 |  | 5.1333 | 80.075 | 82.5 | 82.775 |  | 8.6 | 80.8 | 83.2889 | 83.52 |  | 11.95 | 81.9265 | 84.3359 | 84.5
8.3 |  | 3.4 | 79.9333 | 81.8143 | 82.14 |  | 3.6 | 79.8984 | 82.06 | 82.2 |  | 5 | 79.9 | 82.25 | 82.525 |  | 8.3 | 80.5 | 83.0667 | 83.4 |  | 11.7 | 81.8436 | 84.1574 | 84.3333
8.4 |  | 3.3 | 79.7 | 81.58 | 81.9 |  | 3.55 | 79.6401 | 81.825 | 81.9 |  | 4.9 | 79.575 | 82 | 82.3 |  | 8.2 | 80.2167 | 82.8667 | 83.18 |  | 11.6 | 81.4337 | 83.9193 | 84.0714
8.5 |  | 3.2 | 79.38 | 81.35 | 81.725 |  | 3.5 | 79.3236 | 81.62 | 81.8 |  | 4.8 | 79.2143 | 81.7667 | 82.18 |  | 8 | 79.94 | 82.5714 | 82.8875 |  | 11.2333 | 81.0838 | 83.6575 | 83.9167
8.6 |  | 3.2 | 78.975 | 81.14 | 81.5 |  | 3.4 | 78.9937 | 81.3 | 81.4 |  | 4.7 | 78.8571 | 81.6 | 82 |  | 7.8 | 79.625 | 82.2857 | 82.74 |  | 11 | 80.7981 | 83.4432 | 83.65
8.7 |  | 3.1 | 78.74 | 80.9 | 81.275 |  | 3.3 | 78.6738 | 81.06 | 81.15 |  | 4.6 | 78.65 | 81.4 | 81.7 |  | 7.6 | 79.2 | 82.0857 | 82.4714 |  | 10.7 | 80.5339 | 83.1861 | 83.5
8.8 |  | 3.1 | 78.3714 | 80.68 | 80.9 |  | 3.26 | 78.2989 | 80.825 | 81.025 |  | 4.5 | 78.225 | 81.0333 | 81.36 |  | 7.5 | 78.9 | 81.92 | 82.1 |  | 10.5 | 80.2506 | 82.9576 | 83.2333
8.9 |  | 3.1 | 78 | 80.375 | 80.68 |  | 3.2 | 77.949 | 80.64 | 80.85 |  | 4.4 | 77.94 | 80.8833 | 81.14 |  | 7.2667 | 78.6 | 81.6571 | 81.9778 |  | 10.175 | 79.934 | 82.7291 | 82.9667
9 |  | 3 | 77.55 | 80.15 | 80.42 |  | 3.2 | 77.6241 | 80.3143 | 80.525 |  | 4.3 | 77.55 | 80.7 | 80.9333 |  | 7.1 | 78.3625 | 81.5143 | 81.75 |  | 9.95 | 79.6174 | 82.5113 | 82.8
9.1 |  | 3 | 77.2 | 79.925 | 80.25 |  | 3.1 | 77.1743 | 80.1 | 80.375 |  | 4.2 | 77.2 | 80.425 | 80.7 |  | 7 | 77.98 | 81.25 | 81.6143 |  | 9.8 | 79.2641 | 82.253 | 82.5333
9.2 |  | 2.9 | 76.6444 | 79.8 | 80.125 |  | 3.1 | 76.7344 | 79.9 | 80.1333 |  | 4.1 | 76.85 | 80.2 | 80.4833 |  | 6.8333 | 77.8 | 81.1 | 81.3286 |  | 9.7 | 79.0342 | 82.0663 | 82.3667
9.3 |  | 2.9 | 76.05 | 79.65 | 79.9 |  | 3 | 76.2746 | 79.7 | 79.9 |  | 4 | 76.45 | 79.94 | 80.3 |  | 6.7 | 77.3333 | 80.8333 | 81 |  | 9.58 | 78.7643 | 81.7864 | 82.1
9.4 |  | 2.8 | 75.3429 | 79.4 | 79.7333 |  | 3 | 75.7748 | 79.4375 | 79.8 |  | 3.9 | 76 | 79.8 | 79.975 |  | 6.5333 | 77.0571 | 80.54 | 80.9 |  | 9.4 | 78.5272 | 81.4864 | 81.95
9.5 |  | 2.8 | 74.5714 | 79.3 | 79.6 |  | 2.9 | 75.2749 | 79.26 | 79.4667 |  | 3.9 | 75.6 | 79.48 | 79.85 |  | 6.4 | 76.6333 | 80.3667 | 80.62 |  | 9.175 | 78.1844 | 81.3955 | 81.7833
9.6 |  | 2.8 | 73.1 | 79.1 | 79.44 |  | 2.9 | 74.5751 | 79.05 | 79.325 |  | 3.8 | 75.2 | 79.25 | 79.6714 |  | 6.3 | 76.3 | 80 | 80.4 |  | 9 | 77.7701 | 81.132 | 81.6167
9.7 |  | 2.7 | 71.18 | 79 | 79.2 |  | 2.8 | 73.7754 | 78.8 | 79.12 |  | 3.7 | 74.8 | 79.1 | 79.42 |  | 6.1 | 75.92 | 79.9429 | 80.16 |  | 8.8 | 77.5845 | 80.9294 | 81.3
9.8 |  | 2.7 | 68.2 | 78.8 | 79.1 |  | 2.8 | 72.6758 | 78.7 | 78.9 |  | 3.7 | 74.28 | 78.82 | 79.2 |  | 6 | 75.62 | 79.64 | 79.88 |  | 8.6667 | 77.1846 | 80.6865 | 81.0833
9.9 |  | 2.65 | 63.6 | 78.725 | 79.0111 |  | 2.8 | 71.3262 | 78.5143 | 78.7857 |  | 3.6 | 73.75 | 78.6167 | 79.025 |  | 5.9 | 75.28 | 79.3667 | 79.7 |  | 8.5 | 76.8646 | 80.4266 | 80.8571
10 |  | 2.6 | 56.7857 | 78.575 | 78.9 |  | 2.7 | 69.2436 | 78.4 | 78.6 |  | 3.5 | 73.14 | 78.4 | 78.725 |  | 5.8 | 74.84 | 79.1167 | 79.5 |  | 8.4 | 76.5347 | 80.1978 | 80.6
10.1 |  | 2.6 | 49.7 | 78.5 | 78.7571 |  | 2.7 | 66.078 | 78.3 | 78.38 |  | 3.5 | 72.46 | 78.2429 | 78.6167 |  | 5.7 | 74.46 | 78.95 | 79.3 |  | 8.2 | 76.0515 | 79.9617 | 80.3
10.2 |  |  | 41.3667 | 78.3 | 78.6 |  | 2.7 | 61.7794 | 78.12 | 78.3 |  | 3.4 | 71.58 | 78.1 | 78.4 |  | 5.6 | 74.0667 | 78.7667 | 79.1 |  | 8.05 | 75.7448 | 79.77 | 80.1333
10.3 |  |  | 33.65 | 78.15 | 78.4571 |  |  | 55.3316 | 77.9667 | 78.125 |  | 3.4 | 70.5 | 77.88 | 78.2286 |  | 5.5 | 73.7 | 78.5 | 78.8333 |  | 7.9 | 75.4349 | 79.6153 | 79.9667
10.4 |  |  | 26.5 | 78 | 78.35 |  |  | 48.5838 | 77.875 | 78 |  | 3.3 | 69.025 | 77.725 | 78.02 |  | 5.4 | 73.28 | 78.2333 | 78.525 |  | 7.8 | 75.085 | 79.4268 | 79.7286
10.5 |  |  | 20.96 | 77.875 | 78.15 |  |  | 41.4862 | 77.7 | 77.8 |  | 3.3 | 67.1778 | 77.64 | 77.8875 |  | 5.3 | 72.88 | 78.0667 | 78.3 |  | 7.6 | 74.6851 | 79.0668 | 79.4857
10.6 |  |  | 16.82 | 77.7 | 78 |  |  | 33.4888 | 77.6 | 77.7 |  |  | 64.3714 | 77.36 | 77.725 |  | 5.25 | 72.38 | 77.8 | 78.125 |  | 7.5 | 74.4051 | 78.8726 | 79.3333
10.7 |  |  | 14.5 | 77.6 | 77.8 |  |  | 27.2909 | 77.5 | 77.6 |  |  | 60 | 77.2667 | 77.5 |  | 5.1 | 71.85 | 77.58 | 77.9 |  | 7.3333 | 74.0602 | 78.7069 | 79.0333
10.8 |  |  | 12.85 | 77.475 | 77.7 |  |  | 21.7927 | 77.25 | 77.425 |  |  | 55 | 77.1 | 77.38 |  | 5 | 71.42 | 77.3833 | 77.6 |  | 7.3 | 73.7452 | 78.5269 | 78.76
10.9 |  |  | 11.8167 | 77.3 | 77.6 |  |  | 17.919 | 77.1 | 77.275 |  |  | 48.56 | 76.96 | 77.2 |  | 5 | 70.78 | 77.2167 | 77.45 |  | 7.1667 | 73.4353 | 78.4536 | 78.525
11 |  |  | 11.05 | 77.2 | 77.38 |  |  | 15.4198 | 77.05 | 77.1 |  |  | 41.16 | 76.825 | 77.1 |  | 4.9 | 70.1333 | 77 | 77.3 |  | 7 | 72.9711 | 78.187 | 78.2857
11.1 |  |  | 10.5 | 77 | 77.3 |  |  | 13.7788 | 76.9 | 76.94 |  |  | 34.6 | 76.6333 | 76.92 |  |  | 69.48 | 76.7667 | 77.05 |  | 6.9 | 72.7055 | 77.9537 | 78.1111
11.2 |  |  | 10.05 | 76.875 | 77.18 |  |  | 12.6291 | 76.7 | 76.9 |  |  | 28.7 | 76.4 | 76.75 |  |  | 68.8 | 76.4857 | 76.825 |  | 6.8 | 72.1713 | 77.7204 | 77.9286
11.3 |  |  | 9.7667 | 76.7 | 77 |  |  | 11.821 | 76.575 | 76.7 |  |  | 23.8714 | 76.2667 | 76.525 |  |  | 67.575 | 76.4 | 76.66 |  | 6.55 | 71.8056 | 77.4871 | 77.7625
11.4 |  |  | 9.4333 | 76.6 | 76.9 |  |  | 11.1713 | 76.425 | 76.6 |  |  | 20.1 | 76.1 | 76.375 |  |  | 66.1667 | 76.2333 | 76.525 |  | 6.5 | 71.4857 | 77.2204 | 77.4571
11.5 |  |  | 9.2 | 76.46 | 76.7 |  |  | 10.7464 | 76.28 | 76.4 |  |  | 17.6167 | 75.95 | 76.2 |  |  | 64.425 | 75.9667 | 76.34 |  | 6.4 | 71.0358 | 76.9372 | 77.22
11.6 |  |  | 9.0333 | 76.3 | 76.55 |  |  | 10.3299 | 76.15 | 76.275 |  |  | 15.88 | 75.8 | 76.1 |  |  | 61.65 | 75.8 | 76.15 |  |  | 70.4859 | 76.6872 | 76.95
11.7 |  |  | 8.8 | 76.125 | 76.4 |  |  | 10.0466 | 75.95 | 76.125 |  |  | 14.8 | 75.675 | 75.9 |  |  | 58 | 75.6429 | 75.96 |  |  | 70.1193 | 76.5872 | 76.8
11.8 |  |  | 8.64 | 76 | 76.26 |  |  | 9.7567 | 75.7333 | 76 |  |  | 14 | 75.425 | 75.7 |  |  | 53.075 | 75.4375 | 75.8 |  |  | 69.4861 | 76.354 | 76.6
11.9 |  |  | 8.5 | 75.875 | 76.2 |  |  | 9.4768 | 75.6 | 75.825 |  |  | 13.32 | 75.225 | 75.6 |  |  | 47.46 | 75.2444 | 75.575 |  |  | 68.9291 | 76.1873 | 76.4375
12 |  |  | 8.3714 | 75.725 | 76.0667 |  |  | 9.2219 | 75.5 | 75.6 |  |  | 12.8 | 75.175 | 75.4 |  |  | 41.7667 | 75.075 | 75.35 |  |  | 68.3863 | 75.9207 | 76.225
12.1 |  |  | 8.25 | 75.6 | 75.875 |  |  | 9.097 | 75.4 | 75.5125 |  |  | 12.36 | 75 | 75.225 |  |  | 35.96 | 74.92 | 75.1889 |  |  | 67.7864 | 75.7541 | 76
12.2 |  |  | 8.1 | 75.4 | 75.7333 |  |  | 8.897 | 75.225 | 75.325 |  |  | 11.9429 | 74.82 | 75.175 |  |  | 30.64 | 74.7429 | 75.06 |  |  | 66.8866 | 75.5874 | 75.8
12.3 |  |  | 8 | 75.28 | 75.54 |  |  | 8.6721 | 75.06 | 75.1667 |  |  | 11.58 | 74.6667 | 75.04 |  |  | 26.56 | 74.5727 | 74.9143 |  |  | 65.8202 | 75.3303 | 75.6286
12.4 |  |  | 7.85 | 75.2 | 75.4 |  |  | 8.5222 | 74.85 | 75.02 |  |  | 11.3286 | 74.42 | 74.7 |  |  | 23.4286 | 74.4143 | 74.7143 |  |  | 64.5621 | 75.1075 | 75.46
12.5 |  |  | 7.7 | 75.1 | 75.3 |  |  | 8.3222 | 74.78 | 74.82 |  |  | 11.1 | 74.2714 | 74.6 |  |  | 20.9 | 74.2556 | 74.4 |  |  | 62.6275 | 74.8125 | 75.1429
12.6 |  |  | 7.66 | 74.98 | 75.1 |  |  | 8.1973 | 74.625 | 74.775 |  |  | 10.8143 | 74.1 | 74.4 |  |  | 19.35 | 74.04 | 74.2667 |  |  | 60.228 | 74.6875 | 75
12.7 |  |  | 7.52 | 74.7286 | 74.9286 |  |  | 8.0373 | 74.4 | 74.65 |  |  | 10.6 | 73.8857 | 74.2 |  |  | 18.28 | 73.8 | 74.1 |  |  | 57.0286 | 74.4376 | 74.85
12.8 |  |  | 7.4 | 74.58 | 74.84 |  |  | 7.9473 | 74.28 | 74.5 |  |  | 10.3571 | 73.7 | 74.0333 |  |  | 17.3429 | 73.7 | 73.9 |  |  | 52.6895 | 74.3276 | 74.6833
12.9 |  |  | 7.3 | 74.4833 | 74.62 |  |  | 7.8374 | 74.1 | 74.275 |  |  | 10.16 | 73.6 | 73.85 |  |  | 16.66 | 73.5 | 73.8 |  |  | 47.9304 | 74.1543 | 74.4333
13 |  |  | 7.2 | 74.2 | 74.5 |  |  | 7.6974 | 73.9571 | 74.05 |  |  | 10 | 73.42 | 73.65 |  |  | 16.1 | 73.26 | 73.5 |  |  | 42.6915 | 73.9591 | 74.25
13.1 |  |  | 7.0857 | 74.0833 | 74.3667 |  |  | 7.5475 | 73.8 | 73.88 |  |  | 9.82 | 73.3 | 73.5143 |  |  | 15.65 | 73.1 | 73.3 |  |  | 37.4782 | 73.7433 | 74.0333
13.2 |  |  | 7 | 73.9 | 74.2 |  |  | 7.4975 | 73.65 | 73.725 |  |  | 9.6 | 73.02 | 73.3143 |  |  | 15.2 | 73 | 73.14 |  |  | 33.0934 | 73.4878 | 73.9
13.3 |  |  | 6.9 | 73.7833 | 74.0333 |  |  | 7.3476 | 73.5 | 73.58 |  |  | 9.4 | 72.9 | 73.1167 |  |  | 14.8333 | 72.7333 | 72.94 |  |  | 29.0567 | 73.3878 | 73.7333
13.4 |  |  | 6.8 | 73.6 | 73.85 |  |  | 7.2226 | 73.25 | 73.4 |  |  | 9.24 | 72.7333 | 72.9333 |  |  | 14.6 | 72.5 | 72.8 |  |  | 25.7948 | 73.2878 | 73.6
13.5 |  |  | 6.7 | 73.4167 | 73.7 |  |  | 7.0976 | 73 | 73.2 |  |  | 9.125 | 72.52 | 72.74 |  |  | 14.3 | 72.3333 | 72.5 |  |  | 23.7096 | 73.1307 | 73.5167
13.6 |  |  | 6.6 | 73.26 | 73.44 |  |  | 6.9977 | 72.9 | 73 |  |  | 8.9 | 72.28 | 72.5857 |  |  | 14 | 72.15 | 72.36 |  |  | 22.0956 | 72.8712 | 73.32
13.7 |  |  | 6.575 | 73.05 | 73.3 |  |  | 6.8977 | 72.7 | 72.8 |  |  | 8.8 | 72.0833 | 72.36 |  |  | 13.6857 | 71.975 | 72.24 |  |  | 20.8958 | 72.6212 | 73.1
13.8 |  |  | 6.425 | 72.82 | 73.1 |  |  | 6.8377 | 72.45 | 72.6333 |  |  | 8.675 | 71.9 | 72.1857 |  |  | 13.4 | 71.7667 | 72.075 |  |  | 19.996 | 72.4879 | 72.85
13.9 |  |  | 6.4 | 72.7 | 72.825 |  |  | 6.6978 | 72.2 | 72.4 |  |  | 8.55 | 71.76 | 72.1 |  |  | 13.2333 | 71.5167 | 71.78 |  |  | 19.2361 | 72.2755 | 72.6714
14 |  |  | 6.3 | 72.44 | 72.625 |  |  | 6.5978 | 72 | 72.15 |  |  | 8.4 | 71.4 | 71.8 |  |  | 13 | 71.35 | 71.575 |  |  | 18.6963 | 72.113 | 72.4429
14.1 |  |  | 6.275 | 72.1667 | 72.3667 |  |  | 6.4978 | 71.8333 | 72 |  |  | 8.3 | 71.2714 | 71.52 |  |  | 12.8 | 71.1 | 71.425 |  |  | 18.2463 | 71.9255 | 72.2
14.2 |  |  | 6.2 | 71.825 | 72.0333 |  |  | 6.3979 | 71.6 | 71.7 |  |  | 8.1 | 71 | 71.4 |  |  | 12.64 | 70.9 | 71.22 |  |  | 17.7765 | 71.6698 | 72.0571
14.3 |  |  | 6.1 | 71.42 | 71.65 |  |  | 6.355 | 71.4 | 71.5 |  |  | 8 | 70.8667 | 71.175 |  |  | 12.4 | 70.6667 | 71.0333 |  |  | 17.3466 | 71.4881 | 71.85
14.4 |  |  | 6.05 | 71.05 | 71.25 |  |  | 6.2579 | 71.15 | 71.22 |  |  | 7.88 | 70.7 | 70.9333 |  |  | 12.14 | 70.3833 | 70.725 |  |  | 16.9966 | 71.3381 | 71.66
14.5 |  |  | 6 | 70.38 | 70.6 |  |  | 6.1979 | 70.78 | 70.85 |  |  | 7.7 | 70.4333 | 70.675 |  |  | 11.92 | 70.3 | 70.5333 |  |  | 16.7538 | 71.0882 | 71.425
14.6 |  |  | 5.9 | 69.65 | 69.7667 |  |  | 6.098 | 70.42 | 70.55 |  |  | 7.6 | 70.16 | 70.4333 |  |  | 11.74 | 70 | 70.34 |  |  | 16.4167 | 70.8168 | 71.275
14.7 |  |  | 5.825 | 68.4667 | 68.7 |  |  | 6.098 | 70 | 70.1333 |  |  | 7.5 | 69.8333 | 70.15 |  |  | 11.6 | 69.76 | 70.1333 |  |  | 16.1856 | 70.5882 | 71.0714
14.8 |  |  | 5.8 | 66.8 | 67.1 |  |  | 5.998 | 69.45 | 69.65 |  |  | 7.4 | 69.7 | 69.9111 |  |  | 11.45 | 69.5857 | 69.9 |  |  | 15.9568 | 70.3883 | 70.8429
14.9 |  |  | 5.7167 | 63.95 | 64.3 |  |  | 5.898 | 68.9 | 69.05 |  |  | 7.3 | 69.375 | 69.6 |  |  | 11.2 | 69.4 | 69.62 |  |  | 15.6969 | 70.1383 | 70.6091
15 |  |  | 5.6 | 59.875 | 60 |  |  | 5.898 | 68.15 | 68.3 |  |  | 7.22 | 69.125 | 69.2429 |  |  | 11.05 | 69.1125 | 69.4 |  |  | 15.4769 | 69.9772 | 70.3286
15.1 |  |  | 5.5833 | 54.35 | 54.55 |  |  | 5.7981 | 67.1 | 67.2667 |  |  | 7.2 | 68.7571 | 69.0286 |  |  | 10.9 | 68.9125 | 69.16 |  |  | 15.197 | 69.8313 | 70.1556
15.2 |  |  |  | 47.2 | 48 |  |  | 5.6981 | 65.74 | 65.975 |  |  | 7.1 | 68.42 | 68.54 |  |  | 10.7 | 68.7333 | 68.9667 |  |  | 14.9827 | 69.617 | 69.9143
15.3 |  |  |  | 40.275 | 41 |  |  | 5.6781 | 63.7 | 63.875 |  |  | 6.95 | 68.1 | 68.26 |  |  | 10.575 | 68.4 | 68.725 |  |  | 14.817 | 69.4598 | 69.68
15.4 |  |  |  | 33.6 | 34.1 |  |  | 5.5981 | 60.7 | 60.7 |  |  | 6.9 | 67.56 | 67.9333 |  |  | 10.425 | 68.1333 | 68.4667 |  |  | 14.6257 | 69.1885 | 69.5143
15.5 |  |  |  | 27.5 | 28.1 |  |  |  | 55.9 | 56.05 |  |  | 6.8 | 67.0857 | 67.3857 |  |  | 10.275 | 67.9857 | 68.275 |  |  | 14.3721 | 69.0135 | 69.2556
15.6 |  |  |  | 22 | 23.2 |  |  |  | 50.24 | 50.76 |  |  | 6.7 | 66.45 | 66.88 |  |  | 10.1 | 67.7 | 68 |  |  | 14.1472 | 68.7314 | 69.0571
15.7 |  |  |  | 18.8 | 19.85 |  |  |  | 43.85 | 44.3 |  |  | 6.6 | 65.85 | 66.14 |  |  | 9.975 | 67.4571 | 67.8 |  |  | 13.9172 | 68.5053 | 68.9143
15.8 |  |  |  | 16.4 | 17.7 |  |  |  | 37.3 | 37.9 |  |  | 6.5 | 64.9 | 65.2 |  |  | 9.84 | 67.25 | 67.5333 |  |  | 13.6973 | 68.3886 | 68.5667
15.9 |  |  |  | 15.1333 | 16.48 |  |  |  | 31 | 31.7 |  |  | 6.48 | 63.775 | 64 |  |  | 9.64 | 66.9 | 67.275 |  |  | 13.4973 | 68.2315 | 68.3833
16 |  |  |  | 14.14 | 15.58 |  |  |  | 25.28 | 26.5 |  |  |  | 62 | 62.2857 |  |  | 9.525 | 66.6333 | 66.95 |  |  | 13.3402 | 67.922 | 68.2167
16.1 |  |  |  | 13.5 | 14.94 |  |  |  | 21.2 | 22.4 |  |  |  | 59.3333 | 59.8 |  |  | 9.4 | 66.4 | 66.64 |  |  | 13.1474 | 67.6887 | 68.04
16.2 |  |  |  | 13.1 | 14.46 |  |  |  | 18.35 | 19.6 |  |  |  | 55.8 | 56.3714 |  |  | 9.2833 | 66.06 | 66.375 |  |  | 12.8974 | 67.4555 | 67.8
16.3 |  |  |  | 12.7 | 14.12 |  |  |  | 16.58 | 17.9 |  |  |  | 50.94 | 51.725 |  |  | 9.1 | 65.8429 | 66.1 |  |  | 12.8307 | 67.2055 | 67.54
16.4 |  |  |  | 12.4286 | 13.86 |  |  |  | 15.4 | 16.7 |  |  |  | 45.48 | 46.15 |  |  | 9.0167 | 65.5333 | 65.8333 |  |  | 12.6775 | 66.9389 | 67.275
16.5 |  |  |  | 12.15 | 13.625 |  |  |  | 14.6 | 15.9 |  |  |  | 40 | 40.45 |  |  | 8.9 | 65.2333 | 65.5 |  |  | 12.5308 | 66.6489 | 67.04
16.6 |  |  |  | 12 | 13.4 |  |  |  | 13.92 | 15.3 |  |  |  | 34 | 34.725 |  |  | 8.8 | 64.94 | 65.2 |  |  | 12.2975 | 66.4722 | 66.8
16.7 |  |  |  | 11.8 | 13.22 |  |  |  | 13.525 | 14.85 |  |  |  | 28.6571 | 29.58 |  |  | 8.7 | 64.5 | 64.75 |  |  | 12.1976 | 66.3056 | 66.56
16.8 |  |  |  | 11.7 | 13.075 |  |  |  | 13.125 | 14.4 |  |  |  | 24.48 | 25.4333 |  |  | 8.58 | 64.05 | 64.4333 |  |  | 12.0176 | 65.989 | 66.2833
16.9 |  |  |  | 11.55 | 12.925 |  |  |  | 12.82 | 14.15 |  |  |  | 21.1875 | 22.3 |  |  | 8.5 | 63.6 | 63.9 |  |  | 11.8309 | 65.7557 | 66.075
17 |  |  |  | 11.425 | 12.8 |  |  |  | 12.525 | 13.9 |  |  |  | 19.18 | 20.4333 |  |  |  | 63.1571 | 63.42 |  |  | 11.5977 | 65.5058 | 65.7667
17.1 |  |  |  | 11.2667 | 12.7 |  |  |  | 12.3 | 13.7 |  |  |  | 17.7 | 19.14 |  |  |  | 62.5667 | 62.9 |  |  | 11.4977 | 65.2225 | 65.5
17.2 |  |  |  | 11.175 | 12.6 |  |  |  | 12.1 | 13.46 |  |  |  | 16.78 | 18.1143 |  |  |  | 61.9 | 62.2 |  |  | 11.3977 | 64.9892 | 65.2333
17.3 |  |  |  | 11.1 | 12.5 |  |  |  | 11.975 | 13.32 |  |  |  | 16.0143 | 17.44 |  |  |  | 60.9625 | 61.3667 |  |  | 11.2977 | 64.7225 | 64.96
17.4 |  |  |  | 10.9833 | 12.3833 |  |  |  | 11.8 | 13.1 |  |  |  | 15.475 | 16.85 |  |  |  | 59.9125 | 60.3333 |  |  | 11.1178 | 64.4036 | 64.75
17.5 |  |  |  | 10.875 | 12.3 |  |  |  | 11.575 | 12.925 |  |  |  | 15 | 16.38 |  |  |  | 58.4 | 58.7 |  |  | 10.9728 | 64.1226 | 64.5
17.6 |  |  |  | 10.725 | 12.1 |  |  |  | 11.425 | 12.8 |  |  |  | 14.65 | 16.0286 |  |  |  | 56.0286 | 56.6 |  |  | 10.7978 | 63.7894 | 64.1667
17.7 |  |  |  | 10.625 | 12 |  |  |  | 11.3 | 12.7 |  |  |  | 14.3 | 15.7 |  |  |  | 53.0857 | 53.5 |  |  | 10.7122 | 63.5561 | 63.85
17.8 |  |  |  | 10.575 | 11.9 |  |  |  | 11.2 | 12.525 |  |  |  | 14.125 | 15.375 |  |  |  | 49.2429 | 49.92 |  |  | 10.5979 | 63.2228 | 63.4571
17.9 |  |  |  | 10.475 | 11.84 |  |  |  | 11 | 12.375 |  |  |  | 13.8 | 15.14 |  |  |  | 44.3857 | 45.2 |  |  | 10.3979 | 62.8895 | 63.18
18 |  |  |  | 10.4 | 11.8 |  |  |  | 10.9 | 12.225 |  |  |  | 13.58 | 14.95 |  |  |  | 39.825 | 40.3 |  |  |  | 62.6562 | 63
18.1 |  |  |  | 10.3 | 11.7 |  |  |  | 10.825 | 12.1 |  |  |  | 13.4 | 14.75 |  |  |  | 34.7714 | 35.7667 |  |  |  | 62.2696 | 62.5667
18.2 |  |  |  | 10.2 | 11.6 |  |  |  | 10.7857 | 12.025 |  |  |  | 13.2 | 14.6 |  |  |  | 30.36 | 31.625 |  |  |  | 61.8497 | 62.2333
18.3 |  |  |  | 10.2 | 11.54 |  |  |  | 10.62 | 11.9 |  |  |  | 12.9857 | 14.4 |  |  |  | 26.6 | 27.95 |  |  |  | 61.4898 | 61.7714
18.4 |  |  |  | 10.1 | 11.4 |  |  |  | 10.5 | 11.8 |  |  |  | 12.8 | 14.1857 |  |  |  | 23.7857 | 25.05 |  |  |  | 60.9756 | 61.375
18.5 |  |  |  | 10 | 11.3 |  |  |  | 10.4 | 11.7 |  |  |  | 12.65 | 14 |  |  |  | 21.9 | 23.3 |  |  |  | 60.5399 | 60.95
18.6 |  |  |  | 9.9 | 11.3 |  |  |  | 10.3 | 11.6 |  |  |  | 12.46 | 13.9 |  |  |  | 20.5286 | 22 |  |  |  | 60.04 | 60.43
18.7 |  |  |  | 9.88 | 11.2 |  |  |  | 10.22 | 11.5 |  |  |  | 12.3 | 13.75 |  |  |  | 19.5667 | 21.025 |  |  |  | 59.2401 | 59.7125
18.8 |  |  |  | 9.75 | 11.1 |  |  |  | 10.2 | 11.4 |  |  |  | 12.2167 | 13.5571 |  |  |  | 18.8 | 20.26 |  |  |  | 58.3737 | 58.8
18.9 |  |  |  | 9.7 | 11.0333 |  |  |  | 10.05 | 11.32 |  |  |  | 12 | 13.4 |  |  |  | 18.325 | 19.66 |  |  |  | 57.2071 | 57.6714
19 |  |  |  | 9.6 | 10.98 |  |  |  | 10 | 11.2333 |  |  |  | 11.9286 | 13.3 |  |  |  | 17.8 | 19.25 |  |  |  | 55.7907 | 56.1667
19.1 |  |  |  | 9.6 | 10.9 |  |  |  | 9.9 | 11.2 |  |  |  | 11.72 | 13.1 |  |  |  | 17.38 | 18.85 |  |  |  | 53.7411 | 54.3
19.2 |  |  |  | 9.45 | 10.8 |  |  |  | 9.8 | 11.1 |  |  |  | 11.58 | 12.95 |  |  |  | 17.1 | 18.44 |  |  |  | 50.8249 | 51.45
19.3 |  |  |  | 9.4 | 10.7333 |  |  |  | 9.7 | 11.0125 |  |  |  | 11.4714 | 12.875 |  |  |  | 16.7375 | 18.1667 |  |  |  | 47.5635 | 48
19.4 |  |  |  | 9.3 | 10.64 |  |  |  | 9.6 | 10.9 |  |  |  | 11.3286 | 12.7 |  |  |  | 16.4429 | 17.92 |  |  |  | 43.3928 | 44.14
19.5 |  |  |  | 9.225 | 10.5429 |  |  |  | 9.55 | 10.8125 |  |  |  | 11.225 | 12.6 |  |  |  | 16.2167 | 17.7 |  |  |  | 39.0135 | 39.9286
19.6 |  |  |  | 9.175 | 10.4 |  |  |  | 9.5 | 10.7857 |  |  |  | 11.08 | 12.5 |  |  |  | 16 | 17.4 |  |  |  | 34.9942 | 36.0714
19.7 |  |  |  | 9.025 | 10.3 |  |  |  | 9.4 | 10.62 |  |  |  | 11 | 12.3667 |  |  |  | 15.78 | 17.25 |  |  |  | 31.3948 | 32.425
19.8 |  |  |  | 8.975 | 10.24 |  |  |  | 9.3429 | 10.6 |  |  |  | 10.86 | 12.2 |  |  |  | 15.6143 | 17.025 |  |  |  | 28.0553 | 29.3
19.9 |  |  |  | 8.825 | 10.14 |  |  |  | 9.2 | 10.5 |  |  |  | 10.8 | 12.1286 |  |  |  | 15.36 | 16.8 |  |  |  | 25.6958 | 26.9
20 |  |  |  | 8.8 | 10.06 |  |  |  | 9.12 | 10.4 |  |  |  | 10.6 | 12 |  |  |  | 15.1833 | 16.6 |  |  |  | 23.896 | 25.32
20.1 |  |  |  | 8.7 | 9.96 |  |  |  | 9.04 | 10.3 |  |  |  | 10.525 | 11.9 |  |  |  | 14.9 | 16.4 |  |  |  | 22.6296 | 24.1
20.2 |  |  |  | 8.6 | 9.85 |  |  |  | 9 | 10.3 |  |  |  | 10.42 | 11.7714 |  |  |  | 14.75 | 16.2 |  |  |  | 21.6297 | 23.2
20.3 |  |  |  |  | 9.8 |  |  |  | 8.9 | 10.1375 |  |  |  | 10.4 | 11.7 |  |  |  | 14.5833 | 16 |  |  |  | 20.8965 | 22.5143
20.4 | [8 empty] | 8.8333 | 10.04 |  |  |  | 10.2 | 11.6 |  |  |  | 14.4286 | 15.8 |  |  |  | 20.2966 | 22
20.5 | [8 empty] | 8.7 | 10 |  |  |  | 10.1 | 11.46 |  |  |  | 14.24 | 15.65 |  |  |  | 19.8467 | 21.5
20.6 | [8 empty] | 8.6333 | 9.9 |  |  |  | 10 | 11.3625 |  |  |  | 14.06 | 15.5 |  |  |  | 19.5168 | 21.1333
20.7 | [9 empty] | 9.7286 |  |  |  | 9.95 | 11.3 |  |  |  | 13.9 | 15.3 |  |  |  | 19.1301 | 20.7
20.8 | [13 empty] | 9.825 | 11.2 |  |  |  | 13.7286 | 15.1667 |  |  |  | 18.7969 | 20.42
20.9 | [13 empty] | 9.7714 | 11.1 |  |  |  | 13.6 | 15 |  |  |  | 18.5969 | 20.1333
21 | [13 empty] | 9.68 | 11 |  |  |  | 13.4333 | 14.82 |  |  |  | 18.3469 | 19.9333
21.1 | [13 empty] | 9.6 | 10.9 |  |  |  | 13.3 | 14.7 |  |  |  | 18.1399 | 19.7143
21.2 | [13 empty] | 9.54 | 10.8 |  |  |  | 13.12 | 14.5 |  |  |  | 17.9137 | 19.4333
21.3 | [13 empty] | 9.4 | 10.7 |  |  |  | 12.9333 | 14.325 |  |  |  | 17.6471 | 19.2714
21.4 | [13 empty] | 9.3 | 10.6 |  |  |  | 12.8778 | 14.22 |  |  |  | 17.4638 | 19
21.5 | [18 empty] | 12.7143 | 14.0571 |  |  |  | 17.2138 | 18.8
21.6 | [18 empty] | 12.5455 | 13.9 |  |  |  | 16.9472 | 18.6667
21.7 | [18 empty] | 12.4286 | 13.8143 |  |  |  | 16.7805 | 18.4167
21.8 | [18 empty] | 12.32 | 13.7 |  |  |  | 16.6139 | 18.2
21.9 | [18 empty] | 12.2 | 13.5 |  |  |  | 16.4473 | 18.0167
22 | [18 empty] | 12.0625 | 13.425 |  |  |  | 16.2806 | 17.85
22.1 | [18 empty] | 11.9167 | 13.3 |  |  |  | 16.114 | 17.6714
22.2 | [18 empty] | 11.8 | 13.2667 |  |  |  | 15.8474 | 17.42
22.3 | [18 empty] | 11.7 | 13.1143 |  |  |  | 15.6807 | 17.3
22.4 | [18 empty] | 11.58 | 13 |  |  |  | 15.526 | 17.1
22.5 | [18 empty] | 11.45 | 12.8 |  |  |  | 15.3641 | 16.9571
22.6 | [18 empty] | 11.32 | 12.7 |  |  |  | 15.1975 | 16.84

## Wedge 30 Longitudinal Profiles
Longitudinal Profiles
Wedge 30
SSD=100cm
Field Size=40x30cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 11.0375
-34.9 |  |  |  |  | 11.1556
-34.8 |  |  |  |  | 11.2
-34.7 |  |  |  |  | 11.3143
-34.6 |  |  |  |  | 11.4
-34.5 |  |  |  |  | 11.4875
-34.4 |  |  |  |  | 11.6
-34.3 |  |  |  |  | 11.65
-34.2 |  |  |  |  | 11.7
-34.1 |  |  |  |  | 11.8125
-34 |  |  |  |  | 11.9429
-33.9 |  |  |  |  | 12.075
-33.8 |  |  |  |  | 12.1125
-33.7 |  |  |  |  | 12.2375
-33.6 |  |  |  |  | 12.3
-33.5 |  |  |  |  | 12.4
-33.4 |  |  |  |  | 12.5222
-33.3 |  |  |  |  | 12.6214
-33.2 |  |  |  |  | 12.6929
-33.1 |  |  |  |  | 12.8
-33 |  |  |  |  | 12.85
-32.9 |  |  |  |  | 12.975
-32.8 |  |  |  |  | 13.1
-32.7 |  |  |  |  | 13.2
-32.6 |  |  |  |  | 13.2714
-32.5 |  |  |  |  | 13.4
-32.4 |  |  |  |  | 13.575
-32.3 |  |  |  | 9.9 | 13.7
-32.2 |  |  |  | 10.0143 | 13.8
-32.1 |  |  |  | 10.1 | 13.9286
-32 |  |  |  | 10.15 | 14
-31.9 |  |  |  | 10.275 | 14.1167
-31.8 |  |  |  | 10.3533 | 14.2625
-31.7 |  |  |  | 10.42 | 14.3875
-31.6 |  |  |  | 10.4867 | 14.48
-31.5 |  |  |  | 10.58 | 14.625
-31.4 |  |  |  | 10.6615 | 14.7571
-31.3 |  |  |  | 10.7625 | 14.8875
-31.2 |  |  |  | 10.8 | 15.0143
-31.1 |  |  |  | 10.9 | 15.15
-31 |  |  |  | 10.9875 | 15.2857
-30.9 |  |  |  | 11.1 | 15.4143
-30.8 |  |  |  | 11.2 | 15.55
-30.7 |  |  |  | 11.3 | 15.6857
-30.6 |  |  |  | 11.4 | 15.8125
-30.5 |  |  |  | 11.46 | 15.9429
-30.4 |  |  |  | 11.6 | 16.075
-30.3 |  |  |  | 11.7 | 16.2125
-30.2 |  |  |  | 11.8 | 16.3429
-30.1 |  |  |  | 11.9 | 16.55
-30 |  |  |  | 12.05 | 16.7
-29.9 |  |  |  | 12.12 | 16.875
-29.8 |  |  |  | 12.22 | 17.0833
-29.7 |  |  |  | 12.32 | 17.1857
-29.6 |  |  | 8.9 | 12.5143 | 17.375
-29.5 |  |  | 8.9583 | 12.6 | 17.5
-29.4 |  |  | 9 | 12.675 | 17.7375
-29.3 |  |  | 9.0727 | 12.8125 | 17.8714
-29.2 |  |  | 9.1467 | 12.9429 | 18.1
-29.1 |  |  | 9.2133 | 13.075 | 18.3375
-29 |  |  | 9.28 | 13.2125 | 18.4714
-28.9 |  |  | 9.3875 | 13.3375 | 18.7
-28.8 |  |  | 9.5071 | 13.4 | 18.875
-28.7 |  |  | 9.5786 | 13.6143 | 19.0625
-28.6 |  |  | 9.6 | 13.7 | 19.3
-28.5 |  |  | 9.6556 | 13.7857 | 19.575
-28.4 |  |  | 9.78 | 14.0143 | 19.7625
-28.3 |  | 8.3286 | 9.9067 | 14.15 | 20
-28.2 |  | 8.4 | 9.9733 | 14.275 | 20.1286
-28.1 |  | 8.5 | 10 | 14.4125 | 20.3875
-28 |  | 8.6 | 10.0857 | 14.5375 | 20.6
-27.9 |  | 8.6 | 10.2 | 14.7429 | 20.85
-27.8 |  | 8.7 | 10.2429 | 14.8889 | 21.1143
-27.7 |  | 8.7375 | 10.3 | 14.9818 | 21.2875
-27.6 |  | 8.8 | 10.46 | 15.2 | 21.6
-27.5 |  | 8.8778 | 10.5125 | 15.35 | 21.85
-27.4 |  | 8.9 | 10.6 | 15.6 | 22.2333
-27.3 | 8.5889 | 9 | 10.675 | 15.7273 | 22.54
-27.2 | 8.6 | 9.0375 | 10.8 | 15.85 | 22.94
-27.1 | 8.7 | 9.1 | 10.9 | 16.1 | 23.24
-27 | 8.7 | 9.2 | 11.02 | 16.26 | 23.75
-26.9 | 8.8 | 9.2429 | 11.1 | 16.4375 | 24.3333
-26.8 | 8.88 | 9.3 | 11.2 | 16.6429 | 25.0857
-26.7 | 8.9 | 9.4 | 11.3 | 16.8 | 25.9125
-26.6 | 9 | 9.4375 | 11.4 | 17 | 27.15
-26.5 | 9 | 9.5 | 11.5375 | 17.225 | 28.5857
-26.4 | 9.06 | 9.6 | 11.6714 | 17.4143 | 30.5875
-26.3 | 9.16 | 9.7 | 11.7375 | 17.675 | 33.0143
-26.2 | 9.2 | 9.7714 | 11.8714 | 17.88 | 36.15
-26.1 | 9.3 | 9.8 | 12 | 18.15 | 39.8857
-26 | 9.36 | 9.9286 | 12.1 | 18.4143 | 43.725
-25.9 | 9.4 | 10 | 12.2 | 18.5875 | 48.0714
-25.8 | 9.5 | 10.1 | 12.3286 | 18.85 | 52.375
-25.7 | 9.5875 | 10.2 | 12.4625 | 19.1143 | 56.6625
-25.6 | 9.6 | 10.2556 | 12.6 | 19.375 | 60.2714
-25.5 | 9.7 | 10.35 | 12.7571 | 19.65 | 63.35
-25.4 | 9.78 | 10.4571 | 12.9 | 19.95 | 66.0875
-25.3 | 9.8 | 10.5 | 13.0286 | 20.25 | 67.9571
-25.2 | 9.96 | 10.6091 | 13.1625 | 20.64 | 69.525
-25.1 | 10 | 10.7 | 13.3 | 21.1 | 70.6
-25 | 10.06 | 10.8 | 13.4375 | 21.5625 | 71.525
-24.9 | 10.16 | 10.8625 | 13.5625 | 22.2143 | 72.2571
-24.8 | 10.26 | 11 | 13.775 | 22.95 | 72.95
-24.7 | 10.3 | 11.0286 | 13.9333 | 23.9125 | 73.45
-24.6 | 10.4 | 11.1625 | 14.0571 | 25.15 | 73.8875
-24.5 | 10.46 | 11.3 | 14.275 | 26.9714 | 74.3
-24.4 | 10.56 | 11.3286 | 14.45 | 29.2 | 74.8286
-24.3 | 10.6 | 11.4625 | 14.6571 | 32.4125 | 75.0667
-24.2 | 10.68 | 11.6 | 14.875 | 35.975 | 75.4
-24.1 | 10.8 | 11.7286 | 15.125 | 40.6 | 75.7857
-24 | 10.83 | 11.8 | 15.3143 | 45.2 | 76.1875
-23.9 | 10.9333 | 11.9 | 15.575 | 50.3714 | 76.5
-23.8 | 11.0364 | 12.0286 | 15.75 | 55.5 | 76.7571
-23.7 | 11.13 | 12.1625 | 15.9571 | 60.475 | 77.0875
-23.6 | 11.23 | 12.3 | 16.175 | 64.7 | 77.4
-23.5 | 11.33 | 12.425 | 16.4429 | 68.2625 | 77.6571
-23.4 | 11.43 | 12.6 | 16.8 | 71.025 | 77.9875
-23.3 | 11.53 | 12.7857 | 17.125 | 73.1 | 78.1875
-23.2 | 11.63 | 12.925 | 17.4375 | 74.6625 | 78.46
-23.1 | 11.73 | 13.1 | 17.8714 | 75.7778 | 78.65
-23 | 11.83 | 13.2857 | 18.475 | 76.6571 | 78.9714
-22.9 | 11.93 | 13.4286 | 19.075 | 77.4286 | 79.22
-22.8 | 12.06 | 13.65 | 19.9 | 78.0625 | 79.4
-22.7 | 12.23 | 13.8714 | 21.05 | 78.6714 | 79.68
-22.6 | 12.36 | 14.0286 | 22.5875 | 79.05 | 79.88
-22.5 | 12.53 | 14.28 | 24.65 | 79.4375 | 80.18
-22.4 | 12.66 | 14.55 | 27.5429 | 79.8286 | 80.48
-22.3 | 12.83 | 14.8714 | 31 | 80.225 | 80.6667
-22.2 | 12.96 | 15.1429 | 36.1 | 80.6333 | 80.96
-22.1 | 13.13 | 15.55 | 41.6 | 80.9667 | 81.16
-22 | 13.26 | 16 | 47.5 | 81.24 | 81.35
-21.9 | 13.46 | 16.5714 | 53.9857 | 81.5333 | 81.5667
-21.8 | 13.66 | 17.4 | 60.575 | 81.8667 | 81.7429
-21.7 | 13.8818 | 18.475 | 66.6 | 82.08 | 81.8857
-21.6 | 14.16 | 20.0875 | 71.3429 | 82.3333 | 82.1286
-21.5 | 14.4889 | 22.2571 | 75.3875 | 82.6667 | 82.4
-21.4 | 14.9364 | 25.275 | 78.6 | 82.94 | 82.575
-21.3 | 15.44 | 29.4667 | 80.525 | 83.175 | 82.8125
-21.2 | 16.2 | 34.9 | 82.0714 | 83.3625 | 82.9857
-21.1 | 17.28 | 41 | 83.2 | 83.6133 | 83.175
-21 | 18.84 | 48.4125 | 84.025 | 83.88 | 83.4
-20.9 | 21.2444 | 56.1714 | 84.6857 | 84.1467 | 83.6375
-20.8 | 24.55 | 63.2 | 85.3375 | 84.4133 | 83.7714
-20.7 | 29.02 | 70.2 | 85.775 | 84.66 | 83.9
-20.6 | 34.88 | 75.9143 | 86.1714 | 84.86 | 84.1364
-20.5 | 41.98 | 80.2 | 86.5625 | 85.0143 | 84.32
-20.4 | 49.96 | 83.0429 | 86.9375 | 85.2 | 84.51
-20.3 | 58.24 | 85.4 | 87.2857 | 85.4385 | 84.62
-20.2 | 66.32 | 86.8 | 87.5714 | 85.6545 | 84.825
-20.1 | 73.5636 | 87.7714 | 87.825 | 85.8429 | 85.0429
-20 | 79.2889 | 88.575 | 88.0857 | 86.0571 | 85.2091
-19.9 | 83.8 | 89.15 | 88.35 | 86.2231 | 85.3556
-19.8 | 86.72 | 89.675 | 88.6375 | 86.3714 | 85.475
-19.7 | 88.72 | 90.0714 | 88.9429 | 86.6182 | 85.6125
-19.6 | 90.075 | 90.4625 | 89.075 | 86.7778 | 85.7857
-19.5 | 90.96 | 90.75 | 89.325 | 87 | 85.975
-19.4 | 91.62 | 91.0143 | 89.5857 | 87.1222 | 86.1125
-19.3 | 92.0625 | 91.275 | 89.775 | 87.275 | 86.2429
-19.2 | 92.4 | 91.55 | 89.9 | 87.4714 | 86.45
-19.1 | 92.76 | 91.7571 | 90.0667 | 87.6 | 86.6111
-19 | 93.0222 | 91.975 | 90.2889 | 87.8375 | 86.7286
-18.9 | 93.2333 | 92.18 | 90.4625 | 88.0429 | 86.925
-18.8 | 93.4222 | 92.25 | 90.6 | 88.2 | 87.1
-18.7 | 93.6333 | 92.4571 | 90.8286 | 88.3286 | 87.24
-18.6 | 93.82 | 92.675 | 91.025 | 88.4625 | 87.42
-18.5 | 93.96 | 92.8222 | 91.1 | 88.5875 | 87.62
-18.4 | 94.1125 | 92.9429 | 91.2455 | 88.7286 | 87.8
-18.3 | 94.26 | 93.1714 | 91.4 | 88.8625 | 87.9
-18.2 | 94.4 | 93.3125 | 91.5667 | 88.9875 | 88.0333
-18.1 | 94.5167 | 93.4 | 91.7 | 89.225 | 88.2714
-18 | 94.6556 | 93.4857 | 91.8 | 89.3571 | 88.4
-17.9 | 94.8 | 93.6125 | 91.9286 | 89.4875 | 88.4571
-17.8 | 94.9125 | 93.7429 | 92.1 | 89.625 | 88.5875
-17.7 | 95 | 93.875 | 92.2714 | 89.7571 | 88.7143
-17.6 | 95.0429 | 94.0125 | 92.4 | 89.8875 | 88.85
-17.5 | 95.2091 | 94.1 | 92.5286 | 90.0143 | 88.9857
-17.4 | 95.3 | 94.175 | 92.6556 | 90.15 | 89.1143
-17.3 | 95.35 | 94.3125 | 92.7857 | 90.3714 | 89.3143
-17.2 | 95.45 | 94.4 | 92.825 | 90.5125 | 89.44
-17.1 | 95.55 | 94.475 | 92.95 | 90.6429 | 89.5167
-17 | 95.65 | 94.6 | 93.0857 | 90.775 | 89.6
-16.9 | 95.75 | 94.675 | 93.1167 | 90.9125 | 89.7
-16.8 | 95.8 | 94.8 | 93.2714 | 91.0429 | 89.8
-16.7 | 95.8833 | 94.8833 | 93.4 | 91.1 | 89.9833
-16.6 | 95.9 | 94.925 | 93.5 | 91.2333 | 90.1125
-16.5 | 95.9909 | 95 | 93.6375 | 91.4 | 90.2
-16.4 | 96.09 | 95.0875 | 93.7 | 91.5 | 90.35
-16.3 | 96.1 | 95.1125 | 93.9 | 91.6 | 90.5125
-16.2 | 96.3 | 95.2429 | 94 | 91.7 | 90.6
-16.1 | 96.3 | 95.375 | 94.0714 | 91.8 | 90.6714
-16 | 96.4 | 95.4125 | 94.1 | 92 | 90.8
-15.9 | 96.4 | 95.5429 | 94.2286 | 92.1 | 90.9375
-15.8 | 96.5 | 95.6 | 94.3 | 92.2 | 91.0714
-15.7 | 96.56 | 95.675 | 94.3875 | 92.3 | 91.2
-15.6 | 96.65 | 95.8 | 94.525 | 92.3833 | 91.3
-15.5 | 96.7 | 95.8429 | 94.6 | 92.48 | 91.32
-15.4 | 96.8 | 95.9 | 94.6875 | 92.5667 | 91.5
-15.3 | 96.9 | 96 | 94.725 | 92.7 | 91.6667
-15.2 | 96.9 | 96.0286 | 94.85 | 92.8 | 91.7286
-15.1 | 96.94 | 96.1 | 94.9 | 92.9167 | 91.8625
-15 | 97 | 96.2 | 95 | 93 | 91.9
-14.9 | 97.04 | 96.225 | 95.05 | 93.1 | 92.025
-14.8 | 97.14 | 96.3 | 95.1857 | 93.2 | 92.1571
-14.7 | 97.2 | 96.3857 | 95.3 | 93.3143 | 92.2875
-14.6 | 97.2333 | 96.5 | 95.3 | 93.4 | 92.3143
-14.5 | 97.32 | 96.55 | 95.375 | 93.5 | 92.45
-14.4 | 97.4 | 96.6 | 95.5 | 93.55 | 92.5857
-14.3 | 97.5 | 96.6125 | 95.5429 | 93.6 | 92.6143
-14.2 | 97.5 | 96.7 | 95.6 | 93.7143 | 92.8
-14.1 | 97.52 | 96.7 | 95.7 | 93.85 | 92.9857
-14 | 97.6 | 96.8 | 95.8 | 93.9857 | 93.0143
-13.9 | 97.7 | 96.8429 | 95.8182 | 94.0143 | 93.15
-13.8 | 97.7 | 96.975 | 95.9067 | 94.15 | 93.2
-13.7 | 97.7 | 97 | 95.9733 | 94.2 | 93.4
-13.6 | 97.8 | 97.0429 | 96 | 94.3 | 93.4429
-13.5 | 97.82 | 97.1 | 96.0364 | 94.4 | 93.5
-13.4 | 97.9 | 97.1 | 96.2062 | 94.44 | 93.6
-13.3 | 97.92 | 97.2 | 96.2687 | 94.6 | 93.6375
-13.2 | 98 | 97.2 | 96.3 | 94.6429 | 93.7714
-13.1 | 98.02 | 97.3 | 96.3 | 94.7 | 93.9
-13 | 98.1 | 97.3375 | 96.4333 | 94.8 | 94
-12.9 | 98.1 | 97.4 | 96.5333 | 94.8375 | 94.1
-12.8 | 98.2 | 97.5 | 96.6 | 94.9714 | 94.1375
-12.7 | 98.25 | 97.5286 | 96.6 | 95 | 94.2
-12.6 | 98.3 | 97.6 | 96.7 | 95.0375 | 94.3
-12.5 | 98.3333 | 97.6875 | 96.7429 | 95.1625 | 94.42
-12.4 | 98.4 | 97.7 | 96.8 | 95.3 | 94.5
-12.3 | 98.4 | 97.7 | 96.9 | 95.4 | 94.6
-12.2 | 98.5 | 97.725 | 96.9286 | 95.4625 | 94.7
-12.1 | 98.5 | 97.8 | 97 | 95.5 | 94.8
-12 | 98.6 | 97.8 | 97 | 95.6 | 94.9
-11.9 | 98.6 | 97.88 | 97.1 | 95.7 | 94.9375
-11.8 | 98.6 | 97.9 | 97.1571 | 95.7 | 95
-11.7 | 98.6 | 98 | 97.2 | 95.7875 | 95.1
-11.6 | 98.64 | 98 | 97.225 | 95.825 | 95.1714
-11.5 | 98.7 | 98.04 | 97.3 | 95.9571 | 95.2889
-11.4 | 98.7 | 98.1 | 97.3857 | 96 | 95.3071
-11.3 | 98.8 | 98.1 | 97.4429 | 96.1 | 95.3786
-11.2 | 98.8 | 98.2 | 97.5 | 96.15 | 95.4636
-11.1 | 98.8 | 98.2 | 97.6 | 96.2857 | 95.54
-11 | 98.8667 | 98.2 | 97.6 | 96.3143 | 95.6167
-10.9 | 98.9 | 98.2875 | 97.6286 | 96.4 | 95.7833
-10.8 | 98.96 | 98.3 | 97.7 | 96.5 | 95.8
-10.7 | 99 | 98.3571 | 97.7 | 96.6 | 95.8667
-10.6 | 99 | 98.4 | 97.8 | 96.7 | 95.9556
-10.5 | 99.06 | 98.5 | 97.8 | 96.7 | 96
-10.4 | 99.1 | 98.5 | 97.8778 | 96.7571 | 96.1
-10.3 | 99.1 | 98.5875 | 97.9 | 96.84 | 96.2
-10.2 | 99.2 | 98.6 | 97.9889 | 96.9333 | 96.275
-10.1 | 99.2 | 98.6 | 98 | 97 | 96.4
-10 | 99.2 | 98.6 | 98.05 | 97 | 96.4286
-9.9 | 99.2 | 98.7 | 98.1 | 97.1 | 96.5
-9.8 | 99.3 | 98.7 | 98.1 | 97.1444 | 96.5
-9.7 | 99.3 | 98.7 | 98.18 | 97.2 | 96.6
-9.6 | 99.3 | 98.8 | 98.25 | 97.2 | 96.6
-9.5 | 99.3 | 98.8 | 98.3 | 97.3 | 96.6875
-9.4 | 99.4 | 98.875 | 98.36 | 97.3714 | 96.7143
-9.3 | 99.4 | 98.9 | 98.4 | 97.4 | 96.85
-9.2 | 99.4 | 98.9 | 98.4 | 97.44 | 96.975
-9.1 | 99.5 | 99 | 98.4 | 97.5067 | 97.01
-9 | 99.5 | 99 | 98.5 | 97.5733 | 97.1125
-8.9 | 99.5 | 99 | 98.5 | 97.6 | 97.2
-8.8 | 99.6 | 99.0125 | 98.575 | 97.6357 | 97.2
-8.7 | 99.6 | 99.1 | 98.6 | 97.7071 | 97.27
-8.6 | 99.6 | 99.1 | 98.6857 | 97.7786 | 97.3
-8.5 | 99.6 | 99.1 | 98.7 | 97.8778 | 97.4286
-8.4 | 99.6909 | 99.1 | 98.7714 | 97.9125 | 97.5
-8.3 | 99.7 | 99.1 | 98.8 | 98 | 97.57
-8.2 | 99.7 | 99.2 | 98.8 | 98.06 | 97.6
-8.1 | 99.7143 | 99.2 | 98.88 | 98.1 | 97.6727
-8 | 99.8 | 99.2625 | 98.9 | 98.1 | 97.7
-7.9 | 99.8 | 99.3 | 98.9 | 98.2 | 97.725
-7.8 | 99.8 | 99.3 | 98.9 | 98.2 | 97.8077
-7.7 | 99.8 | 99.3 | 98.9 | 98.3 | 97.8846
-7.6 | 99.8 | 99.3286 | 99 | 98.3 | 97.9533
-7.5 | 99.8 | 99.4 | 99 | 98.3571 | 98.025
-7.4 | 99.86 | 99.4 | 99 | 98.4 | 98.1
-7.3 | 99.9 | 99.4875 | 99.1 | 98.5 | 98.1125
-7.2 | 99.9 | 99.5 | 99.125 | 98.5571 | 98.22
-7.1 | 99.9 | 99.5 | 99.2 | 98.6 | 98.2867
-7 | 99.9 | 99.5 | 99.2 | 98.7 | 98.3
-6.9 | 100 | 99.6 | 99.3 | 98.7 | 98.3143
-6.8 | 100 | 99.6 | 99.3 | 98.7214 | 98.44
-6.7 | 99.9 | 99.6 | 99.3 | 98.7929 | 98.5
-6.6 | 99.975 | 99.6 | 99.3 | 98.8 | 98.5444
-6.5 | 100 | 99.6 | 99.3 | 98.8571 | 98.6
-6.4 | 100 | 99.6 | 99.4 | 98.9 | 98.7
-6.3 | 100 | 99.6125 | 99.4 | 98.9 | 98.7
-6.2 | 100 | 99.7 | 99.4 | 98.9 | 98.7714
-6.1 | 100 | 99.7 | 99.5 | 98.9267 | 98.8
-6 | 100.05 | 99.7 | 99.5 | 98.9933 | 98.8
-5.9 | 100.1 | 99.775 | 99.5 | 99.06 | 98.8714
-5.8 | 100.1 | 99.8 | 99.5 | 99.125 | 98.9
-5.7 | 100.1 | 99.8 | 99.5 | 99.1875 | 98.9
-5.6 | 100.1 | 99.8 | 99.55 | 99.2 | 98.9625
-5.5 | 100.1 | 99.8 | 99.6 | 99.2 | 99
-5.4 | 100.18 | 99.8 | 99.6 | 99.2 | 99.1
-5.3 | 100.2 | 99.8 | 99.6 | 99.2143 | 99.1
-5.2 | 100.2 | 99.8286 | 99.64 | 99.3 | 99.1
-5.1 | 100.2 | 99.9 | 99.7 | 99.3 | 99.1
-5 | 100.2 | 99.9 | 99.7 | 99.4 | 99.2
-4.9 | 100.2 | 99.9 | 99.7 | 99.4429 | 99.2
-4.8 | 100.2 | 99.9 | 99.7 | 99.5 | 99.3
-4.7 | 100.2 | 99.9857 | 99.8 | 99.5 | 99.3
-4.6 | 100.2333 | 100 | 99.8 | 99.5 | 99.3
-4.5 | 100.3 | 100 | 99.9 | 99.5714 | 99.325
-4.4 | 100.2 | 100 | 99.9 | 99.6 | 99.4
-4.3 | 100.2 | 100 | 99.9 | 99.6 | 99.4
-4.2 | 100.2 | 100 | 99.9 | 99.6 | 99.4
-4.1 | 100.2 | 100 | 99.9 | 99.6 | 99.4
-4 | 100.2 | 100 | 99.9 | 99.68 | 99.5
-3.9 | 100.2 | 100 | 99.9 | 99.7 | 99.54
-3.8 | 100.2 | 100 | 99.9 | 99.7 | 99.6
-3.7 | 100.2 | 100 | 99.9571 | 99.7 | 99.6
-3.6 | 100.2 | 100 | 100 | 99.8 | 99.6
-3.5 | 100.2 | 100 | 100 | 99.8 | 99.6286
-3.4 | 100.2 | 100 | 100 | 99.8 | 99.7
-3.3 | 100.2 | 100 | 100 | 99.8 | 99.7
-3.2 | 100.2 | 100 | 99.93 | 99.8375 | 99.7
-3.1 | 100.2 | 100 | 99.9583 | 99.9 | 99.8
-3 | 100.2 | 100 | 100 | 99.9 | 99.8
-2.9 | 100.2 | 100 | 100 | 99.9 | 99.8
-2.8 | 100.2 | 100 | 100 | 99.9 | 99.8
-2.7 | 100.2 | 100 | 100.0455 | 99.8 | 99.8429
-2.6 | 100.2 | 100 | 100.05 | 99.8 | 99.9
-2.5 | 100.2 | 100 | 100 | 99.8625 | 99.9
-2.4 | 100.2 | 100 | 100 | 99.9 | 99.9375
-2.3 | 100.2 | 100 | 100 | 99.9 | 100
-2.2 | 100.2 | 100 | 100 | 99.9875 | 100
-2.1 | 100.1 | 100 | 100 | 100 | 99.9
-2 | 100.1 | 100 | 100 | 100 | 99.9
-1.9 | 100.1 | 100 | 100 | 100 | 99.9
-1.8 | 100.1 | 100 | 100 | 100 | 99.9714
-1.7 | 100.1 | 100 | 100 | 100 | 100
-1.6 | 100.1 | 100 | 100.1 | 100 | 100
-1.5 | 100.1 | 100 | 100.1 | 100 | 100
-1.4 | 100.1 | 99.9778 | 100.1 | 100 | 100.1
-1.3 | 100.1 | 99.9 | 100.1 | 100 | 100
-1.2 | 100.1 | 99.92 | 100.05 | 100 | 100
-1.1 | 100.1 | 100 | 100.0857 | 100 | 100
-1 | 100.1 | 100 | 100.1 | 100 | 100
-0.9 | 100.1 | 100 | 100.1 | 100 | 100
-0.8 | 100.1 | 100 | 100 | 100 | 99.9
-0.7 | 100.1 | 100 | 100 | 100 | 99.96
-0.6 | 100.1 | 100 | 100 | 100 | 100
-0.5 | 100.1 | 100 | 100 | 100 | 100
-0.4 | 100.0875 | 100 | 100 | 100 | 100
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100 | 100
0.4 | 100.0875 | 100 | 100 | 100 | 100
0.5 | 100.1 | 100 | 100 | 100 | 100
0.6 | 100.1 | 100 | 100 | 100 | 100
0.7 | 100.1 | 100 | 100 | 100 | 99.94
0.8 | 100.1 | 100 | 100.0571 | 100 | 99.9
0.9 | 100.1 | 100 | 100.1 | 100 | 100
1 | 100.1 | 100 | 100.1 | 100 | 100
1.1 | 100.1 | 100 | 100.0143 | 100 | 100
1.2 | 100.1 | 100 | 100.0143 | 100 | 100
1.3 | 100.1 | 99.9 | 100.1 | 100 | 100
1.4 | 100.1 | 100 | 100.1 | 100 | 100
1.5 | 100.1 | 100 | 100.1 | 100 | 100
1.6 | 100.1 | 100 | 100.1 | 100 | 100
1.7 | 100.1 | 100 | 100 | 100 | 100
1.8 | 100.1 | 100 | 100 | 100 | 99.9375
1.9 | 100.1 | 100 | 100 | 100 | 99.9
2 | 100.1 | 100 | 100 | 100 | 99.9
2.1 | 100.1 | 100 | 100 | 100 | 99.9
2.2 | 100.2 | 100 | 100 | 99.9875 | 100
2.3 | 100.2 | 100 | 100 | 99.9 | 100
2.4 | 100.2 | 100 | 100 | 99.9 | 99.9
2.5 | 100.2 | 100 | 100.0111 | 99.8625 | 99.9
2.6 | 100.2 | 100 | 100.08 | 99.8 | 99.9
2.7 | 100.2 | 100 | 100 | 99.8 | 99.8
2.8 | 100.2 | 100 | 100 | 99.9 | 99.8
2.9 | 100.2 | 100 | 100 | 99.9 | 99.8
3 | 100.2 | 100 | 100 | 99.9 | 99.8
3.1 | 100.2 | 100 | 100 | 99.9 | 99.775
3.2 | 100.2 | 100 | 100 | 99.8375 | 99.7
3.3 | 100.2 | 100 | 100 | 99.8 | 99.7
3.4 | 100.2 | 100 | 100 | 99.8 | 99.7
3.5 | 100.2 | 100 | 100 | 99.8 | 99.6
3.6 | 100.2 | 100 | 100 | 99.8 | 99.6
3.7 | 100.2 | 100 | 99.9875 | 99.7 | 99.6
3.8 | 100.2 | 100 | 99.9 | 99.7 | 99.6
3.9 | 100.2 | 100 | 99.9 | 99.7 | 99.5
4 | 100.2 | 100 | 99.9 | 99.68 | 99.5
4.1 | 100.2 | 100 | 99.9 | 99.6 | 99.4571
4.2 | 100.2 | 100 | 99.9 | 99.6 | 99.4
4.3 | 100.2 | 100 | 99.9 | 99.6 | 99.4
4.4 | 100.2 | 100 | 99.9 | 99.6 | 99.4
4.5 | 100.3 | 100 | 99.82 | 99.5714 | 99.3
4.6 | 100.2333 | 100 | 99.8 | 99.5 | 99.3
4.7 | 100.2 | 99.95 | 99.7375 | 99.5 | 99.3
4.8 | 100.2 | 99.9 | 99.7 | 99.5 | 99.2625
4.9 | 100.2 | 99.9 | 99.7 | 99.4429 | 99.2
5 | 100.2 | 99.9 | 99.7 | 99.4 | 99.2
5.1 | 100.2 | 99.9 | 99.7 | 99.3 | 99.1
5.2 | 100.2 | 99.8 | 99.68 | 99.3 | 99.1
5.3 | 100.2 | 99.8 | 99.6 | 99.2143 | 99.1
5.4 | 100.18 | 99.8 | 99.6 | 99.2 | 99.1
5.5 | 100.1 | 99.8 | 99.58 | 99.2 | 99
5.6 | 100.1 | 99.8 | 99.6 | 99.2 | 98.9375
5.7 | 100.1 | 99.8 | 99.5 | 99.1875 | 98.9
5.8 | 100.1 | 99.8 | 99.5 | 99.125 | 98.9
5.9 | 100.1 | 99.7444 | 99.5 | 99.06 | 98.8375
6 | 100.05 | 99.7 | 99.5 | 98.9933 | 98.8
6.1 | 100 | 99.7 | 99.5 | 98.9267 | 98.8
6.2 | 100 | 99.7 | 99.42 | 98.9 | 98.7375
6.3 | 100 | 99.675 | 99.4 | 98.9 | 98.7
6.4 | 100 | 99.6 | 99.4 | 98.9 | 98.675
6.5 | 100 | 99.6 | 99.34 | 98.8571 | 98.6
6.6 | 99.975 | 99.6 | 99.3 | 98.8 | 98.5125
6.7 | 99.9 | 99.6 | 99.3 | 98.7929 | 98.5
6.8 | 100 | 99.6 | 99.3 | 98.7214 | 98.41
6.9 | 100 | 99.5857 | 99.3 | 98.7 | 98.375
7 | 99.9 | 99.5 | 99.225 | 98.7 | 98.3
7.1 | 99.9 | 99.5 | 99.2 | 98.6 | 98.2571
7.2 | 99.9 | 99.5 | 99.1625 | 98.5571 | 98.2
7.3 | 99.9 | 99.45 | 99.0286 | 98.5 | 98.1875
7.4 | 99.86 | 99.4 | 99 | 98.4 | 98.0727
7.5 | 99.8 | 99.4 | 99 | 98.3571 | 97.9867
7.6 | 99.8 | 99.3875 | 99 | 98.3 | 97.92
7.7 | 99.8 | 99.3 | 98.9 | 98.3 | 97.8533
7.8 | 99.8 | 99.3 | 98.9 | 98.2 | 97.8
7.9 | 99.8 | 99.3 | 98.9 | 98.2 | 97.7867
8 | 99.8 | 99.2286 | 98.9 | 98.1 | 97.72
8.1 | 99.7143 | 99.2 | 98.9 | 98.1 | 97.6364
8.2 | 99.7 | 99.1625 | 98.8 | 98.06 | 97.6
8.3 | 99.7 | 99.1 | 98.8 | 98 | 97.6
8.4 | 99.6909 | 99.1 | 98.8 | 97.9125 | 97.525
8.5 | 99.6 | 99.1 | 98.7 | 97.8778 | 97.4
8.6 | 99.6 | 99.1 | 98.7 | 97.7786 | 97.3
8.7 | 99.6 | 99.0125 | 98.6 | 97.7071 | 97.3
8.8 | 99.6 | 99 | 98.6 | 97.6357 | 97.2333
8.9 | 99.5 | 99 | 98.5 | 97.6 | 97.2
9 | 99.5 | 99 | 98.5 | 97.5733 | 97.08
9.1 | 99.5 | 98.98 | 98.4167 | 97.5067 | 97
9.2 | 99.4 | 98.9 | 98.4 | 97.44 | 96.9429
9.3 | 99.4 | 98.9 | 98.4 | 97.4 | 96.8125
9.4 | 99.4 | 98.8375 | 98.4 | 97.3714 | 96.7875
9.5 | 99.3 | 98.8 | 98.3 | 97.3 | 96.6571
9.6 | 99.3 | 98.775 | 98.2833 | 97.2 | 96.6
9.7 | 99.3 | 98.7 | 98.2 | 97.2 | 96.5875
9.8 | 99.3 | 98.7 | 98.1 | 97.1444 | 96.5
9.9 | 99.2 | 98.6857 | 98.1 | 97.1 | 96.5
10 | 99.2 | 98.6 | 98.0857 | 97 | 96.4
10.1 | 99.2 | 98.6 | 98 | 97 | 96.4
10.2 | 99.2 | 98.6 | 98 | 96.9333 | 96.3375
10.3 | 99.1 | 98.55 | 97.92 | 96.84 | 96.2125
10.4 | 99.1 | 98.5 | 97.9 | 96.7571 | 96.1
10.5 | 99.06 | 98.4875 | 97.8 | 96.7 | 96.0429
10.6 | 99 | 98.4 | 97.8 | 96.7 | 96
10.7 | 99 | 98.4 | 97.7286 | 96.6 | 95.9385
10.8 | 98.96 | 98.3 | 97.7 | 96.5 | 95.8583
10.9 | 98.9 | 98.2625 | 97.6667 | 96.4 | 95.7667
11 | 98.8667 | 98.2 | 97.6 | 96.3143 | 95.6467
11.1 | 98.8 | 98.2 | 97.5375 | 96.2857 | 95.5133
11.2 | 98.8 | 98.1625 | 97.5 | 96.15 | 95.4308
11.3 | 98.8 | 98.1 | 97.4 | 96.1 | 95.4
11.4 | 98.7 | 98.1 | 97.4 | 96 | 95.3636
11.5 | 98.7 | 98 | 97.3 | 95.9571 | 95.2625
11.6 | 98.64 | 98 | 97.2571 | 95.825 | 95.2
11.7 | 98.6 | 97.98 | 97.2 | 95.7875 | 95.1125
11.8 | 98.6 | 97.9 | 97.1875 | 95.7 | 95.06
11.9 | 98.6 | 97.9 | 97.1 | 95.7 | 94.9125
12 | 98.6 | 97.8 | 97.025 | 95.6 | 94.8833
12.1 | 98.5 | 97.8 | 97 | 95.5 | 94.8
12.2 | 98.5 | 97.7 | 96.9 | 95.4625 | 94.7
12.3 | 98.4 | 97.7 | 96.9 | 95.4 | 94.64
12.4 | 98.4 | 97.7 | 96.8125 | 95.3 | 94.5571
12.5 | 98.3333 | 97.6571 | 96.775 | 95.1625 | 94.4
12.6 | 98.3 | 97.6 | 96.7 | 95.0375 | 94.2833
12.7 | 98.25 | 97.5 | 96.6125 | 95 | 94.2
12.8 | 98.2 | 97.4571 | 96.5357 | 94.9714 | 94.1
12.9 | 98.1 | 97.4 | 96.4667 | 94.8375 | 94.0667
13 | 98.1 | 97.3 | 96.4 | 94.8 | 94
13.1 | 98.02 | 97.2625 | 96.3333 | 94.7 | 93.9
13.2 | 98 | 97.2 | 96.2687 | 94.6429 | 93.775
13.3 | 97.92 | 97.2 | 96.2062 | 94.6 | 93.7
13.4 | 97.9 | 97.1 | 96.14 | 94.44 | 93.575
13.5 | 97.82 | 97.1 | 96.06 | 94.4 | 93.5
13.6 | 97.8 | 97.0125 | 96 | 94.3 | 93.4125
13.7 | 97.7 | 97 | 96 | 94.2 | 93.3857
13.8 | 97.7 | 96.9429 | 95.94 | 94.15 | 93.25
13.9 | 97.7 | 96.9 | 95.8556 | 94.0143 | 93.1143
14 | 97.6 | 96.8 | 95.75 | 93.9857 | 93
14.1 | 97.52 | 96.75 | 95.7 | 93.85 | 92.95
14.2 | 97.5 | 96.7 | 95.6 | 93.7143 | 92.8143
14.3 | 97.5 | 96.6 | 95.575 | 93.6 | 92.6857
14.4 | 97.4 | 96.6 | 95.5 | 93.55 | 92.55
14.5 | 97.32 | 96.5143 | 95.4125 | 93.5 | 92.5
14.6 | 97.2333 | 96.4857 | 95.3 | 93.4 | 92.3857
14.7 | 97.2 | 96.35 | 95.2429 | 93.3143 | 92.25
14.8 | 97.14 | 96.3 | 95.1125 | 93.2 | 92.125
14.9 | 97.04 | 96.2 | 95.0857 | 93.1 | 92
15 | 97 | 96.1625 | 95 | 93 | 91.9625
15.1 | 96.94 | 96.1 | 94.925 | 92.9167 | 91.8286
15.2 | 96.9 | 96 | 94.8 | 92.8 | 91.7
15.3 | 96.9 | 95.9625 | 94.75 | 92.7 | 91.6333
15.4 | 96.8 | 95.9 | 94.625 | 92.5667 | 91.4667
15.5 | 96.7 | 95.8125 | 94.5875 | 92.48 | 91.4
15.6 | 96.65 | 95.775 | 94.5 | 92.3833 | 91.3
15.7 | 96.56 | 95.7 | 94.425 | 92.3 | 91.2
15.8 | 96.5 | 95.63 | 94.3 | 92.2 | 91.1375
15.9 | 96.4 | 95.5125 | 94.2625 | 92.1 | 91
16 | 96.4 | 95.475 | 94.1375 | 92 | 90.7714
16.1 | 96.3 | 95.3429 | 94 | 91.8 | 90.6375
16.2 | 96.3 | 95.2125 | 93.9714 | 91.7 | 90.6
16.3 | 96.1 | 95.1 | 93.8375 | 91.6 | 90.475
16.4 | 96.09 | 95.1 | 93.7 | 91.5 | 90.3429
16.5 | 95.9909 | 95.025 | 93.6 | 91.4 | 90.2125
16.6 | 95.9 | 94.9875 | 93.5375 | 91.2333 | 90.075
16.7 | 95.8833 | 94.85 | 93.4 | 91.1 | 90
16.8 | 95.8 | 94.7857 | 93.2286 | 91.0429 | 89.86
16.9 | 95.75 | 94.7 | 93.1571 | 90.9125 | 89.76
17 | 95.65 | 94.6125 | 93.0143 | 90.775 | 89.6
17.1 | 95.55 | 94.5 | 92.9857 | 90.6429 | 89.5714
17.2 | 95.45 | 94.4125 | 92.85 | 90.5125 | 89.4
17.3 | 95.35 | 94.275 | 92.725 | 90.3714 | 89.3167
17.4 | 95.3 | 94.2 | 92.6875 | 90.15 | 89.0857
17.5 | 95.2091 | 94.1125 | 92.5625 | 90.0143 | 88.95
17.6 | 95.0429 | 94.075 | 92.375 | 89.8875 | 88.8143
17.7 | 95 | 93.8857 | 92.2125 | 89.7571 | 88.7
17.8 | 94.9125 | 93.7125 | 92.0857 | 89.625 | 88.6143
17.9 | 94.8 | 93.5857 | 91.95 | 89.4875 | 88.425
18 | 94.6556 | 93.5 | 91.8333 | 89.3571 | 88.3875
18.1 | 94.5167 | 93.4143 | 91.7222 | 89.225 | 88.2571
18.2 | 94.4 | 93.2857 | 91.5545 | 88.9875 | 88.1
18.3 | 94.26 | 93.15 | 91.4222 | 88.8625 | 87.95
18.4 | 94.1125 | 93.0143 | 91.2 | 88.7286 | 87.7714
18.5 | 93.96 | 92.875 | 91.1167 | 88.5875 | 87.7
18.6 | 93.82 | 92.6571 | 91 | 88.4625 | 87.4
18.7 | 93.6333 | 92.45 | 90.825 | 88.3286 | 87.3
18.8 | 93.4222 | 92.2875 | 90.6286 | 88.2 | 87.0625
18.9 | 93.2333 | 92.14 | 90.5 | 88.0429 | 86.9286
19 | 93.0222 | 91.9667 | 90.27 | 87.8375 | 86.7
19.1 | 92.76 | 91.75 | 90.125 | 87.6 | 86.6714
19.2 | 92.4 | 91.4875 | 89.9375 | 87.4714 | 86.475
19.3 | 92.0625 | 91.3143 | 89.725 | 87.275 | 86.225
19.4 | 91.62 | 91.05 | 89.575 | 87.1222 | 86.175
19.5 | 90.96 | 90.7625 | 89.3857 | 87 | 85.9857
19.6 | 90.075 | 90.4143 | 89.125 | 86.7778 | 85.8125
19.7 | 88.72 | 90.075 | 88.95 | 86.6182 | 85.65
19.8 | 86.72 | 89.6625 | 88.6286 | 86.3714 | 85.4429
19.9 | 83.8 | 89.1857 | 88.4125 | 86.2231 | 85.3222
20 | 79.2889 | 88.525 | 88.125 | 86.0571 | 85.2273
20.1 | 73.5636 | 87.8 | 87.8 | 85.8429 | 85.025
20.2 | 66.32 | 86.7625 | 87.5286 | 85.6545 | 84.7778
20.3 | 58.24 | 85.2714 | 87.2714 | 85.4385 | 84.6727
20.4 | 49.96 | 83.2 | 86.95 | 85.2 | 84.5556
20.5 | 41.98 | 80.2286 | 86.575 | 85.0143 | 84.35
20.6 | 34.88 | 75.8625 | 86.25 | 84.86 | 84.175
20.7 | 29.02 | 70.2 | 85.7714 | 84.66 | 83.9625
20.8 | 24.55 | 63.4429 | 85.3 | 84.4133 | 83.775
20.9 | 21.2444 | 56.1125 | 84.725 | 84.1467 | 83.6
21 | 18.84 | 48.2 | 83.9857 | 83.88 | 83.425
21.1 | 17.28 | 41.2 | 83.175 | 83.6133 | 83.1429
21.2 | 16.2 | 34.7625 | 82.0625 | 83.3625 | 83.0125
21.3 | 15.44 | 29.675 | 80.5375 | 83.175 | 82.775
21.4 | 14.9364 | 25.2571 | 78.3857 | 82.94 | 82.65
21.5 | 14.4889 | 22.2625 | 75.6 | 82.6667 | 82.4286
21.6 | 14.16 | 20 | 71.3875 | 82.3333 | 82.1714
21.7 | 13.8818 | 18.4714 | 66.2714 | 82.08 | 81.9
21.8 | 13.66 | 17.3875 | 60.9 | 81.8667 | 81.7143
21.9 | 13.46 | 16.575 | 54.025 | 81.5333 | 81.5333
22 | 13.26 | 16 | 47.775 | 81.24 | 81.3714
22.1 | 13.13 | 15.5429 | 41.4 | 80.9667 | 81.1333
22.2 | 12.96 | 15.1571 | 35.7375 | 80.6333 | 80.9667
22.3 | 12.83 | 14.8 | 31.2375 | 80.225 | 80.68
22.4 | 12.66 | 14.5286 | 27.4 | 79.8286 | 80.44
22.5 | 12.53 | 14.24 | 24.6857 | 79.4375 | 80.14
22.6 | 12.36 | 14.0714 | 22.525 | 79.05 | 79.84
22.7 | 12.23 | 13.85 | 20.9 | 78.6714 | 79.64
22.8 | 12.06 | 13.6286 | 19.925 | 78.0625 | 79.48
22.9 | 11.93 | 13.3857 | 19.0571 | 77.4286 | 79.2
23 | 11.83 | 13.25 | 18.45 | 76.6571 | 78.975
23.1 | 11.73 | 13.1143 | 17.9 | 75.7778 | 78.675
23.2 | 11.63 | 12.8857 | 17.5 | 74.6625 | 78.45
23.3 | 11.53 | 12.75 | 17.1429 | 73.1 | 78.2333
23.4 | 11.43 | 12.625 | 16.7714 | 71.025 | 77.95
23.5 | 11.33 | 12.4875 | 16.5 | 68.2625 | 77.6
23.6 | 11.23 | 12.3625 | 16.25 | 64.7 | 77.3875
23.7 | 11.13 | 12.2286 | 15.975 | 60.475 | 77.0571
23.8 | 11.0364 | 12.1 | 15.7571 | 55.5 | 76.8
23.9 | 10.9333 | 11.9625 | 15.5286 | 50.3714 | 76.4625
24 | 10.83 | 11.8286 | 15.2875 | 45.2 | 76.1857
24.1 | 10.8 | 11.7 | 15.1143 | 40.6 | 75.8
24.2 | 10.68 | 11.5625 | 14.85 | 35.975 | 75.3875
24.3 | 10.6 | 11.4286 | 14.675 | 32.4125 | 75.075
24.4 | 10.56 | 11.4 | 14.4571 | 29.2 | 74.8
24.5 | 10.46 | 11.2625 | 14.25 | 26.9714 | 74.3143
24.6 | 10.4 | 11.1286 | 14.0875 | 25.15 | 73.9125
24.7 | 10.3 | 11.1 | 13.925 | 23.9125 | 73.4
24.8 | 10.26 | 10.9625 | 13.7286 | 22.95 | 72.8571
24.9 | 10.16 | 10.9 | 13.5875 | 22.2143 | 72.25
25 | 10.06 | 10.8 | 13.4625 | 21.5625 | 71.575
25.1 | 10 | 10.6667 | 13.3286 | 21.1 | 70.5714
25.2 | 9.96 | 10.5636 | 13.1 | 20.64 | 69.4375
25.3 | 9.8 | 10.5 | 12.9625 | 20.25 | 68.05
25.4 | 9.78 | 10.425 | 12.8286 | 19.95 | 66.025
25.5 | 9.7 | 10.4 | 12.7 | 19.65 | 63.4
25.6 | 9.6 | 10.2364 | 12.6333 | 19.375 | 60.3125
25.7 | 9.5875 | 10.1875 | 12.5 | 19.1143 | 56.525
25.8 | 9.5 | 10.1 | 12.3 | 18.85 | 52.6
25.9 | 9.4 | 10.0375 | 12.2286 | 18.5875 | 48.1714
26 | 9.36 | 9.9 | 12.1 | 18.4143 | 43.6
26.1 | 9.3 | 9.8625 | 12.0286 | 18.15 | 39.75
26.2 | 9.2 | 9.7375 | 11.9 | 17.88 | 36.0714
26.3 | 9.16 | 9.7 | 11.7714 | 17.675 | 33.1143
26.4 | 9.06 | 9.6 | 11.7 | 17.4143 | 30.5
26.5 | 9 | 9.5375 | 11.5 | 17.225 | 28.625
26.6 | 9 | 9.5 | 11.4 | 17 | 27.0875
26.7 | 8.9 | 9.4 | 11.3 | 16.8 | 25.9571
26.8 | 8.88 | 9.3375 | 11.22 | 16.6429 | 25.05
26.9 | 8.8 | 9.3 | 11.1429 | 16.4375 | 24.34
27 | 8.7 | 9.1714 | 11 | 16.26 | 23.74
27.1 | 8.7 | 9.1 | 10.9 | 16.1 | 23.2667
27.2 | 8.6 | 9 | 10.76 | 15.85 | 22.9667
27.3 | 8.5889 | 8.975 | 10.7 | 15.7273 | 22.6
27.4 |  | 8.9 | 10.6 | 15.6 | 22.2
27.5 |  | 8.8455 | 10.5429 | 15.35 | 21.9
27.6 |  | 8.8 | 10.4 | 15.2 | 21.5857
27.7 |  | 8.7 | 10.3167 | 14.9818 | 21.325
27.8 |  | 8.6625 | 10.275 | 14.8889 | 21.1286
27.9 |  | 8.6 | 10.15 | 14.7429 | 20.9
28 |  | 8.6 | 10.1 | 14.5375 | 20.625
28.1 |  | 8.5 | 9.9846 | 14.4125 | 20.3571
28.2 |  | 8.44 | 9.9077 | 14.275 | 20.1
28.3 |  | 8.3 | 9.84 | 14.15 | 19.9625
28.4 |  |  | 9.7714 | 14.0143 | 19.7571
28.5 |  |  | 9.7 | 13.7857 | 19.6
28.6 |  |  | 9.5857 | 13.7 | 19.3429
28.7 |  |  | 9.5143 | 13.6143 | 19.075
28.8 |  |  | 9.4467 | 13.4 | 18.9
28.9 |  |  | 9.38 | 13.3375 | 18.6714
29 |  |  | 9.3133 | 13.2125 | 18.475
29.1 |  |  | 9.2467 | 13.075 | 18.3
29.2 |  |  | 9.18 | 12.9429 | 18.1429
29.3 |  |  | 9.1133 | 12.8125 | 17.875
29.4 |  |  | 9 | 12.675 | 17.7
29.5 |  |  | 8.9 | 12.6 | 17.55
29.6 |  |  | 8.9 | 12.5143 | 17.3429
29.7 |  |  |  | 12.32 | 17.2125
29.8 |  |  |  | 12.22 | 17.05
29.9 |  |  |  | 12.12 | 16.9
30 |  |  |  | 12.05 | 16.75
30.1 |  |  |  | 11.9 | 16.5429
30.2 |  |  |  | 11.8 | 16.4125
30.3 |  |  |  | 11.7 | 16.275
30.4 |  |  |  | 11.6 | 16.1429
30.5 |  |  |  | 11.46 | 15.925
30.6 |  |  |  | 11.4 | 15.8
30.7 |  |  |  | 11.3 | 15.7
30.8 |  |  |  | 11.2 | 15.5143
30.9 |  |  |  | 11.1 | 15.3857
31 |  |  |  | 10.9875 | 15.25
31.1 |  |  |  | 10.9 | 15.1143
31.2 |  |  |  | 10.8 | 15.0875
31.3 |  |  |  | 10.7625 | 14.9143
31.4 |  |  |  | 10.6615 | 14.725
31.5 |  |  |  | 10.58 | 14.5875
31.6 |  |  |  | 10.4867 | 14.44
31.7 |  |  |  | 10.42 | 14.3571
31.8 |  |  |  | 10.3533 | 14.3
31.9 |  |  |  | 10.275 | 14.0833
32 |  |  |  | 10.15 | 14
32.1 |  |  |  | 10.1 | 13.9
32.2 |  |  |  | 10.0143 | 13.7714
32.3 |  |  |  | 9.9 | 13.7
32.4 |  |  |  |  | 13.6
32.5 |  |  |  |  | 13.4714
32.6 |  |  |  |  | 13.275
32.7 |  |  |  |  | 13.2
32.8 |  |  |  |  | 13.1
32.9 |  |  |  |  | 13.0429
33 |  |  |  |  | 12.9125
33.1 |  |  |  |  | 12.7857
33.2 |  |  |  |  | 12.6667
33.3 |  |  |  |  | 12.5833
33.4 |  |  |  |  | 12.5
33.5 |  |  |  |  | 12.4
33.6 |  |  |  |  | 12.3375
33.7 |  |  |  |  | 12.2
33.8 |  |  |  |  | 12.175
33.9 |  |  |  |  | 12.0429
34 |  |  |  |  | 11.9125
34.1 |  |  |  |  | 11.8
34.2 |  |  |  |  | 11.75
34.3 |  |  |  |  | 11.7
34.4 |  |  |  |  | 11.5857
34.5 |  |  |  |  | 11.5
34.6 |  |  |  |  | 11.425
34.7 |  |  |  |  | 11.2857
34.8 |  |  |  |  | 11.2
34.9 |  |  |  |  | 11.1222
35 |  |  |  |  | 11.0111

## Wedge 30 Output Factors
Output factors
Wedge 30 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20 | 30
| 3 | 0.888 | 0.9 | 0.908 | 0.915 | 0.922 | 0.928 | 0.932 | 0.937
| 4 | 0.9 | 0.916 | 0.924 | 0.934 | 0.942 | 0.949 | 0.956 | 0.962
| 5 | 0.908 | 0.925 | 0.934 | 0.947 | 0.956 | 0.966 | 0.974 | 0.983
| 7 | 0.919 | 0.938 | 0.951 | 0.965 | 0.978 | 0.992 | 1.001 | 1.014
| 10 | 0.929 | 0.951 | 0.965 | 0.983 | 1 | 1.018 | 1.032 | 1.048
| 15 | 0.937 | 0.961 | 0.977 | 1 | 1.022 | 1.046 | 1.066 | 1.089
| 20 | 0.942 | 0.968 | 0.986 | 1.011 | 1.036 | 1.066 | 1.089 | 1.119
| 30 | 0.95 | 0.977 | 0.996 | 1.026 | 1.056 | 1.094 | 1.123 | 1.162
| 40 | 0.955 | 0.983 | 1.004 | 1.036 | 1.07 | 1.113 | 1.146 | 1.189

## Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 31.7 | 37.7 | 46.8 | 53
0.1 | 35.4667 | 41.2333 | 50.32 | 56.48
0.2 | 41.4 | 47 | 56 | 61.54
0.3 | 51.7 | 56.5333 | 64.1667 | 68.9
0.4 | 60.275 | 65.1 | 71.18 | 75.86
0.5 | 68.34 | 72.2 | 77.52 | 81.44
0.6 | 74.475 | 77.76 | 82.44 | 85.54
0.7 | 79.78 | 82.775 | 86.52 | 88.92
0.8 | 84.225 | 86.42 | 89.6 | 91.7
0.9 | 87.6 | 89.5 | 92.22 | 93.8
1 | 90.25 | 91.9 | 94.1333 | 95.6167
1.1 | 92.575 | 94 | 95.86 | 96.8
1.2 | 94.5 | 95.55 | 97 | 97.8
1.3 | 96 | 96.9 | 98 | 98.6
1.4 | 97.2333 | 97.825 | 98.68 | 99.2
1.5 | 98.04 | 98.72 | 99.3 | 99.6
1.6 | 98.76 | 99.35 | 99.6 | 99.9
1.7 | 99.36 | 99.7 | 99.8 | 100
1.8 | 99.6 | 99.9 | 100 | 100
1.9 | 99.925 | 100 | 100 | 100
2 | 100 | 100 | 100 | 100
2.1 | 99.98 | 100 | 99.9143 | 99.8
2.2 | 99.8 | 99.9 | 99.8 | 99.675
2.3 | 99.68 | 99.8 | 99.5571 | 99.35
2.4 | 99.4667 | 99.625 | 99.36 | 99.12
2.5 | 99.24 | 99.3 | 99.04 | 98.875
2.6 | 98.85 | 98.9286 | 98.75 | 98.525
2.7 | 98.46 | 98.64 | 98.45 | 98.2
2.8 | 98.08 | 98.38 | 98.04 | 97.875
2.9 | 97.7 | 97.92 | 97.82 | 97.46
3 | 97.36 | 97.6714 | 97.4 | 97.3
3.1 | 96.85 | 97.34 | 97.02 | 96.88
3.2 | 96.58 | 97 | 96.7 | 96.4667
3.3 | 96.1 | 96.4714 | 96.275 | 96.125
3.4 | 95.7 | 96.12 | 95.92 | 95.76
3.5 | 95.25 | 95.675 | 95.65 | 95.46
3.6 | 94.8 | 95.3 | 95.275 | 95.075
3.7 | 94.35 | 94.94 | 94.8667 | 94.7667
3.8 | 93.9 | 94.475 | 94.52 | 94.4
3.9 | 93.2714 | 94.2 | 94.1333 | 94.08
4 | 92.9286 | 93.7667 | 93.8667 | 93.7
4.1 | 92.38 | 93.4143 | 93.5 | 93.3
4.2 | 92 | 93 | 93.08 | 92.94
4.3 | 91.48 | 92.5333 | 92.725 | 92.6
4.4 | 91.2286 | 92.24 | 92.48 | 92.2
4.5 | 90.8 | 91.85 | 92.025 | 91.96
4.6 | 90.2571 | 91.4375 | 91.58 | 91.5
4.7 | 89.9 | 91.1 | 91.3 | 91.2
4.8 | 89.2714 | 90.6429 | 91.04 | 90.9
4.9 | 88.88 | 90.3 | 90.62 | 90.56
5 | 88.35 | 89.9 | 90.22 | 90.22
5.1 | 87.95 | 89.6 | 89.8714 | 89.75
5.2 | 87.575 | 89.15 | 89.55 | 89.425
5.3 | 87.06 | 88.7167 | 89.12 | 89.2
5.4 | 86.625 | 88.38 | 88.78 | 88.8
5.5 | 86.24 | 87.96 | 88.52 | 88.46
5.6 | 85.84 | 87.62 | 88.1571 | 88.1
5.7 | 85.34 | 87.14 | 87.75 | 87.7
5.8 | 84.88 | 86.76 | 87.5143 | 87.48
5.9 | 84.625 | 86.35 | 87.06 | 87.075
6 | 84.15 | 85.98 | 86.75 | 86.82
6.1 | 83.64 | 85.76 | 86.34 | 86.4333
6.2 | 83.2 | 85.2 | 85.9857 | 86.1
6.3 | 82.8375 | 84.86 | 85.65 | 85.7
6.4 | 82.44 | 84.5667 | 85.3714 | 85.35
6.5 | 81.9429 | 84.18 | 85 | 85
6.6 | 81.52 | 83.78 | 84.6714 | 84.6571
6.7 | 81.1 | 83.48 | 84.4 | 84.35
6.8 | 80.58 | 83.1 | 83.9571 | 83.92
6.9 | 80.3286 | 82.76 | 83.6 | 83.58
7 | 79.7667 | 82.35 | 83.3 | 83.26
7.1 | 79.44 | 81.9 | 82.9 | 82.9667
7.2 | 79.04 | 81.375 | 82.5143 | 82.64
7.3 | 78.525 | 81.12 | 82.15 | 82.28
7.4 | 78.06 | 80.82 | 81.8667 | 81.94
7.5 | 77.9 | 80.36 | 81.58 | 81.6429
7.6 | 77.4 | 80.02 | 81.32 | 81.4
7.7 | 77.04 | 79.68 | 80.85 | 80.8778
7.8 | 76.5333 | 79.3333 | 80.5 | 80.62
7.9 | 76.2 | 78.95 | 80.2 | 80.28
8 | 75.7571 | 78.525 | 79.825 | 79.9429
8.1 | 75.34 | 78.16 | 79.5667 | 79.68
8.2 | 74.9778 | 77.86 | 79.14 | 79.3
8.3 | 74.65 | 77.4833 | 78.9286 | 79.06
8.4 | 74.2 | 77.3 | 78.55 | 78.7714
8.5 | 73.78 | 76.8 | 78.3143 | 78.4286
8.6 | 73.3714 | 76.3 | 77.9 | 78.0889
8.7 | 73.0667 | 76.06 | 77.4857 | 77.775
8.8 | 72.6714 | 75.725 | 77.2 | 77.4571
8.9 | 72.3286 | 75.4 | 76.9333 | 77.25
9 | 71.88 | 75.12 | 76.675 | 76.8
9.1 | 71.4444 | 74.75 | 76.3 | 76.525
9.2 | 71.2 | 74.4 | 76 | 76.2667
9.3 | 70.6714 | 74.04 | 75.7 | 75.92
9.4 | 70.36 | 73.66 | 75.3714 | 75.5667
9.5 | 70 | 73.32 | 75.02 | 75.225
9.6 | 69.64 | 72.96 | 74.7 | 75
9.7 | 69.3 | 72.58 | 74.375 | 74.6667
9.8 | 68.84 | 72.3333 | 74.08 | 74.3571
9.9 | 68.6 | 72 | 73.82 | 74.025
10 | 68.22 | 71.7 | 73.48 | 73.7
10.1 | 67.8 | 71.36 | 73.125 | 73.3667
10.2 | 67.5143 | 71.05 | 72.8429 | 73.1
10.3 | 67.12 | 70.64 | 72.56 | 72.8571
10.4 | 66.68 | 70.4 | 72.2286 | 72.5429
10.5 | 66.3333 | 70.0667 | 71.88 | 72.28
10.6 | 66.04 | 69.7667 | 71.6 | 71.9
10.7 | 65.62 | 69.4 | 71.28 | 71.68
10.8 | 65.38 | 69.0667 | 70.96 | 71.375
10.9 | 65.02 | 68.7 | 70.6714 | 71.12
11 | 64.6429 | 68.475 | 70.4143 | 70.7143
11.1 | 64.26 | 68.1 | 70.1 | 70.46
11.2 | 63.9 | 67.825 | 69.8167 | 70.1143
11.3 | 63.5778 | 67.4333 | 69.54 | 69.85
11.4 | 63.3429 | 67.1 | 69.28 | 69.55
11.5 | 62.9 | 66.82 | 69.0333 | 69.3
11.6 | 62.5667 | 66.5 | 68.75 | 69.05
11.7 | 62.28 | 66.32 | 68.35 | 68.7
11.8 | 62.02 | 65.8 | 68.1 | 68.48
11.9 | 61.74 | 65.5667 | 67.75 | 68.14
12 | 61.2857 | 65.25 | 67.4857 | 67.84
12.1 | 60.9 | 64.86 | 67.18 | 67.6
12.2 | 60.6333 | 64.5333 | 66.8 | 67.34
12.3 | 60.4 | 64.3 | 66.64 | 66.975
12.4 | 59.95 | 64.075 | 66.24 | 66.68
12.5 | 59.7 | 63.75 | 65.95 | 66.375
12.6 | 59.3571 | 63.38 | 65.6857 | 66.15
12.7 | 59.02 | 63.1 | 65.45 | 65.82
12.8 | 58.8 | 62.74 | 65.1875 | 65.64
12.9 | 58.4 | 62.5 | 64.9 | 65.42
13 | 58.1778 | 62.18 | 64.65 | 65.2
13.1 | 57.82 | 61.9 | 64.2 | 64.8667
13.2 | 57.5 | 61.6 | 63.9333 | 64.55
13.3 | 57.225 | 61.3 | 63.7 | 64.22
13.4 | 56.9667 | 61 | 63.5 | 64
13.5 | 56.65 | 60.75 | 63.2 | 63.7
13.6 | 56.3667 | 60.38 | 62.9 | 63.4
13.7 | 56.1 | 60.22 | 62.6333 | 63.16
13.8 | 55.8 | 59.9 | 62.35 | 62.875
13.9 | 55.48 | 59.56 | 62.1 | 62.64
14 | 55.15 | 59.26 | 61.86 | 62.35
14.1 | 54.7714 | 58.875 | 61.64 | 62.1
14.2 | 54.4857 | 58.65 | 61.35 | 61.88
14.3 | 54.24 | 58.34 | 61.04 | 61.5
14.4 | 54 | 58.12 | 60.9 | 61.3
14.5 | 53.7714 | 57.875 | 60.58 | 61.1
14.6 | 53.4 | 57.575 | 60.2 | 60.8
14.7 | 53.1714 | 57.32 | 60 | 60.625
14.8 | 52.84 | 57.04 | 59.74 | 60.3
14.9 | 52.6333 | 56.8 | 59.48 | 60.0333
15 | 52.375 | 56.56 | 59.1857 | 59.8
15.1 | 52.04 | 56.24 | 59.0125 | 59.56
15.2 | 51.775 | 55.94 | 58.76 | 59.35
15.3 | 51.5 | 55.7 | 58.5 | 59.175
15.4 | 51.2222 | 55.5 | 58.1714 | 58.82
15.5 | 50.9 | 55.2 | 58 | 58.5
15.6 | 50.725 | 54.9 | 57.7333 | 58.3
15.7 | 50.4667 | 54.6 | 57.5 | 58.0429
15.8 | 50.1 | 54.4 | 57.175 | 57.8
15.9 | 49.82 | 54.1 | 56.92 | 57.56
16 | 49.6 | 53.9 | 56.6714 | 57.2
16.1 | 49.375 | 53.6333 | 56.42 | 57.08
16.2 | 49.2 | 53.3167 | 56.1889 | 56.85
16.3 | 48.875 | 53.0333 | 55.96 | 56.58
16.4 | 48.65 | 52.875 | 55.6857 | 56.375
16.5 | 48.35 | 52.6 | 55.52 | 56.08
16.6 | 48.1 | 52.46 | 55.2889 | 55.775
16.7 | 47.78 | 52.1 | 55.05 | 55.56
16.8 | 47.5 | 51.88 | 54.78 | 55.375
16.9 | 47.3 | 51.58 | 54.56 | 55.2
17 | 47.0143 | 51.36 | 54.3 | 54.875
17.1 | 46.9 | 51.06 | 54.1 | 54.64
17.2 | 46.66 | 50.88 | 53.7833 | 54.4
17.3 | 46.3143 | 50.54 | 53.5571 | 54.14
17.4 | 46.075 | 50.4 | 53.325 | 53.9
17.5 | 45.9 | 50.04 | 53 | 53.7429
17.6 | 45.55 | 49.9 | 52.7714 | 53.475
17.7 | 45.3 | 49.5 | 52.5429 | 53.24
17.8 | 45.04 | 49.2 | 52.4 | 52.96
17.9 | 44.82 | 49.1 | 52.15 | 52.7
18 | 44.6 | 48.8 | 51.95 | 52.6
18.1 | 44.3667 | 48.6 | 51.6 | 52.3
18.2 | 44.1286 | 48.3 | 51.3571 | 52.125
18.3 | 43.925 | 48.12 | 51.175 | 51.9
18.4 | 43.7111 | 47.9 | 51 | 51.65
18.5 | 43.4 | 47.76 | 50.7857 | 51.375
18.6 | 43.3143 | 47.46 | 50.62 | 51.22
18.7 | 43.14 | 47.225 | 50.3571 | 51
18.8 | 42.8875 | 47 | 50.1 | 50.7429
18.9 | 42.6 | 46.7667 | 49.8667 | 50.56
19 | 42.4286 | 46.5667 | 49.7 | 50.35
19.1 | 42.16 | 46.325 | 49.4667 | 50.1
19.2 | 41.9333 | 46.12 | 49.28 | 49.975
19.3 | 41.72 | 45.925 | 48.9778 | 49.7
19.4 | 41.4857 | 45.66 | 48.825 | 49.54
19.5 | 41.32 | 45.4 | 48.6333 | 49.225
19.6 | 41.0857 | 45.24 | 48.38 | 49.1
19.7 | 40.92 | 45.1 | 48.2 | 48.8571
19.8 | 40.7167 | 44.9 | 47.98 | 48.6667
19.9 | 40.46 | 44.6 | 47.8 | 48.38
20 | 40.25 | 44.4 | 47.58 | 48.2
20.1 | 40.02 | 44.24 | 47.3286 | 48.0286
20.2 | 39.8 | 43.9 | 47.1 | 47.9
20.3 | 39.64 | 43.7 | 46.8286 | 47.6
20.4 | 39.34 | 43.55 | 46.7 | 47.4333
20.5 | 39.2 | 43.35 | 46.46 | 47.1889
20.6 | 38.9667 | 43.1 | 46.2571 | 46.96
20.7 | 38.78 | 42.95 | 46.0286 | 46.78
20.8 | 38.575 | 42.7167 | 45.9 | 46.675
20.9 | 38.4 | 42.5 | 45.675 | 46.4
21 | 38.1429 | 42.275 | 45.55 | 46.175
21.1 | 37.95 | 42.06 | 45.36 | 46
21.2 | 37.72 | 41.92 | 45.1 | 45.7833
21.3 | 37.6 | 41.66 | 44.88 | 45.6
21.4 | 37.3857 | 41.52 | 44.6714 | 45.38
21.5 | 37.2 | 41.3333 | 44.5167 | 45.2143
21.6 | 36.975 | 41.1 | 44.3 | 45.1
21.7 | 36.8 | 40.85 | 44.075 | 44.88
21.8 | 36.6571 | 40.7333 | 43.9 | 44.6
21.9 | 36.5143 | 40.5 | 43.68 | 44.4143
22 | 36.3 | 40.3667 | 43.4714 | 44.2
22.1 | 36.1 | 40.14 | 43.3 | 44.04
22.2 | 35.9 | 39.92 | 43.08 | 43.875
22.3 | 35.675 | 39.7 | 42.9 | 43.66
22.4 | 35.4667 | 39.575 | 42.7833 | 43.425
22.5 | 35.3 | 39.34 | 42.5 | 43.3
22.6 | 35.16 | 39.14 | 42.325 | 43.14
22.7 | 34.9 | 39 | 42.1 | 42.9
22.8 | 34.8 | 38.8 | 41.925 | 42.74
22.9 | 34.6167 | 38.6 | 41.8 | 42.525
23 | 34.5 | 38.38 | 41.55 | 42.3
23.1 | 34.3 | 38.2 | 41.34 | 42.2
23.2 | 34.05 | 38 | 41.2833 | 41.98
23.3 | 33.925 | 37.9 | 41.08 | 41.86
23.4 | 33.8 | 37.66 | 40.9 | 41.6
23.5 | 33.575 | 37.5 | 40.675 | 41.46
23.6 | 33.4 | 37.4 | 40.5 | 41.325
23.7 | 33.28 | 37.16 | 40.3 | 41.1167
23.8 | 33 | 36.96 | 40.14 | 40.875
23.9 | 32.88 | 36.72 | 40 | 40.76
24 | 32.7 | 36.6 | 39.76 | 40.5333
24.1 | 32.54 | 36.38 | 39.6 | 40.42
24.2 | 32.4 | 36.26 | 39.46 | 40.2
24.3 | 32.24 | 36.0333 | 39.3 | 40.02
24.4 | 32.1 | 35.9 | 39.0714 | 39.9
24.5 | 31.95 | 35.7 | 38.9 | 39.675
24.6 | 31.84 | 35.575 | 38.7429 | 39.55
24.7 | 31.6 | 35.35 | 38.6 | 39.4
24.8 | 31.44 | 35.225 | 38.375 | 39.2
24.9 | 31.26 | 35.0286 | 38.2 | 39
25 | 31.1125 | 34.85 | 38.0714 | 38.8
25.1 | 31 | 34.64 | 37.8 | 38.68
25.2 | 30.7 | 34.5 | 37.78 | 38.54
25.3 | 30.675 | 34.375 | 37.55 | 38.34
25.4 | 30.5 | 34.22 | 37.35 | 38.2
25.5 | 30.3556 | 34.05 | 37.2 | 37.98
25.6 | 30.14 | 33.82 | 37.04 | 37.8667
25.7 | 29.9857 | 33.75 | 36.9 | 37.64
25.8 | 29.8 | 33.5714 | 36.7 | 37.5
25.9 | 29.6875 | 33.4 | 36.6 | 37.38
26 | 29.54 | 33.22 | 36.4 | 37.14
26.1 | 29.3571 | 33.1 | 36.18 | 37
26.2 | 29.2143 | 32.8833 | 36.1 | 36.9
26.3 | 29.1 | 32.7 | 35.9 | 36.7
26.4 | 28.9 | 32.6 | 35.74 | 36.6
26.5 | 28.8 | 32.5 | 35.675 | 36.4
26.6 | 28.675 | 32.275 | 35.45 | 36.26
26.7 | 28.4667 | 32.1 | 35.225 | 36.12
26.8 | 28.3143 | 31.9 | 35.1 | 35.925
26.9 | 28.2 | 31.78 | 35 | 35.7
27 | 28.08 | 31.675 | 34.9 | 35.625
27.1 | 27.92 | 31.54 | 34.68 | 35.5
27.2 | 27.8 | 31.4 | 34.4714 | 35.38
27.3 | 27.625 | 31.2143 | 34.3286 | 35.1
27.4 | 27.5 | 31.1 | 34.2 | 35
27.5 | 27.4 | 30.9 | 34.1 | 34.82
27.6 | 27.2222 | 30.8 | 33.9 | 34.68
27.7 | 27.15 | 30.6 | 33.74 | 34.58
27.8 | 26.925 | 30.44 | 33.6 | 34.45
27.9 | 26.8 | 30.4 | 33.425 | 34.22
28 | 26.6143 | 30.2 | 33.3 | 34
28.1 | 26.5 | 29.975 | 33.22 | 33.9
28.2 | 26.4333 | 29.8714 | 33 | 33.8
28.3 | 26.2 | 29.7 | 32.8 | 33.64
28.4 | 26.1 | 29.6 | 32.7 | 33.4778
28.5 | 26 | 29.5 | 32.6 | 33.34
28.6 | 25.9167 | 29.375 | 32.45 | 33.26
28.7 | 25.8 | 29.14 | 32.3 | 33.1
28.8 | 25.6286 | 29.04 | 32.1 | 32.9
28.9 | 25.58 | 28.9 | 32 | 32.8
29 | 25.3778 | 28.8 | 31.9286 | 32.6875
29.1 | 25.225 | 28.7 | 31.7 | 32.54
29.2 | 25.1 | 28.5 | 31.65 | 32.3429
29.3 | 24.9222 | 28.4 | 31.4 | 32.2
29.4 | 24.8333 | 28.2 | 31.3 | 32.02
29.5 | 24.72 | 28.1 | 31.125 | 31.9
29.6 | 24.5333 | 27.9375 | 31.02 | 31.7857
29.7 | 24.425 | 27.8 | 30.9 | 31.65
29.8 | 24.3333 | 27.6857 | 30.8 | 31.58
29.9 | 24.25 | 27.6 | 30.6 | 31.4
30 | 24.1 | 27.4286 | 30.45 | 31.2375

## SSD80 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 33 | 40.3 | 53.2 | 61.5
0.1 | 36.68 | 44.05 | 56.75 | 65
0.2 | 43 | 50 | 62.3 | 69.975
0.3 | 52.9 | 59.22 | 70.2 | 76.3
0.4 | 62.3 | 67.6 | 77.3 | 82.32
0.5 | 69.475 | 74.65 | 82.4333 | 86.8
0.6 | 76.26 | 80.3 | 86.68 | 90.35
0.7 | 80.85 | 84.54 | 90.075 | 93
0.8 | 85.4 | 88.05 | 92.75 | 94.9333
0.9 | 88.65 | 90.975 | 94.775 | 96.7
1 | 91.325 | 93.3 | 96.25 | 97.8
1.1 | 93.6 | 95.2 | 97.6 | 98.7
1.2 | 95.425 | 96.575 | 98.425 | 99.25
1.3 | 96.7 | 97.775 | 99.125 | 99.7
1.4 | 97.82 | 98.5 | 99.55 | 99.94
1.5 | 98.6 | 99.1286 | 99.9 | 100
1.6 | 99.32 | 99.5333 | 100 | 100
1.7 | 99.8 | 99.9 | 100 | 100
1.8 | 99.9 | 100 | 100 | 99.875
1.9 | 100 | 100 | 99.9 | 99.7
2 | 100 | 100 | 99.7 | 99.475
2.1 | 100 | 99.85 | 99.45 | 99.24
2.2 | 99.775 | 99.62 | 99.1667 | 98.875
2.3 | 99.5143 | 99.45 | 98.85 | 98.55
2.4 | 99.3 | 99.15 | 98.4 | 98.16
2.5 | 98.8143 | 98.75 | 98.12 | 97.8
2.6 | 98.5 | 98.35 | 97.7 | 97.4
2.7 | 98.025 | 97.925 | 97.35 | 97
2.8 | 97.6 | 97.4333 | 96.925 | 96.6
2.9 | 97.1 | 97.25 | 96.475 | 96.2
3 | 96.7 | 96.7333 | 96.05 | 95.8
3.1 | 96.2375 | 96.2571 | 95.6333 | 95.4
3.2 | 95.6875 | 95.85 | 95.325 | 94.9333
3.3 | 95.25 | 95.32 | 94.85 | 94.55
3.4 | 94.6625 | 94.9286 | 94.45 | 94.14
3.5 | 94.25 | 94.5375 | 94.0333 | 93.65
3.6 | 93.7125 | 94.1 | 93.56 | 93.32
3.7 | 93.15 | 93.6 | 93.2 | 92.85
3.8 | 92.65 | 93.15 | 92.6 | 92.45
3.9 | 92.15 | 92.7 | 92.35 | 92.05
4 | 91.6714 | 92.2 | 91.9 | 91.65
4.1 | 91.2 | 91.8 | 91.5 | 91.25
4.2 | 90.7 | 91.3 | 91.05 | 90.85
4.3 | 90.1 | 90.9143 | 90.6333 | 90.4
4.4 | 89.675 | 90.4286 | 90.1 | 90
4.5 | 89.25 | 89.975 | 89.8 | 89.6
4.6 | 88.75 | 89.625 | 89.4 | 89.3
4.7 | 88.1286 | 89.1333 | 88.9 | 88.8
4.8 | 87.6 | 88.75 | 88.5667 | 88.4
4.9 | 87.05 | 88.25 | 88.1 | 88
5 | 86.7 | 87.8667 | 87.75 | 87.55
5.1 | 86.3333 | 87.45 | 87.325 | 87.2333
5.2 | 85.6 | 86.9 | 86.925 | 86.9
5.3 | 85.15 | 86.55 | 86.5286 | 86.4
5.4 | 84.76 | 86.1 | 86.0889 | 86.1
5.5 | 84.225 | 85.6667 | 85.7 | 85.7333
5.6 | 83.8 | 85.2 | 85.3667 | 85.275
5.7 | 83.3 | 84.86 | 85 | 84.95
5.8 | 82.8 | 84.45 | 84.5333 | 84.475
5.9 | 82.45 | 83.94 | 84.1875 | 84.08
6 | 82 | 83.5 | 83.75 | 83.68
6.1 | 81.42 | 83.1 | 83.425 | 83.35
6.2 | 80.925 | 82.7 | 82.875 | 82.9
6.3 | 80.55 | 82.22 | 82.55 | 82.5
6.4 | 80.1 | 81.8 | 82.1333 | 82.2
6.5 | 79.7375 | 81.45 | 81.7875 | 81.8
6.6 | 79.1667 | 81.05 | 81.4125 | 81.45
6.7 | 78.775 | 80.65 | 80.9889 | 81
6.8 | 78.2 | 80.25 | 80.6 | 80.65
6.9 | 77.775 | 79.8 | 80.3 | 80.4
7 | 77.4 | 79.4 | 79.85 | 79.9333
7.1 | 76.92 | 79 | 79.475 | 79.56
7.2 | 76.45 | 78.56 | 79.1333 | 79.15
7.3 | 76.0333 | 78.2 | 78.6625 | 78.75
7.4 | 75.575 | 77.8 | 78.2875 | 78.4
7.5 | 75.1 | 77.4 | 77.875 | 78.05
7.6 | 74.7 | 76.95 | 77.58 | 77.6667
7.7 | 74.25 | 76.525 | 77.2 | 77.3333
7.8 | 74.025 | 76.175 | 76.85 | 76.85
7.9 | 73.3667 | 75.825 | 76.5 | 76.55
8 | 72.975 | 75.35 | 76.05 | 76.175
8.1 | 72.58 | 74.95 | 75.8 | 75.825
8.2 | 72.1 | 74.55 | 75.4 | 75.4667
8.3 | 71.7571 | 74.14 | 75 | 75.1333
8.4 | 71.3 | 73.85 | 74.7 | 74.9
8.5 | 70.95 | 73.45 | 74.4 | 74.5
8.6 | 70.5 | 73.075 | 74.025 | 74.1
8.7 | 70.15 | 72.65 | 73.7 | 73.8
8.8 | 69.72 | 72.25 | 73.3 | 73.5
8.9 | 69.35 | 71.925 | 72.975 | 73.1
9 | 68.9143 | 71.575 | 72.625 | 72.7667
9.1 | 68.4667 | 71.15 | 72.275 | 72.5
9.2 | 68.1333 | 70.9 | 71.85 | 72.1
9.3 | 67.775 | 70.45 | 71.475 | 71.775
9.4 | 67.375 | 70.1 | 71.15 | 71.425
9.5 | 66.96 | 69.75 | 70.78 | 71.05
9.6 | 66.5333 | 69.4 | 70.54 | 70.725
9.7 | 66.15 | 69 | 70.16 | 70.375
9.8 | 65.75 | 68.6 | 69.85 | 70.05
9.9 | 65.375 | 68.25 | 69.5333 | 69.75
10 | 65.05 | 67.9 | 69.15 | 69.275
10.1 | 64.7 | 67.5 | 68.8667 | 69.06
10.2 | 64.3 | 67.225 | 68.5 | 68.72
10.3 | 63.9 | 66.775 | 68.1333 | 68.425
10.4 | 63.5 | 66.525 | 67.85 | 68.0667
10.5 | 63.2 | 66.2333 | 67.5 | 67.775
10.6 | 62.85 | 65.775 | 67.175 | 67.5333
10.7 | 62.4 | 65.425 | 66.85 | 67.175
10.8 | 62.15 | 65.0333 | 66.475 | 66.8
10.9 | 61.7 | 64.775 | 66.225 | 66.525
11 | 61.36 | 64.38 | 65.875 | 66.2
11.1 | 61.05 | 64.15 | 65.525 | 65.875
11.2 | 60.675 | 63.8 | 65.175 | 65.6222
11.3 | 60.325 | 63.4 | 64.925 | 65.2667
11.4 | 59.975 | 63.1 | 64.6 | 65
11.5 | 59.7 | 62.725 | 64.28 | 64.7625
11.6 | 59.35 | 62.4 | 64 | 64.3875
11.7 | 59.025 | 62.1 | 63.65 | 64.0125
11.8 | 58.66 | 61.7 | 63.3 | 63.725
11.9 | 58.3 | 61.4 | 63 | 63.375
12 | 57.9667 | 61.1 | 62.7 | 63.125
12.1 | 57.6 | 60.75 | 62.5 | 62.8125
12.2 | 57.25 | 60.4 | 62.12 | 62.5444
12.3 | 57 | 60.125 | 61.825 | 62.2625
12.4 | 56.6 | 59.75 | 61.5 | 61.9
12.5 | 56.25 | 59.525 | 61.175 | 61.6333
12.6 | 56 | 59.1 | 60.925 | 61.3333
12.7 | 55.6 | 58.8 | 60.65 | 61.05
12.8 | 55.3333 | 58.5143 | 60.325 | 60.75
12.9 | 54.95 | 58.2333 | 60.075 | 60.5
13 | 54.6 | 57.95 | 59.775 | 60.16
13.1 | 54.475 | 57.625 | 59.425 | 59.95
13.2 | 54.125 | 57.3625 | 59.1875 | 59.6
13.3 | 53.75 | 56.95 | 58.9 | 59.275
13.4 | 53.425 | 56.64 | 58.5429 | 59.025
13.5 | 53.2 | 56.3667 | 58.2889 | 58.775
13.6 | 52.8 | 56.15 | 57.95 | 58.5
13.7 | 52.55 | 55.8 | 57.7 | 58.175
13.8 | 52.3 | 55.5 | 57.45 | 57.925
13.9 | 51.95 | 55.2 | 57.15 | 57.6
14 | 51.7 | 55 | 56.875 | 57.35
14.1 | 51.4 | 54.7 | 56.65 | 57.075
14.2 | 51.1 | 54.4 | 56.35 | 56.825
14.3 | 50.775 | 54 | 56 | 56.6
14.4 | 50.525 | 53.7667 | 55.75 | 56.3
14.5 | 50.175 | 53.4667 | 55.525 | 56
14.6 | 49.925 | 53.225 | 55.2667 | 55.75
14.7 | 49.575 | 52.975 | 54.975 | 55.5
14.8 | 49.35 | 52.6375 | 54.75 | 55.25
14.9 | 49.075 | 52.375 | 54.425 | 54.9
15 | 48.8 | 52.125 | 54.18 | 54.7
15.1 | 48.5 | 51.775 | 53.9 | 54.45
15.2 | 48.25 | 51.525 | 53.7 | 54.22
15.3 | 48 | 51.275 | 53.4 | 53.975
15.4 | 47.65 | 51 | 53.2 | 53.725
15.5 | 47.425 | 50.7 | 52.85 | 53.42
15.6 | 47.1833 | 50.45 | 52.6667 | 53.175
15.7 | 46.95 | 50.2 | 52.3 | 52.925
15.8 | 46.62 | 49.95 | 52.1 | 52.6375
15.9 | 46.3 | 49.62 | 51.85 | 52.3833
16 | 46.125 | 49.45 | 51.6 | 52.175
16.1 | 45.875 | 49.125 | 51.35 | 51.925
16.2 | 45.625 | 48.8667 | 51.1 | 51.675
16.3 | 45.375 | 48.625 | 50.86 | 51.3875
16.4 | 45.125 | 48.35 | 50.65 | 51.1
16.5 | 44.8 | 48.125 | 50.325 | 50.95
16.6 | 44.55 | 47.875 | 50.1 | 50.7
16.7 | 44.375 | 47.625 | 49.85 | 50.5
16.8 | 44.05 | 47.375 | 49.6 | 50.3
16.9 | 43.8333 | 47.125 | 49.425 | 49.95
17 | 43.625 | 46.9 | 49.15 | 49.6857
17.1 | 43.42 | 46.6 | 48.95 | 49.525
17.2 | 43.125 | 46.4 | 48.625 | 49.24
17.3 | 42.875 | 46.15 | 48.375 | 49.0167
17.4 | 42.625 | 45.8 | 48.125 | 48.725
17.5 | 42.4 | 45.6 | 47.875 | 48.525
17.6 | 42.15 | 45.3 | 47.625 | 48.275
17.7 | 41.9 | 45.05 | 47.4 | 48.025
17.8 | 41.625 | 44.82 | 47.2 | 47.825
17.9 | 41.4 | 44.575 | 47 | 47.58
18 | 41.14 | 44.325 | 46.725 | 47.35
18.1 | 40.9 | 44.175 | 46.45 | 47.1
18.2 | 40.65 | 43.925 | 46.3 | 46.85
18.3 | 40.4 | 43.7 | 46.05 | 46.7
18.4 | 40.3 | 43.475 | 45.9 | 46.475
18.5 | 40.05 | 43.225 | 45.6 | 46.325
18.6 | 39.8 | 43 | 45.4 | 46.025
18.7 | 39.6 | 42.725 | 45.1 | 45.8
18.8 | 39.4 | 42.58 | 44.975 | 45.625
18.9 | 39.1 | 42.325 | 44.8 | 45.4
19 | 38.95 | 42.1 | 44.575 | 45.15
19.1 | 38.7 | 41.95 | 44.3 | 44.98
19.2 | 38.46 | 41.7 | 44.075 | 44.75
19.3 | 38.3 | 41.425 | 43.825 | 44.5
19.4 | 38.075 | 41.3 | 43.575 | 44.36
19.5 | 37.825 | 41.05 | 43.4 | 44.1
19.6 | 37.675 | 40.8 | 43.175 | 43.85
19.7 | 37.5 | 40.65 | 42.925 | 43.72
19.8 | 37.275 | 40.4 | 42.7833 | 43.4833
19.9 | 37.025 | 40.15 | 42.5333 | 43.2667
20 | 36.775 | 40 | 42.4 | 43
20.1 | 36.6 | 39.75 | 42.18 | 42.9
20.2 | 36.375 | 39.6 | 41.95 | 42.7
20.3 | 36.225 | 39.3667 | 41.7 | 42.45
20.4 | 36.1 | 39.1 | 41.55 | 42.3
20.5 | 35.85 | 38.9 | 41.3333 | 42.05
20.6 | 35.6 | 38.75 | 41.2 | 41.8
20.7 | 35.45 | 38.52 | 40.9 | 41.65
20.8 | 35.2 | 38.3 | 40.75 | 41.5
20.9 | 35 | 38.2 | 40.5 | 41.3
21 | 34.85 | 37.875 | 40.35 | 41.05
21.1 | 34.6 | 37.725 | 40.1 | 40.85
21.2 | 34.475 | 37.475 | 39.975 | 40.7
21.3 | 34.325 | 37.3 | 39.8 | 40.45
21.4 | 34.075 | 37.175 | 39.675 | 40.3
21.5 | 33.9 | 36.925 | 39.4143 | 40.1
21.6 | 33.775 | 36.68 | 39.1667 | 39.8667
21.7 | 33.5 | 36.6 | 38.925 | 39.675
21.8 | 33.3 | 36.375 | 38.775 | 39.525
21.9 | 33.125 | 36.2 | 38.5625 | 39.3
22 | 33 | 36 | 38.4375 | 39.125
22.1 | 32.75 | 35.8 | 38.225 | 38.9667
22.2 | 32.6 | 35.6 | 38 | 38.775
22.3 | 32.4 | 35.375 | 37.875 | 38.575
22.4 | 32.2667 | 35.2 | 37.6857 | 38.325
22.5 | 32.05 | 34.975 | 37.44 | 38.175
22.6 | 31.88 | 34.8333 | 37.25 | 38.05
22.7 | 31.66 | 34.68 | 37.075 | 37.8
22.8 | 31.5 | 34.4333 | 36.9 | 37.7
22.9 | 31.3556 | 34.3 | 36.775 | 37.5
23 | 31.15 | 34.0667 | 36.62 | 37.25
23.1 | 31.025 | 33.9 | 36.4286 | 37.1
23.2 | 30.9 | 33.8 | 36.2222 | 36.95
23.3 | 30.65 | 33.55 | 36.0556 | 36.8
23.4 | 30.52 | 33.4 | 35.9 | 36.6
23.5 | 30.3833 | 33.2333 | 35.6875 | 36.4
23.6 | 30.2167 | 33 | 35.525 | 36.26
23.7 | 29.9875 | 32.9 | 35.3375 | 36.1
23.8 | 29.8625 | 32.7 | 35.125 | 35.9
23.9 | 29.675 | 32.56 | 34.98 | 35.725
24 | 29.5125 | 32.425 | 34.8 | 35.575
24.1 | 29.3857 | 32.2 | 34.65 | 35.425
24.2 | 29.2 | 32.025 | 34.5 | 35.2
24.3 | 29.1 | 31.9 | 34.3 | 35.1
24.4 | 28.9333 | 31.725 | 34.1 | 34.9
24.5 | 28.7 | 31.575 | 34 | 34.7
24.6 | 28.6 | 31.425 | 33.9 | 34.55
24.7 | 28.5 | 31.2 | 33.64 | 34.4
24.8 | 28.25 | 31.05 | 33.475 | 34.25
24.9 | 28.1 | 30.9 | 33.3 | 34.1
25 | 27.95 | 30.75 | 33.2 | 33.94
25.1 | 27.8 | 30.6 | 33.0333 | 33.75
25.2 | 27.65 | 30.45 | 32.8667 | 33.55
25.3 | 27.5 | 30.3 | 32.7 | 33.4
25.4 | 27.325 | 30.15 | 32.5333 | 33.275
25.5 | 27.2 | 29.9 | 32.4 | 33.1
25.6 | 27.1 | 29.75 | 32.2333 | 32.9143
25.7 | 26.9 | 29.7 | 32 | 32.7625
25.8 | 26.8 | 29.45 | 31.9 | 32.6857
25.9 | 26.6 | 29.3 | 31.7 | 32.52
26 | 26.5 | 29.2 | 31.6 | 32.325
26.1 | 26.4 | 29.1 | 31.475 | 32.1667
26.2 | 26.25 | 28.95 | 31.3 | 32.025
26.3 | 26.1 | 28.725 | 31.125 | 31.9
26.4 | 25.95 | 28.6 | 30.98 | 31.725
26.5 | 25.8 | 28.425 | 30.85 | 31.6
26.6 | 25.65 | 28.275 | 30.7 | 31.4
26.7 | 25.5 | 28.125 | 30.5 | 31.3
26.8 | 25.35 | 27.975 | 30.4 | 31.1
26.9 | 25.2 | 27.9 | 30.25 | 31
27 | 25.1 | 27.7 | 30.1 | 30.85
27.1 | 25 | 27.55 | 29.95 | 30.7
27.2 | 24.86 | 27.4 | 29.7 | 30.6
27.3 | 24.7 | 27.3 | 29.6 | 30.4
27.4 | 24.575 | 27.1 | 29.525 | 30.25
27.5 | 24.5 | 27 | 29.375 | 30.1
27.6 | 24.3333 | 26.8 | 29.2 | 30
27.7 | 24.2 | 26.7 | 29.075 | 29.825
27.8 | 24 | 26.6 | 28.9 | 29.675
27.9 | 23.925 | 26.45 | 28.8 | 29.525
28 | 23.8 | 26.3333 | 28.675 | 29.4
28.1 | 23.7 | 26.2 | 28.525 | 29.325
28.2 | 23.525 | 26.075 | 28.38 | 29.175
28.3 | 23.4 | 25.925 | 28.25 | 29.025
28.4 | 23.3 | 25.7833 | 28.1 | 28.875
28.5 | 23.2 | 25.675 | 27.95 | 28.74
28.6 | 23.05 | 25.525 | 27.8 | 28.6
28.7 | 23 | 25.4 | 27.75 | 28.45
28.8 | 22.8 | 25.3 | 27.6 | 28.3
28.9 | 22.625 | 25.1 | 27.45 | 28.2
29 | 22.5667 | 25 | 27.3 | 28
29.1 | 22.425 | 24.875 | 27.2333 | 27.95
29.2 | 22.3 | 24.8 | 27.05 | 27.8
29.3 | 22.2 | 24.65 | 26.9 | 27.66
29.4 | 22.1143 | 24.45 | 26.75 | 27.5571
29.5 | 21.925 | 24.4 | 26.6 | 27.4
29.6 | 21.8111 | 24.3 | 26.5 | 27.225
29.7 | 21.6875 | 24.1 | 26.425 | 27.175
29.8 | 21.5625 | 24 | 26.2429 | 27
29.9 | 21.5 | 23.85 | 26.125 | 26.9
30 | 21.4125 | 23.8 | 26 | 26.75

## SSD120 Wedge 45 Depth Doses
Percentage Depth Doses
Wedge 45
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 20x20cm2 | 40x20cm2
0 | 33.5 | 38.8 | 46.8 | 52.1
0.1 | 38.3 | 43.2 | 50.8 | 56.35
0.2 | 46.1 | 50.6 | 58.35 | 62.6
0.3 | 55.6 | 59.7 | 66.1 | 70.075
0.4 | 64.5 | 67.9 | 73.3 | 76.65
0.5 | 70.85 | 74.325 | 78.7 | 81.5333
0.6 | 76.6 | 79.65 | 83.6 | 85.6875
0.7 | 81.2833 | 83.6667 | 86.725 | 88.7875
0.8 | 85.2 | 87.6 | 89.825 | 91.6
0.9 | 88.3571 | 90 | 92.275 | 93.5857
1 | 91 | 92.4 | 94.2 | 95.34
1.1 | 93.14 | 94.3 | 95.5667 | 96.6125
1.2 | 94.9 | 95.8 | 96.925 | 97.625
1.3 | 96.25 | 96.95 | 97.85 | 98.5
1.4 | 97.54 | 98 | 98.55 | 99.075
1.5 | 98.2625 | 98.6 | 99.24 | 99.5
1.6 | 98.8875 | 99.2667 | 99.6333 | 99.825
1.7 | 99.3111 | 99.5333 | 99.8 | 99.95
1.8 | 99.6 | 99.9 | 99.95 | 100
1.9 | 99.8667 | 100 | 100 | 100
2 | 100 | 100 | 100 | 100
2.1 | 100 | 100 | 99.9 | 99.9
2.2 | 99.98 | 99.9 | 99.8 | 99.7875
2.3 | 99.725 | 99.875 | 99.7 | 99.65
2.4 | 99.5 | 99.6 | 99.5 | 99.375
2.5 | 99.25 | 99.425 | 99.3 | 99.2
2.6 | 99 | 99.225 | 99.075 | 98.9857
2.7 | 98.7 | 98.95 | 98.8 | 98.6375
2.8 | 98.325 | 98.5333 | 98.525 | 98.3778
2.9 | 97.9 | 98.2 | 98.2 | 98.1143
3 | 97.6125 | 97.9333 | 97.875 | 97.7333
3.1 | 97.325 | 97.6 | 97.575 | 97.375
3.2 | 96.8714 | 97.1333 | 97.3 | 97.1
3.3 | 96.5 | 96.85 | 96.975 | 96.825
3.4 | 96 | 96.55 | 96.55 | 96.45
3.5 | 95.6143 | 96.2 | 96.25 | 96.15
3.6 | 95.1778 | 95.85 | 95.825 | 95.8143
3.7 | 94.7 | 95.6333 | 95.575 | 95.55
3.8 | 94.4 | 95.075 | 95.18 | 95.36
3.9 | 93.9 | 94.8 | 94.9 | 94.975
4 | 93.4667 | 94.475 | 94.64 | 94.55
4.1 | 93.1 | 94.05 | 94.3 | 94.225
4.2 | 92.7 | 93.7 | 93.98 | 93.85
4.3 | 92.325 | 93.3 | 93.5 | 93.675
4.4 | 91.94 | 92.9333 | 93.15 | 93.3
4.5 | 91.3667 | 92.575 | 92.8 | 92.9
4.6 | 91.0333 | 92.3 | 92.5 | 92.55
4.7 | 90.65 | 91.85 | 92.3 | 92.325
4.8 | 90.1875 | 91.575 | 91.9333 | 92.02
4.9 | 89.8125 | 91.125 | 91.6333 | 91.65
5 | 89.3889 | 90.85 | 91.225 | 91.3375
5.1 | 89 | 90.48 | 90.975 | 91.075
5.2 | 88.5333 | 90.05 | 90.72 | 90.75
5.3 | 88 | 89.8 | 90.325 | 90.3889
5.4 | 87.74 | 89.4 | 89.9 | 90.2
5.5 | 87.4 | 89.05 | 89.7333 | 89.75
5.6 | 87.025 | 88.7 | 89.35 | 89.44
5.7 | 86.55 | 88.35 | 89 | 89.1
5.8 | 86.15 | 88 | 88.75 | 88.75
5.9 | 85.7143 | 87.65 | 88.5 | 88.4
6 | 85.475 | 87.3 | 88.2 | 88.15
6.1 | 84.95 | 87.1 | 87.7 | 87.925
6.2 | 84.575 | 86.675 | 87.4667 | 87.5667
6.3 | 84.2125 | 86.275 | 87.2 | 87.225
6.4 | 83.675 | 86.025 | 86.85 | 86.9571
6.5 | 83.38 | 85.45 | 86.62 | 86.6875
6.6 | 83.1 | 85.125 | 86.175 | 86.275
6.7 | 82.65 | 84.9 | 85.85 | 86.0833
6.8 | 82.3 | 84.55 | 85.575 | 85.75
6.9 | 81.85 | 84.25 | 85.25 | 85.425
7 | 81.4 | 83.75 | 84.975 | 85.16
7.1 | 80.95 | 83.5 | 84.725 | 84.8286
7.2 | 80.6 | 83.24 | 84.25 | 84.5
7.3 | 80.275 | 82.9 | 84.1 | 84.2286
7.4 | 79.925 | 82.45 | 83.675 | 83.96
7.5 | 79.45 | 82.1 | 83.38 | 83.6375
7.6 | 79.125 | 81.9 | 83 | 83.2
7.7 | 78.6571 | 81.5 | 82.75 | 82.95
7.8 | 78.32 | 81.1667 | 82.375 | 82.6333
7.9 | 77.975 | 80.85 | 82.05 | 82.35
8 | 77.675 | 80.5 | 81.8 | 82.0667
8.1 | 77.22 | 80.175 | 81.5 | 81.8
8.2 | 76.9 | 79.8667 | 81.2 | 81.5
8.3 | 76.45 | 79.525 | 80.95 | 81.06
8.4 | 76.2 | 79.15 | 80.7 | 80.8143
8.5 | 75.7 | 78.6 | 80.3 | 80.575
8.6 | 75.4 | 78.325 | 80 | 80.4
8.7 | 74.9667 | 78.1 | 79.7333 | 79.95
8.8 | 74.72 | 77.8 | 79.3667 | 79.65
8.9 | 74.25 | 77.5 | 79.05 | 79.425
9 | 73.825 | 77.15 | 78.725 | 79.1
9.1 | 73.475 | 76.95 | 78.475 | 78.9
9.2 | 73.15 | 76.675 | 78.225 | 78.425
9.3 | 72.8125 | 76.3 | 77.875 | 78.2
9.4 | 72.38 | 75.925 | 77.525 | 77.9
9.5 | 72.05 | 75.5857 | 77.275 | 77.6
9.6 | 71.8 | 75.275 | 77.0333 | 77.3
9.7 | 71.35 | 74.925 | 76.775 | 77
9.8 | 71.05 | 74.575 | 76.48 | 76.7333
9.9 | 70.7 | 74.325 | 76.2 | 76.4667
10 | 70.35 | 74.0571 | 75.88 | 76.1
10.1 | 69.9444 | 73.65 | 75.6 | 75.9
10.2 | 69.6625 | 73.375 | 75.25 | 75.675
10.3 | 69.2875 | 73.025 | 75 | 75.3333
10.4 | 68.875 | 72.675 | 74.65 | 75.1667
10.5 | 68.625 | 72.425 | 74.4 | 74.75
10.6 | 68.2571 | 72.15 | 74.15 | 74.5
10.7 | 67.9 | 71.8 | 73.9 | 74.18
10.8 | 67.56 | 71.525 | 73.6 | 73.95
10.9 | 67.25 | 71.175 | 73.3 | 73.625
11 | 66.9 | 70.85 | 72.975 | 73.45
11.1 | 66.65 | 70.5667 | 72.725 | 73.05
11.2 | 66.3 | 70.2 | 72.375 | 72.875
11.3 | 65.95 | 69.9667 | 72.1 | 72.625
11.4 | 65.6 | 69.6333 | 71.775 | 72.3
11.5 | 65.3 | 69.375 | 71.575 | 71.94
11.6 | 65 | 69.05 | 71.225 | 71.7
11.7 | 64.6 | 68.7 | 71.05 | 71.45
11.8 | 64.22 | 68.4333 | 70.725 | 71.08
11.9 | 63.9429 | 68.25 | 70.375 | 70.95
12 | 63.7 | 67.9 | 70.225 | 70.55
12.1 | 63.4 | 67.6 | 69.88 | 70.4
12.2 | 63.1 | 67.3 | 69.65 | 70.06
12.3 | 62.75 | 66.95 | 69.4 | 69.92
12.4 | 62.5 | 66.65 | 69.15 | 69.575
12.5 | 62.15 | 66.45 | 68.9 | 69.25
12.6 | 61.9 | 66.2 | 68.55 | 69
12.7 | 61.6 | 65.8 | 68.2 | 68.75
12.8 | 61.2 | 65.5667 | 68.05 | 68.525
12.9 | 60.95 | 65.24 | 67.7 | 68.275
13 | 60.6 | 64.9 | 67.45 | 68
13.1 | 60.3 | 64.675 | 67.2 | 67.72
13.2 | 60 | 64.4 | 66.9 | 67.4667
13.3 | 59.75 | 64.175 | 66.7 | 67.2
13.4 | 59.42 | 63.825 | 66.45 | 67
13.5 | 59.175 | 63.475 | 66.2 | 66.65
13.6 | 58.8875 | 63.225 | 65.975 | 66.3
13.7 | 58.56 | 63.1 | 65.625 | 66.1667
13.8 | 58.2286 | 62.75 | 65.375 | 65.9
13.9 | 57.975 | 62.5 | 65.1 | 65.65
14 | 57.6875 | 62.15 | 64.775 | 65.3
14.1 | 57.3125 | 61.9 | 64.525 | 65.175
14.2 | 57.025 | 61.55 | 64.2 | 64.92
14.3 | 56.6778 | 61.3 | 64.1 | 64.65
14.4 | 56.45 | 60.95 | 63.875 | 64.325
14.5 | 56.25 | 60.7 | 63.525 | 64.06
14.6 | 55.8571 | 60.45 | 63.275 | 63.9
14.7 | 55.6 | 60.2 | 63 | 63.6
14.8 | 55.35 | 59.95 | 62.675 | 63.35
14.9 | 55.15 | 59.62 | 62.5 | 63.025
15 | 54.84 | 59.3667 | 62.2 | 62.9
15.1 | 54.55 | 59.175 | 62 | 62.62
15.2 | 54.2 | 58.825 | 61.85 | 62.325
15.3 | 54 | 58.675 | 61.6 | 62.1429
15.4 | 53.75 | 58.425 | 61.35 | 62
15.5 | 53.5 | 58.15 | 61.1 | 61.68
15.6 | 53.275 | 57.925 | 60.75 | 61.5
15.7 | 53 | 57.65 | 60.5 | 61.3
15.8 | 52.7 | 57.325 | 60.25 | 60.95
15.9 | 52.4 | 57.075 | 60 | 60.7
16 | 52.15 | 56.8333 | 59.825 | 60.5
16.1 | 52 | 56.575 | 59.6 | 60.3
16.2 | 51.65 | 56.325 | 59.325 | 60.05
16.3 | 51.4 | 56.175 | 59.075 | 59.8
16.4 | 51.15 | 55.925 | 58.85 | 59.5
16.5 | 50.9 | 55.575 | 58.675 | 59.2
16.6 | 50.55 | 55.3 | 58.425 | 59.05
16.7 | 50.4 | 55.1 | 58.175 | 58.8
16.8 | 50.175 | 54.8 | 57.925 | 58.55
16.9 | 49.925 | 54.6 | 57.6 | 58.36
17 | 49.675 | 54.4 | 57.44 | 58.0667
17.1 | 49.425 | 54.1 | 57.28 | 57.825
17.2 | 49.175 | 53.9 | 57 | 57.675
17.3 | 48.85 | 53.6 | 56.8 | 57.425
17.4 | 48.675 | 53.475 | 56.6 | 57.225
17.5 | 48.425 | 53.2 | 56.3 | 56.975
17.6 | 48.175 | 52.975 | 56.05 | 56.725
17.7 | 47.925 | 52.725 | 55.8 | 56.56
17.8 | 47.675 | 52.375 | 55.64 | 56.2667
17.9 | 47.5 | 52.225 | 55.4 | 56.1
18 | 47.25 | 51.975 | 55.24 | 55.825
18.1 | 47.04 | 51.825 | 54.9 | 55.675
18.2 | 46.825 | 51.55 | 54.75 | 55.425
18.3 | 46.6286 | 51.25 | 54.5 | 55.275
18.4 | 46.3143 | 50.975 | 54.25 | 55.025
18.5 | 46.05 | 50.825 | 54.1 | 54.78
18.6 | 45.8 | 50.6 | 53.9 | 54.6
18.7 | 45.6556 | 50.4333 | 53.72 | 54.4
18.8 | 45.425 | 50.175 | 53.36 | 54.15
18.9 | 45.1333 | 49.88 | 53.125 | 54
19 | 44.9167 | 49.7 | 53 | 53.65
19.1 | 44.7 | 49.45 | 52.775 | 53.5
19.2 | 44.525 | 49.3 | 52.525 | 53.25
19.3 | 44.25 | 49 | 52.35 | 53.05
19.4 | 44 | 48.7333 | 52.1 | 52.8333
19.5 | 43.8 | 48.55 | 51.8667 | 52.6
19.6 | 43.625 | 48.4 | 51.6667 | 52.4
19.7 | 43.425 | 48.15 | 51.5 | 52.2
19.8 | 43.225 | 47.9 | 51.275 | 51.975
19.9 | 42.975 | 47.7 | 51.05 | 51.825
20 | 42.75 | 47.5 | 50.825 | 51.5667
20.1 | 42.525 | 47.2333 | 50.58 | 51.4
20.2 | 42.375 | 47.1 | 50.45 | 51.2
20.3 | 42.125 | 46.9 | 50.3 | 50.9667
20.4 | 41.9333 | 46.62 | 50 | 50.8
20.5 | 41.8 | 46.4 | 49.8 | 50.6
20.6 | 41.6 | 46.3 | 49.65 | 50.3667
20.7 | 41.35 | 46 | 49.4 | 50.175
20.8 | 41.2 | 45.825 | 49.15 | 50.025
20.9 | 40.95 | 45.575 | 48.9 | 49.775
21 | 40.625 | 45.425 | 48.76 | 49.525
21.1 | 40.575 | 45.175 | 48.5 | 49.3
21.2 | 40.4 | 45.02 | 48.3667 | 49.125
21.3 | 40.1667 | 44.775 | 48.1667 | 48.9
21.4 | 40 | 44.6 | 48 | 48.75
21.5 | 39.7333 | 44.375 | 47.7 | 48.55
21.6 | 39.525 | 44.15 | 47.6 | 48.4
21.7 | 39.3 | 43.975 | 47.475 | 48.1
21.8 | 39.125 | 43.825 | 47.225 | 47.9286
21.9 | 38.975 | 43.575 | 47 | 47.7667
22 | 38.8 | 43.4 | 46.9 | 47.55
22.1 | 38.575 | 43.18 | 46.675 | 47.4
22.2 | 38.44 | 42.925 | 46.425 | 47.25
22.3 | 38.225 | 42.8 | 46.275 | 47
22.4 | 38 | 42.54 | 46.075 | 46.825
22.5 | 37.9 | 42.4 | 45.8625 | 46.65
22.6 | 37.675 | 42.25 | 45.7375 | 46.475
22.7 | 37.525 | 42 | 45.525 | 46.225
22.8 | 37.275 | 41.85 | 45.275 | 46.075
22.9 | 37.1 | 41.7 | 45.1125 | 45.9
23 | 36.925 | 41.5 | 44.975 | 45.675
23.1 | 36.7 | 41.3 | 44.725 | 45.5
23.2 | 36.525 | 41.1 | 44.55 | 45.4
23.3 | 36.38 | 40.85 | 44.3 | 45.2
23.4 | 36.2 | 40.7 | 44.15 | 45
23.5 | 36.0333 | 40.5 | 44 | 44.85
23.6 | 35.875 | 40.3 | 43.775 | 44.6
23.7 | 35.7 | 40.16 | 43.6 | 44.45
23.8 | 35.525 | 40.025 | 43.4 | 44.3
23.9 | 35.3 | 39.775 | 43.2 | 44.06
24 | 35.2 | 39.625 | 42.9857 | 43.9
24.1 | 35 | 39.4 | 42.9 | 43.7
24.2 | 34.8 | 39.225 | 42.7 | 43.525
24.3 | 34.68 | 39.1 | 42.58 | 43.36
24.4 | 34.5 | 38.925 | 42.35 | 43.125
24.5 | 34.3 | 38.675 | 42.2 | 42.975
24.6 | 34.2 | 38.525 | 42.1 | 42.8
24.7 | 33.95 | 38.375 | 41.875 | 42.6714
24.8 | 33.8 | 38.225 | 41.625 | 42.4889
24.9 | 33.6 | 38.025 | 41.5 | 42.325
25 | 33.5 | 37.875 | 41.35 | 42.2
25.1 | 33.4 | 37.7 | 41.2 | 41.95
25.2 | 33.175 | 37.5 | 41.05 | 41.75
25.3 | 33 | 37.325 | 40.8 | 41.6143
25.4 | 32.8 | 37.2 | 40.65 | 41.4667
25.5 | 32.625 | 37 | 40.5 | 41.3
25.6 | 32.575 | 36.875 | 40.35 | 41.1
25.7 | 32.325 | 36.7 | 40.1 | 40.95
25.8 | 32.2 | 36.5 | 39.95 | 40.8
25.9 | 32.025 | 36.3667 | 39.8 | 40.6
26 | 31.9 | 36.2 | 39.65 | 40.5
26.1 | 31.725 | 36 | 39.425 | 40.275
26.2 | 31.6 | 35.9 | 39.3 | 40.125
26.3 | 31.44 | 35.7 | 39.2 | 39.9
26.4 | 31.3 | 35.5 | 38.975 | 39.775
26.5 | 31.1 | 35.3 | 38.825 | 39.6
26.6 | 30.9 | 35.2 | 38.7167 | 39.5125
26.7 | 30.8 | 35.06 | 38.475 | 39.325
26.8 | 30.7 | 34.9 | 38.4 | 39.175
26.9 | 30.55 | 34.76 | 38.2 | 39.025
27 | 30.4 | 34.6 | 38.05 | 38.875
27.1 | 30.25 | 34.4 | 37.9 | 38.7
27.2 | 30.15 | 34.3 | 37.75 | 38.5
27.3 | 30 | 34.1 | 37.6 | 38.425
27.4 | 29.85 | 34 | 37.4 | 38.2
27.5 | 29.7 | 33.8 | 37.2 | 38.025
27.6 | 29.55 | 33.7 | 37.1 | 37.875
27.7 | 29.4 | 33.475 | 37 | 37.8
27.8 | 29.275 | 33.3 | 36.8 | 37.6
27.9 | 29.125 | 33.175 | 36.6 | 37.5
28 | 28.975 | 33.025 | 36.475 | 37.375
28.1 | 28.825 | 32.875 | 36.3 | 37.125
28.2 | 28.675 | 32.8 | 36.15 | 37
28.3 | 28.525 | 32.6 | 36 | 36.825
28.4 | 28.4 | 32.5 | 35.86 | 36.7
28.5 | 28.3 | 32.4 | 35.7 | 36.55
28.6 | 28.1 | 32.1 | 35.575 | 36.3875
28.7 | 28 | 32 | 35.425 | 36.225
28.8 | 27.9 | 31.9 | 35.26 | 36.1
28.9 | 27.725 | 31.7 | 35.2 | 35.925
29 | 27.6 | 31.6 | 34.975 | 35.7875
29.1 | 27.5 | 31.5 | 34.825 | 35.6625
29.2 | 27.4 | 31.35 | 34.675 | 35.5375
29.3 | 27.25 | 31.2 | 34.525 | 35.4125
29.4 | 27.1 | 31 | 34.4375 | 35.275
29.5 | 26.9 | 30.9 | 34.3125 | 35.0625
29.6 | 26.8 | 30.76 | 34.0875 | 34.9444
29.7 | 26.7 | 30.625 | 34 | 34.825
29.8 | 26.6 | 30.475 | 33.9286 | 34.5875
29.9 | 26.4 | 30.325 | 33.62 | 34.5
30 | 26.3 | 30.175 | 33.5333 | 34.35

## Wedge 45 Lateral Profiles
Lateral Profiles
Wedge 45
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:2 cm FieldSize:4x4 cm2 | Depth:2 cm FieldSize:10x10 cm2 | Depth:2 cm FieldSize:20x20 cm2 | Depth:2 cm FieldSize:40x20 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:20x20 cm2 | Depth:5 cm FieldSize:40x20 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:20x20 cm2 | Depth:10 cm FieldSize:40x20 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:20x20 cm2 | Depth:20 cm FieldSize:40x20 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:20x20 cm2 | Depth:30 cm FieldSize:40x20 cm2
-17.9 | [18 empty] | 13.8611 | 17.36
-17.8 | [18 empty] | 14.0744 | 17.56
-17.7 | [18 empty] | 14.2193 | 17.775
-17.6 | [18 empty] | 14.4342 | 17.9667
-17.5 | [18 empty] | 14.6691 | 18.1
-17.4 | [18 empty] | 14.7741 | 18.42
-17.3 | [18 empty] | 15.019 | 18.66
-17.2 | [18 empty] | 15.2339 | 18.8667
-17.1 | [18 empty] | 15.5688 | 19.1
-17 | [18 empty] | 15.6937 | 19.4
-16.9 | [14 empty] | 11.05 | 14.175 |  |  | 15.9686 | 19.62
-16.8 | [14 empty] | 11.3 | 14.4 |  |  | 16.1735 | 19.85
-16.7 | [14 empty] | 11.5 | 14.6 |  |  | 16.4267 | 20.08
-16.6 | [14 empty] | 11.7 | 14.8 |  |  | 16.7933 | 20.3
-16.5 | [14 empty] | 11.8 | 15 |  |  | 17.0182 | 20.575
-16.4 | [14 empty] | 11.94 | 15.14 |  |  | 17.1931 | 20.9
-16.3 | [14 empty] | 12.175 | 15.32 |  |  | 17.468 | 21.175
-16.2 | [14 empty] | 12.38 | 15.5667 |  |  | 17.8928 | 21.5333
-16.1 | [14 empty] | 12.6 | 15.8 |  |  | 18.0928 | 21.76
-16 | [14 empty] | 12.74 | 16.025 |  |  | 18.3727 | 22
-15.9 | [10 empty] | 7.9813 | 10.6833 |  |  | 12.94 | 16.26 |  |  | 18.5726 | 22.42
-15.8 | [10 empty] | 8.198 | 10.8 |  |  | 13.2333 | 16.46 |  |  | 18.9324 | 22.7
-15.7 | [10 empty] | 8.2979 | 10.9571 |  |  | 13.375 | 16.7 |  |  | 19.2723 | 23.04
-15.6 | [10 empty] | 8.4479 | 11.1 |  |  | 13.6714 | 17 |  |  | 19.6322 | 23.35
-15.5 | [10 empty] | 8.4979 | 11.28 |  |  | 13.9 | 17.3 |  |  | 19.942 | 23.62
-15.4 | [6 empty] | 6.5 | 9 |  |  | 8.6178 | 11.4333 |  |  | 14.1 | 17.4 |  |  | 20.2919 | 23.96
-15.3 | [6 empty] | 6.6286 | 9.1 |  |  | 8.7978 | 11.58 |  |  | 14.4571 | 17.58 |  |  | 20.5668 | 24.26
-15.2 | [6 empty] | 6.7 | 9.28 |  |  | 8.9578 | 11.72 |  |  | 14.7 | 17.96 |  |  | 20.9416 | 24.55
-15.1 |  |  | 6.3 | 8.8143 |  |  | 6.8 | 9.325 |  |  | 9.0977 | 11.9 |  |  | 14.875 | 18.2143 |  |  | 21.2415 | 24.9667
-15 |  |  | 6.4 | 8.98 |  |  | 6.9 | 9.5 |  |  | 9.3143 | 12.02 |  |  | 15.2 | 18.5 |  |  | 21.6913 | 25.28
-14.9 |  |  | 6.4875 | 9.0625 |  |  | 7.0857 | 9.6 |  |  | 9.4976 | 12.2 |  |  | 15.5167 | 18.825 |  |  | 22.0512 | 25.7333
-14.8 |  |  | 6.6 | 9.2 |  |  | 7.2 | 9.7 |  |  | 9.6262 | 12.3667 |  |  | 15.7429 | 19.0667 |  |  | 22.4411 | 26.06
-14.7 |  |  | 6.7 | 9.225 |  |  | 7.28 | 9.85 |  |  | 9.8975 | 12.66 |  |  | 16.125 | 19.35 |  |  | 22.8309 | 26.48
-14.6 |  |  | 6.7286 | 9.3333 |  |  | 7.35 | 10 |  |  | 10.0689 | 12.8 |  |  | 16.38 | 19.6667 |  |  | 23.1907 | 26.86
-14.5 |  |  | 6.9 | 9.5 |  |  | 7.5286 | 10.1 |  |  | 10.2475 | 13 |  |  | 16.675 | 20.025 |  |  | 23.7105 | 27.3667
-14.4 |  |  | 7.0444 | 9.6333 |  |  | 7.6 | 10.25 |  |  | 10.3724 | 13.18 |  |  | 17.0667 | 20.28 |  |  | 24.1681 | 27.86
-14.3 |  |  | 7.1 | 9.7 |  |  | 7.75 | 10.4 |  |  | 10.5974 | 13.4 |  |  | 17.3 | 20.7 |  |  | 24.7501 | 28.3333
-14.2 |  |  | 7.2 | 9.8 |  |  | 7.925 | 10.55 |  |  | 10.7773 | 13.575 |  |  | 17.75 | 21.08 |  |  | 25.3898 | 28.78
-14.1 |  |  | 7.3 | 10 |  |  | 8 | 10.75 |  |  | 10.9973 | 13.82 |  |  | 18.06 | 21.35 |  |  | 25.9696 | 29.6
-14 |  |  | 7.4 | 10.14 |  |  | 8.2 | 10.85 |  |  | 11.2472 | 14.02 |  |  | 18.45 | 21.66 |  |  | 26.8004 | 30.3
-13.9 |  |  | 7.5857 | 10.2 |  |  | 8.3571 | 11 |  |  | 11.4971 | 14.32 |  |  | 18.82 | 22.06 |  |  | 27.589 | 31.28
-13.8 |  |  | 7.7 | 10.3 |  |  | 8.5 | 11.2 |  |  | 11.7971 | 14.525 |  |  | 19.2 | 22.44 |  |  | 28.8599 | 32.22
-13.7 |  |  | 7.875 | 10.5 |  |  | 8.7 | 11.3333 |  |  | 11.977 | 14.82 |  |  | 19.6 | 22.8667 |  |  | 30.3079 | 33.64
-13.6 |  |  | 7.9333 | 10.66 |  |  | 8.8167 | 11.525 |  |  | 12.217 | 15 |  |  | 20 | 23.325 |  |  | 32.1996 | 35.5
-13.5 |  |  | 8.1 | 10.8429 |  |  | 9 | 11.725 |  |  | 12.4636 | 15.3 |  |  | 20.5 | 23.84 |  |  | 34.886 | 37.9
-13.4 |  |  | 8.2 | 10.95 |  |  | 9.2 | 11.88 |  |  | 12.7768 | 15.6 |  |  | 20.95 | 24.24 |  |  | 38.813 | 41.6
-13.3 |  |  | 8.3667 | 11.1 |  |  | 9.36 | 12 |  |  | 12.9968 | 15.88 |  |  | 21.5 | 24.75 |  |  | 44.0824 | 46.54
-13.2 |  |  | 8.5 | 11.2667 |  |  | 9.5 | 12.24 |  |  | 13.3467 | 16.26 |  |  | 22.075 | 25.38 |  |  | 51.2652 | 53.06
-13.1 |  |  | 8.625 | 11.4 |  |  | 9.8 | 12.4667 |  |  | 13.6766 | 16.5 |  |  | 22.72 | 26.12 |  |  | 60.2759 | 61.1
-13 |  |  | 8.8 | 11.64 |  |  | 10 | 12.68 |  |  | 13.9965 | 16.7714 |  |  | 23.5 | 26.84 |  |  | 70.072 | 70.4
-12.9 |  |  | 8.9 | 11.8 |  |  | 10.1 | 12.88 |  |  | 14.3298 | 17.1 |  |  | 24.5 | 27.8 |  |  | 80.6678 | 80.3
-12.8 |  |  | 9.175 | 11.9286 |  |  | 10.34 | 13.08 |  |  | 14.6563 | 17.56 |  |  | 25.65 | 28.9 |  |  | 90.7137 | 89.675
-12.7 |  |  | 9.28 | 12.1 |  |  | 10.5667 | 13.3 |  |  | 15.0629 | 17.9143 |  |  | 27.16 | 30.2 |  |  | 99.3605 | 97.7
-12.6 |  |  | 9.42 | 12.34 |  |  | 10.85 | 13.6 |  |  | 15.3962 | 18.25 |  |  | 29.0667 | 32.36 |  |  | 106.0147 | 104.65
-12.5 |  |  | 9.65 | 12.5 |  |  | 11.0857 | 13.875 |  |  | 15.8761 | 18.675 |  |  | 31.9 | 35.1 |  |  | 111.3556 | 109.24
-12.4 |  |  | 9.875 | 12.6667 |  |  | 11.3714 | 14.1333 |  |  | 16.296 | 19.14 |  |  | 36.075 | 38.7 |  |  | 114.8254 | 112.85
-12.3 |  |  | 10.02 | 12.85 |  |  | 11.7 | 14.4 |  |  | 16.8358 | 19.6857 |  |  | 41.82 | 44.5 |  |  | 117.173 | 114.9
-12.2 |  |  | 10.225 | 13.08 |  |  | 12 | 14.75 |  |  | 17.4956 | 20.25 |  |  | 50.4 | 52.78 |  |  | 118.7193 | 116.3
-12.1 |  |  | 10.425 | 13.275 |  |  | 12.3429 | 15.1 |  |  | 18.1205 | 20.9444 |  |  | 60.64 | 62.1 |  |  | 119.752 | 117.54
-12 |  |  | 10.7 | 13.5429 |  |  | 12.7 | 15.46 |  |  | 18.8953 | 21.64 |  |  | 72.12 | 72.8 |  |  | 120.5853 | 118.25
-11.9 |  |  | 10.9714 | 13.775 |  |  | 13.0667 | 15.8 |  |  | 19.8201 | 22.6 |  |  | 83.35 | 84.4 |  |  | 121.102 | 118.8
-11.8 |  |  | 11.2 | 14 |  |  | 13.48 | 16.25 |  |  | 21.0947 | 23.72 |  |  | 95.32 | 95.32 |  |  | 121.4177 | 119.3
-11.7 |  |  | 11.4857 | 14.3333 |  |  | 13.9571 | 16.76 |  |  | 22.6372 | 25.25 |  |  | 104.725 | 104.38 |  |  | 121.801 | 119.7
-11.6 |  |  | 11.84 | 14.6429 |  |  | 14.54 | 17.3667 |  |  | 24.7538 | 27.22 |  |  | 112.4 | 111.6 |  |  | 122.1939 | 119.9
-11.5 |  |  | 12.16 | 14.975 |  |  | 15.2571 | 18.0833 |  |  | 27.968 | 30.375 |  |  | 118.075 | 116.7 |  |  | 122.251 | 120.05
-11.4 |  |  | 12.6 | 15.4 |  |  | 16.2 | 19 |  |  | 32.6518 | 34.72 |  |  | 121.6 | 119.88 |  | 7.2571 | 122.431 | 120.1
-11.3 |  |  | 13.1 | 15.85 |  |  | 17.44 | 20.06 |  |  | 40.21 | 41.5 |  |  | 124.05 | 122.1 |  | 7.46 | 122.551 | 120.2
-11.2 |  |  | 13.6333 | 16.4833 |  |  | 18.8 | 21.6286 |  |  | 50.4207 | 51.2 |  |  | 125.36 | 123.5 |  | 7.55 | 122.651 | 120.4
-11.1 |  |  | 14.45 | 17.3 |  |  | 21.125 | 23.66 |  |  | 62.4644 | 63.25 |  |  | 126.4 | 124.4 |  | 7.6 | 122.7288 | 120.4
-11 |  |  | 15.5 | 18.2 |  |  | 24.4 | 26.7 |  |  | 76.6208 | 77.3286 |  | 5.5 | 127.16 | 125.18 |  | 7.84 | 122.751 | 120.5
-10.9 |  |  | 16.9333 | 19.5833 |  |  | 29.325 | 31.85 |  |  | 91.6971 | 90.3125 |  | 5.54 | 127.6 | 125.6 |  | 7.96 | 122.851 | 120.5
-10.8 |  |  | 19.1 | 21.6 |  |  | 37.2333 | 39.7875 |  |  | 104.1743 | 103.56 |  | 5.6 | 127.94 | 125.9 |  | 8.0571 | 122.851 | 120.5
-10.7 |  |  | 22.2 | 24.6857 |  |  | 48.4 | 50.48 |  |  | 115.171 | 113.925 |  | 5.7 | 128.2 | 126.1 |  | 8.2 | 122.751 | 120.5
-10.6 |  |  | 27.5 | 29.6 |  |  | 61.9 | 63.6 |  |  | 123.969 | 121.95 |  | 5.85 | 128.3667 | 126.3 |  | 8.325 | 122.731 | 120.5
-10.5 |  |  | 36.05 | 38.075 |  |  | 77.85 | 79.4 |  | 3.7 | 129.368 | 127.3 |  | 6 | 128.4333 | 126.3 |  | 8.5 | 122.651 | 120.5
-10.4 |  |  | 48.5125 | 50.1857 |  |  | 94.4 | 94.0714 |  | 3.7 | 132.867 | 130.6 |  | 6.1 | 128.5 | 126.4 |  | 8.6 | 122.591 | 120.5
-10.3 |  |  | 63.9 | 64.34 |  |  | 108.8 | 108.55 |  | 3.8 | 134.646 | 132.6286 |  | 6.1 | 128.5 | 126.32 |  | 8.85 | 122.511 | 120.4714
-10.2 |  |  | 81.84 | 81.6714 |  | 2.9 | 121.75 | 120.4 |  | 3.8667 | 135.916 | 133.8 |  | 6.25 | 128.4 | 126.3 |  | 9.075 | 122.451 | 120.4
-10.1 |  |  | 98.95 | 97.34 |  | 3 | 130.9 | 128.98 |  | 3.9 | 136.766 | 134.6 |  | 6.4 | 128.4 | 126.3 |  | 9.2 | 122.451 | 120.3
-10 |  | 2.8 | 114.9 | 113.4333 |  | 3 | 136.4 | 134.5 |  | 4 | 137.226 | 134.9857 |  | 6.54 | 128.3 | 126.2 |  | 9.4 | 122.351 | 120.2667
-9.9 |  | 2.9 | 127.3429 | 125.84 |  | 3 | 139.7667 | 137.9 |  | 4.1 | 137.466 | 135.2714 |  | 6.7 | 128.2 | 126.1 |  | 9.55 | 122.251 | 120.2
-9.8 |  | 2.9 | 136.9 | 134.8429 |  | 3.1 | 141.8 | 139.5571 |  | 4.16 | 137.466 | 135.3111 |  | 6.8 | 128.1 | 126 |  | 9.66 | 122.101 | 120.1
-9.7 |  | 2.9 | 142.4 | 139.98 |  | 3.1 | 142.8 | 140.46 |  | 4.2 | 137.466 | 135.35 |  | 6.9 | 127.8833 | 125.92 |  | 9.9 | 122.011 | 120
-9.6 |  | 3 | 145.0111 | 142.7 |  | 3.2 | 143.3286 | 141.0857 |  | 4.3 | 137.466 | 135.3 |  | 7.1 | 127.7 | 125.8 |  | 10.1 | 121.876 | 119.9
-9.5 |  | 3 | 146.475 | 144.2 |  | 3.2 | 143.5 | 141.2 |  | 4.4 | 137.326 | 135.2 |  | 7.25 | 127.48 | 125.7 |  | 10.2667 | 121.751 | 119.8
-9.4 |  | 3.1 | 147.2 | 145 |  | 3.3 | 143.5571 | 141.2 |  | 4.46 | 137.116 | 135.0125 |  | 7.4 | 127.35 | 125.575 |  | 10.575 | 121.601 | 119.7
-9.3 |  | 3.1 | 147.3 | 145 |  | 3.3 | 143.4857 | 141.15 |  | 4.6 | 136.966 | 134.8 |  | 7.6 | 127.3 | 125.4333 |  | 10.68 | 121.411 | 119.56
-9.2 |  | 3.1 | 147.3 | 145 |  | 3.4 | 143.3333 | 140.9571 |  | 4.7 | 136.666 | 134.5833 |  | 7.8 | 127.02 | 125.22 |  | 10.9333 | 121.252 | 119.475
-9.1 |  | 3.2 | 147.18 | 144.9667 |  | 3.5 | 143.1429 | 140.75 |  | 4.8 | 136.416 | 134.3 |  | 8 | 126.8 | 124.98 |  | 11.1571 | 121.172 | 119.22
-9 |  | 3.2 | 146.8143 | 144.62 |  | 3.5 | 142.82 | 140.5 |  | 4.85 | 136.016 | 134.0333 |  | 8.1 | 126.675 | 124.8 |  | 11.4 | 121.032 | 119.2
-8.9 |  | 3.3 | 146.46 | 144.25 |  | 3.6 | 142.4 | 140.1333 |  | 5.025 | 135.741 | 133.6 |  | 8.3 | 126.44 | 124.6 |  | 11.62 | 120.902 | 119.1
-8.8 |  | 3.35 | 145.9 | 143.875 |  | 3.7 | 141.8571 | 139.8 |  | 5.1 | 135.366 | 133.45 |  | 8.6 | 126.22 | 124.475 |  | 11.8 | 120.652 | 119.04
-8.7 |  | 3.4 | 145.45 | 143.26 |  | 3.7 | 141.4286 | 139.26 |  | 5.2 | 135.0517 | 133.1143 |  | 8.7 | 125.9833 | 124.24 |  | 12.16 | 120.392 | 118.9
-8.6 |  | 3.5 | 145.0375 | 142.7 |  | 3.8 | 140.9 | 138.9 |  | 5.35 | 134.666 | 132.7 |  | 8.9 | 125.775 | 124.025 |  | 12.56 | 120.352 | 118.8
-8.5 |  | 3.5 | 144.3571 | 142.075 |  | 3.9 | 140.45 | 138.35 |  | 5.5 | 134.216 | 132.4 |  | 9.2 | 125.54 | 123.9 |  | 12.725 | 120.152 | 118.7
-8.4 |  | 3.6 | 143.6333 | 141.48 |  | 4 | 140.025 | 137.88 |  | 5.6 | 133.867 | 132 |  | 9.4 | 125.24 | 123.6 |  | 13 | 120.002 | 118.5
-8.3 |  | 3.7 | 142.9143 | 140.96 |  | 4.1 | 139.48 | 137.35 |  | 5.8 | 133.417 | 131.48 |  | 9.6 | 125.05 | 123.42 |  | 13.28 | 119.7895 | 118.3
-8.2 |  | 3.7 | 142.3 | 140.36 |  | 4.2 | 138.85 | 136.92 |  | 5.9 | 132.967 | 131.05 |  | 9.8 | 124.84 | 123.125 |  | 13.75 | 119.652 | 118.1
-8.1 |  | 3.8 | 141.625 | 139.5667 |  | 4.3 | 138.225 | 136.2 |  | 6.02 | 132.567 | 130.84 |  | 10.15 | 124.5 | 122.92 |  | 13.98 | 119.5234 | 117.9
-8 |  | 3.9 | 140.975 | 138.94 |  | 4.4 | 137.82 | 135.78 |  | 6.3 | 132.067 | 130.4 |  | 10.4 | 124.2 | 122.6 |  | 14.3 | 119.192 | 117.7
-7.9 |  | 4 | 140.18 | 138.5 |  | 4.5 | 137.2 | 135.3 |  | 6.5 | 131.667 | 130.2 |  | 10.6333 | 123.95 | 122.42 |  | 14.78 | 119.052 | 117.6
-7.8 |  | 4.1 | 139.62 | 137.84 |  | 4.68 | 136.5667 | 134.825 |  | 6.6 | 131.307 | 129.74 |  | 11 | 123.55 | 122.2 |  | 15.2 | 118.872 | 117.375
-7.7 |  | 4.1 | 138.8 | 137.175 |  | 4.7667 | 136.025 | 134.22 |  | 6.75 | 130.817 | 129.2 |  | 11.25 | 123.3 | 121.92 |  | 15.725 | 118.553 | 117.16
-7.6 |  | 4.25 | 138.32 | 136.56 |  | 4.9 | 135.52 | 133.65 |  | 7 | 130.407 | 128.8333 |  | 11.6 | 123.06 | 121.8 | 3.1016 | 16.2714 | 118.353 | 116.9333
-7.5 |  | 4.4 | 137.7 | 135.98 |  | 5.0667 | 134.94 | 133.18 |  | 7.225 | 129.968 | 128.3667 |  | 11.9667 | 122.8 | 121.3 | 2.9015 | 17 | 118.253 | 116.64
-7.4 |  | 4.5 | 136.9 | 135.26 |  | 5.2 | 134.275 | 132.54 |  | 7.5 | 129.528 | 127.9 | 2.1 | 12.35 | 122.4 | 121 | 3.0015 | 17.88 | 118.078 | 116.52
-7.3 |  | 4.6 | 136.175 | 134.62 |  | 5.4 | 133.66 | 132.2 |  | 7.7333 | 129.068 | 127.5333 | 2.1 | 12.825 | 122.1333 | 120.78 | 3.0015 | 18.9222 | 117.7197 | 116.42
-7.2 |  | 4.75 | 135.7 | 134.1 |  | 5.5 | 133.12 | 131.38 | 1.3 | 8 | 128.588 | 127.2333 | 2.14 | 13.3 | 121.8 | 120.5667 | 3.1016 | 20.225 | 117.433 | 116.3
-7.1 |  | 4.9 | 135.05 | 133.52 | 1.1 | 5.7333 | 132.7333 | 130.925 | 1.3 | 8.3 | 128.168 | 126.65 | 2.2 | 13.875 | 121.575 | 120.18 | 3.1016 | 22 | 117.253 | 116.025
-7 | 0.9 | 5 | 134.375 | 132.86 | 1.1 | 5.9667 | 132.12 | 130.18 | 1.4 | 8.675 | 127.688 | 126.35 | 2.3 | 14.55 | 121.32 | 119.8667 | 3.1766 | 24.3 | 117.003 | 115.7667
-6.9 | 1 | 5.1333 | 133.76 | 132.22 | 1.1 | 6.125 | 131.425 | 129.85 | 1.4 | 8.95 | 127.128 | 125.9 | 2.3 | 15.4 | 120.7667 | 119.58 | 3.2266 | 27.675 | 116.773 | 115.6
-6.8 | 1 | 5.3667 | 133.1714 | 131.6 | 1.1 | 6.45 | 130.75 | 129.2667 | 1.4 | 9.4 | 126.893 | 125.3667 | 2.3 | 16.5 | 120.52 | 119.2 | 3.3016 | 32.56 | 116.493 | 115.5
-6.7 | 1 | 5.56 | 132.44 | 131 | 1.1 | 6.7 | 130.2667 | 128.82 | 1.5 | 9.8667 | 126.218 | 124.9667 | 2.4 | 17.8 | 120.2 | 118.98 | 3.3016 | 39.2 | 116.179 | 115.275
-6.6 | 1 | 5.7 | 131.74 | 130.32 | 1.1 | 6.975 | 129.8 | 128.24 | 1.5 | 10.3 | 125.929 | 124.58 | 2.4 | 19.5 | 119.84 | 118.75 | 3.4017 | 47.5 | 115.904 | 115.04
-6.5 | 1 | 5.96 | 131.18 | 129.875 | 1.1 | 7.3667 | 129.05 | 127.625 | 1.5 | 11.0333 | 125.449 | 124.2 | 2.5 | 22.025 | 119.4 | 118.45 | 3.5017 | 56.4 | 115.779 | 114.8
-6.4 | 1.1 | 6.1667 | 130.475 | 129.1 | 1.1 | 7.8 | 128.58 | 127.06 | 1.5 | 11.8 | 125.1023 | 123.7 | 2.5 | 25.65 | 119.22 | 118.2 | 3.5268 | 66.1 | 115.629 | 114.5833
-6.3 | 1.1 | 6.5 | 129.9 | 128.58 | 1.2 | 8.35 | 128.0667 | 126.55 | 1.6 | 12.8 | 124.5023 | 123.3 | 2.5667 | 30.75 | 118.9333 | 117.9 | 3.6018 | 75.44 | 115.334 | 114.4
-6.2 | 1.1 | 6.9 | 129.1714 | 127.9667 | 1.2 | 9 | 127.55 | 126.05 | 1.6 | 14.1 | 124.069 | 122.7 | 2.7 | 38 | 118.72 | 117.7 | 3.6018 | 84 | 115.029 | 114.175
-6.1 | 1.1 | 7.35 | 128.6 | 127.34 | 1.2 | 9.8 | 126.9 | 125.58 | 1.6 | 15.8 | 123.529 | 122.3 | 2.8 | 46.9 | 118.3333 | 117.425 | 3.7519 | 90.78 | 114.834 | 113.85
-6 | 1.1 | 7.95 | 128.02 | 126.6 | 1.2 | 10.9 | 126.42 | 125.02 | 1.7 | 18.4 | 123.169 | 122.1 | 2.8 | 57.2 | 118.0667 | 117.1 | 3.8269 | 95.8 | 114.729 | 113.7
-5.9 | 1.1 | 8.8 | 127.3857 | 126.25 | 1.2 | 12.3 | 125.8 | 124.4 | 1.7 | 22.2 | 122.669 | 121.65 | 2.8667 | 67.325 | 117.66 | 116.65 | 3.9019 | 99.64 | 114.554 | 113.6
-5.8 | 1.1 | 9.9 | 126.6714 | 125.7 | 1.2 | 14.6 | 125.25 | 123.98 | 1.8 | 27.9 | 122.1027 | 121.06 | 2.9 | 77.8 | 117.32 | 116.42 | 4.002 | 101.8857 | 114.254 | 113.42
-5.7 | 1.2 | 11.5 | 126.1333 | 125.075 | 1.3 | 17.7 | 124.6143 | 123.6 | 1.8 | 35.8 | 121.72 | 120.6333 | 3 | 86.425 | 117.08 | 116.14 | 4.1021 | 103.6 | 114.034 | 113.16
-5.6 | 1.2 | 14.1 | 125.625 | 124.44 | 1.3 | 22.8 | 124.1 | 123.1 | 1.9 | 46.1 | 121.09 | 120.24 | 3.1 | 93.85 | 116.78 | 115.8 | 4.1021 | 104.7 | 113.8293 | 112.86
-5.5 | 1.2 | 17.8 | 124.95 | 123.84 | 1.3333 | 30.6 | 123.5 | 122.38 | 1.9 | 57.9 | 120.77 | 119.65 | 3.2 | 98.975 | 116.56 | 115.68 | 4.2021 | 105.4333 | 113.655 | 112.7667
-5.4 | 1.2 | 24.0667 | 124.4 | 123.3 | 1.4 | 40.1 | 123.0333 | 121.8667 | 2 | 69.8 | 120.27 | 119.32 | 3.3 | 102.2 | 116.26 | 115.4333 | 4.4022 | 105.8 | 113.43 | 112.5667
-5.3 | 1.2 | 33.65 | 123.8286 | 122.82 | 1.4 | 51.8 | 122.3667 | 121.1 | 2 | 80.3333 | 119.89 | 119.08 | 3.4 | 104.275 | 115.8667 | 115.18 | 4.5023 | 106.35 | 113.255 | 112.4
-5.2 | 1.2 | 45.1333 | 123.15 | 122.22 | 1.4333 | 65.4 | 121.9286 | 120.8 | 2.1 | 90.2 | 119.39 | 118.45 | 3.5 | 105.8 | 115.5333 | 114.95 | 4.6023 | 106.7 | 113.055 | 112.08
-5.1 | 1.3 | 59.3 | 122.54 | 121.675 | 1.5 | 77.6 | 121.46 | 120.3 | 2.1 | 98.1 | 118.87 | 118.08 | 3.5 | 106.875 | 115.3 | 114.58 | 4.8024 | 106.9 | 112.695 | 111.85
-5 | 1.3 | 72.9 | 122.0143 | 121.1 | 1.5 | 88.6 | 120.9571 | 120.02 | 2.2 | 102.9 | 118.57 | 117.8 | 3.7 | 107.35 | 114.88 | 114.225 | 4.9225 | 107.1 | 112.595 | 111.64
-4.9 | 1.38 | 85.45 | 121.5857 | 120.45 | 1.6 | 98.1 | 120.5 | 119.46 | 2.3 | 106.1 | 118.21 | 117.48 | 3.8 | 108.15 | 114.6 | 113.8667 | 5.1025 | 107.3 | 112.255 | 111.4
-4.8 | 1.4 | 95.45 | 121.0333 | 120 | 1.6 | 104.6 | 120.0429 | 118.96 | 2.4 | 108 | 117.871 | 117.05 | 3.9 | 108.4 | 114.4 | 113.56 | 5.2026 | 107.3 | 112.055 | 111.225
-4.7 | 1.4 | 104.32 | 120.34 | 119.6 | 1.625 | 108.4 | 119.5 | 118.55 | 2.5 | 109.35 | 117.391 | 116.6 | 4 | 108.6 | 114.1 | 113.34 | 5.3026 | 107.4 | 111.755 | 110.9667
-4.6 | 1.5 | 109.3 | 119.9 | 119.14 | 1.7 | 110.525 | 118.9571 | 117.96 | 2.6 | 110.16 | 116.971 | 116.2 | 4.1 | 108.7 | 113.8667 | 113.16 | 5.5027 | 107.5 | 111.555 | 110.74
-4.5 | 1.5 | 112.1 | 119.44 | 118.6 | 1.8 | 111.9667 | 118.4 | 117.58 | 2.6 | 110.6 | 116.531 | 115.9 | 4.2 | 108.8 | 113.6 | 112.82 | 5.7028 | 107.56 | 111.3754 | 110.6667
-4.4 | 1.5 | 113.675 | 118.9222 | 118.08 | 1.8 | 113 | 117.9571 | 117.18 | 2.7 | 111 | 116.171 | 115.5 | 4.4 | 108.8 | 113.3 | 112.5 | 5.953 | 107.5 | 111.116 | 110.32
-4.3 | 1.6 | 114.55 | 118.4 | 117.625 | 1.9 | 113.4 | 117.5 | 116.85 | 2.9 | 111.1 | 115.791 | 114.9 | 4.6 | 108.9 | 112.8 | 112.3 | 6.103 | 107.5 | 110.816 | 110.025
-4.2 | 1.6 | 114.9667 | 117.8 | 117.24 | 2 | 113.6 | 117.0667 | 116.38 | 3 | 111.2 | 115.371 | 114.6 | 4.7 | 108.8 | 112.52 | 112.02 | 6.2031 | 107.48 | 110.516 | 109.86
-4.1 | 1.7 | 115.1 | 117.35 | 116.675 | 2.1 | 113.6 | 116.7 | 115.825 | 3.1 | 111.2 | 114.951 | 114.36 | 4.9 | 108.7 | 112.15 | 111.675 | 6.4032 | 107.4 | 110.281 | 109.58
-4 | 1.8 | 115.1 | 116.9 | 116.32 | 2.2 | 113.6 | 116.225 | 115.5 | 3.2667 | 111.2 | 114.671 | 113.94 | 5 | 108.6333 | 111.92 | 111.42 | 6.8034 | 107.3 | 110.176 | 109.4
-3.9 | 1.8 | 115 | 116.45 | 115.94 | 2.3 | 113.6 | 115.6 | 114.88 | 3.4 | 111.2 | 114.211 | 113.54 | 5.35 | 108.5 | 111.725 | 111.12 | 7.1036 | 107.2 | 109.896 | 109.225
-3.8 | 1.9 | 114.8333 | 116 | 115.38 | 2.4 | 113.5 | 115.2571 | 114.52 | 3.6 | 111 | 113.772 | 113.16 | 5.55 | 108.5 | 111.44 | 110.8333 | 7.5038 | 106.98 | 109.576 | 109
-3.7 | 2 | 114.5667 | 115.4429 | 114.84 | 2.5 | 113.4 | 114.92 | 114 | 3.8 | 110.95 | 113.372 | 112.72 | 5.9 | 108.4 | 111.08 | 110.54 | 7.904 | 106.875 | 109.376 | 108.6
-3.6 | 2.1 | 114.3333 | 115 | 114.45 | 2.7 | 113.1 | 114.36 | 113.5 | 4 | 110.8 | 113.0053 | 112.34 | 6.25 | 108.3 | 110.7429 | 110.24 | 8.5293 | 106.8 | 109.176 | 108.4333
-3.5 | 2.2 | 113.9333 | 114.475 | 113.9667 | 2.9 | 112.8 | 113.9429 | 113.18 | 4.35 | 110.55 | 112.6053 | 112.02 | 6.6667 | 108.2 | 110.4 | 109.96 | 9.2046 | 106.8 | 108.876 | 108.2
-3.4 | 2.4 | 113.5333 | 114.08 | 113.3 | 3.1 | 112.525 | 113.5857 | 112.72 | 4.65 | 110.3 | 112.272 | 111.64 | 7.15 | 108 | 110 | 109.74 | 10.1051 | 106.65 | 108.577 | 108.0667
-3.3 | 2.5 | 113.14 | 113.5714 | 112.96 | 3.38 | 112.2667 | 113.04 | 112.4333 | 5 | 110 | 111.822 | 111.3 | 7.8 | 107.9 | 109.74 | 109.38 | 11.2056 | 106.5333 | 108.3237 | 107.98
-3.2 | 2.7333 | 112.9 | 113 | 112.6333 | 3.7 | 111.85 | 112.6 | 112 | 5.5 | 109.8333 | 111.472 | 110.86 | 8.7 | 107.7 | 109.5571 | 109 | 12.7397 | 106.4 | 107.982 | 107.6667
-3.1 | 3.025 | 112.55 | 112.6 | 112.08 | 4.2 | 111.4667 | 112.175 | 111.5333 | 6.1333 | 109.6 | 110.952 | 110.5667 | 9.875 | 107.5 | 109.1 | 108.7 | 15.0075 | 106.28 | 107.782 | 107.45
-3 | 3.4333 | 112.1 | 112.1333 | 111.7333 | 4.75 | 111.2 | 111.7 | 111.1 | 6.9 | 109.3 | 110.672 | 110.18 | 11.4667 | 107.3 | 108.8 | 108.4 | 18.2091 | 106.1 | 107.432 | 107.16
-2.9 | 3.9 | 111.62 | 111.7 | 111.18 | 5.5 | 110.8667 | 111.38 | 110.82 | 8.1 | 109.1667 | 110.292 | 109.75 | 13.8 | 107.1 | 108.44 | 108.1 | 22.5113 | 106 | 107.177 | 106.875
-2.8 | 4.7 | 111.2667 | 111.3 | 110.7667 | 6.5333 | 110.45 | 110.8 | 110.2333 | 9.6 | 108.7 | 109.773 | 109.5 | 16.8 | 107 | 108.2 | 107.875 | 28.9645 | 105.825 | 106.937 | 106.7
-2.7 | 5.7 | 110.75 | 110.8 | 110.38 | 8 | 110 | 110.35 | 109.7 | 11.85 | 108.4667 | 109.453 | 109.075 | 21.8333 | 106.6 | 107.85 | 107.55 | 37.919 | 105.6429 | 106.717 | 106.44
-2.6 | 7.4 | 110.35 | 110.3 | 109.9333 | 10.2 | 109.6333 | 110.0286 | 109.45 | 15.1 | 108.1 | 109.013 | 108.6 | 28.2 | 106.45 | 107.6 | 107.3 | 46.7233 | 105.5 | 106.497 | 106.18
-2.5 | 9.8 | 109.8333 | 109.9667 | 109.6 | 13.45 | 109.3 | 109.65 | 109.0667 | 20.1667 | 107.8667 | 108.6873 | 108.24 | 36.9 | 106.22 | 107.12 | 107.02 | 56.8034 | 105.3556 | 106.2078 | 106.04
-2.4 | 13.9 | 109.56 | 109.3 | 109.1 | 18.6667 | 109.0333 | 109.1714 | 108.72 | 27.4 | 107.5 | 108.248 | 108 | 47.4 | 106.0333 | 106.8429 | 106.625 | 66.9668 | 105.28 | 105.958 | 105.78
-2.3 | 20.25 | 109.2 | 108.8667 | 108.5 | 26.55 | 108.6 | 108.7 | 108.24 | 37 | 107.2667 | 107.973 | 107.7 | 58.3 | 105.8 | 106.6286 | 106.4 | 76.3715 | 105.0429 | 105.858 | 105.54
-2.2 | 29.2 | 108.82 | 108.6167 | 108.16 | 36.1 | 108.2333 | 108.4 | 107.82 | 47.8667 | 106.9 | 107.5286 | 107.25 | 68.5 | 105.6 | 106.34 | 106.1 | 84.2421 | 104.8 | 105.598 | 105.4
-2.1 | 40.1 | 108.25 | 108.15 | 107.9 | 48.1 | 107.8 | 107.8667 | 107.5 | 59.45 | 106.7 | 107.1159 | 106.8 | 78.4 | 105.42 | 106.15 | 105.8667 | 90.0784 | 104.5833 | 105.338 | 105
-2 | 52.925 | 107.8 | 107.72 | 107.42 | 60.2 | 107.35 | 107.46 | 107.1 | 70.7333 | 106.5 | 106.8286 | 106.52 | 86.2 | 105.1333 | 105.825 | 105.6667 | 93.947 | 104.4 | 105.108 | 104.8
-1.9 | 65.2 | 107.4 | 107.3333 | 107.15 | 71 | 107.1 | 107.06 | 106.7333 | 80.475 | 106 | 106.433 | 106.125 | 91.6 | 105.02 | 105.48 | 105.34 | 97.1486 | 104.1833 | 104.798 | 104.62
-1.8 | 77.05 | 107 | 106.78 | 106.66 | 81.575 | 106.8 | 106.78 | 106.3 | 88.1 | 105.8 | 106.1235 | 105.94 | 95.7 | 104.8 | 105.1833 | 105.1 | 98.7494 | 103.975 | 104.558 | 104.375
-1.7 | 86.4 | 106.6 | 106.44 | 106.35 | 89.4667 | 106.5 | 106.5 | 105.9667 | 93.65 | 105.6 | 105.799 | 105.5667 | 97.9 | 104.62 | 104.9 | 104.78 | 100.15 | 103.8625 | 104.398 | 104.12
-1.6 | 93.7 | 106.3667 | 106.1 | 106.12 | 94.6333 | 106.0333 | 105.86 | 105.5 | 96.9667 | 105.125 | 105.414 | 105.24 | 99.6 | 104.45 | 104.5714 | 104.38 | 100.9002 | 103.6 | 104.058 | 103.8333
-1.5 | 97.8667 | 105.8 | 105.72 | 105.56 | 97.8667 | 105.675 | 105.58 | 105.3 | 98.8667 | 104.8 | 105.0026 | 104.92 | 100.6333 | 104.15 | 104.3 | 104.1333 | 101.351 | 103.4167 | 103.8784 | 103.6
-1.4 | 100.3 | 105.5667 | 105.2333 | 105.18 | 99.8667 | 105.3 | 105.2 | 104.95 | 100.2 | 104.525 | 104.774 | 104.48 | 101.26 | 103.9 | 104.1667 | 103.9667 | 101.851 | 103.2 | 103.559 | 103.4
-1.3 | 101.4 | 105.12 | 104.76 | 104.84 | 101 | 104.9 | 104.8 | 104.52 | 101.2333 | 104.4 | 104.4073 | 104.2 | 101.7 | 103.55 | 103.8 | 103.6667 | 102.051 | 103.1 | 103.319 | 103.28
-1.2 | 102.24 | 104.7 | 104.4 | 104.42 | 101.6333 | 104.65 | 104.5 | 104.14 | 101.625 | 103.925 | 104.074 | 103.9 | 101.975 | 103.2 | 103.48 | 103.3667 | 102.051 | 102.92 | 103.059 | 103.0333
-1.1 | 102.5 | 104.325 | 104.06 | 104.05 | 101.9 | 104.2 | 104.0667 | 103.925 | 101.8 | 103.65 | 103.7311 | 103.575 | 102 | 103.04 | 103.075 | 102.98 | 102.051 | 102.7 | 102.7019 | 102.68
-1 | 102.6 | 103.85 | 103.7667 | 103.78 | 102 | 103.7 | 103.65 | 103.48 | 101.9 | 103.3333 | 103.374 | 103.225 | 102 | 102.8667 | 102.8 | 102.64 | 102.051 | 102.42 | 102.459 | 102.55
-0.9 | 102.6 | 103.38 | 103.3 | 103.275 | 102.1 | 103.3 | 103.35 | 102.9833 | 101.9 | 103 | 103.0597 | 102.95 | 101.95 | 102.65 | 102.5 | 102.5 | 101.851 | 102.2 | 102.379 | 102.225
-0.8 | 102.5 | 103.2 | 103.04 | 102.94 | 102.05 | 102.9 | 102.9 | 102.775 | 101.8 | 102.58 | 102.694 | 102.58 | 101.9 | 102.3 | 102.26 | 102.25 | 101.751 | 102.02 | 101.919 | 102.05
-0.7 | 102.3 | 102.775 | 102.6 | 102.62 | 101.9333 | 102.6 | 102.6 | 102.54 | 101.7 | 102.2 | 102.3311 | 102.2 | 101.75 | 101.96 | 101.88 | 102.02 | 101.651 | 101.8 | 101.659 | 101.68
-0.6 | 102.04 | 102.4 | 102.18 | 102.25 | 101.75 | 102.3 | 102.35 | 102.05 | 101.6 | 101.75 | 101.9248 | 101.8 | 101.5333 | 101.7333 | 101.75 | 101.66 | 101.551 | 101.5167 | 101.509 | 101.5
-0.5 | 101.75 | 102 | 101.7667 | 101.7 | 101.5333 | 101.8667 | 101.92 | 101.7333 | 101.4 | 101.5 | 101.475 | 101.52 | 101.4 | 101.5 | 101.44 | 101.42 | 101.251 | 101.2 | 101.279 | 101.3
-0.4 | 101.4 | 101.5 | 101.4 | 101.4667 | 101.35 | 101.5333 | 101.4 | 101.45 | 101.2 | 101.15 | 101.325 | 101.2 | 101.225 | 101.125 | 101.06 | 101.15 | 101.0005 | 101.04 | 101.06 | 100.975
-0.3 | 101.05 | 101.2 | 100.95 | 101.14 | 101 | 101.2333 | 101.1 | 101.04 | 100.9 | 100.9333 | 100.975 | 100.92 | 101 | 100.9 | 100.85 | 100.76 | 100.85 | 100.7 | 100.8267 | 100.6
-0.2 | 100.8 | 100.8 | 100.58 | 100.8 | 100.7 | 100.8333 | 100.75 | 100.78 | 100.4667 | 100.55 | 100.7 | 100.62 | 100.6333 | 100.6667 | 100.38 | 100.5 | 100.55 | 100.4571 | 100.48 | 100.375
-0.1 | 100.55 | 100.35 | 100.3 | 100.32 | 100.3667 | 100.45 | 100.32 | 100.42 | 100.3 | 100.2333 | 100.275 | 100.24 | 100.3333 | 100.3 | 100.2 | 100.2333 | 100.35 | 100.25 | 100.26 | 100.2
0 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100 | 100
0.1 | 99.6 | 99.7 | 99.58 | 99.64 | 99.55 | 99.625 | 99.72 | 99.675 | 99.5 | 99.7333 | 99.5751 | 99.65 | 99.7 | 99.7 | 99.6 | 99.68 | 99.7499 | 99.75 | 99.6401 | 99.76
0.2 | 99.34 | 99.2 | 99.24 | 99.32 | 99.175 | 99.35 | 99.3333 | 99.3 | 99.1 | 99.3 | 99.2002 | 99.2 | 99.4333 | 99.36 | 99.28 | 99.45 | 99.5498 | 99.46 | 99.4402 | 99.45
0.3 | 98.88 | 98.9 | 98.94 | 99.06 | 98.75 | 99.05 | 99.0167 | 98.9286 | 98.9 | 98.9333 | 98.8753 | 99.08 | 99.05 | 99.15 | 98.98 | 99.14 | 99.3497 | 99.2 | 99.2003 | 98.94
0.4 | 98.4 | 98.5667 | 98.4667 | 98.6333 | 98.45 | 98.6 | 98.68 | 98.65 | 98.5 | 98.45 | 98.5254 | 98.75 | 98.7667 | 98.92 | 98.76 | 98.9 | 99.0995 | 98.96 | 98.7805 | 98.8
0.5 | 97.95 | 98.1 | 98.14 | 98.3 | 98.1 | 98.2667 | 98.3333 | 98.3857 | 97.9 | 98.2 | 98.1355 | 98.38 | 98.45 | 98.55 | 98.42 | 98.5 | 98.7494 | 98.7333 | 98.5406 | 98.7
0.6 | 97.45 | 97.7 | 97.8 | 97.9667 | 97.35 | 98.0333 | 97.98 | 97.9667 | 97.5 | 97.82 | 97.8006 | 98.125 | 98.0667 | 98.2 | 98 | 98.2 | 98.2825 | 98.36 | 98.3407 | 98.44
0.7 | 97.0333 | 97.4 | 97.55 | 97.64 | 96.9 | 97.6667 | 97.6667 | 97.58 | 97.1 | 97.55 | 97.6356 | 97.68 | 97.7 | 97.8 | 97.7 | 97.9 | 98.049 | 98.2 | 98.0608 | 98.18
0.8 | 96.4333 | 97 | 97.1333 | 97.2667 | 96.3 | 97.2667 | 97.3 | 97.2333 | 96.6 | 97.26 | 97.2757 | 97.4 | 97.0667 | 97.5 | 97.4 | 97.675 | 97.5988 | 97.85 | 97.7209 | 98.0333
0.9 | 95.9 | 96.6 | 96.7714 | 96.8857 | 95.8 | 97 | 96.875 | 96.98 | 96.1 | 96.95 | 96.9957 | 97.05 | 96.7667 | 97.2 | 97.26 | 97.35 | 97.1486 | 97.6 | 97.561 | 97.76
1 | 95.26 | 96.4 | 96.5 | 96.625 | 95.1 | 96.55 | 96.6 | 96.6333 | 95.2667 | 96.48 | 96.6259 | 96.78 | 96.3 | 96.9 | 97.025 | 97.05 | 96.8484 | 97.3857 | 97.3411 | 97.36
1.1 | 94.2667 | 96 | 96.1571 | 96.34 | 94.2667 | 96.15 | 96.24 | 96.28 | 94.7 | 96.1667 | 96.196 | 96.4 | 95.7 | 96.6 | 96.5333 | 96.7 | 96.2481 | 97.02 | 97.0612 | 97.2
1.2 | 93.225 | 95.5 | 95.7 | 95.9 | 93.3 | 95.7 | 95.88 | 96 | 93.9333 | 95.7 | 95.896 | 96.025 | 94.9 | 96.3 | 96.34 | 96.5 | 95.6478 | 96.6333 | 96.7946 | 96.875
1.3 | 92.0333 | 95.1 | 95.25 | 95.7 | 91.9 | 95.4667 | 95.56 | 95.575 | 92.7 | 95.45 | 95.6511 | 95.8 | 94.3 | 96 | 96.0167 | 96.3 | 95.1476 | 96.35 | 96.5281 | 96.64
1.4 | 90.08 | 94.8 | 94.88 | 95.3 | 90.35 | 95.075 | 95.34 | 95.35 | 91.45 | 95.16 | 95.3262 | 95.5 | 93.3 | 95.55 | 95.72 | 95.9667 | 94.5473 | 96.08 | 96.1615 | 96.34
1.5 | 87.3 | 94.525 | 94.6286 | 94.96 | 88.1 | 94.75 | 94.95 | 94.92 | 89.8 | 94.8333 | 95.0163 | 95.14 | 92.4 | 95.05 | 95.3 | 95.66 | 93.8469 | 95.85 | 95.8867 | 96.1
1.6 | 83.2 | 94 | 94.3 | 94.65 | 84.65 | 94.28 | 94.54 | 94.55 | 87.35 | 94.28 | 94.7163 | 94.9167 | 90.9 | 94.8 | 95.08 | 95.3 | 93.2466 | 95.56 | 95.5818 | 95.8
1.7 | 76.3 | 93.8 | 94 | 94.2 | 79.74 | 93.85 | 94.175 | 94.28 | 83.9 | 94 | 94.2514 | 94.6 | 89.2 | 94.45 | 94.8667 | 95.1 | 91.8459 | 95.18 | 95.3419 | 95.525
1.8 | 68.1 | 93.4 | 93.6125 | 93.98 | 72.4 | 93.55 | 93.9 | 93.9667 | 78.7333 | 93.7 | 93.9365 | 94.12 | 87.05 | 94.15 | 94.58 | 94.75 | 90.2451 | 94.88 | 95.1334 | 95.2
1.9 | 57.45 | 92.9333 | 93.26 | 93.68 | 63.24 | 93.2 | 93.55 | 93.48 | 71.75 | 93.1667 | 93.5566 | 93.85 | 83.3667 | 93.7667 | 94.32 | 94.42 | 88.5443 | 94.44 | 94.9192 | 95.02
2 | 46.8 | 92.6 | 92.96 | 93.32 | 53.8 | 92.85 | 93.1 | 93.32 | 63.2 | 92.86 | 93.2367 | 93.58 | 78.4 | 93.4333 | 93.94 | 94.2 | 86.043 | 94.2 | 94.3623 | 94.7
2.1 | 34.9667 | 92.3 | 92.62 | 92.8 | 43.3 | 92.38 | 92.8 | 92.925 | 53.8 | 92.5 | 92.9435 | 93.3 | 72.2 | 93.1667 | 93.65 | 93.875 | 82.6413 | 93.9 | 94.0909 | 94.475
2.2 | 25.6 | 91.7 | 92.28 | 92.5286 | 32.9 | 92.075 | 92.5 | 92.6333 | 43.575 | 91.875 | 92.6569 | 93.04 | 64.15 | 92.7 | 93.34 | 93.6 | 77.6388 | 93.6 | 93.8875 | 94.2
2.3 | 17.7 | 91.35 | 91.9143 | 92.3 | 24.3 | 91.7 | 92.2 | 92.3 | 33.4333 | 91.7 | 92.4169 | 92.66 | 54.9 | 92.3667 | 92.95 | 93.26 | 71.6858 | 93.24 | 93.5626 | 93.86
2.4 | 12.1 | 90.95 | 91.5333 | 91.925 | 17.35 | 91.35 | 91.82 | 91.9429 | 24.975 | 91.275 | 91.957 | 92.26 | 44.7 | 92.1 | 92.7 | 92.9333 | 64.1321 | 92.9 | 93.4376 | 93.5667
2.5 | 8.72 | 90.5667 | 91.24 | 91.6 | 12.5333 | 90.8 | 91.5444 | 91.6 | 18.4 | 90.8 | 91.6771 | 91.875 | 36.15 | 91.5667 | 92.45 | 92.6286 | 54.9275 | 92.5 | 93.1627 | 93.325
2.6 | 6.5 | 90.25 | 90.775 | 91.25 | 9.3 | 90.5 | 91.2 | 91.36 | 14 | 90.5 | 91.3772 | 91.64 | 27.8 | 91.2667 | 92.08 | 92.32 | 46.2232 | 92.0333 | 92.8029 | 93.025
2.7 | 5.125 | 89.85 | 90.6 | 90.8875 | 7.25 | 90.2 | 90.72 | 91.05 | 10.875 | 90.2 | 91.0772 | 91.325 | 21.5667 | 90.8 | 91.78 | 91.9 | 37.2186 | 91.8 | 92.6129 | 92.7
2.8 | 4.2 | 89.4 | 90.25 | 90.475 | 5.92 | 89.7 | 90.3667 | 90.55 | 8.8667 | 89.6 | 90.7573 | 91.0667 | 16.6667 | 90.4667 | 91.52 | 91.775 | 29.6148 | 91.3333 | 92.3031 | 92.5143
2.9 | 3.55 | 88.9333 | 89.8667 | 90.1667 | 5 | 89.1 | 90.125 | 90.3 | 7.4 | 89.2 | 90.3774 | 90.775 | 13.35 | 90.1 | 91.2 | 91.5 | 23.1616 | 91.1 | 91.9632 | 92.22
3 | 3.1 | 88.7 | 89.45 | 89.925 | 4.275 | 88.7 | 89.8 | 89.95 | 6.45 | 88.85 | 89.9775 | 90.4 | 11.2 | 89.7333 | 90.85 | 91.1333 | 18.2092 | 90.6 | 91.5883 | 91.8143
3.1 | 2.7 | 88.2667 | 89.1 | 89.5143 | 3.85 | 88.5 | 89.425 | 89.6 | 5.6 | 88.3 | 89.7276 | 89.96 | 9.6 | 89.3 | 90.5 | 90.8429 | 14.9075 | 90.25 | 91.2635 | 91.6
3.2 | 2.4 | 87.8 | 88.825 | 89.25 | 3.4 | 88 | 89.0667 | 89.25 | 5.12 | 87.8 | 89.3277 | 89.7 | 8.4 | 88.85 | 90.18 | 90.6 | 12.5396 | 89.9 | 91.0969 | 91.4333
3.3 | 2.2 | 87.4667 | 88.4857 | 88.9 | 3.0333 | 87.6 | 88.84 | 88.975 | 4.6 | 87.6 | 89.0777 | 89.4 | 7.525 | 88.4333 | 89.9333 | 90.2667 | 10.8554 | 89.44 | 90.7637 | 91.175
3.4 | 2.0333 | 86.95 | 88.18 | 88.5625 | 2.8 | 87.15 | 88.5143 | 88.68 | 4.25 | 87 | 88.8278 | 89.05 | 6.8 | 87.975 | 89.58 | 89.9 | 9.8049 | 89.075 | 90.3639 | 90.9
3.5 | 1.9 | 86.5667 | 87.6714 | 88.18 | 2.6333 | 86.6 | 88.1 | 88.32 | 4 | 86.55 | 88.4579 | 88.76 | 6.4 | 87.55 | 89.3 | 89.66 | 8.8044 | 88.66 | 90.0925 | 90.65
3.6 | 1.8 | 86.1 | 87.28 | 87.75 | 2.4 | 86.2 | 87.6714 | 87.94 | 3.7 | 86.1 | 88.153 | 88.4 | 5.9 | 87 | 89.1 | 89.38 | 8.1041 | 88.35 | 89.7641 | 90.1857
3.7 | 1.7333 | 85.6 | 87 | 87.5 | 2.3 | 85.6 | 87.4 | 87.68 | 3.5 | 85.7333 | 87.6981 | 88.1333 | 5.6333 | 86.7667 | 88.65 | 89.02 | 7.4704 | 87.88 | 89.5392 | 90.04
3.8 | 1.7 | 85.1 | 86.7667 | 87.2571 | 2.2 | 85.2 | 87.0125 | 87.3 | 3.2667 | 85.2 | 87.3115 | 87.675 | 5.4 | 86.3 | 88.3 | 88.5857 | 7.2036 | 87.4667 | 89.3199 | 89.7667
3.9 | 1.6 | 84.4333 | 86.4 | 86.9 | 2.05 | 84.575 | 86.7 | 87.05 | 3.1 | 84.7 | 87.0782 | 87.42 | 5.1 | 85.7333 | 87.96 | 88.3857 | 6.8034 | 87.12 | 89.0894 | 89.5
4 | 1.5 | 84 | 86.0571 | 86.575 | 2 | 84 | 86.3667 | 86.7 | 3 | 84.15 | 86.7783 | 87.0667 | 4.9 | 85.26 | 87.6 | 88.08 | 6.3832 | 86.7 | 88.8645 | 89.0667
4.1 | 1.4 | 83.1333 | 85.64 | 86.1667 | 1.85 | 83.425 | 86.05 | 86.34 | 2.9 | 83.55 | 86.4184 | 86.66 | 4.7 | 84.9333 | 87.4 | 87.9111 | 6.103 | 86.14 | 88.5046 | 88.8
4.2 | 1.4 | 82.4667 | 85.3714 | 85.8429 | 1.8 | 82.65 | 85.7625 | 86 | 2.7 | 82.9667 | 86.0785 | 86.5 | 4.45 | 84.5 | 87.1 | 87.4 | 5.9029 | 85.85 | 88.2647 | 88.5571
4.3 | 1.4 | 81.3 | 85 | 85.44 | 1.7 | 81.7 | 85.425 | 85.6857 | 2.6 | 82.25 | 85.7119 | 86.06 | 4.35 | 83.75 | 86.75 | 87.15 | 5.8029 | 85.4 | 87.8523 | 88.3
4.4 | 1.3 | 80.15 | 84.7333 | 85.075 | 1.6 | 80.6 | 85.0778 | 85.3 | 2.5 | 81.6 | 85.412 | 85.7667 | 4.1 | 83.3 | 86.45 | 86.86 | 5.5027 | 84.8 | 87.54 | 87.95
4.5 | 1.3 | 78.35 | 84.45 | 84.8429 | 1.6 | 79.6 | 84.8143 | 85 | 2.4 | 80.8 | 85.1987 | 85.5 | 4 | 82.75 | 86.15 | 86.6143 | 5.3026 | 84.34 | 87.2651 | 87.54
4.6 | 1.2 | 76 | 84 | 84.6286 | 1.5 | 78 | 84.5 | 84.76 | 2.3 | 79.8 | 84.7588 | 85.08 | 3.8 | 82.3 | 85.8714 | 86.28 | 5.1025 | 83.9 | 86.9852 | 87.4
4.7 | 1.2 | 71.65 | 83.72 | 84.2667 | 1.475 | 76 | 84.08 | 84.35 | 2.2 | 78.6 | 84.4589 | 84.8 | 3.7 | 81.525 | 85.5286 | 85.95 | 4.9358 | 83.35 | 86.6511 | 87.1143
4.8 | 1.1 | 65.8 | 83.4 | 83.8 | 1.4 | 72.3 | 83.7333 | 83.9 | 2.1 | 77.35 | 84.179 | 84.4 | 3.6 | 80.9 | 85.1714 | 85.725 | 4.8024 | 83.04 | 86.2905 | 86.72
4.9 | 1.1 | 58.5 | 83.02 | 83.5429 | 1.375 | 67.5 | 83.375 | 83.6333 | 2.1 | 75.45 | 83.839 | 84.14 | 3.5 | 80.2 | 84.92 | 85.3857 | 4.7024 | 82.5625 | 86.0085 | 86.3714
5 | 1.1 | 49.1 | 82.7 | 83.16 | 1.3 | 60.95 | 83.1143 | 83.275 | 2 | 72.7 | 83.3792 | 83.88 | 3.4 | 79.5 | 84.4857 | 85.025 | 4.6023 | 82 | 85.7657 | 86.16
5.1 | 1.1 | 40.2 | 82.3667 | 82.9143 | 1.3 | 52.95 | 82.68 | 82.96 | 1.9 | 69.15 | 83.1392 | 83.5 | 3.3 | 78.4 | 84.32 | 84.6 | 4.5523 | 81.52 | 85.4858 | 85.82
5.2 | 1 | 31.3 | 82.08 | 82.55 | 1.3 | 45.2 | 82.3429 | 82.7 | 1.9 | 63.6 | 82.6793 | 83.2 | 3.2 | 77.3 | 83.9 | 84.4 | 4.4022 | 80.7333 | 85.2059 | 85.5
5.3 | 1 | 23.3 | 81.7 | 82.2 | 1.2 | 36.3 | 82 | 82.275 | 1.8 | 56.9333 | 82.3594 | 82.92 | 3.1 | 75.8 | 83.62 | 84.02 | 4.3022 | 80.2 | 84.7911 | 85.22
5.4 | 1 | 17.1 | 81.3 | 81.85 | 1.2 | 28.5 | 81.7143 | 81.9333 | 1.8 | 49.3 | 82.1128 | 82.56 | 3 | 74.1 | 83.35 | 83.6667 | 4.1021 | 79.28 | 84.5662 | 85.1
5.5 | 1 | 12.7 | 81.05 | 81.5143 | 1.2 | 21.5 | 81.3 | 81.66 | 1.7 | 41.0667 | 81.8395 | 82.175 | 2.9 | 71.6 | 83 | 83.35 | 4.002 | 78.5 | 84.2263 | 84.6625
5.6 | 1 | 10.2 | 80.68 | 81.1 | 1.1 | 16.4 | 80.98 | 81.25 | 1.62 | 33.58 | 81.3796 | 81.86 | 2.8 | 68.25 | 82.6667 | 82.9667 | 3.9019 | 77.62 | 83.8265 | 84.28
5.7 | 0.9 | 8.5 | 80.32 | 80.75 | 1.1 | 12.8 | 80.7 | 80.88 | 1.6 | 26.6 | 80.9598 | 81.5 | 2.8 | 63.4 | 82.3 | 82.75 | 3.8019 | 76.625 | 83.5466 | 83.95
5.8 | 0.9 | 7.4 | 79.95 | 80.5 | 1.1 | 10.7 | 80.36 | 80.66 | 1.5 | 21.1 | 80.6398 | 81.26 | 2.7 | 57.45 | 81.96 | 82.4286 | 3.8019 | 75.05 | 83.2167 | 83.6857
5.9 | 0.9 | 6.6 | 79.7 | 80.2222 | 1.1 | 9.3 | 80 | 80.3 | 1.5 | 16.7333 | 80.4299 | 80.9 | 2.6 | 50.35 | 81.64 | 82.1 | 3.7018 | 73.38 | 82.9069 | 83.44
6 | 0.9 | 6.0667 | 79.3 | 79.7 | 1 | 8.2 | 79.64 | 79.94 | 1.5 | 13.975 | 80.08 | 80.5 | 2.5 | 43.05 | 81.35 | 81.7667 | 3.6018 | 70.75 | 82.6003 | 83.0857
6.1 | 0.9 | 5.675 | 78.975 | 79.5 | 1 | 7.4 | 79.3333 | 79.6 | 1.4 | 12.1667 | 79.7201 | 80.08 | 2.5 | 35.6667 | 80.8667 | 81.5 | 3.5017 | 67.5 | 82.3671 | 82.75
6.2 | 0.9 | 5.3 | 78.72 | 79.12 | 1 | 6.9 | 78.94 | 79.1667 | 1.4 | 10.925 | 79.3802 | 79.86 | 2.4 | 29.3 | 80.54 | 81.2714 | 3.4017 | 62.9 | 82.0072 | 82.5571
6.3 | 0.9 | 5 | 78.32 | 78.76 | 1 | 6.4 | 78.54 | 79 | 1.4 | 10 | 79.0469 | 79.5 | 2.4 | 24.25 | 80.28 | 80.86 | 3.3016 | 57.475 | 81.634 | 82.2429
6.4 | 0.8 | 4.8 | 77.95 | 78.4333 | 1 | 6 | 78.25 | 78.6 | 1.3 | 9.2 | 78.6803 | 79.15 | 2.3 | 20.1 | 79.96 | 80.6 | 3.2016 | 50.94 | 81.3175 | 81.9
6.5 | 0.8 | 4.6 | 77.64 | 78.125 | 0.9 | 5.7 | 77.94 | 78.4 | 1.3 | 8.6 | 78.2947 | 78.85 | 2.3 | 17.4667 | 79.62 | 80.18 | 3.1016 | 44.25 | 81.0176 | 81.55
6.6 | 0.8 | 4.4333 | 77.3 | 77.7857 | 0.9 | 5.4 | 77.55 | 77.92 | 1.3 | 8.14 | 77.8805 | 78.45 | 2.2 | 15.6 | 79.2 | 79.7 | 3.1016 | 37.44 | 80.6677 | 81.2
6.7 | 0.8 | 4.3 | 76.8667 | 77.4 | 0.9 | 5.2 | 77.24 | 77.65 | 1.225 | 7.75 | 77.6806 | 78.05 | 2.2 | 14.25 | 78.9333 | 79.38 | 3.0515 | 31.62 | 80.3179 | 81.04
6.8 | 0.8 | 4.15 | 76.58 | 77.12 | 0.9 | 5 | 76.86 | 77.325 | 1.2 | 7.4 | 77.2557 | 77.725 | 2.1 | 13.3 | 78.56 | 79.0429 | 3.0015 | 26.9667 | 80.068 | 80.82
6.9 | 0.8 | 4.05 | 76.225 | 76.75 | 0.9 | 4.825 | 76.48 | 77 | 1.2 | 7.1 | 77.0208 | 77.5 | 2.1 | 12.5 | 78.2 | 78.7333 | 2.9015 | 22.9 | 79.7681 | 80.3857
7 | 0.8 | 4 | 75.84 | 76.5333 | 0.9 | 4.6 | 76.25 | 76.4667 | 1.2 | 6.875 | 76.5666 | 77.2 | 2 | 11.9 | 77.975 | 78.45 | 2.9015 | 20.2333 | 79.4682 | 79.875
7.1 |  | 3.9 | 75.58 | 76.1 | 0.9 | 4.5 | 75.9667 | 76.2 | 1.2 | 6.6 | 76.2809 | 76.7143 | 2 | 11.3 | 77.5 | 78.06 | 2.8014 | 18.25 | 79.0934 | 79.6429
7.2 |  | 3.8 | 75.275 | 75.8 |  | 4.4 | 75.475 | 75.9333 | 1.2 | 6.375 | 76.0382 | 76.375 | 2 | 10.9 | 77.2333 | 77.72 | 2.7014 | 16.875 | 78.8685 | 79.3
7.3 |  | 3.7 | 74.85 | 75.44 |  | 4.2333 | 75.3 | 75.58 |  | 6.2 | 75.5811 | 76.1143 | 1.9 | 10.5 | 76.8667 | 77.3 | 2.6013 | 15.9 | 78.5186 | 79.02
7.4 |  | 3.6 | 74.5333 | 75.1 |  | 4.1 | 74.9 | 75.1 |  | 6 | 75.2384 | 75.68 |  | 10.2 | 76.44 | 77.05 | 2.6013 | 15 | 78.0887 | 78.68
7.5 |  | 3.5 | 74.24 | 74.8 |  | 4 | 74.55 | 74.88 |  | 5.8 | 74.7813 | 75.42 |  | 9.9 | 76.22 | 76.75 | 2.6013 | 14.35 | 77.7939 | 78.45
7.6 |  | 3.4 | 73.88 | 74.42 |  | 3.9 | 74.18 | 74.55 |  | 5.65 | 74.4939 | 75.15 |  | 9.6 | 75.96 | 76.38 |  | 13.76 | 77.489 | 78.075
7.7 |  | 3.4 | 73.5 | 74.04 |  | 3.8 | 73.8 | 74.12 |  | 5.5 | 74.2314 | 74.6857 |  | 9.3 | 75.575 | 76 |  | 13.4 | 76.9942 | 77.725
7.8 |  | 3.3 | 73.2 | 73.62 |  | 3.7 | 73.5 | 73.84 |  | 5.3 | 73.844 | 74.425 |  | 9.1 | 75.18 | 75.7 |  | 12.9 | 76.8093 | 77.3286
7.9 |  | 3.2 | 72.78 | 73.3143 |  | 3.6 | 73.1 | 73.44 |  | 5.2 | 73.4617 | 74.025 |  | 8.775 | 74.675 | 75.4 |  | 12.56 | 76.4694 | 76.98
8 |  | 3.1667 | 72.5 | 73.075 |  | 3.5 | 72.8 | 73.1333 |  | 5 | 73.0595 | 73.66 |  | 8.6 | 74.48 | 74.98 |  | 12.2286 | 76.1895 | 76.72
8.1 |  | 3.1 | 72.1 | 72.68 |  | 3.4 | 72.3 | 72.64 |  | 4.9 | 72.8018 | 73.3 |  | 8.4 | 74.16 | 74.675 |  | 11.94 | 75.8697 | 76.3
8.2 |  | 3 | 71.7 | 72.35 |  | 3.4 | 72 | 72.4 |  | 4.8 | 72.3819 | 72.96 |  | 8.15 | 73.8 | 74.3333 |  | 11.725 | 75.4199 | 75.92
8.3 |  | 3 | 71.45 | 71.94 |  | 3.3 | 71.65 | 72.0286 |  | 4.7 | 71.9249 | 72.55 |  | 7.85 | 73.38 | 73.8714 |  | 11.4 | 75.07 | 75.7333
8.4 |  | 2.9 | 71.05 | 71.6333 |  | 3.2 | 71.22 | 71.66 |  | 4.6 | 71.5821 | 72.15 |  | 7.6 | 73 | 73.475 |  | 11.0667 | 74.6701 | 75.475
8.5 |  | 2.9 | 70.7 | 71.32 |  | 3.2 | 70.88 | 71.25 |  | 4.425 | 71.2572 | 71.7333 |  | 7.5 | 72.7 | 73.25 |  | 10.85 | 74.4103 | 75.1
8.6 |  | 2.8 | 70.35 | 70.84 |  | 3.1 | 70.52 | 70.85 |  | 4.4 | 70.8823 | 71.4667 |  | 7.4 | 72.3 | 72.775 |  | 10.65 | 74.0704 | 74.75
8.7 |  | 2.8 | 69.94 | 70.55 |  | 3 | 70.14 | 70.5 |  | 4.2 | 70.5574 | 71.04 |  | 7.2 | 71.975 | 72.48 |  | 10.375 | 73.6455 | 74.46
8.8 |  | 2.7 | 69.5 | 70.1 |  | 3 | 69.74 | 70.075 |  | 4.2 | 70.1824 | 70.72 |  | 7.1 | 71.62 | 72.15 |  | 10.2 | 73.3421 | 74.05
8.9 |  | 2.7 | 69.075 | 69.675 |  | 2.9 | 69.35 | 69.8143 |  | 4.1 | 69.7576 | 70.34 |  | 6.9 | 71.24 | 71.7 |  | 10 | 72.9994 | 73.6
9 |  | 2.6333 | 68.68 | 69.3 |  | 2.9 | 68.8 | 69.36 |  | 4 | 69.2827 | 69.8 |  | 6.8 | 70.9 | 71.48 |  | 9.8 | 72.7709 | 73.4143
9.1 |  | 2.6 | 68.15 | 68.74 |  | 2.8 | 68.4143 | 68.92 |  | 3.9 | 68.8578 | 69.475 |  | 6.5667 | 70.45 | 71.12 |  | 9.6 | 72.3311 | 73.06
9.2 |  | 2.6 | 67.5 | 68.16 |  | 2.8 | 67.96 | 68.44 |  | 3.8 | 68.4496 | 69.15 |  | 6.4 | 70.025 | 70.7 |  | 9.4 | 72.0312 | 72.7375
9.3 |  | 2.5 | 66.92 | 67.5 |  | 2.7 | 67.5 | 67.9 |  | 3.7 | 68.183 | 68.6714 |  | 6.3 | 69.7 | 70.2 |  | 9.25 | 71.7047 | 72.24
9.4 |  | 2.5 | 66.2 | 66.875 |  | 2.7 | 66.95 | 67.4 |  | 3.7 | 67.5831 | 68.25 |  | 6.1 | 69.4143 | 69.9 |  | 9.02 | 71.3515 | 72.04
9.5 |  | 2.5 | 65.3 | 65.9 |  | 2.6 | 66.3857 | 66.82 |  | 3.6 | 67.1832 | 67.775 |  | 6 | 69.02 | 69.7 |  | 8.82 | 71.0145 | 71.74
9.6 |  | 2.4 | 63.8 | 64.72 |  | 2.6 | 65.6 | 66.1 |  | 3.5 | 66.8033 | 67.2667 |  | 5.9 | 68.52 | 69.3 |  | 8.575 | 70.5861 | 71.38
9.7 |  | 2.4 | 62.1 | 62.7571 |  | 2.6 | 64.8 | 65.275 |  | 3.4667 | 66.3234 | 66.8 |  | 5.8 | 68.2 | 68.9 |  | 8.5 | 70.172 | 71
9.8 |  | 2.4 | 59.0667 | 60.05 |  | 2.6 | 63.7333 | 64.3 |  | 3.4 | 65.6836 | 66.26 |  | 5.7 | 67.8143 | 68.45 |  | 8.34 | 69.8321 | 70.6714
9.9 |  | 2.3 | 54.56 | 55.9 |  | 2.5 | 62.38 | 62.92 |  | 3.3 | 65.1237 | 65.8 |  | 5.525 | 67.35 | 68.0143 |  | 8.1 | 69.5008 | 70.3
10 |  | 2.3 | 49.0667 | 50.5 |  | 2.5 | 60.68 | 61.1 |  | 3.3 | 64.4839 | 65.16 |  | 5.4 | 66.9429 | 67.66 |  | 8 | 69.2523 | 69.9125
10.1 |  | 2.3 | 42.2 | 44.4667 |  | 2.4 | 57.74 | 58.48 |  | 3.2 | 63.7174 | 64.4 |  | 5.4 | 66.56 | 67.2286 |  | 7.9 | 68.8974 | 69.5
10.2 |  |  | 35.6 | 37.44 |  | 2.4 | 53.85 | 54.5667 |  | 3.15 | 62.9843 | 63.66 |  | 5.2 | 66.06 | 66.75 |  | 7.8 | 68.5726 | 69.24
10.3 |  |  | 28.52 | 31.1 |  |  | 48.75 | 49.38 |  | 3.1 | 61.8445 | 62.7 |  | 5.1 | 65.7 | 66.3889 |  | 7.64 | 68.0442 | 68.76
10.4 |  |  | 22.88 | 25.3 |  |  | 42.62 | 43.7 |  | 3 | 60.4849 | 61.2875 |  | 5.1 | 65.22 | 65.925 |  | 7.5 | 67.7586 | 68.56
10.5 |  |  | 17.8 | 20.6125 |  |  | 36.15 | 37.4 |  | 3 | 58.6853 | 59.55 |  | 4.9333 | 64.8 | 65.6 |  | 7.3 | 67.3531 | 68.18
10.6 |  |  | 14.58 | 17.18 |  |  | 29.9 | 31.0667 |  |  | 56.061 | 57.1 |  | 4.9 | 64.425 | 65.06 |  | 7.1 | 66.9532 | 67.675
10.7 |  |  | 12.35 | 15.0714 |  |  | 24.05 | 25.54 |  |  | 52.027 | 53.5 |  | 4.8 | 63.88 | 64.4667 |  | 7 | 66.5234 | 67.25
10.8 |  |  | 11 | 13.8 |  |  | 19.2571 | 21.175 |  |  | 47.2548 | 48.94 |  | 4.7 | 63.36 | 64 |  | 6.9 | 66.2957 | 66.95
10.9 |  |  | 10.15 | 12.9571 |  |  | 15.7 | 17.9 |  |  | 41.7896 | 43.375 |  | 4.6 | 62.76 | 63.5286 |  | 6.8 | 65.6737 | 66.55
11 |  |  | 9.6 | 12.3 |  |  | 13.5 | 15.775 |  |  | 35.6661 | 37.8 |  |  | 62.06 | 62.84 |  | 6.6167 | 65.4239 | 66.075
11.1 |  |  | 9.175 | 11.9 |  |  | 12.1 | 14.56 |  |  | 30.1125 | 32.35 |  |  | 61.4143 | 62.1333 |  | 6.5 | 65.1073 | 65.6
11.2 |  |  | 8.8429 | 11.55 |  |  | 11.1286 | 13.6 |  |  | 24.8438 | 27.1857 |  |  | 60.68 | 61.44 |  | 6.5 | 64.5142 | 65.3
11.3 |  |  | 8.6667 | 11.26 |  |  | 10.4 | 12.9 |  |  | 20.3949 | 23.3 |  |  | 59.6 | 60.5333 |  | 6.4 | 64.1077 | 64.94
11.4 |  |  | 8.3857 | 11.04 |  |  | 9.95 | 12.42 |  |  | 17.4956 | 20.2667 |  |  | 58.36 | 59.15 |  | 6.3 | 63.7945 | 64.52
11.5 |  |  | 8.2 | 10.92 |  |  | 9.5 | 12.06 |  |  | 15.3762 | 17.98 |  |  | 56.7857 | 57.5 |  |  | 63.1747 | 64.02
11.6 |  |  | 8.1 | 10.7571 |  |  | 9.2333 | 11.8 |  |  | 14.0565 | 16.7 |  |  | 54.2333 | 55.32 |  |  | 62.8149 | 63.6
11.7 |  |  | 8 | 10.6 |  |  | 8.9 | 11.44 |  |  | 13.0968 | 15.72 |  |  | 51.1286 | 52.2286 |  |  | 62.4151 | 63.12
11.8 |  |  | 7.9 | 10.5 |  |  | 8.7167 | 11.2 |  |  | 12.3569 | 15.0571 |  |  | 47.1 | 48.3 |  |  | 61.8552 | 62.65
11.9 |  |  | 7.82 | 10.5 |  |  | 8.5 | 11.02 |  |  | 11.817 | 14.5 |  |  | 42 | 43.58 |  |  | 61.3005 | 62
12 |  |  | 7.7 | 10.4 |  |  | 8.3889 | 10.9 |  |  | 11.4372 | 14.1571 |  |  | 37.3 | 38.7 |  |  | 60.6757 | 61.6
12.1 |  |  | 7.7 | 10.3 |  |  | 8.2143 | 10.7 |  |  | 11.0972 | 13.8 |  |  | 32.3125 | 34.18 |  |  | 60.0426 | 60.9
12.2 |  |  | 7.7 | 10.2167 |  |  | 8.1 | 10.6 |  |  | 10.764 | 13.48 |  |  | 27.7 | 29.8 |  |  | 59.1264 | 60.2
12.3 |  |  | 7.6 | 10.2 |  |  | 8.02 | 10.5 |  |  | 10.4574 | 13.15 |  |  | 23.9 | 26.15 |  |  | 57.9968 | 59.275
12.4 |  |  | 7.6 | 10.1 |  |  | 7.9 | 10.4 |  |  | 10.2225 | 12.925 |  |  | 20.95 | 23.5 |  |  | 56.8173 | 58.06
12.5 |  |  | 7.5 | 10.1 |  |  | 7.82 | 10.35 |  |  | 9.9975 | 12.7 |  |  | 18.9 | 21.56 |  |  | 55.2379 | 56.45
12.6 |  |  | 7.46 | 10 |  |  | 7.76 | 10.2 |  |  | 9.8175 | 12.525 |  |  | 17.45 | 20.2833 |  |  | 53.0788 | 54.36
12.7 |  |  | 7.4 | 9.96 |  |  | 7.7 | 10.2 |  |  | 9.6176 | 12.32 |  |  | 16.45 | 19.4 |  |  | 49.98 | 51.64
12.8 |  |  | 7.36 | 9.9 |  |  | 7.6 | 10.1 |  |  | 9.4976 | 12.1 |  |  | 15.7 | 18.6375 |  |  | 46.3614 | 48.325
12.9 |  |  | 7.3 | 9.8 |  |  | 7.5 | 10 |  |  | 9.3727 | 12 |  |  | 15.1 | 18.08 |  |  | 42.2231 | 44.4
13 |  |  | 7.2 | 9.725 |  |  | 7.4667 | 10 |  |  | 9.1977 | 11.9 |  |  | 14.6429 | 17.6143 |  |  | 37.7516 | 40.075
13.1 |  |  | 7.1 | 9.7 |  |  | 7.4 | 9.9 |  |  | 9.0977 | 11.8143 |  |  | 14.2 | 17.2429 |  |  | 33.5266 | 36.1
13.2 |  |  | 7.1 | 9.6 |  |  | 7.3 | 9.7167 |  |  | 8.9977 | 11.6 |  |  | 13.8857 | 16.9 |  |  | 29.1633 | 32.1667
13.3 |  |  | 7 | 9.6 |  |  | 7.25 | 9.7 |  |  | 8.8978 | 11.525 |  |  | 13.6 | 16.6 |  |  | 25.7497 | 28.975
13.4 |  |  | 6.9 | 9.48 |  |  | 7.15 | 9.6 |  |  | 8.7228 | 11.4 |  |  | 13.3571 | 16.28 |  |  | 23.2107 | 26.3333
13.5 |  |  | 6.9 | 9.4 |  |  | 7.1 | 9.55 |  |  | 8.6578 | 11.3 |  |  | 13 | 16.075 |  |  | 21.1715 | 24.525
13.6 |  |  | 6.8 | 9.3 |  |  | 7 | 9.5 |  |  | 8.5479 | 11.2 |  |  | 12.7889 | 15.8 |  |  | 19.992 | 23.275
13.7 |  |  | 6.7 | 9.2 |  |  | 6.9 | 9.4 |  |  | 8.3979 | 11 |  |  | 12.56 | 15.58 |  |  | 18.8925 | 22.2333
13.8 |  |  | 6.6 | 9.1 |  |  | 6.8 | 9.3 |  |  | 8.3979 | 10.96 |  |  | 12.35 | 15.4 |  |  | 18.2527 | 21.5
13.9 |  |  | 6.6 | 9 |  |  | 6.8 | 9.2 |  |  | 8.198 | 10.85 |  |  | 12.2 | 15.2 |  |  | 17.593 | 20.9667
14 |  |  | 6.4833 | 8.95 |  |  | 6.7 | 9.1167 |  |  | 8.158 | 10.7 |  |  | 12.0667 | 14.96 |  |  | 17.0732 | 20.55
14.1 |  |  | 6.4 | 8.8 |  |  | 6.6125 | 9.02 |  |  | 7.998 | 10.675 |  |  | 11.8286 | 14.775 |  |  | 16.6734 | 20.14
14.2 |  |  | 6.3 | 8.7714 |  |  | 6.5667 | 8.9 |  |  | 7.998 | 10.52 |  |  | 11.7 | 14.6 |  |  | 16.3934 | 19.82
14.3 |  |  | 6.2 | 8.7 |  |  | 6.5 | 8.8 |  |  | 7.8314 | 10.42 |  |  | 11.52 | 14.4 |  |  | 15.9936 | 19.48
14.4 |  |  | 6.14 | 8.5778 |  |  | 6.4 | 8.75 |  |  | 7.6981 | 10.3333 |  |  | 11.4 | 14.3 |  |  | 15.7937 | 19.08
14.5 |  |  | 6 | 8.5 |  |  | 6.3 | 8.7 |  |  | 7.6981 | 10.2 |  |  | 11.15 | 14.12 |  |  | 15.5605 | 18.85
14.6 |  |  | 6 | 8.4 |  |  | 6.2 | 8.56 |  |  | 7.5267 | 10.0714 |  |  | 11.02 | 13.92 |  |  | 15.1939 | 18.64
14.7 |  |  | 5.9 | 8.3333 |  |  | 6.1 | 8.5 |  |  | 7.4981 | 10 |  |  | 10.9 | 13.7333 |  |  | 14.994 | 18.4
14.8 |  |  | 5.8 | 8.2 |  |  | 6 | 8.4 |  |  | 7.3982 | 9.8714 |  |  | 10.8 | 13.625 |  |  | 14.6941 | 18.1
14.9 |  |  | 5.7 | 8.2 |  |  | 5.9667 | 8.3 |  |  | 7.2982 | 9.8 |  |  | 10.6 | 13.48 |  |  | 14.4942 | 17.9
15 |  |  | 5.7 | 8.04 |  |  | 5.9 | 8.2 |  |  | 7.1982 | 9.7 |  |  | 10.5 | 13.3 |  |  | 14.2943 | 17.7
15.1 |  |  | 5.6 | 8 |  |  | 5.8 | 8.1429 |  |  | 7.0982 | 9.6 |  |  | 10.38 | 13.2 |  |  | 14.1543 | 17.5
15.2 | [6 empty] | 5.72 | 8.1 |  |  | 6.9982 | 9.5 |  |  | 10.2 | 13.08 |  |  | 13.9944 | 17.24
15.3 | [6 empty] | 5.7 | 8 |  |  | 6.8983 | 9.4 |  |  | 10 | 12.9 |  |  | 13.8144 | 17.125
15.4 | [6 empty] | 5.5444 | 7.9 |  |  | 6.7983 | 9.3 |  |  | 9.9 | 12.78 |  |  | 13.5279 | 16.9
15.5 | [10 empty] | 6.6983 | 9.2 |  |  | 9.8 | 12.6 |  |  | 13.3696 | 16.68
15.6 | [10 empty] | 6.5983 | 9.1 |  |  | 9.7 | 12.58 |  |  | 13.2547 | 16.5
15.7 | [10 empty] | 6.5234 | 8.9778 |  |  | 9.6 | 12.4 |  |  | 13.0448 | 16.3
15.8 | [10 empty] | 6.3984 | 8.9 |  |  | 9.425 | 12.2857 |  |  | 12.8948 | 16.125
15.9 | [10 empty] | 6.3384 | 8.8 |  |  | 9.3 | 12.12 |  |  | 12.595 | 16
16 | [14 empty] | 9.2 | 11.9333 |  |  | 12.495 | 15.8286
16.1 | [14 empty] | 9.0714 | 11.8 |  |  | 12.395 | 15.7
16.2 | [14 empty] | 8.95 | 11.74 |  |  | 12.1951 | 15.5857
16.3 | [14 empty] | 8.825 | 11.625 |  |  | 12.0952 | 15.42
16.4 | [14 empty] | 8.7 | 11.5 |  |  | 11.8952 | 15.3
16.5 | [14 empty] | 8.68 | 11.4 |  |  | 11.7953 | 15.0667
16.6 | [14 empty] | 8.5 | 11.25 |  |  | 11.6953 | 14.9
16.7 | [14 empty] | 8.4 | 11.1 |  |  | 11.5554 | 14.7
16.8 | [14 empty] | 8.28 | 11 |  |  | 11.4354 | 14.6
16.9 | [14 empty] | 8.2 | 10.8875 |  |  | 11.2705 | 14.4
17 | [18 empty] | 11.1555 | 14.3
17.1 | [18 empty] | 10.9956 | 14.2
17.2 | [18 empty] | 10.8734 | 14.075
17.3 | [18 empty] | 10.7207 | 13.8714
17.4 | [18 empty] | 10.4958 | 13.76
17.5 | [18 empty] | 10.4358 | 13.6
17.6 | [18 empty] | 10.3758 | 13.5
17.7 | [18 empty] | 10.1959 | 13.375
17.8 | [18 empty] | 10.096 | 13.2
17.9 | [18 empty] | 9.996 | 13.04

## Wedge 45 Longitudinal Profiles
Longitudinal Profiles
Wedge 45
SSD=100cm
Field Size=40x20cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 8.4
-34.9 |  |  |  |  | 8.425
-34.8 |  |  |  |  | 8.5571
-34.7 |  |  |  |  | 8.6538
-34.6 |  |  |  |  | 8.7308
-34.5 |  |  |  |  | 8.8
-34.4 |  |  |  |  | 8.8364
-34.3 |  |  |  |  | 8.92
-34.2 |  |  |  |  | 8.9867
-34.1 |  |  |  |  | 9.08
-34 |  |  |  |  | 9.1667
-33.9 |  |  |  |  | 9.2667
-33.8 |  |  |  |  | 9.3
-33.7 |  |  |  |  | 9.4
-33.6 |  |  |  |  | 9.5
-33.5 |  |  |  |  | 9.6
-33.4 |  |  |  |  | 9.7
-33.3 |  |  |  |  | 9.78
-33.2 |  |  |  |  | 9.88
-33.1 |  |  |  |  | 9.9
-33 |  |  |  |  | 10
-32.9 |  |  |  |  | 10.1
-32.8 |  |  |  |  | 10.2
-32.7 |  |  |  |  | 10.3
-32.6 |  |  |  |  | 10.4571
-32.5 |  |  |  |  | 10.54
-32.4 |  |  |  |  | 10.7
-32.3 |  |  |  | 7.525 | 10.8
-32.2 |  |  |  | 7.6 | 10.85
-32.1 |  |  |  | 7.6 | 11
-32 |  |  |  | 7.6625 | 11.1125
-31.9 |  |  |  | 7.7467 | 11.2429
-31.8 |  |  |  | 7.8143 | 11.375
-31.7 |  |  |  | 7.8857 | 11.5
-31.6 |  |  |  | 7.9727 | 11.5375
-31.5 |  |  |  | 8.05 | 11.6625
-31.4 |  |  |  | 8.1375 | 11.8
-31.3 |  |  |  | 8.2385 | 11.9286
-31.2 |  |  |  | 8.3154 | 12.0625
-31.1 |  |  |  | 8.3923 | 12.2
-31 |  |  |  | 8.5 | 12.2857
-30.9 |  |  |  | 8.5571 | 12.4143
-30.8 |  |  |  | 8.6 | 12.58
-30.7 |  |  |  | 8.7 | 12.65
-30.6 |  |  |  | 8.8 | 12.85
-30.5 |  |  |  | 8.9 | 12.9125
-30.4 |  |  |  | 8.98 | 13.0429
-30.3 |  |  |  | 9 | 13.175
-30.2 |  |  |  | 9.1 | 13.4
-30.1 |  |  |  | 9.2 | 13.5375
-30 |  |  |  | 9.3 | 13.6714
-29.9 |  |  |  | 9.44 | 13.8
-29.8 |  |  |  | 9.54 | 13.9375
-29.7 |  |  |  | 9.64 | 14.125
-29.6 |  |  | 6.575 | 9.7333 | 14.3
-29.5 |  |  | 6.6 | 9.8 | 14.4286
-29.4 |  |  | 6.65 | 9.9375 | 14.625
-29.3 |  |  | 6.7857 | 10 | 14.9
-29.2 |  |  | 6.8 | 10.1 | 15.0333
-29.1 |  |  | 6.8333 | 10.2375 | 15.1857
-29 |  |  | 6.9444 | 10.3625 | 15.45
-28.9 |  |  | 7 | 10.4875 | 15.6571
-28.8 |  |  | 7.1 | 10.5818 | 15.7875
-28.7 |  |  | 7.1 | 10.7067 | 16.025
-28.6 |  |  | 7.1714 | 10.83 | 16.2143
-28.5 |  |  | 7.2 | 10.9375 | 16.3875
-28.4 |  |  | 7.3375 | 11.0455 | 16.65
-28.3 |  | 6.1 | 7.4 | 11.1889 | 16.8571
-28.2 |  | 6.1556 | 7.46 | 11.3 | 16.9875
-28.1 |  | 6.2 | 7.5 | 11.5 | 17.25
-28 |  | 6.3 | 7.5357 | 11.6 | 17.5
-27.9 |  | 6.325 | 7.6091 | 11.6714 | 17.7714
-27.8 |  | 6.4 | 7.7 | 11.9 | 18.0286
-27.7 |  | 6.4 | 7.8067 | 12.0429 | 18.3
-27.6 |  | 6.4857 | 7.8733 | 12.2 | 18.5714
-27.5 |  | 6.5125 | 7.95 | 12.3375 | 18.7375
-27.4 |  | 6.6 | 8 | 12.5 | 19.1286
-27.3 | 6.3 | 6.6 | 8.0875 | 12.6333 | 19.525
-27.2 | 6.3857 | 6.7 | 8.2 | 12.8167 | 19.8375
-27.1 | 6.4 | 6.7375 | 8.2571 | 12.9833 | 20.2286
-27 | 6.4 | 6.8 | 8.3 | 13.1571 | 20.775
-26.9 | 6.4857 | 6.8 | 8.38 | 13.35 | 21.3625
-26.8 | 6.5 | 6.9 | 8.5063 | 13.5571 | 22.1
-26.7 | 6.58 | 6.95 | 8.5687 | 13.8 | 23.1
-26.6 | 6.6 | 7 | 8.6667 | 13.9 | 24.2667
-26.5 | 6.6833 | 7.1 | 8.8 | 14.1714 | 26.05
-26.4 | 6.7 | 7.1 | 8.8667 | 14.3286 | 27.9
-26.3 | 6.7462 | 7.16 | 8.9455 | 14.6 | 30.725
-26.2 | 6.8 | 7.2 | 9 | 14.8714 | 34
-26.1 | 6.8 | 7.3143 | 9.0875 | 15.0286 | 37.9
-26 | 6.8889 | 7.4 | 9.2 | 15.25 | 42.3
-25.9 | 7 | 7.4 | 9.2571 | 15.4714 | 46.9
-25.8 | 7 | 7.49 | 9.3875 | 15.8 | 51.6
-25.7 | 7.0857 | 7.5 | 9.525 | 16.0375 | 55.8333
-25.6 | 7.1 | 7.6 | 9.65 | 16.2429 | 59.72
-25.5 | 7.1143 | 7.6625 | 9.7857 | 16.6 | 63.2333
-25.4 | 7.2 | 7.7778 | 9.825 | 16.8667 | 65.86
-25.3 | 7.2857 | 7.8 | 9.9571 | 17.1714 | 68.1
-25.2 | 7.3 | 7.9 | 10.0875 | 17.5625 | 69.8
-25.1 | 7.3571 | 7.9625 | 10.26 | 18 | 71
-25 | 7.4 | 8 | 10.4429 | 18.4714 | 71.9429
-24.9 | 7.5 | 8.1 | 10.5857 | 19.09 | 72.7714
-24.8 | 7.5 | 8.1625 | 10.7125 | 19.86 | 73.45
-24.7 | 7.6 | 8.2 | 10.8429 | 20.9909 | 73.9143
-24.6 | 7.6571 | 8.3 | 11.05 | 22.5 | 74.3875
-24.5 | 7.7 | 8.3571 | 11.225 | 24.2667 | 74.9
-24.4 | 7.7286 | 8.4 | 11.4429 | 26.9333 | 75.2875
-24.3 | 7.8 | 8.525 | 11.575 | 30.1 | 75.6625
-24.2 | 7.9 | 8.65 | 11.725 | 33.6 | 76.0857
-24.1 | 7.9 | 8.7 | 11.9857 | 38.4 | 76.425
-24 | 8 | 8.8143 | 12.175 | 43.5333 | 76.675
-23.9 | 8.0286 | 8.95 | 12.4125 | 49 | 77.075
-23.8 | 8.1 | 9.0857 | 12.575 | 54.4 | 77.4
-23.7 | 8.2 | 9.2143 | 12.8429 | 59.6 | 77.5857
-23.6 | 8.2333 | 9.3 | 13 | 64.1 | 77.9571
-23.5 | 8.3 | 9.3857 | 13.375 | 67.9286 | 78.4
-23.4 | 8.4 | 9.6143 | 13.6429 | 70.875 | 78.5857
-23.3 | 8.4625 | 9.75 | 14 | 73.0875 | 78.8286
-23.2 | 8.6 | 9.8857 | 14.3125 | 74.7857 | 79.1
-23.1 | 8.7 | 10.0125 | 14.6875 | 75.975 | 79.3714
-23 | 8.7625 | 10.1857 | 15.3 | 76.8 | 79.6286
-22.9 | 8.9 | 10.3857 | 15.8714 | 77.6143 | 79.85
-22.8 | 9.025 | 10.5125 | 16.7375 | 78.3286 | 80.1571
-22.7 | 9.1 | 10.6857 | 17.9 | 78.9 | 80.425
-22.6 | 9.2 | 10.8857 | 19.4286 | 79.3125 | 80.6857
-22.5 | 9.35 | 11.125 | 21.6875 | 79.6875 | 80.9714
-22.4 | 9.5 | 11.3857 | 24.5 | 80.1 | 81.1286
-22.3 | 9.6 | 11.65 | 28.4 | 80.42 | 81.4
-22.2 | 9.7333 | 11.9375 | 33.3125 | 80.9 | 81.575
-22.1 | 9.9167 | 12.3125 | 39.125 | 81.1857 | 81.8
-22 | 10.1 | 12.7857 | 45.6571 | 81.4625 | 82.0429
-21.9 | 10.2167 | 13.4 | 52.4375 | 81.675 | 82.25
-21.8 | 10.4 | 14.2 | 59.325 | 82.05 | 82.5
-21.7 | 10.6 | 15.3429 | 65.2 | 82.3143 | 82.675
-21.6 | 10.8 | 16.9 | 70.6571 | 82.6625 | 82.8714
-21.5 | 11.1 | 19.325 | 74.5714 | 82.9429 | 83.0167
-21.4 | 11.6 | 22.5571 | 77.8 | 83.3 | 83.2429
-21.3 | 12.1 | 26.8 | 80.0857 | 83.5714 | 83.5
-21.2 | 12.8 | 32.575 | 81.6286 | 83.7286 | 83.675
-21.1 | 14 | 39.2 | 82.85 | 84 | 83.8833
-21 | 15.5667 | 46.4 | 83.8143 | 84.25 | 84.1
-20.9 | 17.84 | 54.3625 | 84.3857 | 84.425 | 84.22
-20.8 | 21.24 | 62.0143 | 85.05 | 84.6857 | 84.4
-20.7 | 25.8 | 68.8 | 85.5429 | 84.875 | 84.5
-20.6 | 31.68 | 74.3571 | 85.9375 | 85.125 | 84.7
-20.5 | 38.98 | 78.875 | 86.3286 | 85.3429 | 84.8857
-20.4 | 47.0857 | 82.1 | 86.65 | 85.55 | 85.2143
-20.3 | 55.66 | 84.35 | 87.025 | 85.8125 | 85.35
-20.2 | 63.98 | 85.9143 | 87.3286 | 85.9429 | 85.4857
-20.1 | 71.48 | 87 | 87.65 | 86.15 | 85.61
-20 | 77.82 | 87.925 | 87.9 | 86.3 | 85.7143
-19.9 | 82.46 | 88.5286 | 88.175 | 86.5375 | 85.9
-19.8 | 85.64 | 89.05 | 88.3714 | 86.725 | 86.1333
-19.7 | 87.82 | 89.475 | 88.7 | 86.9 | 86.2778
-19.6 | 89.02 | 89.8143 | 88.9375 | 87.0286 | 86.4
-19.5 | 90.08 | 90.1625 | 89.0714 | 87.225 | 86.6727
-19.4 | 90.74 | 90.4 | 89.3167 | 87.5 | 86.88
-19.3 | 91.24 | 90.7625 | 89.5667 | 87.65 | 87
-19.2 | 91.62 | 91.02 | 89.7667 | 87.8571 | 87.175
-19.1 | 91.94 | 91.1857 | 89.9 | 87.9875 | 87.3
-19 | 92.225 | 91.4667 | 90.1286 | 88.15 | 87.4
-18.9 | 92.4429 | 91.6125 | 90.2714 | 88.3 | 87.525
-18.8 | 92.65 | 91.76 | 90.4286 | 88.3875 | 87.7333
-18.7 | 92.9125 | 91.94 | 90.6 | 88.625 | 87.925
-18.6 | 93.0429 | 92.1667 | 90.7375 | 88.75 | 88.0571
-18.5 | 93.2 | 92.375 | 90.925 | 88.9714 | 88.1875
-18.4 | 93.3429 | 92.4125 | 91.1 | 89.1286 | 88.225
-18.3 | 93.4857 | 92.5857 | 91.2375 | 89.35 | 88.4
-18.2 | 93.6375 | 92.775 | 91.3625 | 89.4857 | 88.5
-18.1 | 93.7714 | 92.8125 | 91.5 | 89.6 | 88.7125
-18 | 93.95 | 92.9429 | 91.6571 | 89.74 | 88.8429
-17.9 | 94.1 | 93.075 | 91.8 | 89.82 | 88.975
-17.8 | 94.1286 | 93.2 | 91.9 | 90.0714 | 89.1125
-17.7 | 94.2625 | 93.3429 | 92.0286 | 90.1 | 89.2375
-17.6 | 94.4 | 93.5 | 92.1625 | 90.2143 | 89.3
-17.5 | 94.5286 | 93.5429 | 92.2875 | 90.35 | 89.5
-17.4 | 94.6 | 93.675 | 92.525 | 90.4857 | 89.6333
-17.3 | 94.6714 | 93.8 | 92.6 | 90.6125 | 89.7571
-17.2 | 94.8286 | 93.9375 | 92.6875 | 90.7429 | 89.8
-17.1 | 94.9 | 94 | 92.725 | 90.875 | 89.925
-17 | 95 | 94.1 | 92.8571 | 91.02 | 90.0889
-16.9 | 95.0625 | 94.23 | 92.9875 | 91.1 | 90.2667
-16.8 | 95.2 | 94.3214 | 93.125 | 91.2 | 90.4
-16.7 | 95.3 | 94.3929 | 93.25 | 91.38 | 90.5333
-16.6 | 95.3625 | 94.55 | 93.3 | 91.48 | 90.6333
-16.5 | 95.4 | 94.6467 | 93.4 | 91.68 | 90.7
-16.4 | 95.5 | 94.7154 | 93.5 | 91.7333 | 90.8333
-16.3 | 95.5286 | 94.7923 | 93.6857 | 91.82 | 90.9667
-16.2 | 95.6714 | 94.8 | 93.8143 | 92.02 | 91.1
-16.1 | 95.7 | 94.9 | 93.9 | 92.2 | 91.225
-16 | 95.7769 | 95.0375 | 93.975 | 92.3 | 91.3267
-15.9 | 95.8778 | 95.1 | 94.1125 | 92.4 | 91.3933
-15.8 | 95.9 | 95.1 | 94.2 | 92.4 | 91.55
-15.7 | 96 | 95.2375 | 94.2 | 92.5 | 91.6538
-15.6 | 96.0375 | 95.3 | 94.3 | 92.6 | 91.7333
-15.5 | 96.1625 | 95.4 | 94.4375 | 92.7 | 91.8154
-15.4 | 96.2583 | 95.5 | 94.5625 | 92.86 | 91.8923
-15.3 | 96.3 | 95.5714 | 94.7 | 92.9429 | 92.02
-15.2 | 96.4 | 95.6 | 94.7222 | 93.1 | 92.1615
-15.1 | 96.425 | 95.7 | 94.82 | 93.12 | 92.3125
-15 | 96.5571 | 95.7625 | 94.8867 | 93.2 | 92.4333
-14.9 | 96.6 | 95.9 | 94.98 | 93.3 | 92.5
-14.8 | 96.6286 | 95.925 | 95.0533 | 93.4 | 92.5857
-14.7 | 96.7833 | 96 | 95.1 | 93.5 | 92.8
-14.6 | 96.8 | 96.0875 | 95.1727 | 93.6 | 92.85
-14.5 | 96.9 | 96.1 | 95.32 | 93.6571 | 92.9
-14.4 | 96.98 | 96.1571 | 95.3867 | 93.7875 | 93
-14.3 | 97 | 96.2875 | 95.4533 | 93.9 | 93.0375
-14.2 | 97.04 | 96.4 | 95.52 | 93.9 | 93.1
-14.1 | 97.1 | 96.4571 | 95.5867 | 94.0091 | 93.2
-14 | 97.2 | 96.5 | 95.6571 | 94.1333 | 93.275
-13.9 | 97.22 | 96.6 | 95.7 | 94.2 | 93.4625
-13.8 | 97.3 | 96.6 | 95.7333 | 94.2667 | 93.6
-13.7 | 97.32 | 96.6875 | 95.8 | 94.3625 | 93.7286
-13.6 | 97.4 | 96.7 | 95.8714 | 94.4875 | 93.8
-13.5 | 97.4 | 96.75 | 96 | 94.56 | 93.9
-13.4 | 97.5 | 96.8857 | 96.1 | 94.6267 | 94
-13.3 | 97.6 | 96.9 | 96.1714 | 94.6933 | 94.0625
-13.2 | 97.7 | 96.9429 | 96.2 | 94.76 | 94.1875
-13.1 | 97.7333 | 97 | 96.3 | 94.8267 | 94.3
-13 | 97.8 | 97.0125 | 96.3714 | 94.8933 | 94.35
-12.9 | 97.9 | 97.1429 | 96.4 | 95.02 | 94.4
-12.8 | 97.9 | 97.2 | 96.5 | 95.1533 | 94.5
-12.7 | 98 | 97.3 | 96.5 | 95.2867 | 94.6375
-12.6 | 98 | 97.3375 | 96.6 | 95.3 | 94.7714
-12.5 | 98 | 97.4 | 96.6 | 95.325 | 94.8
-12.4 | 98 | 97.4 | 96.7 | 95.4 | 94.8714
-12.3 | 98.075 | 97.5 | 96.7571 | 95.4 | 95
-12.2 | 98.1 | 97.5375 | 96.8 | 95.475 | 95
-12.1 | 98.2 | 97.6 | 96.825 | 95.5889 | 95.1111
-12 | 98.2 | 97.6 | 96.9571 | 95.69 | 95.2286
-11.9 | 98.2 | 97.6625 | 97.02 | 95.79 | 95.3625
-11.8 | 98.2 | 97.7 | 97.1 | 95.9 | 95.4
-11.7 | 98.2714 | 97.7375 | 97.1 | 95.9 | 95.4
-11.6 | 98.4 | 97.8625 | 97.1 | 95.975 | 95.4571
-11.5 | 98.4 | 97.9 | 97.2 | 96.0143 | 95.5
-11.4 | 98.4 | 97.925 | 97.25 | 96.15 | 95.6
-11.3 | 98.4875 | 98 | 97.3 | 96.2 | 95.6
-11.2 | 98.5 | 98 | 97.3125 | 96.2 | 95.6875
-11.1 | 98.5 | 98.1 | 97.4 | 96.3 | 95.8
-11 | 98.5875 | 98.1 | 97.475 | 96.3 | 95.85
-10.9 | 98.625 | 98.1875 | 97.5125 | 96.42 | 95.9
-10.8 | 98.7 | 98.2 | 97.6 | 96.6 | 96.0333
-10.7 | 98.7 | 98.2 | 97.675 | 96.6 | 96.1
-10.6 | 98.725 | 98.2875 | 97.7 | 96.62 | 96.2
-10.5 | 98.8 | 98.4 | 97.775 | 96.72 | 96.2143
-10.4 | 98.8 | 98.4 | 97.8 | 96.8 | 96.35
-10.3 | 98.86 | 98.475 | 97.875 | 96.9 | 96.4857
-10.2 | 98.9 | 98.5 | 97.9 | 96.9 | 96.5
-10.1 | 98.9 | 98.5 | 97.9375 | 97 | 96.55
-10 | 98.9 | 98.575 | 98 | 97.0375 | 96.6125
-9.9 | 99 | 98.6 | 98 | 97.1 | 96.7
-9.8 | 99 | 98.6 | 98.0375 | 97.1 | 96.775
-9.7 | 99.1 | 98.7 | 98.1 | 97.2 | 96.9
-9.6 | 99.1 | 98.7 | 98.1 | 97.2625 | 97
-9.5 | 99.1 | 98.7 | 98.2 | 97.3 | 97
-9.4 | 99.1571 | 98.8 | 98.2625 | 97.325 | 97.1
-9.3 | 99.2 | 98.8 | 98.3 | 97.4571 | 97.1375
-9.2 | 99.2125 | 98.8857 | 98.3 | 97.5 | 97.2
-9.1 | 99.3 | 98.9 | 98.3571 | 97.6 | 97.2286
-9 | 99.375 | 98.9 | 98.4 | 97.6571 | 97.3
-8.9 | 99.4 | 98.9667 | 98.425 | 97.7 | 97.3286
-8.8 | 99.4 | 99 | 98.5 | 97.725 | 97.4
-8.7 | 99.4545 | 99 | 98.5 | 97.8 | 97.5
-8.6 | 99.5 | 99.075 | 98.525 | 97.8875 | 97.6
-8.5 | 99.5 | 99.1 | 98.6 | 97.9 | 97.7
-8.4 | 99.5 | 99.1214 | 98.6 | 97.9429 | 97.8
-8.3 | 99.5 | 99.1929 | 98.6 | 98 | 97.9
-8.2 | 99.5214 | 99.2 | 98.6 | 98.1 | 97.9
-8.1 | 99.5929 | 99.2 | 98.6857 | 98.1429 | 97.9
-8 | 99.66 | 99.2 | 98.8 | 98.2 | 97.9
-7.9 | 99.7 | 99.2533 | 98.8 | 98.2125 | 98
-7.8 | 99.7 | 99.3 | 98.8 | 98.3 | 98.05
-7.7 | 99.7 | 99.3 | 98.9 | 98.375 | 98.1
-7.6 | 99.7 | 99.3 | 98.9 | 98.4 | 98.2
-7.5 | 99.7 | 99.3125 | 98.9167 | 98.4 | 98.2571
-7.4 | 99.7 | 99.4 | 99 | 98.4 | 98.3
-7.3 | 99.7 | 99.4 | 99 | 98.4 | 98.4143
-7.2 | 99.7 | 99.5 | 99 | 98.5 | 98.5
-7.1 | 99.71 | 99.5 | 99.04 | 98.5625 | 98.5
-7 | 99.8 | 99.5714 | 99.1 | 98.6 | 98.5
-6.9 | 99.8 | 99.6 | 99.1 | 98.6333 | 98.5
-6.8 | 99.8 | 99.6 | 99.2 | 98.7 | 98.5
-6.7 | 99.8 | 99.6 | 99.2 | 98.7833 | 98.6
-6.6 | 99.8 | 99.6 | 99.2 | 98.8 | 98.6429
-6.5 | 99.8 | 99.6 | 99.2083 | 98.8571 | 98.7
-6.4 | 99.8 | 99.6625 | 99.2917 | 98.9333 | 98.8
-6.3 | 99.9 | 99.7 | 99.3 | 99 | 98.8429
-6.2 | 99.9 | 99.7 | 99.3375 | 99 | 98.9
-6.1 | 99.9 | 99.7 | 99.4 | 99 | 98.9
-6 | 100 | 99.7857 | 99.4 | 99 | 98.9
-5.9 | 100 | 99.8 | 99.5 | 99.1 | 99
-5.8 | 100 | 99.8 | 99.5 | 99.1 | 99
-5.7 | 100 | 99.8 | 99.56 | 99.1857 | 99
-5.6 | 100 | 99.8 | 99.6 | 99.2 | 99.06
-5.5 | 100 | 99.8 | 99.6 | 99.2 | 99.1
-5.4 | 100 | 99.8714 | 99.6 | 99.3 | 99.1857
-5.3 | 100.1 | 99.9 | 99.6 | 99.32 | 99.3
-5.2 | 100.1 | 99.9 | 99.6 | 99.4 | 99.3833
-5.1 | 100.1 | 99.9357 | 99.6 | 99.4 | 99.4
-5 | 100.1 | 100 | 99.6 | 99.4 | 99.4
-4.9 | 100.1 | 100 | 99.6 | 99.4 | 99.4
-4.8 | 100.1 | 100 | 99.6 | 99.4 | 99.5
-4.7 | 100.1 | 100 | 99.7 | 99.4 | 99.5
-4.6 | 100.1375 | 100 | 99.7 | 99.4 | 99.5
-4.5 | 100.2 | 100 | 99.8 | 99.5 | 99.5
-4.4 | 100.2 | 100 | 99.8 | 99.5 | 99.5
-4.3 | 100.2 | 100 | 99.8 | 99.5 | 99.6
-4.2 | 100.1 | 100 | 99.8 | 99.5375 | 99.6625
-4.1 | 100.1 | 100 | 99.8 | 99.6 | 99.7
-4 | 100.1 | 100 | 99.8 | 99.6 | 99.7
-3.9 | 100.1 | 100 | 99.8 | 99.6286 | 99.7
-3.8 | 100.1 | 100.01 | 99.8 | 99.7 | 99.7
-3.7 | 100.1 | 100.1 | 99.8 | 99.7 | 99.7
-3.6 | 100.1 | 100.1 | 99.8875 | 99.7 | 99.7375
-3.5 | 100.12 | 100.1 | 99.9 | 99.7 | 99.8
-3.4 | 100.2 | 100.1 | 99.9 | 99.7 | 99.8
-3.3 | 100.2 | 100.1 | 99.9 | 99.7 | 99.8625
-3.2 | 100.2 | 100.1 | 99.9 | 99.7 | 99.9
-3.1 | 100.18 | 100.1 | 99.9 | 99.8 | 99.9125
-3 | 100.1 | 100.1 | 99.9 | 99.8 | 100
-2.9 | 100.1 | 100.1 | 99.9 | 99.8 | 100
-2.8 | 100.1 | 100.1 | 99.9 | 99.8 | 100
-2.7 | 100.1 | 100.1 | 99.9 | 99.9 | 100
-2.6 | 100.1 | 100.1 | 99.9 | 99.9 | 100
-2.5 | 100.1 | 100.1 | 99.9 | 99.9 | 100
-2.4 | 100.1 | 100.1 | 99.9 | 99.9 | 100
-2.3 | 100.1 | 100.1 | 99.9 | 99.9 | 100
-2.2 | 100.1 | 100.1 | 99.9 | 99.9 | 100
-2.1 | 100.1 | 100.1 | 99.925 | 99.9 | 100.0375
-2 | 100.1 | 100.1 | 100 | 99.9 | 100.1
-1.9 | 100.1 | 100.1 | 100 | 99.9 | 100.1
-1.8 | 100.1 | 100.1 | 100 | 99.9 | 100
-1.7 | 100.1 | 100.1 | 100 | 99.9 | 100
-1.6 | 100.1 | 100.1 | 100 | 99.9 | 99.9125
-1.5 | 100.1 | 100.1 | 100 | 99.9 | 99.9
-1.4 | 100 | 100 | 100 | 99.9 | 99.95
-1.3 | 100 | 100 | 100 | 99.9 | 100
-1.2 | 100 | 100 | 100 | 99.9 | 100
-1.1 | 100 | 100 | 100 | 99.9 | 100
-1 | 100 | 100 | 100 | 99.9 | 100
-0.9 | 100 | 100 | 100 | 99.9 | 100
-0.8 | 100 | 100 | 100 | 99.9 | 100
-0.7 | 100 | 100 | 100 | 99.9 | 100
-0.6 | 100 | 100 | 100 | 99.9 | 100
-0.5 | 100 | 100 | 100 | 99.9 | 100
-0.4 | 100 | 100 | 100 | 99.9 | 100
-0.3 | 100 | 100 | 100 | 100 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 99.9875 | 100
0.4 | 100 | 100 | 100 | 99.9 | 100
0.5 | 100 | 100 | 100 | 99.9 | 100
0.6 | 100 | 100 | 100 | 99.9 | 100
0.7 | 100 | 100 | 100 | 99.9 | 100
0.8 | 100 | 100 | 100 | 99.9 | 100
0.9 | 100 | 100 | 100 | 99.9 | 100
1 | 100 | 100 | 100 | 99.9 | 100
1.1 | 100 | 100 | 100 | 99.9 | 100
1.2 | 100 | 100 | 100 | 99.9 | 100
1.3 | 100 | 100 | 100 | 99.9 | 100
1.4 | 100 | 100 | 100 | 99.9 | 99.9143
1.5 | 100.1 | 100.075 | 100 | 99.9 | 99.9
1.6 | 100.1 | 100.1 | 100 | 99.9 | 99.95
1.7 | 100.1 | 100.1 | 100 | 99.9 | 100
1.8 | 100.1 | 100.1 | 100 | 99.9 | 100
1.9 | 100.1 | 100.1 | 100 | 99.9 | 100.0375
2 | 100.1 | 100.1 | 100 | 99.9 | 100.1
2.1 | 100.1 | 100.1 | 99.925 | 99.9 | 100
2.2 | 100.1 | 100.1 | 99.9 | 99.9 | 100
2.3 | 100.1 | 100.1 | 99.9 | 99.9 | 100
2.4 | 100.1 | 100.1 | 99.9 | 99.9 | 100
2.5 | 100.1 | 100.1 | 99.9 | 99.9 | 100
2.6 | 100.1 | 100.1 | 99.9 | 99.9 | 100
2.7 | 100.1 | 100.1 | 99.9 | 99.875 | 100
2.8 | 100.1 | 100.1 | 99.9 | 99.8 | 100
2.9 | 100.1 | 100.1 | 99.9 | 99.8 | 100
3 | 100.1 | 100.1 | 99.9 | 99.8 | 100
3.1 | 100.14 | 100.1 | 99.9 | 99.7875 | 99.9857
3.2 | 100.2 | 100.1 | 99.9 | 99.7 | 99.9
3.3 | 100.2 | 100.1 | 99.9 | 99.7 | 99.825
3.4 | 100.2 | 100.1 | 99.9 | 99.7 | 99.8
3.5 | 100.16 | 100.1 | 99.9 | 99.7 | 99.8
3.6 | 100.1 | 100.1 | 99.8875 | 99.7 | 99.8
3.7 | 100.1 | 100.1 | 99.8 | 99.7 | 99.7
3.8 | 100.1 | 100.0375 | 99.8 | 99.7 | 99.7
3.9 | 100.1 | 100 | 99.8 | 99.6 | 99.7
4 | 100.1 | 100 | 99.8 | 99.6 | 99.7
4.1 | 100.1 | 100 | 99.8 | 99.6 | 99.7
4.2 | 100.1 | 100 | 99.8 | 99.5 | 99.7
4.3 | 100.175 | 100 | 99.8 | 99.5 | 99.6
4.4 | 100.2 | 100 | 99.8 | 99.5 | 99.575
4.5 | 100.2 | 100 | 99.8 | 99.46 | 99.5
4.6 | 100.1 | 100 | 99.7 | 99.4 | 99.5
4.7 | 100.1 | 100 | 99.7 | 99.4 | 99.5
4.8 | 100.1 | 100 | 99.6 | 99.4 | 99.4625
4.9 | 100.1 | 100 | 99.6 | 99.4 | 99.4
5 | 100.1 | 100 | 99.6 | 99.4 | 99.4
5.1 | 100.1 | 99.9667 | 99.6 | 99.4 | 99.4
5.2 | 100.1 | 99.9 | 99.6 | 99.4 | 99.36
5.3 | 100.1 | 99.9 | 99.6 | 99.3 | 99.3
5.4 | 100.02 | 99.9 | 99.6 | 99.3 | 99.2571
5.5 | 100 | 99.8 | 99.6 | 99.2143 | 99.1
5.6 | 100 | 99.8 | 99.6 | 99.2 | 99.0333
5.7 | 100 | 99.8 | 99.56 | 99.15 | 99
5.8 | 100 | 99.8 | 99.5 | 99.1 | 99
5.9 | 100 | 99.8 | 99.5 | 99.0875 | 98.975
6 | 99.9375 | 99.7143 | 99.4 | 99 | 98.9
6.1 | 99.9286 | 99.7 | 99.4 | 99 | 98.9
6.2 | 99.9 | 99.7 | 99.3375 | 99 | 98.9
6.3 | 99.9 | 99.7 | 99.3 | 99 | 98.8125
6.4 | 99.8286 | 99.6875 | 99.2917 | 98.9 | 98.775
6.5 | 99.8 | 99.6 | 99.2083 | 98.8286 | 98.7
6.6 | 99.8 | 99.6 | 99.2 | 98.8 | 98.6125
6.7 | 99.8 | 99.6 | 99.2 | 98.75 | 98.575
6.8 | 99.8 | 99.6 | 99.2 | 98.7 | 98.5
6.9 | 99.8 | 99.6 | 99.1 | 98.7 | 98.5
7 | 99.8 | 99.5 | 99.1 | 98.6 | 98.5
7.1 | 99.74 | 99.5 | 99.04 | 98.5286 | 98.5
7.2 | 99.7 | 99.4375 | 99 | 98.5 | 98.5
7.3 | 99.7 | 99.4 | 99 | 98.4 | 98.4
7.4 | 99.7 | 99.3714 | 99 | 98.4 | 98.35
7.5 | 99.7 | 99.3 | 98.9167 | 98.4 | 98.225
7.6 | 99.7 | 99.3 | 98.9 | 98.4 | 98.1875
7.7 | 99.7 | 99.3 | 98.9 | 98.3429 | 98.1
7.8 | 99.7 | 99.3 | 98.8 | 98.3 | 98.0167
7.9 | 99.675 | 99.2867 | 98.8 | 98.2 | 98
8 | 99.6 | 99.22 | 98.8 | 98.2 | 97.92
8.1 | 99.6 | 99.2 | 98.6857 | 98.1125 | 97.9
8.2 | 99.55 | 99.2 | 98.6 | 98.075 | 97.9
8.3 | 99.5 | 99.2 | 98.6 | 98 | 97.86
8.4 | 99.5 | 99.1562 | 98.6 | 97.9125 | 97.8
8.5 | 99.5 | 99.1 | 98.6 | 97.9 | 97.7
8.6 | 99.5 | 99.1 | 98.525 | 97.85 | 97.6
8.7 | 99.4889 | 99.0286 | 98.5 | 97.8 | 97.5
8.8 | 99.4 | 99 | 98.5 | 97.7 | 97.48
8.9 | 99.4 | 98.9111 | 98.425 | 97.7 | 97.4
9 | 99.4 | 98.9 | 98.4 | 97.625 | 97.3
9.1 | 99.275 | 98.9 | 98.3571 | 97.6 | 97.3
9.2 | 99.2 | 98.9 | 98.3 | 97.5571 | 97.2
9.3 | 99.2 | 98.8 | 98.3 | 97.425 | 97.2
9.4 | 99.1875 | 98.725 | 98.2625 | 97.3 | 97.1
9.5 | 99.1 | 98.7 | 98.2 | 97.3 | 97.0375
9.6 | 99.1 | 98.7 | 98.1 | 97.225 | 97
9.7 | 99.0333 | 98.7 | 98.1 | 97.2 | 96.9
9.8 | 99 | 98.6 | 98.0375 | 97.1571 | 96.8375
9.9 | 99 | 98.6 | 98 | 97.1 | 96.7125
10 | 98.9 | 98.5125 | 98 | 97 | 96.6
10.1 | 98.9 | 98.5 | 97.9375 | 96.9714 | 96.6
10.2 | 98.9 | 98.5 | 97.9 | 96.9 | 96.5
10.3 | 98.9 | 98.5 | 97.875 | 96.88 | 96.45
10.4 | 98.8 | 98.4 | 97.8 | 96.8 | 96.3143
10.5 | 98.8 | 98.35 | 97.775 | 96.7 | 96.2857
10.6 | 98.7571 | 98.2143 | 97.7 | 96.6 | 96.2
10.7 | 98.7 | 98.2 | 97.675 | 96.6 | 96.1
10.8 | 98.6875 | 98.2 | 97.6 | 96.58 | 96.1
10.9 | 98.6 | 98.2 | 97.5125 | 96.38 | 95.95
11 | 98.6 | 98.1 | 97.475 | 96.3 | 95.8143
11.1 | 98.5 | 98.1 | 97.4 | 96.3 | 95.7857
11.2 | 98.5 | 98.025 | 97.3125 | 96.2 | 95.7
11.3 | 98.425 | 98 | 97.3 | 96.2 | 95.625
11.4 | 98.4 | 97.9625 | 97.25 | 96.1125 | 95.5875
11.5 | 98.4 | 97.9 | 97.2 | 96 | 95.5
11.6 | 98.3375 | 97.9 | 97.1 | 95.9429 | 95.5
11.7 | 98.3 | 97.7714 | 97.1 | 95.9 | 95.4
11.8 | 98.2 | 97.7 | 97.1 | 95.8857 | 95.4
11.9 | 98.2 | 97.6 | 97.02 | 95.8 | 95.325
12 | 98.2 | 97.6 | 96.9571 | 95.76 | 95.2
12.1 | 98.2 | 97.6 | 96.825 | 95.65 | 95.08
12.2 | 98.1375 | 97.5 | 96.8 | 95.48 | 95
12.3 | 98.1 | 97.5 | 96.7571 | 95.4 | 94.9556
12.4 | 98 | 97.44 | 96.7 | 95.4 | 94.9
12.5 | 98 | 97.4 | 96.6 | 95.3867 | 94.8
12.6 | 98 | 97.375 | 96.6 | 95.32 | 94.7375
12.7 | 97.94 | 97.2429 | 96.5 | 95.2562 | 94.6
12.8 | 97.9 | 97.2 | 96.5 | 95.1867 | 94.5
12.9 | 97.84 | 97.1 | 96.4 | 95.0533 | 94.4429
13 | 97.8 | 97.0429 | 96.3714 | 94.96 | 94.3125
13.1 | 97.7667 | 97 | 96.3 | 94.8867 | 94.2857
13.2 | 97.7 | 96.975 | 96.2 | 94.7533 | 94.15
13.3 | 97.6143 | 96.9 | 96.1714 | 94.7 | 94.025
13.4 | 97.52 | 96.8143 | 96.1 | 94.6867 | 94
13.5 | 97.4333 | 96.7857 | 96 | 94.5533 | 93.9
13.6 | 97.4 | 96.7 | 95.8714 | 94.4625 | 93.8286
13.7 | 97.36 | 96.625 | 95.8 | 94.4 | 93.7
13.8 | 97.3 | 96.6 | 95.7333 | 94.3333 | 93.6
13.9 | 97.26 | 96.5571 | 95.7 | 94.2333 | 93.5375
14 | 97.16 | 96.5 | 95.6571 | 94.1 | 93.3
14.1 | 97.1 | 96.4875 | 95.5867 | 94.0231 | 93.1714
14.2 | 97.06 | 96.4 | 95.52 | 93.9125 | 93.1
14.3 | 97 | 96.325 | 95.4533 | 93.8875 | 93
14.4 | 97 | 96.1875 | 95.3867 | 93.7571 | 93
14.5 | 96.9 | 96.1 | 95.32 | 93.625 | 92.9429
14.6 | 96.8 | 96.025 | 95.1727 | 93.6 | 92.8125
14.7 | 96.7 | 96 | 95.1 | 93.5167 | 92.7857
14.8 | 96.6571 | 95.9571 | 95.0533 | 93.42 | 92.6
14.9 | 96.6 | 95.8333 | 94.98 | 93.3 | 92.5
15 | 96.5 | 95.8 | 94.8867 | 93.2 | 92.4
15.1 | 96.4625 | 95.7 | 94.82 | 93.2 | 92.28
15.2 | 96.4 | 95.6286 | 94.7222 | 93.1 | 92.0933
15.3 | 96.3 | 95.6 | 94.7 | 93 | 92.0267
15.4 | 96.2833 | 95.5 | 94.5625 | 92.825 | 91.96
15.5 | 96.2 | 95.4375 | 94.4375 | 92.7429 | 91.89
15.6 | 96.0714 | 95.3 | 94.3 | 92.6 | 91.7867
15.7 | 96 | 95.2 | 94.2 | 92.5 | 91.6533
15.8 | 95.9222 | 95.1375 | 94.2 | 92.4 | 91.54
15.9 | 95.8 | 95.1 | 94.1125 | 92.38 | 91.42
16 | 95.8 | 95.0714 | 93.975 | 92.28 | 91.2938
16.1 | 95.7273 | 94.9375 | 93.9 | 92.18 | 91.2313
16.2 | 95.6125 | 94.8222 | 93.8143 | 91.98 | 91.1333
16.3 | 95.5571 | 94.72 | 93.6857 | 91.8 | 91
16.4 | 95.5 | 94.6467 | 93.5 | 91.8 | 90.8667
16.5 | 95.4 | 94.58 | 93.4 | 91.64 | 90.7667
16.6 | 95.4 | 94.5133 | 93.3 | 91.46 | 90.7
16.7 | 95.2625 | 94.4273 | 93.25 | 91.36 | 90.5667
16.8 | 95.1286 | 94.35 | 93.125 | 91.1667 | 90.4333
16.9 | 95.1 | 94.2625 | 92.9875 | 91.1 | 90.3
17 | 95 | 94.1375 | 92.8571 | 91.1 | 90.1333
17.1 | 94.9167 | 94 | 92.725 | 90.8375 | 89.9
17.2 | 94.8625 | 93.8714 | 92.6875 | 90.8 | 89.8571
17.3 | 94.7 | 93.8 | 92.6 | 90.6 | 89.725
17.4 | 94.6 | 93.7125 | 92.525 | 90.55 | 89.6
17.5 | 94.5625 | 93.575 | 92.2875 | 90.3286 | 89.4714
17.6 | 94.4286 | 93.4429 | 92.1625 | 90.1833 | 89.3375
17.7 | 94.2 | 93.3125 | 92.0286 | 90.1 | 89.2
17.8 | 94.1625 | 93.15 | 91.9 | 90.05 | 89.0714
17.9 | 94.0286 | 93.1 | 91.8 | 89.8857 | 88.9375
18 | 93.9 | 92.975 | 91.6571 | 89.72 | 88.8125
18.1 | 93.8 | 92.8429 | 91.5 | 89.62 | 88.7714
18.2 | 93.6714 | 92.7125 | 91.3625 | 89.45 | 88.55
18.3 | 93.5167 | 92.575 | 91.2375 | 89.3143 | 88.4143
18.4 | 93.375 | 92.4429 | 91.1 | 89.1714 | 88.2875
18.5 | 93.1667 | 92.3125 | 90.925 | 88.95 | 88.1571
18.6 | 93.075 | 92.18 | 90.7375 | 88.8143 | 88.1
18.7 | 92.9429 | 91.98 | 90.6 | 88.675 | 87.9714
18.8 | 92.6375 | 91.7857 | 90.4286 | 88.4571 | 87.7333
18.9 | 92.45 | 91.6 | 90.2714 | 88.325 | 87.5875
19 | 92.2333 | 91.4714 | 90.1286 | 88.1875 | 87.4625
19.1 | 91.94 | 91.22 | 89.9 | 88.0571 | 87.3375
19.2 | 91.54 | 90.9667 | 89.7667 | 87.85 | 87.1
19.3 | 91.18 | 90.75 | 89.5667 | 87.675 | 86.9714
19.4 | 90.68 | 90.38 | 89.3167 | 87.4625 | 86.8
19.5 | 90.04 | 90.225 | 89.0714 | 87.2571 | 86.6167
19.6 | 89.06 | 89.775 | 88.9375 | 87.1 | 86.4625
19.7 | 87.78 | 89.4714 | 88.7 | 86.8625 | 86.2889
19.8 | 85.68 | 89.075 | 88.3714 | 86.7286 | 86.08
19.9 | 82.64 | 88.55 | 88.175 | 86.5222 | 85.9143
20 | 77.98 | 87.9429 | 87.9 | 86.3429 | 85.775
20.1 | 71.7 | 87.075 | 87.65 | 86.1375 | 85.5667
20.2 | 64.18 | 85.875 | 87.3286 | 86.0125 | 85.45
20.3 | 55.94 | 84.3571 | 87.025 | 85.775 | 85.3143
20.4 | 47.0286 | 82.075 | 86.65 | 85.5857 | 85.1714
20.5 | 39.2 | 79.1 | 86.3286 | 85.325 | 84.95
20.6 | 31.8 | 74.5667 | 85.9375 | 85.15 | 84.75
20.7 | 26 | 68.6333 | 85.5429 | 84.8857 | 84.5
20.8 | 21.3 | 61.9 | 85.05 | 84.625 | 84.38
20.9 | 17.9 | 54.3714 | 84.3857 | 84.375 | 84.2
21 | 15.5 | 46.4875 | 83.8143 | 84.2429 | 84.1
21.1 | 13.9667 | 39 | 82.85 | 83.9375 | 83.92
21.2 | 12.9 | 32.6429 | 81.6286 | 83.6857 | 83.7
21.3 | 12.12 | 26.9375 | 80.0857 | 83.55 | 83.5429
21.4 | 11.58 | 22.4 | 77.8 | 83.3286 | 83.275
21.5 | 11.18 | 19.3571 | 74.5714 | 82.9625 | 83
21.6 | 10.84 | 16.9375 | 70.6571 | 82.6143 | 82.875
21.7 | 10.64 | 15.3 | 65.2 | 82.35 | 82.6
21.8 | 10.44 | 14.2 | 59.325 | 82.0625 | 82.475
21.9 | 10.24 | 13.3875 | 52.4375 | 81.7143 | 82.2857
22 | 10.04 | 12.8 | 45.6571 | 81.45 | 82
22.1 | 9.94 | 12.3143 | 39.125 | 81.2 | 81.85
22.2 | 9.7667 | 11.9125 | 33.3125 | 80.8875 | 81.5857
22.3 | 9.6 | 11.625 | 28.4 | 80.4 | 81.4125
22.4 | 9.48 | 11.3714 | 24.5 | 80.0875 | 81.15
22.5 | 9.3833 | 11.15 | 21.6875 | 79.6857 | 80.9
22.6 | 9.2167 | 10.9286 | 19.4286 | 79.3 | 80.6286
22.7 | 9.1 | 10.75 | 17.9 | 78.8857 | 80.45
22.8 | 8.9625 | 10.55 | 16.7375 | 78.2429 | 80.15
22.9 | 8.9 | 10.3286 | 15.8714 | 77.6 | 79.8286
23 | 8.7875 | 10.1857 | 15.3 | 76.9 | 79.6714
23.1 | 8.6625 | 10.05 | 14.6875 | 75.9 | 79.4
23.2 | 8.5286 | 9.9143 | 14.3125 | 74.8125 | 79.1286
23.3 | 8.5 | 9.7857 | 14 | 73.025 | 78.8714
23.4 | 8.4 | 9.6 | 13.6429 | 70.8714 | 78.6
23.5 | 8.3286 | 9.4143 | 13.375 | 67.9125 | 78.4143
23.6 | 8.2714 | 9.2857 | 13 | 64.1 | 77.9571
23.7 | 8.2 | 9.15 | 12.8429 | 59.7667 | 77.6
23.8 | 8.1 | 9.0143 | 12.575 | 54.5308 | 77.4143
23.9 | 8.0625 | 8.9 | 12.4125 | 49 | 77.0625
24 | 8 | 8.85 | 12.175 | 43.4 | 76.6714
24.1 | 7.9 | 8.725 | 11.9857 | 38.3333 | 76.35
24.2 | 7.84 | 8.6 | 11.725 | 33.7667 | 76.1
24.3 | 7.8 | 8.55 | 11.575 | 29.7 | 75.6714
24.4 | 7.7625 | 8.425 | 11.4429 | 26.825 | 75.275
24.5 | 7.7 | 8.3 | 11.225 | 24.3267 | 74.85
24.6 | 7.6875 | 8.3 | 11.05 | 22.3933 | 74.3857
24.7 | 7.6 | 8.225 | 10.8429 | 21.0231 | 73.9
24.8 | 7.5143 | 8.2 | 10.7125 | 19.9 | 73.45
24.9 | 7.5 | 8.1 | 10.5857 | 19.08 | 72.8
25 | 7.425 | 8.0286 | 10.4429 | 18.5 | 71.9429
25.1 | 7.4 | 8 | 10.26 | 18.05 | 71.0429
25.2 | 7.3 | 7.9 | 10.0875 | 17.5714 | 69.7571
25.3 | 7.2143 | 7.8375 | 9.9571 | 17.175 | 68.15
25.4 | 7.2 | 7.8 | 9.825 | 16.9 | 66.08
25.5 | 7.15 | 7.7091 | 9.7857 | 16.5429 | 63.4
25.6 | 7.1 | 7.6 | 9.65 | 16.275 | 60
25.7 | 7.1 | 7.5429 | 9.525 | 16 | 56.02
25.8 | 7 | 7.5 | 9.3875 | 15.75 | 51.48
25.9 | 6.9571 | 7.4111 | 9.2571 | 15.5 | 46.74
26 | 6.9 | 7.3857 | 9.2 | 15.2286 | 42.12
26.1 | 6.825 | 7.3 | 9.0875 | 15.0714 | 37.76
26.2 | 6.7533 | 7.2143 | 9 | 14.8 | 33.88
26.3 | 6.7 | 7.19 | 8.9455 | 14.6143 | 30.5
26.4 | 6.7 | 7.1 | 8.8667 | 14.3714 | 28.0143
26.5 | 6.6231 | 7.1 | 8.8 | 14.1 | 25.925
26.6 | 6.6 | 7.0143 | 8.6667 | 13.925 | 24.2
26.7 | 6.6 | 6.9857 | 8.5687 | 13.78 | 23.1
26.8 | 6.525 | 6.85 | 8.5063 | 13.55 | 22.1
26.9 | 6.5 | 6.8 | 8.38 | 13.3714 | 21.375
27 | 6.4 | 6.8 | 8.3 | 13.1333 | 20.7714
27.1 | 6.4 | 6.7 | 8.2571 | 13 | 20.25
27.2 | 6.3143 | 6.6375 | 8.2 | 12.86 | 19.85
27.3 |  | 6.6 | 8.0875 | 12.7 | 19.5286
27.4 |  | 6.6 | 8 | 12.45 | 19.1375
27.5 |  | 6.5429 | 7.95 | 12.3 | 18.75
27.6 |  | 6.5 | 7.8733 | 12.1714 | 18.5
27.7 |  | 6.4 | 7.8067 | 12 | 18.2286
27.8 |  | 6.4 | 7.7 | 11.8714 | 18.0714
27.9 |  | 6.3545 | 7.6091 | 11.7375 | 17.8
28 |  | 6.3 | 7.5357 | 11.6125 | 17.5286
28.1 |  | 6.2125 | 7.5 | 11.4714 | 17.275
28.2 |  | 6.1889 | 7.46 | 11.3375 | 17.0143
28.3 |  |  | 7.4 | 11.2 | 16.825
28.4 |  |  | 7.3375 | 11.1167 | 16.6625
28.5 |  |  | 7.2 | 11 | 16.3571
28.6 |  |  | 7.1714 | 10.7929 | 16.225
28.7 |  |  | 7.1 | 10.7214 | 15.9875
28.8 |  |  | 7.1 | 10.6462 | 15.825
28.9 |  |  | 7 | 10.5111 | 15.6286
29 |  |  | 6.9444 | 10.4 | 15.4625
29.1 |  |  | 6.8333 | 10.3 | 15.1571
29.2 |  |  | 6.8 | 10.1 | 15
29.3 |  |  | 6.7857 | 10.0375 | 14.825
29.4 |  |  | 6.65 | 9.9 | 14.6286
29.5 |  |  | 6.6 | 9.8 | 14.5
29.6 |  |  | 6.575 | 9.7 | 14.3625
29.7 |  |  |  | 9.6 | 14.1125
29.8 |  |  |  | 9.6 | 13.9
29.9 |  |  |  | 9.4 | 13.8
30 |  |  |  | 9.32 | 13.675
30.1 |  |  |  | 9.24 | 13.5
30.2 |  |  |  | 9.1 | 13.3714
30.3 |  |  |  | 9.04 | 13.175
30.4 |  |  |  | 8.96 | 13.0125
30.5 |  |  |  | 8.9 | 12.975
30.6 |  |  |  | 8.8 | 12.8429
30.7 |  |  |  | 8.7 | 12.7125
30.8 |  |  |  | 8.6 | 12.55
30.9 |  |  |  | 8.5286 | 12.3857
31 |  |  |  | 8.4875 | 12.3
31.1 |  |  |  | 8.37 | 12.1833
31.2 |  |  |  | 8.28 | 12.0286
31.3 |  |  |  | 8.2133 | 11.9
31.4 |  |  |  | 8.1 | 11.7625
31.5 |  |  |  | 8.0167 | 11.7
31.6 |  |  |  | 7.9385 | 11.5
31.7 |  |  |  | 7.8583 | 11.475
31.8 |  |  |  | 7.78 | 11.3429
31.9 |  |  |  | 7.7133 | 11.2125
32 |  |  |  | 7.7 | 11.1429
32.1 |  |  |  | 7.6 | 11
32.2 |  |  |  | 7.58 | 10.9
32.3 |  |  |  | 7.5 | 10.8
32.4 |  |  |  |  | 10.7
32.5 |  |  |  |  | 10.5
32.6 |  |  |  |  | 10.4167
32.7 |  |  |  |  | 10.34
32.8 |  |  |  |  | 10.24
32.9 |  |  |  |  | 10.1
33 |  |  |  |  | 10
33.1 |  |  |  |  | 9.94
33.2 |  |  |  |  | 9.86
33.3 |  |  |  |  | 9.8
33.4 |  |  |  |  | 9.7
33.5 |  |  |  |  | 9.6
33.6 |  |  |  |  | 9.5
33.7 |  |  |  |  | 9.4
33.8 |  |  |  |  | 9.3
33.9 |  |  |  |  | 9.2375
34 |  |  |  |  | 9.13
34.1 |  |  |  |  | 9.0417
34.2 |  |  |  |  | 8.9615
34.3 |  |  |  |  | 8.9
34.4 |  |  |  |  | 8.9
34.5 |  |  |  |  | 8.79
34.6 |  |  |  |  | 8.6933
34.7 |  |  |  |  | 8.6267
34.8 |  |  |  |  | 8.525
34.9 |  |  |  |  | 8.4
35 |  |  |  |  | 8.4

## Wedge 45 Output Factors
Output factors
Wedge 45 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15 | 20
| 3 | 0.89 | 0.903 | 0.91 | 0.918 | 0.924 | 0.93 | 0.933
| 4 | 0.903 | 0.917 | 0.926 | 0.936 | 0.944 | 0.951 | 0.956
| 5 | 0.91 | 0.926 | 0.936 | 0.949 | 0.958 | 0.967 | 0.973
| 7 | 0.921 | 0.94 | 0.952 | 0.967 | 0.98 | 0.991 | 1
| 10 | 0.93 | 0.951 | 0.966 | 0.984 | 1 | 1.017 | 1.029
| 15 | 0.938 | 0.962 | 0.977 | 1 | 1.02 | 1.042 | 1.059
| 20 | 0.942 | 0.967 | 0.985 | 1.009 | 1.033 | 1.059 | 1.08
| 30 | 0.949 | 0.975 | 0.995 | 1.022 | 1.051 | 1.084 | 1.11
| 40 | 0.953 | 0.98 | 1.001 | 1.03 | 1.061 | 1.098 | 1.127

## Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=100cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 30.9 | 36.9 | 41.9 | 49.8
0.1 | 35 | 40.95 | 45.8714 | 53.425
0.2 | 41.18 | 47.15 | 51.3 | 58.8
0.3 | 51.36 | 56.5 | 60.325 | 66.18
0.4 | 60.54 | 65.1 | 68.3333 | 74
0.5 | 68.5 | 71.8 | 75 | 79.6429
0.6 | 74.4 | 77.64 | 80.2 | 84.15
0.7 | 79.6 | 82.4 | 84.4 | 87.6
0.8 | 84.125 | 86.1333 | 88.1 | 90.64
0.9 | 87.6333 | 89.38 | 90.75 | 93.025
1 | 90.42 | 91.84 | 92.96 | 94.9
1.1 | 92.74 | 93.88 | 94.8 | 96.2571
1.2 | 94.7 | 95.42 | 96.4 | 97.48
1.3 | 96.15 | 96.7833 | 97.4 | 98.35
1.4 | 97.4 | 97.88 | 98.4 | 98.925
1.5 | 98.225 | 98.7143 | 98.95 | 99.425
1.6 | 99.15 | 99.2 | 99.375 | 99.675
1.7 | 99.45 | 99.5 | 99.8 | 99.925
1.8 | 99.8 | 99.95 | 99.9 | 100
1.9 | 99.9 | 100 | 100 | 100
2 | 100 | 100 | 100 | 100
2.1 | 100 | 99.9 | 99.9 | 99.9
2.2 | 100 | 99.9 | 99.8 | 99.6167
2.3 | 99.8 | 99.75 | 99.7 | 99.475
2.4 | 99.5667 | 99.5 | 99.44 | 99.2
2.5 | 99.3857 | 99.2 | 99.1 | 99.0333
2.6 | 99.1 | 98.95 | 98.9 | 98.72
2.7 | 98.65 | 98.6 | 98.5 | 98.42
2.8 | 98.3 | 98.3 | 98.2125 | 98.02
2.9 | 97.9 | 97.8667 | 97.85 | 97.72
3 | 97.5 | 97.62 | 97.55 | 97.3333
3.1 | 97.125 | 97.18 | 97.2429 | 97.06
3.2 | 96.7 | 96.8 | 96.8143 | 96.55
3.3 | 96.2333 | 96.4571 | 96.34 | 96.3
3.4 | 95.8 | 96.0286 | 96 | 95.925
3.5 | 95.2 | 95.7 | 95.72 | 95.6667
3.6 | 94.85 | 95.25 | 95.34 | 95.3429
3.7 | 94.4333 | 94.9 | 95.02 | 94.94
3.8 | 94.06 | 94.5143 | 94.6333 | 94.55
3.9 | 93.52 | 94.15 | 94.075 | 94.12
4 | 93.1 | 93.8 | 93.8333 | 93.775
4.1 | 92.54 | 93.3 | 93.4667 | 93.525
4.2 | 92.1 | 92.9571 | 93.1 | 93.075
4.3 | 91.75 | 92.72 | 92.74 | 92.8
4.4 | 91.3 | 92.25 | 92.3 | 92.4
4.5 | 90.6667 | 91.7 | 92 | 92.0333
4.6 | 90.3667 | 91.3667 | 91.55 | 91.6
4.7 | 90.02 | 91.04 | 91.26 | 91.15
4.8 | 89.48 | 90.64 | 90.9 | 90.9
4.9 | 89.0667 | 90.275 | 90.54 | 90.55
5 | 88.55 | 89.9286 | 90.16 | 90.2333
5.1 | 88 | 89.5 | 89.8375 | 89.86
5.2 | 87.72 | 89.1 | 89.5 | 89.4143
5.3 | 87.25 | 88.6429 | 89.1 | 89.075
5.4 | 86.82 | 88.4 | 88.8 | 88.72
5.5 | 86.56 | 87.9714 | 88.45 | 88.4667
5.6 | 86.12 | 87.6 | 88.1 | 88.2
5.7 | 85.65 | 87.2 | 87.6333 | 87.74
5.8 | 85.075 | 86.92 | 87.275 | 87.3286
5.9 | 84.6 | 86.4 | 87.05 | 87.1
6 | 84.15 | 86.075 | 86.58 | 86.6571
6.1 | 83.825 | 85.65 | 86.18 | 86.32
6.2 | 83.4 | 85.34 | 85.9571 | 85.95
6.3 | 83 | 84.9857 | 85.6429 | 85.6857
6.4 | 82.45 | 84.625 | 85.2 | 85.32
6.5 | 82.075 | 84.2667 | 84.675 | 84.9375
6.6 | 81.82 | 83.98 | 84.4 | 84.64
6.7 | 81.4375 | 83.5333 | 84.2 | 84.3571
6.8 | 80.9 | 83.075 | 83.8 | 84
6.9 | 80.4286 | 82.68 | 83.45 | 83.62
7 | 80.1 | 82.2286 | 83.1 | 83.2667
7.1 | 79.65 | 81.95 | 82.6667 | 82.95
7.2 | 79.1 | 81.55 | 82.275 | 82.575
7.3 | 78.775 | 81.175 | 82.1 | 82.34
7.4 | 78.425 | 80.9 | 81.7333 | 81.94
7.5 | 77.8333 | 80.35 | 81.3 | 81.6
7.6 | 77.5 | 80.1 | 81.05 | 81.34
7.7 | 77.1 | 79.72 | 80.7667 | 80.95
7.8 | 76.66 | 79.35 | 80.26 | 80.6
7.9 | 76.2667 | 78.96 | 79.9667 | 80.2333
8 | 75.9 | 78.725 | 79.6333 | 79.86
8.1 | 75.55 | 78.34 | 79.18 | 79.7
8.2 | 75.125 | 77.9571 | 79.05 | 79.3
8.3 | 74.8 | 77.65 | 78.8 | 78.975
8.4 | 74.3857 | 77.24 | 78.35 | 78.675
8.5 | 73.96 | 76.875 | 78 | 78.4
8.6 | 73.45 | 76.45 | 77.65 | 78.1
8.7 | 73.05 | 76.175 | 77.3 | 77.7
8.8 | 72.8333 | 75.94 | 77 | 77.2667
8.9 | 72.55 | 75.6 | 76.7 | 77.0667
9 | 72.1 | 75.0429 | 76.25 | 76.58
9.1 | 71.8 | 74.8 | 75.9333 | 76.3667
9.2 | 71.4667 | 74.5 | 75.65 | 76.0571
9.3 | 71.075 | 74.1 | 75.34 | 75.78
9.4 | 70.8 | 73.875 | 75 | 75.4667
9.5 | 70.26 | 73.525 | 74.6714 | 75.1
9.6 | 69.9 | 73.2333 | 74.46 | 74.86
9.7 | 69.54 | 72.8125 | 74.15 | 74.625
9.8 | 69.15 | 72.52 | 73.8167 | 74.3
9.9 | 68.7 | 72.16 | 73.5 | 74.0167
10 | 68.3 | 71.85 | 73.1714 | 73.7
10.1 | 68.15 | 71.55 | 72.775 | 73.2667
10.2 | 67.6333 | 71.22 | 72.5 | 73.02
10.3 | 67.3 | 70.92 | 72.3 | 72.64
10.4 | 67.1 | 70.5667 | 71.95 | 72.25
10.5 | 66.7 | 70.15 | 71.6 | 72.1
10.6 | 66.35 | 69.75 | 71.25 | 71.84
10.7 | 65.9 | 69.5667 | 70.88 | 71.475
10.8 | 65.54 | 69.2143 | 70.55 | 71.225
10.9 | 65.26 | 68.8571 | 70.2 | 71
11 | 64.95 | 68.68 | 70.0143 | 70.75
11.1 | 64.5 | 68.3 | 69.7 | 70.2889
11.2 | 64.2 | 67.98 | 69.2857 | 69.98
11.3 | 63.92 | 67.675 | 69.0667 | 69.75
11.4 | 63.58 | 67.28 | 68.78 | 69.4
11.5 | 63.36 | 66.925 | 68.5 | 69.15
11.6 | 63 | 66.7143 | 68.2 | 68.95
11.7 | 62.58 | 66.32 | 67.8 | 68.48
11.8 | 62.15 | 66.04 | 67.5333 | 68.225
11.9 | 62 | 65.8 | 67.2 | 68.0429
12 | 61.6 | 65.35 | 66.8857 | 67.66
12.1 | 61.325 | 65.1 | 66.72 | 67.35
12.2 | 60.975 | 64.8 | 66.3125 | 67.14
12.3 | 60.65 | 64.4333 | 66.0667 | 66.825
12.4 | 60.175 | 64.2 | 65.8333 | 66.5667
12.5 | 59.8 | 63.92 | 65.5 | 66.1667
12.6 | 59.65 | 63.575 | 65.25 | 66
12.7 | 59.3143 | 63.3 | 65 | 65.6
12.8 | 59.0286 | 62.95 | 64.7 | 65.35
12.9 | 58.6 | 62.6667 | 64.32 | 65.16
13 | 58.4 | 62.4 | 64.025 | 64.8333
13.1 | 58.05 | 62.1 | 63.7667 | 64.5667
13.2 | 57.76 | 61.78 | 63.5333 | 64.3333
13.3 | 57.44 | 61.5667 | 63.28 | 64
13.4 | 57.2667 | 61.2 | 62.95 | 63.725
13.5 | 56.8 | 60.9667 | 62.6 | 63.4
13.6 | 56.5 | 60.6 | 62.3667 | 63.15
13.7 | 56.225 | 60.3 | 62.0571 | 62.925
13.8 | 55.9667 | 60.075 | 61.9 | 62.65
13.9 | 55.725 | 59.84 | 61.5 | 62.2667
14 | 55.325 | 59.6 | 61.28 | 62.075
14.1 | 55.05 | 59.3 | 60.9 | 61.84
14.2 | 54.7875 | 58.95 | 60.7 | 61.64
14.3 | 54.5 | 58.62 | 60.45 | 61.2857
14.4 | 54.25 | 58.3 | 60.1333 | 61.04
14.5 | 53.875 | 58.05 | 59.9667 | 60.8
14.6 | 53.6 | 57.8 | 59.575 | 60.55
14.7 | 53.4333 | 57.5 | 59.35 | 60.32
14.8 | 53.2 | 57.275 | 59.1286 | 59.9
14.9 | 52.7857 | 56.92 | 58.88 | 59.74
15 | 52.5 | 56.7625 | 58.62 | 59.6
15.1 | 52.225 | 56.425 | 58.3444 | 59.26
15.2 | 52.05 | 56.2 | 58.04 | 59
15.3 | 51.76 | 55.95 | 57.8 | 58.775
15.4 | 51.55 | 55.62 | 57.58 | 58.5
15.5 | 51.1667 | 55.3667 | 57.325 | 58.2143
15.6 | 50.8 | 55.2 | 57.1 | 57.95
15.7 | 50.675 | 54.8 | 56.7714 | 57.7
15.8 | 50.35 | 54.64 | 56.5 | 57.5
15.9 | 50.1 | 54.3667 | 56.3 | 57.2667
16 | 49.95 | 54.1 | 56.05 | 56.94
16.1 | 49.7 | 53.8 | 55.8 | 56.7429
16.2 | 49.35 | 53.6 | 55.5 | 56.38
16.3 | 49.175 | 53.3333 | 55.225 | 56.2
16.4 | 48.85 | 53.0857 | 55.1 | 56
16.5 | 48.6 | 52.8333 | 54.8125 | 55.74
16.6 | 48.4 | 52.48 | 54.575 | 55.58
16.7 | 48.08 | 52.3 | 54.3571 | 55.3286
16.8 | 47.76 | 52.05 | 54.05 | 55
16.9 | 47.58 | 51.8 | 53.8 | 54.8333
17 | 47.32 | 51.62 | 53.6 | 54.6
17.1 | 47.1 | 51.32 | 53.4 | 54.35
17.2 | 46.84 | 51.0571 | 53.15 | 54.04
17.3 | 46.62 | 50.82 | 52.9 | 53.8
17.4 | 46.38 | 50.425 | 52.5 | 53.55
17.5 | 46.0429 | 50.2 | 52.3 | 53.4
17.6 | 45.8286 | 49.98 | 52.12 | 53.12
17.7 | 45.58 | 49.75 | 51.92 | 52.98
17.8 | 45.4 | 49.52 | 51.68 | 52.6429
17.9 | 45.1286 | 49.325 | 51.425 | 52.4286
18 | 44.84 | 49.15 | 51.1 | 52.14
18.1 | 44.62 | 48.85 | 50.95 | 52.0111
18.2 | 44.45 | 48.6 | 50.6667 | 51.78
18.3 | 44.24 | 48.38 | 50.4 | 51.5714
18.4 | 43.95 | 48.2333 | 50.28 | 51.3571
18.5 | 43.7 | 47.925 | 50.1143 | 51.1
18.6 | 43.5 | 47.72 | 49.85 | 50.9
18.7 | 43.3 | 47.475 | 49.52 | 50.6
18.8 | 43.15 | 47.2 | 49.4 | 50.45
18.9 | 42.9 | 47.04 | 49.1 | 50.2
19 | 42.675 | 46.9 | 48.9 | 50.075
19.1 | 42.325 | 46.55 | 48.7 | 49.825
19.2 | 42.2 | 46.3 | 48.46 | 49.66
19.3 | 41.9333 | 46.18 | 48.225 | 49.4625
19.4 | 41.64 | 45.95 | 48.075 | 49.2
19.5 | 41.5 | 45.7 | 47.84 | 48.9
19.6 | 41.3 | 45.45 | 47.66 | 48.72
19.7 | 41.125 | 45.3333 | 47.4 | 48.5
19.8 | 40.96 | 45.14 | 47.2 | 48.275
19.9 | 40.7571 | 44.8714 | 46.9 | 48.025
20 | 40.6 | 44.6 | 46.78 | 47.8
20.1 | 40.3 | 44.4 | 46.5333 | 47.7
20.2 | 40.075 | 44.22 | 46.275 | 47.475
20.3 | 39.8714 | 44 | 46.15 | 47.2333
20.4 | 39.68 | 43.8 | 46 | 46.975
20.5 | 39.4 | 43.66 | 45.65 | 46.8714
20.6 | 39.24 | 43.425 | 45.35 | 46.6667
20.7 | 38.94 | 43.22 | 45.2333 | 46.4
20.8 | 38.8 | 43.06 | 44.98 | 46.275
20.9 | 38.6 | 42.76 | 44.8 | 46.1
21 | 38.4 | 42.525 | 44.7 | 45.825
21.1 | 38.2 | 42.3333 | 44.5 | 45.5571
21.2 | 38.025 | 42.1 | 44.3 | 45.4
21.3 | 37.775 | 41.875 | 43.95 | 45.25
21.4 | 37.55 | 41.7 | 43.74 | 45.04
21.5 | 37.3857 | 41.48 | 43.66 | 44.8429
21.6 | 37.2 | 41.375 | 43.55 | 44.6
21.7 | 37.1 | 41.14 | 43.2 | 44.4
21.8 | 36.86 | 40.9667 | 43.06 | 44.25
21.9 | 36.66 | 40.72 | 42.8571 | 44.1
22 | 36.475 | 40.58 | 42.7143 | 43.85
22.1 | 36.225 | 40.3286 | 42.44 | 43.62
22.2 | 36.1 | 40.2 | 42.2 | 43.48
22.3 | 35.975 | 40 | 42.1 | 43.3
22.4 | 35.775 | 39.76 | 41.9 | 43.1
22.5 | 35.5 | 39.56 | 41.8 | 42.9
22.6 | 35.3333 | 39.3 | 41.52 | 42.725
22.7 | 35.18 | 39.2 | 41.425 | 42.56
22.8 | 35 | 39 | 41.18 | 42.38
22.9 | 34.8 | 38.875 | 41 | 42.16
23 | 34.6667 | 38.65 | 40.8 | 41.9
23.1 | 34.56 | 38.425 | 40.55 | 41.825
23.2 | 34.3667 | 38.3 | 40.4111 | 41.62
23.3 | 34.1667 | 38.1 | 40.18 | 41.425
23.4 | 33.94 | 37.8 | 40 | 41.225
23.5 | 33.8429 | 37.68 | 39.8 | 41.025
23.6 | 33.625 | 37.5 | 39.6 | 40.9
23.7 | 33.3833 | 37.4 | 39.54 | 40.7
23.8 | 33.2 | 37.2429 | 39.3 | 40.5
23.9 | 33.1 | 37.1 | 39.14 | 40.38
24 | 32.9 | 36.7833 | 38.98 | 40.1
24.1 | 32.8 | 36.6 | 38.7 | 39.9667
24.2 | 32.65 | 36.5 | 38.6 | 39.78
24.3 | 32.425 | 36.3 | 38.425 | 39.6
24.4 | 32.26 | 36.1 | 38.2 | 39.4
24.5 | 32.025 | 35.9 | 38.1 | 39.3167
24.6 | 31.9 | 35.725 | 37.925 | 39.1
24.7 | 31.75 | 35.6 | 37.7333 | 38.8833
24.8 | 31.7 | 35.4 | 37.62 | 38.8
24.9 | 31.45 | 35.4 | 37.4714 | 38.65
25 | 31.3 | 35.15 | 37.275 | 38.42
25.1 | 31.12 | 34.95 | 37.05 | 38.225
25.2 | 31 | 34.7778 | 36.825 | 38.1
25.3 | 30.82 | 34.6 | 36.6 | 38
25.4 | 30.675 | 34.44 | 36.5 | 37.85
25.5 | 30.44 | 34.25 | 36.4 | 37.5714
25.6 | 30.325 | 34.14 | 36.24 | 37.48
25.7 | 30.2429 | 33.9 | 36 | 37.26
25.8 | 30.1 | 33.775 | 35.9 | 37.1
25.9 | 29.975 | 33.65 | 35.76 | 36.925
26 | 29.8 | 33.4857 | 35.5571 | 36.7
26.1 | 29.65 | 33.3 | 35.4 | 36.6
26.2 | 29.4571 | 33.125 | 35.275 | 36.4778
26.3 | 29.25 | 33.1 | 35.025 | 36.3
26.4 | 29.1 | 32.8333 | 34.9 | 36.1
26.5 | 29 | 32.7333 | 34.75 | 35.95
26.6 | 28.84 | 32.5 | 34.5667 | 35.8
26.7 | 28.625 | 32.32 | 34.4 | 35.65
26.8 | 28.575 | 32.225 | 34.2667 | 35.5
26.9 | 28.4333 | 32.05 | 34.1 | 35.28
27 | 28.3 | 31.9 | 33.925 | 35.16
27.1 | 28.15 | 31.725 | 33.775 | 35
27.2 | 28 | 31.5333 | 33.64 | 34.875
27.3 | 27.8333 | 31.4 | 33.5 | 34.725
27.4 | 27.7 | 31.2714 | 33.3 | 34.6
27.5 | 27.58 | 31.1286 | 33.2 | 34.4
27.6 | 27.46 | 31 | 33.0143 | 34.25
27.7 | 27.3 | 30.9 | 32.9 | 34.0833
27.8 | 27.14 | 30.7 | 32.7857 | 33.9
27.9 | 26.95 | 30.525 | 32.7 | 33.8
28 | 26.88 | 30.4 | 32.5 | 33.7
28.1 | 26.7 | 30.25 | 32.3222 | 33.55
28.2 | 26.575 | 30.0667 | 32.2 | 33.325
28.3 | 26.475 | 30 | 32.0143 | 33.28
28.4 | 26.35 | 29.9 | 31.9 | 33.1143
28.5 | 26.2 | 29.7 | 31.7 | 32.95
28.6 | 26.1 | 29.55 | 31.6111 | 32.825
28.7 | 26 | 29.4 | 31.5 | 32.6
28.8 | 25.8 | 29.26 | 31.4 | 32.5
28.9 | 25.7143 | 29.1 | 31.2 | 32.4
29 | 25.55 | 29 | 31 | 32.22
29.1 | 25.4625 | 28.9 | 30.9 | 32.1
29.2 | 25.3 | 28.7286 | 30.7 | 31.92
29.3 | 25.2 | 28.5 | 30.6 | 31.8
29.4 | 25.1 | 28.4 | 30.48 | 31.7
29.5 | 24.95 | 28.3 | 30.325 | 31.525
29.6 | 24.8 | 28.22 | 30.125 | 31.4
29.7 | 24.65 | 28.1 | 30.1 | 31.3
29.8 | 24.5 | 27.925 | 29.9 | 31.1
29.9 | 24.4 | 27.775 | 29.75 | 30.95
30 | 24.3 | 27.7 | 29.5833 | 30.74

## SSD80 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=80cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 32.1 | 40.1 | 47.4 | 58
0.1 | 36.15 | 43.8 | 51.075 | 61.74
0.2 | 42.02 | 49.9 | 56.4 | 66.46
0.3 | 52.76 | 59.5 | 65.275 | 73.325
0.4 | 61.7 | 67.2 | 72.6 | 79.8
0.5 | 69.2714 | 74.32 | 78.95 | 84.88
0.6 | 75.7 | 79.78 | 83.8 | 88.65
0.7 | 80.8857 | 84.4714 | 87.6 | 91.6
0.8 | 85.05 | 88.11 | 90.8 | 93.9
0.9 | 88.3778 | 90.93 | 93.1 | 95.8
1 | 91.375 | 93.1875 | 95.1 | 97.2
1.1 | 93.4778 | 95.05 | 96.6 | 98.125
1.2 | 95.22 | 96.52 | 97.6 | 98.84
1.3 | 96.7143 | 97.68 | 98.55 | 99.4
1.4 | 97.85 | 98.4636 | 99.2 | 99.825
1.5 | 98.66 | 99.1 | 99.55 | 100
1.6 | 99.35 | 99.5 | 99.9 | 100
1.7 | 99.82 | 99.8 | 99.94 | 100
1.8 | 99.94 | 99.9 | 100 | 99.9833
1.9 | 100 | 100 | 100 | 99.8
2 | 100 | 99.9833 | 99.86 | 99.6
2.1 | 99.9375 | 99.8 | 99.725 | 99.35
2.2 | 99.825 | 99.6 | 99.42 | 98.96
2.3 | 99.6 | 99.4 | 99.12 | 98.6429
2.4 | 99.35 | 99.1 | 98.76 | 98.375
2.5 | 98.8 | 98.7 | 98.38 | 97.925
2.6 | 98.475 | 98.25 | 98.08 | 97.58
2.7 | 98.1 | 97.96 | 97.65 | 97.24
2.8 | 97.675 | 97.525 | 97.26 | 96.84
2.9 | 97.2857 | 97.14 | 96.86 | 96.4182
3 | 96.8 | 96.74 | 96.46 | 95.95
3.1 | 96.2143 | 96.26 | 96.04 | 95.51
3.2 | 95.7143 | 95.78 | 95.6 | 95.18
3.3 | 95.26 | 95.4 | 95.2 | 94.78
3.4 | 94.7 | 94.94 | 94.6 | 94.3875
3.5 | 94.25 | 94.5 | 94.3 | 93.9857
3.6 | 93.7 | 94 | 93.9 | 93.5
3.7 | 93.1571 | 93.58 | 93.4 | 93.08
3.8 | 92.74 | 93.18 | 93 | 92.6889
3.9 | 92.2429 | 92.7 | 92.5 | 92.15
4 | 91.66 | 92.28 | 92.1 | 91.76
4.1 | 91.12 | 91.8 | 91.7 | 91.38
4.2 | 90.725 | 91.38 | 91.25 | 90.92
4.3 | 90.3 | 90.92 | 90.88 | 90.5333
4.4 | 89.78 | 90.46 | 90.42 | 90.175
4.5 | 89.3 | 90.02 | 89.96 | 89.7
4.6 | 88.7 | 89.5667 | 89.52 | 89.36
4.7 | 88.1667 | 89.24 | 89.1333 | 88.95
4.8 | 87.8 | 88.74 | 88.72 | 88.44
4.9 | 87.3 | 88.26 | 88.32 | 88.04
5 | 86.92 | 87.86 | 87.8333 | 87.6333
5.1 | 86.275 | 87.44 | 87.55 | 87.24
5.2 | 85.84 | 86.9667 | 87.1143 | 86.85
5.3 | 85.35 | 86.6 | 86.7 | 86.46
5.4 | 84.875 | 86.1 | 86.2 | 86.1
5.5 | 84.36 | 85.7 | 85.8333 | 85.7
5.6 | 83.96 | 85.3 | 85.42 | 85.3
5.7 | 83.44 | 84.9 | 85.02 | 85
5.8 | 82.96 | 84.5 | 84.62 | 84.5
5.9 | 82.55 | 83.9 | 84.26 | 84.1
6 | 82.06 | 83.6 | 83.86 | 83.7
6.1 | 81.7 | 83.1333 | 83.44 | 83.3
6.2 | 81.18 | 82.7 | 83.04 | 82.8
6.3 | 80.65 | 82.4333 | 82.56 | 82.5
6.4 | 80.1 | 81.86 | 82.26 | 82.14
6.5 | 79.76 | 81.44 | 81.8 | 81.74
6.6 | 79.35 | 81.04 | 81.4 | 81.34
6.7 | 78.82 | 80.65 | 81 | 80.94
6.8 | 78.3 | 80.2333 | 80.7 | 80.62
6.9 | 77.9 | 79.8333 | 80.3 | 80.16
7 | 77.5 | 79.3667 | 79.9 | 79.78
7.1 | 77 | 79.0333 | 79.4 | 79.46
7.2 | 76.62 | 78.6 | 79.0333 | 79.06
7.3 | 76.18 | 78.25 | 78.6667 | 78.66
7.4 | 75.7571 | 77.8333 | 78.36 | 78.2667
7.5 | 75.2 | 77.3667 | 77.9333 | 77.96
7.6 | 74.7857 | 77.0333 | 77.54 | 77.6
7.7 | 74.4 | 76.65 | 77.14 | 77.2
7.8 | 74.05 | 76.15 | 76.74 | 76.8
7.9 | 73.62 | 75.8667 | 76.34 | 76.3833
8 | 73.14 | 75.45 | 75.94 | 76.1
8.1 | 72.7333 | 75.05 | 75.6667 | 75.7
8.2 | 72.3 | 74.6667 | 75.25 | 75.32
8.3 | 71.8667 | 74.28 | 74.9 | 75.02
8.4 | 71.54 | 73.9 | 74.58 | 74.6667
8.5 | 71.05 | 73.5 | 74.14 | 74.3333
8.6 | 70.74 | 73.2 | 73.84 | 74.0222
8.7 | 70.3 | 72.8 | 73.54 | 73.6333
8.8 | 69.8667 | 72.4 | 73.18 | 73.24
8.9 | 69.5 | 72 | 72.78 | 72.92
9 | 69.04 | 71.625 | 72.38 | 72.6
9.1 | 68.72 | 71.28 | 71.98 | 72.24
9.2 | 68.36 | 70.94 | 71.74 | 71.88
9.3 | 67.875 | 70.54 | 71.4 | 71.48
9.4 | 67.58 | 70.15 | 71.06 | 71.18
9.5 | 67.25 | 69.7429 | 70.82 | 70.88
9.6 | 66.76 | 69.48 | 70.3 | 70.5
9.7 | 66.425 | 69.08 | 69.9 | 70.2
9.8 | 66.04 | 68.68 | 69.6 | 69.82
9.9 | 65.7 | 68.34 | 69.3 | 69.52
10 | 65.3 | 67.98 | 69 | 69.22
10.1 | 64.8 | 67.68 | 68.54 | 68.86
10.2 | 64.5 | 67.28 | 68.16 | 68.52
10.3 | 64 | 66.88 | 67.84 | 68.22
10.4 | 63.8 | 66.56 | 67.62 | 67.82
10.5 | 63.3429 | 66.18 | 67.24 | 67.46
10.6 | 63.1 | 65.85 | 66.86 | 67.22
10.7 | 62.72 | 65.62 | 66.5714 | 66.86
10.8 | 62.3 | 65.22 | 66.25 | 66.575
10.9 | 61.92 | 64.82 | 65.96 | 66.18
11 | 61.56 | 64.56 | 65.6571 | 65.96
11.1 | 61.24 | 64.18 | 65.25 | 65.66
11.2 | 61.02 | 63.8 | 64.8429 | 65.36
11.3 | 60.65 | 63.5167 | 64.53 | 64.98
11.4 | 60.22 | 63.22 | 64.32 | 64.7
11.5 | 59.85 | 62.84 | 63.9727 | 64.38
11.6 | 59.52 | 62.52 | 63.66 | 64
11.7 | 59.18 | 62.14 | 63.28 | 63.76
11.8 | 58.875 | 61.78 | 62.96 | 63.46
11.9 | 58.5 | 61.54 | 62.66 | 63.14
12 | 58.18 | 61.18 | 62.36 | 62.8
12.1 | 57.82 | 60.88 | 62.06 | 62.5
12.2 | 57.46 | 60.6 | 61.68 | 62.225
12.3 | 57.16 | 60.24 | 61.3818 | 61.8857
12.4 | 56.78 | 59.9 | 61.16 | 61.6571
12.5 | 56.52 | 59.56 | 60.82 | 61.2857
12.6 | 56.1571 | 59.29 | 60.49 | 61.06
12.7 | 55.875 | 58.96 | 60.225 | 60.7
12.8 | 55.6444 | 58.66 | 59.92 | 60.3375
12.9 | 55.15 | 58.32 | 59.54 | 60.125
13 | 54.8429 | 57.96 | 59.28 | 59.84
13.1 | 54.575 | 57.7714 | 59.04 | 59.54
13.2 | 54.275 | 57.44 | 58.74 | 59.24
13.3 | 53.96 | 57.0714 | 58.38 | 59.04
13.4 | 53.65 | 56.7857 | 58.14 | 58.68
13.5 | 53.3333 | 56.56 | 57.76 | 58.45
13.6 | 53.04 | 56.26 | 57.54 | 58.2
13.7 | 52.7 | 55.98 | 57.25 | 57.9
13.8 | 52.4667 | 55.68 | 57 | 57.6
13.9 | 52.2 | 55.42 | 56.6333 | 57.3
14 | 51.96 | 55.075 | 56.3 | 57.1
14.1 | 51.55 | 54.74 | 56.0667 | 56.8
14.2 | 51.34 | 54.4143 | 55.88 | 56.5
14.3 | 51 | 54.2 | 55.55 | 56.22
14.4 | 50.76 | 53.9 | 55.28 | 55.9333
14.5 | 50.4667 | 53.7 | 55.04 | 55.7333
14.6 | 50.18 | 53.3 | 54.6667 | 55.425
14.7 | 49.9333 | 53 | 54.54 | 55.16
14.8 | 49.6 | 52.7 | 54.24 | 54.9333
14.9 | 49.3667 | 52.5 | 53.975 | 54.66
15 | 49.06 | 52.3 | 53.725 | 54.3333
15.1 | 48.7333 | 52 | 53.45 | 54.0333
15.2 | 48.44 | 51.7 | 53.16 | 53.8
15.3 | 48.225 | 51.4 | 52.86 | 53.5
15.4 | 47.9 | 51.1 | 52.56 | 53.3
15.5 | 47.65 | 50.8 | 52.26 | 53.05
15.6 | 47.3 | 50.6 | 52.02 | 52.8
15.7 | 47.05 | 50.3167 | 51.8 | 52.525
15.8 | 46.8 | 50.04 | 51.6 | 52.25
15.9 | 46.6 | 49.82 | 51.34 | 52
16 | 46.36 | 49.54 | 51.02 | 51.7
16.1 | 46.12 | 49.32 | 50.74 | 51.5
16.2 | 45.8 | 49.04 | 50.52 | 51.2667
16.3 | 45.5167 | 48.74 | 50.32 | 51
16.4 | 45.32 | 48.52 | 50.02 | 50.78
16.5 | 45 | 48.32 | 49.72 | 50.54
16.6 | 44.86 | 48 | 49.52 | 50.28
16.7 | 44.525 | 47.7333 | 49.32 | 50.04
16.8 | 44.24 | 47.52 | 49.12 | 49.825
16.9 | 43.98 | 47.2667 | 48.84 | 49.54
17 | 43.8 | 47.06 | 48.52 | 49.28
17.1 | 43.5333 | 46.78 | 48.2667 | 49.1
17.2 | 43.28 | 46.48 | 47.98 | 48.78
17.3 | 43 | 46.2667 | 47.78 | 48.4833
17.4 | 42.74 | 46.0667 | 47.54 | 48.3
17.5 | 42.6 | 45.6667 | 47.3 | 48.1
17.6 | 42.2571 | 45.425 | 47.06 | 47.9
17.7 | 42.05 | 45.24 | 46.78 | 47.6
17.8 | 41.78 | 44.9667 | 46.58 | 47.4
17.9 | 41.525 | 44.7 | 46.28 | 47.2
18 | 41.35 | 44.48 | 46.08 | 46.9
18.1 | 41.125 | 44.3 | 45.88 | 46.8
18.2 | 40.84 | 44.125 | 45.6667 | 46.5
18.3 | 40.64 | 43.8 | 45.4 | 46.3
18.4 | 40.42 | 43.7 | 45.2 | 46.1
18.5 | 40.3 | 43.4 | 44.9 | 45.9
18.6 | 40.02 | 43.2 | 44.8 | 45.6
18.7 | 39.7857 | 42.9 | 44.5 | 45.4
18.8 | 39.525 | 42.7 | 44.35 | 45.2
18.9 | 39.375 | 42.5 | 44.05 | 45
19 | 39.14 | 42.3 | 43.9 | 44.74
19.1 | 39 | 42.1 | 43.66 | 44.54
19.2 | 38.8 | 41.84 | 43.46 | 44.3
19.3 | 38.4857 | 41.6 | 43.16 | 44.06
19.4 | 38.3 | 41.42 | 43 | 43.86
19.5 | 38.06 | 41.22 | 42.86 | 43.66
19.6 | 37.85 | 40.92 | 42.56 | 43.475
19.7 | 37.675 | 40.74 | 42.36 | 43.28
19.8 | 37.4 | 40.52 | 42.16 | 43.08
19.9 | 37.2333 | 40.32 | 41.96 | 42.8286
20 | 37 | 40.02 | 41.78 | 42.53
20.1 | 36.8333 | 39.95 | 41.48 | 42.41
20.2 | 36.62 | 39.7714 | 41.3 | 42.2
20.3 | 36.4667 | 39.5333 | 41.16 | 41.95
20.4 | 36.2429 | 39.3 | 40.88 | 41.825
20.5 | 35.94 | 39.0667 | 40.68 | 41.61
20.6 | 35.8111 | 38.94 | 40.48 | 41.3625
20.7 | 35.6667 | 38.64 | 40.28 | 41.2286
20.8 | 35.42 | 38.4625 | 40.08 | 40.96
20.9 | 35.2667 | 38.28 | 39.88 | 40.74
21 | 35.02 | 38.08 | 39.68 | 40.54
21.1 | 34.8667 | 37.8833 | 39.48 | 40.42
21.2 | 34.625 | 37.7 | 39.3 | 40.1818
21.3 | 34.4 | 37.5286 | 39.1 | 39.98
21.4 | 34.26 | 37.3 | 38.9 | 39.86
21.5 | 34 | 37 | 38.7286 | 39.66
21.6 | 33.88 | 36.9 | 38.54 | 39.46
21.7 | 33.6667 | 36.7 | 38.25 | 39.275
21.8 | 33.48 | 36.52 | 38.1222 | 39.12
21.9 | 33.275 | 36.32 | 37.94 | 38.9
22 | 33.1 | 36.12 | 37.74 | 38.64
22.1 | 32.94 | 35.925 | 37.5182 | 38.44
22.2 | 32.72 | 35.74 | 37.2778 | 38.3
22.3 | 32.58 | 35.54 | 37.1 | 38.2
22.4 | 32.42 | 35.34 | 36.9125 | 37.88
22.5 | 32.2571 | 35.14 | 36.82 | 37.7
22.6 | 32.1 | 34.94 | 36.62 | 37.6
22.7 | 31.9444 | 34.8 | 36.425 | 37.38
22.8 | 31.72 | 34.6556 | 36.22 | 37.18
22.9 | 31.4714 | 34.45 | 36.02 | 37.08
23 | 31.3 | 34.2714 | 35.9 | 36.88
23.1 | 31.225 | 34.1286 | 35.725 | 36.68
23.2 | 31.1 | 33.96 | 35.525 | 36.48
23.3 | 30.9429 | 33.76 | 35.32 | 36.3
23.4 | 30.74 | 33.5444 | 35.2 | 36.16
23.5 | 30.5429 | 33.46 | 34.9857 | 36
23.6 | 30.4 | 33.16 | 34.76 | 35.8
23.7 | 30.2714 | 33.06 | 34.6857 | 35.6
23.8 | 30.1 | 32.9 | 34.52 | 35.4
23.9 | 29.8857 | 32.76 | 34.29 | 35.3
24 | 29.7 | 32.58 | 34.18 | 35.1
24.1 | 29.55 | 32.39 | 33.98 | 34.9
24.2 | 29.4125 | 32.28 | 33.79 | 34.8
24.3 | 29.2889 | 32.08 | 33.68 | 34.6
24.4 | 29.12 | 31.88 | 33.48 | 34.5
24.5 | 28.9143 | 31.74 | 33.28 | 34.3
24.6 | 28.7143 | 31.55 | 33.09 | 34.16
24.7 | 28.7 | 31.39 | 32.98 | 33.95
24.8 | 28.5143 | 31.28 | 32.78 | 33.82
24.9 | 28.3 | 31.08 | 32.59 | 33.62
25 | 28.16 | 30.89 | 32.4909 | 33.42
25.1 | 28.0667 | 30.7778 | 32.4 | 33.32
25.2 | 27.85 | 30.5714 | 32.1 | 33.12
25.3 | 27.6 | 30.4375 | 32.02 | 32.92
25.4 | 27.5 | 30.26 | 31.9 | 32.82
25.5 | 27.4 | 30.16 | 31.7 | 32.66
25.6 | 27.28 | 29.96 | 31.52 | 32.5
25.7 | 27.1 | 29.8 | 31.4 | 32.3
25.8 | 27 | 29.66 | 31.22 | 32.15
25.9 | 26.875 | 29.46 | 31.02 | 32
26 | 26.6667 | 29.3 | 30.92 | 31.9
26.1 | 26.5 | 29.1429 | 30.72 | 31.78
26.2 | 26.4 | 29 | 30.6 | 31.6
26.3 | 26.28 | 28.8556 | 30.4 | 31.44
26.4 | 26.06 | 28.7375 | 30.28 | 31.3
26.5 | 25.9 | 28.6 | 30.175 | 31.125
26.6 | 25.82 | 28.46 | 30.04 | 31
26.7 | 25.7 | 28.3429 | 29.84 | 30.9
26.8 | 25.6 | 28.18 | 29.7 | 30.725
26.9 | 25.45 | 28 | 29.54 | 30.5333
27 | 25.3 | 27.88 | 29.4 | 30.4333
27.1 | 25.2 | 27.78 | 29.26 | 30.3
27.2 | 25 | 27.6 | 29.1125 | 30.1
27.3 | 24.8667 | 27.48 | 28.975 | 30
27.4 | 24.78 | 27.3 | 28.825 | 29.8
27.5 | 24.6 | 27.18 | 28.7167 | 29.66
27.6 | 24.5 | 27.1 | 28.52 | 29.5
27.7 | 24.34 | 26.9 | 28.4 | 29.4
27.8 | 24.24 | 26.8 | 28.22 | 29.26
27.9 | 24.1 | 26.6 | 28.12 | 29.1
28 | 24 | 26.5 | 28 | 28.92
28.1 | 23.8 | 26.4 | 27.82 | 28.88
28.2 | 23.75 | 26.275 | 27.7 | 28.68
28.3 | 23.56 | 26.075 | 27.52 | 28.58
28.4 | 23.45 | 26 | 27.4333 | 28.4
28.5 | 23.32 | 25.9 | 27.3 | 28.2286
28.6 | 23.2 | 25.7 | 27.1 | 28.14
28.7 | 23.06 | 25.5 | 27 | 27.94
28.8 | 22.96 | 25.4 | 26.88 | 27.9
28.9 | 22.82 | 25.3 | 26.78 | 27.7
29 | 22.72 | 25.2 | 26.6 | 27.6286
29.1 | 22.575 | 25 | 26.5 | 27.5
29.2 | 22.475 | 24.85 | 26.4 | 27.3625
29.3 | 22.325 | 24.76 | 26.2667 | 27.2
29.4 | 22.22 | 24.66 | 26.1 | 27.1
29.5 | 22.1 | 24.56 | 25.9667 | 27
29.6 | 21.98 | 24.46 | 25.9 | 26.8167
29.7 | 21.9 | 24.26 | 25.75 | 26.7
29.8 | 21.8 | 24.16 | 25.6 | 26.6
29.9 | 21.62 | 24 | 25.5 | 26.46
30 | 21.5 | 23.9 | 25.325 | 26.36

## SSD120 Wedge 60 Depth Doses
Percentage Depth Doses
Wedge 60
SSD=120cm
Detector: IBA Dosimetry CC13 Ionization chamber
Detector offset of -1.8mm (0.6*rcav) applied
Depth [cm] | 4x4cm2 | 10x10cm2 | 15x15cm2 | 15x40cm2
0 | 32.8 | 38.5 | 43.1 | 49.6
0.1 | 37 | 43.08 | 47.6 | 53.98
0.2 | 45.8 | 51.26 | 55.16 | 60.64
0.3 | 55.38 | 60 | 63.48 | 68.1667
0.4 | 63.6333 | 67.9833 | 71.04 | 74.76
0.5 | 71.0333 | 74 | 76.8667 | 80.18
0.6 | 76.6 | 79.4 | 81.7 | 84.48
0.7 | 81.34 | 83.7 | 85.3 | 88.08
0.8 | 85.32 | 87.2667 | 88.7 | 90.76
0.9 | 88.3 | 90.1 | 91.18 | 92.98
1 | 91.1 | 92.35 | 93.38 | 94.74
1.1 | 93.1 | 94.0667 | 95.14 | 96.1167
1.2 | 94.9 | 95.85 | 96.3 | 97.2
1.3 | 96.4 | 96.9 | 97.4667 | 98.1
1.4 | 97.4 | 97.875 | 98.1 | 98.7
1.5 | 98.4 | 98.7 | 98.7667 | 99.3
1.6 | 99 | 99.1667 | 99.35 | 99.6
1.7 | 99.45 | 99.6 | 99.7333 | 99.8
1.8 | 99.76 | 99.8 | 99.9 | 99.9833
1.9 | 99.9 | 99.95 | 100 | 100
2 | 100 | 100 | 100 | 100
2.1 | 100 | 100 | 100 | 99.9
2.2 | 99.96 | 99.9 | 99.9 | 99.82
2.3 | 99.86 | 99.82 | 99.7333 | 99.62
2.4 | 99.68 | 99.62 | 99.55 | 99.34
2.5 | 99.375 | 99.42 | 99.4 | 99.125
2.6 | 99.16 | 99.22 | 99.1 | 98.84
2.7 | 98.78 | 98.94 | 98.86 | 98.64
2.8 | 98.46 | 98.64 | 98.6 | 98.38
2.9 | 98.16 | 98.25 | 98.1667 | 97.98
3 | 97.7833 | 98.04 | 97.9 | 97.74
3.1 | 97.46 | 97.68 | 97.6 | 97.38
3.2 | 96.94 | 97.22 | 97.2 | 97.08
3.3 | 96.48 | 96.88 | 96.9 | 96.86
3.4 | 96.2 | 96.48 | 96.5 | 96.6
3.5 | 95.9667 | 96.24 | 96.2 | 96.22
3.6 | 95.5 | 95.94 | 95.8 | 95.8
3.7 | 95 | 95.58 | 95.5 | 95.6182
3.8 | 94.6 | 95.25 | 95.3167 | 95.15
3.9 | 94.1 | 94.82 | 94.94 | 94.83
4 | 93.8 | 94.46 | 94.64 | 94.4667
4.1 | 93.3 | 94.1 | 94.14 | 94.2
4.2 | 92.8 | 93.82 | 93.8 | 93.75
4.3 | 92.4333 | 93.48 | 93.44 | 93.4286
4.4 | 92.1 | 93.12 | 93.04 | 93.22
4.5 | 91.62 | 92.66 | 92.82 | 92.9545
4.6 | 91.22 | 92.32 | 92.44 | 92.66
4.7 | 90.82 | 92.0667 | 92.24 | 92.28
4.8 | 90.48 | 91.6 | 91.84 | 91.91
4.9 | 90.02 | 91.225 | 91.45 | 91.58
5 | 89.72 | 90.9 | 91.2 | 91.1875
5.1 | 89.26 | 90.55 | 90.86 | 90.9
5.2 | 88.78 | 90.16 | 90.62 | 90.5667
5.3 | 88.34 | 89.74 | 90.24 | 90.3333
5.4 | 87.9 | 89.52 | 89.92 | 89.94
5.5 | 87.6333 | 89.24 | 89.56 | 89.6667
5.6 | 87.175 | 88.95 | 89.3 | 89.34
5.7 | 86.74 | 88.48 | 89.06 | 88.98
5.8 | 86.45 | 88.2 | 88.66 | 88.74
5.9 | 85.8 | 87.9 | 88.28 | 88.3
6 | 85.4 | 87.5 | 87.78 | 87.98
6.1 | 85.2 | 87.1 | 87.5 | 87.76
6.2 | 84.85 | 86.8333 | 87.2 | 87.3667
6.3 | 84.3667 | 86.42 | 86.9 | 87
6.4 | 83.9667 | 86.15 | 86.6 | 86.85
6.5 | 83.625 | 85.74 | 86.3 | 86.44
6.6 | 83.2333 | 85.38 | 85.9 | 86.22
6.7 | 82.8333 | 85.08 | 85.6 | 85.92
6.8 | 82.45 | 84.72 | 85.3 | 85.52
6.9 | 82.1 | 84.34 | 85.1 | 85.22
7 | 81.8 | 84.08 | 84.7 | 84.95
7.1 | 81.2 | 83.64 | 84.5167 | 84.72
7.2 | 80.95 | 83.28 | 84.1 | 84.425
7.3 | 80.5667 | 82.94 | 83.75 | 84.0857
7.4 | 80.1 | 82.7 | 83.4 | 83.79
7.5 | 79.7 | 82.3 | 83.0333 | 83.49
7.6 | 79.3 | 81.94 | 82.7667 | 83.1375
7.7 | 78.9 | 81.72 | 82.38 | 82.74
7.8 | 78.6 | 81.35 | 82.1333 | 82.46
7.9 | 78.24 | 80.94 | 81.82 | 82.22
8 | 77.84 | 80.6667 | 81.58 | 81.8333
8.1 | 77.38 | 80.275 | 81.2667 | 81.5625
8.2 | 77.04 | 80.02 | 80.94 | 81.24
8.3 | 76.72 | 79.8333 | 80.65 | 81.04
8.4 | 76.3667 | 79.4333 | 80.38 | 80.76
8.5 | 75.88 | 79.0333 | 80.04 | 80.36
8.6 | 75.52 | 78.7333 | 79.72 | 79.98
8.7 | 75.24 | 78.44 | 79.34 | 79.78
8.8 | 74.88 | 78.1333 | 79.04 | 79.58
8.9 | 74.58 | 77.6667 | 78.76 | 79.24
9 | 74.08 | 77.48 | 78.45 | 78.88
9.1 | 73.68 | 77.14 | 78.2 | 78.65
9.2 | 73.3 | 76.66 | 77.82 | 78.38
9.3 | 73.1 | 76.4 | 77.52 | 78.1
9.4 | 72.8 | 76.14 | 77.22 | 77.75
9.5 | 72.4 | 75.75 | 76.96 | 77.5333
9.6 | 72 | 75.36 | 76.55 | 77.18
9.7 | 71.7 | 75.1 | 76.225 | 76.78
9.8 | 71.2 | 75 | 76.02 | 76.54
9.9 | 71 | 74.6 | 75.72 | 76.24
10 | 70.6 | 74.1 | 75.4 | 75.98
10.1 | 70.225 | 73.8 | 75.12 | 75.58
10.2 | 69.9 | 73.6 | 74.86 | 75.34
10.3 | 69.575 | 73.3 | 74.55 | 75.1
10.4 | 69.3 | 72.9 | 74.22 | 74.85
10.5 | 68.95 | 72.6 | 73.9714 | 74.62
10.6 | 68.6 | 72.2857 | 73.7 | 74.3
10.7 | 68.3333 | 72.06 | 73.36 | 73.96
10.8 | 67.9 | 71.7429 | 73.07 | 73.66
10.9 | 67.5333 | 71.4333 | 72.78 | 73.42
11 | 67.2333 | 71.1429 | 72.56 | 73.16
11.1 | 66.8667 | 70.8125 | 72.17 | 72.86
11.2 | 66.6 | 70.42 | 71.87 | 72.52
11.3 | 66.2333 | 70.125 | 71.57 | 72.26
11.4 | 65.92 | 69.92 | 71.28 | 71.975
11.5 | 65.64 | 69.54 | 71.0667 | 71.7333
11.6 | 65.2333 | 69.23 | 70.7429 | 71.36
11.7 | 64.98 | 69.02 | 70.5 | 71.16
11.8 | 64.6333 | 68.73 | 70.2 | 70.96
11.9 | 64.28 | 68.48 | 69.875 | 70.58
12 | 64.06 | 68.18 | 69.6429 | 70.4143
12.1 | 63.76 | 67.88 | 69.4167 | 70.2333
12.2 | 63.36 | 67.58 | 69.0667 | 69.8143
12.3 | 62.98 | 67.2 | 68.8091 | 69.58
12.4 | 62.8 | 66.95 | 68.58 | 69.275
12.5 | 62.46 | 66.74 | 68.32 | 69.04
12.6 | 62.16 | 66.38 | 68.08 | 68.78
12.7 | 61.76 | 66.08 | 67.8 | 68.4625
12.8 | 61.4667 | 65.9111 | 67.4429 | 68.25
12.9 | 61.2 | 65.54 | 67.22 | 67.9
13 | 60.9 | 65.28 | 67.02 | 67.6429
13.1 | 60.6 | 64.91 | 66.7571 | 67.4
13.2 | 60.2 | 64.65 | 66.38 | 67.15
13.3 | 60 | 64.46 | 66.1 | 66.9
13.4 | 59.7 | 64.2 | 65.7429 | 66.6111
13.5 | 59.5 | 63.8125 | 65.43 | 66.3125
13.6 | 59.2167 | 63.56 | 65.24 | 66.1429
13.7 | 58.84 | 63.24 | 65.04 | 65.9
13.8 | 58.44 | 62.96 | 64.8 | 65.7375
13.9 | 58.22 | 62.66 | 64.38 | 65.4
14 | 57.93 | 62.48 | 64.1 | 65.1
14.1 | 57.63 | 62.1429 | 63.9 | 64.85
14.2 | 57.33 | 61.82 | 63.7 | 64.5857
14.3 | 57.03 | 61.625 | 63.375 | 64.2
14.4 | 56.82 | 61.28 | 63.12 | 64.02
14.5 | 56.44 | 61 | 62.8333 | 63.74
14.6 | 56.2364 | 60.74 | 62.66 | 63.52
14.7 | 55.9333 | 60.54 | 62.42 | 63.36
14.8 | 55.6545 | 60.28 | 62.2 | 63.06
14.9 | 55.44 | 59.9833 | 61.88 | 62.7286
15 | 55.2 | 59.76 | 61.6 | 62.52
15.1 | 54.9571 | 59.4 | 61.4 | 62.32
15.2 | 54.61 | 59.2 | 61.1 | 62.0333
15.3 | 54.31 | 59.02 | 60.9 | 61.825
15.4 | 54.04 | 58.74 | 60.6 | 61.64
15.5 | 53.78 | 58.45 | 60.3 | 61.34
15.6 | 53.46 | 58.1 | 60.1 | 61.08
15.7 | 53.3 | 57.875 | 59.8 | 60.84
15.8 | 53 | 57.65 | 59.6167 | 60.58
15.9 | 52.8286 | 57.4 | 59.32 | 60.34
16 | 52.5429 | 57.1 | 59.12 | 60.08
16.1 | 52.2 | 56.9 | 58.92 | 59.84
16.2 | 52 | 56.6 | 58.6 | 59.65
16.3 | 51.7 | 56.325 | 58.42 | 59.34
16.4 | 51.4375 | 56.0667 | 58.12 | 59.2
16.5 | 51.22 | 55.88 | 57.92 | 58.96
16.6 | 50.875 | 55.66 | 57.74 | 58.72
16.7 | 50.72 | 55.38 | 57.52 | 58.42
16.8 | 50.52 | 55.1667 | 57.14 | 58.26
16.9 | 50.22 | 54.94 | 57.02 | 58.05
17 | 50.02 | 54.68 | 56.74 | 57.7667
17.1 | 49.72 | 54.44 | 56.42 | 57.6667
17.2 | 49.52 | 54.225 | 56.225 | 57.2667
17.3 | 49.32 | 53.9667 | 56.025 | 57
17.4 | 49.02 | 53.7 | 55.725 | 56.8667
17.5 | 48.74 | 53.4 | 55.5333 | 56.58
17.6 | 48.42 | 53.18 | 55.36 | 56.425
17.7 | 48.12 | 52.98 | 55.08 | 56.15
17.8 | 48.1 | 52.74 | 54.78 | 55.9
17.9 | 47.84 | 52.54 | 54.6 | 55.65
18 | 47.52 | 52.26 | 54.5 | 55.5
18.1 | 47.32 | 52.06 | 54.2 | 55.3333
18.2 | 47 | 51.8 | 53.9 | 55.1
18.3 | 46.9 | 51.56 | 53.8 | 54.8
18.4 | 46.62 | 51.36 | 53.6 | 54.55
18.5 | 46.44 | 51.12 | 53.3 | 54.35
18.6 | 46.15 | 50.86 | 53.1 | 54.15
18.7 | 45.94 | 50.66 | 52.8 | 53.9
18.8 | 45.68 | 50.42 | 52.6 | 53.64
18.9 | 45.48 | 50.175 | 52.4 | 53.4667
19 | 45.24 | 50.0333 | 52.1 | 53.24
19.1 | 44.94 | 49.76 | 52 | 53.0667
19.2 | 44.84 | 49.6 | 51.8167 | 52.8
19.3 | 44.64 | 49.375 | 51.525 | 52.64
19.4 | 44.38 | 49.175 | 51.24 | 52.44
19.5 | 44.14 | 48.96 | 51.075 | 52.25
19.6 | 43.95 | 48.68 | 50.8 | 52.05
19.7 | 43.75 | 48.475 | 50.6 | 51.8
19.8 | 43.55 | 48.24 | 50.4 | 51.45
19.9 | 43.3 | 47.98 | 50.2 | 51.36
20 | 43.1 | 47.74 | 50 | 51.12
20.1 | 42.9 | 47.64 | 49.8 | 50.9
20.2 | 42.6 | 47.425 | 49.6 | 50.6333
20.3 | 42.4167 | 47.14 | 49.4 | 50.52
20.4 | 42.22 | 46.94 | 49.15 | 50.375
20.5 | 42.02 | 46.74 | 49 | 50.22
20.6 | 41.82 | 46.54 | 48.75 | 49.9
20.7 | 41.62 | 46.34 | 48.6 | 49.72
20.8 | 41.42 | 46.18 | 48.3 | 49.575
20.9 | 41.2 | 45.9 | 48.16 | 49.375
21 | 41 | 45.74 | 47.9 | 49.175
21.1 | 40.8 | 45.54 | 47.76 | 48.875
21.2 | 40.65 | 45.3 | 47.56 | 48.72
21.3 | 40.4 | 45.16 | 47.35 | 48.56
21.4 | 40.2333 | 45 | 47.2 | 48.375
21.5 | 40 | 44.7429 | 47.02 | 48.2
21.6 | 39.85 | 44.45 | 46.72 | 47.975
21.7 | 39.6 | 44.2 | 46.52 | 47.7
21.8 | 39.4667 | 44.1 | 46.32 | 47.5
21.9 | 39.3 | 43.9 | 46.12 | 47.4
22 | 39.12 | 43.7 | 45.92 | 47.2
22.1 | 38.825 | 43.5 | 45.62 | 46.94
22.2 | 38.725 | 43.3 | 45.52 | 46.8
22.3 | 38.5 | 43.2091 | 45.42 | 46.58
22.4 | 38.35 | 42.93 | 45.12 | 46.44
22.5 | 38.1 | 42.72 | 45.02 | 46.24
22.6 | 38 | 42.61 | 44.72 | 46.04
22.7 | 37.68 | 42.33 | 44.52 | 45.84
22.8 | 37.6 | 42.15 | 44.4 | 45.64
22.9 | 37.44 | 41.92 | 44.22 | 45.44
23 | 37.24 | 41.76 | 44.08 | 45.24
23.1 | 37 | 41.62 | 43.82 | 45.15
23.2 | 36.9 | 41.42 | 43.68 | 44.94
23.3 | 36.7 | 41.26 | 43.46 | 44.72
23.4 | 36.5 | 41.12 | 43.32 | 44.5
23.5 | 36.3 | 40.88 | 43.15 | 44.36
23.6 | 36.1 | 40.6636 | 42.9857 | 44.16
23.7 | 36 | 40.54 | 42.8333 | 43.96
23.8 | 35.9 | 40.325 | 42.54 | 43.76
23.9 | 35.6 | 40.1 | 42.45 | 43.6
24 | 35.5 | 39.9286 | 42.22 | 43.45
24.1 | 35.35 | 39.75 | 41.96 | 43.25
24.2 | 35.075 | 39.56 | 41.8 | 43.05
24.3 | 34.92 | 39.3444 | 41.66 | 42.8667
24.4 | 34.72 | 39.2 | 41.46 | 42.7
24.5 | 34.62 | 39.1 | 41.3 | 42.6
24.6 | 34.475 | 38.92 | 41.1714 | 42.4
24.7 | 34.26 | 38.72 | 40.92 | 42.2
24.8 | 34.06 | 38.6 | 40.72 | 42.0667
24.9 | 33.875 | 38.42 | 40.62 | 41.9
25 | 33.7333 | 38.12 | 40.4143 | 41.7
25.1 | 33.6 | 38 | 40.18 | 41.5
25.2 | 33.42 | 37.9 | 40 | 41.3667
25.3 | 33.22 | 37.6857 | 39.88 | 41.1
25.4 | 33.1 | 37.475 | 39.68 | 41
25.5 | 33 | 37.3429 | 39.54 | 40.9
25.6 | 32.82 | 37.16 | 39.4 | 40.7
25.7 | 32.62 | 36.98 | 39.24 | 40.6
25.8 | 32.42 | 36.86 | 39 | 40.4
25.9 | 32.3 | 36.73 | 38.9 | 40.1625
26 | 32.12 | 36.56 | 38.7 | 40.0286
26.1 | 32 | 36.36 | 38.55 | 39.8
26.2 | 31.82 | 36.175 | 38.425 | 39.6286
26.3 | 31.78 | 35.9857 | 38.2333 | 39.48
26.4 | 31.56 | 35.83 | 38.1 | 39.35
26.5 | 31.4 | 35.6889 | 37.8333 | 39.1429
26.6 | 31.22 | 35.56 | 37.7333 | 39.0125
26.7 | 31.1 | 35.375 | 37.6 | 38.82
26.8 | 30.92 | 35.2429 | 37.42 | 38.6667
26.9 | 30.725 | 35.08 | 37.3 | 38.5
27 | 30.7 | 34.9143 | 37.1333 | 38.38
27.1 | 30.52 | 34.7333 | 36.9 | 38.28
27.2 | 30.32 | 34.6 | 36.8 | 38.08
27.3 | 30.1667 | 34.44 | 36.65 | 37.9
27.4 | 30.1 | 34.28 | 36.5 | 37.78
27.5 | 29.925 | 34.1 | 36.35 | 37.5667
27.6 | 29.8 | 34 | 36.2 | 37.4
27.7 | 29.64 | 33.84 | 36.0667 | 37.225
27.8 | 29.5 | 33.64 | 35.8667 | 37.15
27.9 | 29.4 | 33.54 | 35.7 | 37
28 | 29.2 | 33.34 | 35.6 | 36.8
28.1 | 29.0833 | 33.24 | 35.4 | 36.68
28.2 | 29 | 33.14 | 35.3 | 36.5
28.3 | 28.75 | 32.94 | 35.14 | 36.4
28.4 | 28.7 | 32.75 | 34.98 | 36.3
28.5 | 28.54 | 32.64 | 34.8 | 36.15
28.6 | 28.4 | 32.44 | 34.7 | 36
28.7 | 28.2 | 32.35 | 34.55 | 35.7667
28.8 | 28.06 | 32.2 | 34.35 | 35.6667
28.9 | 27.96 | 32.06 | 34.25 | 35.5
29 | 27.8 | 31.9 | 34.1 | 35.4
29.1 | 27.7 | 31.76 | 33.96 | 35.2333
29.2 | 27.6 | 31.66 | 33.8 | 35.12
29.3 | 27.5 | 31.5 | 33.7 | 34.9
29.4 | 27.36 | 31.36 | 33.5 | 34.8
29.5 | 27.2 | 31.16 | 33.36 | 34.66
29.6 | 27.06 | 31.06 | 33.2 | 34.5333
29.7 | 26.96 | 30.9 | 33.1 | 34.4333
29.8 | 26.86 | 30.82 | 32.9667 | 34.2
29.9 | 26.7 | 30.7 | 32.82 | 34.06
30 | 26.5 | 30.6 | 32.7 | 33.9667

## Wedge 60 Lateral Profiles
Lateral Profiles
Wedge 60
SSD=100cm
Scanning direction: Crossline
Detector: IBA Dosimetry CC13 Ionization chamber
| Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%] |  |  |  | Dose[%]
Off axis position [cm] | Depth:2 cm FieldSize:4x4 cm2 | Depth:2 cm FieldSize:10x10 cm2 | Depth:2 cm FieldSize:15x15 cm2 | Depth:2 cm FieldSize:40x15 cm2 | Depth:5 cm FieldSize:4x4 cm2 | Depth:5 cm FieldSize:10x10 cm2 | Depth:5 cm FieldSize:15x15 cm2 | Depth:5 cm FieldSize:40x15 cm2 | Depth:10 cm FieldSize:4x4 cm2 | Depth:10 cm FieldSize:10x10 cm2 | Depth:10 cm FieldSize:15x15 cm2 | Depth:10 cm FieldSize:40x15 cm2 | Depth:20 cm FieldSize:4x4 cm2 | Depth:20 cm FieldSize:10x10 cm2 | Depth:20 cm FieldSize:15x15 cm2 | Depth:20 cm FieldSize:40x15 cm2 | Depth:30 cm FieldSize:4x4 cm2 | Depth:30 cm FieldSize:10x10 cm2 | Depth:30 cm FieldSize:15x15 cm2 | Depth:30 cm FieldSize:40x15 cm2
-14.7 | [18 empty] | 12.8333 | 18.18
-14.6 | [18 empty] | 13.02 | 18.4
-14.5 | [18 empty] | 13.2571 | 18.625
-14.4 | [18 empty] | 13.5 | 18.775
-14.3 | [18 empty] | 13.6571 | 19.02
-14.2 | [18 empty] | 13.9 | 19.25
-14.1 | [18 empty] | 14.05 | 19.5
-14 | [18 empty] | 14.4 | 19.7
-13.9 | [14 empty] | 10.2449 | 14.905 |  |  | 14.6375 | 20
-13.8 | [14 empty] | 10.4348 | 15.105 |  |  | 14.8 | 20.18
-13.7 | [14 empty] | 10.628 | 15.2251 |  |  | 14.95 | 20.5
-13.6 | [14 empty] | 10.7946 | 15.5252 |  |  | 15.3 | 20.68
-13.5 | [14 empty] | 10.9545 | 15.6552 |  |  | 15.55 | 20.9
-13.4 | [14 empty] | 11.1944 | 15.9303 |  |  | 15.825 | 21.16
-13.3 | [14 empty] | 11.2944 | 16.1054 |  |  | 16.075 | 21.625
-13.2 | [11 empty] | 11.1968 |  |  | 11.5514 | 16.3387 |  |  | 16.275 | 21.8
-13.1 | [10 empty] | 7.3833 | 11.3301 |  |  | 11.7442 | 16.5305 |  |  | 16.55 | 22.075
-13 | [10 empty] | 7.4 | 11.4967 |  |  | 11.994 | 16.8556 |  |  | 16.8333 | 22.325
-12.9 | [10 empty] | 7.68 | 11.5967 |  |  | 12.1939 | 17.1057 |  |  | 17.22 | 22.65
-12.8 | [6 empty] | 5.9 | 9.6048 |  |  | 7.8 | 11.7966 |  |  | 12.3738 | 17.2457 |  |  | 17.45 | 23.04
-12.7 | [6 empty] | 6.05 | 9.7549 |  |  | 7.9429 | 11.9299 |  |  | 12.6687 | 17.4915 |  |  | 17.82 | 23.35
-12.6 |  |  | 5.8 | 9.5 |  |  | 6.1 | 9.8649 |  |  | 8.025 | 12.0965 |  |  | 12.8936 | 17.7059 |  |  | 18.2 | 23.6667
-12.5 |  |  | 5.84 | 9.6 |  |  | 6.2667 | 9.9717 |  |  | 8.12 | 12.2965 |  |  | 13.1135 | 18.056 |  |  | 18.52 | 23.9
-12.4 |  |  | 5.94 | 9.7 |  |  | 6.3143 | 10.1051 |  |  | 8.35 | 12.3965 |  |  | 13.4183 | 18.2661 |  |  | 18.7 | 24.3
-12.3 |  |  | 6 | 9.825 |  |  | 6.4 | 10.2051 |  |  | 8.475 | 12.5364 |  |  | 13.5932 | 18.5061 |  |  | 19.04 | 24.6
-12.2 |  |  | 6.1 | 9.975 |  |  | 6.5 | 10.3552 |  |  | 8.6571 | 12.7821 |  |  | 13.9331 | 18.8062 |  |  | 19.4 | 24.95
-12.1 |  |  | 6.2 | 10.0333 |  |  | 6.66 | 10.4552 |  |  | 8.8 | 12.9363 |  |  | 14.1929 | 19.1286 |  |  | 19.82 | 25.3
-12 |  |  | 6.3 | 10.2 |  |  | 6.7571 | 10.6053 |  |  | 8.975 | 13.1296 |  |  | 14.4928 | 19.3636 |  |  | 20.1 | 25.6714
-11.9 |  |  | 6.4 | 10.275 |  |  | 6.86 | 10.7387 |  |  | 9.2 | 13.3962 |  |  | 14.7426 | 19.7495 |  |  | 20.48 | 26.05
-11.8 |  |  | 6.4714 | 10.3833 |  |  | 6.9571 | 10.8555 |  |  | 9.425 | 13.5711 |  |  | 15.0725 | 20.0867 |  |  | 21 | 26.4714
-11.7 |  |  | 6.6 | 10.5 |  |  | 7.1 | 11.0255 |  |  | 9.575 | 13.7161 |  |  | 15.3352 | 20.3568 |  |  | 21.3 | 26.9
-11.6 |  |  | 6.7 | 10.6 |  |  | 7.225 | 11.1056 |  |  | 9.8333 | 13.946 |  |  | 15.6922 | 20.6569 |  |  | 21.7 | 27.3
-11.5 |  |  | 6.8 | 10.78 |  |  | 7.4 | 11.3557 |  |  | 10.025 | 14.1959 |  |  | 16.042 | 21.007 |  |  | 22.08 | 27.7
-11.4 |  |  | 6.9 | 10.9 |  |  | 7.6 | 11.5058 |  |  | 10.175 | 14.4459 |  |  | 16.3419 | 21.3571 |  | 8.4833 | 22.7 | 28.2
-11.3 |  |  | 7 | 11.08 |  |  | 7.7 | 11.7059 |  |  | 10.42 | 14.6958 |  |  | 16.7416 | 21.8273 |  | 8.625 | 23.1571 | 28.68
-11.2 |  |  | 7.16 | 11.2 |  |  | 7.88 | 11.8059 |  |  | 10.6 | 14.9207 |  |  | 17.0665 | 22.074 |  | 8.8 | 23.6 | 29.2
-11.1 |  |  | 7.3 | 11.3 |  |  | 8.025 | 11.956 |  |  | 10.8875 | 15.179 |  |  | 17.3913 | 22.4575 |  | 8.9 | 24.225 | 29.8
-11 |  |  | 7.45 | 11.45 |  |  | 8.2 | 12.2061 |  |  | 11.15 | 15.4156 |  | 6.4 | 17.834 | 22.8076 |  | 9 | 24.9 | 30.32
-10.9 |  |  | 7.6 | 11.6 |  |  | 8.375 | 12.3812 |  |  | 11.4167 | 15.633 |  | 6.5 | 18.2709 | 23.2877 |  | 9.2 | 25.5333 | 31.025
-10.8 |  |  | 7.75 | 11.8 |  |  | 8.525 | 12.5563 |  |  | 11.6667 | 15.9704 |  | 6.7 | 18.7307 | 23.7479 |  | 9.4 | 26.4 | 31.8
-10.7 |  |  | 7.8571 | 11.9571 |  |  | 8.7333 | 12.7314 |  |  | 11.98 | 16.2204 |  | 6.7 | 19.1705 | 24.188 |  | 9.525 | 27.4 | 32.76
-10.6 |  |  | 7.95 | 12.1 |  |  | 8.92 | 12.9815 |  |  | 12.25 | 16.4953 |  | 6.8 | 19.6188 | 24.6582 |  | 9.675 | 28.64 | 33.875
-10.5 |  |  | 8.125 | 12.2667 |  |  | 9.1167 | 13.1399 |  |  | 12.5 | 16.8202 |  | 7 | 20.11 | 25.0833 |  | 9.825 | 30.0667 | 35.275
-10.4 |  |  | 8.3 | 12.4 |  |  | 9.3333 | 13.4317 |  | 4.5 | 12.8667 | 17.1951 |  | 7.1 | 20.7396 | 25.6086 |  | 10 | 32.2 | 37.2
-10.3 |  |  | 8.4833 | 12.65 |  |  | 9.525 | 13.6068 |  | 4.6 | 13.25 | 17.47 |  | 7.225 | 21.2894 | 26.2488 |  | 10.125 | 34.95 | 39.55
-10.2 |  |  | 8.65 | 12.8 |  | 3.6 | 9.775 | 13.8069 |  | 4.7 | 13.6 | 17.9449 |  | 7.4 | 22.014 | 26.8804 |  | 10.3 | 38.6667 | 42.92
-10.1 |  |  | 8.86 | 13 |  | 3.7 | 10.02 | 14.1071 |  | 4.7 | 13.9333 | 18.2698 |  | 7.5 | 22.7636 | 27.6892 |  | 10.55 | 44.18 | 47.85
-10 |  | 3.5307 | 9.05 | 13.1 |  | 3.7 | 10.275 | 14.4072 |  | 4.8 | 14.35 | 18.5947 |  | 7.6 | 23.6216 | 28.5345 |  | 10.7 | 51.6333 | 54.925
-9.9 |  | 3.5973 | 9.2 | 13.3 |  | 3.8 | 10.6 | 14.6473 |  | 4.9 | 14.7667 | 19.1945 |  | 7.8 | 24.7376 | 29.5098 |  | 10.925 | 61.84 | 63.8
-9.8 |  | 3.6639 | 9.425 | 13.6 |  | 3.8 | 10.82 | 14.9325 |  | 5 | 15.3 | 19.5277 |  | 7.9 | 26.1869 | 30.9603 |  | 11.18 | 73.0333 | 74.6
-9.7 |  | 3.6972 | 9.66 | 13.85 |  | 3.9 | 11.1 | 15.2476 |  | 5.1 | 15.75 | 20.0657 |  | 8.1 | 28.111 | 32.5909 |  | 11.32 | 86.35 | 86.4
-9.6 |  | 3.6972 | 9.875 | 14.08 |  | 3.9 | 11.44 | 15.5703 |  | 5.1333 | 16.3 | 20.5608 |  | 8.3 | 30.4048 | 34.8116 |  | 11.525 | 99.6333 | 97.9
-9.5 |  | 3.6972 | 10.1333 | 14.3 |  | 4 | 11.8 | 16.008 |  | 5.3 | 16.9333 | 21.1439 |  | 8.4 | 33.8831 | 38.146 |  | 11.76 | 110.9 | 108.76
-9.4 |  | 3.7971 | 10.375 | 14.5286 |  | 4 | 12.2 | 16.3368 |  | 5.4 | 17.6 | 21.8937 |  | 8.6 | 38.9405 | 42.8476 |  | 11.9571 | 120.775 | 117.6
-9.3 |  | 3.8638 | 10.6 | 14.8 |  | 4.1 | 12.7 | 16.7584 |  | 5.5 | 18.4 | 22.6935 |  | 8.7667 | 46.2569 | 49.5165 |  | 12.3 | 128.48 | 124.8
-9.2 |  | 3.8971 | 10.92 | 15.15 |  | 4.2 | 13.1 | 17.2753 |  | 5.55 | 19.5 | 23.5933 |  | 9 | 56.1319 | 59.034 |  | 12.525 | 133.75 | 129.7
-9.1 |  | 3.997 | 11.2333 | 15.475 |  | 4.2 | 13.7 | 17.7839 |  | 5.7 | 20.7 | 24.7596 |  | 9.1667 | 69.0655 | 70.0233 |  | 12.76 | 137.32 | 132.9
-9 |  | 3.997 | 11.6333 | 15.8667 |  | 4.3 | 14.3 | 18.4092 |  | 5.8 | 22.3 | 26.3925 |  | 9.4 | 82.9985 | 83.2706 |  | 13.025 | 139.3 | 135
-8.9 |  | 4.0969 | 12.02 | 16.25 |  | 4.4 | 15.15 | 19.2297 |  | 5.9 | 24.45 | 28.3752 |  | 9.6 | 97.7512 | 96.8323 |  | 13.26 | 140.5 | 136.3
-8.8 |  | 4.1969 | 12.5 | 16.825 |  | 4.5 | 16.1 | 20.2101 |  | 6.1 | 27.6 | 31.1111 |  | 9.85 | 111.5438 | 109.6033 |  | 13.5 | 141.5 | 137
-8.7 |  | 4.1969 | 13.18 | 17.46 |  | 4.6 | 17.4 | 21.3607 |  | 6.2 | 32.2 | 35.6898 |  | 10.05 | 123.3784 | 121.44 |  | 13.85 | 141.975 | 137.5
-8.6 |  | 4.2968 | 14 | 18.125 |  | 4.6 | 19.2 | 23.1116 |  | 6.3 | 39.15 | 42.1379 |  | 10.3 | 132.8732 | 129.6034 |  | 14.2 | 142.125 | 137.6
-8.5 |  | 4.3967 | 15.04 | 19.2 |  | 4.8 | 21.7167 | 25.3927 |  | 6.5 | 49.5 | 52.1601 |  | 10.5 | 139.23 | 135.765 |  | 14.45 | 142.2 | 137.7
-8.4 |  | 4.3967 | 16.55 | 20.48 |  | 4.8 | 25.35 | 28.8344 |  | 6.6 | 63.86 | 64.7015 |  | 10.7333 | 143.1287 | 139.7264 |  | 14.7 | 142.075 | 137.6
-8.3 |  | 4.4966 | 18.5 | 22.5 |  | 4.95 | 30.9 | 33.957 |  | 6.8 | 79.5 | 79.8058 |  | 11 | 145.6021 | 141.887 |  | 15.075 | 141.9 | 137.48
-8.2 |  | 4.5716 | 21.55 | 25.4 |  | 5.1 | 39.65 | 42.2926 |  | 7 | 97.3 | 96.3924 |  | 11.3333 | 147.1062 | 143.208 |  | 15.45 | 141.8 | 137.25
-8.1 |  | 4.6965 | 26.675 | 30.075 |  | 5.2 | 52.95 | 54.2271 |  | 7.1 | 114.2333 | 112.1679 |  | 11.575 | 147.8546 | 143.988 |  | 15.8429 | 141.55 | 137.0333
-8 |  | 4.7464 | 34.9 | 38 |  | 5.3 | 69.2 | 69.8849 |  | 7.4 | 128.4 | 126.0838 |  | 11.85 | 148.126 | 144.248 |  | 16.2333 | 141.175 | 136.875
-7.9 |  | 4.8963 | 47.84 | 49.8 |  | 5.4333 | 88.0429 | 86.8434 |  | 7.6 | 140.6 | 137.5105 |  | 12.2 | 148.226 | 144.248 |  | 16.5714 | 140.925 | 136.58
-7.8 |  | 4.9629 | 65.55 | 66.22 |  | 5.6 | 107.18 | 105.678 |  | 7.8 | 148.7 | 145.0782 |  | 12.55 | 148.226 | 144.248 |  | 17.16 | 140.275 | 136.25
-7.7 |  | 5.0962 | 84.96 | 85.9333 |  | 5.7 | 124.2333 | 122.4617 |  | 8 | 153.36 | 149.4573 |  | 12.8667 | 148.101 | 144.148 |  | 17.8 | 139.95 | 135.9333
-7.6 |  | 5.2211 | 105.56 | 105.44 |  | 5.9 | 140.58 | 136.9685 |  | 8.2 | 155.8333 | 152.0903 |  | 13.25 | 147.826 | 143.908 | 3.1 | 18.4 | 139.62 | 135.575
-7.5 |  | 5.296 | 125.64 | 123.5 |  | 6.05 | 151.0667 | 148.574 |  | 8.475 | 157.36 | 153.306 |  | 13.6667 | 147.306 | 143.598 | 3.2 | 19.16 | 139.125 | 135.325
-7.4 |  | 5.4209 | 142.9 | 139.9167 |  | 6.2 | 158.55 | 155.0028 |  | 8.8 | 157.8 | 153.7893 | 2.2011 | 14.1 | 146.8124 | 143.148 | 3.2 | 20.125 | 138.65 | 135.05
-7.3 |  | 5.5958 | 155.25 | 151.75 |  | 6.4 | 162.1875 | 158.4123 |  | 9.1 | 157.8 | 153.856 | 2.4012 | 14.65 | 146.307 | 142.7337 | 3.2 | 21.34 | 138.18 | 134.55
-7.2 |  | 5.6957 | 163.3 | 159.425 |  | 6.6 | 163.95 | 160.3133 | 1.6 | 9.3667 | 157.8 | 153.856 | 2.4012 | 15.25 | 145.7413 | 142.107 | 3.3 | 22.85 | 137.8 | 134.05
-7.1 |  | 5.8956 | 167.34 | 163.125 | 1.1994 | 6.8 | 164.5667 | 160.88 | 1.6 | 9.7 | 157.55 | 153.656 | 2.4012 | 15.9 | 145.227 | 141.747 | 3.3 | 25 | 137.25 | 133.8
-7 | 1.2 | 6.0455 | 169.2 | 164.8 | 1.2994 | 7.1 | 164.6 | 160.88 | 1.6 | 10.05 | 157.025 | 153.1893 | 2.5013 | 16.625 | 144.588 | 141.187 | 3.4 | 27.85 | 137 | 133.4
-6.9 | 1.2 | 6.2453 | 169.3 | 165 | 1.2994 | 7.3 | 164.52 | 160.855 | 1.6 | 10.5 | 156.35 | 152.656 | 2.5013 | 17.6 | 144.128 | 140.647 | 3.5 | 32 | 136.35 | 133.04
-6.8 | 1.2 | 6.3952 | 169.3 | 165 | 1.2994 | 7.6 | 164.12 | 160.4514 | 1.7 | 10.95 | 155.625 | 151.857 | 2.6013 | 18.75 | 143.508 | 140.187 | 3.6 | 37.5 | 135.8 | 132.4
-6.7 | 1.2 | 6.645 | 168.9 | 164.75 | 1.2994 | 7.95 | 163.4 | 159.7657 | 1.7 | 11.45 | 154.8 | 151.097 | 2.6013 | 20.4 | 142.8787 | 139.6718 | 3.6 | 44.95 | 135.15 | 131.86
-6.6 | 1.2 | 6.7949 | 168.1 | 164.0143 | 1.2994 | 8.4 | 162.6 | 159.08 | 1.7 | 12.12 | 154.125 | 150.4903 | 2.6763 | 22.5333 | 142.1623 | 139.046 | 3.7 | 54.78 | 134.75 | 131.425
-6.5 | 1.3 | 7.0947 | 167.18 | 163.1 | 1.2994 | 8.7 | 161.5 | 158.079 | 1.7 | 12.9 | 153.225 | 149.657 | 2.7014 | 25.32 | 141.454 | 138.346 | 3.8 | 65.94 | 134.05 | 130.925
-6.4 | 1.3 | 7.4944 | 166.1 | 161.9625 | 1.3993 | 9.2 | 160.7 | 157.179 | 1.8 | 13.725 | 152.44 | 148.8974 | 2.8014 | 29.4333 | 140.8798 | 137.846 | 3.8 | 77.5 | 133.6 | 130.525
-6.3 | 1.3 | 7.7942 | 164.64 | 161 | 1.3993 | 9.8667 | 159.58 | 156.278 | 1.8 | 14.9 | 151.5 | 148.233 | 2.8014 | 35.7 | 140.03 | 137.2835 | 3.9 | 88.6 | 132.975 | 130.1
-6.2 | 1.3 | 8.1939 | 163.4 | 159.925 | 1.3993 | 10.6 | 158.6 | 155.1275 | 1.8 | 16.52 | 150.6 | 147.358 | 2.9015 | 44.1 | 139.355 | 136.6054 | 3.9 | 98.1 | 132.6 | 129.55
-6.1 | 1.3 | 8.7934 | 162.2 | 158.8 | 1.3993 | 11.6 | 157.2 | 154.177 | 1.9 | 18.7 | 149.8 | 146.258 | 3.0015 | 55.6 | 138.731 | 135.945 | 4 | 105.925 | 132.06 | 129.0111
-6 | 1.3 | 9.4929 | 161 | 157.4571 | 1.3993 | 12.9 | 156.0667 | 153.0513 | 1.9 | 21.6333 | 148.8 | 145.4585 | 3.0015 | 67 | 138.231 | 135.265 | 4.1 | 112.1 | 131.6 | 128.55
-5.9 | 1.3 | 10.4921 | 159.66 | 156.1 | 1.3993 | 14.6333 | 155.2 | 151.776 | 1.9 | 26.2 | 147.7875 | 144.609 | 3.1016 | 80.3333 | 137.4977 | 134.645 | 4.2 | 115.9 | 130.86 | 128.1143
-5.8 | 1.3 | 11.7912 | 157.95 | 154.75 | 1.4992 | 17.2333 | 153.7143 | 150.775 | 2 | 33.4 | 146.76 | 143.759 | 3.2016 | 91.5 | 136.7606 | 133.845 | 4.3 | 118.4 | 130.4167 | 127.7
-5.7 | 1.3 | 13.7897 | 156.74 | 153.525 | 1.4992 | 21.3 | 152.6 | 149.775 | 2 | 43.35 | 145.8 | 142.959 | 3.3016 | 102.1 | 135.932 | 133.344 | 4.4 | 120.05 | 129.7 | 127.1375
-5.6 | 1.4 | 16.7874 | 155.48 | 152.2333 | 1.4992 | 27 | 151.5286 | 148.624 | 2.1 | 55.45 | 144.7 | 141.9342 | 3.3016 | 110.2 | 135.1211 | 132.644 | 4.5 | 120.9 | 129.3667 | 126.58
-5.5 | 1.4 | 21.6837 | 154.2 | 151.2 | 1.5492 | 36.6333 | 150.2 | 147.474 | 2.1 | 69.45 | 143.7571 | 141.24 | 3.4017 | 116.2 | 134.613 | 131.9773 | 4.6 | 121.475 | 128.9333 | 126.075
-5.4 | 1.4 | 29.6277 | 152.8 | 149.85 | 1.5992 | 48.7 | 149.36 | 146.473 | 2.2 | 84.25 | 143.02 | 140.3314 | 3.5017 | 119.6 | 133.813 | 131.394 | 4.7 | 121.62 | 128.025 | 125.6
-5.3 | 1.4 | 40.9193 | 151.4222 | 148.6 | 1.5992 | 62.5 | 147.925 | 145.073 | 2.25 | 97.25 | 142.2 | 139.235 | 3.6018 | 121.7667 | 133.193 | 130.844 | 4.9 | 121.8 | 127.26 | 125.15
-5.2 | 1.4 | 55.9913 | 149.9333 | 147.35 | 1.5992 | 78.6 | 146.775 | 144.0434 | 2.3 | 108.05 | 140.9 | 138.2603 | 3.7018 | 123.06 | 132.384 | 130.043 | 5 | 121.8 | 126.75 | 124.65
-5.1 | 1.5 | 71.5463 | 148.725 | 145.9 | 1.6991 | 93.6 | 145.625 | 142.871 | 2.4 | 116.65 | 140.05 | 137.4943 | 3.8019 | 123.8333 | 131.534 | 129.323 | 5.1 | 121.7 | 126.5 | 123.975
-5 | 1.5 | 89.0332 | 147.55 | 144.6 | 1.6991 | 106.3 | 144.225 | 141.9543 | 2.4333 | 122 | 138.9 | 136.511 | 3.9019 | 124.1 | 130.755 | 128.743 | 5.2 | 121.5 | 125.6 | 123.525
-4.9 | 1.5667 | 103.4223 | 146.375 | 143.5 | 1.7991 | 116.7 | 143.1167 | 140.77 | 2.5 | 125.1667 | 138.1 | 135.561 | 4.002 | 124.1 | 130.135 | 128.018 | 5.3 | 121.32 | 125.1 | 123.15
-4.8 | 1.6 | 116.113 | 145.1143 | 142.5 | 1.7991 | 124.2 | 142.1 | 139.8033 | 2.6 | 126.85 | 137 | 134.662 | 4.1354 | 124.1 | 129.285 | 127.543 | 5.4 | 120.85 | 124.5 | 122.6
-4.7 | 1.6 | 124.7565 | 143.7429 | 141.2714 | 1.899 | 127.8 | 141 | 138.7094 | 2.7 | 127.7 | 135.96 | 133.912 | 4.3022 | 123.9 | 128.536 | 126.742 | 5.65 | 120.6 | 123.8 | 121.95
-4.6 | 1.65 | 130.302 | 142.3429 | 139.95 | 1.9324 | 130.1 | 139.78 | 137.6403 | 2.8 | 128 | 135 | 132.862 | 4.4022 | 123.6 | 127.976 | 126.182 | 5.8 | 120.2571 | 123.34 | 121.7
-4.5 | 1.7 | 132.5673 | 141 | 138.7 | 1.999 | 131 | 138.42 | 136.368 | 2.8 | 128.1 | 134 | 132.062 | 4.5023 | 123.2 | 127.236 | 125.422 | 5.9667 | 120.06 | 122.775 | 121.0333
-4.4 | 1.7 | 133.5 | 139.8 | 137.5 | 2.0989 | 131.2 | 137.5 | 135.368 | 3 | 128 | 133.18 | 131.163 | 4.7024 | 122.7 | 126.587 | 124.642 | 6.1 | 119.7 | 122.25 | 120.25
-4.3 | 1.8 | 133.7 | 138.66 | 136.46 | 2.1489 | 131.125 | 136.275 | 134.247 | 3.1 | 127.9 | 131.96 | 130.138 | 4.9024 | 122.4 | 125.937 | 124.166 | 6.3 | 119.24 | 121.55 | 119.9
-4.2 | 1.9 | 133.7 | 137.425 | 135.3 | 2.1989 | 131.1 | 135.1 | 133.2333 | 3.2 | 127.5 | 131.175 | 129.363 | 5.0025 | 122.1 | 125.1203 | 123.416 | 6.6 | 118.8 | 121.075 | 119.525
-4.1 | 1.9 | 133.6 | 136 | 134.1 | 2.2988 | 130.5 | 133.9 | 132.1946 | 3.3667 | 126.8 | 130.12 | 128.363 | 5.2026 | 121.5 | 124.318 | 122.916 | 6.85 | 118.5 | 120.55 | 118.85
-4 | 1.925 | 132.9667 | 135 | 133.0333 | 2.3988 | 130 | 132.95 | 131.0058 | 3.5 | 126.2 | 129.2429 | 127.504 | 5.4027 | 121 | 123.613 | 122.141 | 7.0667 | 118.1 | 120.1 | 118.4333
-3.9 | 2 | 132.051 | 133.9 | 131.9 | 2.4987 | 129.35 | 131.7333 | 130.165 | 3.6667 | 125.6667 | 128.3857 | 126.6783 | 5.7028 | 120.65 | 122.889 | 121.621 | 7.4333 | 117.75 | 119.56 | 118.0667
-3.8 | 2.1 | 131.102 | 132.45 | 130.7333 | 2.5987 | 128.4 | 130.8 | 128.864 | 3.85 | 125 | 127.44 | 125.924 | 6.003 | 120.15 | 122.299 | 121.24 | 7.75 | 117.4 | 118.9 | 117.55
-3.7 | 2.2 | 130.003 | 131.4 | 129.8 | 2.8236 | 127.75 | 129.76 | 127.964 | 4.1 | 124.4 | 126.6 | 125.164 | 6.3031 | 119.75 | 121.739 | 120.44 | 8.1 | 117.14 | 118.56 | 117.1333
-3.6 | 2.3 | 129.403 | 130.4 | 128.5 | 2.9485 | 127.075 | 128.5 | 127.124 | 4.2333 | 123.85 | 125.5286 | 124.1448 | 6.57 | 119.2333 | 121.0395 | 119.84 | 8.675 | 116.74 | 118.12 | 116.75
-3.5 | 2.4 | 128.304 | 129.25 | 127.4 | 3.1318 | 126.05 | 127.6 | 126.163 | 4.6 | 123.2667 | 124.78 | 123.265 | 7.1036 | 118.8 | 120.34 | 119.19 | 9.5 | 116.2667 | 117.6 | 116.3
-3.4 | 2.6 | 127.5043 | 128.125 | 126.5 | 3.3983 | 125.4 | 126.6 | 125.313 | 4.9 | 122.6 | 124.1 | 122.565 | 7.6038 | 118.3 | 119.84 | 118.74 | 10.35 | 115.95 | 116.975 | 115.86
-3.3 | 2.8 | 126.68 | 127.2286 | 125.62 | 3.6982 | 124.65 | 125.74 | 124.362 | 5.3333 | 122.1 | 123.1 | 121.965 | 8.3041 | 117.85 | 119.34 | 118.159 | 11.7 | 115.55 | 116.7 | 115.3571
-3.2 | 3 | 125.606 | 126.2 | 124.5 | 4.0979 | 123.8 | 125 | 123.612 | 5.8 | 121.2 | 122.35 | 121.165 | 9.3713 | 117.35 | 118.7743 | 117.689 | 13.4 | 115.2 | 116.2 | 114.9
-3.1 | 3.3 | 124.806 | 125.35 | 123.8 | 4.4644 | 123.1 | 124 | 122.761 | 6.4667 | 120.6 | 121.7 | 120.541 | 10.4052 | 117.1 | 118.041 | 117.039 | 15.5 | 114.75 | 115.62 | 114.45
-3 | 3.7 | 123.907 | 124.36 | 123.1 | 5.1474 | 122.45 | 123 | 121.961 | 7.3 | 119.9 | 121.1 | 119.866 | 12.1394 | 116.7 | 117.421 | 116.589 | 18.8 | 114.4333 | 115.1714 | 114.16
-2.9 | 4.2 | 123.0413 | 123.35 | 122.1 | 5.8971 | 121.6 | 122.175 | 120.9603 | 8.475 | 119.5 | 120.175 | 118.791 | 14.7074 | 115.9 | 116.9745 | 116.089 | 23.3 | 113.9 | 114.68 | 113.62
-2.8 | 5.1 | 122.308 | 122.4333 | 121.15 | 7.1964 | 120.95 | 121.3667 | 120.3029 | 10.16 | 118.8 | 119.4 | 118.391 | 17.7839 | 115.525 | 116.3295 | 115.539 | 30 | 113.35 | 114.1 | 113.2
-2.7 | 6.1667 | 121.459 | 121.6 | 120.5 | 8.8456 | 120.2 | 120.6 | 119.46 | 12.8 | 118.1 | 118.525 | 117.5862 | 23.2116 | 115.1 | 115.7545 | 115.038 | 38.4 | 112.95 | 113.64 | 112.7
-2.6 | 7.9 | 120.6595 | 120.775 | 119.5 | 11.0945 | 119.5 | 119.74 | 118.8342 | 16.1 | 117.6 | 117.95 | 117.0003 | 30.1151 | 114.35 | 115.192 | 114.388 | 48.7 | 112.62 | 113.2 | 112.1714
-2.5 | 10.8 | 119.71 | 119.7833 | 118.7 | 14.9426 | 118.65 | 118.975 | 118.009 | 21.4 | 116.8 | 117.28 | 116.367 | 39.3196 | 113.85 | 114.623 | 113.663 | 59.4 | 112.25 | 112.66 | 111.68
-2.4 | 15.2 | 118.961 | 119.04 | 118 | 20.9895 | 117.8 | 118.3 | 117.359 | 29.1 | 116.1667 | 116.56 | 115.6956 | 50.8588 | 113.275 | 113.968 | 113.1951 | 70.6 | 111.7 | 112.1556 | 111.2286
-2.3 | 22.18 | 118.211 | 118 | 117.05 | 29.2854 | 117.1667 | 117.22 | 116.608 | 39.3 | 115.5 | 115.92 | 114.867 | 61.4307 | 112.9 | 113.323 | 112.5978 | 80.4 | 111.36 | 111.88 | 110.675
-2.2 | 32.1 | 117.437 | 117.1429 | 116.0875 | 39.247 | 116.2 | 116.56 | 115.808 | 51.6 | 114.85 | 115.1 | 114.1503 | 73.4367 | 112.35 | 112.794 | 112.1513 | 88 | 110.95 | 111.34 | 110.4333
-2.1 | 44.0333 | 116.646 | 116.3667 | 115.6 | 51.7075 | 115.8 | 115.9 | 114.9075 | 63.4 | 114.25 | 114.4 | 113.6474 | 83.1416 | 111.9 | 112.344 | 111.737 | 94.4 | 110.4 | 110.7 | 109.925
-2 | 58.1 | 115.713 | 115.52 | 114.7 | 65.2674 | 114.9667 | 115 | 114.1945 | 75.1667 | 113.6 | 113.5714 | 112.968 | 91.1456 | 111.3 | 111.6154 | 111.137 | 98.6 | 109.95 | 110.2 | 109.4
-1.9 | 70.2 | 114.814 | 114.9167 | 113.8286 | 77.2614 | 114.1 | 114.05 | 113.357 | 85.15 | 113 | 112.9 | 112.148 | 96.7984 | 110.7 | 110.9644 | 110.337 | 101.6 | 109.5 | 109.8 | 108.86
-1.8 | 82.3 | 114.0643 | 114 | 113.1714 | 87.4563 | 113.1667 | 113.1 | 112.6562 | 93.25 | 112.2 | 112.2 | 111.368 | 100.5837 | 110.05 | 110.4783 | 109.8941 | 103.1 | 109 | 109.1 | 108.46
-1.7 | 92.2 | 113.065 | 113.26 | 112.4 | 94.8193 | 112.4 | 112.5 | 111.936 | 98.55 | 111.6 | 111.3667 | 110.848 | 102.851 | 109.5 | 109.695 | 109.336 | 104.1333 | 108.4 | 108.42 | 108.1
-1.6 | 99.3 | 112.316 | 112.3 | 111.3 | 100.55 | 111.6667 | 111.7 | 111.2055 | 102.05 | 110.8333 | 110.6714 | 110.1884 | 104.252 | 109.1 | 109.112 | 108.8789 | 104.6 | 108.1 | 107.975 | 107.65
-1.5 | 103.2333 | 111.616 | 111.5286 | 110.75 | 103.348 | 110.825 | 111.05 | 110.455 | 104.05 | 110.1 | 110.14 | 109.644 | 104.952 | 108.5667 | 108.6793 | 108.236 | 105 | 107.65 | 107.55 | 107.2
-1.4 | 105.3 | 110.717 | 110.7571 | 110 | 104.8975 | 109.9 | 110.275 | 109.8217 | 104.85 | 109.575 | 109.575 | 108.844 | 105.353 | 107.95 | 108.0031 | 107.686 | 105.2 | 107.1 | 107.2 | 106.75
-1.3 | 106.14 | 110.217 | 109.9 | 109.4 | 105.4803 | 109.3 | 109.575 | 109.0148 | 105.4 | 108.9 | 108.625 | 108.369 | 105.428 | 107.3 | 107.406 | 107.216 | 105.2 | 106.65 | 106.6 | 106.22
-1.2 | 106.5 | 109.218 | 109.02 | 108.7 | 105.747 | 108.6667 | 108.7 | 108.404 | 105.5333 | 108.05 | 107.95 | 107.629 | 105.453 | 106.7 | 106.9864 | 106.836 | 105.1 | 106.225 | 106.0333 | 105.6
-1.1 | 106.5 | 108.269 | 108.15 | 107.95 | 105.747 | 107.8 | 107.9667 | 107.654 | 105.58 | 107.6 | 107.325 | 106.9979 | 105.353 | 106.2 | 106.327 | 106.135 | 104.9 | 105.7333 | 105.45 | 105.3
-1 | 106.5 | 107.5857 | 107.4667 | 107.2 | 105.747 | 107.15 | 107.15 | 107.054 | 105.4 | 106.72 | 106.5 | 106.22 | 105.0863 | 105.8 | 105.797 | 105.535 | 104.5667 | 105.1333 | 105 | 104.8
-0.9 | 106.0667 | 106.92 | 106.72 | 106.375 | 105.4137 | 106.35 | 106.46 | 106.203 | 105 | 106.2 | 105.8 | 105.65 | 104.6853 | 105 | 105.147 | 105.0017 | 104.1 | 104.75 | 104.55 | 104.3333
-0.8 | 105.525 | 105.821 | 106.0667 | 105.675 | 104.9147 | 105.55 | 105.7 | 105.528 | 104.7 | 105.4333 | 105.1 | 104.99 | 104.302 | 104.6 | 104.398 | 104.385 | 103.6 | 104.325 | 103.9 | 103.65
-0.7 | 105.125 | 105.221 | 105.28 | 105 | 104.4147 | 104.9 | 105.025 | 104.7857 | 104.1 | 104.6 | 104.58 | 104.25 | 103.802 | 104 | 103.848 | 103.935 | 103.2 | 103.62 | 103.45 | 103.25
-0.6 | 104.175 | 104.522 | 104.3857 | 104.25 | 103.848 | 104.25 | 104.2 | 104.232 | 103.7333 | 104 | 103.9333 | 103.69 | 103.352 | 103.5 | 103.398 | 103.4546 | 102.95 | 103 | 103.05 | 102.875
-0.5 | 103.7 | 103.722 | 103.65 | 103.475 | 103.348 | 103.4333 | 103.575 | 103.5091 | 103.1 | 103.425 | 103.2 | 103.0708 | 102.901 | 102.8 | 102.849 | 102.884 | 102.5 | 102.7 | 102.65 | 102.45
-0.4 | 102.8 | 103.0563 | 103.1 | 102.86 | 102.799 | 102.6 | 102.8 | 102.951 | 102.6 | 102.7333 | 102.6333 | 102.371 | 102.551 | 102.3 | 102.249 | 102.284 | 102.1 | 102.275 | 102.05 | 101.95
-0.3 | 102.3 | 102.273 | 102.22 | 102.1 | 101.9157 | 102.2 | 102.22 | 102.1081 | 101.9 | 101.975 | 101.8333 | 101.7996 | 102.051 | 101.725 | 101.849 | 101.754 | 101.5 | 101.85 | 101.46 | 101.64
-0.2 | 101.35 | 101.424 | 101.56 | 101.375 | 101.674 | 101.3 | 101.4 | 101.451 | 101.6 | 101.6 | 101.25 | 101.2281 | 101.201 | 101.2 | 100.9248 | 101.294 | 101 | 101.25 | 101.08 | 101.2
-0.1 | 100.6333 | 100.724 | 100.66 | 100.7 | 100.75 | 100.6 | 100.85 | 100.6833 | 100.9 | 100.8 | 100.65 | 100.571 | 100.5 | 100.5667 | 100.37 | 100.5835 | 100.5333 | 100.825 | 100.6 | 100.66
0 | 100 | 100.0001 | 100 | 100 | 100.0001 | 100 | 100 | 100 | 100 | 100 | 100 | 99.9997 | 100 | 100 | 100 | 99.9998 | 100 | 100 | 100 | 100
0.1 | 99.3 | 99.2256 | 99.3 | 99.4 | 99.2504 | 99.2 | 99.375 | 99.3782 | 99.5 | 99.3 | 99.3 | 99.3716 | 99.4748 | 99.7333 | 99.2504 | 99.6457 | 99.5 | 99.66 | 99.525 | 99.52
0.2 | 98.5333 | 98.5761 | 98.52 | 98.76 | 98.5507 | 98.55 | 98.8 | 98.7494 | 98.7 | 98.7667 | 98.5 | 98.7593 | 98.6493 | 98.9 | 98.8006 | 99.083 | 99 | 99.2125 | 99 | 99.15
0.3 | 97.5667 | 97.8932 | 97.8 | 98.06 | 97.7511 | 97.85 | 98.075 | 98.0991 | 98.15 | 98.2 | 98 | 98.172 | 98.1991 | 98.2 | 98.3009 | 98.6128 | 98.5 | 98.7 | 98.6 | 98.88
0.4 | 96.9 | 97.277 | 97.3 | 97.48 | 97.1181 | 97.1 | 97.45 | 97.5488 | 97.3 | 97.4333 | 97.5333 | 97.4922 | 97.5488 | 97.8 | 97.7178 | 97.9827 | 97.9 | 98.3 | 98.05 | 98.26
0.5 | 96.2 | 96.6275 | 96.5286 | 96.775 | 96.3019 | 96.45 | 96.7 | 96.9485 | 96.5 | 96.9 | 96.9 | 97.0723 | 96.8484 | 97.2333 | 97.0515 | 97.5325 | 97.3333 | 97.6 | 97.4 | 97.78
0.6 | 95.45 | 95.8781 | 95.86 | 96 | 95.4856 | 95.9 | 95.98 | 96.2481 | 96 | 96.3 | 96.15 | 96.2725 | 96.1481 | 96.6667 | 96.5318 | 97.0323 | 96.9 | 97.08 | 97.1 | 97.45
0.7 | 94.525 | 95.1786 | 95.25 | 95.45 | 94.8026 | 95.2 | 95.35 | 95.4477 | 95.3 | 95.5 | 95.6 | 95.6926 | 95.7479 | 96.1 | 96.002 | 96.4822 | 96.4333 | 96.4 | 96.5857 | 96.9
0.8 | 93.9 | 94.3959 | 94.54 | 94.76 | 93.7532 | 94.6333 | 94.8 | 95.0975 | 94.4 | 94.95 | 95.02 | 95.2128 | 95.1476 | 95.6 | 95.4523 | 96.032 | 95.7 | 96 | 96.05 | 96.38
0.9 | 93 | 93.6698 | 93.925 | 94.1857 | 92.8536 | 93.8 | 93.9833 | 94.3972 | 93.56 | 94.3 | 94.25 | 94.433 | 94.2471 | 95.1 | 94.8526 | 95.3318 | 94.9 | 95.7 | 95.65 | 96.1286
1 | 91.85 | 93.0302 | 93.275 | 93.525 | 92.054 | 93.1 | 93.48 | 93.7219 | 92.75 | 93.7333 | 93.575 | 93.9332 | 93.4467 | 94.2667 | 94.4028 | 94.9316 | 94.1 | 95.18 | 95.34 | 95.4
1.1 | 90.85 | 92.3058 | 92.625 | 92.98 | 90.7879 | 92.4333 | 92.7286 | 93.0065 | 91.825 | 93.1667 | 93 | 93.4233 | 92.8965 | 93.85 | 93.9031 | 94.4565 | 93.3 | 94.46 | 94.6 | 95.05
1.2 | 89.7 | 91.6562 | 91.88 | 92.375 | 89.2554 | 91.85 | 92.1333 | 92.4963 | 90.3 | 92.36 | 92.5667 | 92.7735 | 91.846 | 93.2 | 93.2534 | 93.9027 | 92.85 | 94.075 | 94.05 | 94.56
1.3 | 87.8 | 90.8319 | 91.34 | 91.65 | 87.9061 | 91.1333 | 91.5 | 92.046 | 89.1 | 91.9 | 91.85 | 92.1403 | 90.7954 | 92.5 | 92.6737 | 93.206 | 92.1 | 93.7 | 93.55 | 94.125
1.4 | 85.8 | 90.2323 | 90.6667 | 91.1 | 86.1569 | 90.5 | 90.8 | 91.3742 | 87.8 | 91.1 | 91.2 | 91.6238 | 89.7949 | 92.1 | 92.3538 | 92.7509 | 91.275 | 93.04 | 92.95 | 93.74
1.5 | 82.96 | 89.6661 | 89.98 | 90.45 | 83.3583 | 89.9 | 90.1 | 90.612 | 86 | 90.55 | 90.66 | 90.949 | 88.5443 | 91.5 | 91.8541 | 92.2507 | 90.4 | 92.5333 | 92.75 | 93.1
1.6 | 78.3 | 89.0332 | 89.4 | 89.8333 | 79.6602 | 89.2333 | 89.45 | 90.125 | 83.4 | 90 | 90.1 | 90.5741 | 86.9935 | 91 | 91.083 | 91.7906 | 89.3 | 92.04 | 92.075 | 92.65
1.7 | 71.7 | 88.4337 | 88.625 | 89.15 | 74.3378 | 88.6 | 89.05 | 89.5114 | 79.8 | 89.2 | 89.4 | 89.9343 | 84.8424 | 90.2 | 90.5947 | 91.2104 | 88.1 | 91.4 | 91.6667 | 92.25
1.8 | 62.8 | 87.5343 | 87.95 | 88.65 | 67.3663 | 87.9 | 88.4 | 88.8644 | 74.6333 | 88.8 | 88.7714 | 89.3745 | 82.2411 | 89.7 | 90.005 | 90.7102 | 86.6 | 90.92 | 91.1 | 91.85
1.9 | 53.7 | 86.9015 | 87.48 | 87.98 | 59.2454 | 87.325 | 87.74 | 88.3442 | 67.925 | 88 | 88.2 | 88.7746 | 78.7394 | 89.2667 | 89.6052 | 90.2551 | 84.6 | 90.4 | 90.65 | 91.275
2 | 42.65 | 86.3019 | 86.85 | 87.3571 | 49.7751 | 86.5 | 87.1375 | 87.7439 | 59.725 | 87.4 | 87.5429 | 88.3176 | 74.1371 | 88.4 | 88.9305 | 89.7299 | 82.0667 | 89.8 | 90.12 | 90.7
2.1 | 33 | 85.7357 | 86.2 | 86.7 | 39.5302 | 85.84 | 86.46 | 87.1936 | 50.3 | 86.9 | 87 | 87.7549 | 67.6672 | 87.9 | 88.4558 | 89.2011 | 78.7 | 89.2571 | 89.7 | 90.4
2.2 | 23.4333 | 85.1028 | 85.8 | 86.225 | 30.2849 | 85.5 | 85.9 | 86.5933 | 40.725 | 86.1667 | 86.55 | 87.1251 | 60.5303 | 87.3 | 87.8161 | 88.7296 | 74.3 | 88.8 | 89.25 | 89.86
2.3 | 16.6 | 84.3367 | 85 | 85.6333 | 21.7558 | 84.6 | 85.4 | 85.943 | 31.8 | 85.55 | 85.95 | 86.6086 | 51.6258 | 86.95 | 87.3563 | 88.1294 | 68.1 | 88.4 | 88.7 | 89.4
2.4 | 11.5 | 83.8871 | 84.5571 | 85.15 | 15.6922 | 84.1 | 84.6667 | 85.4427 | 23.8 | 84.9 | 85.45 | 85.9254 | 41.981 | 86.3333 | 86.7066 | 87.6492 | 60.4 | 87.6833 | 88.1333 | 88.9
2.5 | 8.2 | 83.2376 | 83.85 | 84.6 | 11.5609 | 83.3 | 84.2 | 84.8424 | 17.5 | 84.3 | 84.825 | 85.5006 | 33.917 | 85.8 | 86.1569 | 87.2576 | 52.5 | 87.18 | 87.625 | 88.45
2.6 | 6.3 | 82.6047 | 83.35 | 84.05 | 8.6957 | 82.9 | 83.6 | 84.3822 | 13.3 | 83.8 | 84.15 | 84.9757 | 26.4132 | 85.1333 | 85.7071 | 86.7718 | 44.0333 | 86.85 | 87.1 | 87.95
2.7 | 4.9 | 82.0385 | 82.74 | 83.6 | 6.9299 | 82.3 | 83 | 83.7669 | 10.45 | 82.9667 | 83.7 | 84.4759 | 20.1434 | 84.8 | 85.1974 | 86.3288 | 36 | 86.2333 | 86.56 | 87.55
2.8 | 4 | 81.4639 | 82.3 | 83.05 | 5.5972 | 81.75 | 82.6 | 83.1916 | 8.5 | 82.5 | 83.15 | 83.926 | 15.6078 | 84.1 | 84.7976 | 85.6952 | 28.3 | 85.525 | 86.1333 | 87.0667
2.9 | 3.4 | 80.9393 | 81.85 | 82.54 | 4.7476 | 81.1 | 82.1333 | 82.8414 | 7.1 | 81.9 | 82.6 | 83.3762 | 12.6063 | 83.2 | 84.1437 | 85.2284 | 22.1 | 84.925 | 85.75 | 86.6
3 | 2.95 | 80.3897 | 81.4 | 82 | 4.048 | 80.6 | 81.55 | 82.4412 | 6.35 | 81.3 | 82.1 | 82.8263 | 10.6053 | 82.825 | 83.6382 | 84.7683 | 17.8 | 84.525 | 85.4333 | 86.15
3.1 | 2.6 | 79.7402 | 80.75 | 81.6 | 3.6482 | 80.1 | 81.1 | 81.7837 | 5.5 | 80.9 | 81.6 | 82.3765 | 9.1045 | 82.3 | 83.3183 | 84.2781 | 14.7 | 84 | 85 | 85.8
3.2 | 2.3667 | 79.2406 | 80.15 | 81.16 | 3.2984 | 79.45 | 80.6 | 81.4074 | 4.9333 | 80.3667 | 81.175 | 81.9266 | 8.1041 | 81.8 | 82.8336 | 83.7279 | 12.45 | 83.35 | 84.42 | 85.4
3.3 | 2.2 | 78.7409 | 79.7857 | 80.6333 | 2.9485 | 78.9333 | 80 | 80.7904 | 4.5 | 79.65 | 80.66 | 81.5017 | 7.2703 | 81.3667 | 82.3188 | 83.2027 | 10.8 | 83 | 83.95 | 84.76
3.4 | 2 | 78.1414 | 79.3 | 80.075 | 2.6987 | 78.3 | 79.46 | 80.2651 | 4.2 | 79.275 | 80.2 | 81.0369 | 6.6033 | 80.7 | 81.7592 | 82.8476 | 9.7 | 82.4 | 83.42 | 84.35
3.5 | 1.9 | 77.4419 | 78.75 | 79.6286 | 2.4987 | 77.7333 | 79.2 | 79.8399 | 3.8333 | 78.575 | 79.75 | 80.5436 | 6.3031 | 80.1667 | 81.3793 | 82.5075 | 8.8 | 81.9 | 83.05 | 83.875
3.6 | 1.8 | 76.9423 | 78.3 | 79.15 | 2.3988 | 77.05 | 78.625 | 79.4398 | 3.7 | 78 | 79.2667 | 80.0371 | 5.9029 | 79.6 | 80.9595 | 82.0473 | 7.9667 | 81.4 | 82.68 | 83.575
3.7 | 1.7333 | 76.3927 | 77.75 | 78.7 | 2.1989 | 76.45 | 78.1 | 78.9395 | 3.45 | 77.325 | 78.9333 | 79.5273 | 5.4694 | 79 | 80.3998 | 81.5021 | 7.6 | 80.8 | 82.2857 | 83.34
3.8 | 1.7 | 75.8431 | 77.34 | 78.16 | 2.0989 | 75.7667 | 77.575 | 78.4593 | 3.28 | 76.7 | 78.2667 | 79.1274 | 5.2026 | 78.42 | 79.96 | 81.102 | 7.175 | 80.4 | 81.7857 | 82.7
3.9 | 1.6 | 75.0187 | 76.8 | 77.7 | 1.999 | 75.15 | 77.1667 | 77.999 | 3.1 | 76.1 | 77.875 | 78.6275 | 4.9525 | 77.9 | 79.4802 | 80.6769 | 6.7667 | 79.75 | 81.34 | 82.4
4 | 1.5 | 74.2943 | 76.3 | 77.3 | 1.899 | 74.5 | 76.78 | 77.4959 | 2.95 | 75.6 | 77.275 | 78.1027 | 4.8024 | 77.35 | 79.0891 | 80.2517 | 6.375 | 79.45 | 80.86 | 81.9
4.1 | 1.5 | 73.6114 | 75.9 | 76.675 | 1.7991 | 73.7 | 76.2857 | 76.9985 | 2.8 | 74.9 | 76.9 | 77.7028 | 4.6023 | 76.8 | 78.6035 | 79.8266 | 6.2 | 78.85 | 80.3 | 81.4429
4.2 | 1.4 | 72.7954 | 75.3667 | 76.15 | 1.7491 | 72.9333 | 75.7143 | 76.6383 | 2.7 | 74.2 | 76.45 | 77.2279 | 4.4022 | 76.1333 | 78.061 | 79.3264 | 5.975 | 78.3 | 79.9 | 81.2
4.3 | 1.4 | 71.7212 | 74.85 | 75.85 | 1.6991 | 72.1 | 75.24 | 76.1381 | 2.6 | 73.45 | 76 | 76.7581 | 4.2021 | 75.55 | 77.7112 | 78.8763 | 5.7 | 77.65 | 79.5 | 80.6571
4.4 | 1.4 | 70.3472 | 74.5 | 75.4 | 1.5992 | 71.125 | 74.8 | 75.7379 | 2.5 | 72.6333 | 75.6 | 76.4781 | 4.1021 | 75.05 | 77.2214 | 78.2261 | 5.6 | 77.2 | 79.075 | 80.2286
4.5 | 1.3 | 68.5985 | 73.85 | 74.9 | 1.5992 | 70 | 74.34 | 75.1376 | 2.4 | 71.8 | 75 | 75.8783 | 4.002 | 74.15 | 76.8187 | 77.926 | 5.4 | 76.7333 | 78.5667 | 79.88
4.6 | 1.3 | 66.3003 | 73.4 | 74.5 | 1.4992 | 68.3 | 73.8 | 74.7974 | 2.3 | 70.8 | 74.6 | 75.4284 | 3.8686 | 73.6 | 76.4868 | 77.5258 | 5.2 | 76.1 | 78.14 | 79.4444
4.7 | 1.2 | 62.3199 | 72.975 | 73.9 | 1.4992 | 66.2 | 73.2 | 74.3372 | 2.2 | 69.7 | 74.06 | 74.9586 | 3.7018 | 72.9 | 75.9049 | 77.0457 | 5.1 | 75.62 | 77.675 | 79
4.8 | 1.2 | 57.0572 | 72.55 | 73.4667 | 1.3993 | 63.2 | 72.85 | 73.8119 | 2.1 | 68.3667 | 73.6 | 74.5787 | 3.6018 | 72.2667 | 75.4123 | 76.6505 | 4.9 | 75 | 77.4 | 78.6333
4.9 | 1.2 | 50.2623 | 72.1333 | 73.0333 | 1.3993 | 58.8 | 72.32 | 73.3867 | 2 | 66.5 | 73.14 | 74.2121 | 3.5017 | 71.7 | 74.8959 | 76.2504 | 4.9 | 74.6 | 77.1 | 78.2
5 | 1.1 | 42.4082 | 71.6333 | 72.55 | 1.3327 | 53.2 | 71.9 | 72.8793 | 2 | 64.1 | 72.62 | 73.629 | 3.4017 | 70.7 | 74.3379 | 75.6752 | 4.7 | 73.95 | 76.6 | 77.6857
5.1 | 1.1 | 34.5407 | 71.1 | 72.125 | 1.2994 | 46.3 | 71.28 | 72.3362 | 1.95 | 60.6 | 72.05 | 73.1791 | 3.3016 | 69.9 | 74.003 | 75.3651 | 4.6 | 73.3667 | 76.1 | 77.35
5.2 | 1.1 | 26.8798 | 70.65 | 71.675 | 1.2994 | 39 | 70.9571 | 71.996 | 1.9 | 56 | 71.54 | 72.7459 | 3.2016 | 68.8 | 73.6132 | 74.825 | 4.45 | 72.55 | 75.76 | 77
5.3 | 1.1 | 20.3181 | 70.2 | 71.175 | 1.1994 | 31.3 | 70.5667 | 71.4357 | 1.8 | 50.2 | 71.24 | 72.1794 | 3.1016 | 67.4 | 73.1968 | 74.3998 | 4.3 | 71.9 | 75.25 | 76.54
5.4 | 1.1 | 15.1487 | 69.75 | 70.75 | 1.1994 | 24.7 | 69.95 | 71.0555 | 1.8 | 43.25 | 70.84 | 71.7295 | 3.0015 | 65.7 | 72.6637 | 73.9996 | 4.2 | 71.24 | 74.775 | 76
5.5 | 1 | 11.7912 | 69.2 | 70.2333 | 1.1994 | 18.7 | 69.45 | 70.4352 | 1.7 | 36.4 | 70.22 | 71.3796 | 3.0015 | 63.3 | 72.1389 | 73.6245 | 4.1333 | 70.4 | 74.525 | 75.68
5.6 | 1 | 9.4929 | 68.68 | 69.75 | 1.0995 | 14.6 | 69.025 | 70.1351 | 1.7 | 29.35 | 69.74 | 70.8464 | 2.9015 | 60.1 | 71.6642 | 73.0743 | 4.1 | 69.4 | 73.95 | 75.4143
5.7 | 1 | 8.044 | 68.25 | 69.34 | 1.0995 | 11.6 | 68.54 | 69.7549 | 1.65 | 23.3333 | 69.275 | 70.3999 | 2.8014 | 55.8 | 71.2644 | 72.6909 | 4 | 68.4 | 73.45 | 74.9
5.8 | 1 | 6.9947 | 67.78 | 68.82 | 1.0995 | 9.9 | 68.0333 | 69.0345 | 1.6 | 18.8667 | 68.7 | 69.9134 | 2.7264 | 50.5 | 70.8146 | 72.2241 | 3.85 | 67.06 | 73.05 | 74.42
5.9 | 1 | 6.3952 | 67.36 | 68.32 | 1.0995 | 8.6 | 67.6 | 68.5743 | 1.5 | 15.1 | 68.26 | 69.3202 | 2.7014 | 44.4 | 70.3648 | 71.7039 | 3.75 | 65.38 | 72.65 | 74.1
6 | 0.9 | 5.9455 | 66.78 | 67.8 | 1.0995 | 7.7 | 67.02 | 68.059 | 1.5 | 12.85 | 67.88 | 68.9303 | 2.6013 | 38.1 | 69.8401 | 71.2523 | 3.7 | 63.05 | 72.3 | 73.7444
6.1 | 0.9 | 5.5458 | 66.3 | 67.35 | 0.9995 | 7.05 | 66.52 | 67.5838 | 1.5 | 11.25 | 67.3 | 68.4805 | 2.5013 | 32.15 | 69.5153 | 70.8236 | 3.6 | 60.325 | 71.8 | 73.1857
6.2 | 0.9 | 5.2461 | 65.75 | 66.8 | 0.9995 | 6.55 | 66.04 | 67.1335 | 1.5 | 10.16 | 66.7 | 67.8556 | 2.4012 | 26.7 | 68.9322 | 70.2734 | 3.5 | 56.32 | 71.4 | 72.64
6.3 | 0.9 | 4.9962 | 65.2 | 66.275 | 0.9995 | 6.15 | 65.45 | 66.5933 | 1.4 | 9.4 | 66.36 | 67.4608 | 2.4012 | 22.2 | 68.5057 | 69.8233 | 3.4 | 51.4 | 70.95 | 72.25
6.4 | 0.9 | 4.7964 | 64.6143 | 65.78 | 0.9995 | 5.8 | 64.95 | 66.0531 | 1.4 | 8.825 | 65.8571 | 66.9309 | 2.3512 | 18.7 | 68.016 | 69.3231 | 3.3 | 45.775 | 70.4 | 71.76
6.5 | 0.9 | 4.6965 | 64.1143 | 65.25 | 0.9995 | 5.55 | 64.375 | 65.4828 | 1.4 | 8.25 | 65.2 | 66.3525 | 2.3011 | 16.3 | 67.4913 | 68.773 | 3.3 | 39.8 | 69.875 | 71.3
6.6 | 0.9 | 4.4966 | 63.46 | 64.5 | 0.9995 | 5.3 | 63.8333 | 64.9825 | 1.3 | 7.85 | 64.8 | 65.8062 | 2.3011 | 14.7 | 66.9998 | 68.4228 | 3.3 | 34.34 | 69.45 | 70.92
6.7 | 0.9 | 4.3967 | 62.72 | 64 | 0.89955 | 5.1 | 63.2 | 64.3072 | 1.3 | 7.45 | 64.16 | 65.3314 | 2.2011 | 13.5 | 66.4668 | 68.0727 | 3.2 | 29 | 69.04 | 70.5286
6.8 | 0.8 | 4.2968 | 61.9667 | 63.2 | 0.89955 | 4.96 | 62.44 | 63.7319 | 1.3 | 7.15 | 63.58 | 64.8314 | 2.101 | 12.5 | 66.0336 | 67.5225 | 3.1333 | 24.66 | 68.64 | 70
6.9 | 0.8 | 4.1469 | 61.1 | 62.24 | 0.89955 | 4.8 | 61.7 | 63.0715 | 1.2 | 6.9 | 63.12 | 64.2531 | 2.101 | 11.9 | 65.4273 | 66.9723 | 3.1 | 21.25 | 68.2 | 69.5778
7 | 0.8 | 3.997 | 59.8 | 61.15 | 0.89955 | 4.65 | 61 | 62.2061 | 1.2 | 6.7 | 62.4 | 63.5485 | 2.101 | 11.3 | 65.0175 | 66.3971 | 3 | 18.9667 | 67.66 | 69.1429
7.1 |  | 3.997 | 58.1333 | 59.52 |  | 4.5 | 60 | 61.2306 | 1.2 | 6.4 | 61.775 | 63.007 | 2.001 | 10.85 | 64.5344 | 65.922 | 2.9 | 17.34 | 67.2 | 68.7
7.2 |  | 3.8971 | 55.7333 | 57.275 |  | 4.4 | 58.9 | 60.1634 | 1.2 | 6.3 | 61.2 | 62.3822 | 2.001 | 10.5 | 63.9347 | 65.5552 | 2.9 | 16 | 66.7375 | 68.3
7.3 |  | 3.7971 | 52.1 | 53.7 |  | 4.2 | 57.2 | 58.6293 |  | 6.1 | 60.475 | 61.5824 | 1.9009 | 10.15 | 63.4683 | 65.1217 | 2.9 | 15.1 | 66.4 | 67.8
7.4 |  | 3.6972 | 47.3667 | 49.26 |  | 4.175 | 55.1 | 56.5883 |  | 5.9 | 59.42 | 60.7826 | 1.9009 | 9.8 | 62.9935 | 64.4882 | 2.8 | 14.42 | 65.88 | 67.46
7.5 |  | 3.5973 | 41.45 | 43.65 |  | 4.05 | 51.875 | 53.6769 |  | 5.75 | 58.55 | 59.9114 |  | 9.5 | 62.3438 | 63.9613 | 2.7 | 13.7 | 65.4 | 66.9333
7.6 |  | 3.5973 | 35.15 | 37.6 |  | 3.975 | 47.56 | 49.4914 |  | 5.6 | 57.35 | 58.7083 |  | 9.3 | 61.9191 | 63.4711 | 2.6 | 13.275 | 64.975 | 66.22
7.7 |  | 3.4974 | 28.55 | 31.475 |  | 3.9 | 42.0125 | 44.4472 |  | 5.4 | 55.68 | 57.0586 |  | 9.075 | 61.3027 | 62.7709 |  | 12.8667 | 64.475 | 65.775
7.8 |  | 3.4974 | 22.9 | 25.72 |  | 3.8 | 36.1 | 38.8695 |  | 5.3 | 53.3 | 54.9843 |  | 8.8 | 60.6697 | 62.3608 |  | 12.5 | 63.95 | 65.45
7.9 |  | 3.3975 | 17.68 | 21 |  | 3.7 | 30.0333 | 33.1766 |  | 5.2 | 50.05 | 51.9052 |  | 8.6 | 60.0034 | 61.7456 |  | 12.0857 | 63.65 | 65.05
8 |  | 3.3975 | 14.06 | 17.64 |  | 3.6 | 24.44 | 27.6138 |  | 5 | 46.08 | 47.9006 |  | 8.4 | 59.3703 | 61.006 |  | 11.8857 | 63.0625 | 64.55
8.1 |  | 3.3975 | 11.62 | 15.38 |  | 3.6 | 19.45 | 22.9115 |  | 4.9 | 41.0667 | 43.1377 |  | 8.15 | 58.5957 | 60.3201 |  | 11.58 | 62.54 | 64.0667
8.2 |  | 3.3975 | 10.1 | 13.82 |  | 3.6 | 15.4 | 19.3097 |  | 4.8 | 35.32 | 37.7892 |  | 8 | 57.7112 | 59.4198 |  | 11.26 | 61.95 | 63.6
8.3 |  | 3.3975 | 9.18 | 12.98 |  | 3.5 | 12.9333 | 16.8084 |  | 4.7333 | 29.95 | 32.3908 |  | 7.8667 | 56.6716 | 58.3945 |  | 11.1 | 61.54 | 63.0667
8.4 |  | 3.2975 | 8.6 | 12.4 |  | 3.5 | 11.3 | 15.0276 |  | 4.6333 | 24.65 | 27.6921 |  | 7.6 | 55.2724 | 57.099 |  | 10.8333 | 60.8714 | 62.7
8.5 |  | 3.2975 | 8.1 | 11.9625 |  | 3.5 | 10.2 | 13.927 |  | 4.5667 | 20.15 | 23.5933 |  | 7.5 | 53.4733 | 55.5185 |  | 10.6 | 60.36 | 62
8.6 |  | 3.2975 | 7.78 | 11.625 |  | 3.5 | 9.46 | 13.1566 |  | 4.425 | 16.825 | 20.3941 |  | 7.3 | 51.1145 | 53.2177 |  | 10.44 | 59.7125 | 61.3667
8.7 |  | 3.2975 | 7.575 | 11.3333 |  | 3.4 | 8.85 | 12.6563 |  | 4.4 | 14.4667 | 18.1948 |  | 7.1 | 47.7761 | 50.1967 |  | 10.18 | 59.14 | 60.9
8.8 |  | 3.2975 | 7.35 | 11.2 |  | 3.4 | 8.44 | 12.2347 |  | 4.34 | 12.92 | 16.7952 |  | 7 | 43.464 | 46.3154 |  | 10.0333 | 58.45 | 60.3
8.9 |  | 3.2975 | 7.2 | 11 |  | 3.4 | 8.1667 | 11.906 |  | 4.3 | 11.92 | 15.7955 |  | 6.9 | 39.0005 | 41.994 |  | 9.9 | 57.6 | 59.42
9 |  | 3.1976 | 7.1 | 10.9 |  | 3.4 | 7.875 | 11.6201 |  | 4.25 | 11.25 | 14.9957 |  | 6.7 | 34.0116 | 37.5125 |  | 9.675 | 56.6714 | 58.48
9.1 |  | 3.1976 | 7 | 10.8 |  | 3.3 | 7.6667 | 11.3657 |  | 4.2 | 10.6 | 14.4959 |  | 6.6 | 29.3853 | 32.911 |  | 9.4667 | 55.56 | 57.425
9.2 |  | 3.1976 | 6.9 | 10.7 |  | 3.3 | 7.5 | 11.1681 |  | 4.2 | 10.2 | 14.0627 |  | 6.5 | 25.0304 | 28.9096 |  | 9.325 | 53.95 | 56.06
9.3 |  | 3.0977 | 6.875 | 10.6 |  | 3.3 | 7.375 | 11.0055 |  | 4.1 | 9.85 | 13.6961 |  | 6.4 | 21.3294 | 25.4685 |  | 9.02 | 52.16 | 54.26
9.4 |  | 3.0977 | 6.8 | 10.58 |  | 3.2 | 7.2 | 10.9222 |  | 4.1 | 9.5 | 13.4461 |  | 6.25 | 18.9655 | 23.1578 |  | 8.9 | 49.5143 | 51.94
9.5 |  | 3.0977 | 6.8 | 10.5 |  | 3.2 | 7.1 | 10.8054 |  | 4 | 9.24 | 13.1962 |  | 6.1 | 17.0515 | 21.3871 |  | 8.775 | 46.24 | 48.88
9.6 |  | 3.0644 | 6.8 | 10.5 |  | 3.2 | 7 | 10.7054 |  | 4 | 9.1 | 12.9296 |  | 6.1 | 15.6922 | 20.1067 |  | 8.7 | 42.45 | 45.16
9.7 |  | 2.9977 | 6.7 | 10.4333 |  | 3.1 | 6.925 | 10.6053 |  | 3.9 | 8.85 | 12.7214 |  | 6 | 14.7177 | 19.2064 |  | 8.48 | 37.9 | 41.2
9.8 |  | 2.9977 | 6.7 | 10.4 |  | 3.1 | 6.9 | 10.5553 |  | 3.9 | 8.7 | 12.5464 |  | 5.9333 | 14.068 | 18.6062 |  | 8.3 | 33.425 | 37.3
9.9 |  | 2.9977 | 6.7 | 10.4 |  | 3.1 | 6.9 | 10.5053 |  | 3.8 | 8.5 | 12.4365 |  | 5.9 | 13.5333 | 18.026 |  | 8.2 | 29.4 | 33.3
10 |  | 2.9977 | 6.7 | 10.3 |  | 3 | 6.8 | 10.4052 |  | 3.8 | 8.4 | 12.2465 |  | 5.8 | 13.0792 | 17.6202 |  | 8.0625 | 25.72 | 29.94
10.1 |  | 2.8978 | 6.6 | 10.3 |  | 3 | 6.72 | 10.3052 |  | 3.7 | 8.25 | 12.1215 |  | 5.7 | 12.7437 | 17.2057 |  | 7.9 | 22.9 | 27.5
10.2 |  |  | 6.6 | 10.3 |  | 2.925 | 6.7 | 10.3052 |  | 3.7 | 8.1 | 11.9966 |  | 5.7 | 12.4224 | 16.9557 |  | 7.8143 | 20.5 | 25.36
10.3 |  |  | 6.58 | 10.2 |  |  | 6.6333 | 10.2051 |  | 3.675 | 8 | 11.8966 |  | 5.6 | 12.074 | 16.6055 |  | 7.7 | 18.85 | 23.9
10.4 |  |  | 6.5 | 10.16 |  |  | 6.6 | 10.2051 |  | 3.6 | 7.9 | 11.7966 |  | 5.5 | 11.8541 | 16.2911 |  | 7.6 | 17.85 | 22.8
10.5 |  |  | 6.4 | 10.1 |  |  | 6.5333 | 10.1051 |  |  | 7.9 | 11.6967 |  | 5.4 | 11.5942 | 16.1721 |  | 7.5 | 16.9 | 21.9857
10.6 |  |  | 6.4 | 10 |  |  | 6.5 | 10.005 |  |  | 7.8 | 11.5467 |  | 5.3 | 11.3277 | 15.8553 |  | 7.45 | 16.3286 | 21.3
10.7 |  |  | 6.3 | 9.9571 |  |  | 6.4 | 10.005 |  |  | 7.7 | 11.4967 |  | 5.3 | 11.1944 | 15.6385 |  | 7.3 | 15.9 | 20.9
10.8 |  |  | 6.2 | 9.8 |  |  | 6.4 | 9.9049 |  |  | 7.6 | 11.3967 |  | 5.2 | 10.9945 | 15.4852 |  | 7.25 | 15.4 | 20.45
10.9 |  |  | 6.16 | 9.8 |  |  | 6.3 | 9.9049 |  |  | 7.5 | 11.2718 |  | 5.1 | 10.8146 | 15.2337 |  | 7.1333 | 15 | 20.15
11 |  |  | 6.1 | 9.7 |  |  | 6.2 | 9.8049 |  |  | 7.4 | 11.1968 |  |  | 10.6572 | 15.005 |  | 7.1 | 14.7 | 19.825
11.1 |  |  | 6 | 9.6333 |  |  | 6.1 | 9.7049 |  |  | 7.3 | 11.0968 |  |  | 10.4948 | 14.905 |  | 7 | 14.45 | 19.3714
11.2 |  |  | 5.9 | 9.5 |  |  | 6.1 | 9.6048 |  |  | 7.3 | 10.9969 |  |  | 10.3663 | 14.762 |  | 6.9 | 14.2 | 19.2
11.3 |  |  | 5.86 | 9.4 |  |  | 6 | 9.5333 |  |  | 7.2 | 10.8969 |  |  | 10.1749 | 14.6192 |  | 6.8 | 13.9 | 18.9333
11.4 |  |  | 5.8 | 9.3 |  |  | 5.9 | 9.4047 |  |  | 7.0667 | 10.7969 |  |  | 10.0783 | 14.4048 |  | 6.7333 | 13.64 | 18.6667
11.5 |  |  | 5.7 | 9.2333 |  |  | 5.8 | 9.3332 |  |  | 7 | 10.6969 |  |  | 9.975 | 14.3048 |  |  | 13.42 | 18.4667
11.6 |  |  | 5.6 | 9.1 |  |  | 5.8 | 9.2296 |  |  | 6.88 | 10.597 |  |  | 9.7951 | 14.1297 |  |  | 13.2 | 18.2
11.7 |  |  | 5.5 | 9.0833 |  |  | 5.7 | 9.1045 |  |  | 6.8 | 10.497 |  |  | 9.7351 | 14.0047 |  |  | 12.96 | 18.0667
11.8 |  |  | 5.475 | 9 |  |  | 5.6 | 9.0045 |  |  | 6.725 | 10.397 |  |  | 9.5752 | 13.9046 |  |  | 12.84 | 17.8667
11.9 |  |  | 5.4 | 8.9 |  |  | 5.55 | 8.9473 |  |  | 6.675 | 10.2971 |  |  | 9.4096 | 13.7046 |  |  | 12.66 | 17.66
12 |  |  | 5.3 | 8.8 |  |  | 5.42 | 8.8544 |  |  | 6.6 | 10.1971 |  |  | 9.2953 | 13.5795 |  |  | 12.5 | 17.5
12.1 |  |  | 5.2 | 8.7 |  |  | 5.3833 | 8.7544 |  |  | 6.4 | 10.0971 |  |  | 9.262 | 13.5045 |  |  | 12.3 | 17.32
12.2 |  |  | 5.2 | 8.6 |  |  | 5.3 | 8.7043 |  |  | 6.35 | 9.9971 |  |  | 9.0954 | 13.3044 |  |  | 12.1 | 17.1
12.3 |  |  | 5.1 | 8.5 |  |  | 5.2143 | 8.621 |  |  | 6.3 | 9.8972 |  |  | 8.9955 | 13.2044 |  |  | 11.975 | 16.9
12.4 |  |  | 5 | 8.4571 |  |  | 5.2 | 8.5243 |  |  | 6.2 | 9.7972 |  |  | 8.8123 | 13.0901 |  |  | 11.8 | 16.72
12.5 |  |  | 5 | 8.4 |  |  | 5.1 | 8.4709 |  |  | 6.1 | 9.6472 |  |  | 8.7956 | 13.0043 |  |  | 11.675 | 16.58
12.6 |  |  | 4.96 |  |  |  | 5 | 8.3542 |  |  | 6 | 9.5973 |  |  | 8.6207 | 12.9043 |  |  | 11.525 | 16.4429
12.7 | [6 empty] | 4.9857 | 8.2899 |  |  | 5.9 | 9.4973 |  |  | 8.4957 | 12.7042 |  |  | 11.4 | 16.2286
12.8 | [6 empty] | 4.9 | 8.2041 |  |  | 5.875 | 9.3973 |  |  | 8.4158 | 12.6042 |  |  | 11.25 | 16.025
12.9 | [10 empty] | 5.8 | 9.3173 |  |  | 8.3158 | 12.4041 |  |  | 11.175 | 15.8429
13 | [10 empty] | 5.7 | 9.2174 |  |  | 8.1959 | 12.3041 |  |  | 11.02 | 15.7
13.1 | [10 empty] | 5.6 | 9.0974 |  |  | 8.076 | 12.2541 |  |  | 10.9 | 15.55
13.2 | [10 empty] | 5.5 | 8.9974 |  |  | 7.996 | 12.104 |  |  | 10.775 | 15.5
13.3 | [14 empty] | 7.8711 | 11.9373 |  |  | 10.7 | 15.3
13.4 | [14 empty] | 7.7961 | 11.8039 |  |  | 10.6 | 15.2
13.5 | [14 empty] | 7.6762 | 11.6039 |  |  | 10.38 | 15.0571
13.6 | [14 empty] | 7.5962 | 11.5539 |  |  | 10.3 | 14.9143
13.7 | [14 empty] | 7.3963 | 11.5038 |  |  | 10.1 | 14.8
13.8 | [14 empty] | 7.2964 | 11.3038 |  |  | 9.95 | 14.675
13.9 | [14 empty] | 7.1964 | 11.2287 |  |  | 9.9 | 14.5333
14 | [18 empty] | 9.7667 | 14.4
14.1 | [18 empty] | 9.675 | 14.32
14.2 | [18 empty] | 9.525 | 14.12
14.3 | [18 empty] | 9.4 | 14.1
14.4 | [18 empty] | 9.3 | 13.9143
14.5 | [18 empty] | 9.2 | 13.68
14.6 | [18 empty] | 9.1 | 13.5
14.7 | [18 empty] | 8.925 | 13.48

## Wedge 60 Longitudinal Profiles
Longitudinal Profiles
Wedge 60
SSD=100cm
Field Size=40x15cm2
Scanning direction: Inline
Detector: IBA Dosimetry CC13 Ionization chamber
Dose[%]
Off axis position [cm] | Depth:2 cm | Depth:5 cm | Depth:10 cm | Depth:20 cm | Depth:30 cm
-35 |  |  |  |  | 7.775
-34.9 |  |  |  |  | 7.8125
-34.8 |  |  |  |  | 7.9375
-34.7 |  |  |  |  | 8
-34.6 |  |  |  |  | 8.1
-34.5 |  |  |  |  | 8.1375
-34.4 |  |  |  |  | 8.2
-34.3 |  |  |  |  | 8.3
-34.2 |  |  |  |  | 8.3286
-34.1 |  |  |  |  | 8.4
-34 |  |  |  |  | 8.4875
-33.9 |  |  |  |  | 8.525
-33.8 |  |  |  |  | 8.6571
-33.7 |  |  |  |  | 8.7778
-33.6 |  |  |  |  | 8.8125
-33.5 |  |  |  |  | 8.9375
-33.4 |  |  |  |  | 9
-33.3 |  |  |  |  | 9.1
-33.2 |  |  |  |  | 9.1286
-33.1 |  |  |  |  | 9.2
-33 |  |  |  |  | 9.3
-32.9 |  |  |  |  | 9.4286
-32.8 |  |  |  |  | 9.5
-32.7 |  |  |  |  | 9.6
-32.6 |  |  |  |  | 9.7154
-32.5 |  |  |  |  | 9.7923
-32.4 |  |  |  |  | 9.8563
-32.3 |  |  |  | 7.0375 | 9.92
-32.2 |  |  |  | 7.1 | 9.9867
-32.1 |  |  |  | 7.1 | 10.16
-32 |  |  |  | 7.2 | 10.32
-31.9 |  |  |  | 7.2 | 10.3867
-31.8 |  |  |  | 7.3 | 10.4533
-31.7 |  |  |  | 7.3375 | 10.54
-31.6 |  |  |  | 7.5 | 10.6733
-31.5 |  |  |  | 7.5333 | 10.7667
-31.4 |  |  |  | 7.6 | 10.8857
-31.3 |  |  |  | 7.6571 | 10.975
-31.2 |  |  |  | 7.7 | 11.1556
-31.1 |  |  |  | 7.8286 | 11.2
-31 |  |  |  | 7.9 | 11.3375
-30.9 |  |  |  | 8 | 11.4625
-30.8 |  |  |  | 8.0625 | 11.5875
-30.7 |  |  |  | 8.1 | 11.625
-30.6 |  |  |  | 8.26 | 11.7571
-30.5 |  |  |  | 8.3 | 11.8875
-30.4 |  |  |  | 8.3875 | 12.05
-30.3 |  |  |  | 8.48 | 12.2571
-30.2 |  |  |  | 8.58 | 12.3
-30.1 |  |  |  | 8.6333 | 12.425
-30 |  |  |  | 8.7857 | 12.6143
-29.9 |  |  |  | 8.8 | 12.875
-29.8 |  |  |  | 8.9 | 12.925
-29.7 |  |  |  | 8.9375 | 13.0571
-29.6 |  |  | 6.2 | 9.0714 | 13.1875
-29.5 |  |  | 6.3 | 9.2 | 13.4143
-29.4 |  |  | 6.3375 | 9.2286 | 13.6
-29.3 |  |  | 6.4 | 9.3625 | 13.7
-29.2 |  |  | 6.5 | 9.4875 | 13.825
-29.1 |  |  | 6.5 | 9.525 | 14.0429
-29 |  |  | 6.5385 | 9.6571 | 14.175
-28.9 |  |  | 6.6133 | 9.7875 | 14.42
-28.8 |  |  | 6.68 | 9.925 | 14.62
-28.7 |  |  | 6.7438 | 10 | 14.7857
-28.6 |  |  | 6.8111 | 10.0714 | 15.08
-28.5 |  |  | 6.9 | 10.2133 | 15.1667
-28.4 |  |  | 6.9 | 10.3467 | 15.44
-28.3 |  | 5.9 | 6.9778 | 10.44 | 15.6167
-28.2 |  | 5.925 | 7.0125 | 10.5154 | 15.8
-28.1 |  | 6 | 7.1 | 10.6692 | 16.1
-28 |  | 6 | 7.175 | 10.78 | 16.2833
-27.9 |  | 6.0818 | 7.2125 | 10.9333 | 16.56
-27.8 |  | 6.1 | 7.3429 | 11.08 | 16.76
-27.7 |  | 6.2 | 7.4 | 11.2083 | 17.06
-27.6 |  | 6.2286 | 7.5 | 11.2917 | 17.36
-27.5 |  | 6.3 | 7.5 | 11.425 | 17.62
-27.4 |  | 6.3 | 7.5625 | 11.5571 | 17.96
-27.3 | 6.1 | 6.3714 | 7.6 | 11.775 | 18.26
-27.2 | 6.1909 | 6.4 | 7.72 | 11.9133 | 18.66
-27.1 | 6.2077 | 6.4375 | 7.825 | 12.0467 | 19.08
-27 | 6.2846 | 6.5 | 7.9 | 12.25 | 19.56
-26.9 | 6.3 | 6.6 | 7.9214 | 12.4571 | 20.14
-26.8 | 6.32 | 6.6 | 7.9929 | 12.675 | 20.85
-26.7 | 6.3867 | 6.6625 | 8.075 | 12.8143 | 21.8
-26.6 | 6.4 | 6.7 | 8.1778 | 13.0143 | 22.98
-26.5 | 6.4214 | 6.8 | 8.2667 | 13.1875 | 24.6
-26.4 | 6.4929 | 6.8 | 8.3 | 13.4143 | 26.78
-26.3 | 6.5375 | 6.84 | 8.4 | 13.6 | 29.6286
-26.2 | 6.6 | 6.9 | 8.45 | 13.85 | 32.95
-26.1 | 6.6 | 6.9571 | 8.575 | 14.025 | 37
-26 | 6.7 | 7.04 | 8.6125 | 14.2429 | 41.35
-25.9 | 6.7 | 7.1 | 8.7429 | 14.525 | 46.0714
-25.8 | 6.7286 | 7.1429 | 8.8 | 14.725 | 50.7625
-25.7 | 6.8 | 7.2 | 8.9 | 15 | 55.4857
-25.6 | 6.82 | 7.3 | 8.9375 | 15.3 | 59.575
-25.5 | 6.9 | 7.3 | 9.0714 | 15.5857 | 62.8
-25.4 | 6.9 | 7.4 | 9.2 | 15.85 | 65.6286
-25.3 | 6.9571 | 7.4 | 9.3375 | 16.2429 | 67.75
-25.2 | 7 | 7.5 | 9.4 | 16.5 | 69.5286
-25.1 | 7.1 | 7.5 | 9.5 | 16.9875 | 70.7857
-25 | 7.1 | 7.6 | 9.6375 | 17.55 | 71.7
-24.9 | 7.1875 | 7.6 | 9.825 | 18.1 | 72.5286
-24.8 | 7.225 | 7.7 | 10 | 18.8286 | 73.2429
-24.7 | 7.3 | 7.725 | 10.1375 | 19.9625 | 73.75
-24.6 | 7.3 | 7.8 | 10.2625 | 21.2 | 74.3429
-24.5 | 7.3571 | 7.8875 | 10.4 | 23.0667 | 74.8429
-24.4 | 7.4 | 7.925 | 10.625 | 25.6 | 75.25
-24.3 | 7.4143 | 8 | 10.7571 | 28.8 | 75.4857
-24.2 | 7.5 | 8.0857 | 10.975 | 32.8 | 75.85
-24.1 | 7.5 | 8.1143 | 11.15 | 37.7 | 76.3286
-24 | 7.5143 | 8.25 | 11.3571 | 42.9857 | 76.65
-23.9 | 7.6 | 8.3 | 11.575 | 48.6 | 77
-23.8 | 7.6 | 8.3857 | 11.75 | 53.84 | 77.275
-23.7 | 7.7 | 8.525 | 12.0143 | 59.15 | 77.4714
-23.6 | 7.7 | 8.65 | 12.275 | 64 | 77.7
-23.5 | 7.7857 | 8.7857 | 12.475 | 67.8 | 78.1125
-23.4 | 7.8143 | 8.8286 | 12.8 | 70.6286 | 78.3714
-23.3 | 7.9 | 9.05 | 13.0714 | 72.76 | 78.7
-23.2 | 7.975 | 9.1857 | 13.475 | 74.34 | 78.9125
-23.1 | 8.0125 | 9.3125 | 13.85 | 75.6 | 79.2429
-23 | 8.1429 | 9.4857 | 14.4286 | 76.56 | 79.6
-22.9 | 8.2857 | 9.6 | 14.9857 | 77.38 | 79.775
-22.8 | 8.3125 | 9.8125 | 15.85 | 78 | 80.0429
-22.7 | 8.4429 | 9.9857 | 16.95 | 78.6125 | 80.2
-22.6 | 8.575 | 10.175 | 18.5125 | 79.0857 | 80.475
-22.5 | 8.7125 | 10.4333 | 20.7714 | 79.5 | 80.725
-22.4 | 8.8375 | 10.6 | 23.675 | 79.9143 | 81
-22.3 | 8.9714 | 10.9 | 27.4625 | 80.325 | 81.1571
-22.2 | 9.1 | 11.16 | 32.3286 | 80.7 | 81.425
-22.1 | 9.2375 | 11.5333 | 38.1 | 80.9571 | 81.5875
-22 | 9.3714 | 12 | 44.675 | 81.225 | 81.8571
-21.9 | 9.5 | 12.6 | 51.5 | 81.6 | 82.1
-21.8 | 9.675 | 13.4167 | 58.2 | 81.9286 | 82.3
-21.7 | 9.925 | 14.52 | 64.425 | 82.1875 | 82.4571
-21.6 | 10.175 | 16 | 69.5571 | 82.475 | 82.6
-21.5 | 10.45 | 18.38 | 73.75 | 82.75 | 82.9
-21.4 | 10.85 | 21.7875 | 77.0375 | 83.0143 | 83.1
-21.3 | 11.425 | 26.0286 | 79.3 | 83.275 | 83.3
-21.2 | 12.2286 | 31.58 | 81 | 83.525 | 83.4
-21.1 | 13.3 | 38 | 82.2 | 83.7143 | 83.66
-21 | 14.9375 | 45.1286 | 83.0875 | 83.975 | 83.8
-20.9 | 17.25 | 52.8875 | 83.7571 | 84.2143 | 84.08
-20.8 | 20.6444 | 61 | 84.4 | 84.4143 | 84.2857
-20.7 | 25.4727 | 67.3857 | 84.9125 | 84.6286 | 84.475
-20.6 | 31.2636 | 73.225 | 85.2875 | 84.9143 | 84.6
-20.5 | 38.1 | 78 | 85.7 | 85.0875 | 84.775
-20.4 | 46.075 | 81.0857 | 85.9857 | 85.28 | 84.9714
-20.3 | 54.65 | 83.4714 | 86.3875 | 85.56 | 85.1
-20.2 | 62.8 | 84.9833 | 86.7 | 85.7286 | 85.225
-20.1 | 70.1 | 86 | 86.9571 | 85.98 | 85.3571
-20 | 76.475 | 86.8 | 87.225 | 86.18 | 85.575
-19.9 | 81 | 87.55 | 87.5 | 86.38 | 85.7364
-19.8 | 84.2125 | 88.0333 | 87.8286 | 86.7167 | 85.9188
-19.7 | 86.39 | 88.45 | 88.0875 | 86.8167 | 86.1063
-19.6 | 87.925 | 88.8625 | 88.3 | 86.9 | 86.3
-19.5 | 88.9429 | 89.25 | 88.45 | 87.0143 | 86.5
-19.4 | 89.62 | 89.5714 | 88.6571 | 87.2 | 86.6333
-19.3 | 90.15 | 89.875 | 88.8 | 87.4714 | 86.7714
-19.2 | 90.6267 | 90.05 | 89.0286 | 87.6125 | 86.9182
-19.1 | 90.8933 | 90.1875 | 89.3 | 87.7333 | 87.1
-19 | 91.1636 | 90.375 | 89.4 | 87.8667 | 87.1909
-18.9 | 91.4154 | 90.7143 | 89.6143 | 88.05 | 87.3636
-18.8 | 91.6545 | 90.8875 | 89.75 | 88.2571 | 87.5111
-18.7 | 91.9125 | 91.0143 | 89.9714 | 88.475 | 87.6286
-18.6 | 92.0375 | 91.2 | 90.2143 | 88.6 | 87.7625
-18.5 | 92.2154 | 91.3 | 90.4 | 88.7125 | 87.9
-18.4 | 92.4333 | 91.4286 | 90.5 | 88.8 | 88.0286
-18.3 | 92.6 | 91.65 | 90.6286 | 88.95 | 88.2
-18.2 | 92.7125 | 91.775 | 90.85 | 89.1125 | 88.3333
-18.1 | 92.8429 | 91.9125 | 90.9857 | 89.2429 | 88.5286
-18 | 92.975 | 92.0429 | 91.1143 | 89.45 | 88.6625
-17.9 | 93.0125 | 92.175 | 91.2 | 89.6125 | 88.8
-17.8 | 93.1857 | 92.325 | 91.275 | 89.7429 | 88.88
-17.7 | 93.3 | 92.5429 | 91.4 | 89.875 | 89.0571
-17.6 | 93.4125 | 92.6 | 91.6 | 90 | 89.2
-17.5 | 93.5375 | 92.7143 | 91.75 | 90.0571 | 89.3
-17.4 | 93.7429 | 92.8 | 91.8125 | 90.2625 | 89.4
-17.3 | 93.8 | 92.9 | 91.9429 | 90.3875 | 89.6
-17.2 | 93.9 | 93.0375 | 92.075 | 90.625 | 89.7
-17.1 | 93.9625 | 93.1714 | 92.2 | 90.7 | 89.8
-17 | 94 | 93.3 | 92.3286 | 90.7 | 89.8571
-16.9 | 94.1286 | 93.4375 | 92.4625 | 90.825 | 89.9875
-16.8 | 94.2625 | 93.5714 | 92.5875 | 90.9 | 90.05
-16.7 | 94.3 | 93.6 | 92.7 | 91.05 | 90.3
-16.6 | 94.4 | 93.625 | 92.7556 | 91.2125 | 90.4857
-16.5 | 94.4 | 93.75 | 92.8 | 91.3429 | 90.5125
-16.4 | 94.5 | 93.9 | 93.0143 | 91.475 | 90.6429
-16.3 | 94.6667 | 94.0308 | 93.15 | 91.5889 | 90.7857
-16.2 | 94.76 | 94.1062 | 93.2667 | 91.72 | 90.8429
-16.1 | 94.8 | 94.1688 | 93.3538 | 91.8533 | 90.975
-16 | 94.84 | 94.2385 | 93.4333 | 91.9867 | 91.1
-15.9 | 95.04 | 94.3222 | 93.5125 | 92.06 | 91.14
-15.8 | 95.1 | 94.4 | 93.575 | 92.125 | 91.2
-15.7 | 95.14 | 94.4571 | 93.68 | 92.1875 | 91.3
-15.6 | 95.2286 | 94.5875 | 93.8133 | 92.3067 | 91.375
-15.5 | 95.3 | 94.625 | 93.9467 | 92.44 | 91.5714
-15.4 | 95.3571 | 94.7571 | 94.04 | 92.5733 | 91.7
-15.3 | 95.4875 | 94.8875 | 94.1067 | 92.6533 | 91.8286
-15.2 | 95.625 | 94.9125 | 94.1733 | 92.72 | 91.9625
-15.1 | 95.7 | 95 | 94.24 | 92.7867 | 92.1
-15 | 95.7875 | 95.075 | 94.31 | 92.9067 | 92.1286
-14.9 | 95.825 | 95.1 | 94.4143 | 93.0333 | 92.2
-14.8 | 95.9 | 95.2083 | 94.5 | 93.15 | 92.3
-14.7 | 95.975 | 95.2917 | 94.575 | 93.2 | 92.45
-14.6 | 96 | 95.4286 | 94.7 | 93.31 | 92.6
-14.5 | 96.0429 | 95.5 | 94.7429 | 93.4125 | 92.7
-14.4 | 96.175 | 95.6 | 94.875 | 93.5 | 92.75
-14.3 | 96.2125 | 95.7 | 95 | 93.575 | 92.8857
-14.2 | 96.3 | 95.7375 | 95.0429 | 93.7 | 93
-14.1 | 96.3857 | 95.8 | 95.1 | 93.7667 | 93.1
-14 | 96.4 | 95.8 | 95.2 | 93.85 | 93.1833
-13.9 | 96.4429 | 95.9 | 95.2375 | 93.92 | 93.24
-13.8 | 96.5 | 95.925 | 95.3 | 94.02 | 93.3
-13.7 | 96.6 | 96.05 | 95.3 | 94.1 | 93.4
-13.6 | 96.6429 | 96.1 | 95.4 | 94.1286 | 93.5714
-13.5 | 96.7 | 96.2 | 95.5 | 94.2833 | 93.7
-13.4 | 96.8 | 96.2125 | 95.6 | 94.38 | 93.75
-13.3 | 96.8375 | 96.3 | 95.7 | 94.48 | 93.8
-13.2 | 96.9 | 96.3125 | 95.8 | 94.6 | 93.8143
-13.1 | 96.9 | 96.4 | 95.8 | 94.7 | 93.95
-13 | 97 | 96.5 | 95.9 | 94.7 | 94
-12.9 | 97.0625 | 96.56 | 95.96 | 94.8 | 94.06
-12.8 | 97.1 | 96.6 | 96.06 | 94.8375 | 94.16
-12.7 | 97.2 | 96.7 | 96.1 | 94.9625 | 94.3
-12.6 | 97.2625 | 96.7333 | 96.2 | 95.1 | 94.3833
-12.5 | 97.4 | 96.8 | 96.2 | 95.2 | 94.5
-12.4 | 97.4 | 96.8714 | 96.2833 | 95.2625 | 94.5625
-12.3 | 97.4143 | 96.9 | 96.3 | 95.3 | 94.6875
-12.2 | 97.5 | 96.9 | 96.3 | 95.325 | 94.7
-12.1 | 97.5 | 97 | 96.4286 | 95.4571 | 94.7833
-12 | 97.6071 | 97.1 | 96.5 | 95.5 | 94.8143
-11.9 | 97.6786 | 97.1 | 96.6 | 95.6 | 94.9
-11.8 | 97.7467 | 97.12 | 96.625 | 95.6667 | 94.975
-11.7 | 97.8 | 97.2 | 96.7 | 95.725 | 95.1
-11.6 | 97.8 | 97.2833 | 96.7 | 95.8571 | 95.2
-11.5 | 97.8 | 97.3 | 96.8 | 95.9 | 95.2714
-11.4 | 97.8583 | 97.4 | 96.85 | 96.0125 | 95.4
-11.3 | 97.9625 | 97.4 | 96.9 | 96.1 | 95.4375
-11.2 | 98.0583 | 97.48 | 97 | 96.175 | 95.5714
-11.1 | 98.1 | 97.5 | 97.0857 | 96.2 | 95.6
-11 | 98.1 | 97.5 | 97.1333 | 96.3 | 95.7
-10.9 | 98.1375 | 97.6 | 97.2 | 96.3714 | 95.7571
-10.8 | 98.2714 | 97.7 | 97.2 | 96.4143 | 95.8
-10.7 | 98.3 | 97.7286 | 97.2857 | 96.5 | 95.875
-10.6 | 98.3857 | 97.8 | 97.3 | 96.6 | 95.9875
-10.5 | 98.4 | 97.8 | 97.4 | 96.6 | 96.1
-10.4 | 98.4 | 97.9 | 97.4833 | 96.7 | 96.1364
-10.3 | 98.4 | 97.9 | 97.6 | 96.7375 | 96.2231
-10.2 | 98.4571 | 98 | 97.6 | 96.8 | 96.3
-10.1 | 98.5 | 98.1 | 97.675 | 96.8 | 96.3667
-10 | 98.6 | 98.1 | 97.7 | 96.9 | 96.4313
-9.9 | 98.6 | 98.1143 | 97.7 | 96.9714 | 96.4938
-9.8 | 98.6875 | 98.2 | 97.775 | 97 | 96.5
-9.7 | 98.7 | 98.2 | 97.8 | 97.1 | 96.5533
-9.6 | 98.7 | 98.3 | 97.9 | 97.1625 | 96.6867
-9.5 | 98.7875 | 98.3 | 97.9 | 97.2 | 96.76
-9.4 | 98.8 | 98.3 | 98 | 97.3 | 96.8267
-9.3 | 98.8 | 98.4125 | 98.0286 | 97.3571 | 96.8933
-9.2 | 98.9 | 98.5 | 98.1 | 97.4 | 96.9
-9.1 | 98.9 | 98.5 | 98.1 | 97.4286 | 96.9333
-9 | 98.9 | 98.5125 | 98.2 | 97.5 | 97
-8.9 | 99 | 98.6 | 98.2 | 97.5571 | 97.05
-8.8 | 99.04 | 98.6 | 98.2 | 97.7 | 97.1857
-8.7 | 99.1 | 98.7 | 98.2286 | 97.7833 | 97.3143
-8.6 | 99.1 | 98.7 | 98.3 | 97.8 | 97.4
-8.5 | 99.1 | 98.8 | 98.4 | 97.82 | 97.4
-8.4 | 99.14 | 98.8 | 98.4333 | 97.9 | 97.4125
-8.3 | 99.2 | 98.8 | 98.5 | 97.9 | 97.5
-8.2 | 99.2 | 98.9 | 98.5286 | 97.9 | 97.575
-8.1 | 99.2 | 98.9 | 98.6 | 97.9833 | 97.7
-8 | 99.2 | 98.9 | 98.6 | 98 | 97.7667
-7.9 | 99.3 | 98.9 | 98.7 | 98 | 97.8125
-7.8 | 99.3 | 99 | 98.7 | 98.1143 | 97.9
-7.7 | 99.3 | 99 | 98.7 | 98.2 | 98
-7.6 | 99.32 | 99.1 | 98.7 | 98.3 | 98
-7.5 | 99.4 | 99.1 | 98.75 | 98.3 | 98
-7.4 | 99.4625 | 99.1714 | 98.8 | 98.375 | 98.1
-7.3 | 99.5 | 99.2 | 98.8125 | 98.4 | 98.15
-7.2 | 99.5 | 99.2 | 98.9 | 98.4429 | 98.2
-7.1 | 99.5 | 99.3 | 98.9714 | 98.5 | 98.3
-7 | 99.5 | 99.3 | 99 | 98.5267 | 98.3
-6.9 | 99.6 | 99.3 | 99.03 | 98.5933 | 98.3857
-6.8 | 99.6 | 99.3 | 99.1 | 98.66 | 98.4
-6.7 | 99.6 | 99.3 | 99.1 | 98.7267 | 98.4
-6.6 | 99.7 | 99.4 | 99.1 | 98.7933 | 98.4
-6.5 | 99.7 | 99.4 | 99.1 | 98.8 | 98.5
-6.4 | 99.7 | 99.4 | 99.16 | 98.8 | 98.5
-6.3 | 99.7 | 99.4444 | 99.2 | 98.85 | 98.5
-6.2 | 99.7 | 99.5556 | 99.2071 | 98.9 | 98.6
-6.1 | 99.775 | 99.6 | 99.2786 | 99 | 98.6
-6 | 99.8 | 99.6 | 99.3 | 99 | 98.7
-5.9 | 99.8 | 99.6429 | 99.3 | 99 | 98.78
-5.8 | 99.8 | 99.7 | 99.4 | 99 | 98.8
-5.7 | 99.9 | 99.7 | 99.4 | 99.0286 | 98.9
-5.6 | 99.9 | 99.7 | 99.4 | 99.1 | 98.9667
-5.5 | 99.9 | 99.7 | 99.4545 | 99.1 | 99
-5.4 | 99.9 | 99.7 | 99.5 | 99.2 | 99.1
-5.3 | 99.9 | 99.7 | 99.575 | 99.2625 | 99.1
-5.2 | 99.9 | 99.7 | 99.6 | 99.3 | 99.1
-5.1 | 100 | 99.7667 | 99.6 | 99.3 | 99.2
-5 | 100 | 99.8 | 99.6 | 99.3 | 99.2
-4.9 | 100 | 99.8 | 99.6 | 99.3 | 99.2
-4.8 | 100 | 99.8625 | 99.6 | 99.3154 | 99.3
-4.7 | 100.06 | 99.9 | 99.6571 | 99.3923 | 99.3571
-4.6 | 100.06 | 99.9 | 99.7 | 99.475 | 99.4
-4.5 | 100.1 | 99.9 | 99.7 | 99.5 | 99.4
-4.4 | 100.1 | 99.9 | 99.7 | 99.5 | 99.4
-4.3 | 100.1 | 99.9 | 99.7 | 99.58 | 99.4
-4.2 | 100.1 | 99.9 | 99.7 | 99.6 | 99.5
-4.1 | 100.1 | 99.9 | 99.7667 | 99.6 | 99.5667
-4 | 100.06 | 100 | 99.8 | 99.6 | 99.6
-3.9 | 100 | 100 | 99.8 | 99.6 | 99.6429
-3.8 | 100 | 100 | 99.8 | 99.64 | 99.7
-3.7 | 100 | 100 | 99.8 | 99.7 | 99.7
-3.6 | 100 | 100 | 99.8 | 99.7 | 99.7
-3.5 | 100.1 | 100 | 99.8 | 99.8 | 99.7
-3.4 | 100.1 | 100 | 99.9 | 99.8 | 99.7
-3.3 | 100.1 | 100 | 99.9 | 99.8 | 99.7
-3.2 | 100.1 | 100 | 99.9 | 99.8 | 99.7
-3.1 | 100.1 | 100 | 99.9 | 99.8 | 99.7
-3 | 100.1 | 100 | 99.9 | 99.8 | 99.775
-2.9 | 100.1 | 100 | 99.9 | 99.8 | 99.8
-2.8 | 100.1 | 100 | 99.9 | 99.8 | 99.8
-2.7 | 100.1 | 100 | 99.9 | 99.8 | 99.9
-2.6 | 100.1 | 100 | 99.9 | 99.88 | 99.9667
-2.5 | 100.1 | 100 | 99.9125 | 99.8 | 100
-2.4 | 100.1 | 100 | 100 | 99.8 | 100
-2.3 | 100.1 | 100 | 100 | 99.8857 | 100
-2.2 | 100.1 | 100 | 100 | 99.9 | 99.9
-2.1 | 100.1 | 100 | 100 | 99.9 | 99.9
-2 | 100.1 | 100 | 100 | 99.9143 | 99.9
-1.9 | 100.1 | 100 | 100 | 100 | 100
-1.8 | 100.1 | 100 | 100 | 100 | 100
-1.7 | 100.1 | 100.0714 | 100 | 100 | 100.1
-1.6 | 100.1 | 100.1 | 100 | 100 | 100.0625
-1.5 | 100.1 | 100.1 | 100 | 100 | 100
-1.4 | 100.1 | 100.1 | 100 | 100 | 100
-1.3 | 100.1 | 100.1 | 100 | 100 | 100
-1.2 | 100.1 | 100.1 | 100 | 100 | 100
-1.1 | 100.0533 | 100.04 | 100 | 100 | 100
-1 | 100 | 100 | 100 | 100 | 100
-0.9 | 100 | 100 | 100 | 100 | 100
-0.8 | 100 | 100 | 100 | 100 | 99.9
-0.7 | 100 | 100 | 100 | 100 | 100
-0.6 | 100 | 100 | 100 | 100 | 99.9
-0.5 | 100 | 100 | 100 | 100 | 99.9
-0.4 | 100 | 100 | 100 | 100.0625 | 99.98
-0.3 | 100 | 100 | 100 | 100.1 | 100
-0.2 | 100 | 100 | 100 | 100 | 100
-0.1 | 100 | 100 | 100 | 100 | 100
0 | 100 | 100 | 100 | 100 | 100
0.1 | 100 | 100 | 100 | 100 | 100
0.2 | 100 | 100 | 100 | 100 | 100
0.3 | 100 | 100 | 100 | 100.1 | 100
0.4 | 100 | 100 | 100 | 100.0625 | 99.98
0.5 | 100 | 100 | 100 | 100 | 99.9
0.6 | 100 | 100 | 100 | 100 | 99.9
0.7 | 100 | 100 | 100 | 100 | 100
0.8 | 100 | 100 | 100 | 100 | 99.9
0.9 | 100 | 100 | 100 | 100 | 100
1 | 100.05 | 100 | 100 | 100 | 100
1.1 | 100.1 | 100 | 100 | 100 | 100
1.2 | 100.1 | 100.1 | 100 | 100 | 100
1.3 | 100.1 | 100.1 | 100 | 100 | 100
1.4 | 100.1 | 100.1 | 100 | 100 | 100
1.5 | 100.1 | 100.1 | 100 | 100 | 100
1.6 | 100.1 | 100.1 | 100 | 100 | 100.0625
1.7 | 100.1 | 100.1 | 100 | 100 | 100.1
1.8 | 100.1 | 100 | 100 | 100 | 100
1.9 | 100.1 | 100 | 100 | 100 | 100
2 | 100.1 | 100 | 100 | 99.9143 | 99.9
2.1 | 100.1 | 100 | 100 | 99.9 | 99.9
2.2 | 100.1 | 100 | 100 | 99.9 | 99.9
2.3 | 100.1 | 100 | 100 | 99.8857 | 100
2.4 | 100.1 | 100 | 100 | 99.8 | 100
2.5 | 100.1 | 100 | 99.9429 | 99.8 | 100
2.6 | 100.1 | 100 | 99.9 | 99.88 | 99.9667
2.7 | 100.1 | 100 | 99.9 | 99.8 | 99.9
2.8 | 100.1 | 100 | 99.9 | 99.8 | 99.8
2.9 | 100.1 | 100 | 99.9 | 99.8 | 99.8
3 | 100.1 | 100 | 99.9 | 99.8 | 99.775
3.1 | 100.1 | 100 | 99.9 | 99.8 | 99.7
3.2 | 100.1 | 100 | 99.9 | 99.8 | 99.7
3.3 | 100.1 | 100 | 99.9 | 99.8 | 99.7
3.4 | 100.0571 | 100 | 99.9 | 99.8 | 99.7
3.5 | 100.075 | 100 | 99.8444 | 99.8 | 99.7
3.6 | 100 | 100 | 99.8 | 99.7 | 99.7
3.7 | 100 | 100 | 99.8 | 99.7 | 99.7
3.8 | 100 | 100 | 99.8 | 99.64 | 99.7
3.9 | 100 | 100 | 99.8 | 99.6 | 99.6429
4 | 100.02 | 99.975 | 99.8 | 99.6 | 99.6
4.1 | 100.1 | 99.9 | 99.8 | 99.6 | 99.5667
4.2 | 100.1 | 99.9 | 99.7333 | 99.6 | 99.5
4.3 | 100.1 | 99.9 | 99.7 | 99.58 | 99.4
4.4 | 100.1 | 99.9 | 99.7 | 99.5 | 99.4
4.5 | 100.1 | 99.9 | 99.7 | 99.5 | 99.4
4.6 | 100.1 | 99.9 | 99.7 | 99.475 | 99.4
4.7 | 100.1 | 99.9 | 99.7 | 99.3923 | 99.3571
4.8 | 100 | 99.8286 | 99.6 | 99.3154 | 99.3
4.9 | 100 | 99.8 | 99.6 | 99.3 | 99.2
5 | 100 | 99.8 | 99.6 | 99.3 | 99.2
5.1 | 100 | 99.7375 | 99.6 | 99.3 | 99.2
5.2 | 99.9 | 99.7 | 99.6 | 99.3 | 99.1
5.3 | 99.9 | 99.7 | 99.5125 | 99.2625 | 99.1
5.4 | 99.9 | 99.7 | 99.4833 | 99.2 | 99.1
5.5 | 99.9 | 99.7 | 99.4 | 99.1 | 99
5.6 | 99.9 | 99.7 | 99.4 | 99.1 | 98.9667
5.7 | 99.9 | 99.7 | 99.4 | 99.0286 | 98.9
5.8 | 99.8125 | 99.7 | 99.4 | 99 | 98.8
5.9 | 99.8 | 99.7 | 99.3333 | 99 | 98.78
6 | 99.8 | 99.6 | 99.3 | 99 | 98.7
6.1 | 99.8 | 99.6 | 99.2111 | 99 | 98.6
6.2 | 99.7 | 99.525 | 99.2 | 98.9 | 98.6
6.3 | 99.7 | 99.5 | 99.2 | 98.85 | 98.5
6.4 | 99.7 | 99.4 | 99.1909 | 98.8 | 98.5
6.5 | 99.7 | 99.4 | 99.1 | 98.8 | 98.5
6.6 | 99.7 | 99.3875 | 99.1 | 98.7933 | 98.4
6.7 | 99.625 | 99.3 | 99.1 | 98.7267 | 98.4
6.8 | 99.6 | 99.3 | 99.0778 | 98.66 | 98.4
6.9 | 99.6 | 99.3 | 99 | 98.5933 | 98.3857
7 | 99.5 | 99.3 | 99 | 98.5267 | 98.3
7.1 | 99.5 | 99.3 | 99 | 98.5 | 98.3
7.2 | 99.5 | 99.2 | 98.9 | 98.4429 | 98.2
7.3 | 99.5 | 99.2 | 98.85 | 98.4 | 98.15
7.4 | 99.5 | 99.2 | 98.8 | 98.375 | 98.1
7.5 | 99.4 | 99.1 | 98.7857 | 98.3 | 98
7.6 | 99.3 | 99.1 | 98.7 | 98.3 | 98
7.7 | 99.3 | 99 | 98.7 | 98.2 | 98
7.8 | 99.3 | 99 | 98.7 | 98.1143 | 97.9
7.9 | 99.3 | 98.9 | 98.65 | 98 | 97.8125
8 | 99.2 | 98.9 | 98.6 | 98 | 97.7667
8.1 | 99.2 | 98.9 | 98.6 | 97.9833 | 97.7
8.2 | 99.2 | 98.9 | 98.5571 | 97.9 | 97.575
8.3 | 99.2 | 98.84 | 98.5 | 97.9 | 97.5
8.4 | 99.1 | 98.8 | 98.4 | 97.9 | 97.4125
8.5 | 99.1 | 98.8 | 98.3286 | 97.82 | 97.4
8.6 | 99.1 | 98.7375 | 98.3 | 97.8 | 97.4
8.7 | 99.1 | 98.6833 | 98.2625 | 97.7833 | 97.3143
8.8 | 99 | 98.6 | 98.2 | 97.7 | 97.1857
8.9 | 99 | 98.6 | 98.2 | 97.5571 | 97.05
9 | 98.9143 | 98.5 | 98.2 | 97.5 | 97
9.1 | 98.9 | 98.5 | 98.1286 | 97.4286 | 96.9333
9.2 | 98.9 | 98.5 | 98.1 | 97.4 | 96.9
9.3 | 98.8 | 98.4 | 98 | 97.3571 | 96.8933
9.4 | 98.8 | 98.3429 | 98 | 97.3 | 96.8267
9.5 | 98.7333 | 98.3 | 97.9 | 97.2 | 96.76
9.6 | 98.7 | 98.2857 | 97.9 | 97.1625 | 96.6867
9.7 | 98.7 | 98.2 | 97.8375 | 97.1 | 96.5533
9.8 | 98.625 | 98.2 | 97.8 | 97 | 96.5
9.9 | 98.6 | 98.1 | 97.7 | 96.9714 | 96.4938
10 | 98.5571 | 98.1 | 97.7 | 96.9 | 96.4313
10.1 | 98.5 | 98.1 | 97.6125 | 96.8 | 96.3667
10.2 | 98.4875 | 98.06 | 97.6 | 96.8 | 96.3
10.3 | 98.4 | 97.9 | 97.55 | 96.7375 | 96.2231
10.4 | 98.4 | 97.8875 | 97.5 | 96.7 | 96.1364
10.5 | 98.4 | 97.8 | 97.4 | 96.6 | 96.1
10.6 | 98.3286 | 97.8 | 97.3143 | 96.6 | 95.9875
10.7 | 98.3 | 97.8 | 97.3 | 96.5 | 95.875
10.8 | 98.2 | 97.7 | 97.2 | 96.4143 | 95.8
10.9 | 98.1714 | 97.6 | 97.2 | 96.3714 | 95.7571
11 | 98.1 | 97.54 | 97.1 | 96.3 | 95.7
11.1 | 98.1 | 97.5 | 97.0143 | 96.2 | 95.6
11.2 | 98 | 97.44 | 96.9833 | 96.175 | 95.5714
11.3 | 98 | 97.4 | 96.9 | 96.1 | 95.4375
11.4 | 97.8857 | 97.4 | 96.8857 | 96.0125 | 95.4
11.5 | 97.8143 | 97.3 | 96.8 | 95.9 | 95.2714
11.6 | 97.8 | 97.25 | 96.725 | 95.8571 | 95.2
11.7 | 97.7583 | 97.2 | 96.6875 | 95.725 | 95.1
11.8 | 97.68 | 97.1 | 96.6 | 95.6667 | 94.975
11.9 | 97.6133 | 97.1 | 96.6 | 95.6 | 94.9
12 | 97.6 | 97.1 | 96.5 | 95.5 | 94.8143
12.1 | 97.5143 | 97 | 96.4 | 95.4571 | 94.7833
12.2 | 97.5 | 96.9167 | 96.3286 | 95.325 | 94.7
12.3 | 97.45 | 96.9 | 96.3 | 95.3 | 94.6875
12.4 | 97.4 | 96.8429 | 96.22 | 95.2625 | 94.5625
12.5 | 97.3222 | 96.8 | 96.2 | 95.2 | 94.5
12.6 | 97.3 | 96.7 | 96.2 | 95.1 | 94.3833
12.7 | 97.2 | 96.7 | 96.1 | 94.9625 | 94.3
12.8 | 97.1286 | 96.6 | 96.1 | 94.8375 | 94.16
12.9 | 97.1 | 96.6 | 96 | 94.8 | 94.06
13 | 97 | 96.5 | 95.9 | 94.7 | 94
13.1 | 96.9286 | 96.4 | 95.8 | 94.7 | 93.95
13.2 | 96.9 | 96.3714 | 95.8 | 94.6 | 93.8143
13.3 | 96.8714 | 96.3 | 95.7 | 94.48 | 93.8
13.4 | 96.8 | 96.275 | 95.6375 | 94.38 | 93.75
13.5 | 96.7125 | 96.2 | 95.4857 | 94.2833 | 93.7
13.6 | 96.675 | 96.1 | 95.4 | 94.1286 | 93.5714
13.7 | 96.6 | 96.0143 | 95.3125 | 94.1 | 93.4
13.8 | 96.5125 | 95.9 | 95.28 | 94.02 | 93.3
13.9 | 96.475 | 95.9 | 95.2 | 93.92 | 93.24
14 | 96.4 | 95.8625 | 95.2 | 93.85 | 93.1833
14.1 | 96.4 | 95.8 | 95.1125 | 93.7667 | 93.1
14.2 | 96.3 | 95.7 | 95.075 | 93.7 | 93
14.3 | 96.2429 | 95.7 | 95 | 93.575 | 92.8857
14.4 | 96.1125 | 95.5625 | 94.9125 | 93.5 | 92.75
14.5 | 96.075 | 95.5 | 94.775 | 93.4125 | 92.7
14.6 | 96 | 95.4333 | 94.7 | 93.31 | 92.6
14.7 | 96 | 95.3571 | 94.6143 | 93.2 | 92.45
14.8 | 95.8875 | 95.25 | 94.4857 | 93.15 | 92.3
14.9 | 95.8 | 95.1 | 94.4 | 93.0333 | 92.2
15 | 95.725 | 95.0375 | 94.3 | 92.9067 | 92.1286
15.1 | 95.6875 | 95 | 94.1733 | 92.7867 | 92.1
15.2 | 95.5571 | 94.9778 | 94.1067 | 92.72 | 91.9625
15.3 | 95.5 | 94.8571 | 94.04 | 92.6533 | 91.8286
15.4 | 95.3833 | 94.725 | 93.9733 | 92.5733 | 91.7
15.5 | 95.3 | 94.6875 | 93.9067 | 92.44 | 91.5714
15.6 | 95.2571 | 94.5571 | 93.78 | 92.3067 | 91.375
15.7 | 95.18 | 94.425 | 93.675 | 92.1875 | 91.3
15.8 | 95.1 | 94.4 | 93.6125 | 92.125 | 91.2
15.9 | 94.98 | 94.2833 | 93.4933 | 92.06 | 91.14
16 | 94.88 | 94.2 | 93.37 | 91.9867 | 91.1
16.1 | 94.8 | 94.1333 | 93.3 | 91.8533 | 90.975
16.2 | 94.7 | 94.1 | 93.3 | 91.72 | 90.8429
16.3 | 94.7 | 94.1 | 93.0857 | 91.5889 | 90.7857
16.4 | 94.5333 | 93.9182 | 92.95 | 91.475 | 90.6429
16.5 | 94.4 | 93.7364 | 92.8143 | 91.3429 | 90.5125
16.6 | 94.4 | 93.6875 | 92.7 | 91.2125 | 90.4857
16.7 | 94.325 | 93.6 | 92.6625 | 91.05 | 90.3
16.8 | 94.3 | 93.5375 | 92.6 | 90.9 | 90.05
16.9 | 94.1625 | 93.5 | 92.5 | 90.825 | 89.9875
17 | 94.0286 | 93.2714 | 92.3625 | 90.7 | 89.8571
17.1 | 94 | 93.1375 | 92.2375 | 90.7 | 89.8
17.2 | 93.8625 | 93 | 92.1 | 90.625 | 89.7
17.3 | 93.8 | 92.9714 | 91.9714 | 90.3875 | 89.6
17.4 | 93.7 | 92.8286 | 91.8375 | 90.2625 | 89.4
17.5 | 93.575 | 92.675 | 91.7125 | 90.0571 | 89.3
17.6 | 93.4429 | 92.6 | 91.5714 | 90 | 89.2
17.7 | 93.3125 | 92.5125 | 91.4375 | 89.875 | 89.0571
17.8 | 93.175 | 92.375 | 91.3125 | 89.7429 | 88.88
17.9 | 93.0429 | 92.2429 | 91.1857 | 89.6125 | 88.8
18 | 92.9125 | 92.025 | 91.1 | 89.45 | 88.6625
18.1 | 92.85 | 91.875 | 91.0143 | 89.2429 | 88.5286
18.2 | 92.7 | 91.7429 | 90.7857 | 89.1125 | 88.3333
18.3 | 92.5833 | 91.6125 | 90.65 | 88.95 | 88.2
18.4 | 92.4 | 91.4857 | 90.5143 | 88.8 | 88.0286
18.5 | 92.2062 | 91.35 | 90.3857 | 88.7125 | 87.9
18.6 | 92.0187 | 91.2143 | 90.2 | 88.6 | 87.7625
18.7 | 91.8714 | 91.0857 | 89.9286 | 88.475 | 87.6286
18.8 | 91.6636 | 90.95 | 89.7714 | 88.2571 | 87.5111
18.9 | 91.4385 | 90.7286 | 89.55 | 88.05 | 87.3636
19 | 91.2583 | 90.375 | 89.4143 | 87.8667 | 87.1909
19.1 | 90.8417 | 90.1571 | 89.2857 | 87.7333 | 87.1
19.2 | 90.5923 | 90.025 | 89.05 | 87.6125 | 86.9182
19.3 | 90.2 | 89.8143 | 88.825 | 87.4714 | 86.7714
19.4 | 89.6556 | 89.55 | 88.675 | 87.2 | 86.6333
19.5 | 88.9125 | 89.2625 | 88.4625 | 87.0143 | 86.5
19.6 | 87.9143 | 88.8714 | 88.3286 | 86.9 | 86.3
19.7 | 86.35 | 88.4 | 88.1 | 86.8167 | 86.1063
19.8 | 84.24 | 88.05 | 87.825 | 86.7167 | 85.9188
19.9 | 80.9875 | 87.5143 | 87.5571 | 86.38 | 85.7364
20 | 76.3857 | 86.9 | 87.3 | 86.18 | 85.575
20.1 | 70.125 | 85.96 | 86.925 | 85.98 | 85.3571
20.2 | 62.8125 | 85 | 86.6571 | 85.7286 | 85.225
20.3 | 54.6857 | 83.6 | 86.4 | 85.56 | 85.1
20.4 | 46.1 | 81.1667 | 85.9875 | 85.28 | 84.9714
20.5 | 37.88 | 78 | 85.6571 | 85.0875 | 84.775
20.6 | 30.8167 | 73.5 | 85.2 | 84.9143 | 84.6
20.7 | 24.9833 | 67.8 | 84.8857 | 84.6286 | 84.475
20.8 | 20.53 | 60.7 | 84.35 | 84.4143 | 84.2857
20.9 | 17.3 | 52.8429 | 83.7 | 84.2143 | 84.08
21 | 14.9286 | 45.4 | 83.05 | 83.975 | 83.8
21.1 | 13.3125 | 37.85 | 82.1429 | 83.7143 | 83.66
21.2 | 12.2 | 31.2667 | 81.0125 | 83.525 | 83.4
21.3 | 11.3889 | 26.0286 | 79.3 | 83.275 | 83.3
21.4 | 10.8636 | 21.7 | 77.0143 | 83.0143 | 83.1
21.5 | 10.44 | 18.56 | 73.8375 | 82.75 | 82.9
21.6 | 10.15 | 16.175 | 69.55 | 82.475 | 82.6
21.7 | 9.9 | 14.48 | 64.3571 | 82.1875 | 82.4571
21.8 | 9.6714 | 13.4 | 58.3125 | 81.9286 | 82.3
21.9 | 9.5375 | 12.58 | 51.45 | 81.6 | 82.1
22 | 9.3 | 11.9857 | 44.6 | 81.225 | 81.8571
22.1 | 9.2 | 11.5571 | 38.125 | 80.9571 | 81.5875
22.2 | 9.1375 | 11.18 | 32.4667 | 80.7 | 81.425
22.3 | 8.9 | 10.9333 | 27.7 | 80.325 | 81.1571
22.4 | 8.775 | 10.68 | 23.75 | 79.9143 | 81
22.5 | 8.6429 | 10.4 | 20.9286 | 79.5 | 80.725
22.6 | 8.6 | 10.1857 | 18.65 | 79.0857 | 80.475
22.7 | 8.475 | 10.0125 | 16.9714 | 78.6125 | 80.2
22.8 | 8.3429 | 9.8667 | 15.825 | 78 | 80.0429
22.9 | 8.2125 | 9.68 | 15.05 | 77.38 | 79.775
23 | 8.175 | 9.5143 | 14.3714 | 76.56 | 79.6
23.1 | 8.0429 | 9.2857 | 13.8625 | 75.6 | 79.2429
23.2 | 8 | 9.15 | 13.4714 | 74.34 | 78.9125
23.3 | 7.875 | 9.0143 | 13.0429 | 72.76 | 78.7
23.4 | 7.8 | 8.8857 | 12.775 | 70.6286 | 78.3714
23.5 | 7.7143 | 8.75 | 12.5143 | 67.8 | 78.1125
23.6 | 7.7 | 8.6143 | 12.25 | 64 | 77.7
23.7 | 7.7 | 8.4857 | 11.975 | 59.15 | 77.4714
23.8 | 7.6143 | 8.4 | 11.7571 | 53.84 | 77.275
23.9 | 7.6 | 8.3 | 11.55 | 48.6 | 77
24 | 7.55 | 8.2143 | 11.2875 | 42.9857 | 76.65
24.1 | 7.5 | 8.1 | 11.1571 | 37.7 | 76.3286
24.2 | 7.5 | 8.05 | 10.95 | 32.8 | 75.85
24.3 | 7.45 | 8 | 10.775 | 28.8 | 75.4857
24.4 | 7.4 | 7.9 | 10.5625 | 25.6 | 75.25
24.5 | 7.3 | 7.8571 | 10.4286 | 23.0667 | 74.8429
24.6 | 7.3 | 7.8 | 10.2 | 21.2 | 74.3429
24.7 | 7.3 | 7.7 | 10.0625 | 19.9625 | 73.75
24.8 | 7.2571 | 7.7 | 9.9375 | 18.8286 | 73.2429
24.9 | 7.2 | 7.6286 | 9.8 | 18.1 | 72.5286
25 | 7.1 | 7.6 | 9.6714 | 17.55 | 71.7
25.1 | 7.1 | 7.5 | 9.5375 | 16.9875 | 70.7857
25.2 | 7.025 | 7.4857 | 9.4 | 16.5 | 69.5286
25.3 | 6.9875 | 7.4 | 9.3 | 16.2429 | 67.75
25.4 | 6.9 | 7.3857 | 9.2375 | 15.85 | 65.6286
25.5 | 6.9 | 7.3 | 9.1 | 15.5857 | 62.8
25.6 | 6.85 | 7.3 | 8.9714 | 15.3 | 59.575
25.7 | 6.8 | 7.2 | 8.9 | 15 | 55.4857
25.8 | 6.7571 | 7.1143 | 8.8 | 14.725 | 50.7625
25.9 | 6.7 | 7.1 | 8.7 | 14.525 | 46.0714
26 | 6.6625 | 7 | 8.6429 | 14.2429 | 41.35
26.1 | 6.6 | 7 | 8.6 | 14.025 | 37
26.2 | 6.6 | 6.9 | 8.475 | 13.85 | 32.95
26.3 | 6.5 | 6.9 | 8.4 | 13.6 | 29.6286
26.4 | 6.5 | 6.8 | 8.3125 | 13.4143 | 26.78
26.5 | 6.4533 | 6.7833 | 8.2 | 13.1875 | 24.6
26.6 | 6.3867 | 6.7 | 8.2 | 13.0143 | 22.98
26.7 | 6.32 | 6.6333 | 8.1 | 12.8143 | 21.8
26.8 | 6.3 | 6.6 | 8.0286 | 12.675 | 20.85
26.9 | 6.2867 | 6.5714 | 7.9538 | 12.4571 | 20.14
27 | 6.22 | 6.5 | 7.9 | 12.25 | 19.56
27.1 | 6.2 | 6.4 | 7.8333 | 12.0467 | 19.08
27.2 | 6.1333 | 6.4 | 7.7 | 11.9133 | 18.66
27.3 | 6.1 | 6.3375 | 7.6286 | 11.775 | 18.26
27.4 |  | 6.3 | 7.6 | 11.5571 | 17.96
27.5 |  | 6.3 | 7.5 | 11.425 | 17.62
27.6 |  | 6.2 | 7.4375 | 11.2917 | 17.36
27.7 |  | 6.2 | 7.4 | 11.2083 | 17.06
27.8 |  | 6.1625 | 7.3 | 11.08 | 16.76
27.9 |  | 6.1 | 7.2429 | 10.9333 | 16.56
28 |  | 6.05 | 7.2 | 10.78 | 16.2833
28.1 |  | 6 | 7.1 | 10.6692 | 16.1
28.2 |  | 6 | 7.0429 | 10.5154 | 15.8
28.3 |  |  | 7 | 10.44 | 15.6167
28.4 |  |  | 6.9 | 10.3467 | 15.44
28.5 |  |  | 6.8667 | 10.2133 | 15.1667
28.6 |  |  | 6.8 | 10.0714 | 15.08
28.7 |  |  | 6.775 | 10 | 14.7857
28.8 |  |  | 6.7125 | 9.925 | 14.62
28.9 |  |  | 6.6467 | 9.7875 | 14.42
29 |  |  | 6.5727 | 9.6571 | 14.175
29.1 |  |  | 6.5 | 9.525 | 14.0429
29.2 |  |  | 6.4375 | 9.4875 | 13.825
29.3 |  |  | 6.4 | 9.3625 | 13.7
29.4 |  |  | 6.3714 | 9.2286 | 13.6
29.5 |  |  | 6.3 | 9.2 | 13.4143
29.6 |  |  | 6.2 | 9.0714 | 13.1875
29.7 |  |  |  | 8.9375 | 13.0571
29.8 |  |  |  | 8.9 | 12.925
29.9 |  |  |  | 8.8 | 12.875
30 |  |  |  | 8.7857 | 12.6143
30.1 |  |  |  | 8.6333 | 12.425
30.2 |  |  |  | 8.58 | 12.3
30.3 |  |  |  | 8.48 | 12.2571
30.4 |  |  |  | 8.3875 | 12.05
30.5 |  |  |  | 8.3 | 11.8875
30.6 |  |  |  | 8.26 | 11.7571
30.7 |  |  |  | 8.1 | 11.625
30.8 |  |  |  | 8.0625 | 11.5875
30.9 |  |  |  | 8 | 11.4625
31 |  |  |  | 7.9 | 11.3375
31.1 |  |  |  | 7.8286 | 11.2
31.2 |  |  |  | 7.7 | 11.1556
31.3 |  |  |  | 7.6571 | 10.975
31.4 |  |  |  | 7.6 | 10.8857
31.5 |  |  |  | 7.5333 | 10.7667
31.6 |  |  |  | 7.5 | 10.6733
31.7 |  |  |  | 7.3375 | 10.54
31.8 |  |  |  | 7.3 | 10.4533
31.9 |  |  |  | 7.2 | 10.3867
32 |  |  |  | 7.2 | 10.32
32.1 |  |  |  | 7.1 | 10.16
32.2 |  |  |  | 7.1 | 9.9867
32.3 |  |  |  | 7.0375 | 9.92
32.4 |  |  |  |  | 9.8563
32.5 |  |  |  |  | 9.7923
32.6 |  |  |  |  | 9.7154
32.7 |  |  |  |  | 9.6
32.8 |  |  |  |  | 9.5
32.9 |  |  |  |  | 9.4286
33 |  |  |  |  | 9.3
33.1 |  |  |  |  | 9.2
33.2 |  |  |  |  | 9.1286
33.3 |  |  |  |  | 9.1
33.4 |  |  |  |  | 9
33.5 |  |  |  |  | 8.9375
33.6 |  |  |  |  | 8.8125
33.7 |  |  |  |  | 8.7778
33.8 |  |  |  |  | 8.6571
33.9 |  |  |  |  | 8.525
34 |  |  |  |  | 8.4875
34.1 |  |  |  |  | 8.4
34.2 |  |  |  |  | 8.3286
34.3 |  |  |  |  | 8.3
34.4 |  |  |  |  | 8.2
34.5 |  |  |  |  | 8.1375
34.6 |  |  |  |  | 8.1
34.7 |  |  |  |  | 8
34.8 |  |  |  |  | 7.9375
34.9 |  |  |  |  | 7.8125
35 |  |  |  |  | 7.775

## Wedge 60 Output Factors
Output factors
Wedge 60 R
SSD=95cm
Depth=5cm
Detector: IBA Dosimetry CC13 Ionization chamber
Field size X [cm]
Field Size Y [cm] |  | 3 | 4 | 5 | 7 | 10 | 15
| 3 | 0.887 | 0.899 | 0.906 | 0.914 | 0.921 | 0.926
| 4 | 0.899 | 0.913 | 0.922 | 0.933 | 0.941 | 0.949
| 5 | 0.907 | 0.923 | 0.933 | 0.945 | 0.956 | 0.966
| 7 | 0.917 | 0.936 | 0.949 | 0.965 | 0.978 | 0.991
| 10 | 0.927 | 0.949 | 0.963 | 0.983 | 1 | 1.018
| 15 | 0.936 | 0.96 | 0.977 | 1 | 1.022 | 1.046
| 20 | 0.941 | 0.965 | 0.984 | 1.01 | 1.037 | 1.066
| 30 | 0.948 | 0.975 | 0.996 | 1.025 | 1.057 | 1.094
| 40 | 0.952 | 0.981 | 1.002 | 1.035 | 1.069 | 1.11